## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US07458943 | B1 | 2008-12-02 | Spine tilt monitor with biofeedback | 2007-06-25 | The present invention relates to a system and method for continually monitoring the spine of a user via device containing electromechanical transducer, processor, delay, transmitter, and a feedback component. Through the device and method of employing, the user will be able to learn and sustain a proper posture, thereby avoiding future back pain. | 1. A method for preventing back pain, comprising the steps of: attaching a spine monitoring device to a user;setting said device by inputting a most suitable spinal condition or position during work periods;detecting a change in the spine to said user, wherein said change is forward flexion;determining whether said change is sufficient to cause a negative effect on spine;If change is sufficient, then steps a)-b), a) transmitting a signal to said user; andb) adjusting the spine of said user;andcontinuing to monitor said user's spine,wherein said change is sufficient when said change is more than around +/−5% of said most suitable condition. 2. The method for preventing back pain of claim 1, wherein attaching said spine monitoring device occurs at said user's ear. 3. The method of preventing back pain of claim 1, wherein said condition or position includes spine angle, pressure placed upon the spine, force on the spine, acceleration by which the spine changes said angle, or velocity by which the spine changes. 4. The method of preventing back pain of claim 1, wherein setting said device occurs via user interfaces. 5. The method of preventing back pain of claim 1, wherein transmitting said signal comprises transmitting an audible signal, visual signal, or sensory signal. 6. The method of preventing back pain of claim 1, wherein said negative effect can be selected from the group consisting of neck pain, back pain, spinal disk injury, spinal inflammation, and sciatica. | The Hong Kong Polytechnic University | A61B005/117 A61B005/103
US20080281298 | A1 | 2008-11-13 | Electronic support system for biological data sensor | 2006-02-07 | An electronic support system for controlling a biological data sensor and related methods of use are disclosed herein. | 1. An electronic support system for controlling an implanted biological data sensor, comprising: a) a primary module comprised of: i) a main controller having a data acquisition interface and a telemetry interface; ii) a data acquisition unit in communication with the main controller through the data acquisition interface and configured to receive data transmitted from a biological data sensor; iii) a first telemetry unit in communication with the main controller through the telemetry interface; iv) an internal power unit; andb) a remote module comprised of a remote charging unit and a second telemetry unit,wherein the first telemetry unit is configured to communicate with second telemetry unit. 2. The electronic support system of claim 1, wherein the primary module is implantable into a test subject 3. The electronic support system of claim 1, wherein the biological data sensor is a biological analyte sensor. 4. The electronic support system of claim 3, wherein the analyte is glucose. 5. The electronic support system of claim 1, wherein the main controller comprises a microcontroller and a memory capacity. 6. The electronic support system of claim 5, wherein the memory capacity is capable of storing at least 24 hours of sampled data. 7. The electronic support system of claim 5, wherein the microcontroller is programmed with an instruction set. 8. The electronic support system of claim 5, wherein the microcontroller is an 8 or 16 bit microcontroller. 9. The electronic support system of claim 1, wherein the data acquisition interface comprises a serial peripheral interface. 10. The electronic support system of claim 1, wherein the telemetry interface comprises an RS-232 serial port. 11. The electronic support system of claim 1, wherein the data acquisition unit is configured to receive sample data from a photodiode array. 12. The electronic support system of claim 1, wherein the data acquisition unit comprises a current integrator in communication with the implantable biological sensor. 13. The electronic support system of claim 12, further comprising an analog to digital converter in communication with the current integrator. 14. The electronic support system of claim 1, wherein the data acquisition unit comprises a current integrating analog to digital converter. 15. The electronic support system of claim 14, wherein the analog/digital converter provides digitized voltage values having a level of precision in the range of from 8 bits to 128 bits. 16. The electronic support system of claim 15, wherein the analog/digital converter provides digitized voltage values having a level of precision in the range of from 16 bits to 24 bits. 17. The electronic support system of claim 1, wherein the first telemetry unit comprises a RF transmitter. 18. The electronic support system of claim 17, wherein the RF transmitter is configured to communicate with a second telemetry unit according to IEEE 802.15.4 wireless protocol. 19. The electronic... | - | A61K009/22 A61B005/055
US20080287750 | A1 | 2008-11-20 | Ergonomic probes | 2007-05-01 | An ergonomic device for intraoperative tissue characterization is provided, having a gripping handle and a sensor head, arranged with an angle α between them, wherein in absolute value, α>10 degrees, and α may be adjusted in one, two, or three planes. The device may further comprise a light fixture, configured for lighting the tissue, as it is characterized, a marking module, configured for marking the tissue, as it is characterized, a vacuum system, for improved contact between a sensing surface and the tissue, and a transparent frame, which enables an operator to observe the tissue, as it is characterized. The device may be operative with at least one sensor, selected from the group consisting of an optical sensor, an X-ray sensor, an RF sensor, a MW sensor, an infrared thermography sensor, an ultrasound sensor, an MR sensor, an impedance sensor, a temperature sensor, a biosensor, a chemical sensor, a radioactive-emission sensor, a nonirradiative RF sensor, and a mechanical sensor. | 1. An ergonomic device for intraoperative tissue characterization, comprising: a body, which comprises: proximal and distal portions, with respect to the tissue, as it is characterized;a gripping handle, at the distal portion; anda sensor head, at the proximal portion, the sensor head comprising at least one sensor for tissue characterization,wherein the gripping handle and the sensor head are arranged at an angle α, wherein in absolute value, α>10 degrees. 2. The ergonomic device of claim 1, wherein the angle α between the gripping handle and the sensor head is adjustable. 3. The ergonomic device of claim 1, wherein the angle α between the gripping handle and the sensor head is adjustable in two planes. 4. The ergonomic device of claim 1, wherein the angle α between the gripping handle and the sensor head is adjustable in three planes. 5. The ergonomic device of claim 1, wherein the sensor head further includes a vacuum system, for improved contact between a sensing surface and the tissue to be characterized. 6. The ergonomic device of claim 1, wherein the sensor head further includes a light fixture, configured for lighting the tissue, as it is characterized. 7. The ergonomic device of claim 1, wherein the sensor head further includes a marking module, configured for marking the tissue, as it is characterized. 8. The ergonomic device of claim 1, wherein the sensor head further includes a transparent frame, which surrounds the at least one sensor and enables an operator to observe the tissue, as it is characterized, through it. 9. The device of claim 1, wherein the at least one sensor is an irradiative sensor, selected from the group consisting of an optical sensor, an X-ray sensor, an RF sensor, a MW sensor, an infrared thermography sensor, and an ultrasound sensor. 10. The device of claim 1, wherein the at least one sensor is selected from the group consisting of an MR sensor, an impedance sensor, a temperature sensor, a biosensor, a chemical sensor, a radioactive-emission sensor, a nonirradiative RF sensor, and a mechanical sensor. 11. The device of claim 1, and further comprising a plurality of sensors. 12. The device of claim 1, wherein the at least one sensor includes at least two different types of sensors. 13. The device of claim 1, wherein the at least one sensor includes at least two different types of sensors, selected from the group consisting of optical sensors, X-ray sensors, RF sensors, MW sensors, infrared thermography sensors, ultrasound sensors, MR sensors, impedance sensors, temperature sensors, biosensors, chemical sensors, radioactive-emission sensors, mechanical sensors, and nonirradiative RF sensors. 14. An ergonomic device for intraoperative tissue characterization, comprising: a body, which comprises: proximal and distal portions, with respect to the tissue, as it is characterized;a gripping handle, at the distal portion; anda sensor head, at the proximal portion, the sensor head comprising: at least one sen... | Dune Medical Devices Ltd. | A61B005/05 A61B005/00
US20080287746 | A1 | 2008-11-20 | SYSTEM AND METHOD FOR COMMUNICATING HEALTH CARE ALERTS VIA AN INTERACTIVE PERSONAL HEALTH RECORD | 2007-05-16 | An automated system is described for presenting a patient with an interactive personal health record (PHR) capable of delivering individualized alerts based on comparison of evidence-based, best standards of care to information related to the patient's actual medical care. This will assure the patient that the PHR is providing him or her with optimal health care. A health care organization collects and processes a wide spectrum of medical care information, including clinical data relating to a patient and condition-specific medical reference data, in order to generate and deliver customized alerts, including Care Considerations and personalized wellness alerts, directly to the patient via an online interactive personal health record (PHR). In addition to aggregating patient-specific medical record and alert information, the PHR also solicits the patient's input for tracking of alert follow-up actions and allows the health care organization to track alert outcomes. | 1. A method of providing a customized alert to an individual patient via a personal health record comprising: establishing a set of rules from available medical standards, at least one such rule defining an expected mode of care given a particular set of clinical data;collecting medical care information relating to the patient using an automated system with access to at least one source of data, including claims data reflecting clinical information relating to the patient obtained from at least one health care provider and submitted in connection with a claim under a health plan;organizing the collected medical care information into a medical data file for the patient and storing the medical data file, the medical data file comprising patient clinical data, the patient clinical data indicating an actual mode of care provided to the patient;applying the set of rules to the contents of the medical data file to identify at least one alert based on the patient clinical data, and storing an indicator of the at least one alert in the medical data file;providing a set of one or more password-protected, personal Web pages for the patient, the set of Web pages including a display of certain elements of the patient's healthcare history automatically populated based on the contents of the medical data file; andissuing the at least one alert to the patient, via the set of Web pages, automatically-generated based on the presence of the alert indicator in the medical data file, the at least one alert providing an explanation of circumstances underlying the at least one alert and a suggestion for patient follow up. 2. The method of claim 1 wherein the at least one alert is generated when the patient's actual care as indicated by the patient clinical data is inconsistent with an expected mode of care defined by at least one of the rules. 3. The method of claim 1 wherein issuing the at least one alert further comprises providing to the patient, via the set of Web pages, access to relevant health reference information specifically pertaining to the circumstances underlying the at least one alert. 4. The method of claim 1 further comprising: providing to the patient access to an interactive health risk assessment questionnaire;receiving a response to the questionnaire from the patient; andstoring information derived from the response in the medical data file. 5. The method of claim 4 wherein access to the questionnaire is provided via the set of Web pages. 6. The method of claim 1 further comprising providing to the patient via the set of Web pages an input section corresponding to the at least one alert, the input section capable of receiving an indication from the patient that the patient has complied with the suggestion for patient follow up. 7. The method of claim 6 further comprising storing information derived from the indication in the medical data file. 8. The method of claim 1 further comprising: revising the set of rules based on changes in availab... | - | A61B005/00
US20080255465 | A1 | 2008-10-16 | Heart-function information display including color-range scalogram content | 2007-12-26 | A method employing electronic processing for presenting a heart-function informational display including the steps of (a) gathering, over a selected time interval, a person's ECG and heart-sound signals, (b) electronically sampling and digitizing such gathered signals, and (c) electronically processing those signals to create a scalogram wherein colors which lie within a selected range of colors are employed to indicate energy-based content reflected in the signals. | 1. A method employing electronic processing for presenting a heart-function informational display comprising gathering, over a selected time interval, a person's ECG and heart-sound signals,electronically sampling and digitizing such gathered signals, andelectronically processing the sampled and digitized signals to create a visually presentable wavelet scalogram thereof in the form of a plural-axis time-and-frequency-content graphical energy display, with energy content present at each time-frequency location in the display represented by a predetermined, associated color which lies within a selected range of colors. 2. The method of claim 1 which further comprises, prior to said creating, electronically processing the sampled and digitized signals to detect selected heart-function behaviors reflected therein. 3. The method of claim 1, wherein the created plural-axis display includes time and frequency axes. 4. The method of claim 2, wherein the created plural-axis display includes time and frequency axes. 5. The method of claim 1, wherein the created plural-axis display includes time, frequency and energy-level axes. 6. The method of claim 2, wherein the created plural-axis display includes time, frequency and energy-level axes. 7. A method employing electronic processing for presenting heart-function information comprising gathering over a selected time interval a person's ECG and heart-sound signals, andelectronically processing those signals to create a visually presentable graphic display which, in relation to at least a portion of the mentioned time interval, and utilizing colors which lie within a selected range of colors, simultaneously, utilizing plural display axes, time frequency and energy content of plural locations distributed within the gathered signals. 8. The method of claim 7, wherein said electronic processing includes signal sampling, digitizing and wavelet transforming. 9. The method of claim 7, wherein said creating involves producing a scalogram display. 10. A method employing electronic processing for presenting a heart-function informational display comprising gathering over a selected time interval a person's ECG and heart-sound signals,electronically sampling and digitizing such gathered signals, andelectronically processing those signals to create a scalogram wherein colors which lie within a selected range of colors are employed to indicate energy-based content reflected in the signals. | Inovise Medical, Inc. | A61B005/044
US20080243009 | A1 | 2008-10-02 | METHOD OF CONTROLLING INFLATION OF A CUFF IN BLOOD PRESSURE DETERMINATION | 2007-03-30 | The present application discloses a method of calculating an initial inflation pressure during a blood pressure determination using an NIB system. The cuff provided with the system is inflated towards a default initial inflation pressure and a plurality of oscillometric pulses are obtained during inflation. A quick systolic pressure is estimated from a pre-defined function having a physiologically- expected shape of an oscillometric envelope fitted to oscillometric data obtained during the inflation. In an embodiment, the parameters within the function are specifically found by fitting a plurality of oscillometric pulse amplitudes along with their corresponding cuff pressures obtained during inflation to the pre-defined function. The cuff is inflated up to a calculated initial inflation pressure, which is found from the estimated quick systolic pressure. After the cuff is brought to the initial inflation pressure, deflation is begun for determining the actual systolic and diastolic pressures for output to a user. | 1. A method of calculating an initial inflation pressure, while monitoring blood pressure, comprising the steps of: (a) inflating a blood pressure cuff;(b) monitoring for the presence of oscillometric pulses during the inflation of the blood pressure cuff;(c) estimating a quick systolic blood pressure for a patient based on the oscillometric pulses detected during the inflation of the blood pressure cuff; and(d) defining the initial inflation pressure as a predetermined pressure above the estimated systolic pressure;wherein the systolic blood pressure is estimated by curve fitting with a pre-defined function having a physiologically-expected shape of an oscillometric envelope, the curve fitting being done by using a plurality of oscillometric pulses along with corresponding cuff pressures obtained during an early part of inflation. 2. A method as in claim 1, wherein the step of inflating the blood pressure cuff comprises increasing cuff pressure at a rate such that the initial inflation pressure, is defined before the cuff pressure reaches a default initial inflation pressure. 3. A method as in claim 1, wherein the step of monitoring for the presence of oscillometric pulses, comprising the steps of: (a) receiving a plurality of oscillations from a pressure transducer during the inflation of the blood pressure cuff; and(b) filtering a cuff pressure waveform to extract the oscillometric pulses. 4. A method as in claim 3, wherein while monitoring for the presence of oscillometric pulses, further comprises defining an oscillometric envelope using amplitudes of the oscillometric pulses along with corresponding cuff pressures. 5. A method as in claim 1, wherein estimating the systolic blood pressure, comprises the steps of: (a) finding the function parameters within the pre-defined function with curve fitting using oscillometric pulses along with corresponding cuff pressures obtained during the inflation of the cuff;(b) obtaining a mean arterial pressure and a diastolic pressure from the fitted curve; and(c) estimating the systolic pressure from the estimated mean arterial pressure and a diastolic pressure. 6. A method as in claim 5, wherein estimating systolic blood pressure from the estimated mean arterial pressure and diastolic pressure, further comprises the steps of: (a) selecting at least one oscillometric pulse from the plurality of oscillometric pulses obtained during the inflation;(b) calibrating the selected oscillometric pulse with the obtained mean arterial pressure and diastolic pressure; and(c) estimating the systolic pressure from the calibrated oscillometric pulse. 7. A method as in claim 5, wherein estimating systolic blood pressure from the estimated mean arterial pressure and diastolic pressure, further comprises estimating systolic pressure using any mathematical formula relation of mean arterial pressure and diastolic pressure with systolic pressure. 8. A method as in claim 5, wherein the quick systolic pressure is estimated... | GENERAL ELECTRIC COMPANY | A61B005/02
US07450985 | B2 | 2008-11-11 | Head restraint system for medical research, diagnosis and operation | 2007-05-16 | The present invention relates to a head restraint kit for use with a medical head device for immobilizing the head of a patient in medical research, diagnosis and operation. The head restraint kit is particularly suitable for use with a head coil for magnetic resonance imaging. The head restraint kit contains a head rest pillow for providing support to the head of a patient, a neck cushion for providing neck support, and one or more wedge cushions for providing further support for the head. The head restraint kit may also contain one or more compact square cushions which can be inserted into a space between the head and the medical head device for additional head support. The head restraint kit may further contain a cloth strap placed over the forehead of a patient for protecting and securing the head against movement. | 1. A neuroimaging head coil and head restraint cushion assembly, comprising: a neuroimaging head coil having an inner bottom and side surfaces;a head rest pillow, comprising: a curved first end positioned on top of the inner bottom surface of the neuroimaging head coil;a flat second end positioned outside the neuroimaging head coil; andfirst and second parallel side walls extending from the curved first end to the flat second end, wherein the first and second side walls extend from inside the neuroimaging head coil to outside the neuroimaging head coil;a neck cushion positioned outside the neuroimaging head coil adjacent to the flat second end of the head rest pillow, the neck cushion having a curved top surface for providing neck support;two flat-faced wedge cushions positioned at least partially inside the neuroimaging head coil for holding headphones against a patient's head, wherein each wedge cushion is wedge-shaped having a first end thicker than a second end opposite the first end, wherein the first end of a first wedge cushion is adjacent to the first side wall of the head rest pillow, and wherein the first end of a second wedge cushion is adjacent to the second side wall of the head rest pillow;a first compact cushion positioned inside the neuroimaging head coil adjacent to the first side wall of the head rest pillow and adjacent to the first wedge cushion to cradle side and top portions of a patient's head;a second compact cushion positioned inside the neuroimaging head coil adjacent to the second side wall of the head rest pillow and adjacent to the second wedge cushion to cradle side and top portions of a patient's head; anda sensory feedback head strap adhered to the neuroimaging head coil and placed over a forehead of a patient to remind the patient of a preferred head position within the neuroimaging head coil. 2. The neuroimaging head coil and head restraint cushion assembly of claim 1, wherein said first and second compact cushions are rectangular. 3. The neuroimaging head coil and head restraint cushion assembly of claim 1, wherein said first and second compact cushions are square. 4. The neuroimaging head coil and head restraint cushion assembly of claim 1, wherein the pillow and cushions further comprise an anti-soiling coating. 5. The neuroimaging head coil and head restraint cushion assembly of claim 1, wherein the neuroimaging head coil is a magnetic resonance imaging head coil. | - | A61B005/05 A47B007/00 A61F005/37
US07611483 | B2 | 2009-11-03 | Indicator metrics for infection monitoring | 2007-04-19 | A method for monitoring an infection in proximity to an implantable medical device includes monitoring an indicator of infection in proximity to an implanted medical device; determining whether the indicator crosses a first threshold indicative of infection for a first period of time; and issuing a first alert if the indicator crosses the first threshold for the first period of time. The method further includes determining whether the indicator crosses a second threshold indicative of infection for a second period of time. A value associated with the second threshold is less indicative of an infection than a comparable value associated with the first threshold. The second period of time is greater than the first period of time. The method further includes issuing a second alert if the indicator crosses the second threshold for the second period of time. The first and second alert may be the same or different. | 1. A method for alerting that an infection is likely to be present in proximity to an implanted medical device, the method comprising: using the implanted medical device to monitor an indicator of infection in proximity to the implanted medical device; determining whether a metric of the indicator crosses a first threshold indicative of infection for a first period of time; determining whether the metric of the indicator crosses a second threshold indicative of infection for a second period of time, the second threshold being less indicative of an infection than the first threshold and the second period of time being greater than the first period of time; andissuing an alert if (i) the metric of the indicator crosses the first threshold indicative of infection for the first period of time, or(ii) the metric of the indicator crosses the second threshold indicative of infection for the second period of time. 2. The method of claim 1, wherein the first or second threshold is an absolute value metric. 3. The method of claim 1, wherein the first or second threshold is a percent deviation from a mean value metric. 4. The method of claim 1, wherein monitoring the indicator of infection comprises monitoring the indicator in discrete time intervals at a first frequency range. 5. The method of claim 4, further comprising: determining whether the metric of the monitored indicator crosses a third threshold for a third period of time; andincreasing the frequency with which the indicator is monitored if the metric of the monitored indicator crosses the third threshold for the third period of time. 6. The method of claim 1, wherein determining whether the metric of the indicator of infection crosses the second threshold comprises comparing information regarding the indicator obtained in discrete time intervals. 7. The method of claim 6, wherein issuing the alert comprises issuing the alert if the indicator crosses the second threshold in consecutive discrete time intervals. 8. The method of claim 1, wherein determining whether the metric of the indicator crosses the first threshold for the first period of time comprises determining whether the metric of the indicator crosses the threshold for a predetermined percentage of the first time period, and wherein the alert is issued if the indicator crosses the first threshold for the predetermined percentage of the first time period. 9. The method of claim 1, wherein determining whether the metric of the indicator crosses the second threshold for the second period of time comprises determining whether the metric of the indicator crosses the threshold for a predetermined percentage of the second time period, and wherein the alert is issued if the metric of the indicator crosses the threshold for a predetermined percentage of the second time period. 10. The method of claim 1, wherein monitoring the indicator of infection comprises monitoring temperature in proximity to the implanted medical device. 11. The me... | Medtronic, Inc. | A61M031/00
US20080312555 | A1 | 2008-12-18 | Devices and methods for glucose measurement using rechargeable battery energy sources | 2008-05-28 | An analyte measurement device has a housing and a visual display on said housing. The visual display has at lease one visual indicator position next to a corresponding marking on the housing. A processor drives the visual display and runs software that is modifiable to provide a variable user interface on the visual display. A rechargeable battery is provided. | 1. An analyte measurement device comprising: a housing;a visual display on said housing, said visual display having at lease one visual indicator position next to a corresponding marking on the housing; anda processor driving the visual display, wherein the processor runs software that is modifiable to provide a variable user interface on the visual display.a rechargeable battery. 2. A method of periodically obtaining blood samples, comprising: periodically using an electrically powered portable lancing aid to puncture a body part and obtaining a blood sample from the body part; at least partially depleting an energy source of the portable lancing aid; and periodically connecting the portable lancing aid to a charging station and thereby recharging the energy source of the portable lancing aid. 3. An analyte measurement device, comprising: a housing;a visual display on the housing;a penetrating member configured to be coupled to an actuator;a port for recharging a power source positioned in the housing. 4. The device of claim 3, wherein the power source is a battery. 5. The device of claim 3, wherein the power source is a rechargeable battery. 6. The device of claim 3, wherein in response to an input at the visual display user interface an electrical input is provided to the power source. 7. The device of claim 4, wherein the battery and the actuator provide for the creation of a wound tract that remains open for a sufficient time to permit a flow of a body fluid to a sample chamber in the housing. 8. The device of claim 3, wherein the power supply is an electric motor. 9. The device of claim 3, wherein the visual display is accessible from outside the housing. 10. The device of claim 4, wherein the visual display provides an indication of battery status. 11. The device of claim 5, wherein an electric motor comprises a direct current motor. 12. The device of claim 1, wherein the device is configured to hold a disposable having a plurality of penetrating members. 13. The device of claim 4, further comprising: a charge level indicator that displays the electrical charge level of the battery 14. The system of claim 3, wherein the actuator comprises a magnetic system. 15. The system of claim 3, wherein the actuator comprises a piezoelectric system. 16. A skin lancing device, comprising: a housing;a visual display on said housing, said visual displaying having at least one visual indicator position next to a corresponding marking on the housing;a processor driving the visual display, wherein the processor runs software that is modifiable to provide a variable user interface on the visual display;a penetrating member coupled to an actuator; anda port for recharging a battery positioned in the housing. 17. The device of claim 16, wherein the power source is a rechargeable battery. 18. The device of claim 16, wherein in response to an input at the visual display user interface an electrical input is provided to the battery. 20. The devi... | - | A61B005/151
US20080281213 | A1 | 2008-11-13 | CARDIAC COUPLED RESPIRATION COORDINATION SYSTEM | 2008-05-16 | A series of feed-forward applications are designed to coordinate heart/respiration/muscle and brain to produce a better quality of life endpoint. The applications will be used to treat anxiety, sleep disorders, arrhythmias and enhanced performance during athletic training. The coordination may be used as part of a diagnostic tool. | 1. A method comprising the steps of: coordinating a physiological activity of a human being with the cardiac cycle;measuring a physiological variable for the human being, the physiological variable having a characteristic response to the coordination of the cardiac cycle with the physiological activity; andcomparing the measurement of the physiological variable with a benchmark to determine if the human being has a normal or abnormal response. 2. The method of claim 1 in which the benchmark is a normal of the physiological variable derived from a class of human beings. 3. The method of claim 2 in which the physiological activity is the respiratory cycle of the human being. 4. The method of claim 3 in which the step of coordinating further comprises establishing a breathing rate of between 0.05 and 0.17 Hz. 5. The method of claim 2 in which the physiological variable comprises the cardiac cycle, and in which measuring further comprises measuring the T-wave alternans of the cardiac cycle. 6. The method of claim 2 in which measuring further comprises measuring at least one of the periodicity, frequency, magnitude, phase angle, and regularity of the physiological variable. 7. The method of claim 2 in which the physiological variable is at least one of cardiac cycle, R-R interval, blood pressure, and breathing rate. 8. The method of claim 2 in which measuring further comprises measuring the extent of coordination of the physiological variable with the physiological activity. 9. The method of claim 8 in which the physiological variable is at least one of the cardiac cycle, R-R interval, and blood pressure. 10. The method of claim 2 in which measuring further comprises measuring the extent of coordination of the physiological variable with the cardiac cycle. 11. The method of claim 10 in which the physiological variable is at least one of blood pressure and breathing rate. 12. The method of claim 2 in which coordinating the physiological activity with the cardiac cycle comprises matching a peak of inspiration or expiration with a point in the cardiac cycle. 13. The method of claim 2, in which coordinating further comprises measuring the cardiac cycle with an instrument and generating a signal coordinated with the cardiac cycle. 14. The method of claim 2 further comprising the steps of: finding an indication of the extent of coordination of the cardiac cycle with the physiological activity; andusing the indication of the extent of coordination to improve the coordination. 15. The method of claim 2 further comprising the step of monitoring a physiological parameter that is responsive to the coordination of the cardiac cycle with the physiological parameter. 16. The method of claim 15 further comprising the step of producing a signal indicative of a desired response of the physiological parameter. 17. The method of claim 2 in which the physiological activity is at least one of a hand movement, a whole body movement, and a response to an aud... | - | A61B005/02
US20080281190 | A1 | 2008-11-13 | SURGICAL PROCEDURES | 2008-07-21 | A method of locating and identifying, during surgery, target regions within a body intended for excision of suspect tissue and/or removal of diseased organs or foreign objects, resides in implantation within the body proximate the suspect tissue or object prior to surgery of one or more passive integrated transponder tags and, at the time of and/or during surgery, scanning of the body with a radio frequency scanner or reader that activates the tag or tags and provides the surgeon with one or more signals indicative of the approximate location and unique identification of each of the tags, thereby to aid the surgeon in performance of the surgery. Verification of the success of the surgical procedure is obtained following surgery by scanning the site for absence of the tag or tags and/or by scanning the excised tissue for presence of the tag or tags. | 1. A method of locating within a body target tissue, organs and objects intended for surgery comprising the steps of providing at least one implantable transponder,providing a transponder implantation device,providing a transponder locating device,prior to surgery, implanting a transponder within the body proximate the target tissue, organ or object, andat the time of surgery, scanning the body for the transponder with the transponder locating device. 2. A method as set forth in claim 1 including the step of implanting a plurality of implantable transponders proximate the target tissue, organ or object prior to surgery and, at the time of surgery, scanning the body for the transponders with the transponder locating device. 3. A method as set forth in claim 1 including the step of scanning the body with the transponder locating device following surgery. 4. A method as set forth in claim 1 including the step of scanning surgically excised matter with the transponder locating device following surgery. 5. A method as set forth in claim 1 including the steps of scanning the surgical site and surgically excised matter with the transponder locating device following surgery. 6. A method of locating and identifying within a body target tissue, organs and objects intended for surgery comprising the steps of providing at least one implantable transponder containing an identifier,providing a transponder identification reading device,prior to surgery implanting a transponder within the body proximate the target tissue, organ or object, andat the time of surgery, locating and identifying the transponder by scanning the body for the transponder with the transponder identification reading device. 7. A method as set forth in claim 6 including the steps of providing a plurality of implantable transponders each containing an identifier, implanting the plurality of implantable transponders within the body proximate the target tissue, organ or object prior to surgery and, at the time of surgery, scanning the body for the transponders with the transponder identification reading device. 8. A method as set forth in claim 6 including the steps of scanning the body and/or and surgically excised matter with the transponder identification reading device following surgery. 9. A method as set forth in claim 6 wherein the transponder comprises a radio frequency responsive passive integrated transponder and the transponder identification reading device comprises a radio frequency transmitter and receiver. 10. A procedure for use in surgically removing target tissue, organs and objects from a body comprising the steps of prior to surgery, implanting at least one implantable transponder containing an identifier within the body proximate the target tissue, organ or object,at the time of surgery, locating and identifying the target tissue, organ or object, by scanning the body with transponder locating and identifier reading apparatus, andfollowing surgery, scanning one or... | Health Beacons, Inc. | A61B005/05 A61B019/00
US20080281170 | A1 | 2008-11-13 | Method for Detecting Critical Trends in Multi-Parameter Patient Monitoring and Clinical Data Using Clustering | 2006-10-17 | A physiological data analysis component (10) determines a condition of an individual. The physiological data analysis component (10) includes an input component (12) that receives a plurality of different physiological parameters of the individual. A classification component (20) of the physiological data analysis component (10) maps these parameters to a multi-dimensional space having a plurality of regions corresponding to two or more conditions. The classification component (20) determines the condition of the individual based on the region the physiological parameters mapped within. An output component (24) of the physiological data analysis component (10) conveys the condition of the individual to a user of the physiological data analysis component (10). | 1. A physiological data analysis component that determines a condition of an individual, comprising: an input component that receives a plurality of different physiological parameters of the individual;a classification component that maps the plurality of physiological parameters to a multi-dimensional space having a plurality of regions corresponding to two or more conditions and determines the condition of the individual based on the region the physiological parameters mapped within; andan output component that conveys the condition to a user of the component. 2. The physiological data analysis component as set forth in claim 1, wherein the classification component maps two or more sets of physiological parameters obtained at different time intervals and predicts a future condition of the individual based on a trend derived from the mappings. 3. The physiological data analysis component as set forth in claim 2, wherein the classification component performs a time-series analysis to determine the trend. 4. The physiological data analysis component as set forth in claim 2, wherein the classification component generates the trend by connecting two or more mappings through a vector and extrapolating subsequent mapping. 5. The physiological data analysis component as set forth in claim 2, wherein the physiological parameters mapped to the multi-dimensional space include one or more of the following: temperature;heart rate;respiration rate;systolic blood pressure; andwhite blood cell count. 6. The physiological data analysis component as set forth in claim 1, wherein the classification component maps the physiological parameters to the multi-dimensional space through one or more of the following techniques: clustering, k-means, k-medoids, Expectation Maximization (EM), neural networks, hierarchical methods, probabilistic analysis, statistic analysis, a priori knowledge, classifiers, support vector machines, distance measures, expert systems, Bayesian belief networks, fuzzy logic, pattern recognition, interpolation, extrapolation, data fusion engines, look-up tables and polynomial expansion. 7. The physiological data analysis component as set forth in claim 1, wherein the physiological data includes two or more of heart rate, blood pressure, blood oxygen level, core body temperature, heart electrical activity, white blood count, and hormone level. 8. The physiological data analysis component as set forth in claim 1, wherein the classification component defines one or more regions of stability within the multi-dimensional space by mapping physiological parameters indicative of a stable condition to the multi-dimensional space and labelling these regions as stable. 9. The physiological data analysis component as set forth in claim 1, wherein the classification component defines one or more regions of instability within the multi-dimensional space by mapping physiological parameters indicative of an unstable condition to the multi-dimensional spa... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/00
US07549959 | B2 | 2009-06-23 | Stimulation arrangement for measurement of physiological signal reactivity | 2007-05-03 | The invention relates to a method and apparatus for assessing the reactivity observable in a certain physiological signal, especially the EEG signal, of a subject. In order to obtain a compact measurement arrangement that enables reliable assessment of changes in the reactivity of the subject even during a long time period, a conventional monitoring device is provided with a stimulation module configured to supply stimuli to the subject through a measurement probe that the device uses to receive physiological signal data from the subject. The measurement probe may comprise, for example, an arm cuff through which the device monitors the blood pressure of the subject. | 1. A method for assessing physiological signal reactivity of a subject, the method comprising: receiving physiological signal data from a subject through at least a first measurement probe and a second measurement probe attached to the subject, wherein the first measurement probe includes a plurality of electrodes that all receive the physiological signal data from the subject;supplying a stimulus to the subject through at least one electrode of the first measurement probe, wherein the at least one electrode of the first measurement probe also receives physiological signal data from the subject; anddetermining whether reactivity caused by the stimulus is present in physiological signal data received through the second measurement probe. 2. A method according to claim 1, wherein the determining includes determining whether reactivity caused by the stimulus is present in the physiological signal data received through the second measurement probe, in which the second measurement probe comprises a set of EEG electrodes for receiving EEG signal data from the subject. 3. A method according to claim 2, wherein the supplying includes supplying the stimulus through the first measurement probe, in which the stimulus is an electrical stimulus and the plurality of electrodes of the first measurement probe comprises a set of ECG electrodes. 4. A method according to claim 1, wherein the supplying includes supplying the stimulus through the first measurement probe, in which the stimulus is an electrical stimulus and the plurality of electrodes of the first measurement probe comprises a set of EEG electrodes. 5. A method according to claim 1, wherein the supplying includes supplying the stimulus through the first measurement probe, in which the stimulus is an electrical stimulus and the plurality of electrodes of the first measurement probe comprises a set of ECG electrodes. 6. A method according to claim 1, further comprising repeating the supplying and determining; anddefining whether changes occur over time in the reactivity of the subject. 7. A method according to claim 6, further comprising optimizing the strength of the stimulus for the subject. 8. A method for assessing physiological signal reactivity of a subject, the method comprising: receiving physiological signal data from a subject through a first measurement probe and a second measurement probe attached to the subject;supplying a stimulus to the subject through the first measurement probe, wherein the supplying includes pressurizing a blood pressure probe, the first measurement probe comprising the blood pressure probe, wherein the first measurement probe also receives physiological signal data from the subject; anddetermining whether reactivity caused by the stimulus is present in physiological signal data received through the second measurement probe. 9. A method according to claim 8, wherein the pressurizing is carried out in connection with the subject's blood pressure measurement. 10.... | The General Electric Company | A61B005/00 A61B005/02
US07470396 | B2 | 2008-12-30 | Device and method for establishing an artificial isolated circulation in a target area of a human or animal body | 2003-09-25 | A device for establishing an isolated perfusion including a pump arrangement, a venous catheter, and an arterial catheter, for establishing an artificial circulation in a target area of a human or animal body, which artificial circulation is isolated from the systemic circulation. The device has first means for feeding an analysis gas into the artificial circulation, and second means for monitoring whether a blood exchange takes place between the artificial circulation and the systemic circulation. | 1. A device comprising a pump arrangement, at least one venous catheter, and at least one arterial catheter, for establishing and maintaining an artificial circulation in a target area of a human or animal body, said artificial circulation being isolated from the blood circulation of the systemic body, wherein the device further comprises an oxygenator with an analysis gas delivery line into which an analysis gas can be introduced for feeding the analysis gas into said artificial circulation, and at least one gas sensor for the analysis gas for monitoring whether a blood exchange takes place between said artificial circulation and said systemic circulation and wherein the at least one gas sensor is arranged in an air outlet line of the oxygenator. 2. The device according to claim 1, further comprising at least a second gas sensor is arranged in a respiratory mask for the air exhaled from the body. 3. The device according to claim 1, further comprising at least a second gas sensor arranged in said artificial circulation. 4. The device according to claim 1, further comprising at least a second gas sensor which is arranged in the analysis gas delivery line. 5. The device according to claim 1, wherein a control unit is provided which is connected to said at least one gas sensor and to said pump arrangement, for switching off said artificial circulation in the event of a blood exchange between said artificial circulation and said systemic circulation. 6. The device according to claim 4, further comprising a control unit connected to the further gas sensor. 7. The device according to claim 1, wherein said analysis gas comprises laughing gas (N2O), and said at least one gas sensor comprises an N2O sensor. 8. The device according to claim 4, wherein said analysis gas comprises laughing gas (N2O), and said further gas sensor comprises an N2O sensor. 9. The device according to claim 7, wherein said at least one gas sensor has an N2O sensitivity of 1 to 1000 ppm. 10. The device according to claim 8, wherein said further gas sensor has an N2O sensitivity of 1 to 1000 ppm. 11. A device comprising a pump arrangement, at least one venous catheter, and at least one arterial catheter, for establishing and maintaining an artificial circulation in a target area of a human or animal body, said artificial circulation being isolated from the systemic blood circulation of the body, further comprising first means for feeding an analysis gas into said artificial circulation, and second means for monitoring whether said analysis gas passes from said artificial circulation into said systemic circulation, said first means comprising an oxygenator with a gas delivery line into which said analysis gas can be introduced, and said second means comprising a gas sensor arranged in an air outlet line of said oxygenator. | Jostra AG | A61M001/00 A61M037/00 A61M005/00 A61B005/08
US07463759 | B2 | 2008-12-09 | System and method for vascular border detection | 2006-06-08 | A system and method use the frequency spectrum of a radio frequency (RF) signal backscattered from vascular tissue to identify at least one border on a vascular image. A data gathering device is connected to a computing device and a transducer via a catheter. RF data generated via the transducer is provided to the computing device via the data-gathering device. The computing device includes: at least one data storage device for storing a plurality of tissue types and parameters related thereto, and at least one application (e.g., a characterization application, a gradient-border application, a frequency-border application and/or an active-contour application). The characterization application converts the RF data into the frequency domain and identifies a plurality of associate parameters. The identified parameters are compared to the parameters stored in the data storage device to identify the corresponding tissue type and thereafter determine at least one border on a vascular image. | 1. A method of identifying borders on an image of a vascular object, comprising: acquiring RF signal segment data backscattered from vascular tissue wherein individual ones of the RF signal segments represent an angular portion of the image;transforming said RF signal segment data from time domain data to RF signal segment frequency domain data;identifying a plurality of parameters of the RF signal segment frequency domain data;using said plurality of parameters and previously stored data to identity at least a first portion and a second portion of said RF signal segment data that corresponds to a blood portion and a non-blood portion of said vascular tissue, respectively; anddetermining a blood to non-blood border on said image of said vascular object in accordance with said first portion and said second portion of said RF signal segment data. 2. The method of claim 1, wherein the non-blood portion of said vascular tissue corresponds to a plaque portion of said vascular tissue. 3. The method of claim 1, wherein said step of determining said blood to non-blood border further comprises determining the luminal border on said image of said vascular object in accordance with said first and second portions of said RF signal segment data. 4. The method of claim 3 further comprising using said plurality of parameters and previously stored data to identify at least a third portion of said RF signal segment data that corresponds to at least a third tissue type of said vascular tissue, said third tissue type being selected from a list of tissue types consisting of medial tissue and adventitial tissue; and determining the medial-adventitial border on said image of said vascular object in accordance with said second and third portion of said transformed RF data. 5. The method of claim 4, wherein said third tissue type corresponds to both medial and adventitial tissues. 6. The method of claim 1, wherein said step of determining said border further comprises determining said border on said image of said vascular object in accordance with gradient information, said gradient information being derived from said RF signal segment data. 7. The method of claim 6, wherein said gradient information further comprises gradient-force data and gradient-border data. 8. The method of claim 1, further comprising the step of using other-border data from at least one other image of said vascular object to approximate said blood to non-blood border on said image of said vascular object. 9. The method of claim 8, wherein said step of determining said border further comprises determining said border on said image of said vascular object in accordance with gradient information, said gradient information being derived from said RF signal segment data. 10. The method of claim 1, further comprising the step of filtering said first and second portions of said RF signal frequency domain data before they are used to determine said blood to non-blood border. 11. The method of c... | The Cleveland Clinic Foundation | G06K009/00 A61B005/02
US07468044 | B2 | 2008-12-23 | Device, system and method for determining in vivo body lumen conditions | 2002-01-16 | Provided are a system and method for in vivo and in situ detection of body lumen conditions. The system comprises at least one interaction chamber for containing an endo-luminal sample, the interaction chamber comprising at least one indicator; at least one light source for illuminating the interaction chamber; and at least one optical detector for detecting in vivo optical changes occurring in the interaction chamber. The reaction between the indicator and sample may result in an optical change, which is detected and possibly imaged by the optical detector. | 1. A method for determining body lumen conditions in-vivo, the method comprising: inserting into a body lumen an in-vivo imaging device having an interaction chamber, an imager and an optical window, wherein said interaction chamber and said imager are positioned within said imaging device behind said optical window and said imager captures images of said interaction chamber;allowing a first endo-luminal sample to enter through a first opening of the in-vivo interaction chamber, the interaction chamber having immobilized therein an indicator configured to react with at least the first endo-luminal sample, said reaction occurring within the interaction chamber, and the reaction resulting in an optical change;detecting the optical change;discharging the first endo-luminal sample through a second opening of the interaction chamber; andreplacing the first sample in the interaction chamber with a new sample. 2. The method according to claim 1 comprising imaging the interaction chamber with an optical system. 3. The method according to claim 1 comprising imaging the optical changes in the interaction chamber. 4. The method according to claim 1 comprising illuminating said interaction chamber wherein at least a portion of the interaction chamber is transparent in a wavelength of illumination. 5. The method according to claim 1, comprising transmitting images to an external receiver. 6. The method according to claim 1, comprising pumping the endo-luminal sample into the interaction chamber. 7. The method according to claim 1 comprising capturing an image of a gastrointestinal wall. 8. The method according to claim 1 wherein said imager captures images of said body lumen in addition to said interaction chamber. 9. The method according to claim 1 wherein said imager captures images of optical changes that occur within the interaction chamber. | Given Imaging Ltd. | A61B005/00 A61B001/06 B65D081/00
US07440796 | B2 | 2008-10-21 | System and method for three-dimensional visualization of conductivity and current density distribution in electrically conducting object | 2003-12-24 | System for visualizing conductivity and current density distributions including a plurality of current injecting devices for injecting currents into a measuring object, an MRI scanner for measuring one directional component of a magnetic flux density due to each of the currents injected into a measuring object, an operating part for controlling the current injecting devices so as to inject currents of different directions into the measuring object, and calculating a conductivity distribution and a current density distribution inside of the measuring object by using the one directional component of a magnetic flux density, and displaying means for visualizing the conductivity and current density distributions calculated by the operating part, thereby permitting to visualize the conductivity and the current density of the measuring object more accurately. | 1. A system for visualizing conductivity and current density distributions comprising: a plurality of current injecting devices for injecting currents into a measured object;an MRI scanner for measuring one directional component of a magnetic flux density due to each of the currents injected into the measured object;an operating part for controlling the current injecting devices so as to inject currents of different directions into the measured object, and calculating a conductivity distribution and a current density distribution inside of the measured object by using the one directional component of a magnetic flux density wherein the operating part calculates an inside voltage and a surface voltage of the measured object for a discretionary conductivity, and calculates the conductivity distribution by using the calculated inside voltage and the one directional component of the magnetic flux density; anddisplaying means for visualizing the conductivity distribution and current density distributions calculated by the operating part. 2. The system as claimed in claim 1, wherein the current injecting device includes; an electrode,an insulating container with the electrode attached to one side, the insulating container having an electrolyte substance, anda wire for supplying the current to the electrode. 3. The system as claimed in claim 1, wherein the operating part controls the current injecting devices such that one pair of the current injecting devices are selected in succession, and the selected pair of the current injecting devices supply the current to the measured object. 4. The system as claimed in claim 1, wherein the operating part multiplies or divides a constant to the conductivity distribution according to a ratio of the calculated surface voltage to a measured surface voltage. 5. The system as claimed in claim 1, wherein, if an absolute value of a difference of the discretionary conductivity and the conductivity distribution is greater than a preset value, the operating part calculates a inside voltage and a surface voltage for the conductivity distribution, and calculates a new conductivity distribution by using the inside voltage for the second conductivity and the one directional component of the magnetic flux density. 6. The system as claimed in claim 1, wherein the operating part determines that the conductivity distribution is a true conductivity, if the absolute value of the difference of the discretionary conductivity and the conductivity distribution is smaller than the present value. 7. The system as claimed in claim 1, wherein the operating part substitutes the inside voltage and the one direction component of the magnetic flux density into the following equation, and subjects the equation to a line integral, to obtain the seoend conductivity distribution:1μ0⁡[∇2⁢Bz1⋮∇2⁢BzN]=[∂V1∂y-∂V1∂x⋮⋮∂VN∂y-∂VN∂x]⁡[∂σ∂x∂σ∂y]Where, μo denotes a magnetic permeability of the free space. 8. The system as claimed in claim 1, wherein t... | Kohwang Foundation, Kohwang Board of Trustee | A61B005/05
US07430905 | B2 | 2008-10-07 | Leg strength calculation device, leg strength calculation method, and program | 2006-02-28 | An acceleration sensor measures accelerations of the subject which act in directions of axes x, y, and z orthogonal to each other when the subject walks. A leg strength calculation unit calculates leg strength of the subject based on values of the accelerations measured by the acceleration sensor. | 1. A leg strength calculation device detachably worn by a subject, comprising: an acceleration sensor configured to measure accelerations of the subject which act in directions of axes x, y, and z orthogonal to each other when the subject walks; anda leg strength calculation unit configured to calculate leg strength of the subject based on values of the accelerations measured by the acceleration sensor. 2. The leg strength calculation device according to claim 1, wherein the axis z is an axis in a vertical direction;the leg strength calculation unit calculates, based on a ratio of a first acceleration value to a second acceleration value, the leg strength of the subject corresponding to the ratio;the first acceleration value is a value of the acceleration acting in the direction of the axis z measured by the acceleration sensor; andthe second acceleration value is a composite value of the accelerations acting in the directions of the axes x, y, and z measured by the acceleration sensor. 3. The leg strength calculation device according to claim 2, wherein the leg strength calculation unit calculates the leg strength of the subject corresponding to a ratio of the first acceleration value to the second acceleration value based on an approximate expression which is obtained based on a relation between a specific ratio and leg strength of the large number of subjects; andthe specific ratio is a ratio of an acceleration acting in the direction of the axis z to a composite value of accelerations acting in the directions of the axes x, y, and z of a large number of subjects. 4. The leg strength calculation device according to any one of claims 2 and 3, wherein the leg strength calculation unit calculates, based on a ratio of the first acceleration value to the second acceleration, the leg strength of the subject corresponding to the ratio value when the second acceleration value is larger than a threshold value. 5. The leg strength calculation device according to claim 1, wherein the leg strength calculation unit calculates, based on a specific value, the leg strength of the subject corresponding to the specific value; andthe specific value is a composite value of accelerations acting in the directions of the axes x, y, and z measured by the acceleration sensor. 6. The leg strength calculation device according to claim 1, wherein the leg strength calculation unit calculates, based on a specific value, the leg strength of the subject corresponding to the specific value; andthe specific value which includes weight of the subject and a composite value of accelerations acting in the directions of the axes x, y, and z measured by the acceleration sensor. 7. The leg strength calculation device according to any one of claims 5 and 6, wherein the leg strength calculation unit calculates the leg strength of the subject corresponding to the specific value based on an approximate expression which is obtained based on a relation between the specific value of a... | Sanyo Electric Co., Ltd. Non-Profit Organization Jukunen Taiiku Daigaku Research Center | A61B005/22
US07442190 | B2 | 2008-10-28 | Contact assessment of balloon catheters | 2005-05-13 | A method and system for determining contact assessment includes the steps of positioning a catheter at a tissue treatment site, where the catheter has a proximal end portion and a distal end portion, the proximal end portion defining at least one fluid inlet port and at least one fluid outlet port, an expandable membrane defining a cooling chamber a coolant injection lumen in fluid communication with the at least one fluid inlet port and the cooling chamber, a coolant return lumen in fluid communication with the at least one fluid outlet port and the cooling chamber, the coolant injection tube, the cooling chamber, and the primary coolant return lumen defining a fluid pathway and a temperature sensor located near the coolant return lumen; measuring an internal temperature of the chamber, and modifying the position of the catheter in response to the measured temperature. The method and system can also use a measured impedance to determine contact assessment. | 1. A method for determining contact assessment comprising: positioning a catheter at a tissue treatment site, the catheter comprising: a proximal end portion and a distal end portion, the proximal end portion defining at least one fluid inlet port and at least one fluid outlet port;an expandable membrane defining a cooling chamber;a coolant injection lumen in fluid communication with the at least one fluid inlet port and the cooling chamber;a coolant return lumen in fluid communication with the at least one fluid outlet port and the cooling chamber, the coolant injection tube, the cooling chamber, and the primary coolant return lumen defining a fluid pathway; and,a temperature sensor located near the coolant return lumen;circulating a cryogenic fluid through the catheter;measuring an internal temperature of the chamber; and,modifying the position of the catheter in response to the measured temperature. 2. The method of claim 1, wherein the catheter is repositioned when the measured temperature is higher than −80 degrees Celsius. 3. The method of claim 1, wherein the catheter is repositioned when the measured temperature is in the range of −60 to −75 degrees Celsius. 4. A method for determining contact assessment comprising: positioning a catheter at a tissue treatment site, the catheter comprising: an elongate body defining an injection lumen and an exhaust lumen;an expandable membrane defining a cooling chamber disposed at a point along the elongate body, the cooling chamber in fluid communication with the injection lumen and the exhaust lumen;a first electrode located on the distal side of the expandable member; anda second electrode located on the proximal side of the expandable member;injecting an electrical current between the first and second electrodes;measuring an impedance between the first and second electrodes; andassessing contact between the expandable membrane and the tissue treatment site based at least in part on the measured impedance. 5. The method of claim 4, wherein the catheter system further comprises a control unit. 6. The method of claim 4, further comprising modifying the position of the catheter in response to the measured impedance. | CryoCath Technologies Inc. | A61B018/18 A61B005/02
US07440794 | B2 | 2008-10-21 | Optical system for measuring metabolism in a body and imaging method | 2005-01-19 | In an optical measurement system and imaging method adapted to measure in vivo information in a living body without harming the living body, light rays of a plurality of wavelengths which are modulated in intensity with a plurality of different frequencies are irradiated on a plurality of irradiation positions on the surface of a living body, and time-variable changes in living body transmitting light intensity levels corresponding to the respective wavelengths and the respective irradiation positions are measured at different positions on the surface of the living body. Light is utilized to image the results of the measurements, in which the measuring time is shortened by estimating fluctuation attributable to the living body, and the presence or absence of a change in measured signal can be decided easily by displaying an estimation signal and a measured signal at a time. | 1. An optical system for measuring change in cerebral hemodynamic movement comprising: a plurality of light irradiators for irradiating light rays on a head portion of a subject;a plurality of light receivers for receiving light rays reflected within said subject;a computing unit for calculating spatial distribution of signals of an object to be measured on the basis of signals detected by said light receivers; anda display unit for displaying said spatial distribution;wherein plural said light receivers adjacent to a common said light irradiator, are all disposed on positions distant by a substantially same distance from said common said light irradiator, where said distance is not less than 10 mm and not larger than 50 mm. 2. An optical system for measuring change in cerebral hemodynamic movement comprising: a plurality of light irradiators for irradiating light rays on a head portion of a subject;a plurality of light receivers for receiving light rays reflected within said subject;a computing unit for calculating spatial distribution of signals of an object to be measured on the basis of signals detected by said light receivers; anda display unit for displaying said spatial distribution;wherein plural said light receivers adjacent to a common said light irradiator, are all disposed on positions distant by a substantially same distance from said common said light irradiator, where said distance is not less than 10 mm and not larger than 50 mm;wherein one of said light receivers detect light rays irradiated by some of said light irradiators adjacent to said one of said light receivers; andwherein light rays irradiated by said one of said light irradiators are detected by said some of said light receivers adjacent to said one of said light irradiators. 3. An optical system for measuring change in cerebral hemodynamic movement comprising: a plurality of light irradiators for irradiating light rays on a head portion of a subject;a plurality of light receivers for receiving light rays reflected within said subject;a computing unit for calculating spatial distribution of signals of an object to be measured on the basis of signals detected by said light receivers; anda display unit for displaying said spatial distribution;wherein said plurality of light irradiators and said plurality of light receivers are alternately arranged on said head portion of said subject in lattice; andwherein plural said light receivers adjacent to a common said light irradiator, are all disposed on positions distant by a substantially same distance from said common said light irradiator, where said distance is not less than 10 mm and not larger than 50 mm. 4. An optical system according to claim 1, wherein said display unit displays the signals of said object in order of measuring time or as a dynamic image. 5. An optical system according to claim 1, further comprising: a modulator for modulating light rays irradiated by said light irradiators in intensity with desired ... | Hitachi, Ltd. | A61B005/00
US07438686 | B2 | 2008-10-21 | Apparatus and method for monitoring for disordered breathing | 2003-04-21 | The present invention relates to a device and method for monitoring for sleep disordered breathing or other types of disordered breathing such as Cheyne-Stokes breathing. More specifically, a device and method for detecting disordered breathing is provided that monitors a physiological parameter, which becomes cyclical due to apnea-hyperpnea (or arousal) alternation and provides the basis for the determination of a number of breathing disorder metrics. | 1. A method of detecting disordered breathing in an implantable medical device in a patient, comprising: monitoring a physiologic characteristic of a patient that is a direct indicator of heart rate with at least one implantable sensor that provides an output signal indicative of the heart rate;identifying a cyclical variation of the heart rate, the cyclical variation comprising a first period corresponding to a first heart rate alternating with a second period corresponding to a second heart rate greater than the first heart rate;comparing said cyclical variation to threshold criteria indicative of a sleep disordered breathing event; anddeclaring the presence of the disordered breathing episode in the event that there is a positive correlation between the comparison of the cyclical variation and the threshold. 2. A method according to claim 1, wherein the at least one sensor is an electrogram circuit and said electrogram circuit further comprises at least a pair of electrodes electrically coupled to the implantable medical device. 3. A method according to claim 2, further comprising an impedance sensor coupled to a sensing circuit of the implantable medical device. 4. A method according to claim 1, wherein the at least one sensor is an impedance sensor. 5. A method according to claim 1, wherein said at least one sensor is a QT interval measurement means. 6. A method according to claim 5, further comprising withholding a declaration of the presence of the disordered breathing event until a preset number of cardiac cycles have occurred. 7. A method according to claim 1, wherein a predetermined sensing period is defined and the aggregate number of declared disordered breathing events is divided by the sensing period so that an index value is calculated for said sensing period. 8. A method according to claim 1, further comprising providing an anti-disordered breathing electrical stimulation therapy to a portion of tissue of said patient. 9. The method of claim 1 wherein identifying a cyclical variation of the heart rate comprises: determining a mean heart rate;setting an upper rate threshold and a lower rate threshold in response to the determined mean heart rate;detecting a first crossing of the lower rate threshold;detecting a first crossing in an upper direction of the upper rate threshold subsequent to the first crossing of the lower rate threshold;detecting a second crossing in a downward direction of the upper rate threshold subsequent to the first upper rate threshold crossing;detecting a second crossing of the lower rate threshold subsequent to the second upper rate threshold crossing; andidentifying the cyclical variation in heart rate in response to the detected second crossing. 10. The method of claim 9 wherein detecting the cyclical variation in heart rate further comprises: setting a cycle onset time in response to the first crossing of the lower rate threshold,setting a cycle end time in response to the second crossing of the ... | Medtronic, Inc. | A61B005/0205
US07462150 | B1 | 2008-12-09 | System and method for evaluating impaired glucose tolerance and diabetes mellitus within a patient using an implantable medical device | 2006-06-09 | Techniques are described for use by a pacemaker or implantable cardioverter/defibrillator (ICD) or other implantable medical device. The techniques are provided for evaluating the likelihood that the patient, in which the device is implanted, has impaired glucose tolerance (IGT) or diabetes mellitus. Briefly, a value representative of sleep quality of the patient is detected and then the likelihood that the patient has IGT or diabetes mellitus is determined based on the sleep quality value. In this regard, it has been found that a decrease in overall sleep quality is associated with an increased likelihood of IGT or diabetes mellitus, which is in turn associated with an increased risk of mortality. Hence, sleep quality may be used as a proxy for evaluating the likelihood that the patient has IGT or diabetes mellitus and for assessing associated mortality risk. | 1. A method for use with an implantable medical device for use within a patient, the method comprising: detecting a value representative of sleep quality of the patient; andevaluating the likelihood that the patient has an abnormal glucose level medical condition, the evaluation performed based on the sleep quality value. 2. The method of claim 1 wherein the abnormal glucose level medical condition includes one or more of impaired glucose intolerance (IGT) and diabetes mellitus. 3. The method of claim 1 wherein detecting sleep quality within the patient includes detecting sleep efficiency (SE) associated with a period of sleep. 4. The method of claim 3 wherein detecting sleep efficiency (SE) includes detecting a total quality sleep time (TQST) associated with a period of sleep;detecting a total amount of time the patient spent in the sleep position (TSSP) during the period of sleep; andcalculating a ratio of total quality sleep time (TQST) to time spent in sleep position (TSSP). 5. The method of claim 4 wherein detecting total quality sleep time (TQST) includes: detecting the onset (SO) of the period of sleep;detecting a final awakening (FA) from the period of sleep;calculating the total sleep period (TSP) as the time from sleep onset (SO) to final awakening (FA);detecting any time the patient spends moving (TSM) during the total sleep period (TSP);detecting any temporary awake time (TAT) during the total sleep period (TSP); andcalculating total quality sleep time (TQST) as total sleep period (TSP) less any intervening time the patient spends moving (TSM) during the total sleep period (TSP) and less any temporary awake time (TAT) during the total sleep period (TSP). 6. The method of claim 5 wherein detecting sleep onset (SO) and detecting final awakening (FA) is performed, in part, based on changes in one or more of (a) heart rate variability, (b) cardiac signal morphology, (c) respiration rate and (d) respiration depth, each derived from an intracardiac electrogram (IEGM) of the patient sensed using implanted leads, and (e) blood pressure sensed using an implanted blood pressure sensor. 7. The method of claim 5 wherein detecting any time the patient spends moving (TSM) is performed, in part, based on changes in one or more of (a) the morphology of an IEGM of the patient sensed using implanted leads, (b) paced depolarization integral (PDI) sensed using implanted leads, and (c) patient movement sensed using an implanted activity sensor. 8. The method of claim 5 wherein detecting any temporary awake time (TAT) is performed, in part, based on changes in one or more of (a) heart rate variability derived from an IEGM sensed using implanted leads, (b) patient motion sensed using an implanted activity sensor and (c) blood pressure sensed using an implanted blood pressure sensor. 9. The method of claim 5 wherein detecting total time of the patient spent in sleep position (TSSP) is performed, in part, based on patient posture sensed using an impla... | Pacesetter, Inc. | A61B005/00
US07433730 | B1 | 2008-10-07 | Systems and methods for reconstructing information using a Duncan and Horn formulation of the Kalman filter for regularization | 2004-03-11 | Systems and methods for generating regularized solutions for inverse problems, where the regularization is performed using the Duncan-Horn formulation of a Kalman filter. The various alternative embodiments of the invention are applicable to many different types of inverse problems. In one electrocardiography implementation, electrical potentials are measured within a chamber of the heart and these measurements are processed to generate a mapping of the electrical activity at the interior surface of the heart chamber (these potentials at the heart s surface resulted in the potentials that were measured in the heart chamber). The generation of the mapping of the potentials at the heart s surface is regularized by using a Duncan-Horn formulation of a Kalman filter. The regularization may be based upon multiple steps in time, and the number of time steps may be selectable. | 1. A method comprising: receiving signals corresponding to a first parameter;performing an inversion of the received signals, including regularizing the inversion by a Duncan-Horn formulation of a Kalman filter, to generate a mapping of a second parameter; andproviding the mapping to a user. 2. The method of claim 1, wherein the first parameter comprises electrocardial potentials at a distance from a surface of a heart, and wherein the second parameter comprises electrocardial potentials at the surface of the heart. 3. The method of claim 2, wherein the first parameter electrocardial potentials are measured at positions within a chamber of the heart. 4. The method of claim 2, wherein the first parameter electrocardial potentials are measured at positions external to the heart. 5. The method of claim 2, wherein the first parameter electrocardial potentials are measured simultaneously by a multi-electrode probe. 6. The method of claim 5, wherein the first parameter electrical potentials are measured simultaneously by regularly spaced electrodes on the multi-electrode probe. 7. The method of claim 2, wherein the first parameter electrocardial potentials are measured using a multi-sensor probe including elements selected from the group consisting of superconductive quantum interference devices, magnetometers and electrometer amplifier based sensors. 8. The method of claim 1, wherein the Duncan-Horn formulation of the Kalman filter is based upon multiple steps in time. 9. The method of claim 8, wherein the Duncan-Horn formulation of the Kalman filter is based upon a selectable number of steps in time. 10. The method of claim 1, wherein providing the mapping to a user comprises displaying a graphical image display of the mapping to the user. 11. A system comprising: a data processor;a data input interface coupled to the data processor; anda data output interface coupled to the data processor;wherein the processor is configured to receive input data corresponding to a first parameter from the data input interface,perform regularization using a Duncan-Horn formulation of a Kalman filter in generating output data corresponding to a second parameter, andprovide the output data to the data output interface. 12. The system of claim 11, wherein the first parameter comprises electrocardial potentials at a distance from a surface of a heart, and wherein the second parameter comprises electrocardial potentials at the surface of the heart, and wherein the data processor is configured to generate a mapping of the electrocardial potentials at the surface of the heart. 13. The system of claim 12, further comprising a multi-electrode probe coupled to the data input interface and configured to measure the first parameter electrocardial potentials. 14. The system of claim 13, wherein the multi-electrode probe is configured to measure the first parameter electrocardial potentials at positions within a chamber of the heart. 15. The system of claim 13, where... | - | A61B005/04
US08868147 | B2 | 2014-10-21 | Method and apparatus for controlling positioning of a noninvasive analyzer sample probe | 2008-05-21 | The invention relates generally to a probe interface method and apparatus for use in conjunction with an optical based noninvasive analyzer. More particularly, an algorithm controls a sample probe position and attitude relative to a skin sample site before and/or during sampling. For example, a sample probe head of a sample module is controlled by an algorithm along the normal-to-skin-axis. Preferably, the sample probe head is positioned in terms of 3-D location in the x-, y-, and z-axes and is attitude orientated in terms of pitch, yaw, and roll. Further, attitude of the probe head is preferably orientated prior to contact of the sample probe head with the tissue sample using indicators, such as non-contact distance feedback from capacitance sensor, contacting or non-contacting optical sensors, and/or contact electrical sensors. | 1. A noninvasive spectroscopic analyzer configured to determine an analyte property from a sample site of a body part, comprising: a sample probe having a sample probe tip;a controller for controlling attitude of said sample probe tip relative to localized curvature of the sample site based on an input;a first motor operating through a first pivot point positioned on a first concentric ring to move a second concentric ring to control pitch of said sample probe tip; anda second motor operating through a second pivot point positioned on the second concentric ring to move a third concentric ring to control roll of said sample probe tip; andwherein the controller adjusts the attitude including at least pitch and roll of said sample probe relative to the localized curvature of the sample site by controlling the first and second motors;wherein the input is configured to receive an input signal providing said controller with information about said attitude of said sample probe tip relative to the sample site. 2. The noninvasive spectroscopic analyzer of claim 1, wherein attitude of said sample probe tip relative to the sample site is defined in terms of an x-, y-, and z-axis coordinate system, wherein said x-axis is defined along a length of said body part, wherein said y-axis is defined across said body part, and wherein said z-axis is normal to a plane formed by said x- and y-axes. 3. The noninvasive spectroscopic analyzer of claim 2, wherein said controller controls said attitude of said sample probe tip along all of said x-, y-, and z-axes. 4. The noninvasive spectroscopic analyzer of claim 2, wherein said controller controls said attitude of said sample probe tip along a normal-to-sample site axis, wherein said normal-to-sample site axis is not said z-axis. 5. The noninvasive spectroscopic analyzer of claim 2, wherein pitch comprises rotation of said sample probe tip about said y-axis, wherein roll comprises rotation of said sample probe tip about said x-axis. 6. The noninvasive spectroscopic analyzer of claim 2, further comprising: a third motor for controlling movement of said sample probe tip along a normal-to-skin-axis, wherein said normal-to-skin axis is tangential to a plane defined by said x- and y-axes and is not aligned with an axis corresponding to that of a force attributable to gravity. 7. The noninvasive spectroscopic analyzer of claim 6, wherein the third motor is positioned at a periphery of the third concentric ring to move the sample probe tip along the normal-to-skin-axis. 8. The noninvasive spectroscopic analyzer of claim 1, wherein said input is further configured to receive said input signal comprising a plurality of capacitance signals, wherein each of said plurality of capacitance signals provides information on a distance between a localized section of said sample probe tip and the sample site. 9. The noninvasive spectroscopic analyzer of claim 1, wherein said input signal comprises a conductive contact signal, where... | GLT Acquisition Corp. | A61B005/1455 A61B005/06 A61B005/145 A61B005/00
US08172155 | B2 | 2012-05-08 | Thermal sensation determining apparatus and method, and air-conditioning control apparatus using thermal sensation determination result | 2008-03-11 | A thermal sensation determining apparatus includes a skin temperature detecting unit configured to detect a skin temperature of a peripheral site of a subject's body, a variation state determining unit configured to determine a variation state of the skin temperature during a first interval to obtain a variation state determination result, a gradient detecting unit configured to detect a gradient of the skin temperature during a second interval to determine the gradient to obtain a gradient detection result, and a thermal sensation determining unit configured to determine the subject's thermal sensation using the variation state determination result and the gradient detection result to obtain a thermal sensation determination result classified into levels ranging from “hot” to “cold” based on a “moderate” temperature. | 1. A thermal sensation determining apparatus comprising: a skin temperature detecting unit configured to detect a skin temperature of a peripheral site of a subject's body;a variation state determining unit configured to (a) determine a variation state determination result of the skin temperature during a first interval to be a first state if a difference between a maximum value and a minimum value of the skin temperature during the first interval is smaller than a temperature threshold or if the difference between the maximum value and the minimum value is equal to or greater than the temperature threshold and a time difference in detection time between the maximum value and the minimum value is equal to or greater than a first time threshold, (b) determine the variation state determination result to be a second state if the difference between the maximum value and the minimum value is equal to or greater than the temperature threshold and the time difference is smaller than the first time threshold and equal to or greater than a second time threshold which is smaller than the first time threshold, and (c) determine the variation state determination result to be a third state if the difference between the maximum value and the minimum value is equal to or greater than the temperature threshold and the time difference is smaller than the second time threshold;a gradient detecting unit configured to detect a gradient of the skin temperature during a second interval longer than the first interval to (d) determine a gradient detection result to be a first gradient if the gradient of the skin temperature is equal to or greater than a first gradient threshold, (e) determine the gradient detection result to be a second gradient if the gradient of the skin temperature is equal to or greater than a second gradient threshold smaller than the first gradient threshold and is smaller than the first gradient threshold, and (f) determine the gradient detection result to be a third gradient if the gradient of the skin temperature is smaller than the second gradient threshold; anda thermal sensation determining unit configured to determine the subject's thermal sensation using an old thermal sensation determination result, the variation state determination result and the gradient detection result to obtain a new thermal sensation determination result classified into levels ranging from “hot” to “cold” based on a “moderate” temperature. 2. An air-conditioning control apparatus comprising: a skin temperature detecting unit configured to detect a skin temperature of a peripheral site of a subject's body;a variation state determining unit configured to (a) determine a variation state determination result of the skin temperature during a first interval to be a first state if a difference between a maximum value and a minimum value of the skin temperature during the first interval is smaller than a temperature threshold or if the difference between the maximum value ... | Kabushiki Kaisha Toshiba | F23N005/20 G05D023/00 F24F003/14 G01K003/00
US08112139 | B2 | 2012-02-07 | Skin screw electrode | 2008-02-04 | Electrodes providing excellent recording and physical stability. Electrodes that include a plurality of small teeth that possess a novel design shape and orientation are disclosed. The shallow and relatively long teeth preferably run parallel to the rim of the electrode that presses against the patient's skin. When the electrode is twisted onto skin, the teeth penetrate nearly horizontally under the stratum corneum. The electrodes cause minimal discomfort to the patient since the teeth do not extend to the pain fibers which are located in deeper layers of the skin. The electrodes may house a diversity of electronic components to enable numerous experimental and medical implementations. The electrodes may also be used wirelessly without electrode leads. The electrodes may be fabricated using precision photo-chemical etching techniques that are well known in the art. An electrode installation device that preferably employs the electrodes is also disclosed. | 1. An electrode for attachment to skin of a patient comprising: an electrode body, wherein said electrode body is a cylinder having a vertical wall with a first rim at a top of said cylinder and a second rim at a bottom of said cylinder, said first rim of said cylinder having a plurality of teeth, wherein each tooth includes a shaft that initially extends vertically from said first rim and then turns to run roughly horizontally to said first rim within the vertical plane defined by the wall of said cylinder and;a cap that is connected to said second rim of said electrode body. 2. The electrode of claim 1, wherein said teeth are located around the entirety of said first rim. 3. The electrode of claim 1, wherein said shaft of said teeth are between about 0.005 and about 0.05 inches long. 4. The electrode of claim 1, wherein said teeth are adapted to penetrate the stratum corneum of said skin, but not as deep as the layer containing pain fibers. 5. The electrode of claim 1, wherein said teeth extend vertically from about 0.002″ to about 0.005″ from said first rim. 6. The electrode of claim 1, wherein said teeth are angled away from said first rim from about 2° to about 5°. 7. The electrode of claim 1, further comprising a recessed area in said first rim at a base of each of said plurality of said teeth, wherein each recessed area is adapted to accommodate any hair located on the surface of said skin. 8. The electrode of claim 1, wherein said electrode is fabricated from a conductive material selected from the group consisting of stainless steel, copper alloy, conductive polymer, gold alloy, and platinum. 9. The electrode of claim 1, wherein said teeth are electroplated or coated with a silver-silver chloride, gold, iridium, or conductive metal oxide. 10. The electrode of claim 1, wherein said electrode body and said cap are magnetic. 11. The electrode of claim 10, further comprising an electrical lead wire connected to said cap. 12. The electrode of claim 1, further comprising an amplifier, a telemetric circuit, and a transmitter. 13. The electrode of claim 12, wherein said amplifier, telemetric circuit, and transmitter are located in the interior of said electrode body. 14. The electrode of claim 13, wherein said electrode is adapted to transmit information to a receiver wirelessly. 15. The electrode of claim 1, wherein said electrode body comprises: a first portion in the form of a cylinder that includes said teeth on a proximal end of said first portion;a second portion that is distal to said first portion, said second portion comprising a constraining wall, said constraining wall including an alignment notch along its outside face, wherein a distal aspect of said first portion is adapted to fit snugly over a proximal aspect of said second portion to the level of said alignment notch. 16. An array of electrodes according to claim 1, wherein said electrodes are interconnected via wire leads, wherein said wire leads are adapted to tr... | University of Pittsburgh—of the Commonwealth System of Higher Education | A61B005/0478
US08043289 | B2 | 2011-10-25 | Methods and systems for focused bipolar tissue ablation | 2008-05-27 | Large tissue regions are treated using pairs of electrode arrays. The electrode arrays may be concave and disposed in tissue so that their concave portions are opposed to each other. Axial conductors may be provided extending from the arrays and toward each other in order to increase the heating of tissues lying along the axis between the deployed electrode arrays. By properly spacing the electrode arrays apart and selecting the diameters of the arrays, desired volumes of tissue may be treated, typically with a bipolar, radiofrequency current. | 1. A tissue ablation probe comprising: a shaft having a distal end and a proximal end;a first array of electrodes mounted on the shaft and configured for movement between a retracted configuration and a deployed configuration, wherein movement to the deployed configuration comprises deployment of at least one electrode of the first array from the shaft in an initial distal direction followed by a radially outward eversion so as to place an end face of the at least one electrode of the first array at least partially in the proximal direction;a second array of electrodes mounted on the shaft and configured for movement between a retracted configuration and a deployed configuration, wherein movement to the deployed configuration comprises deployment of at least one electrode of the second array from the shaft in an initial proximal direction followed by everting in the radially outward direction so as to place an end face of the at least one electrode of the second array at least partially in the distal direction;a first connector coupled to the shaft and electrically connected to the first electrode array, the first connector configured to connect to one pole of a power supply; anda second connector coupled to the shaft and electrically connected to the second electrode array, the second connector configured to connect to a second pole of the power supply. 2. The tissue ablation probe of claim 1, wherein the first array of electrodes and second array of electrodes are configured to necrose a volume of tissue therebetween when energy is applied to the first array of electrodes and second array of electrodes. 3. The tissue ablation probe of claim 2, wherein an inner portion of the volume of tissue is necrosed prior to an outer portion of the volume of tissue. 4. The tissue ablation probe of claim 1, further comprising a first axial conductor coupled to the first array of electrodes and a second axial conductor coupled to the second array of electrodes. 5. The tissue ablation probe of claim 4, where the first axial conductor and the second axial conductor are aligned on a common axis line. 6. The tissue ablation probe of claim 1, wherein in the deployed configuration the first array of electrodes and the second array of electrodes have an outer diameter within a range of 1 cm to 6 cm. 7. The tissue ablation probe of claim 6, wherein the outer diameter of the first array of electrodes is the same as the outer diameter of the second array of electrodes. 8. The tissue ablation probe of claim 1, wherein the first electrode array and the second electrode array are spaced at least 2 cm from each other. 9. The tissue ablation probe of claim 1, wherein the first electrode array and the second electrode array each has a concave face when deployed. 10. A tissue ablation probe comprising: a shaft having a distal end and a proximal end;a first array of electrodes mounted on the shaft and configured for movement between a retracted configuration and a dep... | Boston Scientific Scimed, Inc. | A61B018/14
US08285015 | B2 | 2012-10-09 | Simultaneous acquisition of differing image types | 2008-07-02 | A system in one embodiment includes an image forming device for forming an image from an area of interest containing different image components; an illumination device for illuminating the area of interest with light containing multiple components; at least one light source coupled to the illumination device, the at least one light source providing light to the illumination device containing different components, each component having distinct spectral characteristics and relative intensity; an image analyzer coupled to the image forming device, the image analyzer decomposing the image formed by the image forming device into multiple component parts based on type of imaging; and multiple image capture devices, each image capture device receiving one of the component parts of the image. A method in one embodiment includes receiving an image from an image forming device; decomposing the image formed by the image forming device into multiple component parts based on type of imaging; receiving the component parts of the image; and outputting image information based on the component parts of the image. Additional systems and methods are presented. | 1. A system, comprising: an image analyzer for coupling to an image forming device adapted for insertion in a human body, the image analyzer decomposing an image formed by the image forming device into multiple component parts based on type of imaging;multiple image capture devices, each image capture device receiving one of the component parts of the image; anda processing device coupled to the image capture devices,wherein the processing device is configured to acquire and process image data from the multiple image capture devices in real tune and provide a new composite image based on the component parts. 2. The system of claim 1, further comprising at least one light source for providing light to an illumination device coupled to the image forming device for delivering illumination into the area of interest. 3. The system of claim 2, further comprising the image forming device and an illumination channel, the illumination channel being configured such that light from the illumination channel is focused onto a smaller portion of the area of interest than that illuminated by the illumination device. 4. The system of claim 3, further comprising a spectrometer equipped with a capture device for receiving light from the image analyzer originating from the smaller portion of the area of interest that is illuminated by the focused light of the illumination channel. 5. The system of claim 2, wherein the at least one light source includes a first light source for generating ultraviolet, far red, near infrared or infrared light and a second light source for generating visible white light. 6. The system of claim 2, wherein the at least one light source includes a broadband visible light source and at least one other light source for inducing tissue autofluorescence outside of a visible spectrum or fluorescence from the tissue outside of a visible spectrum in the presence of a fluorescence contrast agent. 7. The system of claim 2, wherein the at least one light source includes a visible light source and at least one other light source for generating broadband or multiple narrow bands in the far red and near infrared (NIR) spectral region. 8. The system of claim 1, wherein the image forming device includes a single endoscopic or endoscopic microscope or a device that can provide endoscopic imaging from the microscopic to the macroscopic level. 9. The system of claim 1, wherein composite image information based on multiple component parts of the image is output. 10. The system of claim 1 wherein the image forming device is a handheld or robotically held device that can view an interior or exterior of a human body. 11. The system of claim 10, wherein the image forming device provides spatial resolution capable of viewing tissue organization and structure at a cellular level. 12. The system of claim 10, wherein the image forming device provides a user-selectable spatial resolution. 13. A system, comprising: an image analyzer for coupling to an im... | Lawrence Livermore Natioonal Security, LLC | G06K009/00 A61B006/00
US09283026 | B2 | 2016-03-15 | Assessment of electrode coupling for tissue ablation | 2006-12-06 | An electrode catheter (14) and a method for assessing electrode-tissue contact and coupling are disclosed. An exemplary electrode catheter (14) comprises an electrode (20) adapted to apply electrical energy. A measurement circuit (42) is adapted to measure impedance between the electrode (20) and ground as the electrode approaches a target tissue (24). A processor (50) determines a contact and coupling condition for the target tissue (24) based at least in part on reactance of the impedance measured by the measurement circuit (42). In another exemplary embodiment, the electrode catheter (14) determines the contact and coupling condition based at least in part on a phase angle of the impedance. | 1. A coupling sensing system comprising: a measurement circuit adapted to measure impedance between an electrode of a catheter adapted to apply electrical energy and ground as the electrode approaches a target tissue;a first data structure including a plurality of sets of empirically predetermined electrical coupling conditions indicative of empirical amounts of ablative energy passing between a testing electrode and a corresponding plurality of types of tissue during ablation, wherein each set of said plurality of sets of empirically predetermined electrical coupling conditions is related to electrical conductivity and mechanical compliance of each type of tissue of the corresponding plurality of types of tissue, the corresponding plurality of types of tissue including the target tissue; anda processor configured to make a determination of an electrical coupling condition indicative of an amount of ablative energy that will pass between said electrode and the target tissue upon application of electrical ablation energy to said electrode, wherein said determination is based at least in part on a difference between the impedance measured by said measurement circuit and a baseline value in blood proximate to the target tissue, and wherein said processor is configured to make a comparison of said difference with one set of said plurality of sets of empirically predetermined electrical coupling conditions corresponding to the target tissue; andan output device for indicating the electrical coupling condition to a user prior to application of said ablation energy to said electrode. 2. The coupling sensing system of claim 1, wherein a level of said electrical coupling condition directly corresponds to a reactance component of the impedance measured. 3. The coupling sensing system of claim 1, wherein a level of said electrical coupling condition directly corresponds to a phase angle component of the impedance measured. 4. The coupling sensing system of claim 1, wherein the output device includes analog output corresponding to the electrical coupling condition. 5. The coupling sensing system of claim 1, wherein the measurement circuit includes a resistance-capacitance-inductance (RCL) meter. 6. The coupling sensing system of claim 1, wherein the measurement circuit includes a phase detection circuit. 7. The coupling sensing system of claim 1, wherein the electrode is adapted to apply said electrical energy at a first power for the determination of the electrical coupling condition and apply said electrical energy at a second power, wherein said electrical energy at said second power is said ablation energy. 8. The coupling sensing system of claim 7, further comprising a constant current source to apply said electrical energy for said electrical coupling condition. 9. The coupling sensing system of claim 1, wherein each of said plurality of sets of empirically predetermined electrical coupling conditions further comprises: a plurality of electrica... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B018/04 A61B005/053 A61B018/12 A61B018/14 A61B019/00 A61B018/00
US10264993 | B2 | 2019-04-23 | Sample scanning and analysis system and methods for using the same | 2007-06-21 | A system and methods for scanning and analyzing one or more characteristics of a sample utilizing electromagnetic radiation is described. More particularly, the system and methods utilize an electromagnetic radiation source connected to a transmitter and an analyzer connected to a receiver. A sample to be analyzed is placed between the transmitter and receiver and a frequency sweep of electromagnetic radiation is transmitted through the sample to create a series of spectral data sets, which are developed into a composite spectrogram by the analyzer and processed to determine one or more characteristics of the sample. A magnetic field can alternatively be applied around the transmitter, receiver and sample to enhance some characteristic analysis applications and to make other characteristic analysis applications possible. Samples include inert and living items and the characteristics include a wide variety of different applications. Although pattern recognition is utilized in some applications to match a present composite spectrogram with previously detected spectrograms, the present composite spectrogram can also be analyzed based on pattern components to make characteristic determinations without per-use calibration and without utilizing any pattern matching. | 1. A portable sample scanning and analysis system comprising: a sample housing including a human finger-sized housing for holding a sample to be scanned;a scanner for scanning the sample, the scanner including a transmitter for transmitting a series of electromagnetic signals, each of the electromagnetic signals having a transmitted amplitude over a plurality of sample points within a range of frequencies, and a receiver for receiving the series of electromagnetic signals as the series of electromagnetic signals are reflected from or transmitted through the sample at substantially the same time as the series of electromagnetic signals are being transmitted, and generating a series of spectral data sets representing at least a scanned amplitude of each of the electromagnetic signals over the plurality of sample points;magnets that produce a non-varying magnetic field around any portion of the sample to be scanned, the magnetic field representing substantially all of the magnetic field applied by the system around the sample at the same time the electromagnetic signals are transmitted and received, the transmitter and the receiver being held by the human finger-sized housing on either side of the sample and in close proximity to the magnets and the sample; andan analyzer including a memory and a processor for generating one or more composite spectrograms from the series of spectral data sets, analyzing the composite spectrograms to recognize one or more patterns within the composite spectrograms representing one or more characteristics of the sample, and making a determination regarding the characteristics based on the patterns. 2. The system of claim 1, wherein the sample housing includes at least one node connected to the transmitter and the receiver, the node being held in a fixed position relative to the sample. 3. The system of claim 2, wherein the sample housing further includes a base formed from a dielectric material having a low dissipation factor in which the sample is positioned for scanning, and wherein the housing forms a shield around at least a portion of the base. 4. The system of claim 2, wherein the shape, orientation and material composition of the node is selected to enhance the composite spectrograms. 5. The system of claim 2, wherein the sample housing includes a transmitter node connected to the transmitter and a receiver node connected to the receiver and a base formed to hold the transmitter node and the receiver node in physical contact with the sample. 6. The system of claim 1, wherein the sample housing includes a transmitter node connected to the transmitter and a receiver node connected to the receiver, and wherein the receiver node only receives the series of signals transmitted from the transmitter node. 7. The system of claim 1, wherein the sample housing includes one or more transmitter nodes connected to the transmitter and one or more receiver nodes connected to the receiver, and wherein a single transmitte... | RF SCIENCE & TECHNOLOGY INC. | G01N033/48 G01N033/50 A61B005/05 G06F019/00 A61B005/145 A61B005/00 G06F019/12
US20080312542 | A1 | 2008-12-18 | MULTI-SENSOR ARRAY FOR MEASURING BLOOD PRESSURE | 2008-06-13 | A sensor for monitoring a patient's blood pressure, the sensor including a housing unit with a back surface and which includes: a pair of electrodes mounted on the back surface; an optical system mounted on the back surface and including at least one light source that emits optical radiation near 570 nm and at least one photodetector; a first amplifier which generates an analog electrical waveform from the electrical signals from the electrodes; a second amplifier that generates an analog optical waveform from the optical signal from the photodetector; analog-to-digital converter circuitry configured to receive the analog electrical waveform and generate a digital electrical waveform therefrom and to receive the analog optical waveform and generate a digital optical waveform therefrom; and a processor programmed to receive the digital electrical and optical waveforms and determine a pulse transit time for the patient which is a measure of a separation in time of a first feature of the digital electrical waveform and a second feature of the digital optical waveform and to use the pulse transit time to determine a blood pressure value for the patient. | 1. A sensor for monitoring a patient's blood pressure, the sensor comprising a housing unit configured to be worn on a patient's arm, the housing unit having a back surface and comprising: a pair of electrodes mounted on said back surface and positioned to contact the patient's skin to generate electrical signals when the housing is worn on the patient's arm;an optical system mounted on said back surface and comprising at least one light source that emits optical radiation near 570 nm and at least one photodetector, wherein the at least one light source and the at least one photodetector are positioned to be adjacent to the patient's skin when the housing is worn on the patient's arm, wherein optical system is configured to generate an optical signal by detecting optical radiation emitted by the at least one light source and reflected from blood vessels underneath the patient's skin;a first amplifier configured to receive the electrical signals from the electrodes and generate an analog electrical waveform therefrom;a second amplifier configured to receive the optical signal from the photodetector and generate an analog optical waveform therefrom;analog-to-digital converter circuitry configured to receive the analog electrical waveform and generate a digital electrical waveform therefrom and to receive the analog optical waveform and generate a digital optical waveform therefrom; anda processor programmed to receive the digital electrical and optical waveforms and determine a pulse transit time for the patient which is a measure of a separation in time of a first feature of the digital electrical waveform and a second feature of the digital optical waveform and to use the pulse transit time to determine a blood pressure value for the patient. 2. The sensor of claim 1, further comprising a short-range wireless transceiver configured to receive the blood pressure value for the patient and transmit it to a remote receiver. 3. The sensor of claim 1, further comprising an armband attached to the housing unit and for attaching the housing unit to the patient's arm. 4. The sensor of claim 1, wherein the optical system comprises an array of optical modules mounted on said back surface and positioned to be adjacent to the patient's skin when the housing is worn on the patient's arm, wherein each optical module of the array of optical modules comprises a light source that emits optical radiation near 570 nm and a photodetector configured to generate an optical signal by detecting optical radiation emitted by the light source within that optical module and reflected from blood vessels underneath the patient's skin and wherein the at least one light source and the at least one photodetector are among the light sources and photodetectors of the array. 5. The sensor of claim 1, wherein the optical modules are linearly arranged on said back surface. 6. The sensor of claim 1, wherein the optical modules are arranged in a circular pattern on said back surfac... | TRIAGE WIRELESS, INC. | A61B005/021
US20080255476 | A1 | 2008-10-16 | Methods and devices for treating obesity | 2007-11-13 | Methods and devices for treating obesity are provided. The consumption of calorie dense, lipid rich, or fatty foods is discouraged through the modulation of a subjects gallbladder function or output. Disclosed are devices and methods for delivering devices within the gallbladder and associated ducts and vasculature; other methods involve implanting devices on or around the gallbladder and associated ducts and vasculature. Further treatments involve the use of energy, surgery, or chemicals to alter the function of the gallbladder and biliary system. | 1. A method of treating obesity comprising: Permanently or temporarily lowering one or more of the following aspects of a subjects gallbladder: motility, ejection fraction, turnover rate, bile flow rate; wherein said lowered aspect causes the consumption of certain amounts and types of food to result in physical discomfort in the subject. 2. The method in claim 1 further comprising the step of applying energy to a portion of the gallbladder or cystic duct. 3. The method in claim 1 further comprising the step of causing a stricture or stenosis in at least a porting of the gallbladder or cystic duct. 4. The method in claim 1 further comprising the step of inserting an implant within the gallbladder or cystic duct wherein said device at least partially limits the flow of bile or interferes with the contraction of the gallbladder. 5. The method in claim 1 further comprising the step of surgically or chemically altering or damaging the tissue of the gallbladder or cystic duct. 6. The method in claim 1 further comprising the step of causing neuropathic damage or paralyzing a portion of the cystic duct. 7. The method in claim 1 further comprising the step of diminishing the capacity of duodenum to deliver CCK or inhibiting the capacity of the gallbladder to absorb CCK. 8. The method in claim 1 further comprising the step pacing at least a portion of the biliary system. 9. The method in claim 1 further comprising the step of causing neuropathic damage or paralyzing a portion of the gallbladder. 10. The method in claim 1 further comprising the step of monitoring an aspect of the biliary system. 11. The method in claim 1 further comprising the step of lowering the evacuation fraction of the gallbladder to between 80 and 10 percent. 12. The method in claim 1 further comprising the step of placing on the gallbladder a clip-like device operable to constrict or prevent the expansion of at least a portion of the gall bladder. 13. The method in claim 1 further comprising the step of injecting the gallbladder with botulism toxin. 14. The method in claim 1 further comprising the step of inserting one or more artificial gallstones in the gallbladder. 15. The method in claim 1 further comprising the step of implanting a drug or agent eluting patch, stent, or device on or within the gallbladder. 16. The method in claim 1 further comprising the step changing the orientation and shape of the gallbladder and cystic duct by torsion. 17. A method of encouraging a subject to eat a diet limited in fat comprising the steps of: accessing an internal or external surface of a cystic duct; and causing or allowing the temporary or permanent stricture or stenosis of a cystic duct such that flow through the duct is at least partially restricted. 18. The method in claim 17 further comprising the step of causing at least partial torsion of the cystic duct thereby restricting flow therethrough. 19. A method of modulating gall bladder ejection frac... | - | A61F002/04 A61N005/067 A61B005/107
US20080243217 | A1 | 2008-10-02 | CARDIAC STIMULATION APPARATUS | 2007-11-08 | Cardiac stimulation apparatus (10) such as a lead comprising one or more electrodes (17, 19 and 21) for electrical contact with heart or other tissue. An adhesive substance is provided for adhesively attaching the electrodes in position in relation to the heart or other tissue to provide a non-traumatic fixing procedure. The electrodes (17, 19, 21) are positioned on a support (15) to which the adhesive substance is applied attached. | 1. Cardiac stimulation apparatus comprising an electrode support, an electrode provided on the support for electrical contact with internal tissue within the body of a patient, and means for releasably attaching the support in position in relation to the internal tissue, an electrical lead defining an electroconductive path for delivering stimulation to the electrode, the support being forcibly removable from the internal tissue and also collapsible upon application of a pulling force to the electrical lead to facilitate withdrawal of the support from the body of the patient along with the electrical lead. 2. A cardiac stimulation apparatus according to claim 1 wherein the support is releasably attachable by adhesive attachment comprising an adhesive substance on the support for the electrode. 3. A cardiac stimulation apparatus according to claim 2 wherein the support comprises a wafer. 4. A cardiac stimulation apparatus according to claim 2 wherein the support comprises an adhesive gel pad. 5. A cardiac stimulation apparatus according to claim 1 wherein the electrode is provided at a distal end thereof and a proximal end thereof adapted to be connected to the source of electrical stimulation. 6. A cardiac stimulation lead comprising an electrode support at one end of the lead, first and second electrodes mounted in spaced relation on the electrode support, and a third electrode mounted on the electrode support in spaced relation to said first and second electrodes, said first and second electrodes being dimensioned for cardiac pacing and the third electrode having a larger electrical contact area than the first and second electrodes for cardiac defibrillation, and means for providing an electroconductive path between each electrode and the other end of the lead for electrical connection to a source of electrical stimulation, means for releasably attaching the support in position in relation to internal tissue within the body of a patient, the support being forcibly removable from the internal tissue and also collapsible upon application of a pulling force to the cardiac stimulation lead to facilitate withdrawal of the support from the body of the patient. 7. A cardiac stimulation lead according to claim 6 wherein the third electrode has a large surface area and extends for at least a substantial part around the first and second electrodes. 8. A cardiac stimulation lead according to claim 6 further comprising a further electrode having a larger electrical contact area than the first and second electrodes so that bipolar cardiac defibrillation can be performed. 9. A cardiac stimulation lead according to claim 6 wherein the electrode support comprises a wafer. 10. A cardiac stimulation lead according to claim 9 wherein the wafer is adapted to be placed against part of the heart or other tissue requiring stimulation with the necessary electrode or electrodes in electrical contact therewith. 11. A cardiac stimulation lead according to claim... | - | A61N001/00 A61B005/04
US20080249384 | A1 | 2008-10-09 | System and Method for Estimating the Glucose Concentration in Blood | 2005-09-05 | This application relates to glucose monitoring systems for continuously measuring the glucose concentration in a patients blood. The system is adapted to communicate with one or more sensors (10, 20) for transcutaneous insertion into a patient and for producing sensor signals related to the glucose concentration. The system comprises an electronic calculator unit and a display (14) for displaying the measured glucose concentration. The electronic calculator unit further comprises means for calculating an estimate of the uncertainty, i.e. the degree of accuracy of the glucose measurement, and the display is configured for displaying an interval representing said uncertainty (15). | 1. A system for measuring glucose concentration comprising a transcutaneous sensor, an electronic calculator unit and a display for displaying the measured glucose concentration, characterized in that the electronic calculator unit has means for calculating an estimate of the uncertainty of the glucose measurement; and that the display is configured for displaying an interval representing said uncertainty. 2. A system according to claim 1 characterized in that the display is configured for graphical representation of the interval. 3. A system according to claim 2, characterized in that the display is configured for being able to produce various types of graphical representations of the interval. 4. A system according to claim 3, characterized in that the sensor comprises an electronic circuit configured for being able to communicate with the electronic calculator unit. 5. A system according to claim 4, characterized in that the electronic circuit in the sensor contains calibration information. 6. A system according to claim 5, characterized in that the electronic calculator unit comprises data storage for calibration information. 7. A system according to claim 6, characterized in that the data storage is configured for receiving information used in the electronic calculator unit for iteratively estimating the uncertainty of the blood-glucose measurement. 8. A system according to claim 7, characterized in that the apparatus has means for generating said data/information. 9. A system according to claim 8, characterized in that said means comprise a test-strip glucose-measurement device. 10. A system according to claim 8, characterized in that said means comprise a further transcutaneous sensor. 11. A system according to claim 1, characterized in that the apparatus comprises transmitter and receiver circuits for wireless communication of said data/information between the separate parts of the apparatus. 12. A system according to claim 1, characterized in that the apparatus comprises an electronic alarm circuit; and that the apparatus comprises means by which a user is able to adjust the threshold values of the alarm circuit. 13. A system according to claims 1, characterized in that the apparatus is configured for being able to communicate with a unit intended for dosing a medicament. 14. A system according to claim 1, characterized in that it comprises means for dosing and administering a medicament. 15. A method of estimating the glucose concentration in blood by means of a sensor inserted subcutaneously in live tissue and by means of associated electronic calculator circuits, characterized in that an estimate is provided of the uncertainty of the glucose-concentration measurement; and that a result is displayed on a display comprising the display of an interval that represents the estimated uncertainty. 16. A method according to claim 15, characterized in that the uncertainty is reduced by performing valid calibration measurements. ... | NOVO Nordisk A/S | A61B005/055
US20080243025 | A1 | 2008-10-02 | Medical Device | 2004-12-23 | In a medical device and a method for operating the medical device, it is first determined whether a patient, at whom a medical measurement is to be made, satisfied specified criteria that will ensure comparability of the measurement results obtained from the patient. Only when the specified criteria had been satisfied is an electrical bio-impedance signal obtained from the patient. The cardiac component of the electrical bio-impedance signal is extracted, and is analyzed to identify a change in a medical condition of the patient. | 1.-37. (canceled) 38. A method for detecting a change in a medical condition of a patient, comprising the steps of: determining that a patient satisfies specified criteria ensuring comparability of measurements results obtained from the patient;only when said patient satisfies said specified criteria, obtaining electrical bio-impedance signals from the patient respectively as separated points in time;extracting a cardiac component of each electrical bio-impedance signal, said cardiac component of each electrical bio-impedance signal being associated with a medical condition of the patient; andanalyzing said cardiac components of the respective electrical bio-impedance signal to identify a change in said medical condition. 39. A method as claimed in claim 38 comprising determining that the patient satisfies said specified criteria by determining when the patient is in a specified body position. 40. A method as claimed in claim 39 comprising, after determining that the patient is in said specified body position, delaying obtaining said electrical bio-impedance signal from the patient for a predetermined period of time. 41. A method as claimed in claim 38 comprising determining that the patient satisfies said specified criteria by sensing a heart rate of the patient, determining whether the heart rate is a within a predetermined range, and obtaining said electrical bio-impedance signal from the patient only if the heart rate is within said predetermined range. 42. A method as claimed in claim 38 comprising determining that the patient satisfies said specified criteria by sensing a activity level of the patient, determining whether the activity level is a within a predetermined range, and obtaining said electrical bio-impedance signal from the patient only if the activity level is within said predetermined range. 43. A method as claimed in claim 38 comprising electronically storing the respective electrical bio-impedance signals. 44. A method as claimed in claim 38 wherein the step of obtaining said electrical bio-impedance signals from the patient comprises measuring an intrathoracic impedance of the patient. 45. A method as claimed in claim 44 wherein the step of measuring the intrathoracic impedance of the patient comprises: applying an excitation current pulse between a first electrode located within the heart of the patient, and a second electrode; andmeasuring said electrical bio-impedance in tissue located between said first electrode and said second electrode in response to said excitation current pulse. 46. A method as claimed in claim 44 comprising analyzing said cardiac components of the respective electrical bio-impedance signals to detect insipient pulmonary edema. 47. A method as claimed in claim 38 comprising extracting said cardiac component of each electrical bio-impedance signal by: applying an excitation current pulse between a first electrode located at a first position within the heart of the patient and a second electr... | - | A61B005/053
US20080275349 | A1 | 2008-11-06 | MONITORING, PREDICTING AND TREATING CLINICAL EPISODES | 2008-05-01 | Apparatus (10) is provided that includes at least one sensor (30), configured to sense a physiological parameter of a subject (12) and to sense large body movement of the subject (12), an output unit (24), and a control unit (14). The control unit (14) is configured to monitor a condition of the subject (12) by analyzing the physiological parameter and the sensed large body movement, and to drive the output unit (24) to generate an alert upon detecting a deterioration of the monitored condition. Other embodiments are also described. | 1. Apparatus comprising: at least one sensor, configured to sense a physiological parameter of a subject and to sense large body movement of the subject;an output unit; anda control unit, configured to: monitor a condition of the subject by analyzing the physiological parameter and the sensed large body movement, anddrive the output unit to generate an alert upon detecting a deterioration of the monitored condition. 2. The apparatus according to claim 1, wherein the control unit is configured to determine an activity level of the subject based on sensed large body movements of the subject, and to monitor the condition of the subject by analyzing the physiological parameter in combination with the activity level of the subject. 3. The apparatus according to claim 1, wherein the physiological parameter is a respiratory rate of the subject, and wherein the at least one sensor is configured to sense the respiratory rate. 4. The apparatus according to claim 1, wherein the physiological parameter is a heart rate of the subject, and wherein the at least one sensor is configured to sense the heart rate. 5. The apparatus according to claim 1, wherein the physiological parameter is a blood oxygen level of the subject, and wherein the at least one sensor is configured to sense the blood oxygen level. 6. The apparatus according to claim 1, wherein the sensor comprises a pulse oximeter. 7. The apparatus according to claim 1, wherein the at least one sensor comprises a first sensor configured to sense the physiological parameter, and a second sensor configured to sense the large body movement. 8. The apparatus according to claim 1, wherein the at least one sensor comprises a same sensor that senses both the physiological parameter and the large body movement. 9. The apparatus according to claim 1, wherein the at least one sensor is configured to sense the physiological parameter by deriving the physiological parameter from the large body movement. 10. The apparatus according to claim 1, wherein the control unit is configured to: receive a specified range of values for the physiological parameter, anddrive the output unit to generate the alert only upon finding that the sensed physiological parameter falls outside the specified range over 50% of the times it is sensed during a period having a duration of at least 30 seconds. 11. The apparatus according to claim 1, wherein the control unit is configured to: receive a specified range of values for the physiological parameter,calculate a representative value of the physiological parameter responsively to sensing the physiological parameter at least once every 10 seconds during a period having a duration of at least 30 seconds, anddrive the output unit to generate the alert only upon finding that the representative value of the physiological parameter falls outside the specified range during the period. 12. The apparatus according to claim 1, wherein the condition includes pressure sores of the subject, ... | EARLYSENSE LTD. | A61B005/02
US20080312553 | A1 | 2008-12-18 | INTRACORPOREAL PRESSURE MEASUREMENT DEVICES AND METHODS | 2008-05-22 | The invention relates to devices, systems, and methods for the measurement of a pressure within a body that is adjusted to compensate for variations in local atmospheric pressure. A pressure measurement system can include an implantable target pressure sensor, an implantable internal reference pressure sensor located remotely from the target pressure sensor, an external reference pressure sensor configured to transmit a telemetric signal that is indicative of the local atmospheric pressure, and at least one condition indicator. The implantable medical device system further includes a controller configured to determine a correlation factor based on a signal from the implantable reference pressure sensor and the signal from the external reference pressure sensor. | 1. A pressure measurement system, comprising: an implantable target pressure sensor;an implantable internal reference pressure sensor located remotely from the target pressure sensor;an external reference pressure sensor configured to wirelessly transmit a local atmospheric pressure signal; anda controller in communication with the implantable target pressure sensor and the implantable internal reference pressure sensor, the controller configured to: determine a correlation factor between a signal from the implantable internal reference pressure sensor and the local atmospheric pressure signal from the external reference pressure sensor at times when the external reference pressure sensor is in communication with the controller;determine a calibrated reference pressure based on a signal from the implantable internal reference pressure sensor and a correlation factor at times when the external reference pressure sensor is not in communication with the controller; anddetermine a relative physiological target pressure valueat times when the external reference pressure sensor is in communication with the controller, by adjusting the signal from the target pressure sensor by the signal from the external reference pressure sensor; andat times when the external reference pressure sensor is not in communication with the controller, by adjusting the signal from the target pressure sensor by a calibrated reference pressure. 2. The pressure measurement system of claim 1, further including at least one condition indicator in communication with the controller. 3. The pressure measurement system of claim 2, wherein the at least one condition indicator includes a body motion sensor. 4. The pressure measurement system of claim 2, wherein the at least one condition indicator includes a body posture sensor. 5. The pressure measurement system of claim 2, wherein the at least one condition indicator includes a time of day sensor. 6. The pressure measurement system of claim 2, wherein the at least one condition indicator includes a pulmonary function sensor. 7. The pressure measurement system of claim 1, wherein the controller is configured to determine a correlation factor between a signal from the implantable internal reference pressure sensor and the local atmospheric pressure signal from the external reference pressure sensor by subtracting the value of the external reference pressure sensor from the value of the implantable internal reference pressure sensor. 8. The pressure measurement system of claim 1, wherein the correlation factor is stored in a calibration matrix. 9. The pressure measurement system of claim 2, wherein the controller is configured to categorize a state of a patient based on a signal from the at least one condition indicator. 10. The pressure measurement system of claim 2, wherein the controller is configured to evaluate the rate of change of a signal from the at least one condition indicator. 11. A pressure measurement system comp... | - | A61B005/03
US20080275368 | A1 | 2008-11-06 | Apparatus For Detecting Axial Force In The Digestive System | 2004-08-04 | An apparatus and a method for measuring deformations and force applied to a probe are disclosed. The system may be a mechanical system a physical system or a biological system such as e.g. a bodily hollow system. The apparatus comprises an elongated elastic probe, a conducting medium attached to or contained by the probe, two or more electrodes being electrically connected by the conducting medium, the electrodes being attached to the probe, and the apparatus furthermore comprising means for measuring an electrical parameter, such as the potential difference between at least two of the number of electrodes, the measured electrical parameter being indicative of a deformation of the probe in at least the longitudinal direction of the elongated probe. The force applied to the probe may be determined from a pre-calibration of the electrical parameter-force relationship of the probe. | 1-35. (canceled) 36. An apparatus for measuring the deformation of a system, the apparatus comprising: an elongated elastic probe,a conducting medium attached to or contained by the probe, andtwo or more electrodes being electrically connected by the conducting medium, the electrodes being attached to the probe,wherein the apparatus furthermore comprising means for measuring an electrical parameter between at least two of the number of electrodes, the measured electrical parameter being indicative of a deformation of the probe in at least the longitudinal direction of the elongated probe. 37. An apparatus according to claim 36 comprising four or more electrodes, wherein at least two of the four or more electrodes are measuring electrodes comprising means for measuring the electrical potential between them, and wherein other at least two of the four or more electrodes are generating electrodes comprising means for generating an AC-field between the measuring electrodes. 38. An apparatus according to claim 36, wherein the measured electrical parameter is indicative of a force of a certain magnitude being applied to the probe. 39. An apparatus according to claim 38, wherein the force or the deformation are deduced from the measured electrical parameter by means of a pre-determined calibration function. 40. An apparatus according to claim 36, wherein the apparatus further comprising timer means for determining a timing of a change of the measured electrical parameter. 41. An apparatus according to claim 36, wherein the more than two electrodes are placed along a reference curve of the probe. 42. An apparatus according to claim 41, where the reference curve is a longitudinal axis extending along the elongation of the elongate probe, and thereby the more than two electrodes are placed along the longitudinal axis. 43. An apparatus according to claim 36, wherein the conducting medium is a liquid medium serving as an electrolyte for conducting the electric current between the electrodes. 44. An apparatus according to claim 43, wherein the liquid medium is a liquid preferably non-harmful to the bodily hollow system or the engineered structure being stimulated, such as an acid like HCl in the stomach, or such as bile salts in the small intestine, or such as water with NaCl in the esophagus. 45. An apparatus according to claim 36, wherein the conducting medium is a solid medium, such as compounds including at least one substance selected from the group of: soft metals, polymers, ceramics, composites and natural materials. 46. An apparatus according to claim 36, further comprising at least one inflatable balloon situated between a proximal end and a distal end of the probe, and the apparatus comprising means for passing an inflating fluid, preferably a liquid, from the proximal end to the balloon, and where the apparatus optionally is provided with means for measuring at least one physical properties of the balloon. 47. An apparatus according to c... | Ditens A/S | A61B005/00
US20080306398 | A1 | 2008-12-11 | BRAIN WAVE DETECTING APPARATUS | 2008-06-05 | A brain wave detecting apparatus includes a plurality of brain wave detecting portions that are arranged on a side surface of a head and detect brain waves of the head, a selecting portion selecting one of the plurality of brain wave detecting portions on the basis of brain wave signals transmitted from the plurality of brain wave detecting portions, and a transmitting portion transmitting information about the brain wave signal transmitted from said one of the plurality of brain wave detecting portions. | 1. A brain wave detecting apparatus comprising: a plurality of brain wave detecting portions that are arranged on a side surface of a head and detect brain waves of the head;a selecting portion selecting one of the plurality of brain wave detecting portions on the basis of brain wave signals transmitted from the plurality of brain wave detecting portions; anda transmitting portion transmitting information about the brain wave signal transmitted from said one of the plurality of brain wave detecting portions. 2. The brain wave detecting apparatus as claimed in claim 1, wherein the plurality of the brain wave detecting portions are arranged on the side surface of the head in a vertical direction. 3. The brain wave detecting apparatus as claimed it claim 1, wherein the plurality of the brain wave detecting portions are arranged on the side surface of the head in a lateral direction. 4. The brain wave detecting apparatus as claimed in claim 1, wherein the plurality of the brain wave detecting portions are arranged in vertical and lateral directions on the side surface of the head so as to form a cross shape. 5. The brain wave detecting apparatus as claimed in claim 1, further comprising a hat on which the plurality of the brain wave detecting portions are arranged. | FUJITSU COMPONENT LIMITED | A61B005/04
US20080287781 | A1 | 2008-11-20 | Registration Methods and Apparatus | 2005-03-07 | A method, system and computer program for generating a registered image of a body part of a patient for use in an image guided surgical procedure is described. The method includes attaching a marker detectable by a tracking system to the body part before carrying out surgical steps. The tracking system has a reference frame and the position of the marker in the reference frame is detected. A first image of the body part is captured using an imaging system. An indication of the position of the first image relative to the reference frame of the tracking system is obtained. The first image is then mapped into registration with the position of the body part. The system includes an imaging system for capturing images of the body part. A tracking system detects a marker and determines the position of the marker in the reference frame of the tracking system. A marker is attachable to the body part and detectable by the tracking system. A computer control system is configured to obtain an indication of the position of the first image relative to the reference frame of the tracking system and to map the first image into registration with the position of the body part. | 1. A method for generating a registered image of a body part of a patient for use in a computer aided surgical procedure, the method comprising: attaching a marker detectable by a tracking system to the body part prior to any surgical steps of the surgical procedure, the tracking system having a reference frame;detecting the position of the marker in the reference frame;capturing at least a first image of the body part using an imaging system;obtaining an indication of the position of the first image relative to the reference frame of the tracking system; anddetermining a mapping to bring the first image into registration with the position of the body part. 2. The method of claim 1, and further comprising the step of mapping the first image into registration with the position of the body part in the reference frame of the tracking system. 3. The method of claim 1, wherein the step of obtaining an indication of the position of the at least first image relative to the reference frame of the tracking system includes the step of detecting the position in the reference frame of the tracking system of a further marker attached to a part of the imaging system using the tracking system. 4. The method of claim 1, wherein the first image includes the marker and at least a part of the body part, and wherein the position of the marker is detected when the first image is captured thereby providing the indication. 5. The method of claim 2, and further comprising the step of displaying the registered image during the computer aided surgical procedure. 6. The method of claim 5, wherein the surgical procedure is an orthopaedic procedure. 7. The method of claim 1, wherein the step of attaching the marker includes implanting the marker in a bone of the patient. 8. The method of claim 7, wherein the step of implanting the marker includes percutaneously implanting the marker. 9. The method of claim 1, wherein the marker is wirelessly detectable at radio frequencies by the tracking system. 10. The method of claim 1, wherein the imaging system is an X-ray system. 11. The method of claim 11, wherein the position of the marker is detected with the patient standing. 12. The method of claim 1, wherein the marker is wirelessly tracked using a magnetic tracking system. 13. The method of claim 10, wherein the step of obtaining an indication of the position of the at least first image relative to the reference frame of the tracking system includes the step of determining the position of an X-ray detector in the reference frame of the tracking system. 14. The method of claim 10, and further comprising the step of capturing a second image of the body part using the X-ray system, and wherein the second image is in a second direction different to a first direction in which the first image was captured. 15. The method of claim 14, wherein the step of capturing the second image includes moving the patient relative to the X-ray system. 16. The method of claim 14, where... | DEPUY INTERNATIONAL LIMITED | A61B005/05
US20080242948 | A1 | 2008-10-02 | Effective low-profile health monitoring or the like | 2007-03-30 | A system, method, computer program product, and carrier are described for obtaining data from occupational or leisure intercommunication at least between a device and a user or receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user; and applying one or more data extraction criteria to the health-status-indicative data surreptitiously captured from the interaction between the device and the user or signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication. | 1-27. (canceled) 28. A method comprising: obtaining data from occupational or leisure intercommunication at least between a device and a user; andsignaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication. 29. (canceled) 30. The method of claim 28 in which obtaining data from occupational or leisure intercommunication at least between a device and a user comprises: receiving an identifier of the user with the data from the occupational or leisure intercommunication. 31-33. (canceled) 34. The method of claim 28 in which signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication comprises: causing a module to process at least a natural language expression from the data. 35. (canceled) 36. The method of claim 28 in which signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication comprises: applying one or more normalcy criteria to at least a portion of the data from the occupational or leisure intercommunication. 37. (canceled) 38. The method of claim 28 in which signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication comprises: generating an expression indicating a composition of matter at least partly based on the data. 39-43. (canceled) 44. The method of claim 28 in which signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication comprises: removing a portion of the data from the occupational or leisure intercommunication. 45-47. (canceled) 48. The method of claim 28 in which obtaining data from occupational or leisure intercommunication at least between a device and a user comprises: configuring a data acquisition mode of the device in response to information about the user. 49. (canceled) 50. The method of claim 28 in which obtaining data from occupational or leisure intercommunication at least between a device and a user comprises: obtaining auditory information in the data from the occupational or leisure intercommunication. 51-54. (canceled) 55. A system comprising: means for obtaining data from occupational or leisure intercommunication at least between a device and a user; andmeans for signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication. 56. The system of claim 55 in which the means for obtaining data from occupational or leisure intercommunication at least between a device and a user comprises: means for receiving the data from the occupational or leisure intercommunication remotely. 57. The system of cl... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00
US20080249394 | A1 | 2008-10-09 | METHOD FOR IMPROVED ROTATIONAL ALIGNMENT IN JOINT ARTHROPLASTY | 2008-04-03 | A method for improved rotational alignment of the bones in joint surgery is described. The method involves the tracking of the relative motion of a third bone in with respect to the movement of the first and second bone. In one aspect of the invention, the motion of the patella is used to derive the axis of rotation of the femoral and tibial components in total knee arthroplasty, either alone or in combination with other techniques. | 1. A method of determining the rotational alignment axis of bone joints involving a first and a second bone, comprising: tracking the position of a third bone with reference to the first or the second bone;obtaining a trajectory of the third bone, wherein the trajectory is generated by the movement of the bone joint; andcalculating an axis of alignment of the first or second bone at the joint using the relative movement of the third bone with reference to either the first or the second bone. 2. The method of claim 1, wherein the tracking comprises using imaging to track the third bone with reference to the first or the second bone. 3. The method of claim 2, wherein the imaging comprises X-ray imaging, computed tomography imaging, fluoroscopic imaging, magnetic resonance imaging, or ultrasound imaging. 4. The method of claim 1, wherein the tracking comprises using a tracker attached to the third bone. 5. The method of claim 4, wherein the tracker is an optical sensor. 6. The method of claim 4, wherein the tracker is an electromagnetic sensor. 7. The method of claim 4, wherein the tracker is a mechanical sensor. 8. The method of claim 1, wherein the first bone is the femur, the second bone is the tibia and the third bone is the patella. 9. The method of claim 1, wherein the first bone is the femur and the second bone is the pelvis and the third bone is the tibia. 10. The method of claim 1 wherein the first bone is the femur, the second bone is the tibia and the third bone is the talus. 11. A method of determining the rotational alignment axis of the femur at the knee joint, comprising: tracking the position of the patella within the femoral groove as the joint is flexed or extended, using a tracker attached to the patella;obtaining the trajectory of the patella with reference to the femoral groove as the knee is flexed or extended; andcalculating the axis of rotational alignment of the femur at the knee joint from data indicating a curve of motion of the patella by fitting an appropriate function to the data. 12. The method of claim 11, further comprising: projecting the fitted function on to a cross-sectional plane perpendicular to the anatomical axis of the femur. 13. The method of claim 11 where the data is projected on to a plane perpendicular to the mechanical axis of the femur. 14. The method of claim 11, wherein the fitting function is a spline function. 15. The method of claim 11, wherein the fitting function is a polynomial function. 16. The method of claim 11, wherein the data is filtered using a Fourier filtering scheme. 17. A method of determining the rotational alignment axis of the tibia at the knee joint, comprising: tracking the position of the patella with reference to the tibia as the joint is flexed or extended, using a tracker attached to the patella;obtaining the trajectory of the patella with reference to proximal end of the tibia as the knee is flexed or extended; andcalculating the axis of rotational alignmen... | The Board of Trustees of the Leland Stanford Junior University | A61B005/00 A61B017/56
US20080298549 | A1 | 2008-12-04 | Web-Based Expert System for Educational Therapy Planning | 2006-11-30 | The invention consists of an expert system supporting non-experts in their capacity to develop, implement, monitor and evolve therapy plans. Therapy plans in this application are educational and non-medical. The purpose of the therapy plan is to guide the caregiver or therapist in what to do for a client and how to do it. The client is a person with a previously diagnosed condition (psychological, cognitive, or physical) which responds to a teaching process. The plans are developed based on expert system analysis of user input. This analysis is reflective of decision making processes in the fields of both special needs education and rehabilitative medicine, including speech language pathology, psychology, physiotherapy, and occupational therapy. | 1. An automated expert system-based method for planning, tailoring and evolving therapy programs using a web or wireless delivery system, comprising: a) collecting user input from one or more users concerning a client diagnosed with a medical or psychological condition amenable to an educational therapy program and concerning preferences and priorities of persons responsible for said client, said persons including: parents, caregivers, therapists, educators;b) analyzing said user input using an expert algorithm to create a client profile;c) providing a treatment plan based on said client profile, said treatment plan consisting of treatment activities and a set of priority-ranked therapy goals, said priority ranking determined by said expert algorithm; andd) repeating steps a) to c) on an ongoing basis to track client progress and provide updated treatment plans for said client. 2. The method of claim 1, wherein said treatment plan includes a combination of support plans for preparing and introducing activities and strategy plans for conducting and practicing activities. 3. The method of claim 2, wherein said strategy plans further include multimedia demonstrations of activities. 4. The method of claim 1, wherein said user is capable of overriding rankings determined by said expert algorithm. 5. The method of claim 1, further including a step e) of analyzing client progress, feedback, and treatment plans for multiple clients to develop improvements to said expert algorithm. 6. An automated expert system for planning, tailoring and evolving therapy programs, comprising: a) a collection unit for collecting user input from one or more users concerning a client diagnosed with a medical or psychological condition amenable to an educational therapy program and concerning preferences and priorities of persons responsible for said client, said persons including: parents, caregivers, therapists, educators;b) an analysis unit for analyzing said user input using an expert algorithm to create a client profile and for providing a treatment plan based on said client profile, said treatment plan consisting of treatment activities and a set of priority-ranked developmental goals, said priority ranking determined by said expert algorithm; andd) a tracking unit, operating on an ongoing basis to track client progress and provide updated treatment plans for said client. 7. The system of claim 6, further including a database unit for storing articles, media and other related information for access by users of the system. 8. The system of claim 6, further including a library of demonstration media, including audio and visual media, for demonstrating and assisting the user in conducting planned activities. 9. The system of claim 6, wherein said system is implementable as a web-based client-server system. 10. The system of claim 6, further including a data-mining unit for analyzing client progress, feedback, and treatment plans for multiple clients to develop im... | VIRTUAL EXPERT CLINICS INC. | A61N005/10 A61B005/00
US20080275326 | A1 | 2008-11-06 | SENSOR FOR MONITORING A CONDITION OF A PATIENT | 2007-05-01 | A sensor may include a substrate having a sensing portion defining a sensor thereon and a circuit mounting portion defining at least one electrically conductive pad that is electrically connected to the sensor. The sensor may be configured to produce a signal indicative of a condition of the patient. An anisotropic medium may be disposed on the circuit mounting portion and may be electrically conductive in a direction through the medium and electrically insulating in directions along the medium. An electrical circuit may be mechanically mounted to the circuit mounting portion of the first substrate via the anisotropic medium with at least one electrically conductive terminal juxtaposed over the at least one electrically conductive pad. The anisotropic medium may establish local electrical contact between the at least one electrically conductive terminal and the at least one electrically conductive pad. | 1. A sensor comprising: a first substrate having a sensing portion configured to be percutaneously inserted into a patient and an extracorporeal circuit mounting portion, the sensing portion defining a sensor thereon and the circuit mounting portion having at least one electrically conductive pad formed thereon that is electrically connected to the sensor, the sensor configured to produce a signal indicative of a condition of the patient,an anisotropic medium disposed on the circuit mounting portion, the medium configured to be electrically conductive in a direction through the medium and electrically insulating in directions along the medium,an electrical circuit having at least one electrically conductive terminal, the circuit being mechanically mounted to the circuit mounting portion of the first substrate via the anisotropic medium with the at least one electrically conductive terminal juxtaposed over the at least one electrically conductive pad, the anisotropic medium establishing local electrical contact between the at least one electrically conductive terminal and the at least one electrically conductive pad. 2. The sensor of claim 1 wherein the first substrate is one of a flexible substrate and a rigid substrate. 3. The sensor of claim 1 wherein the anisotropic medium is an anisotropic adhesive tape, the tape being configured to mechanically bond the electrical circuit to the circuit mounting portion of the first substrate, the tape further being electrically insulating along a plane of the tape and electrically conductive through the tape in a direction generally perpendicular to the plane of the tape. 4. The sensor of claim 1 wherein the anisotropic medium is an anisotropic elastomer configured to mechanically bond the circuit to the circuit mounting portion of the first substrate, the anisotropic elastomer being electrically insulating along a plane generally parallel to opposing surfaces of the electrical circuit and the circuit mounting portion of the first substrate, and electrically conductive through the elastomer in a direction generally perpendicular to the plane. 5. The sensor of claim 1 wherein the first substrate defines a first number of electrically conductive pads on the circuit mounting portion thereof, and the electrical circuit has a second number of electrical terminals, and wherein at least some of the second number of electrical terminals align with at least some of the number of electrically conductive pads when the electrical circuit is mounted to the circuit mounting portion of the first substrate. 6. The sensor of claim 1 wherein the electrical circuit comprises: a second substrate defining the at least one electrical terminal, anda number of electrical components mounted to the second substrate and electrically interconnected to form a sensor control circuit, the sensor control circuit being electrically connected to the at least one electrical terminal. 7. The sensor of claim 6 wherein the sensor control ... | - | A61B005/00
US20080319307 | A1 | 2008-12-25 | Method for medical imaging using fluorescent nanoparticles | 2007-06-19 | A first method for medically imaging a patient includes obtaining a carrier medium containing fluorescent nanoparticles. The carrier medium containing the fluorescent nanoparticles is disposed inside the patient at a site where the carrier medium will migrate to specific tissue of known shape. The nanoparticles are activated to fluoresce. Image data of the patient is obtained including the specific tissue after the specific tissue has picked-up the fluorescing fluorescent nanoparticles. The known shape of the specific tissue is identified in the image data. An image representation of the image data is created using the identified known shape as a fiducial. An image of the image representation of the image data is displayed. In a second method, the carrier medium is injected inside tissue of the patient at a site where the carrier medium will be localized in the injected tissue, and the injecting deposits the medium in a predetermined shape in the tissue. | 1. A method for medically imaging a patient comprising: a) obtaining a carrier medium containing fluorescent nanoparticles;b) disposing the carrier medium containing the fluorescent nanoparticles inside the patient at a site where the carrier medium will migrate to, and be picked up by, specific tissue of the patient, wherein the specific tissue has a known shape;c) activating the fluorescent nanoparticles to fluoresce;d) obtaining image data of the patient including the specific tissue after the specific tissue has picked-up the fluorescing fluorescent nanoparticles;e) identifying the known shape of the specific tissue in the image data;f) creating an image representation of the image data using the identified known shape as a fiducial; andg) displaying an image of the image representation of the image data. 2. The method of claim 1, wherein the image data is magnetic resonance (MRI) image data. 3. The method of claim 1, wherein the image data is computerized tomography (CT) image data. 4. The method of claim 1, wherein the image data is positive electron emission (PET) images 5. The method of claim 1, wherein the image data is fluoroscope image data. 6. The method of claim 1, wherein steps e) and f) are performed by a computer. 7. The method of claim 6, wherein step e) is performed by the computer using a segmentation subroutine. 8. The method of claim 1, also including the steps of: h) obtaining additional image data of the patient including the specific tissue;i) creating an image representation of the additional image data; andj) displaying an image of the image representation of the additional image data registered and overlaid with the displayed image of the image representation of the image data. 9. The method of claim 8, wherein the additional image data and the image data are obtained from different imaging modalities. 10. The method of claim 8, wherein the image data is pre-acquired image data and wherein the additional image data is real-time image data. 11. A method for medically imaging a patient comprising: a) obtaining a carrier medium containing fluorescent nanoparticles;b) injecting the carrier medium containing the fluorescent nanoparticles inside tissue of the patient at a site where the carrier medium will be localized in the injected tissue, wherein the injecting deposits the carrier medium in a predetermined shape in the injected tissue;c) activating the fluorescent nanoparticles to fluoresce;d) obtaining image data of the patient including the injected tissue;e) identifying the predetermined shape of the injected carrier medium in the image data;f) creating an image representation of the image data using the identified predetermined shape as a fiducial; andg) displaying an image of the image representation of the image data. 12. The method of claim 11, wherein the image data is magnetic resonance (MRI) image data. 13. The method of claim 11, wherein the image data is computerized tomography (CT) image data. 14... | Ethicon Endo-Surgery, Inc. | A61B005/055
US20080255707 | A1 | 2008-10-16 | Method and system for categorizing blood glucose tests at test time in a portable device or later in a downloading program and then analyzing the categories separately | 2007-04-10 | A four-part method for analyzing data of a patient's blood glucose (BG) tests or other parameters e.g. carbohydrate intakes and discrete insulin doses (boluses): 1. An LCD-button on the portable glucose meter or insulin pump, that allows flagging or assigning a category to each blood glucose test at test-time. Examples of categories: pre-meal BG tests, post-meal BG tests. 2. In the computer-based downloading program, features for changing the boundaries of the meal-oriented time-intervals by the faster methods of: a) mouse-dragging the time-boundary lines on a scatter chart, b) placing stretchable rectangles around the data points of interest. 3. Features in the computer-based downloading program, for digitally identifying the time of the meal and digitally assigning the category to the pre-meal BG tests and to the post-meal BG tests. 4. Features in the computer-based downloading program, for statistically analyzing and graphically-displaying the differently-categorized data separately or together in combinations. | 1. A method for manually assigning digital categories to blood glucose tests (BG's) in a diabetic patient's portable blood glucose meter, insulin pump or other portable device by the patient at the time of testing by manual activation of a physical button or a selector on an LCD screen, wherein the categories may include but are not limited to pre-meal, post-meal, disregard, or other categories determined by the user, or wherein the button or selector has no label but its meaning may be understood by the patient and the practitioner, wherein the activation of the button or selector may allow one of two values or may allow a text string. 2. The method of claim 1 wherein all the methods described for BG's may also be used for meal boluses, correction boluses, and carbohydrate intakes. 3. (canceled) 4. The method of claim 1 wherein the category is a variable having two possible values, indicating whether the button or selector was pressed or not pressed. 5. The method of claim 4 wherein the category variable is binary. 6. The method of claim 1 wherein there is a plurality of category variables each having two possible values, indicating that the selector was activated or not activated. 7. The method of claim 1 wherein the category variable or plurality of category variables are short text strings. 8. The method of claim 7 wherein each category chosen for each variable is chosen from a fixed drop-down list in the LCD screen of the portable device. 9. The method of claim 7 wherein each category is chosen from a drop-down list in the LCD screen that can be added to through an alphanumeric input feature in the portable device. 10. The method of claim 9 wherein the portable device has alphanumeric input capability by virtue of being built into a palm device or a blackberry device. 11. (canceled) 12. (canceled) 13. An improvement the computer-based programs that download blood glucose meters, insulin pumps and other devices (downloading programs), for analyzing and graphically-displaying a plurality of differently-categorized data separately, all together, or together in combinations smaller than the total, wherein said combinations may employ differently-colored and/or differently-shaped markers for the BG's of different categories, and wherein mouse-based computer interface methods may be used to select BG's into time intervals. 14. The method of claim 13 wherein the category variables in the downloading program are character-strings. 15. The method of claim 14 wherein there are a plurality of chart data series, each of which is provided with a selector in which the selections are made for categories of a BG for inclusion in the said chart series. 16. The method of claim 15 wherein one of the category selections is “none” or some other phrasing indicating that no categories were assigned to those BG's by the patient. 17. The method of claim 15 wherein the selector for each chart data series is comprised of a drop-down list, showing the c... | - | G05D007/00 A61B005/145 A61K009/22 A61M005/142
US20080249396 | A1 | 2008-10-09 | Method And Apparatus For Determining Indications Helping The Diagnosis Of Orthopedical Diseases | 2007-10-17 | Apparatus for determining indications helping the diagnosis of orthopedic diseases comprising a section detecting and acquiring signals from a body under examination or from a part thereof wherein the detecting section is a unit detecting images by nuclear magnetic resonance and it is integrated in a section for processing acquired images as regards image data and/or resonance signals, which processing section defines, from image data and/or resonance signals, values of one or more different numerical parameters indicating the presence or absence of an orthopedic disease and/or a measure of the evolution condition of said orthopedic disease.;The invention comprises also a method for determining indications helping the diagnosis of orthopedic diseases. | 1. Apparatus for determining indications helping the diagnosis of orthopedic diseases comprising: a section detecting and acquiring signals from a body under examination or from a part thereof wherein the detecting section includes a detection unit for detecting images by nuclear magnetic resonance and said unit is integrated in a section for processing acquired images as regards image data and/or resonance signals, which processing section defines, from image data and/or resonance signals, values of one or more different numerical parameters indicating the presence or absence of an orthopedic disease and/or a measure of the evolution condition of said orthopedic disease. 2. Apparatus according to claim 1, wherein the detecting section acquires images by nuclear magnetic resonance, the detecting section including MRI scanning means that are shaped and have such a dimension to acquire nuclear magnetic resonance images only of some limited parts of a body under examination and/or only of some limited anatomical regions or just only of some limited part of said anatomical regions. 3. Apparatus according to claim 2, wherein said MRI scanning means are shaped and have such a dimension that allow the acquisition of MRI images from the anatomical region of orthopedical interest and particularly from the region of the backbone, particularly of the backbone in the lumbosacral and/or cervical portion. 4. Apparatus according to claim 1, wherein the processing section comprises a sub-section with means for determining geometric parameters of main muscle skeletal structures of orthopedic interest and particularly of the backbone, particularly of vertebrae, disc structures and/or spinal canal. 5. Apparatus according to claim 4, wherein said geometric parameters are detected at different tilting degrees of the backbone, particularly with the backbone in a substantially horizontal position, i.e. at resting condition, and/or in a substantially vertical position, i.e. under load and/or in intermediate positions between the horizontal and vertical position of the backbone. 6. Apparatus according to claim 1, wherein the processing section comprises a sub-section with means for determining geometric parameters concerning the reciprocal space position of main structures of the backbone, particularly of vertebrae, disc structures and/or spinal canal in a substantially horizontal position, i.e. at resting condition, and/or in a substantially vertical position, i.e. under load and/or in intermediate positions between the horizontal and vertical position of the backbone. 7. Apparatus according to claim 1, wherein the processing section comprises a sub-section with means for comparing said numerical parameters with a reference value for discriminating the presence/absence of a pathologic condition and/or for the comparison with a reference scale for determining a parameter indicating the evolution degree of the pathologic condition, which reference value for discrimi... | ESAOTE S.p.A. | A61B005/05
US20080281232 | A1 | 2008-11-13 | Device for Measuring the Diameter of an Aortic Annulus | 2006-03-09 | The invention relates to a device for measuring the diameter of an aortic annulus. The distal end of the device is equipped with an expandable cylindrical measuring head (1) which is designed to be introduced through the aortotomy into the outflow tract of the left ventricle of a heart, while the proximal end of the device is equipped with a means (5) for controlling and measuring the constant, adjustable measurement of the expansion of the head, said measuring head and control and measuring means being connected by a rod (4). | 1. Device for measuring the diameter of an aortic annulus comprising, at the distal end thereof, an expansible cylindrical measuring head (1, 1′) which is adapted in order to be introduced via aortotomy into the outflow chamber of a left ventricle, and, at the proximal end thereof, a means (5, 5′) for controlling and measuring the expansion of the head with a constant and adjustable measuring force, and a connecting rod (4, 4′) between the measuring head and the control and measurement means. 2. Device according to claim 1, characterised in that the end (2, 2′) of the measuring head (5, 5′) is non-invasive. 3. Device according to claim 1, characterised in that the measuring head (1, 1′) has a length greater than 10 cm and is adapted for measuring an aortic annulus. 4. Device according to claim 1, characterised in that the length of the connecting rod (4, 4′) between the measuring head (1, 1′) and the control and measurement mechanism (5, 5′) is greater than 8 cm and is adapted to allow measurement from outside the thorax of a patient. 5. Device according to claim 1, characterised in that it is adapted in order to measure a diameter of at least between 18 mm and 32 mm, with a measuring force of at least between 3 N and 10 N. 6. Device according to claim 1, characterised in that the means (5) for controlling and measuring the expansion of the head is a micrometric screw with a torque limiter. 7. Device according to claim 1, characterised in that the measuring head (1) comprises three blades (3) which are elongate in a manner parallel with the longitudinal axis of the device and which are radially movable relative to the longitudinal axis of the device. 8. Device according to claim 1, characterised in that the measuring head (1′) comprises a band (3′) which is radially expansible relative to the longitudinal axis of the device. 9. Device according to claim 1, characterised in that it is constituted by materials which are suitable for use as a surgical instrument. 10. Method for measuring the diameter of a natural aortic annulus of a stopped heart, with extra-corporal circulation and being accessible via an aortotomy, according to which, via the aortotomy, a measuring device according claim 1 whose measuring head is retracted is introduced into the outflow chamber of the left ventricle, and, using the control and measurement means, the measuring head is expanded until a predefined limit measuring force is reached, then the diameter indicated by the control and measurement means is recorded and, after the measuring head has been retracted, the measuring device is withdrawn. 11. Device according to claim 2, characterised in that the measuring head (1, 1′) has a length greater than 10 cm and is adapted for measuring an aortic annulus. 12. Device according to claim 2, characterised in that the length of the connecting rod (4, 4′) between the measuring head (1, 1′) and the control and measurement mechanism (5, 5′) is greater than 8 cm and is ad... | ASSISTANCE PUBLIQUE - HOSPAUX DE PARIS | A61B005/00
US07729741 | B2 | 2010-06-01 | Method of multi-region contrast-enhanced magnetic resonance arteriography | 2007-12-25 | The present invention is a technique of, and system for, imaging vascular anatomy over distance considerably greater than the maximum practical field of view of a magnetic resonance imaging system while using substantially one contrast agent injection. The technique and system of the present invention acquires image data of a plurality of image volumes which are representative of different portions of the patient's body. The image data of each image volume includes image data which is representative of the center of k-space. The acquisition of image data which is representative of the center of k-space is correlated with a concentration of contrast agent in the artery(ies) residing in the image volume being substantially greater than the concentration of contrast agent in veins and background tissue adjacent to the artery(ies). This provides preferential enhancement of arteries relative to adjacent veins and background tissue for each acquisition, wherein each acquisition is representative of a different portion of the arterial system (e.g., abdominal aorta, femoral, popliteal, and tibial arteries). | 1. A method of multiregion contrast-enhanced magnetic resonance imaging, the method comprising: positioning a subject with respect to a magnetic resonance imaging system so that magnetic resonance imaging data of a first region of interest may be obtained, wherein the first region of interest includes a first portion of the subject's arterial system;introducing contrast material into the subject's vasculature via one injection, thereby causing an arterial phase of the injected contrast material, which phase occurs in the first region of interest and later in a second region of interest;obtaining magnetic resonance imaging data from the first region of interest, wherein such obtaining is so timed relative to the introduction of contrast material that magnetic resonance imaging data representative of the center of k-space is obtained from the first region of interest during the arterial phase of the injected contrast material in the first region of interest;repositioning the subject with respect to the magnetic resonance imaging system so that magnetic resonance imaging data of the second region of interest may be obtained, wherein the second region of interest includes a second portion of the subject's arterial system that is arterially downstream of the first portion of the subject's arterial system; andobtaining magnetic resonance imaging data of the second region of interest, wherein such obtaining is so timed relative to (a) the introduction of contrast material or (b) the obtaining of magnetic resonance imaging data of the first region of interest, that magnetic resonance imaging data representative of the center of k-space is obtained from the second region of interest during the arterial phase of the injected contrast material in the second region of interest. 2. The method of claim 1 further comprising forming an image of the first portion of the subject's arterial system using the magnetic resonance imaging data of the first region of interest. 3. The method of claim 2 further comprising forming an image of the second portion of the subject's arterial system using the magnetic resonance imaging data of the second region of interest. 4. The method of claim 3 further comprising combining the images of the first and second portions of the subject's arterial system to form a mosaic image. 5. The method of claim 1 wherein introducing comprises injecting a bolus of the contrast material into the subject's vasculature. 6. The method of claim 1 wherein obtaining magnetic resonance imaging data from the first region of interest comprises obtaining imaging data representative of the center of k-space prior to obtaining imaging data representative of the periphery of k-space. 7. The method of claim 1 wherein obtaining magnetic resonance imaging data from the first region of interest comprises obtaining imaging data representative of the center of k-space after obtaining imaging data representative of the periphery of k-space. 8. The method of... | Martin R. Prince | A61B005/055
US07785257 | B2 | 2010-08-31 | System and process for non-invasive collection and analysis of physiological signals | 2008-07-07 | A system and method for detecting, monitoring and analyzing physiological characteristics. Signals from a subject are acquired from a suite of sensors, such as temperature, carbon dioxide, humidity, light, movement, electromagnetic and vibration sensors, in a passive, non-invasive manner. The signals are processed and physiological characteristics are isolated for analysis. The system and method are to analyze sleep patterns, as well as to prevent bed sores or detect conditions such as illness, restless leg syndrome, periodic leg movement, sleep walking, or sleep apnea. However, numerous other applications of the invention are also disclosed. | 1. A non-invasive system for assessing physiological characteristics of a subject, said system comprising: a network of sensors configured to provide electronic output signals indicative of at least one sensed condition, said network including at least one passive sensor selected from the group consisting of vibration, temperature, carbon dioxide, humidity, electromagnetic and light sensors, and wherein said network of sensors further comprises a movable support including at least one sensor of the network of sensors;data acquisition circuitry for collecting the electronic signals output from said network of sensors;a processor configured to process the electronic signals, said processor including an amplifier configured to amplify at least one of the electronic signals, and first and second filters configured to receive the amplified electronic signals to produce two processed electronic signals, wherein said first filter includes a band pass filter and said second filter includes a low pass filter and a high pass filter, and wherein said first filter outputs a first signal corresponding to a respiratory rate on a breath-by-breath basis of the subject and said second filter outputs a second signal corresponding to a heart rate on a beat-by-beat basis of the subject; andan output device for outputting the processed signals into human-readable data indicative of at least one condition of a subject,wherein the human-readable data comprises a subject movement percentage. 2. The system of claim 1 wherein said system is adapted to assess the subject's quality of sleep and wherein: said data acquisition circuitry is configured to collect the electronic signals output from said network of sensors indicative of heart rate and/or respiratory rate;said first filter receives signals indicative of chest movement and said second filter receives signals indicative of heart movement; andsaid processor includes an algorithm to generate a heart rate and a respiratory rate based on processed electronic signals output at or below about 50 Hz. 3. The system of claim 2, further comprising an algorithm for determining sleep staging and aspects of sleep quality based on the heart rate, the respiratory rate, a variability of the heart rate, a variability of the respiratory rate and a subject movement percentage. 4. The system of claim 1, further comprising a microcontroller for converting the collected electronic signals into a bit stream to be logged into a data file for immediate analysis or for later review. 5. The system of claim 4, wherein said network of sensors includes at least one detector for sensing motion, the bit stream includes electronic signals from said at least detector and said processor processes the bit stream to provide a positional map of the subject. 6. The system of claim 1, wherein said amplifier comprises an instrumentation amplifier that provides buffering from said network of sensors and an infinite common mode rejection ratio. 7. The ... | Home Guardian LLC | A61B005/00
US07467004 | B2 | 2008-12-16 | System, method and apparatus for surgical patient table | 2005-05-12 | A patient table includes a cradle for slidably receiving a board carrying a patient, a cradle supporting member for slidably receiving the cradle, a locker for preventing the board from sliding off the cradle, and a stopper for preventing the cradle from sliding off the cradle supporting member. A method of transferring a patient to an imaging system from a second table uses the patient table to move the board from the second table to the patient table to the imaging system. An RF coil system usable with the table includes a coil, a coil cable, an intermediate box for receiving the coil cable, and a system cable extending from the intermediate box and connected to the imaging system. | 1. A patient table comprising: a cradle for linearly receiving and transferring a board carrying a patient;a cradle supporting member for slidably receiving the cradle in a first horizontal direction;a locker within the cradle for positioning the board and patient on the table, and for preventing the board from sliding off the cradle in the first horizontal direction;a stopper within the cradle for preventing the cradle from sliding off the cradle supporting member. 2. The patient table of claim 1 wherein the locker includes a locker manual release mechanism. 3. The patient table of claim 2 wherein the locker includes a spring tab biased to engage with a notch within the board, wherein pulling on the locker manual release mechanism releases the spring tab from the board. 4. The patient table of claim 1 wherein docking the patient table with a second table disengages the locker from the board. 5. The patient table of claim 4 wherein the locker includes a spring tab biased to engage with a notch within the board, and further including a pusher knob protruding from the cradle and operatively linked to the spring tab, wherein, docking the patient table with a second table pushes the pusher knob inwardly within the cradle and releases the spring tab from the board. 6. The patient table of claim 1 wherein the locker includes a first locker and a second locker, the first locker including a locker manual release mechanism, and wherein docking the patient table with a second table disengages the second locker from the board. 7. The patient table of claim 1 wherein the stopper is disengaged from the cradle supporting member when the patient table abuts a system for receiving the cradle. 8. The patient table of claim 7 wherein the stopper includes a slidable block biased to protrude outwardly from the cradle into a notch within the cradle supporting member, the stopper further including a block moving rod protruding outwardly from the cradle and operatively linked to the slidable block, wherein, when the block moving rod is pushed inwardly within the cradle, slidable block is pulled out of the notch. 9. The patient table of claim 1 wherein the stopper is disengaged from the cradle supporting member when the board is fully supported upon the cradle. 10. The patient table of claim 9 wherein the stopper includes a slidable block biased to protrude outwardly from the cradle into a notch within the cradle supporting member, the stopper further including a flipper protruding upwardly from the cradle and operatively linked to the slidable block, wherein, when the board is pushed over the flipper, the slidable block is pulled out of the notch. 11. The patient table of claim 1 wherein the stopper includes a first stopper disengaging from the cradle supporting member when the patient table abuts a system for receiving the cradle, and a second stopper disengaging from the cradle supporting member when the board is fully supported upon the cradle. 12. The pa... | General Electric Company | A61B005/05 A47B071/00 A47B013/00
US07818155 | B2 | 2010-10-19 | Esophageal waveform analysis for detection and quantification of reflux episodes | 2008-02-11 | A system and method for automatically analyzing impedance and pH data from an esophageal probe includes a data collection system that collects and stores the output from the sensors for a certain period of time to locate reflux episodes in the waveforms. The data analysis system uses wavelet analysis to assist in locating bolus entry and exit points in the waveforms and to smooth waveforms for additional analysis. It also distinguishes between candidate acid reflux episodes and candidate non-acid reflux episodes for determining actual acid reflux episodes and non-acid reflux episodes, and it distinguishes between, and provides different signal processing for, signals from patients that have healthy esophageal tissue and those that have diseased esophageal tissue. | 1. A method of identifying and classifying reflux episodes in a plurality of impedance versus time waveforms obtained simultaneously at a plurality of different locations spaced apart from each other along a length of an esophagus between the pharynx and the stomach along with a pH versus time curve obtained in the esophagus, wherein the impedance versus time waveforms include proximal waveforms from the locations closest to the pharynx, distal waveforms from the locations closest to the stomach, and mid-channel waveforms between the proximal waveforms and the distal waveforms, said method comprising: applying a smoothing algorithm to the impedance versus time waveforms;locating candidate transition points and candidate end points in the smoothed waveforms to indicate candidate reflux episodes;determining whether a pH drop occurs in the pH versus time curve contemporaneous with a candidate transition point to thereby indicate whether the candidate reflux episode associated with such candidate transition point is acidic or non-acidic;determining a gastric contents impedance threshold for acid reflux episodes;determining a gastric contents impedance threshold for non-acid reflux episodes;determining if there is impedance of acid gastric contents associated with the candidate acid reflux episodes by comparing values associated with impedance at the candidate transition points for candidate acid reflux episodes to the gastric contents impedance threshold for acid reflux episodes;determining if there is impedance of non-acid gastric contents associated with the candidate non-acid reflux episodes by comparing values associated with impedance at the candidate transition points for candidate non-acid reflux episodes to the gastric contents impedance threshold for non-acid reflux episodes;for candidate acid reflux episodes wherein there is impedance of acid gastric contents, determining if there is a propagation direction of corresponding candidate transition points in any of the other impedance versus time waveforms starting from a distal waveform;for candidate non-acid reflux episodes wherein there is impedance of non-acid gastric contents, determining if there is a propagation direction of corresponding candidate transition points in any of the other impedance versus time waveforms starting from a distal waveform;classifying candidate acid reflux episodes as actual acid reflux episodes based at least in part on determination that there is impedance of acid gastric contents and on determination that there is propagation direction; andclassifying candidate non-acid reflux episodes as actual non-acid reflux episodes based at least in part on determination that there is impedance of non-acid gastric contents and on determination that there is propagation direction. 2. The method of claim 1 including, for candidate acid reflux episodes wherein there is impedance of acid gastric contents, determining if there is a propagation direction of corresponding cand... | Sandhill Scientific, Inc. | G06G007/58
US07445605 | B2 | 2008-11-04 | Detection of apex motion for monitoring cardiac dysfunction | 2006-03-30 | Systems, devices and methods for detecting and monitoring cardiac dysfunction. The devices include motion sensors for detecting signals representative of the total movement of the heart, and of the apex of the heart in particular. | 1. A method for detecting an ischemic condition in a patient, the method comprising: sensing the motion of the ventricular apex of the heart in two or more directional axes with at least one motion sensor implanted at the heart;processing signals generated by the at least one motion sensor, the processing comprising relating the signals from two or more directional axes of movement to at least one predetermined baseline value of the movement of the apex of the heart; andwherein the location of an ischemic condition of the heart is determined by the deflection of said path that the ventricular apex traverses over time, wherein the deflection is in a direction away from affected myocardium of the heart. 2. The method according to claim 1, wherein the sensing is performed with a single multiple dimensional motion sensor that provides two or three-dimensional measurements. 3. The method according to claim 1, further comprising producing an output, wherein the output is a cardiac pacing signal. 4. The method according to claim 1, further comprising producing an output, wherein the output is the delivery of a therapeutic agent to the heart. 5. The method according to claim 1, further comprising producing an output, wherein the output is an alarm signal produced when a processed signal is determined to be outside a predetermined threshold range. 6. The method according to claim 1, further comprising performing an electrocardiogram of the heart with another sensor. 7. The method according to claim 1, wherein the at least one sensor is implanted endocardially. 8. The method according to claim 1, wherein the at least one sensor is implanted epicardially. 9. The method of claim 1, wherein said at least one motion sensor is a rotational motion sensor. 10. The method of claim 1, wherein the at least one motion sensor is adapted to be sensitive to acceleration in axes perpendicular to the sensor's lead. 11. A method for detecting an ischemic condition in a patient, the method comprising: sensing the motion of the ventricular apex of the heart in two or more directional axes with at least one motion sensor implanted at the heart;processing signals generated by the at least one motion sensor at apical motion frequencies to 4 Hz, the processing comprising relating the signals from two or more directional axes of movement to at least one predetermined baseline value of the movement of the apex of the heart,wherein the location of an ischemic event is determined by the deflection of said path that the ventricular apex traverses over time, wherein the deflection is in a direction away from affected myocardium of the heart. 12. The method of claim 11, wherein the at least one motion sensor is a multiple-dimensional sensor that provides two or three-dimensional measurements. 13. The method of claim 11, wherein said at least one motion sensor comprises a multiple-dimensional sensor that provides two or three-dimensional measurement, and a rotational motion ... | The Board of Trustees of the Leland Stanford Junior University | A61B005/103 A61B005/02 A61N001/00
US07431703 | B2 | 2008-10-07 | Apparatus and method for measuring and monitoring range of motion of the lumbar spine | 2004-03-12 | An apparatus and method for detecting, monitoring, measuring and recording the motion of a patient's spine is disclosed. The apparatus includes a vertical member attached to a belt that fits around the patient's waist. A first sensor for detecting and monitoring flexion and extension of the patient's spine in the midsagittal plane is mounted to the lower end of the vertical member. A second sensor for detecting and monitoring lateral bending and movement of the patient's spine in the frontal plane is mounted to the upper end of the vertical member. | 1. An apparatus for monitoring the movement of a patient's spine comprising: an elongated member adapted to be disposed longitudinally adjacent to the patient's spine and further adapted to be flexible in the midsagittal plane and substantially inflexible in the frontal plane;a first sensor mounted to the elongated member and disposed to monitor flexion and extension motion of the patient's spine in the midsagittal plane, wherein the first sensor includes at least one strain gage; anda second sensor mounted to the elongated member and disposed to monitor lateral motion of the patient's spine in the frontal plane, wherein the second sensor is an optical sensor. 2. An apparatus for monitoring the movement of a patient's spine comprising: an elongated member adapted to be disposed longitudinally adjacent to the patient's spine and further adapted to be flexible in the midsagittal plane and substantially inflexible in the frontal plane;a first sensor mounted to the elongated member and disposed to monitor flexion and extension motion of the patient's some in the midsagittal plane; anda second sensor mounted to the elongated member and disposed to monitor lateral motion of the patient's some in the frontal plane, wherein the second sensor is an optical sensor, and further wherein the second sensor is an optical mouse sensor. 3. An apparatus for monitoring the movement of patient's spine comprising: an elongated member adapted to be disposed longitudinally adjacent to the patient's spine and further adapted to be flexible in the midsagittal plane and substantially inflexible in the frontal plane, wherein the elongated member has a first end and a second end opposite the first end;a first sensor mounted to the elongated member and disposed to monitor flexion and extension motion of the patient's spine in the midsagittal plane, wherein the first sensor is mounted to the elongated member near the first end; anda second sensor mounted to the elongated member and disposed to monitor lateral motion of the patient's spine in the frontal plane, wherein the second sensor is mounted to the elongated member near the second end. 4. An apparatus for monitoring the movement of a patient's spine comprising: an elongated member adapted to be disposed longitudinally adjacent to the patient's spine and further adapted to be flexible in the midsagittal plane and substantially inflexible in the frontal plane;a first sensor mounted to the elongated member and disposed to monitor flexion and extension motion of the patient's spine in the midsagittal plane, wherein the first sensor is adapted to be disposed along the patient's spine at approximately the location of the 1st sacral vertebrae; anda second sensor mounted to the elongated member and disposed to monitor lateral motion of the patient's spine in the frontal plane, wherein the second sensor is adapted to be disposed along the patient's spine at approximately the location of the 12th thoracic vertebrae. 5. An appa... | - | A61B005/117 A61B005/103
US07450983 | B2 | 2008-11-11 | Automated brain MRI and CT prescriptions in Talairach space | 2004-03-18 | Brain Magnetic Resonance Imaging (MRI), Computerized Tomotography (CT), or other diagnostic modalities may employ a three-step procedure during initial (“scout”) cranial pre-scans that corrects for patient positioning (i.e., roll, yaw and pitch) to reduce inter- and intra-patient variation, thereby enhancing the diagnostic and comparative value of subsequent detail scans even across different diagnostic platforms. In MRI, for instance, locating the saggital sinus (SS) and optimizing a line to bisect the brain through this SS may be automated to correct for roll and yaw. By then identifying the contour of the corpus callosum in a lateral saggital scout scan, the Talairach anterior commissure (AC)—posterior commissure (PC) reference line may be found for correcting pitch. Prescription of detailed scans are improved, especially when the three-step correction is repeated periodically identifying the need to repeat a detailed scan or to adjust the coordinates of a subsequent scan. | 1. A method of identifying a head position of a patient undergoing diagnostic imaging, comprising: obtaining a diagnostic image of a patient's head with a diagnostic imaging machine;performing an automated image processing operation to determine coordinates of a Talairach anterior commissure (AC)—posterior commissure (PC) reference line within the diagnostic image by identifying a line passing through a hard palate on the diagnostic image and utilizing the identified line to approximate the Talairach AC-PC reference line as about 12 degrees more extended than the hard palate in a lateral midline image; anddefining a coordinate system of the diagnostic image with reference to the Talairach AC-PC reference line. 2. The method of claim 1, further comprising: prescribing a subsequent scan based on the identified Talairach AC-PC reference line and coordinate system; andrepeating the automated image processing operation to determine current coordinates of the Talairach AC-PC reference for accommodating changes in patient head position since the previous position determination. 3. The method of claim 1, wherein obtaining a diagnostic image comprises obtaining a roll and yaw corrected thin section midsagittal magnetic resonance imaging (MRI) image of the patient's head. 4. The method of claim 3, wherein obtaining a midsagittal MRI image of a patient's head further comprises: obtaining at least two scout views;identifying midline features to permit correction of roll and yaw; andobtaining a midsagittal MRJ image based on identified midline features. 5. The method of claim 4, wherein obtaining a midsagittal MRI image of a patient's head when identifying midline features to permit correction of roll and yaw further comprises: performing at least one rapid scan operatively configured to accentuate venous blood flow in the superior sagittal sinus (SSS) in a plane selected from a group consisting of coronal scan, axial scan and oblique scan of the patient's head;identifying the SSS in cross-section in the at least one rapid scan;identifying a line that bisects the brain with the line passing through the SSS cross-section; anddefining an attitude correction selected from a group consisting of roll correction and yaw correction corresponding to the selected plane for subsequent scans based on the identified line that bisects the brain. 6. The method of claim 1, wherein obtaining a diagnostic image comprises obtaining a lateral computerized tomography (CT) scout image. 7. The method of claim 6, wherein obtaining a lateral CT scout image of a patient's head further comprises physically adjusting patient's head position relative to a scanner that obtains the lateral CT scout image for minimizing roll and yaw visually. 8. A program product embodied in a tangible, recordable media, comprising: (a) a program configured to receive a thin section diagnostic image of a patient's brain from a CT diagnostic imaging machine and to determine coordinates of a Talairach... | University of Cincinnati | A61B005/05
US08139829 | B2 | 2012-03-20 | Diagnostic imaging support system and diagnostic imaging support program | 2006-02-07 | A diagnostic imaging support system includes input means (15) for setting a characteristic quantity for performing a predetermined threshold processing for a bone region in an image of a subject; control means (10) for reading out the image from storage means (11) which stores the image of the subject, extracting a diagnostic region including the bone region from the read image on the basis of the characteristic quantity set by the input means, and calculating, for the extracted diagnostic region, structure analysis information of the bone by use of component identification information representing a bone portion component and a component other than the bone portion; and display means (14) for displaying the calculated structure analysis information of the bone while relating it to the image of the subject. | 1. A diagnostic imaging support system comprising: storage means for storing an image of a subject; andcontrol means for reading out the image from storage means, extracting a diagnostic region including a bone region from the read image on a basis of a characteristic quantity previously set, extracting a bone portion component and a component other than the bone portion from the diagnostic region, and calculating structure analysis information of the bone by utilizing at least one of a run length or density of a region that corresponds to the component other than the bone portion. 2. A diagnostic imaging support system according to claim 1, wherein the structure analysis information is the shape symmetry of a region representing one of a density representing the bone portion, or a portion other than the bone portion, of each of first and second regions, the first and second regions being at least two regions divided from the extracted diagnostic region. 3. A diagnostic imaging support system according to claim 1, wherein the display means displays the bone region of the subject such that the bone component and the component other than the bone can be identified by the characteristic quantity. 4. A diagnostic imaging support system according to claim 1, wherein the display means performs display, while changing a method of projecting the bone portion of the subject depending on whether the degree of progress of osteoporosis or the degree of healthiness of osteoporosis is displayed. 5. A diagnostic imaging support system according to claim 1, wherein the extraction means determines a joint portion of the bone portion on the basis of density; andthe display means displays the joint portion of the bone in an identifiable manner on the basis of results of the determination. 6. A diagnostic imaging support system according to claim 1, further comprising means for generating a ray-sum image of a spine region of the subject as viewed from the front, the image showing normalized densities of the spine region, whereinthe display means simultaneously displays the generated ray-sum image and the characteristic quantity of each tomogram of the subject. 7. A diagnostic imaging support system according to claim 1, further comprising means for calculating, as a total average, an average of characteristic quantities corresponding to a plurality of bone regions of the subject and for calculating, as an excluded average, an average of the characteristic quantities excluding characteristic quantities corresponding to bone regions whose densities are lower than the average density of the plurality of bone regions, whereinthe display means simultaneously displays the calculated total average and excluded average. 8. A diagnostic imaging support system according to claim 1, further comprising means for calculating, as a total average, an average of characteristic quantities corresponding to a plurality of bone regions of a healthy person, wherein the display me... | Hitachi Medical Corporation | G06K009/00
US07951357 | B2 | 2011-05-31 | Implantable power sources and sensors | 2005-07-13 | A protein is provided, including a glucose binding site, cyan fluorescent protein (CFP), and yellow fluorescent protein (YFP). The protein is configured such that binding of glucose to the glucose binding site causes a reduction in a distance between the CFP and the YFP. Substance monitoring apparatus (210) is also provided, including a semi-permeable barrier (212), adapted to be implanted in a body of a subject and to allow passage therethrough of a substance, while inhibiting passage therethrough of immune cells; and microorganisms (214), disposed within the semi-permeable barrier (212) so as to produce a measurable response to a level of the substance. A sensor (220) is adapted to measure the measurable response and not to measure a response of any mammalian cells that may be disposed within the semi-permeable barrier (212). | 1. Apparatus for detecting a concentration of a substance in a subject, the apparatus comprising: a fluorescence resonance energy transfer (FRET) measurement device; anda housing adapted to be implanted in the subject, the housing comprisingcells genetically engineered to produce, in situ, a FRET protein having a FRET complex comprising a fluorescent protein donor, a fluorescent protein acceptor, and a binding site for the substance. 2. The apparatus according to claim 1, wherein the protein that the cells are genetically engineered to produce comprises: a glucose binding site;cyan fluorescent protein (CFP); andyellow fluorescent protein (YFP),wherein the protein is configured such that binding of glucose to the glucose binding site changes a distance between the CFP and the YFP. 3. The apparatus of claim 2, wherein the protein is encoded by an isolated nucleic acid fragment having the nucleotide sequence in SEQ ID NO: 1. 4. The apparatus according to claim 1, comprising an inducer that is adminsterable to the subject and that regulates protein expression of the cells. 5. The apparatus according to claim 1, wherein the cells comprise beta or liver cells. 6. The apparatus according to claim 1, wherein the substance includes glucose, and wherein the binding site comprises a binding site for glucose. 7. The apparatus according to claim 1, wherein the protein comprises a protein selected from the group consisting of blue fluorescent protein (BFP), green fluorescent protein (GFP), and GFP/Rhodamine. 8. The apparatus according to claim 1, wherein the housing comprises a chamber which holds the cells. 9. The apparatus according to claim 1, wherein the cells are held directly within the housing, with no intervening chamber to hold the cells. 10. The apparatus according to claim 1, wherein the cells are disposed within a matrix. 11. The apparatus according to claim 1, wherein the cells are disposed in a suspension. 12. The apparatus according to claim 1, wherein the housing comprises a membrane that prevents passage therethrough of cells of the subject's body. 13. The apparatus according to claim 12, wherein the membrane is configured to prevent passage therethrough of the FRET protein. 14. The apparatus according to claim 1, wherein the cells are configured such that the protein produced by the genetically-engineered cells becomes positioned on the membranes of the cells. 15. The apparatus according to claim 14, wherein the protein comprises a targeting peptide that is configured to direct the protein to the membranes of the cells. 16. The apparatus according to claim 15, wherein the sequence of the targeting peptide is SEQ ID NO: 2. 17. The apparatus according to claim 1, comprising a drug capable of killing the cells, the drug being configured to be administered to the subject subsequently to the implantation of the housing in the subject. 18. The apparatus according to claim 1, wherein the fluorescence resonance energy transfer (FRET)... | Glusense Ltd. | A61B005/00 C07H021/04 C07K014/00 C12N015/85
US08204576 | B2 | 2012-06-19 | Medical guiding system | 2008-06-04 | A medical guiding system of the invention includes: a detection portion for detecting a position and an orientation of a medical instrument and positions of feature points of a subject; a storage portion for storing reference image data having positional information on an organ of a human body; and a guide image creation portion for creating a guide image indicating the position of the medical instrument with respect to the subject from the reference image data. The guide image creation portion calculates the position of the medical instrument with respect to the subject, based on the position and the orientation of the medical instrument and the positions of feature points of the subject which are detected by the detection portion, and creates a guide image by correcting the calculated position of the medical instrument, using the positions of the feature points. | 1. A medical guiding system comprising: a detection portion configured to detect at least one of a position and an orientation of a medical instrument and at least one of positions and orientations of a plurality of feature points of a subject in a real space;a storage portion configured to store reference image data having anatomical positional information on at least one organ of a human body; anda guide image creation portion configured: to define a mapping for converting at least one of the position and the orientation of the medical instrument in the real space into at least one of a position and an orientation in the reference image data based on information on: at least one of positions and orientations of the plurality of feature points of the subject in the real space, andat least one of positions and orientations of a plurality of reference image feature points which correspond to the plurality of feature points in the reference image data, andto create a guide image indicating a position of the medical instrument or an observation position by the medical instrument based on the mapping and the reference image data,to correct the position of the medical instrument or the observation position by the medical instrument which the guide image indicates, based on information on: at least one of positions and orientations of a plurality of in-vivo feature points in a body of the subject in the real space, andat least one of positions and orientations of one or more correction feature points which anatomically correspond to the in-vivo feature points in the reference image data, andto change anatomical positions of the in-vivo feature points to be used in the correction depending on an anatomical region for which the medical instrument is used. 2. The medical guiding system according to claim 1, wherein the in-vivo feature points are two, andthe correction corrects the calculated position of the medical instrument or the calculated observation position by the medical instrument by correcting a scale size of the mapping based on comparison of a distance between the two in-vivo feature points and a distance between two correction feature points. 3. The medical guiding system according to claim 2, wherein the correction moves the position of the medical instrument or the observation position by the medical instrument which the guide image indicates, such that one of the in-vivo feature points coincides with one of the reference image feature points which corresponds to the one of the in-vivo feature points, and correct the calculated position of the medical instrument or the observation position by the medical instrument. 4. The medical guiding system according to claim 3, wherein presence or absence of execution of the correction of the scale size of the mapping is switched depending on the anatomical region for which the medical instrument is used. 5. The medical guiding system according to claim 1, wherein the correction moves the position... | Olympus Medical Systems Corp. | A61B005/05
US08801621 | B2 | 2014-08-12 | Method, system and program product for analyzing pulse wave data | 2008-03-27 | A pulse wave data analyzing method includes successively detecting bottom and top peak values of pulse wave data along a time axis, calculating successive bottom-to-top amplitude values along the time axis, and comparing first and second peak-to-peak amplitude values occurring in succession along the time axis. If the ratio of the second peak-to-peak amplitude value to the first peak-to-peak amplitude value is smaller than a preset threshold, the bottom and top peak values related to the second peak-to-peak amplitude value are classified as temporarily erased data. The second peak-to-peak amplitude value is compared with a third peak-to-peak amplitude value occurring immediately thereafter and, if the ratio between the second and third peak-to-peak amplitude values is larger than the threshold, the temporarily erased data is restored. If the ratio between the second and third peak-to-peak amplitude values is not larger than the threshold, the temporarily erased data is completely erased. | 1. A pulse wave data analyzing method for extracting biometric information from pulse wave data taken from a living body, said pulse wave data analyzing method conducted by using a processor, the method comprising the steps of: successively detecting bottom peak values and top peak values of the pulse wave data obtained by continuously measuring a pulse wave for a specific period of time along a time axis;pairing a detected bottom peak value with a corresponding detected adjacent top peak value, respectively, in a successive manner along the time axis to generate a plurality of successive pairs;calculating bottom-to-top amplitude difference values representing a difference between the bottom and top peak values of each successive pair along the time axis to define a first maximum difference value, a second maximum difference value, and a third maximum difference value, wherein the first, second, and third maximum differences values obtained from the corresponding successive pairs occur in succession along the time axis;comparing the first maximum difference value and the second maximum difference value;classifying the bottom and top peak values corresponding to the second maximum difference value as temporarily erasable data if a ratio of the first maximum difference value to the second maximum difference value is greater than a predetermined threshold;comparing the second maximum difference value with the third maximum difference value;restoring the temporarily erasable data as data to be used for pulse wave data analysis if a ratio between the second maximum difference value and the third maximum difference value is larger than the predetermined threshold; andcompletely erasing the temporarily erasable data as corresponding to noise if the ratio between the second maximum difference value and the third maximum difference value is not larger than the predetermined threshold. 2. The pulse wave data analyzing method according to claim 1 further comprising the step of calculating one of a top peak interval and a bottom peak interval equivalent to an RR interval observed on an electrocardiogram from the pulse wave data from which the noise has been removed. 3. The pulse wave data analyzing method according to claim 1 further comprising the step of calculating a bottom-to-top amplitude value from the pulse wave data from which the noise has been removed. 4. The pulse wave data analyzing method according to claim 1, further comprising, after classifying the bottom and top peak values related to the second peak-to-peak amplitude value as the temporarily erased data, the step of adopting a bottom-to-top amplitude value calculated from bottom and top peak values detected in succession to the temporarily erased bottom and top peak values along the time axis as a new second peak-to-peak amplitude value and comparing the second peak-to-peak amplitude value with the first peak-to-peak amplitude value at least once, and adopting this bottom-to-top amplitud... | Konica Minolta Sensing, Inc. | A61B005/02 A61B005/00 A61B005/021 A61B005/024 A61B005/0285
US09380966 | B2 | 2016-07-05 | Tissue retractor oximeter | 2008-05-24 | A retractor has an oximeter sensor at its tip, which allows measuring of oxygen saturation of a tissue being retracted by the retractor. The tip includes one or more openings for at least one source and detector. A specific implementation is a spinal nerve root retractor with an oximeter sensor. | 1. A device comprising: a shaft, comprising a first bend, wherein from a distal end of the shaft to the first bend, the shaft extends in a first direction, and from the first bend to a proximal end of the shaft, the shaft extends in a second direction, which is different from the first direction;a handle, coupled to a proximal end of the shaft, extending in a second direction;a tip, coupled to a distal end of the shaft, comprising a retractor portion and an oximeter sensor on a bottom surface of the tip, wherein the retractor portion comprises a blade having a first edge and a second edge, opposite to the first edge, a first thickness between the first and second edges, the first edge has a first length that is greater than the first thickness, the first edge is closer to the bottom surface than the second edge, and the blade extends from the bottom surface in a third direction away from an axis that passes through the shaft;a first optical fiber;a second optical fiber, wherein the first optical fiber passes through a channel in the shaft and a distal end of the first optical fiber is coupled to a first sensor opening of the tip, and the second optical fiber passes through the channel in the shaft and a distal end of the second optical fiber is coupled to a second sensor opening of the tip,the first sensor opening is positioned on the bottom surface of the tip a first distance from the first edge, the second sensor opening is positioned on the bottom surface of the tip a second distance from the first edge, and the first distance is the same as the second distance; anda system unit comprising a display, processor, signal emitter circuit, signal detector circuit, and a receptacle to couple to proximal ends of the first and second optical fibers, wherein the signal emitter circuit sends a signal through the first optical fiber and the signal detector circuit receives the signal from the second optical fiber. 2. The device of claim 1 wherein the system unit comprises a power management circuit, a battery recharger circuit, and a rechargeable battery comprising lithium ion battery chemistry, when the power management circuit detects a low battery condition, the power management circuit causes a warning to show on the display,when the power management circuit detects a specified amount of time of inactivity has occurred, darkening the display, wherein the specified amount of time of inactivity is programmable by a user, andwhen the system unit detects that an oxygen saturation level measurement falls below a threshold value, displaying a warning message on the screen, wherein warning lights are positioned on a handle of the retractor, and the system unit causes the warning lights of the retractor to light. 3. The device of claim 1 wherein the first sensor opening and second sensor opening are separated by a distance from at least about 1.5 millimeters to about five millimeters. 4. The device of claim 1 wherein the console comprises a modem for tran... | ViOptix, Inc. | A61B005/145 A61B005/1459 A61B017/02 A61B005/00
US09314192 | B2 | 2016-04-19 | Detection and compensation method for monitoring the place of activity on the body | 2006-12-05 | A measuring system comprises a sensor arranged to be attached to a subject for obtaining a measured value representing a physical or a physiological quantity of the subject. The measuring system further comprises a microprocessor for deriving a subject-related value from the measured value. The sensor is arranged to be attached at one of a plurality of positions on the subject. The measuring system further comprises a microprocessor for establishing the position of the sensor on the subject. The microprocessor for deriving the subject-related value is arranged for deriving the subject-related value also in dependence on the position of the sensor on the subject. | 1. A measuring system comprising: a sensor arranged to be attached at any one of a plurality of positions on a body of a subject for obtaining a measured value representing a physical or a physiological quantity of the subject; anda processor coupled to said sensor, said processor being arranged to: determine one of the plurality of positions on the subject to which the sensor is attached by analyzing the measured value for features that are position-dependent, andderive a subject-related value from the measured value, where the derivation of the subject-related value also depends on the one of the plurality of positions of the sensor on the subject. 2. The measuring system as claimed in claim 1, wherein the sensor is arranged to be attached at the plurality of positions that include: a wrist, a lower arm, an upper arm, a lower leg, an upper leg, a waist, a chest, a neck, and a head. 3. The measuring system as claimed in claim 1, wherein the subject-related value comprises an activity parameter of the subject. 4. The measuring system as claimed in claim 3, wherein the activity parameter comprises energy expenditure. 5. The measuring system as claimed in claim 3, wherein the activity parameter represents a degree of activity of a body part associated with the one of the plurality of positions on the subject where the sensor is attached. 6. The measuring system as claimed in claim 1, wherein the subject-related value comprises at least one of temperature, ECG, or acceleration. 7. The measuring system as claimed in claim 6, wherein the acceleration is a tri-axial acceleration. 8. The measuring system as claimed in claim 1, wherein the processor is further arranged to select a subset of a predefined set of further physical and/or physiological quantities of the subject in dependence on the one of the plurality of positions of the sensor on the subject, and the sensor is configured to generate a further measured value for each quantity in the selected subset. 9. The measuring system as claimed in claim 1, wherein the processor is further arranged to obtain, from the sensor, the measured value or a plurality of measured values measured during a time interval, and wherein the processor is further arranged to determine any one of the plurality of positions on the subject in dependence on the measured value or measured values measured during the time interval. 10. The measuring system as claimed in claim 9, wherein the processor is arranged to determine any one of the plurality of positions on the subject on the basis of a predefined set of rules related to the measured value or measured values measured during the time interval. 11. The measuring system as claimed in claim 9, wherein the processor is arranged to determine any one of the plurality of positions on the subject using a pattern recognition of a signal derived from the measured values measured during the time interval. 12. The measuring system as claimed in claim 9, wherein the proce... | Koninklijke Philips N.V. | A61B005/11 A61B005/06 A61B005/00
US20080269589 | A1 | 2008-10-30 | Apparatus for the Detection of Heart Activity | 2006-07-14 | The invention relates to heart measurement and heart monitoring, in particular the measurement of mechanical heart activity, and includes a method and apparatus to using doppler radar to transmit an electromagnetic signal of a certain frequency into, and detect a reflected signal from out of, the chest of the individual, to processing the detected signal to produce an output signal representing the rate of change of the doppler signal associated with the reflected signal and to identify from the output signal a group of at least one characteristic point of the output signal, and further to calculate at least one parameter representative of heart activity, this calculation based on the at least one identified characteristic point. The apparatus provides a system for monitoring which is particularly suitable for use in the home and which does not require repeated use of impedance cardiograms which are inappropriate for use by untrained personnel. | 1. A method to detect mechanical heart activity of an individual using doppler radar comprising: transmitting an electromagnetic signal of a certain frequency into, and detecting a reflected signal from out of, the chest of the individual,transmitting a signal representative of the detected signal to a processing system,processing, by processing system the detected signal to produce an output signal representing the rate of change of the doppler signal associated with the reflected signal, the rate of change with respect to time, wherein the method further comprises the steps of:identifying, by the processing system from the output signal a group of at least one characteristic point of the output signal,and further calculating, by the processing system at least one parameter representative of heart activity, the calculation based on the at least one identified characteristic point,wherein the signal representative of the detected signal is transmitted wirelessly to a computer with a connection to the world wide web and from said computer through the world wide web to the processing system. 2. (canceled) 3. A wearable apparatus to detect mechanical heart activity of an individual using doppler radar, comprising: a transducer to transmit electromagnetic signals of a certain frequency into the chest of the individual, and to detect reflected signals from out of the chest,and to transmit a signal representative of the detected signal to a processing system, which system is arranged to use the received signal to calculate an output signal representing the rate of change of the doppler signal associated with the reflected signal, the rate of change with respect to time,identify from the output signal a group of at least one characteristic point of the output signal,and calculate at least one parameter representative of heart activity, the calculation based on the at least one identified characteristic point,wherein the apparatus comprises means for transmitting the signal representative of the detected signal wirelessly to a computer with a connection to the world wide web, wherein said computer is adapted to further transmit the signal representative of the detected signal through the world wide web to the processing system. 4. A processing system, for receiving a signal transmitted from a wearable apparatus to detect mechanical heart activity of an individual using doppler radar, the system arranged to receive a signal representative of a reflected electromagnetic signal detected from out of the chest of an individual, and further arranged to: calculate an output signal representing the rate of change of the doppler signal associated with the reflected signal, the rate of change with respect to time,identify from the output signal a group of at least one characteristic point of the output signal,and calculate at least one parameter representative of heart activity, the calculation based on the at least one identified characteristic point,wherein t... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/05
US20080269568 | A1 | 2008-10-30 | METHODS FOR VALIDATING A RADIOTHERAPY PLAN BY A DOSIMETRY SERVICE | 2007-04-25 | Described are various methods for validating a patient's radiotherapy plan. One possible method comprises receiving a dose distribution at a dosimetry service, receiving dose data at the dosimetry service from a radiation detection device at a treatment site removed from the dosimetry service, the data having been gathered in accordance with the dose distribution, and evaluating the dose data in comparison to the dose distribution to prepare a validation report. | 1. A method for validating a radiotherapy plan, the method comprising: receiving a dose distribution at a dosimetry service;receiving dose data at the dosimetry service from a radiation detection device at a treatment site removed from the dosimetry service, the data having been gathered in accordance with the dose distribution; andevaluating the dose data in comparison to the dose distribution to prepare a validation report. 2. The method of claim 1 further comprising: sending a validation report from the dosimetry service to the person requesting the validation. 3. The method of claim 1 wherein the radiation detection device is an electronic portal imaging device or radiation-sensitive film. 4. The method of claim 3 wherein the electronic portal imaging device is at least one of an ion chamber array, a diode array, and an amorphous silicon array. 5. The method of claim 3 wherein the radiation-sensitive film is radiochromic film. 6. The method of claim 1 wherein the dose data is received electronically. 7. A method for validating a radiotherapy plan, the method comprising the steps of: receiving recorded radiation information at a dosimetry service,receiving a dose distribution at the dosimetry service;measuring the recorded radiation information; andcomparing the measured recorded radiation information to the dose distribution for preparation of a validation report. 8. The method of claim 7 further comprising: sending a film configured to record radiation information from the dosimetry service to a treatment site removed from the dosimetry service. 9. The method of claim 8 further comprising: loading the film into a phantom by the dosimetry service before sending the film to the treatment site. 10. The method of claim 8 further comprising: providing the treatment site with instructions for exposing the film; andsending a validation report from the dosimetry service to the treatment site or the person requesting the validation. 11. The method of claim 10 wherein the instructions direct the treatment site to place the film into a phantom for irradiation. 12. The method of claim 8 wherein the film is radiochromic film. 13. The method of claim 12 wherein the radiochromic film is made using lithium pentacosa-10,12-diynoate filamentary particle technology. 14. The method of claim 12 wherein the film comprises at least one control film, at least one calibration film, and at least one test film corresponding to each radiation field to be evaluated as part of the dose distribution. 15. The method of claim 14 wherein the films are packaged by the dosimetry service to protect the films from harmful environmental exposure. 16. The method of claim 7 wherein the dose distribution is from an intensity modulated radiation therapy treatment plan or a stereotacetic radiosurgery plan. 17. The method of claim 7 wherein measuring the recorded information comprises scanning the film to create a film scan image and converting the film scan image to a ... | ISP INVESTMENTS INC. | A61B005/00
US20080319327 | A1 | 2008-12-25 | BODY-WORN SENSOR FEATURING A LOW-POWER PROCESSOR AND MULTI-SENSOR ARRAY FOR MEASURING BLOOD PRESSURE | 2008-06-25 | A system is described that continuously measures a patient's blood pressure over a length of time. The system features a sensor assembly featuring a flexible cable configured to wrap around a portion of a patient's arm. The flexible cable features a back surface that includes at least two electrodes that are positioned to contact the patient's skin to generate electrical signals. It additionally features an optical sensor that includes at least one light source and at least one photodetector. These components form an optical sensor that is configured to generate an optical signal by detecting optical radiation emitted by the at least one light source and reflected from a blood vessel underneath the patient's skin. | 1. A system for measuring a patient's blood pressure over a length of time, the system comprising a sensor assembly featuring a flexible cable configured to wrap around a portion of a patient's arm, the flexible cable having a back surface and comprising: at least two electrodes, mounted on the back surface and positioned to contact the patient's skin to generate electrical signals when the sensor assembly wraps around a portion of the patient's arm;an optical sensor, mounted on the back surface and comprising at least one light source and at least one photodetector, wherein the at least one light source and at least one photodetector are positioned to be adjacent to the patient's skin when the sensor assembly wraps around a portion of the patient's arm, wherein the optical sensor is configured to generate an optical signal by detecting optical radiation emitted by the at least one light source and reflected from a blood vessel underneath the patient's skin;the system further comprising a controller configured to be worn on the patient's body, and configured to connect to the sensor assembly through a connector, the controller comprising: i) an analog-signal processing circuit comprising a first amplifier configured to receive the electrical signals from the electrodes and generate an analog electrical waveform therefrom, and a second amplifier configured to receive the optical signal from the photodetector and generate an analog optical waveform therefrom, and further comprising an analog-to-digital converter configured to receive the analog electrical waveform and generate a digital electrical waveform therefrom, and to receive the analog optical waveform and generate a digital optical waveform therefrom; ii) a central processing circuit configured to receive the digital electrical and optical waveforms and determine a pulse transit time which is a measure of a separation in time of a first feature of the digital electrical waveform and a second feature of the digital optical waveform, and to use the pulse transit time to determine a blood pressure value for a patient; and, iii) a power-regulating circuit configured to manage power supplied to the analog-signal processing circuit and central processing circuit. 2. The system of claim 1, wherein the flexible cable comprises a rectangular cross section. 3. The system of claim 2, wherein the flexible cable comprises a polymer base. 4. The system of claim 3, wherein the flexible cable comprises a first set of metal pads for mounting the at least one light source, and a second set of metal pads for mounting the at least one photodetector. 5. The system of claim 1, wherein the flexible cable further comprises at least one connector that mates to a connector comprised by a disposable electrode. 6. The system of claim 1, wherein at least one electrode is adhered to the flexible cable with an adhesive. 7. The system of claim 1, wherein the flexible cable comprises a first connector in electrical c... | Triage Wireless, Inc. | A61B005/021
US20080294025 | A1 | 2008-11-27 | Anti-Inflammatory Signal Molecules and Exercise | 2008-06-02 | The present invention provides methods for assessing the effect of exercise on an individual by measuring changes in the concentration of a compound in the individual's blood. Specifically, the invention measures changes in the concentration of an anti-inflammatory signal molecule or endogenous morphine. Anti-inflammatory signal molecules include, for example, adrenocorticotropin (ACTH), cortisol, and interleukin-10 (IL-10).;A first embodiment of the invention provides a method for assessing the effect of exercise on an individual comprising the steps of measuring a concentration of at least one compound in the individual's blood before exercise, measuring a concentration of the at least one compound in the individual's blood after exercise, and comparing the concentrations of the at least one compound. | 1. (canceled) 2. (canceled) 3. (canceled) 4. (canceled) 5. (canceled) 6. A method for treating at least one of a disease and a condition in an individual comprising: measuring a concentration of at least one anti-inflammatory signal molecule in the individual's blood before exercise;directing the individual in an exercise;measuring a concentration of the at least one anti-inflammatory signal molecule in the individual's blood after exercise;comparing the concentrations of the at least one anti-inflammatory signal molecule; andmodifying the exercise to increase a concentration of the at least one anti-inflammatory signal molecule. 7. (canceled) 8. The method of claim 6, wherein the at least one anti-inflammatory signal molecule is selected from a group consisting of adrenocorticotropin (AC TH), interleukin-10 (L-10), and morphine. 9. The method of claim 8, wherein measuring a concentration of morphine comprises assaying a biological response induced by the μ3 opiate receptor. 10. The method of claim 9, wherein the biological response comprises at least one of a change in intracellular calcium concentration and nitric oxide release. 11. The method of claim 6, wherein the exercise is at least one selected from a group consisting of riding an exercise bicycle, walking on stairs, and bouncing on a trampoline. 12. The method of claim 6, wherein the at least one of a disease and a condition is characterized by at least one of increased blood lymphocytes, increased cerebrospinal fluid lymphocytes, increased immunoglobulin synthesis, increased cytokine levels, increased acute phase protein production, inflammatory processes, sepsis conditions, and viral infection. 13. The method of claim 6, wherein the disease is at least one of cardiovascular disease, arthritis, pericarditis, vasculitis, lupus, bronchitis, phrenitis, HIV infections, schizophrenia, Alzheimer's Disease, Chrone's Disease, cholitis, diabetes mellitus, autoimmune diseases, toxoplasmosis, asthma, and multiple sclerosis. 14. A method of increasing a concentration of at least one anti-inflammatory signal molecule in an individual's blood including an exercise regimen comprising: measuring the individual's resting heart rate;directing the individual in a first aerobic exercise until the individual's heart rate stabilizes;directing the individual to rest until the resting heart rate is achieved;directing the individual in a second aerobic exercise until the individual's heart rate is about 10 beats per minute greater than the stabilized heart rate achieved during the first aerobic exercise;directing the individual to rest until the resting heart rate is achieved;directing the individual in a third aerobic exercise such that the heart rate achieved in the second aerobic exercise is achieved in less time;directing the individual to rest until the resting heart rate is achieved;directing the individual in a fourth aerobic exercise until the individual's heart rate is about 10 beats per m... | Lifewaves International, Inc. | A61B005/145 A63B071/00
US20080269579 | A1 | 2008-10-30 | System for Monitoring Changes in an Environmental Condition of a Wearer of a Removable Apparatus | 2007-04-30 | A system for detecting at least one of an environmental condition and a physical condition of a user, the system including a removable apparatus that may be placed in contact with the user when a change in at least one environmental condition is expected, a probe as part of the removable apparatus for detecting the change in the at least one environmental condition, a power source as part of the removable apparatus for powering the probe, a monitoring unit, and a communication system for communicating between the removable apparatus and the remote monitoring unit. | 1. A system for detecting at least one of an environmental condition and a physical condition of a user, the system comprising: a) a removable apparatus that may be placed in contact with the user when a change in the at least one environmental condition is expected;b) a probe as part of the removable apparatus for detecting the change in the at least one environmental condition;c) a power source as part of the removable apparatus for powering the probe;d) a monitoring unit; ande) a communication system for communicating between the removable apparatus and the remote monitoring unit. 2. The system according to claim 1, wherein the removable apparatus is a mouthpiece. 3. The system according to claim 1, wherein the removable apparatus is a patch that is attachable and removable to a part of the user's body wherein the patch does not interfere with an activity the user is participating in. 4. The system according to claim 1, wherein the removable apparatus is at least one of a piece of equipment and an article of clothing. 5. The system according to claim 1, wherein the probe comprises a computer chip. 6. The system according to claim 1, wherein the probe comprises a temperature probe. 7. The system according to claim 1, wherein the probe comprises a velocity probe to detect at least one of acceleration, deceleration, and a rate of change thereof. 8. The system according to claim 1, wherein the probe comprises an infrared spectroscopy to determine oxygen/deoxygenate hemoglobin saturation levels in the user. 9. The system according to claim 1, wherein the probe comprises an impedance/conductance measurement device to determine a change in sweating ability. 10. The system according to claim 1, further comprises a tracking device as part of at least one of the apparatus and the probe to monitor a location of the apparatus. 11. The system according to claim 10, further comprises a GPS system to locate the tracking device. 12. The system according to claim 1, wherein the power source is rechargeable. 13. The system according to claim 1, wherein the power source comprises a wire in a loop configuration to allow inductive charging of the probe. 14. The system according to claim 1, further comprises a warning device that provides information when a change in the environmental condition exceeds a predetermined limit. 15. The system according to claim 14, wherein the warning is device is at least one of part of the apparatus, part of the remote unit, and a second remote unit. 16. The system according to claim 14, wherein the warning device comprises at least one of an audible warning device and a visible warning device. 17. The system according to claim 1, further comprises a storage device to retain information about a change in the user's environmental condition. 18. The system according to claim 17, wherein the storage device is located at least one of as part of the removable apparatus and at a remote unit. 19. The system according to c... | - | A61B005/053 A61B005/00 G01P015/00
US20080269835 | A1 | 2008-10-30 | METHOD AND APPARATUS FOR DETECTION OF NERVOUS SYSTEM DISORDERS | 2006-12-12 | Systems and methods for detecting and/or treating nervous system disorders, such as seizures, are disclosed. Certain embodiments of the invention relate generally to implantable medical devices (IMDs) adapted to detect and treat nervous system disorders in patients with an IMD. Certain embodiments of the invention include detection of seizures based upon comparisons of long-term and short-term representations of physiological signals. Other embodiments include prediction of seizure activity based upon analysis of physiological signal levels. An embodiment of the invention monitors the quality of physiological signals, and may be able to compensate for signals of low signal quality. A further embodiment of the invention includes detection of seizure activity following the delivery of therapy. | 1. A method of detecting a precursor to a second neurological event following a first neurological event, the method comprising: acquiring EEG signal data following the first neurological event, the EEG signal data comprising a stream of data values;determining a short-term representation of the EEG signal data based on data values acquired over a first sample interval, the first sample interval occurring after a predetermined lock-out period following the first neurological event, the first sample interval including relatively recently acquired data values;determining a long-term representation of the EEG signal data based on data values acquired over a second sample interval, the second sample interval being longer than the first sample interval;determining an event monitoring parameter which approximates a ratio of the short-term representation to the long-term representation;comparing the event monitoring parameter to a low level threshold, the low level threshold having a value of less than about 0.5; anddetecting a precursor to a neurological event when the event monitoring parameter decreases below the low level threshold. 2. The method of claim 1 wherein the neurological event is a seizure. 3. The method of claim 1 wherein the minimum ratio threshold is less than about 0.2. 4. The method of claim 1 wherein the minimum ratio threshold is less than about 0.1. 5. The method of claim 1 where the minimum ratio threshold is about 0.08. 6. The method of claim 1 wherein the lock-out period begins upon detection of the first neurological event. 7. The method of claim 1 wherein the lock-out period begins upon termination of the first neurological event. 8. The method of claim 1 wherein the lock-out period begins when the event monitoring parameter decreases below a high level threshold after first rising above the high level threshold. 9. The method of claim 8 wherein high level threshold has a value greater than about 2. 10. The method of claim 1 wherein the lock-out period begins upon completion of therapy delivery. 11. The method of claim 10 wherein completion of therapy delivery occurs at the end of a period of electrical stimulation. 12. The method of claim 1 wherein the data values acquired over the second sample interval are comprised of the short-term representation values. 13. The method of claim 12 wherein the short-term representation values are downsampled prior to being used as data values for determining the long-term representation. 14. The method of claim 1 further comprising acquiring at least a minimum amount of EEG signal data following the lock-out period. 15. The method of claim 7 further comprising acquiring at least a minimum amount of EEG signal data following the termination of the first neurological event. 16. The method of claim 7 further comprising acquiring at least a minimum amount of EEG signal data following the lock-out period. 17. The method of claim 1 further comprising acquiring EEG signal data fo... | MEDTRONIC, INC. | A61N001/36 A61B005/0484
US20080270188 | A1 | 2008-10-30 | Graphical display of patient data | 2007-04-25 | A dynamic graphical display for presenting patient data received from a medical device automatically sorts the patient data into a Venn diagram type format. In one embodiment, the graphical display includes two or more graphic objects representing different types of physiological parameters or events derived from the physiological parameters. The graphic objects define at least a first section corresponding to a first subset of patient data, a second section corresponding to a second subset of patient data, and a third section corresponding to a conjunction of the first and second subsets of patient data. In some embodiments, a user may select one of the sections of the graphical display to obtain further details about the respective subset of patient data or conjunction of subsets of data. | 1. A method comprising: receiving patient data from a medical device;generating a graphical user interface comprising a plurality of graphic objects defining at least: a first section corresponding to a first subset of patient data;a second section corresponding to a second subset of patient data; anda third section corresponding to a conjunction of the first and second subsets of patient data; andpresenting the graphical user interface via a display. 2. The method of claim 1, wherein at least one of the first, second or third sections are selectable by a user. 3. The method of claim 1, further comprising: receiving user input selecting one of the first, second or third sections; andpresenting output data corresponding to a respective subset of patient data within a selected one of the first, second or third sections. 4. The method of claim 1, wherein receiving patient data from the medical device comprises receiving patient data from an implantable medical device via wireless telemetry. 5. The method of claim 1, wherein receiving patient data from a medical device comprises receiving physiological parameter measurements from the medical device, the method further comprising deriving events from the physiological parameter measurements for inclusion in at least one of the first or second subsets of patient data. 6. The method of claim 5, wherein the physiological parameter comprises at least one of: electrocardiogram measures, electroencephalogram measures, blood pressure measurements, electrical activity of a heart of a patient, minute ventilation, cardiovascular pressure measures, venous oxygen levels, blood pressure patient activity level or body temperature. 7. The method of claim 5, wherein deriving events from the physiological parameter measurements comprises comparing at least one of the physiological parameter measurements to a threshold value. 8. The method of claim 5, wherein the physiological parameter measurements comprise measurements of a first type of physiological parameter, wherein deriving events comprises comparing a trend in the measurement of the first type of physiological parameter to a template. 9. The method of claim 8, wherein the trend indicates a rate of change of the measurement of the first type of physiological parameter over a period of time. 10. The method of claim 1, wherein the physiological parameter measurements comprise measurements of a first type of physiological parameter and a second type of physiological parameter, wherein deriving events comprises analyzing a first frequency component of a waveform of the measurement of the first type of physiological parameter and comparing the first frequency component to a second frequency component of a waveform template. 11. The method of claim 1, further comprising receiving an indication from a patient and generating a patient marker for inclusion in at least one of the first or second subsets of patient data upon receiving the indication from the pati... | Medtronic, Inc. | A61B005/00
US20080275327 | A1 | 2008-11-06 | Three-Dimensional Adhesive Device Having a Microelectronic System Embedded Therein | 2006-03-09 | Accordingly, the present invention relates to a three-dimensional adhesive device to be attached to the body surface of a mammal comprising a microelectronic sensing system characterized by (a) a three-dimensional adhesive body made of a pressure sensitive adhesive having an upper surface and a bottom surface;(b) a microelectronic system embedded in the body of the pressure sensitive adhesive;(c) one or more cover layer(s) attached to the upper surface; and(d) optionally a release liner releasable attached to the bottom surface of the adhesive device.;Suitably the microelectronic system is a microelectronic sensing system capable of sensing physical input such as pressure, vibration, sound, electrical activity (e.g. from muscle activity), tension, blood-flow, moisture, temperature, enzyme activity, bacteria, pH, blood sugar, conductivity, resistance, capacitance, inductance or other chemical, biochemical, biological, mechanical or electrical properties. | 1: A three-dimensional adhesive device to be attached to the body surface of a mammal comprising an microelectronic system characterized by(a) a three-dimensional adhesive body made of a pressure sensitive adhesive having an upper surface and a bottom surface;(b) a microelectronic system embedded in the body of the pressure sensitive adhesive;(c) one or more cover layer(s) attached to the upper surface; and(d) optionally a release liner releasable attached to the bottom surface of the adhesive device. 2. An adhesive device according to claim 1 wherein the micro electronic system is a microelectronic sensing system. 3. The adhesive device according to claim 1 wherein the three-dimensional adhesive body has an essentially planar bottom surface adapted to adhere to the body surface of a mammal and an smooth upper surface, the adhesive device being thickest at the central part of the body and thinnest at the periphery of the body. 4. The adhesive device according to claim 1 wherein the thickness of the adhesive device at the periphery is less than 50% of the thickness of the adhesive device where it is thickest, suitably the thickness at the periphery is less than 25% of the thickness of the adhesive device where it is thickest, preferred the thickness at the periphery is less than 10% of the thickness of the adhesive device where it is thickest and most preferred the thickness at the periphery is less than 5% of the thickness of the adhesive device where it is thickest. 5. The adhesive device according to claim 4 wherein the thickness at the periphery of the adhesive device is more than 0.05 mm, suitably between 0.05 and 0.4 mm and the thickness where the adhesive device is thickest, is between 0.5 and 15 mm, suitably between 1 and 5 mm. 6. The adhesive device according to claim 1 wherein the angle between the bottom surface of the adhesive device and a line drawn from any point on the circumference of the and the point at the upper surface where the adhesive body is thickest is below 60 degrees, preferably below 45 degrees and most preferably below 30 degrees. 7. The adhesive device according to claim 1 wherein the pressure sensitive adhesive making up the three-dimensional adhesive body is a mouldable thermoplastic pressure sensitive adhesive having a flexibility enabling the adhesive device to conform to the curvature of body parts while retaining its adhesive properties even under movements. 8. The adhesive device according to claim 1 wherein the pressure sensitive adhesive is an adhesive based on polymers selected from block-copolymers such as styrene-block-copolymers, and hydrogenated styrene-block-copolymers, amorphous poly-alpha-olefins (APAO), polyacrylics, polyvinylethers, polyurethanes, polyelhylenevinylacetate, silicone, or from the group of hydrogel pressure sensitive adhesives. 9. The adhesive device according to claim 8 wherein the block-copolymers such as styrene-block-copolymers, and hydrogenated styrene-block-copolymers is ... | - | A61B005/04 B32B033/00 B23P017/00 B28B003/02
US07467009 | B2 | 2008-12-16 | Cardiac rhythm management system with noise detector | 2005-04-20 | A system, method, or device determines whether noise is present on a sampled and/or digitized sensed intrinsic cardiac signal based on a moving count of turning/inflection points of the signal. If noise is detected, the manner in which the cardiac signal is acquired, or the manner in which the device operates in response to the acquired cardiac signal (or both) is altered to reduce the risk of erroneously detecting noise as a heart depolarization and, therefore, inappropriately triggering or withholding therapy. | 1. A device for association with a heart using at least one electrode, the device comprising: a cardiac signal detector coupled to the first electrode, and including a detector output providing a sampled cardiac signal;a signal processor circuit to determine, over a predetermined plurality of cardiac signal samples, whether an evaluation sample of the cardiac signal is a turning point with respect to previous and subsequent samples, and to deem a portion of the cardiac signal to be noisy if a number of turning points exceeds a threshold value for the predetermined plurality of cardiac signal samples; andwherein, if the cardiac signal is deemed noisy, the cardiac signal detector is additionally coupled to a different second electrode for detecting depolarizations, and wherein the cardiac signal detector includes independent first and second sensing channels that are respectively coupled to the first and second electrodes for independently sensing a depolarization at each of the first and second electrodes to corroborate depolarizations sensed at the first electrode using corresponding depolarizations independently sensed at the second electrode. 2. The device of claim 1, in which the signal processor operates to determine first and second directions of the cardiac signal preceding and following the evaluation sample, respectively, and to deem the evaluation sample to be a turning point if the first direction is different from the second direction and each of the first and second directions manifest a slope of a magnitude that exceeds a corresponding first and second threshold value before deeming the evaluation sample to be a turning point. 3. The device of claim 1, in which the previous sample is taken at a first predetermined number of periodic samples away from the evaluation sample and the subsequent sample is taken at a second predetermined number of periodic samples away from the evaluation sample. 4. The device of claim 3, in which the signal processor operates to vary the first and second predetermined numbers. 5. The device of claim 1, in which the previous and subsequent samples respectively immediately precede and succeed the evaluation sample. 6. A system including the device of claim 1, and further including the first electrode, and in which the first electrode includes at least one of an intravascular electrode, an intracardiac electrode, an epicardial electrode, a housing electrode, a header electrode, and a skin surface electrode. 7. A system including the device of claim 1, and further including a user interface, remote from and communicatively coupled to the signal processor to receive an indication of whether the portion of the cardiac signal is deemed noisy. 8. The device of claim 1, in which the signal processor circuit includes an N-bit first first-in-first-out (FIFO) storage device. 9. The device of claim 8, in which the first FIFO includes a corresponding counter to sum the N-bits that have a predetermined binary va... | Cardiac Pacemakers, Inc. | A61B005/0428
US07805179 | B2 | 2010-09-28 | Method of examining dynamic cardiac electromagnetic activity and detection of cardiac functions using results thereof | 2007-06-08 | A method of examining cardiac electromagnetic activity over a heart for diagnosing the cardiac functions of the heart is disclosed. The method may include constructing a phase diagram of electromagnetic signals over a heart by collecting sets of time-dependent magnetic signals, determining the zeroth and the first derivations of each set of the magnetic signals at a given time, and categorizing the zeroth and the first derivations of the magnetic signals in either of the four phases: (+, +), (−, −), (+, −), (−, +). The method may also include monitoring a wave propagation of the magnetic signals. | 1. A method of examining cardiac electromagnetic activity of a heart, the method comprising: performing a magnetic measurement at a plurality of measurement positions of a subject to obtain a plurality of sets of spatially distributed, time-dependent magnetic signals, corresponding to the plurality of measurement positions;constructing a fitting curve to each set of the spatially distributed, time-dependent magnetic signals;determining zeroth and first derivations of the fitting curve each set of the spatially distributed, time-dependent magnetic signals at a given time;categorizing the zeroth and the first derivations of the each set of the spatially distributed, time-dependent magnetic signals in one of four phases : (+, +), (−, −), (+, −), (−, +); andmapping the resulting phase with regions of the heart. 2. The method of claim 1, wherein the magnetic signals comprise a magnetic field component (Bz) normal to a surface of the subject. 3. The method of claim 1, wherein the magnetic signals are two-dimensionally or three-dimensionally distributed over a heart of the subject. 4. The method of claim 1, wherein the given time of the each set of the magnetic signals is at a turning point (TP) of the fitting curve at which a second derivation of the each set of the spatially distributed, time-dependent magnetic signals is zero. 5. The method of claim 4, wherein the plurality of sets of spatially distributed, time-dependent magnetic signals is representative of an intramyocardial, electrical behavior of said subject, and exhibits features of at least a P-wave, a Q-wave, a R-wave, a S-wave and a T-wave, and the given time is at the turning point of a ST segment of the each set of the spatially distributed, time-dependent magnetic signals. 6. The method of claim 5, wherein the each set of the magnetic signals is offset by zeroing an interval of the magnetic signals before the P-wave. 7. The method of claim 4, wherein the first derivation of the each set of the magnetic signals is calculated at about 0.01 to about 0.15 second after the turning point. 8. A method of examining cardiac electromagnetic activity for diagnosing cardiac functions of a heart, the method comprising: performing a non-contact magnetic measurement over a subject's thorax region at a plurality of measurement positions to obtain a plurality of sets of spatially distributed, time-varying magnetic signals of the heart of the subject, corresponding to the plurality of the measurement positions, wherein each set of the spatially distributed, time-varying magnetic signals is representative of an intramyocardial, electrical behavior of the subject and comprises features of at least a P-wave, a Q-wave, a R-wave, a S-wave and a T-wave;offsetting the each set of the spatially distributed, time-varying magnetic signals;constructing a fitting curve of polynomial function to each set of the off-set, spatially distributed, time-varying magnetic signals;identifying a time point on the fitting... | - | A61B005/05
US07438688 | B2 | 2008-10-21 | Apparatus and method for measuring pulse rate | 2006-08-28 | In an apparatus adapted to measure a pulse rate of a subject, a plurality of band pass filters, each of which has a center frequency that is different from the others are provided. The band pass filters are connected in parallel to one another so as to be adapted to receive a first signal corresponding to a pulse wave of the subject and be operable to output a plurality of second signals each of which is obtained by filtering the first signal with each center frequency. A first processor is operable to obtain a fundamental frequency of the first signal. A selector is operable to select one of the second signals which is output from one of the band pass filters having such a center frequency that is closest to the fundamental frequency. A second processor is operable to calculate the pulse rate from the selected one of the second signals. | 1. An apparatus, adapted to measure a pulse rate of a subject, comprising: a plurality of band pass filters, each of which has a center frequency that is different from the others, the band pass filters being connected in parallel to one another so as to be adapted to receive a first signal corresponding to a pulse wave of the subject and be operable to output a plurality of second signals each of which is obtained by filtering the first signal with each center frequency;a first processor, operable to obtain a fundamental frequency of the first signal;a selector, operable to select one of the second signals which is output from one of the band pass filters having such a center frequency that is closest to the fundamental frequency; anda second processor, operable to calculate the pulse rate from the selected one of the second signals. 2. The apparatus as set forth in claim 1, wherein: the first processor is operable to perform a fast Fourier transform processing to obtain the fundamental frequency. 3. The apparatus as set forth in claim 2, wherein: the first processor is operable to determine a frequency having a maximum amplitude as the fundamental frequency. 4. The apparatus as set forth in claim 1, wherein: the second processor is operable to calculate the pulse rate on the basis of an interval between time points at which a voltage of the selected one of the second signals either decreases to zero or increases to zero. 5. The apparatus as set forth in claim 4, further comprising: a sound generator, operable to generate a sound at every one of the time points. 6. The apparatus as set forth in claim 1, further comprising: a noise separator, operable to separate noise superimposed on the first signal. 7. The apparatus as set forth in claim 6, wherein: the first signal is at least one of a red light signal and an infrared light signal supplied from a pulse photometer. 8. A method for measuring a pulse rate of a subject, comprising: providing a plurality of band pass filters, each of which has a center frequency that is different from the others, the band pass filters being connected in parallel to one another so as to be adapted to receive a first signal corresponding to a pulse wave of the subject and be operable to output a plurality of second signals each of which is obtained by filtering the first signal with each center frequency;obtaining a fundamental frequency of the first signal;selecting one of the second signals which is output from one of the band pass filters having such a center frequency that is closest to the fundamental frequency; andcalculating the pulse rate from the selected one of the second signals. 9. The method as set forth in claim 8, wherein: a fast Fourier transform processing is performed to obtain the fundamental frequency. 10. The method as set forth in claim 9, wherein: a frequency having a maximum amplitude which is obtained by the fast Fourier transform processing is determined as the fundamental frequenc... | Nihon Kohden Corporation | A61B005/02
US07435221 | B1 | 2008-10-14 | System and method for detecting abnormal respiration based on intracardiac electrogram signals using a pattern recognition device | 2006-11-10 | Techniques are provided for detecting abnormal respiration within a patient based upon intracardiac electrogram (IEGM) signals or other electrical cardiac signals. Briefly, abnormal respiration is detected using a pattern recognition trained to discriminate normal and abnormal respiration based on morphological parameters and interval-based parameters extracted from the IEGM signals. In addition, techniques are described for distinguishing among different cardiac rhythm types within the patient while using one or more pattern classifiers or other pattern recognition devices. | 1. A method for detecting abnormal respiration using an implantable medical device, the method comprising: sensing cardiac electrical signals within a patient in which the device is implanted;determining a cardiac rhythm type of the patient;identifying cardiac cycles within the electrical signals;detecting one or more parameters associated with the cardiac cycles and as a function of the cardiac rhythm type; anddetecting abnormal respiration by applying the parameters of the cardiac cycles to a device trained to discriminate between normal and abnormal respiration for the cardiac rhythm type. 2. The method of claim 1 wherein detecting abnormal respiration by applying the parameters of the cardiac cycles to the trained device is performed to discriminate between normal respiration and apneic/hypopneic respiration. 3. The method of claim 1 wherein detecting abnormal respiration by applying the parameters of the cardiac cycles to the trained device is performed to discriminate among normal respiration, obstructive apneic/hypopneic respiration and central apneic/hypopneic respiration. 4. The method of claim 1 wherein detecting one or more parameters includes detecting interval-based parameters associated with the cardiac cycles. 5. The method of claim 4 wherein the interval-based parameters include one or more of R-R, P-R, P-P and QT intervals. 6. The method of claim 1 wherein the morphological parameters include one or more of vPPI, tPDI, aPDI, ffPDI, and fftPDI. 7. The method of claim 1 wherein detecting parameters associated with the cardiac cycles includes extracting time-domain features from the parameters, the time-domain features being applied to the trained device. 8. The method of claim 7 wherein the time-domain features include mean and standard deviations of the parameters. 9. The method of claim 1 wherein detecting parameters associated with the cardiac cycles includes frequency features from the parameters, the frequency-domain features being applied to the trained device. 10. The method of claim 1 wherein detecting abnormal respiration by applying the parameters of the cardiac cycles to the trained device further includes evaluating a severity of abnormal respiration. 11. The method of claim 1 wherein detecting abnormal respiration by applying the parameters of the cardiac cycles to the trained device is performed by applying only those parameters detected within a particular epoch of time. 12. The method of claim 1 further including detecting whether the patient is asleep and wherein detecting abnormal respiration by applying the parameters of the cardiac cycles to the trained device is performed only using parameters detected while the patient is asleep. 13. The method of claim 1 wherein the implantable device is equipped to perform one or more of bi-ventricular and bi-atrial sensing and wherein detecting abnormal respiration is performed by applying parameters derived from cardiac cycles extracted from one or more of bi-v... | Pacesetter, Inc. | A61B005/08 A61B005/02 A61B005/04
US07468031 | B2 | 2008-12-23 | Medical measuring apparatus | 2007-01-12 | This invention discloses a medical measuring apparatus that includes a medical measuring body and a transparent display unit. The medical measuring body is used for measuring at least one physiological information such as body temperature, pulse, blood pressure, blood sugar, blood oxygen, uric acid, cholesterol, pH value, body fat or bone density. The transparent display unit includes two corresponding display surfaces for displaying the physiological information, and the two display surface are mirrored or inverted, so that the two display surfaces can display identical physiological information to provide different view angles from different directions. | 1. A medical measuring apparatus, comprising: a medical measuring body for measuring at least one physiological metric and obtaining physiological information; anda transparent display unit having a first and second opposing display surfaces, said transparent display unit for displaying a positive image and an inverted image, said positive image being viewed on said first display surface and said inverted image being viewed on said second display surface thereby providing said positive image to be viewed from said second display surface through said transparent unit for displaying physiological information. 2. The medical measuring apparatus of claim 1, wherein said physiological information is shown in a form of image, figure or word. 3. The medical measuring apparatus of claim 1, wherein said display surfaces of said transparent display unit are mirrored or inverted to display identical physiological information on said display surfaces. 4. The medical measuring apparatus of claim 1, wherein said transparent display unit has at least one light source for assisting the illumination and display. 5. The medical measuring apparatus of claim 1, wherein said transparent display unit includes the light source with at least one color light for the illumination and display, so that said transparent display unit produces a color change. 6. The medical measuring apparatus of claim 1, wherein said transparent display unit has a transparency ranging from 0% to 100%. 7. The medical measuring apparatus of claim 1, wherein said transparent display unit includes a lens with at least one color. 8. The medical measuring apparatus of claim 1, wherein said physiological information includes body temperature, pulse, blood pressure, blood sugar, blood oxygen, uric acid, cholesterol, pH value, body fat or bone density. | Health & Life Co., Ltd. | A61B005/00
US07452337 | B2 | 2008-11-18 | Hand-held hearing screener apparatus | 2004-01-15 | A hearing screener apparatus is disclosed and may include a housing and a testing probe operatively coupled to the housing. The testing probe may generate electrical signals based on otoacoustic emissions of the inner ear of a test subject, when the testing probe is inserted into the ear canal of the test subject. The hearing screener apparatus may also include a housing, a digital signal processor (DSP) within the housing, and a testing probe operatively coupled to the housing. The DSP may generate measurement data based on the electrical signals. At least one microphone may be mounted with the testing probe for generating the electrical signals based on the otoacoustic emissions. The testing probe may be vibrationally isolated from the housing. The testing probe may be elastically coupled to the housing. The hearing screener apparatus may also include an isolation body elastically coupled between the testing probe and the housing. | 1. A hearing screener apparatus comprising: a housing;a digital signal processor (DSP) within the housing;a testing probe operatively coupled to the housing, wherein the testing probe generates electrical signals based on otoacoustic emissions, when the testing probe is inserted into the ear canal of a test subject, and wherein the DSP generates measurement data based on the electrical signals; andan isolation body elastically coupled between the testing probe and the housing. 2. The hearing screener apparatus according to claim 1, comprising a display for displaying the generated measurement data, said display arranged in the housing. 3. The hearing screener apparatus according to claim 2, wherein the display comprises an LCD display. 4. The hearing screener apparatus according to claim 1, comprising at least one microphone mounted with the testing probe for generating the electrical signals based on the otoacoustic emissions. 5. The hearing screener apparatus according to claim 1, wherein the testing probe is vibrationally isolated from the housing. 6. The hearing screener apparatus according to claim 1, wherein the testing probe is elastically coupled to the housing. 7. The hearing screener apparatus according to claim 1, comprising an ear tip mounted on the testing probe for acoustically sealing the ear canal of a test subject, when the testing probe is inserted into the ear canal of the test subject. 8. The hearing screener apparatus according to claim 1, wherein the generated electrical signals are used for distortion-product otoacoustic emissions (DPOAE) testing. 9. The hearing screener apparatus according to claim 1, wherein the housing comprises at least one control button for controlling operation of the hearing screener apparatus. 10. The hearing screener apparatus according to claim 1, wherein the housing comprises at least one infrared connection interface for communicating at least a portion of the generated measurement data to at least one peripheral device communicatively coupled to the hearing screener apparatus. 11. The hearing screener apparatus according to claim 10, wherein the at least one infrared connection interface communicates at least a portion of the generated measurement data to a printer which is communicatively coupled to the hearing screener apparatus. 12. The hearing screener apparatus according to claim 11, wherein the printer comprises a printer with infrared communication capabilities. | Etymotic Research, Inc. | A61B005/00
US07462152 | B2 | 2008-12-09 | Method and system utilizing SpO | 2006-09-07 | A method and system for operating a non-invasive blood pressure monitor that utilizes an SpO2 plethysmographic signal to reduce the time required to obtain an estimation of a patient's blood pressure. During operation of the NIBP monitor, the NIBP monitor utilizes the SpO2 plethysmographic signal to determine a timing period and a deflation period for each pulse associated with the patient's heartbeat. Upon receiving an oscillation pulse, the NIBP monitor determines the oscillation amplitude during the timing period and deflates the blood pressure cuff during the deflation period immediately following the timing period. Preferably, the deflation period occurs during the same oscillation pulse used to calculate the oscillation pulse amplitude to decrease the amount of time required to obtain a blood pressure estimate from the patient. | 1. A method of monitoring blood pressure in a patient, the method comprising the steps of: providing a non-invasive blood pressure (NIBP) monitor having a selectively inflatable and deflatable blood pressure cuff and at least one pressure transducer for detecting oscillation pulses from the patient;acquiring a waveform of the patient, the waveform having a series of time-spaced pulses each representing a patient heartbeat;analyzing the waveform to determine a timing period and a deflation period based upon the waveform;inflating the pressure cuff to a first pressure level;detecting the beginning of one of the oscillation pulses from the pressure transducer;determining the amplitude of the oscillation pulse at the first pressure level during the timing period following the beginning of the oscillation pulse; anddeflating the blood pressure cuff during the deflation period to a second pressure level, wherein the deflation period follows the timing period. 2. The method of claim 1 wherein the timing period is the amount of time from a diastolic point of the patient heartbeat to the peak of the pulse. 3. The method of claim 2 wherein the deflation period is less than the amount of time from the peak of the pulse to the diastolic point of the next pulse. 4. The method of claim 1 wherein the deflation of the blood pressure cuff occurs during the deflation period independent of the pressure reduction during deflation of the pressure cuff. 5. The method of claim 1 wherein the waveform is a plethysmographic waveform received from a pulse oximeter sensor. 6. The method of claim 1 wherein the blood pressure cuff is deflated to the second pressure level during a portion of a single oscillation pulse. 7. The method of claim 1 wherein the amplitude of only a single oscillation pulse is detected at each pressure level. 8. The method of claim 1 further comprising the steps of: deflating the pressure cuff in a series of pressure steps each determined by the deflation period;determining the amplitude of a single oscillometric pulse at each pressure level; anddetermining the systolic pressure, mean arterial pressure and diastolic pressure for the patient based upon the amplitude of the oscillation pulses determined at each pressure level. 9. A method of monitoring the blood pressure of a patient, the method comprising the steps of: providing a non-invasive blood pressure (NIBP) monitor having a selectively inflatable and deflatable blood pressure cuff and at least one pressure transducer for detecting oscillation pulses from the patient;acquiring a plethysmographic waveform of the patient, the plethysmographic waveform having a series of time-spaced plethysmographic pulses each representing a patient heartbeat;analyzing the plethysmographic waveform to determine a timing period;inflating the blood pressure cuff to an initial inflation pressure;detecting the beginning of an oscillation pulse from the pressure transducer at a first pressure level of the blood... | The General Electric Company | A61B005/02
US07452336 | B2 | 2008-11-18 | Interactive neural training device | 2004-09-03 | An apparatus and method for rehabilitating patients and assessing tactile skills in response to visual stimuli generated from a portable neural training device. The device generates a visual output to elicit tactile response (hand movement) from the patient. The device preferably automatically adjusts the difficulty of a session in response to demonstrated patient skill level (i.e. combination of correct responses, speed, and accuracy of responses). The invention evaluates and challenges the patient to control tactile response, overcome visual perception difficulties, or practice simple multi-tasking or reasoning skills. The invention also provides information which is valuable to those administering the rehabilitation program or assessing patient skill and/or progress. By way of example the neural training device is described with an embodiment having discrete fixed keys and an embodiment utilizing a touch sensitive display screen. | 1. An apparatus for assisting in the evaluation and rehabilitation of an individual having a neurological impairment, comprising: (a) a touch pad having a plurality of primary touch sensitive keys arranged in a matrix;(b) a secondary touch sensitive key positioned in the approximate center of each primary touch sensitive key;(c) a visual stimulus emitting source positioned in the approximate center of each said secondary touch sensitive key upon which user tactile response is to be directed;(d) a microprocessor and memory for storing executable programming; and(e) programming associated with said microprocessor for carrying out the operations of (i) providing visual stimuli to a user,(ii) monitoring position and timing of user tactile feedback in response to said visual stimuli, and lowering registered positional accuracy of tactile feedback in response to detecting user contact with the primary touch sensitive key when responding to the visual stimuli on said secondary touch-sensitive key, and from which incorrect tactile feedback is distinguished in response to detecting user contact outside of either the secondary or primary touch sensitive key,(iii) controlling presentation of said visual stimuli to the user and for processing said tactile feedback. 2. An apparatus as recited in claim 1, further comprising means for passively restraining a limb to restrict which limb user tactile feedback will be collected. 3. An apparatus as recited in claim 1, further comprising programming for carrying out the operations of: generating one or more patterns of visual stimuli;enabling one or more of said touch sensitive keys to be responsive to tactile feedback from the user in response to said visual stimuli; andaltering the number and position of said touch sensitive keys which are responsive to tactile feedback from the user. 4. An apparatus as recited in claim 1, further comprising programming for supporting an administrator mode in which access is provided to performance statistics and for establishing device operating characteristics. 5. An apparatus as recited in claim 4, wherein said administrative mode programming is configured for providing access to said programming in said microprocessor for controlling operation of said apparatus. 6. An apparatus as recited in claim 1, further comprising programming for providing a responsive mode in which the user provides a tactile response to each visual stimuli. 7. An apparatus as recited in claim 6, wherein said responsive mode comprises: providing said visual stimuli;measuring response time and positional precision of the user's tactile feedback; andadjusting the number and position of said touch sensitive keys based on said response time and precision of user tactile feedback. 8. An apparatus as recited in claim 1, further comprising programming for providing a predictive mode in which response to the stimuli requires the user to extend a sequence. 9. An apparatus as recited in claim 8, wherein sa... | Interactive Neuro Technology, Inc. | A61B005/00 A61B005/04 A61B005/103 A61B005/117 A61B003/00 A61N001/00
US07432949 | B2 | 2008-10-07 | Mobile videoimaging, videocommunication, video production (VCVP) system | 2004-08-20 | A mobile self-powered videoimaging, video communication, video production (VCVP) system designed specifically for health care industry that provides high-resolution audio, video and data communications, production and recording capabilities at hospital operating room/procedure room or field environments for transmission to other remote locations. The VCVP system generally comprises a mobile platform with a plurality of cameras, at least one being mounted on an extensible boom for overhead imaging of surgical procedures. An array of video production equipment is rack-mounted inside the platform, as is an array of network teleconferencing equipment. The mobile imaging system may be parked at a convenient location in an operating room or other environment and controlled by a single operator. As the surgery proceeds, the operator can multiplex the camera signals into a live high resolution video/audio feed that is networked in real time for teleconferencing, and/or recorded to a hard drive or in any known format such as Mini-DV, S-VHS, VHS and DVD as desired. | 1. A mobile videoimaging station, comprising: a wheeled cabinet;a height-adjustable and length-adjustable camera support comprising a mast mounted to and extending upward from said cabinet, and a telescoping boom pivotally mounted at one end to said mast;a first video camera having remotely controlled pan, tilt and zoom functions mounted distally on another end of said boom;a second video camera mounted on said mast;controls mounted on said wheeled cabinet for remotely controlling the pan, tilt and zoom functions of said first video camera;a plurality of displays for displaying outputs from said first video camera and said second video camera;a video production system including a switch for compiling a single video feed from said first video camera and said second video camera; anda teleconferencing system for network conferencing of said video feed in real time;a power supply mounted inside the cabinet;whereby said videoimaging station may be controlled by a single operator to produce said video feed. 2. The mobile videoimaging station according to claim 1, wherein said plurality of video displays comprises a mini-display panel including a plurality of displays for displaying outputs from the first video camera and the second video camera to an operator. 3. The mobile videoimaging station according to claim 2, wherein said plurality of video displays further comprises an active broadcast preview monitor for displaying the live video feed compiled from said first video camera and said second video camera to onlookers. 4. The mobile videoimaging station according to claim 1, wherein said teleconferencing system comprises a codec for networked videoconferencing. 5. The mobile videoimaging station according to claim 4, wherein said video production system comprises a VHS/Mini-DV recorder for recording said video feed. 6. The mobile videoimaging station according to claim 5, wherein said video production system comprises an optical media recorder for recording said video feed. 7. The mobile videoimaging station according to claim 1, wherein said power supply comprises a rechargeable battery power supply. 8. A mobile videoimaging station, comprising: a wheeled cabinet having a work surface;a height-adjustable mast extending vertically from the work surface of said cabinet;a length-adjustable boom extending from said mast;a plurality of remotely-controlled video cameras, including at least one said camera having remotely controlled pan, tilt and zoom functions mounted distally on one end of said boom;controls mounted on the work surface of said mobile platform for remotely controlling the pan, tilt and zoom functions of said at least one camera;a video production system mounted inside the cabinet for mixing, producing and recording a single video feed from said plurality of cameras; anda teleconferencing system mounted inside the cabinet for network conferencing of said video production in real time; a rechargeable battery power supply;whereby s... | - | H04N007/14 A61B005/00
US07457656 | B2 | 2008-11-25 | Medical image management system | 2004-11-30 | Medical image management system to allow any conventional Internet browser to function as a medical workstation. The system may be used to convert medical images from a plurality of image formats to browser compatible format. The Invention also relates to methods to manipulate digital medical images in such a way that multiple imaging modalities from multiple different vendors can be assembled into a database of Internet standard web pages without loss of diagnostic information. | 1. A method of managing medical information comprising: receiving at least one digital medical image in a format that is incompatible with viewing in an Internet web browser;predetermining without user input at least one image display setting of the at least one digital medical image to permit medical diagnosis by the user of the at least one digital medical image having the at least one image display setting applied thereto;converting the at least one digital medical image using the predetermined at least one image display setting,wherein the converted at least one digital medical image is compatible with viewing in an Internet web browser and retains its original resolution. 2. The method of claim 1, wherein predetermining the at least one image display setting comprises determining a region of interest of the at least one digital medical image; andcropping the at least one digital medical image around the region of interest. 3. The method of claim 2, further comprising determining at least one of a brightness and a contrast of the region of interest to form at least one cropped and adjusted image. 4. The method of claim 3, further comprising displaying a plurality of cropped and adjusted images side by side in the Internet web browser. 5. The method of claim 2, further comprising grouping a plurality of cropped images to form a cropped movie. 6. The method of claim 5, further comprising determining at least one of the brightness and the contrast of the plurality of cropped images to form a cropped and adjusted movie. 7. The method of claim 5, further comprising determining a frame speed of the cropped movie to simulate real-time motion of an organ. 8. The method of claim 5, further comprising displaying a plurality of cropped movies side by side in the Internet web browser. 9. The method of claim 1, wherein predetermining the at least one image display setting comprises determining a magnification of the at least one digital medical image. 10. The method of claim 1, wherein predetermining the at least one image display setting comprises determining at least one of a brightness and a contrast of the at least one digital medical image to form at least one adjusted image. 11. The method of claim 10, further comprising displaying a plurality of adjusted images side by side in the Internet web browser. 12. The method of claim 1, wherein predetermining the at least one image display setting comprises grouping a plurality of digital medical images to form a movie. 13. The method of claim 12, further comprising determining at least one of the brightness and the contrast of the plurality of digital medical images to form an adjusted movie. 14. The method of claim 12, further comprising determining a frame speed of the movie to simulate real-time motion of an organ. 15. The method of claim 12, further comprising displaying a plurality of movies side by side in the Internet web browser. 16. The method of claim 1, wherein the at least one d... | Heart Imaging Technologies, LLC | A61B005/05 G06F007/00 G06K009/00 G06K009/54
US07471972 | B2 | 2008-12-30 | Sensor head for use with implantable devices | 2004-12-22 | The present invention provides a sensor head for use in an implantable device that measures the concentration of an analyte in a biological fluid which includes: a non-conductive body; a working electrode, a reference electrode and a counter electrode, wherein the electrodes pass through the non-conductive body forming an electrochemically reactive surface at one location on the body and forming an electronic connection at another location on the body, further wherein the electrochemically reactive surface of the counter electrode is greater than the surface area of the working electrode; and a multi-region membrane affixed to the nonconductive body and covering the working electrode, reference electrode and counter electrode. In addition, the present invention provides an implantable device including at least one of the sensor heads of the invention and methods of monitoring glucose levels in a host utilizing the implantable device of the invention. | 1. A sensor for use in a glucose measuring device, the sensor comprising: a first electrode, a second electrode, and a non-conductive body located between the first electrode and the second electrode, wherein the first electrode and the second electrode each form an electrochemically reactive surface at one end of the sensor and an electronic connection at another end of the sensor, wherein an electrochemically reactive surface area of the second electrode is greater than an electrochemically reactive surface area of the first electrode; anda multi-region membrane covering the first electrode and the second electrode, wherein the multi-region membrane comprises an immobilized enzyme domain comprising an enzyme in at least a portion thereof, and wherein the immobilized enzyme is deactivated over the electrochemical reactive surface area of second electrode. 2. The sensor of claim 1, wherein a portion of the multi-region membrane covering the second electrode further comprises an additional domain capable of selectively blocking glucose from an area over the electrochemically reactive surface of the second electrode. 3. The sensor of claim 1, wherein the multi-region membrane comprises an oxygen antenna domain. 4. The sensor of claim 1, wherein the multi-region membrane comprises a first region distant from the electrochemically reactive surfaces, a second region less distant from the electrochemically reactive surfaces than the first region and a third region less distant from the electrochemically reactive surfaces than the second region. 5. The sensor of claim 4, wherein the first region comprises a cell impermeable domain. 6. The sensor of claim 4, wherein the second region includes a glucose exclusion domain that does not substantially cover the electrochemically reactive surface of the first electrode. 7. The sensor of claim 4, wherein the third region comprises the immobilized enzyme domain distant from the electrochemically reactive surfaces and an electrolyte-containing domain less distant from the electrochemically reactive surfaces. 8. The sensor of claim 4, wherein the third region comprises the immobilized enzyme domain distant from the electrochemically reactive surfaces and a resistance domain more distant from the electrochemically reactive surfaces than the immobilized enzyme domain. 9. The sensor of claim 4, wherein the first region is permeable to oxygen and glucose. 10. The sensor of claim 4, wherein the second region is permeable to oxygen and interferes with glucose transport across the membrane, further wherein the second region does not cover the electroehemically reactive surface of the first electrode. 11. The sensor of claim 1, wherein the first electrode comprises a working electrode. 12. The sensor of claim 11, wherein the second electrode comprises a reference electrode or counter electrode. 13. A sensor for use in a glucose measuring device, the sensor comprising: a first electrode, a second electrode, and... | DexCom, Inc. | A61B005/05
US08346331 | B2 | 2013-01-01 | Deception detection and query methodology for determining deception via neuroimaging | 2008-04-18 | Functional near-infrared (fNIR) neuroimaging is used to assess credibility, detect deception, and implement a query methodology for determining deception via neuroimaging. Oxygenation levels of portions of the brain are imaged via fNIR spectroscopy and utilized to determine if the subject is telling a lie or a truth. In an example configuration, oxygenation levels in the inferior and/or middle prefrontal cortical areas of the brain, such as the bilateral dorsolateral prefrontal and/or inferior frontal cortex, are measured to determine if a subject is lying relative to telling the truth. An example system includes a portable, flexible, belt like sensing device that is positioned proximate the subjects scalp. Sensed neural activity is transmitted either through wired or wireless means, to a processor for analysis of the sensed neural activity. The query methodology utilizes an attestation assertion that mitigates variance in brain responses due to the length or form of a question. | 1. A method comprising: measuring near-infrared optical energy indicative of oxygenation levels in a blood supply of a cortex of a brain of a subject to obtain a plurality of neural responses of the subject each time the subject answers a question of a respective plurality of questions, wherein at least one of the plurality of questions is repeated a plurality of times;measuring near-infrared optical energy indicative of oxygenation levels in the blood supply of the cortex of the brain of the subject to obtain a neural response of the subject associated with the subject's response to a visual cue of an affirmation attestation, wherein the affirmation attestation is indicative of a query as to the truthfulness of a response to a previous one of the plurality of questions;comparing the oxygenation levels indicated by the measured near-infrared optical energy with a baseline oxygenation level, wherein the baseline oxygenation level is indicative of one of a known truthful response or a known false response; andif the oxygenation levels indicated by the measured near-infrared optical energy differs from the baseline oxygenation level, determining that the subject conducted at least one of deception in response to a question or malingered in response to a question. 2. The method of claim 1, further comprising: positioning a sensing device proximate to a scalp of the subject;providing near-infrared optical energy directed toward the scalp; andin response to providing near-infrared optical energy, receiving reflected near-infrared optical energy indicative of the oxygenation level in the blood supply of the cortex of the brain of the subject. 3. The method of claim 2, wherein the near-infrared optical energy is provided at a relatively constant intensity. 4. The method of claim 1, further comprising: converting the received near-infrared optical energy to electrical energy; andproviding a signal indicative of the electrical energy for determining if the subject conducted deception. 5. The method of claim 1, wherein the baseline oxygenation level is an oxygenation level associated with the subject telling a truth. 6. The method of claim 1, further comprising determining the subject conducted deception if the oxygenation level indicated by the received near-infrared optical energy is greater than the baseline oxygenation level. 7. The method of claim 1, wherein the near-infrared optical energy comprises a wavelength in the range of 700 to 900 nanometers, inclusive. 8. The method of claim 1, wherein the cortex of the brain comprises at least one of a dorsolateral prefrontal cortex and an inferior frontal cortex. 9. The method of claim 1, wherein the oxygenation level in the blood supply of the cortex is indicative of at least one of deoxygenated hemoglobin and oxygenated hemoglobin. 10. A system comprising: a sensing device configured to: sense near-infrared optical energy indicative of oxygenation levels in a blood supply of a cortex of a brain o... | Drexel University Philadelphia Health & Education Corporation | A61B005/1455 A61B005/04
US07862520 | B2 | 2011-01-04 | Body fluid sampling module with a continuous compression tissue interface surface | 2008-06-20 | A sampling module includes a housing and a penetrating member positioned in the housing. An attachment device is coupled to the housing. The attachment device is tiltably attached to a tissue interface that permits the tissue interface surface to tilt. The tissue interface includes an opening for the penetrating member to pass there through. | 1. A sampling module, comprising: a housing;a penetrating member positioned in the housing;an attachment device coupled to the housing, the attachment device being tiltably attached to a tissue interface that permits the tissue interface surface to tilt forward or backward, the tissue interface including an opening for the penetrating member to pass there through;a driver that includes at least: a human interface that communicates at least one output to a user and an input from the user to actuate the penetrating member;a first processor that is fast and high power to drive the penetrating member:a second processor that is slow and low power to process the output to and input from the user and to control the power to the first processor. 2. The module of claim 1, wherein the tissue interface has a contoured geometric configuration. 3. The module of claim 1, wherein the tissue interface has a geometric configuration that matches a geometry of a finger tip. 4. The module of claim 3, wherein the tissue interface applies pressure to the finger tip when urged against the finger tip. 5. The module of claim 1, wherein the tissue interface has an ergonomically contoured pressure surface configured to conform to a finger tip. 6. The module of claim 1, further comprising: a sample reservoir positioned in the housing and adapted to receive a fluid sample from a wound created by the penetrating member as it penetrates a tissue site positioned at the tissue interface surface. 7. The module of claim 1, wherein the tissue interface includes a tissue interface member and a retainer and the attachment device attaches the retainer to the housing. 8. The module of claim 1, wherein the attachment device includes a compliant member. 9. The module of claim 1, wherein the attachment device is configured to provide that the tissue interface is free to tilt to a within a predetermined angle range relative to the housing. 10. A sampling module, comprising: a housing with an input port and a continuous compression surface that applies substantially uniform pressure against a dermal tissue target site;a penetrating member positioned in the housing;a sample reservoir coupled to an input port;a driver that includes at least: a human interface that communicates at least one output to a user and an input from the user to actuate the penetrating member;a first processor that is fast and high power to drive the penetrating member;a second processor that is slow and low power to process the output to and input from the user and to control the power to the first processor. 11. The module of claim 10, wherein the continuous compression surface is an ergonomically contoured surface. 12. The module of claim 11, wherein the ergonomically contoured surface has a geometry that matches that of a finger tip. 13. The module of claim 10, wherein the continuous pressure surface has an elliptical continuous contoured compression surface. | Pelikan Technologies, Inc. | A61B005/00 A61B017/14 A61B017/32
US08706213 | B2 | 2014-04-22 | Evaluation of sympathetic tone | 2006-03-02 | The invention relates to a method of determining the sympathetic tone including the steps of: measuring an applied stimulation at a threshold value of the stimulation in one or more sympathetic tone-neutral points and measuring an applied stimulation at the same threshold value in one or more sympathetic tone-dependent points. The invention further relates to a system for applying and measuring a stimulation, and the use of a system for applying and measuring a stimulation for determining the sympathetic tone including the steps of: measuring an applied stimulation at a threshold value of the stimulation at one or more sympathetic tone-neutral points and measuring an applied stimulation at the same threshold value of the stimulation at one or more sympathetic tone-dependent points. | 1. A method of quantitative and/or qualitative determination of sympathetic tone and/or level of stress in a human, said method including: a) storage of a calibration threshold value and a stimulation threshold value, the calibration threshold value being a quantitative measure of a nociception threshold value of a stimulation at which pain is perceived by said human in a sympathetic tone-neutral point on a human body and the stimulation threshold value being a quantitative measure of a nociception threshold value of a stimulation at which pain is perceived by said human in a sympathetic tone-dependent point on the human body, and subsequently;b) calculation of an indication value of sympathetic tone and/or level of stress by comparing the stimulation threshold value with the calibration threshold value, wherein the indication value of sympathetic tone is a measure of the sympathetic tone and/or level of stress in the human, wherein the indication value of the sympathetic tone is a mathematical combination of the calibration threshold value and the stimulation threshold value. 2. The method according to claim 1, wherein the calibration threshold value is a pre-determined or pre-established constant value which is stored permanently. 3. The method according to claim 1, wherein the calibration threshold value is zero, wherein the indication value is identical to the stimulation threshold value or is a function of the stimulation threshold value. 4. The method according to claim 1, wherein the calibration threshold value and the stimulation threshold value are measured substantially simultaneously. 5. The method according to claim 4, wherein the indication value is a function of the ratio between the calibration threshold value and the stimulation threshold value. 6. The method according to claim 4, wherein the indication value is a function of the difference between the calibration threshold value and the stimulation threshold value. 7. The method according to claim 1, wherein a general elevation of the calibration threshold value in combination with a reduction of the stimulation threshold value is an indication of stress. 8. The method according to claim 1, wherein a normal level of the calibration threshold value and a reduction of the stimulation threshold value is an indication of stress. 9. The method according to claim 1, wherein the stimulation threshold value being more than 80% of the calibration threshold value indicates that there is no significant level of stress in the human,the stimulation threshold value being between more than 55% and at most 80% of the calibration threshold value indicates that there is a low but significant level of stress in the human,the stimulation threshold value being between more than 30% and at most 55% of the calibration threshold value indicates that there a intermediate and significant level of stress in the human, andthe stimulation threshold value being at most 30% of the calibration threshold... | ULL Meter A/S | A61N001/36
US08781544 | B2 | 2014-07-15 | Multiple wavelength optical sensor | 2008-03-26 | A multiple wavelength optical sensor has an emitter configured to radiate light having a plurality of wavelengths into a tissue site. The emitter comprises a plurality of LEDs configured in an array and connected within an electrical grid. A detector is configured to receive the light after absorption by pulsatile blood flow within the tissue site. The detector generates a sensor signal capable of being processed by a patient monitor so as to derive oxygen saturation, carboxyhemoglobin, methemoglobin and total hemoglobin. | 1. A physiological sensor comprising: an emitter that emits light having a plurality of wavelengths;a detector that generates an output signal responsive to the emitted light after absorption by tissue, the detector comprising a plurality of Si photodiodes and at least one InGaAs photodiode connected in parallel;an attachment assembly that removably attaches the emitter and the detector to tissue;a spacer that provides a predetermined gap between the emitter and tissue when the emitter is attached to tissue; anda light scattering medium disposed in an optical path between the emitter and tissue;the spacer and the light scattering medium providing at least a substantially uniform illumination of tissue by the emitted light for at least one of the wavelengths,wherein the at least one InGaAs photodiode is mounted interspersed with the plurality of Si photodiodes. 2. The physiological sensor according to claim 1 wherein the light scattering medium comprises glass beads mixed with an encapsulant disposed proximate the spacer. 3. The physiological sensor according to claim 2 wherein the light scattering medium further comprises microspheres mixed with an epoxy disposed proximate the emitter. 4. The physiological sensor according to claim 3 wherein the emitter comprises an array of at least eight light emitting diodes emitting light generally centered around eight unique wavelengths. 5. The physiological sensor according to claim 4 wherein the emitter comprises an array of at least thirteen light emitting diodes emitting light generally centered around at least twelve unique wavelengths. 6. The physiological sensor according to claim 1 wherein the detector comprises two Si photodiodes and four InGaAs photodiodes all connected in parallel. 7. The physiological sensor according to claim 4 wherein the light emitting diodes emit light within a first range of about 620-905 nm. and within a second range of about 1040-1270 nm. 8. The physiological sensor according to claim 1 wherein the detector includes at least two InGaAs photodiodes mounted between two Si photodiodes. 9. A physiological sensor comprising: an emitter configured to radiate light having a plurality of wavelengths into a tissue site;the emitter comprises a plurality of LEDs disposed within an emitter ceramic substrate;a detector configured to receive the light after absorption by pulsatile blood flow within the tissue site;the detector generates a sensor signal capable of being processed by a patient monitor so as to derive total hemoglobin (Hbt);the detector comprises a plurality of photodiodes disposed within a detector ceramic substrate; anda first set of the photodiodes is responsive to a first set of the wavelengths and a second set of the photodiodes is responsive to a second set of the wavelengths,wherein the plurality of photodiodes comprises a plurality of Si photodiodes and at least one InGaAs photodiode connected in parallel, and wherein the at least one InGaAs photodiode is ... | Cercacor Laboratories, Inc. | A61B005/00
US08744154 | B2 | 2014-06-03 | System and method for acquiring magnetic resonance imaging (MRI) data | 2006-09-22 | A method for acquiring image data from a patient with a magnetic resonance imaging (MRI) system. The proposed method comprises the steps of: a) predefining a number of scan geometries for acquiring the image data from at least one region of interest (ROI) relative to the patient, b) performing at least one scan for acquiring the image data in accordance with at least one of the predefined scan geometries, c) analysing in the image data a position of the region of interest to detect a deviation from the at least one predefined scan geometry, d) changing the at least one predefined scan geometry if said deviation exceeds a predetermined threshold value, and e) repeating steps b) to d) until a predetermined number of scans has been performed. Thus, by means of the proposed method the utility of such predefined scan geometries is greatly enhanced. | 1. A method for acquiring image data from a patient with a magnetic resonance imaging system, comprising the steps of: a) predefining a scan geometry for acquiring the image data from at least one region of interest relative to the patient,b) storing a plurality of scans to be performed in a memory of a controller of a magnetic resonance scanner, the scans to be performed of the region of interest using the defined scan geometry, each scan including a plurality of frames,c) with the controller, controlling the magnetic resonance scanner to perform each of the scans stored in the memory using the defined scan geometry and including: 1) for each scan after a first one, sequentially acquiring resonance data for each frame,2) reconstructing the resonance data for a Nth frame into a Nth frame image, where N is a positive integer,3) registering the Nth frame image with a reference frame image to determine a geometry defined transformation matrix, the reference frame image determined from the plurality of scans and being indicative of a dedicated position detection of the scan geometry,4) correcting the defined scan geometry with the transform matrix when a deviation as measured by the transform matrix exceeds a predetermined threshold,5) repeating steps c1-c4 for a plurality of subsequent frames. 2. The method according to claim 1, wherein step c3 further comprises applying said geometry transformation matrix to the defined scan geometry. 3. The method according to claim 1, wherein the geometry transformation matrix is derived by registering the Nth frame image relative to an anatomical reference. 4. The method according to claim 1, wherein an anatomy scouting scan is performed prior to step b) to generate a scout image for detecting anatomical landmarks of the patient, and in that the region of interest is defined relative to said anatomical landmarks. 5. The method according to claim 1, wherein a dedicated position detection is performed prior to step a) for detecting a position of the patient, and in that the region of interest is defined relative to said patient position. 6. The method according to claim 5, wherein said dedicated position detection comprises at least one of using an orbital MR navigator, using fiducial markers, and using external optical detection means. 7. A magnetic resonance imaging system, comprising: a magnet/coil device for subjecting a patient to a magnetic field and for acquiring image data from the patient, anda control system in connection with the magnet/coil device, said control system which in one imaging session includes a processor configured to control the magnet/coil device to: perform a set of diagnostic imaging scans sharing a selected scan geometry and acquire diagnostic image data from each of the diagnostic imaging scans, each of the diagnostic imaging scans registered with a reference frame image, the reference frame image determined from the set of diagnostic imaging scans and being indicative of a ded... | Koninklijke Philips N.V. | A61B005/055
US08715201 | B2 | 2014-05-06 | Detection of body sounds | 2006-10-25 | Disclosed is a method for detecting and monitoring body sounds in humans and animals, in which bioacoustic sensors and analyzers that are mounted downstream are used for the stationary or mobile long-term monitoring of intensive care patients' respiration, for example. The patients' lung sounds are detected and stored along with measured data which are available right away especially for the early detection of diseases and acute disturbances. Adequately monitoring intestinal sounds makes it possible to evaluate peristalsis and detect mechanical/paralytic ileus early on. An early warning system for the clinical sector immediately generates signals allowing doctors and nurses to take rapid action in case of an emergency. The inventive apparatus requires a maximum of only three bioacoustic sensors (12), each of which can be fixed to a point of an object body (K) facing the object, a maximum of one sensor (14) for recording surrounding noises, a maximum of four separable channels (11) for recording and transmitting sound signals or sound data detected by the sensors (12), and devices for supplying power and forwarding, converting, storing, and displaying sequences of signals or data on or in a recorder or a computer unit (20). | 1. A method for detecting and monitoring body sounds in humans and animals, in which bioacoustics sensors and analyzers that are mounted downstream are used, wherein, the bioacoustics sensors are fixed to a maximum of three points of an object body,a maximum of one additional bioacoustics sensor is provided for recording surrounding noises,the bioacoustics sensors continuously detect sound signals and/or sound data, within a selectable time segment,sequences of the sound signals and/or sound data are digitized, stored, and converted into a plurality of frequency-dependent spectral curves,a plurality of node-defined matching lines are continuously allocated to each corresponding frequency-dependent spectral curve, andeach node-defined matching line is compared with the corresponding frequency-dependent spectral curve for analyzable and signalizable identification of specific sound symptoms. 2. The method according to claim 1, wherein a conversion time window with resolutions of approximately 0.1 s is selected for converting the digitized sequences of the sound signals and/or sound data into the plurality of frequency-dependent spectral curves. 3. The method according to claim 1, wherein a Fourier transform or a fast Fourier transform of the data is affected in a spectral range, one spectral curve per conversion time window originating, with which amplitudes are represented over a frequency, transmitted and recorded. 4. The method according to claim 1, wherein from each spectral curve, at least two frequency ranges are selected and median values are determined, which correspond to nodes of the sequences of the sound signals and/or sound data, and that the allocated matching lines are developed on a basis of the nodes. 5. The method according to claim 1, wherein, when viewing the frequency-dependent spectral curves and the node-defined matching lines in an X, Y-coordinate axis system, each node value of the node-defined matching lines is an Y-amount at a frequency point X and that an analysis by comparing Y-values of the corresponding frequency-dependent spectral curve and of the corresponding node-defined matching line is effected by means of threshold values. 6. The method according to claim 1, wherein, for the frequency-dependent spectral curves, two partial frequency ranges are selected from a band of 150 Hz to 1,600 Hz, whose scope is determined by a multiple of an admissible signal peak width, which for humans lies in the range of 5 times to 0.5 times the admissible signal width. 7. The method according to claim 6, wherein the multiple of the admissible signal peak width corresponds to approximately 100 Hz to 200 Hz. 8. The method according to claim 1, wherein a control sensor for the oxygen content of the blood is fixed on the object body. 9. The method according to claim 1, wherein a detection of wheezing events from the frequency-dependent spectral curves is by an application of plausibility criteria based on secured measuring resul... | Heinen + Lowenstein GmbH & Co. KG | A61B005/08
US09802045 | B2 | 2017-10-31 | Method for measuring stable and reproducible electrode-tissue impedance | 2007-10-25 | The present invention is a method for measuring stable and reproducible electrode-tissue impedance, comprising preconditioning an electrode-tissue interface. Further aspect of the invention is a stimulation system for a visual prosthesis generating a stimulation signal to precondition the electrode-tissue interface, comprising a computer; software, loaded in the computer, adapted to perform a stimulating method for a visual prosthesis having a plurality of electrodes; a video processing unit; and an implanted neuron-stimulator. | 1. A method for measuring stable and reproducible electrode-tissue impedance, comprising: providing a stimulation system including a computer; software, loaded in the computer, adapted to perform a stimulating method for a visual prosthesis having a plurality of electrodes;a video processing unit; andan implanted neuron-stimulator including at least one electrode;preconditioning the electrode-tissue interface by delivering a first stimulation signal, including charge balanced biphasic square wave current pulses for a duration suitable to precondition the electrode interface to the electrode, the duration being between 1 second and 10 minutes;stopping preconditioning at the end of the duration;measuring an impedance on the electrode at similar electrode-tissue interface conditions to the preconditioning step, after said preconditioning step within a predetermined delay;transmitting the impedance wirelessly to a computer; andestimating perceptual thresholds based on the impedance. 2. The method for measuring stable and reproducible electrode-tissue impedance according to claim 1 wherein the electrode-tissue impedance is measured in vivo. 3. The method for measuring stable and reproducible electrode-tissue impedance according to claim 1 wherein preconditioning of the electrode-tissue interface is carried out one electrode at a time or in groups of electrodes. 4. The method for measuring stable and reproducible electrode-tissue impedance according to claim 1 wherein preconditioning of the electrode-tissue interface is carried out for all electrodes simultaneously. 5. The method for measuring stable and reproducible electrode-tissue impedance according to claim 4 further comprising varying a combination of variable pulse-width amplitude and frequency is used to control preconditioning of the electrode-tissue interface. 6. The method for measuring stable and reproducible electrode-tissue impedance according to claim 1, comprising preconditioning at 10-300 μA at 15-120 Hz with 0.1-2.0 ms pulse width. 7. The method for measuring stable and reproducible electrode-tissue impedance according to claim 1, comprising preconditioning at 30-150 μA at 30-120 Hz with 0.2-1.0 ms pulse width. 8. The method for measuring stable and reproducible electrode-tissue impedance according to claim 1, comprising: preconditioning at 80-120 μA at 40-60 Hz with 0.4-0.6 ms pulse width. 9. The method for measuring stable and reproducible electrode-tissue impedance according to claim 1, wherein the step of measure is measuring with charge balanced square wave current pulses. | Second Sight Medical Products, Inc. | A61N001/36
US09763597 | B2 | 2017-09-19 | Local MRI breast coil and method of use | 2008-04-23 | A local breast coil is designed to be tightly coupled to the natural shape of the pendant breast to provide high SNR for diagnostic MR imaging applications, while still providing access for interventional procedures through openings in the coil. In accordance with one configuration, the coil has a symmetrical design, such that the coil can be rotated about the breast to position an opening in the coil proximate to a desired portion of the breast without incurring registration or other alignment or artifact errors due to inhomogeneous B1 excitation. Furthermore, the coil is designed to facilitate medial and lateral imaging of the breast. The coil is combined with a patient support system that facilitates rotation of the coil and interchangeability of the coil to match the configuration of the coil to the position of the breast being imaged. | 1. A patient support system configured to support a patient during an MRI process using an MRI system, the patient support system comprising: a bed having a surface configured to receive the patient to support the patient during the MRI process;a first RF coil coupled to the bed comprising: a frame comprising: an opening at a first end having a first cross-sectional area configured to receive a body part to be imaged using the MRI system;a frame body extending along a central axis of the frame from the first end toward a second end having a second cross-sectional area smaller than the first cross sectional area, the frame body rotatable about rotational axis extending along the central axis;a plurality of openings formed in the frame body to provide access from an exterior of the frame body to the body part arranged in an interior of the frame;a conductor coupled to the frame and extending about the frame to form a tapered solenoid shape comprising: a first loop arranged proximate to the first end of the frame and extending in a first plane substantially transverse to the central axis of the frame;a second loop arranged proximate to the second end of the frame and extending in a second plane substantially parallel to the first plane and transverse to the central axis of the frame;at least one leg connecting the first loop and the second loop;wherein at least a portion of the plurality of openings formed in the frame body is accessible between the first loop and the second loop;wherein the conductor forming the solenoid shape and the frame body forming the plurality of openings are symmetrical about the central axis such that rotation of the frame body and conductor about the rotational axis does not incur registration and artifact errors due to inhomogeneous B1 excitation before and after rotation; andwherein the first coil is configured to operate as both a transmit coil and a receive coil and wherein the first coil creates a homogenous B1 excitation field and a homogenous reception field. 2. The patient support system of claim 1 further comprising a second RF coil coupled to the bed comprising, the second RF coil: a frame comprising: an opening at a first end having a first cross-sectional area configured to receive a body part to be imaged using the MRI system;a frame body extending along a central axis of the frame from the first end toward a second end having a second cross-sectional area smaller than the first cross sectional area;a plurality of openings formed in the frame body to provide access from an exterior of the frame body to the body part arranged in an interior of the frame;a conductor coupled to the frame and extending about the frame to form a tapered solenoid shape comprising: a first loop arranged proximate to the first end of the frame and extending in a first plane substantially transverse to the central axis of the frame;a second loop arranged proximate to the second end of the frame and extending in a second plane substan... | Wisconsin Alumni Research Foundation | A61B005/05 A61B005/055 G01R033/34 G01R033/341 G01R033/3415
US20080255432 | A1 | 2008-10-16 | Patient Monitor with User-Defined Monitored Parameters | 2006-10-30 | In a medical monitoring system (10), one or more biometric monitors (20, 21, 22, 23, 24, 26, 66) collect data of a plurality of monitored biometric parameters. An expression evaluator (72) evaluates a user-defined expression incorporating one or more of the monitored biometric parameters to generate data for a user-defined biometric parameter. A further processing component (32, 50, 52, 54, 55, 56, 57, 58, 60, 62, 88) performs processing and display of data of at least one parameter selected from the monitored bio-metric parameters and the user-defined biometric parameter. The further processing component operates in the same way regardless of whether the at least one parameter is selected from the monitored biometric parameters or the user-defined biometric parameter. | 1. A medical monitoring system comprising: one or more biometric monitors configured to collect samples of a plurality of monitored biometric parameters;an expression evaluator configured to evaluate a user-defined expression incorporating one or more of the monitored biometric parameters to generate samples of a user-defined biometric parameters; anda further processing component configured to perform processing and display of samples of at Least one parameter selected from the monitored biometric parameters and the user-defined biometric parameter, the further processing component operating in the same way regardless of whether the at least one parameter is selected from the monitored biometric parameters or the user-defined biometric parameter. 2. The medical monitoring system as set forth in claim 1, wherein the user-defined expression evaluated by the expression evaluator further incorporates at least one patient-specific input as a monitored biometric parameter. 3. The medical monitoring system as set forth in claim 2, wherein the at least one patient-specific input is selected from a group consisting of patient weight, patient age, patient gender, patient disease type, and patient disease severity. 4. The medical monitoring system as set forth in claim 1, wherein the user-defined expression evaluated by the expression evaluator further incorporates at least one clinical monitored biometric parameter received from a network. 5. The medical monitoring system as set forth in claim 4, wherein the at least one clinical monitored biometric parameter received from the network is selected from a group consisting of blood gas data, or lab results obtained via an input/output (I/O) portion from the network. 6. The medical monitoring system as set forth in claim 1, wherein the user-defined biometric parameter includes two or more biometric parameters defined by corresponding two or more user-defined expressions each incorporating one or more of the monitored biometric parameters. 7. The medical monitoring system as set forth in claim 1, wherein the monitored biometric parameters include at least two monitored biometric parameters selected from the group consisting of: heart rate, blood pressure, temperature, blood oxygen saturation (SpO2), and electrocardiographic (ECG) data. 8. The medical monitoring system as set forth in claim 1, wherein an alarm is configured to activate responsive to one or more of (i) the monitored biometric parameters, (ii) the user-defined biometric parameters, (iii) patient-specific input parameters, or (iv) clinical monitored biometric parameters, satisfying a pre-selected alarm criterion. 9. The medical monitoring system as set forth in claim 1, wherein the further processing component comprises: a display configured to display trend samples of the at least one parameter selected from the monitored biometric parameters and the user-derived biometric parameters respective to time. 10. The medical monitoring system ... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/0205
US20080281164 | A1 | 2008-11-13 | APPARATUS AND METHOD FOR A PATIENT MONITOR | 2007-05-11 | A patient monitoring apparatus is disclosed herein. The patient monitoring apparatus includes a controller, and a speaker operatively connected to the controller. The speaker is configured to selectively produce a sound adapted to audibly entertain, distract and/or soothe a patient such that the patient can be conveniently and accurately monitored. A corresponding method for monitoring one or more of a patient's vital signs is also provided. | 1. A patient monitoring apparatus comprising: a controller; anda speaker operatively connected to the controller, wherein the speaker is configured to selectively produce a sound adapted to audibly entertain, distract and/or soothe a patient such that the patient can be conveniently and accurately monitored. 2. The patient monitoring apparatus of claim 1, wherein the patient monitoring device is selected from the group consisting of a blood pressure monitor, an electrocardiogram, an electroencephalogram, a thermometer, and a pulse oximeter. 3. The patient monitoring apparatus of claim 2, wherein the speaker is configured to selectively produce music adapted to audibly entertain, distract and/or soothe the patient. 4. The patient monitoring apparatus of claim 2, further comprising a display operatively connected to the controller, wherein the display is configured to selectively produce an image adapted to visually entertain, distract and/or soothe the patient such that the patient can be conveniently and accurately monitored. 5. The patient monitoring apparatus of claim 2, further comprising a patient interface coupling the controller with the patient. 6. The patient monitoring apparatus of claim 5, wherein the patient interface includes a sensor. 7. The patient monitoring apparatus of claim 2, further comprising an Internet connection coupled with the controller. 8. The patient monitoring apparatus of claim 2, further comprising an input port coupled with the controller. 9. A patient monitoring apparatus comprising: a controller; anda display operatively connected to the controller, wherein the display is configured to selectively produce an image adapted to visually entertain, distract and/or soothe a patient such that the patient can be conveniently and accurately monitored. 10. The patient monitoring apparatus of claim 9, wherein the patient monitoring device is selected from the group consisting of a blood pressure monitor, an electrocardiogram, an electroencephalogram, a thermometer, and a pulse oximeter. 11. The patient monitoring apparatus of claim 10, further comprising a patient interface coupling the controller with the patient. 12. The patient monitoring apparatus of claim 11, wherein the patient interface includes a sensor. 13. The patient monitoring apparatus of claim 12, further comprising a speaker operatively connected to the controller, wherein the speaker is configured to selectively produce a sound adapted to audibly entertain, distract and/or soothe the patient such that the patient can be conveniently and accurately monitored. 14. The patient monitoring apparatus of claim 13, wherein the speaker is configured to selectively produce music adapted to audibly entertain, distract and/or soothe the patient. 15. The patient monitoring apparatus of claim 14, further comprising an Internet connection coupled with the controller. 16. The patient monitoring apparatus of claim 15, further comprising an input port coupled... | GENERAL ELECTRIC COMPANY | A61B005/00
US07454248 | B2 | 2008-11-18 | Method, apparatus and product for acquiring cardiac images | 2004-01-30 | A method for acquiring cardiac information from a patient having a pacer for pacing a heart rhythm, an abnormal EKG, or an abnormal heartbeat is disclosed. The method includes: placing a signal injection device proximate the pacer of the patient and injecting a signal across a skin barrier of the patient toward the pacer; in response to the signal received at the pacer, pacing the patient's heart in a fixed asynchronous pacing mode; and, acquiring cardiac information relating to the patient's fixed asynchronously paced heart. | 1. A method for acquiring a cardiac image from a patient having a pacer for pacing a heart rhythm, an abnormal EKG, or an abnormal heartbeat, the method comprising: receiving a gated electrocardiogram signal having local maxima and minima values and trigger points;determining for a period of time the time between each trigger point and the local maxima or minima associated therewith;in response to the trigger point occurring at the associated local maxima or minima, calculating a zero time differential for a corrected trigger for gating;in response to the trigger point occurring prior to the associated local maxima or minima, calculating a time delay for the corrected trigger;in response to the trigger point occurring after the associated local maxima or minima, calculating a time advancement for the corrected trigger;sending the corrected trigger to a cardiac image acquisition device for gating, wherein the gating is substantially synchronized with the local maxima or minima of the gated electrocardiogram signal;placing a signal injection device proximate the pacer of the patient and injecting a signal across a skin barrier of the patient toward the pacer;in response to the signal received at the pacer, pacing the patient's heart in a fixed asynchronous pacing mode; andacquiring a cardiac image of the patient's fixed asynchronously paced heart. 2. The method of claim 1, wherein the signal injection device is a magnetic signal injection device comprising at least one of a fixed magnet and an electromagnet. 3. The method of claim 2, further comprising: in response to the magnetic signal received at the pacer, activating a switch at the pacer for pacing the patient in a fixed asynchronous pacing mode. 4. The method of claim 2, wherein the magnetic signal injection device is adapted to produce a magnetic signal having a signal strength at the pacer equal to or greater than about 90 Gauss when placed at a distance equal to or less than about 2 inches from the pacer. 5. The method of claim 2, wherein the magnetic signal injection device has an outside dimension equal to or greater than an outside dimension of the pacer. 6. The method of claim 1, wherein the fixed asynchronous pacing mode includes a ventricular, an atrial, or a dual chamber asynchronous pacing mode. 7. A method for acquiring cardiac information from a patient having a pacer for pacing a heart rhythm, an abnormal EKG, or an abnormal heartbeat, the method performed using an apparatus having electrocardiogram-gated acquisition and cardiac imaging capabilities for use with a patient having a pacer, the apparatus comprising an electrocardiograph, a cardiac scanner in signal communication with the electrocardiograph, an interface board in signal communication intermediate the electrocardiograph and the cardiac scanner, and a storage medium, readable by a processing circuit, storing instructions for execution by the processing circuit for performing aspects of the method, the method com... | GE Medical Systems Global Technology, LLC Jasbir S. Sra | A61N001/00 A61B006/00 A61B005/04
US07791340 | B2 | 2010-09-07 | Magnetic resonance imaging apparatus and method that produce a virtual total cylindrical maximum real measurement volume | 2008-06-13 | A magnetic resonance apparatus has a basic field magnet that generates a basic magnetic field with a homogeneity region of the basic magnetic field in which the basic magnetic field is homogeneous and with a maximum real measurement volume contained in the homogeneity region. The apparatus has a movable patient bed and a control unit that controls the movable patient bed. The maximum measurement volume thereby exhibits a cylindrical shape. A virtual total measurement volume that is greater than the maximum real cylindrical measurement volume can be generated by the control unit together with the movable patient bed. | 1. A magnetic resonance apparatus comprising: a magnetic resonance data acquisition unit having a basic field magnet that generates a basic magnetic field having a homogeneity region containing a smaller, homogenous, cylindrical, maximum real measurement volume;a patient bed that is movable through said cylindrical maximum real measurement volume; said patient bed being configured to support a patient thereon; anda control unit configured to control movement of said patient bed through said cylindrical maximum real measurement volume acquisition of magnetic resonance data by said magnetic resonance data acquisition unit in order to produce a virtual total cylindrical measurement volume composed of multiple volumes each equal to said cylindrical maximum real measurement volume, by coordinating movement of said patient bed with acquisition of magnetic resonance data in order to acquire said magnetic resonance data from different, adjacent regions of a patient on the patient bed the different, adjacent regions having an extent within a plurality of multiples of said cylindrical maximum real measurement volume. 2. A magnetic resonance apparatus as claimed in claim 1 wherein said basic field magnet comprises a plurality of concentric magnet coils arranged around a longitudinal axis of said cylindrical maximum real measurement volume. 3. A magnetic resonance apparatus as claimed in claim 1, wherein said basic field magnet comprises shielding coils configured for active shielding of magnetic fields outside of said basic field magnet. 4. A magnetic resonance apparatus as claimed in claim 1 wherein said patient bed is movable in a longitudinal direction through said cylindrical real measurement volume. 5. A magnetic resonance apparatus as claimed in claim 1 wherein said control unit is configured to move said patient bed successively to a plurality of positions in said basic field magnet, and is configured to produce said virtual total measurement volume from at least two successive individual magnetic resonance data acquisitions respectively occurring at two of said successive positions. 6. A method for implementing a magnetic resonance examination with a magnetic resonance apparatus comprising a basic field magnet that generates a basic magnetic field having a homogeneity region containing a smaller, homogenous, cylindrical maximum real measurement volume, and a movable patient bed that is movable through said cylindrical maximum real measurement volume, said method comprising the steps of: successively moving a subject on the patient bed through at least two successive positions in said cylindrical real measurement volume; andobtaining respective magnetic resonance data acquisitions at said successive positions to produce a virtual total cylindrical measurement volume that is composed of multiple volumes each equal to said cylindrical maximum real measurement volume, by coordinating movement of said patient bed with acquisition of magnetic resonan... | Siemens Aktiengesellschaft | G01V003/00
US07460998 | B2 | 2008-12-02 | Voice connection system between humans and animals | 2003-10-24 | A voice connection system comprises a sensor matrix sending to a control unit signals indicating an animal's status in terms of stimuli, feelings, events, actions or behavior of said animal. The control unit selects—depending on the signals received—a voice message recorded into memory means associated to the unit, and activates a loudspeaker belonging to the system, in order to issue said voice message. The system further comprises voice recognition means, which activate pulse generators depending on the content of a voice message uttered by the human user. | 1. Voice connection system between humans and animals, comprising: sensor means designed to be positioned on an animal for converting pulses detected on the animal's body into first electric signals indicating a status of said animal,processing means operatively associated with the sensor means, comprising memory means into which human voice messages corresponding to different statuses of the animal are recorded,loudspeaker means operatively connected to the processing means, the processing means receiving first electric signals coming from said sensor means and activating said loudspeaker means in order to issue a voice message selected in said memory means, in function of the aforesaid first electric signals received,speech recognition means for sending to the processing means second electric signals representing the content of voice messages uttered by a human user, andpulse-generating means, which receive from said processing means said second electric signals representing the content of the voice messages uttered by the human user, and which send to the animal's brain corresponding pulses,whereinsaid processing means includes a neural network control system,said sensor means comprises first and second electroencephalographic sensors and an electromyographic sensor,at least the first and second electroencephalographic sensors are placed close to respective ears of the animal, or close to its occipital-temporal region, and the electromyographic sensor is placed on the animal's neck,and wherein said pulse-generating means is constituted by a generator (10) of radioelectric waves, that converts said second electric signals into radioelectric waves having different frequencies and amplitudes, which are sent directly to the animal's brain. 2. System according to claim 1, wherein said generator (10) works with radiofrequency, so as to modulate a stead-state frequency at about 15 KHz with signals varying from 300 Hz to 4 KHz. 3. System according to claim 1, wherein said neural network control system is programmed for enabling a human/animal interactive self-learning procedure, where in particular the human user can correct or acknowledge with his/her voice messages the correctness of the voice messages issued by said loudspeaker means, and/orthe animal can hear the human voice and the corresponding radioelectric waves simultaneously, thus associating the two stimuli. 4. System according to claim 1, wherein said sensor means, said processing means, said loudspeaker means, said speech recognition means and said pulse-generating means are integrated into a collar. 5. System according to claim 1, wherein said speech recognition means are used as means integrated and/or complementary to said sensor means, in order to improve the interpretation of the animal's status as detected through said sensor means. 6. System according to claim 1, wherein said system develops in time, through an evolutional process, a language which is the animal's own language... | S.I.SV.EL. Societa' Italiana per lo Sviluppo Dell'Elettronica S.P | G10L021/00 A61B005/00
US20080269582 | A1 | 2008-10-30 | Methods and devices for surgical drains with sensors | 2008-04-28 | The present invention is directed to a system and method for postoperative monitoring of the condition of a tissue or organ utilizing sensors that may be embedded in various types of surgical drains. The system is comprised of a probe and a monitoring unit. The probe may include a surgical drain with fluid draining channels housing one or more sensors to measure various parameters of the adjacent tissue. The monitoring unit which controls the sensors of the probe may include a processor to process, record and display the measured parameters. This system may be valuable for monitoring transplanted organs and tissue grafts during the critical postoperative period when most of the clinical complications, such as vascular thrombosis, may occur. | 1. A surgical drain system for draining wound fluids and sensing a parameter of a tissue in a patient's body comprising: a) a surgical drain configured to be implanted in a patient's body, to rest against at least one tissue in the patient's body, to house at least one sensor, and to drain wound fluids from the vicinity of the tissue, comprising:i. a drain body configured to rest against the tissue within the patient's body;ii. a first surface located on an outer side of the drain body;iii. one or more draining grooves along substantially the length of the drain body;iv. a first sensor integrated with the first surface, configured to sense a parameter of the tissue proximate to the first surface; andb) a tube in fluid communication with the surgical drain configured to transport the drained wound fluids out of the body. 2. The system of claim 1, wherein the parameter sensed is selected from the group comprising: oxygen partial pressure, percent oxygen saturation, hemoglobin concentration, blood perfusion, pH, NADH concentration, humidity, biochemical composition, bilirubin concentration, amylase concentration, pus, intestinal content, drug concentration, temperature and pressure. 3. The system of claim 1, wherein the first sensor detects the level of oxygenation of the tissue. 4. The system of claim 1, wherein the first sensor detects the hemoglobin content in the tissue. 5. The system of claim 1, wherein the sensor includes at least one optical fiber. 6. The surgical drain of claim 1, further including a transmitting element configured to deliver energy to the tissue proximate the first surface. 7. The system of claim 1, wherein the first sensor is configured to sense the parameter by sensing energy that is returned from the tissue after having been transmitted to the tissue. 8. The system of claim 1, further comprising a second sensor integrated with the first surface, configured to detect a parameter of the tissue that is different from the parameter sensed by the first sensor. 9. The system of claim 1, wherein the first sensor is embedded within the surgical drain behind material that is optically transparent. 10. The system of claim 1, further including display configured to depict data corresponding to the parameter sensed by the first sensor. 11. The system of claim 1, wherein the surgical drain further includes a second surface located on an outer side of the drain body different from the first surface and a second sensor integrated with the second surface, configured to sense the same parameter of a tissue proximate to the second surface which is different from the tissue sensed by the first sensor. 12. The system of claim 11, further including a processor in communication with the first and the second sensors, and configured to compare the difference between the parameter sensed by the first and the second sensors. 13. A surgical drain system for draining wound fluids and sensing a parameter of a tissue in a patient's body c... | - | A61B005/05 A61M027/00
US20080312510 | A1 | 2008-12-18 | WELLNESS PROGRAMS, INCLUDING COMPUTER IMPLEMENTED WELLNESS PROGRAMS | 2007-06-14 | Aspects of the wellness programs are disclosed. For example, one aspect of the invention is directed toward a computing system that includes a health question generation component configured to provide a set of health questions to each individual in a group. The system further includes a health input tracking component configured to track health inputs received from each individual and to prompt individuals to provide input for one or more health questions to which selected individual(s) have not responded. The system still further includes a health reporting component configured to report information related to health profiles. In selected embodiments, the system can include a readiness question generation component configured to provide a set of readiness questions to one or more individuals, wherein the set of readiness questions are formatted to assess a readiness of the one or more individuals to take a selected action. | 1. A method in a computing system for administering at least a portion of a wellness program, the method comprising: providing a set of health questions to each individual in a group of individuals, the set of health questions being formatted to determine a health profile of each individual that responds to the set of health questions;receiving health inputs from the individuals in response to the set of health questions;analyzing the health inputs to determine a health profile for each individual in a portion of individuals that respond to the set of health questions;determining one or more individuals in the portion of individuals to provide a set of readiness questions based at least in part on a portion of the health profile of the one or more individuals, the set of readiness questions being formatted to assess a readiness of the one or more individual to take a selected action; andproviding the set of readiness questions to the one or more individuals. 2. The method of claim 1, further comprising: receiving readiness input from at least one individual in response to the set of readiness questions;analyzing the readiness input to determine the readiness of the at least one individual to take the selected action; andprovide information to the at least one individual related to the readiness of the at least one individual to take the selected action. 3. The method of claim 1, further comprising reporting information related to the health profile of the individuals in the portion of individuals that respond to at least a portion of the set of health questions. 4. The method of claim 1 wherein: providing the set of health questions includes providing the set of health questions via at least one of a phone and a computer generated display; andreceiving the health inputs includes receiving health inputs via at least one of a phone and a computer input device. 5. The method of claim 1 wherein providing the set of readiness questions includes providing the set of readiness questions via at least one of a phone and a computer generated display, and wherein the method further comprises: receiving readiness input from at least one individual in response to the set of readiness questions via at least one of a phone and a computer input device;analyzing the readiness input to determine the readiness of the at least one individual to take the selected action; andproviding information to the at least one individual related to the readiness of the at least one individual to take the selected action. 6. The method of claim 1, further comprising: tracking health inputs received from a selected individual;determining one or more health questions in the set of health questions to which the selected individual has not responded; andprompting the selected individual to provide input for the one or more health questions to which the selected individual has not responded. 7. The method of claim 1, further comprising formulating the set of readiness questions ... | - | A61B005/00
US20080306347 | A1 | 2008-12-11 | System and Method for a Telephone Feedback System for Fitness Programs | 2005-11-23 | A diet management system and method thereof is disclosed, which aids in the facilitation of a diet and fitness feedback and tracking. In particular, a communication module, text-to-speech converter, VOX, and server are used in the diet management system. Several communication devices are employed that allows the dieter a broad choice of communicating dietary information. Moreover, the processing and storing of dietary reports are maintained by a network server. | 1. A diet management system for providing feedback to a dieter comprising: a communication module configured to receive dietary information from said dieter; a server configured to store said dietary information and provide dietary reports in response to a dieters request. 2. A diet management system of claim 1, wherein said communication module receives said dietary information in the form of a voice or text message from dieter. 3. A diet management system of claim 2, wherein said the communication module includes a Voice to Text converter for translating voice messages received from the dieter into the stored dietary information. 4. The diet management system of claim 1, wherein the communication module is configured to receive exercise information for a dieter. 5. The diet management system of claim 1, wherein the server is configured to update the dieter regarding the dieter's diet progress. 6. The diet management system of claim 5, wherein the server provides diet updates by sending an electronic message to the dieter 7. The diet management system of claim 5, wherein the server provides diet updates by initiating a phone call with the dieter. 8. The diet management system of claim 4, wherein the server provides instantaneous reports to a dieter based on a combination of the dieter's stored dietary information and stored exercise information. 9. A diet management system of claim 8, wherein the dietary reports is converted to voice through said TTS module and communicated to the dieter by telephone or cellular phone. 10. A diet management system of claim 9, wherein the stored dietary information is accessible by the dieter through a computer network. 11. A diet management system of claim 10, wherein the server comprises a web server that is connected to the Internet. 12. A diet management system of claim 9, wherein the voice is synthesized and morphed for reinforcement and branding of a particular external fitness resource. 13. A diet management system of claim 12, wherein the voice is synthesized and morphed to report into a foreign language. 14. A diet management method for providing feedback to a dieter comprising the steps of: receiving dietary information from the dieter and communicating the dietary information to a server;storing said dietary information and processing dietary reports in response to a dieter's request by the server. 15. A diet management method of claim 14, wherein said method further comprises the step of: converting the report into voice using a Text-to-Speech (TTS) module. 16. A diet management method of claim 15, wherein said method further comprises the step of: synthesizing said converted voice response using a VOX module. 17. A diet management method of claim 16, wherein said method further comprises the step of: playing back the synthesized voice response to said dieter. 18. A diet management method of claim 14, wherein said method further comprises the step of: synthesizing said voice to report... | - | A61B005/00
US20080255429 | A1 | 2008-10-16 | SIMULATED SECONDARY GAIN MODEL WITH CONTINUOUS FEEDBACK TO DIAGNOSTIC TESTING PARAMETERS AND INTERPRETATION IN A MEDICAL LEGAL SETTING | 2007-04-11 | The invention generally relates to compositions and methods for the use of a simulated secondary gain model for use in medical legal setting where there is dispute about the validity of a medical patient's injuries. This model is used to ascertain the false positive rate of a given diagnostic test or suite of same and is then used to reevaluate the test on an ongoing basis after changes have been made to reduce the false positive rate. | 1. A method for minimizing the false positive rate of diagnostic tests in a medical legal environment. 2. The method of claim 1 wherein secondary gain is simulated in a group of subjects without disease by instructing these subjects to feign illness so as to produce a false positive result on a diagnostic test. 3. The method of claim 1 wherein the false positive test data produced by this trial is used to alter test parameters to reduce the prospective false positive rate of future diagnostic tests. 4. The method of claim 1 wherein the false positive test data produced by this trial is used to alter test interpretation to reduce the prospective false positive of future diagnostic tests. 5. The method of claim 1, wherein software or algorithms are used to continuously minimize the false positive rate of diagnostic tests in a medical legal environment. 6. The method of claim 1, wherein a prize or reward is provided to those subjects successfully producing a false positive diagnostic test result. 7. The method of claim 6, wherein the prize or reward provided is deemed by the subjects to be significant enough to produce competitive behavior. 8. The method of claim 7, wherein the prize or reward produces a significant number of false positive diagnostic test results. 9. The method of claim 1 wherein the false positive test data produced by this trial is used to eliminate a certain diagnostic test from being used to settle medical legal disputes over the veracity of injuries or illness. 10. The method of claim 1 wherein the false positive test data produced by this trial is used to improve a certain diagnostic test for the purposes of using the test to settle medical legal disputes over the veracity of injuries or illness. 11. The method of claim 1 wherein the false positive test data produced by this trial is used to eliminate a certain diagnostic test from being used to perform accurate diagnosis in a medial legal setting. 12. The method of claim 1 wherein the false positive test data produced by this trial is used to improve a certain diagnostic test for the purposes of performing accurate diagnosis in a medical legal setting. 13. The method of claim 1 wherein the false positive test data produced by this trial is used to alter test implementation to reduce the prospective false positive of future diagnostic tests. 14. The method of claim 1 wherein the false positive test data produced by this trial is used to alter test machinery to reduce the prospective false positive of future diagnostic tests. 15. The method of claim 1 wherein the false positive test data produced by this trial is used to alter test computer hardware or software to reduce the prospective false positive of future diagnostic tests. | - | A61B005/00
US20080306362 | A1 | 2008-12-11 | DEVICE AND SYSTEM FOR MONITORING CONTENTS OF PERSPIRATION | 2007-06-05 | Described are devices and systems for monitoring a content of perspiration. The device comprises a memory storing an index of values corresponding to amounts of a content in perspiration, a detection arrangement detecting the content in the perspiration, a processing unit determining an amount of the content in the perspiration, and a feedback arrangement providing feedback data indicative of a value corresponding to the determined amount. | 1. A device, comprising: a memory storing an index of values corresponding to amounts of a content in perspiration;a detection arrangement detecting the content in the perspiration;a processing unit determining an amount of the content in the perspiration; anda feedback arrangement providing feedback data indicative of a value corresponding to the determined amount. 2. The device according to claim 1, further comprising: a sweating agent inducing the perspiration when applied to skin. 3. The device according to claim 2, wherein the sweating agent is pilocarpine. 4. The device according to claim 2, further comprising: a housing having first and second faces, the sweating agent being disposed on the first face. 5. The device according to claim 4, wherein the second face includes an engagement arrangement removably engaging at least one of the memory, the processing unit and the feedback arrangement. 6. The device according to claim 1, further comprising: an adhesive coupling the device to skin. 7. The device according to claim 1, wherein the detection arrangement includes electrical contacts delivering a current to the perspiration, the processing unit computing a conductivity of the perspiration as a function of the current. 8. The device according to claim 7, wherein the values correspond to conductivities and hydration levels. 9. The device according to claim 8, wherein the feedback data indicates the hydration level. 10. The device according to claim 1, wherein the detection arrangement includes a sensor. 11. The device according to claim 10, wherein the sensor transmits a signal to the processing unit indicative of the amount of the content. 12. The device according to claim 1, further comprising: a power source. 13. The device according to claim 12, wherein the power source is one of an alkaline battery, a rechargeable battery and a solar cell. 14. The device according to claim 1, wherein the feedback arrangement includes at least one of a liquid crystal display, a light-emitting diode, an audible alarm, a mechanical alarm and a wireless communication arrangement. 15. The device according to claim 14, wherein the feedback data of the wireless communication arrangement is included in a wireless signal transmitted thereby. 16. The device according to claim 1, wherein a predetermined value from the index of values is a threshold value. 17. The device according to claim 16, wherein the feedback arrangement provides further feedback data indicating that the value is one of greater and less than the threshold value. 18. The device according to claim 1, wherein the content is one of a salt, an alcohol and a mineral. 19. A system, comprising: a computing device; anda wearable measuring device measuring an amount of a content of perspiration and transmitting feedback data indicative of the amount to the computing device. 20. The system according to claim 19, wherein the wearable measuring device includes a memory storing an index of... | - | A61B005/00
US20080294012 | A1 | 2008-11-27 | MONITORING PHYSIOLOGICAL CONDITIONS | 2007-05-22 | A method for monitoring physiological conditions of an individual including capturing images according to predefined capture criteria of one or more body regions of the individual on an ongoing basis; transforming the captured images in response to capture conditions present during image capture to provide a normalized image; calculating wellness parameters based on the normalized image; storing system data, including the normalized image, capture parameters, normalization transforms, or wellness parameters, or combinations thereof, monitoring and analyzing the wellness parameters in view of both current and prior semantic data to identify potential changes in physiological conditions; and providing alerts indicative of changes in said physiological conditions of at least one image. | 1. A method for monitoring physiological conditions of an individual comprising; capturing images according to predefined capture criteria of one or more body regions of the individual on an ongoing basis;transforming the captured images in response to capture conditions present during image capture to provide a normalized image;calculating wellness parameters based on the normalized image;storing system data, including the normalized image, capture parameters, normalization transforms, or wellness parameters, or combinations thereof;monitoring and analyzing the wellness parameters in view of both current and prior semantic data to identify potential changes in physiological conditions; andproviding alerts indicative of changes in said physiological conditions of at least one image. 2. The method according to claim 1 wherein the predefined capture criteria for an individual are determined by associated privacy settings, semantic data, the well-being parameters, and the capture parameters, either individually or in combination. 3. The method according to claim 1 wherein the predefined capture criteria include lighting conditions, or the pose, or orientation of an individual. 4. The method according to claim 1 wherein the semantic data includes personal data concerning physical characteristics, physiological conditions data of the individual or family, personal behavior, or activity records. 5. The method according to claim 1 wherein the step of analyzing the wellness parameters includes testing for false positives. 6. The method according to claim 1 wherein semantic data concerning the individuals is utilized in the test for false positives. 7. The method according to claim 1 further including capturing a current image and normalizing such image and further comparing the current normalized image with a previous image to identify potential changes in the individuals physiological condition. 8. The method according to claim 1 wherein the wellness parameters relate specifically to the individual. 9. The method according to claim 1 wherein an individual who is being monitored provides at least a portion of the wellness parameters. 10. The method according to claim 1 wherein the method is applied to more than one individual and one of the individuals is a lead user who has system control and access privileges. 11. The method according to claim 10 where the step of capturing images includes the detection and identification of the individual. 12. The method according to claim 11 wherein the step of capturing images of the individual continues automatically after the individual is identified as a known subject. 13. The method according to claim 11 further including tracking of an individual. 14. The method according to claim 1 wherein the step of analyzing the wellness parameters includes applying algorithms to look for trend-lines or changes in the wellness parameters, or in the normalized images, or both. 15. The method according to claim ... | - | A61B005/00
US20080269633 | A1 | 2008-10-30 | Sonification of Level of Consciousness of a Patient | 2006-12-22 | The level of consciousness of a patient (i.e. the hypnotic state and/or the level of anaesthesia) is represented aurally. A measure of the patient's level of consciousness, such as the BIS index value, is obtained, and an audio signal is synthesized from that measure, and then output through a speaker. Both the volume and pitch of the audio signal may vary according to the BIS value being represented, so that a clinician can obtain an indication of the level of consciousness of the patient simply by listening to the sonification of the BIS value. The audio signal may comprise first and second audio components, with the first component representing the previous measure and the second component representing the current measure. The amplitude or pitch of the audio signal may be modulated to represent analgesia or paralysis of the patient. | 1. A method of indicating level of consciousness of a patient aurally, comprising the steps of: (a) obtaining a measure representative of the patient's level of consciousness,(b) synthesizing an audio signal from the measure,(c) outputting the audio signal,wherein at least one audible characteristic of the audio signal is dependent upon the patient's level of consciousness; and(d) repeating steps (a) to (c) at intervals of time. 2. A method as claimed in claim 1, wherein the measure representative of the patient's level of consciousness is a BIS value. 3. A method as claimed in claim 2, wherein the pitch of the audio signal is dependent on the BIS value. 4. A method as claimed in claim 2, wherein the amplitude of the audio signal is dependent on the BIS value. 5. A method as claimed in claim 4 wherein the amplitude of the audio signal is low if the BIS value is within a predetermined range, but audibly increases according to the extent to which the BIS value is outside the predetermined range. 6. A method as claimed in claim 1, wherein the audio signal comprises first and second audio components spaced in time. 7. A method as claimed in claim 6, wherein the second audio component is an audio signal synthesized from the current measure, and the first audio component is an audio signal synthesized from the previous measure. 8. A method as claimed in claim 1, further including the steps of: (a) detecting level of paralysis of the patient, and(b) modifying the audio signal according to the detected level of paralysis,whereby at least one audible characteristic of the modified audio signal is dependent upon the patient's paralysis. 9. A method as claimed in claim 1, further including the steps of: (a) detecting level of analgesia of the patient, and(b) modifying the audio signal according to the detected level of analgesia,wherein at least one audible characteristic of the modified audio signal is dependent upon the patient's analgesia. 10. A method as claimed in claim 8, when wherein the step of modifying the audio signal comprises modulation of the amplitude and/or frequency of the audio signal. 11. A method as claimed in claim 1 wherein the audio signal is output as one of a plurality of audio signals representing other physiological parameters of the patient, and wherein the audio signal is aurally distinguishable from the other one(s) of the plurality of audio signals. 12. A method as claimed in claim 1 wherein the audio signal is output via an earpiece. 13. A method of representing level of consciousness of a patient aurally, comprising the steps of: (a) obtaining measures representative of the patient's level of consciousness at spaced intervals of time,(b) synthesizing an audio signal having at least first and second audio components, the second audio component being dependent on a current measure, and the first audio component being dependent on the next previous measure, whereby at least one audible characteristic of each compone... | THE UNIVERSITY OF QUEENSLAND | A61B005/0476
US20080281181 | A1 | 2008-11-13 | Combination of Multi-Modality Imaging Technologies | 2005-05-13 | Different applications for multiple imaging technologies are provided to advance patient care. In one aspect, a 3-D imaging device is used to identify the location of target region in a patient, such as a tumor or aneurysm, for instance. The location information is then used to control a device used in performing a surgical or other intervention, further imaging or a diagnostic or therapeutic procedure (920). The device can operate fully automatically, as a robot, or can assist a manual procedure performed by a physician. In another aspect, the invention provides a technique for obtaining an improved image of the vasculature in a patient (1020). In another aspect, images from multiple imaging technologies are combined or fused (1120) to achieve synergistic benefits. | 1. A method for positioning a device for accessing a target region in a patient, comprising: determining (910) at least one of a particular path and a particular viewpoint for accessing the target region based on at least one 3-D image of the target region obtaining by imaging the target region; andcontrolling (920) the device for accessing the target region according to the at least one of a particular path and a particular viewpoint. 2. The method of claim 1, wherein: the device for accessing the target region comprises a radiographic imaging system. 3. The method of claim 1, wherein: the at least one 3-D image of the target region is obtained from a 3-D imaging system employing at least one of computed tomography (CT), magnetic resonance imaging (MRI), positron emission tomography (PET), and ultrasound. 4. The method of claim 1, wherein: the at least one 3-D image of the target region comprises a fusion of different images from different 3-D imaging systems employing at least one of computed tomography (CT), magnetic resonance imaging (MRI), positron emission tomography (PET), and ultrasound. 5. The method of claim 1, wherein: the at least one of a particular path and a particular viewpoint for accessing the target region is determined manually by a physician by viewing the at least one 3-D image. 6. The method of claim 1, wherein: the at least one of a particular path and a particular viewpoint for accessing the target region is determined automatically using software that evaluates the at least one 3-D image. 7. The method of claim 1, wherein: the device for accessing the target region comprises a robot. 8. The method of claim 1, wherein: the device for accessing the target region is used in performing at least one of: (a) a surgical or other intervention, (b) further imaging, and (c) a diagnostic and/or therapeutic procedure; andthe device operates automatically, as a robot. 9. The method of claim 1, wherein: the device for accessing the target region is used in performing at least one of: (a) a surgical or other intervention, (b) further imaging, and (c) a diagnostic and/or therapeutic procedure; andthe device assists a manual procedure performed by a physician. 10. The method of claim 1, further comprising: arranging in a controlled relationship to one another, (a) an imaging device for imaging the target region, (b) the device for accessing the target region, and (c) a patient support. 11. The method of claim 1, further comprising: imaging the target region using a CT system to obtain data for reducing the number of angiograms that are needed to visualize an aneurysm. 12. The method of claim 1, wherein the target region is in the head or brain of the patient, further comprising: delivering a contrast agent in the aorta rather than in arteries in the head or brain prior to imaging the target region. 13. The method of claim 1, wherein the target region is anywhere in the body of the patient, further comprising: delivering a con... | THE RESEARCH FOUNDATION OF STATE UNIVERSITY OF NEW YORK | A61B005/05 A61B006/00
US20080319329 | A1 | 2008-12-25 | Manual-driven inflation-powered electronic blood pressure measuring apparatus | 2007-06-25 | A manual-driven inflation-powered electronic blood pressure measuring apparatus includes an inflation unit to generate pressurized air by a user through hand squeezing, a power generation unit to receive the pressurized air to generate and output electric power, an electronic blood pressure measurement unit to receive the electric power output from the power generation unit and the pressurized air from the inflation unit to inflate an air pouch, and a sensor to detect a person's blood pressure signals in cooperation with the inflation pouch, and outputs a measurement result through an output unit. | 1. A manual-driven inflation-powered electronic blood pressure measuring apparatus, comprising: an inflation unit to generate pressurized air by a user through hand squeezing;a power generation unit to receive the pressurized air from the inflation unit and generate and output electric power; andan electronic blood pressure measurement unit to receive the electric power output from the power generation unit and inflate an air pouch with the pressurized air generated by the inflation unit, and detect a person's blood pressure signal through the air pouch and a sensor, and output a measurement result through an output unit. 2. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the power generation unit has a generator to receive the pressurized air to generate AC power signals. 3. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the power generation unit has a power regulator to regulate the electric power generated by the power generation unit. 4. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 3, wherein the power regulator is a rectifier. 5. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 3, wherein the power regulator is voltage stabilizer. 6. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the power generation unit has a power storage to store the electric power generated by the power generation unit. 7. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the air pouch has an air release path. 8. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the inflation unit has a check valve. 9. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the output unit is a display device. 10. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the output unit is a lighting device which turns on and off light to output the measurement result. 11. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the output unit is a lighting device which generates blinking light to output the measurement result. 12. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the output unit is a sound output device which broadcasts voice to output the measurement result. 13. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the output unit is a sound output device which makes audio broadcast to output the measurement result. 14. The manual-driven inflation-powered electronic blood pressure measuring apparatus of claim 1, wherein the pressurized air received by the electronic blood pr... | K-JUMP HEALTH CO., LTD. | A61B005/02 A61M005/178
US20080243010 | A1 | 2008-10-02 | SYSTEM AND METHOD FOR FACILITATING PROPER CUFF USE DURING NON-INVASIVE BLOOD PRESSURE MEASUREMENT | 2007-03-30 | A non-invasive blood pressure system comprises a blood pressure cuff having an identifier and an aperture such that the identifier can be read through the aperture when the cuff is in a desired configuration. Another comprises a processor in operative communication with the cuff and configured to initiate a blood pressure determination when the identifier can be read through the aperture, while a method provides such blood pressure cuffs. | 1. A non-invasive blood pressure system, comprising: a blood pressure cuff having an identifier and a non-insertion aperture such that said identifier can be read through said aperture when said cuff is in a desired configuration. 2. The system of claim 1, wherein said identifier comprises one or more bar codes. 3. The system of claim 1, wherein said identifier comprises one or more radio-frequency identification tags. 4. The system of claim 3, wherein said cuff comprises radio-frequency blocking material except about said aperture. 5. The system of claim 1, wherein said identifier can be read through said aperture only when said cuff is in said desired configuration. 6. The system of claim 1, wherein said desired configuration comprises a desired degree of wrap of said cuff. 7. The system of claim 1, wherein said desired configuration comprises a desired sizing of said cuff for a particular patient. 8. The system of claim 1, wherein said is a visual aperture. 9. The system of claim 1, wherein said aperture is a line-of-sight aperture. 10. The system of claim 1, wherein at least one of said identifier and said aperture are provided on a surface of said cuff. 11. A non-invasive blood pressure system, comprising: a blood pressure cuff having an identifier a non-insertion aperture such that said identifier can be read through said aperture when said cuff is in a desired configuration; anda processor in operative communication with said cuff and configured to initiate a blood pressure determination when said identifier is read through said aperture. 12. The system of claim 11, wherein said identifier comprises one or more bar codes. 13. The system of claim 11, wherein said identifier comprises one or more radio-frequency identification tags. 14. The system of claim 13, wherein said cuff comprises radio-frequency blocking material except about said aperture. 15. The system of claim 11, wherein said identifier can be read through said aperture only when said cuff is in said desired configuration. 16. The system of claim 11, wherein said processor is configured to automatically initiate said determination when said identifier is read through said aperture. 17. A method of determining blood pressure, comprising: providing a blood pressure cuff having an identifier and a non-insertion aperture such that said identifier can be read through said aperture when said cuff is in a desired configuration. 18. The method of claim 17, wherein said identifier comprises one or more bar codes. 19. The method of claim 17, wherein said identifier comprises one or more radio-frequency identification tags. 20. A method of determining blood pressure, comprising: providing a blood pressure cuff having an identifier and a non-insertion aperture such that said identifier can be read through said aperture when said cuff is in a desired configuration: andinitiating a blood pressure determination only when said identifier is read through said aperture. | GENERAL ELECTRIC COMPANY | A61B005/022
US07438687 | B2 | 2008-10-21 | Patient monitoring system with blood pressure measurement capacity | 2005-08-12 | A patient monitoring system including a sensing-recording device that attaches to a patient and a probe utilized by attending personnel at a trauma site. After an initial interaction between the probe and the sensing-recording device, the sensing-recording device samples one or more physiological parameters and stores time stamped readings in a first memory. The probe is adapted to take other measurements including pulse data, download the data from the first memory of the sensing-recording device, manipulate the collective data and return data to a second memory in the sensing-recording device. The probe is also adapted to combine data from the sensing-recording device and its pulse data to provide a blood pressure measurement. | 1. A system for obtaining a measurement of a patient's blood pressure comprising: A) sensing-recording means for attachment to the patient including a memory, sensing means for measuring a predetermined physiological parameter, and means for storing time-stamped samples of the predetermined physiological parameter in said memory as a first data set, andB) a probe for being applied with an application pressure to a measurement site remote from the site of said sensing-recording means, said probe including: i) means for generating a signal representing a pulse wave form at the measurement site,ii) a probe memory,iii) means for storing times-stamped samples of the signal in said probe memory as a second data set,iv) means for retrieving the time-stamped samples from said memory in said sensing-recording means,v) means for correlating the measurement times for the first and second data sets,vi) means for generating a representation of the patient's blood pressure based upon the information in said correlated, time-stamped data sets. 2. A system as recited in claim 1 wherein said probe comprises: i) rigid housing means for being brought into proximity to an artery at the measurement site,ii) pressure transducer means mounted to one side of said housing means for generating an electrical signal that is a function of pressure at an input thereto,iii) flexible membrane means mounted to the other side of said housing for forming a sealed, fluid-filled cavity about said passage means at said housing means, andiv) passage means formed through said housing means for connection to said pressure transducer means input whereby when said flexible membrane means is brought into contact with the skin at the measurement site, said pressure transducer means produces a representation of a wave-form signal corresponding to the pressure changes at the measurement site. 3. A system as recited in claim 2 wherein said flexible membrane means includes a thin, flexible elastomer formed into a closed-end channel and a peripheral flange attachment to said other side of said rigid housing means. 4. A system as recited in claim 3 wherein said rigid housing means forms a chamber for receiving said pressure transducer means, said apparatus including a controlled leak means for long-term pressure equalization between said cavity and said chamber. 5. A system as recited in claim 1 wherein said sensing means in said sensing-recording means includes means for measuring pulse oximetry whereby the first data set stored in said sensing-recording means memory represents blood oxygenation. 6. A system as recited in claim 5 wherein during a measurement said probe is applied to the measurement site with an application pressure and said pressure transducer means generates a composite signal representing the application pressure and the pressure at the measurement site, said probe including processing means for storing data representing the pulse wave form produced at the measurement sit... | Nova Technology Corporation | A61B005/02
US07460896 | B2 | 2008-12-02 | In vivo device and method for collecting oximetry data | 2004-07-29 | An in vivo device and a method for collecting and measuring oximetry data in a body lumen, such as using a swallowable capsule. The in vivo device may include instruments for collecting oximetry data, for example an illuminator and an oximetry data detector. | 1. An autonomous in-vivo device comprising a unit to measure oxygenation of hemoglobin in a blood stream, said device when inside a body not including any physical connection to an additional device, the unit comprising: at least two cavities limiting backscatter or sidescatter light access, wherein a first of said at least two cavities includes either an illumination source or a detector recessed therein and a second of said at least two cavities includes the other of said illumination source or said detector recessed therein,said unit arranged such that light emitted by the illumination source is detectable by the detector when said device is inside a body. 2. The device according to claim 1, comprising a filter to filter sidescatter light. 3. The device according to claim 1, comprising an oximetry data detector. 4. The device according to claim 1, further comprising an optical dome behind which are situated an imager, an optical element and an illumination device to illuminate in vivo sites. 5. The device according to claim 1,. comprising a wireless transmitter. 6. The device according to claim 1, comprising an instrument to collect oximetry data. 7. The device according to claim 1, wherein said cavity is recessed into one side of the device. 8. The device according to claim 1, wherein the illuminator and detector are arranged for obtaining reflectance measurements. 9. The device according to claim 1, wherein said at least one cavity comprises a non-reflective surface to absorb backscatter or sidescatter light. 10. The device according to claim 1, wherein said in-vivo device is a capsule. 11. A system for measuring the oxygenation of hemoglobin in a blood stream, the system comprising: an autonomous in vivo device that, when inside a body does not include any physical connection to an additional device, comprisingan instrument to collect oximetry data;at least two cavities limiting backscatter or sidescatter light access, wherein a first of said at least two cavities includes either an illumination source or a detector recessed therein and a second of said at least two cavities includes the other of said illumination source or said detector recessed therein,said unit arranged such that light emitted by the illumination source is detectable by the detector when said device is inside a body; andan external receiving unit to receive transmitted information. 12. The system according to claim 11, wherein the receiving unit includes at least a component selected from the group consisting of: a recorder unit, a storage unit and a workstation. 13. The system according to claim 11, comprising a controller to calculate oximetric readings from diffused light received by the in-vivo device. 14. A method of in vivo measuring of oximetry, the method comprising: inserting an autonomous in vivo device into a body lumen, said device when inside the body lumen not including any physical connection to an additional device;illuminating tissue of said... | Given Imaging Ltd. | A61B005/1455
US07442401 | B2 | 2008-10-28 | Method, set, and apparatus for obtaining prints of a part of the human body | 2005-01-18 | Apart of a human body is caused to contact the surface layer of a recording medium over which a plurality of metallic grains with an outside size of 200 nm or less are distributed. Then, secretions from the skin surface of the body part are caused to adhere to the surface layer of the recording medium to take the print of the body part. If light is irradiated to the recording medium, specific optical characteristics resulting from the surface structure of the recording medium are obtained, and therefore the color of the recording medium varies between a region having secretions and a region having no secretions. This renders it possible to record a visible print on the recording medium. | 1. A method for obtaining prints of the shape and/or patterns of at least a part of a human body, comprising the steps of: causing said part to contact a surface layer of a recording medium over which a plurality of metallic grains with an outside size of 200 nm or less are distributed; andcausing secretions from a skin surface of said part to adhere to the surface layer of said recording medium to take said print. 2. The method as set forth in claim 1, wherein said surface layer of said recording medium is a layer in which said metallic grains are provided in alumina minute holes obtained by anodizing a material that contains aluminum as its main component. 3. The method as set forth in claim 1, wherein said surface layer is coated with a layer having a refractive index different from that of said secretions, after said secretions adhere to said surface layer of said recording medium. 4. The method as set forth in claim 1, wherein said recording medium is optically read after said secretion adheres to said surface layer of said recording medium,a digital image representing an adhesion status of said secretions is generated, andsaid digital data is stored on a predetermined storage medium. | FUJIFILM Corporation | A61B005/117
US07447537 | B1 | 2008-11-04 | System and method for augmentation of endoscopic surgery | 2000-06-16 | Image processing is used to determine information about the position of a designated object, particularly where the object is difficult to view or locate. In particular, use in endoscopic surgery can determine positional information about an anatomical feature within a patient's body. The positional information is then used to position or reposition a surgical instrument in relation to the designated object. An instrument placed in relation to the designated object is capable of sending information about the object to a computer. Image processing generates images and determines positional information, which can be used as input to robotic devices or rendered to a human user. | 1. A method of controlling the position of a surgical instrument inside a patient's body the comprising steps of: inserting a first surgical instrument into the patient's body, the instrument having a means for transmitting an image out of the patient's body;designating an anatomical feature of interest;transmitting an image of the designated anatomical feature out of the patient's body;determining precise 3D positional information about the designated anatomical feature of interest relative to the first surgical instrument;using the positional information to reposition the first surgical instrument to a desired positional relationship relative to the anatomical feature, where the positional information about the anatomical feature of interest relative to the first surgical instrument is obtained by manipulating a graphics object superimposed on an image of the anatomical feature. 2. A method of controlling the position of a surgical instrument inside a patient's body, as in claim 1, where the positional information about the feature of interest relative to the first surgical instrument is obtained by means of image processing. 3. A method of controlling the position of a surgical instrument inside a patient's body, as in claim 1, where the positional information is used to reposition a second surgical instrument. 4. A method of controlling the position of a surgical instrument inside a patient's body, as in claim 1, further comprising steps of: designating a desired position of an anatomical feature relative to images transmitted out of the patient's body by the first surgical instrument;moving the first surgical instrument to a vantage point form which the designated anatomical feature is at the desired position in images transmitted by the first surgical instrument, the first surgical instrument remaining at a constant distance from the designated anatomical feature. 5. A method of controlling the position of a surgical instrument inside a patient's body, as in claim 1, further comprising steps of: designating a desired direction of motion of an anatomical feature relative to images transmitted out of the patient's body by the first surgical instrument;moving the first surgical instrument so that it remains at a constant distance from the designated anatomical feature while causing the motion of the designated feature in images transmitted from the first surgical instrument to move in the desired direction. 6. A method of controlling the position of a surgical instrument inside a patient's body, comprising steps of: inserting a first surgical instrument into the patient's body, the instrument having a means for transmitting an image out of the patient's body;designating an increment of motion along an axis of view of the first surgical instrument, the axis defined by a line from the vantage point of the surgical instrument to the point on the patient's anatomy that appears in the center of the image transmitted by the first surgical instr... | International Business Machines Corporation | A61B005/05
US07762953 | B2 | 2010-07-27 | Systems and methods for non-invasive physiological monitoring of non-human animals | 2007-10-31 | This invention provides monitoring garments for non-invasively monitoring physiological parameters in un-restrained and/or restrained animals, such as monkeys, rabbits, dogs, horses, and the like. The invention also includes methods and systems for collecting and processing monitoring data and methods for recognizing apneas and other respiratory events, periods of restfulness and wakefulness, stereotypical behavior and other indicators of dysphoric states, periods of emesis, and occurrence of barking and coughing. | 1. A method of identifying emesis in an animal comprising the steps of: receiving monitoring data from the animal, said monitoring data comprising signals reflecting abdominal (AB) dimensions and signals reflecting rib cage (RC) dimensions, said AB and RC signals varying with respiratory motions;determining a characteristic pattern in said AB and RC signals using a processor, said characteristic pattern comprisingan increase in an average AB dimension from a baseline AB dimension and a decrease in an average RC dimension from a baseline RC dimension, the increase and decrease occurring abruptly within a few breaths and persisting for several breaths,a decrease in said average AB dimension towards said baseline AB dimension and an increase in said average RC dimension towards said baseline RC dimension, said RC dimension increase and said AB dimension decrease occurring within five (5) breaths, andone or more synchronous breath induced changes in AB and RC dimensions, the amplitude RC dimension change being greater than a baseline RC dimension; andidentifying a period of emesis when said characteristic pattern is determined. 2. The method of claim 1, wherein said increase in average AB dimension and said decrease in average RC dimension occurs within five (5) breaths, and wherein said RC and AB dimension changes persist for at least ten (10) breaths. 3. A method of identifying cough in an animal comprising the steps of: receiving monitoring data from the animal, said monitoring data comprising signals reflecting abdominal (AB) dimensions and signals reflecting rib cage (RC) dimensions, said AB and RC signals varying with respiratory motions, and signals reflecting vocalizations;determining a characteristic pattern in said AB, RC and vocalization signals using a processor, said characteristic pattern comprisingan inspiratory period with a steady inspiration rate and lasting for a first period of time,an expiratory period with a steady expiratory rate and lasting for a second period of time, wherein said inspiratory first period of time is greater than said expiratory second period of time and said inspiratory rate is lower than said expiratory rate,determining a ratio of said inspiratory time to said expiratory time;a vocalization occurring during expiration; andidentifying a cough event when said characteristic pattern is determined. 4. The method of claim 3 wherein said ratio of said inspiratory time to said expiratory time is at least 2. 5. A method of identifying bark in a barking animal comprising the steps of: receiving monitoring data from the animal, said monitoring data comprising signals reflecting abdominal (AB) dimensions and signals reflecting rib cage (RC) dimensions, said AB and RC signals varying with respiratory motions, and signals reflecting vocalizations;determining a characteristic pattern in said AB, RC and vocalization signals using a processor, said characteristic pattern comprisingan inspiratory period with a steady ins... | Adidas AG | A61B005/00 A61B005/08
US07457660 | B2 | 2008-11-25 | Eliminating interface artifact errors in bioimpedance measurements | 2003-11-28 | A system for diagnosing the possibility of disease in one of a first body part and a second substantially similar body part by impedance measurements is described. The system includes an impedance module for calculating impedances of corresponding segments of the first and second body parts from current and voltage signals. The measured impedances are corrected to account for the effect of stray impedances arising from non-body part sources. | 1. A system for diagnosing the possibility of disease in one of a first body part and a second substantially similar body part by impedance measurements, the system comprising: a first body part module for injecting a first current into the first body part and for receiving a corresponding first voltage signal;a second body part module for injecting a second current into the second body part and for receiving a corresponding second voltage signal;a correction module for obtaining a first correction factor for the first body part and a second correction factor for the second body part, the first and second correction factors accounting for impedances inherent in non-body part sources, the correction module including a magnitude correction module for calculating first and second magnitude correction factors, and a phase correction module for calculating first and second phase correction factors, where the first correction factor is composed of the first magnitude correction factor and the first phase correction factor, and where the second correction factor is composed of the second magnitude correction factor and the second phase correction factor; andan impedance module for calculating a first impedance from the first current, the first voltage signal and the first correction factor, and for calculating a second impedance from the second current, the second voltage signal and the second correction factor,wherein the first and second impedances are used to diagnose the possibility of disease. 2. The system of claim 1, wherein the non-body part sources that give rise to inherent impedances include skin covering the first and second body parts and system components. 3. The system of claim 1, wherein the correction module includes a magnitude correction table to calculate the first and second magnitude correction factors, the magnitude correction table containing calibration impedance magnitude (¦Zcal¦) data and associated magnitude correction factor (Cmag(¦Zcal¦)) data. 4. The system of claim 3, further comprising a calibration apparatus to form the magnitude correction table. 5. The system of claim 4, wherein the calibration apparatus includes an electrical model of the first body part, a variable interface resistance and the impedance module. 6. The system of claim 5, wherein the impedance module includes a bipolar unit for calculating a bipolar impedance, ZBP, from a bipolar voltage measurement made by the first body part module on the first body part, a magnitude of the bipolar impedance used by the correction module to obtain the first correction factor. 7. The system of claim 6, wherein the correction module uses the magnitude correction table and the magnitude of the bipolar impedance to obtain the first magnitude correction factor, which is given by Cmag(¦ZBP¦). 8. The system of claim 7, wherein an uncorrected first impedance, Zraw, is calculated by the impedance module from the first current and the first voltage signal. 9. The syst... | Z-Tech (Canada) Inc. | A61B005/00
US07462156 | B2 | 2008-12-09 | Replacement aortic valve leaflets and related technology | 2005-07-29 | Replacement aortic heart leaflets and method of making the same. The replacement aortic heart leaflets of the present invention include first and second segments. The second segment is integral with, and extends from, the first segment. A virtual line L having a predetermined length corresponding to the distance between two known aortic commissures separates the first and second segments. In another aspect of the invention, the first segment has a virtual line AM of predetermined length. In another aspect of the present invention, a leaflet-measuring device is described. In another aspect of the invention, a standard set of replacement aortic leaflets is described. Methods are provided to provide replacement aortic leaflets based on traced outlines obtained from a surgically opened aortic ring. | 1. A measuring device for measuring inter-commissural dimension L and the length parameter AM, said device comprises: a generally T-shaped member, wherein said T-shaped member comprises: an L distance measuring member, wherein said L distance measuring member has first and second opposite ends and a middle section between said first and second opposite ends, wherein said L distance measuring member is less than about 40 mm in length, and wherein numeric measurement indicia are located at said first and second opposite ends of said L distance measuring member; andan AM measuring member, wherein said AM measuring member has first and second opposite ends, wherein said first end of said AM measuring member is connected to said middle section of said L distance measuring member thereby defining said T-shaped member, wherein said AM distance measuring member is less than about 40 mm in length and wherein further numeric measurement indicia is located at said second end of said AM measuring member. 2. The measuring device according to claim 1, wherein said L distance measuring member is less than about 30 mm in length. 3. The measuring device according to claim 1, wherein said L distance measuring member is about 25 mm in length. 4. The measuring device according to claim 1, wherein said AM measuring member is less than about 30 mm in length. 5. The measuring device according to claim 1, wherein said AM measuring member is about 25 mm in length. 6. A measuring device for measuring inter-commissural dimension L and the length parameter AM, said device comprising: a T-shaped member having a hollow vertical bore and a horizontal arm, wherein said horizontal arm is used for measuring inter-commissural dimension L;an AM measuring shaft for measuring the AM length parameter, said measuring shaft is housed in said bore, and wherein said measuring shaft has a lower end; andmeasurement indicia displayed on said shaft for measuring said AM parameter. 7. The device according to claim 6, wherein said shaft defines a lower end, wherein said indicia are displayed on said lower end of said shaft. 8. The device according to claim 6, further comprising a fixator member. 9. The device according to claim 6, wherein said shaft is movable within said bore, and further wherein said fixator holds said shaft in place inside said bore. | - | A61B005/103 A61B001/00 A61F002/24
US07850619 | B2 | 2010-12-14 | Apparatus and method for breathing pattern determination using a non-contact microphone | 2005-07-21 | A method is provided for analyzing respiration of a subject. Using a non-contact microphone, a raw signal indicative of air-flow sounds of the respiration is generated. The raw signal is analyzed to determine a first set of one or more parameters of the respiration. An algorithm is applied to the first set of parameters to derive a second set of one or more estimated parameters of the respiration that are not generally directly measurable in the raw signal. Other embodiments are also described. | 1. Apparatus for analyzing respiration of a subject, comprising: a non-contact microphone that does not contact the subject, adapted to generate a raw signal representing airflow sounds that are generated at the microphone by airflow of the respiration; anda control unit, adapted to: analyze the airflow sounds, by deriving an expiratory airflow sound signal from the raw signal, and analyzing the expiratory airflow sound signal, to determine a first set of one or more parameters of the respiration, andapply an algorithm to the first set of parameters to derive a second set of one or more estimated parameters of the respiration, including an inspiration time of the subject. 2. The apparatus according to claim 1, wherein the control unit is adapted to set the inspiration time equal to a phenomenological function of the first set of one or more parameters. 3. The apparatus according to claim 1, wherein the first set of parameters includes a measure of breathing amplitude of the respiration, and wherein the control unit is adapted to analyze the airflow sounds by integrating the airflow sounds for a breath of the respiration to determine the measure of the breathing amplitude. 4. The apparatus according to claim 1, wherein the control unit is adapted to analyze a parameter selected from the group consisting of: the first and second sets of parameters, to derive at least one additional breathing-related parameter of the subject selected from the group consisting of: breathing amplitude, a geometrical property of airflow of the subject, a characteristic of the airflow indicative of pursed lips breathing, a characteristic of the breathing indicative of relaxed breathing, a characteristic of the breathing indicative of passive elastic recoil of lungs of the subject, a characteristic of breathing with effort, and a characteristic of breathing during which the lungs of the subject undergo a functional change. 5. The apparatus according to claim 1, comprising an output generator, wherein the control unit is adapted to drive the output generator to guide the subject to perform breathing in a plurality of respiration phases determined at least in part responsively to the inspiration time. 6. The apparatus according to claim 1, wherein the non-contact microphone comprises a non-contact microphone of a consumer electronics device capable of performing at least one function that does not facilitate analyzing respiration of the subject. 7. The apparatus according to claim 1, wherein the non-contact microphone is integrated into a medical device in fluid communication with respiration-related airflow of the subject. 8. The apparatus according to claim 7, wherein the medical device comprises a drug inhaler. 9. The apparatus according to claim 1, comprising an output generator, wherein the control unit is adapted to drive the output generator to generate a real-time indication for the subject that indicates whether expiration has been detected. 10. The appar... | Intercure Ltd. | A61B005/08
US07455644 | B2 | 2008-11-25 | Carbon dioxide sensor and airway adapter incorporated in the same | 2004-02-18 | In a sensor for detecting a carbon dioxide gas in an expiration gas of a living body, an airway case is adapted to be disposed below nostrils of the living body, and formed with an airway passage extending across an optical axis of a light beam emitted from a right emitter of the sensor. A mouth guide is adapted to be disposed in front of a mouth of the living body so as to define a space communicated with the airway passage. The mouth guide is pivotably supported on the airway case. A retainer is adapted to retain an oxygen supply tube on the airway adapter body in such an attitude that an oxygen gas supplied from prongs of the oxygen supply tube is not directly injected into the nostrils. | 1. An airway adapter adapted to be attached to a light emitter of a sensor for detecting a carbon dioxide gas in an expiration gas of a living body, the airway adapter comprising: an airway case, adapted to be disposed below nostrils of the living body, and formed with an airway passage configured to extend across an optical axis of a light beam emitted from the light emitter of the sensor at a position between the nostrils and a mouth of the living body; anda mouth guide, adapted to be disposed in front of the mouth of the living body so as to define a space communicated with the airway passage, the mouth guide being supported on the airway case so as to be pivotable about an axis that extends substantially parallel to the optical axis. 2. The airway adapter as set forth in claim 1, wherein a shaft member is integrally molded with the mouth guide, and fitted into a hole formed in the airway case, so that the mouth guide is pivoted about the hole. 3. The airway adapter as set forth in claim 2, wherein the shaft member is formed with a flexible material so as to have a size which is no less than the size of the hole. 4. The airway adapter as set forth in claim 2, wherein at least one of the airway case and the mouth guide is formed with an elastic material, so as to generate an elastic force directed in an axial direction of the shaft member. 5. The airway adapter as set forth in claim 2, wherein the shaft member is extending in a first direction substantially parallel with a plane represented by the face of the living body, and the mouth guide is pivotable about the shaft member in a second direction perpendicular to the first direction. 6. The airway adapter as set forth in claim 1, further comprising an inlet member, adapted to be inserted into at least one of the nostrils having a passage for guiding a nasal expiration gas to the airway passage, the inlet member being formed with a vent hole communicating the passage and an exterior of the inlet member. 7. The airway adapter as set forth in claim 6, wherein: the passage of the inlet member is defined by a pair of tube members adapted to be inserted into the nostrils and a junction at which the tube members are merged; andthe vent hole is formed at the junction. 8. The airway adapter as set forth in claim 7, wherein the vent hole is arranged such that a flow of a gas discharged from the vent hole is not substantially interfered by the living body. 9. The airway adapter as set forth in claim 8, wherein the vent hole is arranged so as not to oppose to a face of the living body. 10. A sensor for detecting a carbon dioxide gas in an expiration gas of a living body, comprising: a photo emitter;a photo receiver; andan airway adapter, which supports the photo emitter and the photo receiver such that a light beam emitted from the photo emitter is received by the photo receiver, the airway adapter comprising:an airway case, adapted to be disposed below nostrils of the living body, and formed wi... | Nihon Kohden Corporation | A61B005/08
US07959567 | B2 | 2011-06-14 | Device to enable quick entry of caloric content | 2007-10-22 | A monitoring apparatus that includes a sensor device and an I/O device in communication with the sensor device that generates derived data using the data from the sensor device. The derived data cannot be directly detected by the associated sensors. Alternatively, an apparatus that includes a wearable sensor device and an I/O device in communication with the sensor device that includes means for displaying information and a dial for entering information. Alternatively, an apparatus for tracking caloric consumption and caloric expenditure data that includes a sensor device and an I/O device in communication with the sensor device. The sensor device includes a processor programmed to generate data relating to caloric expenditure from sensor data. Alternatively, an apparatus for tracking caloric information for an individual that utilizes a plurality of classification identifiers for classifying meals consumed by the individual, each of the classification identifiers having a corresponding caloric amount. | 1. An apparatus for tracking caloric consumption and caloric expenditure of an individual, comprising: an I/O device adapted to enable a user to enter information relating to foods consumed by said individual, said information comprising a user's estimate of the relative size of a meal in relation to said user's normal-sized meal; anda processor in communication with said I/O device, said processor for receiving said user's estimate and programmed to generate data corresponding to the caloric content of said meal based on at least one of (a) aggregated data of the caloric content associated with other users' estimates of the relative size of meals in relation to said other users' normal-sized meals, and (b) a previous estimate of said user of the relative size of a previous meal in relation to said user's normal-sized meals. 2. The apparatus of claim 1, wherein processor is further programmed to cause said I/O device to provide the user with an interface enabling said user to enter said information relating to foods consumed by said individual. 3. The apparatus of claim 2, wherein said interface provides the user with a list of selectable identifiers corresponding to user's estimate. 4. The apparatus of claim 3, wherein said selectable identifiers comprise small, medium, large, and extra large. 5. The apparatus of claim 3, wherein said selectable identifiers relate to the user's estimate of the relative density of a meal in relation to said user's normal meal. 6. The apparatus of claim 3, wherein said selectable identifiers comprise a point value corresponding to said user's estimate. 7. The apparatus of claim 1, wherein said meal comprises a snack. 8. The apparatus of claim 1, wherein said processor is programmed to adjust data corresponding to the caloric content of said meal based on said user's weight and information relating to calories expended by said individual. | BodyMedia, Inc. | A61B005/00
US08038628 | B2 | 2011-10-18 | Torque device for a sensor guide wire | 2007-05-24 | The invention relates to a torque device for a sensor guide wire having a sensor provided at a distal portion and a male connector provided at the proximal end, which torque device comprises a grip body and a cap adapted to be joined to the grip body, and a number of chuck segments provided on the cap or the grip body, wherein the torque device is a one-way device defining an insertion direction for the sensor guide wire and wherein the chuck segments have free ends which are directed in the insertion direction. | 1. A torque device having a lumen adapted for reception of a portion of a sensor guide wire, comprising: a grip body;a cap adapted to be joined to the grip body;a number of chuck segments provided on the cap or the grip body;wherein the torque device is a one-way device defining an intended insertion direction for the sensor guide wire and wherein the chuck segments have free ends which are directed in the insertion direction; anda mechanism configured to allow introduction of the sensor guide wire in the intended insertion direction and prevent introduction of the sensor guide wire in an opposite direction;wherein a proximal end of the cap is inserted into a distal end of the grip body. 2. A torque device according to claim 1, wherein the joining of the grip body and the cap is accomplished by a threaded portion which is provided on the cap and which matches with a threaded portion provided on the grip body. 3. A torque device according to claim 1, wherein the chuck segments are provided on the cap and wherein each of the chuck segments is provided with a shoulder that matches with a transition step provided inside the grip body. 4. A torque device according to claim 1, wherein the torque device has a lumen and comprises a pair of flexible lips, which extend into the lumen of the torque device and whose tips are directed essentially in the intended insertion direction, to allow introduction of the sensor guide wire in the intended insertion direction and prevent introduction of the sensor guide wire in the opposite direction. 5. A torque device according to claim 1, wherein the torque device is a separate device which is movable along the length of the sensor guide wire and which can be clamped and attached at an arbitrary position along the sensor guide wire. 6. A torque device according to claim 1, wherein the torque device constitutes a female connector adapted to accommodate a male connector provided at the proximal end of the sensor guide wire. 7. A torque device according to claim 1, wherein a weakened section adjoins to the chuck segments, wherein the weakened section has a smaller outer diameter than portions adjacent the weakened section in distal and proximal directions of the torque device. 8. A torque device according to claim 1, wherein a proximal end of the grip body includes a hole which is smaller than a hole in the distal end of the grip body. 9. A torque device having a lumen adapted for reception of a portion of a sensor guide wire, comprising: a grip body;a cap adapted to be joined to the grip body;a number of chuck segments provided on the cap;wherein the torque device is a one-way device defining an intended insertion direction for the sensor guide wire and wherein the chuck segments have free ends which are directed in the insertion direction; anda mechanism configured to allow introduction of the sensor guide wire in the intended insertion direction and prevent introduction of the sensor guide wire in an opposite ... | Radi Medical Systems AB | A61M025/00
US08092381 | B2 | 2012-01-10 | Threshold training system | 2007-05-31 | A system for increasing the fitness level of a fitness enthusiast. The system includes a personalized set of intensity zones corresponding to increasing levels of exercise intensity. The personalization of the system is accomplished through determining a threshold level based on oxygen consumption or a numerical rating of perception of effort or lactate or metabolism that is unique to each individual. From this determination, an anchor point is created upon which the heart rate values for each of eight zones is based. Each zone corresponds to a multiplier that when factored in to the amount time spent in each zone by the individual, yields a total training load value. | 1. A method for creating on a chart a personalized set of intensity zones of an individual, the method comprising: a. measuring blood lactate level of an individual while the individual exercises, and calculating therefrom a threshold heart rate corresponding to a lactic acid metabolic marker;b. determining an anchor point based on said threshold heart rate;c. providing: i. a computer readable memory for storage of said anchor point; andii. a processor in communication with said memory that contains instructions that determine a heart rate range as a percentage of said anchor point for each of a plurality of heart rate intensity zones, wherein said heart rate range comprises an upper heart rate and a lower heart rate, and wherein one of said intensity zones comprises a threshold zone beyond which zones of higher intensity correspond to increases in said metabolic marker;d. determining a multiplier associated with each said intensity zone; ande. printing a chart comprising personalized information regarding said intensity zones and said anchor point. 2. The of method claim 1 wherein said chart comprises major subdivisions that correspond to said intensity zones, and one of said major subdivisions is further divided into a plurality of minor subdivisions. 3. The method of claim 2 wherein each of said major subdivisions corresponds to a color. 4. The method of claim 1 wherein said individual's heart rate is within said heart rate range for an amount of time, said system further comprising the step of multiplying said amount of time by said multiplier to determine a training load. 5. The method of 1 wherein for intensity zones above said threshold zone said multiplier is substantially exponentially related thereto. 6. The method of claim 5 wherein intensity zones below said threshold zone have a substantially linear correspondence to said multiplier. 7. The method of claim 1 wherein metabolic markers relate to lactic acid level or oxygen exchange rates. 8. A method of creating on a chart a personalized set of intensity zones for an individual, the method comprising the steps of: a. determining a threshold heart rate of an individual as the individual partakes in an activity, wherein said threshold heart rate is determined by measuring a blood lactate level of the individual and determining a corresponding lactic acid metabolic marker;b. determining an anchor point based on said threshold heart rate;c. providing: i. a computer readable memory for storage of said anchor point; andii. providing a processor in communication with said memory and that contains instructions that generate based on said anchor point a personalized heart rate range for each of a plurality of heart rate intensity zones including one threshold zone beyond intensity zones correspond to an increase in lactic acid buildup, and wherein each said intensity zone corresponds to a multiplier; andd. printing a chart comprising said personalized heart rate range and said anchor poi... | Heart Zones USA | A61B005/00
US08543189 | B2 | 2013-09-24 | Method and apparatus for electromagnetic navigation of a magnetic stimulation probe | 2008-04-17 | A system and method is disclosed that can be used to track and navigate an instrument relative to a patient. The system can include an electromagnetic tracking system to track an electromagnetic stimulation probe, such as a transcranial magnetic stimulation probe. The system can, according to various embodiments, provide a tracking device on the probe, track the coil of the probe, provide a field relative to the probe, and determine the position of the patient based upon the field produced by the probe. | 1. A system to navigate a procedure for a patient, comprising: an electromagnetic tracking system including an electromagnetic localizer system having at least a first conductive coil to be configured with a first set of parameters to generate a localization field;a magnetic stimulation probe configured to be external to the patient and having at least the first conductive coil that is configured to be switched to operate as the electromagnetic localizer system to generate the localization field with the first set of parameters and configured to produce a probe electromagnetic field having an electromagnetic focal region with a second set of parameters, wherein the probe electromagnetic field is configured to induce a current in a conductive material of the patient; anda processor system configured to determine the position of the electromagnetic focal region relative to the patient based at least in part on the localization field produced by the electromagnetic localizer system;wherein the probe electromagnetic field of the magnetic stimulation probe is operable to induce a current in a conductive material including inducing the therapeutic current to stimulate a selected portion of the patient at the electromagnetic focal region. 2. The system of claim 1, wherein the magnetic stimulation probe operates to produce the probe electromagnetic field to induce an electric stimulation within a conductive material, wherein the magnetic stimulation probe is energized to stimulate the selected portion of the patient. 3. The system of claim 2, wherein the first set of parameters includes a current at a first frequency that can be driven through the first conductive coil to produce the localizer electromagnetic field; and wherein the second set of parameters includes a current at a second frequency that can be driven through the first conductive coil to produce the probe electromagnetic field. 4. The system of claim 3, wherein the first frequency is less than the second frequency. 5. The system of claim 4, wherein the first frequency is about two kilohertz and the second frequency is about four kilohertz. 6. The system of claim 1, further comprising: a dynamic reference frame configured to be associated with the patient;wherein the electromagnetic localizer field is sensed by the dynamic reference frame. 7. The system of claim 1, further comprising: an imaging device to obtain image data of the patient. 8. The system of claim 7, further comprising: a display device operable to display the image data of the patient;wherein the processor system is operable to determine a location of an icon displayed with the display device illustrating a position of the focal region where the probe electromagnetic field is operable to induce a current in the conductive material. 9. The system of claim 8, wherein the illustrated position is illustrated in a brain of the patient. 10. The system of claim 8, further comprising: a memory system operable to store a proce... | Medtronic Navigation, Inc. | A61B005/05
US08643675 | B2 | 2014-02-04 | Method and a system for interactive probing and annotating medical images using profile flags | 2006-10-18 | A method (100) of and a system (800) for probing a data of interest on the basis of an object comprised in a related image data, includes an instantiating step (105) for instantiating a probe, an applying step (115) for applying the probe to a location on a surface of the object, and a determining step (120) for determining a profile from the data of interest on the basis of the location of the probe. The probe is a virtual tool for extracting information from the data of interest. This tool is navigated on the basis of a view rendered from the related image data. The probe further determines the scope of probing. Thus, the method (100) and the system (800) enable the user to extract the desired information from the data of interest. | 1. A method of probing data of interest on a basis of an object comprised in a related image data, the method comprising: acquiring, with an image acquisition apparatus unit, the related image data representing the object;rendering, on an associated display, an image of the object from the related image data;with a processor, instantiating a probe in the displayed image;applying the probe to a location on a surface of the object in the displayed image; anddetermining, with the processor, a profile from the data of interest associated with the object on a basis of the location of the probe relative to the surface of the object in the displayed image. 2. The method as claimed in claim 1 wherein the profile represents a property of the object along at least one trajectory extending into the object from the location on the surface of the object. 3. The method as claimed in claim 2 further including with the probe defining: a direction of the trajectory;a depth of the trajectory; and/or a plurality of trajectories. 4. The method as claimed in claim 3 further comprising visualizing the profile on the associated display. 5. The method as claimed in claim 4 wherein the probe comprises a flag on which a graph depicting the property of the object vs depth below the surface of the object. 6. The method as claimed in claim 4 wherein the probe defines a plurality of parallel trajectories arranged in a line to define a cutting plane for visualizing a cross section of the object. 7. The method as claimed in claim 2 further comprising moving the probe along the surface of the object in the displayed image for browsing the property along trajectories from different locations on the object. 8. The method as claimed in claim 4 further comprising rotating the probe to rotate the cutting plane. 9. The method as claimed in claim 1, further comprising: applying a second probe to a second location on the surface of the object in the displayed image; anddetermining a second profile on the basis of the second location. 10. The method as claimed in claim 1 further comprising; annotating the probe;storing the image and the annotated probe;displaying the stored image in association with the annotated probe. 11. A non-transitory computer-readable medium carrying computer code that, when executed by a computer, controls the computer to perform the method of claim 1. 12. The method of claim 1, wherein the related image data is magnetic resonance imaging data, computed tomography data, X-ray data, ultrasound data, positron emission tomography data, or single photon emission computed tomography data. 13. The method as claimed in claim 1, wherein the object is an internal anatomical organ or structure. 14. A system for probing properties of an object , the system comprising: a display device;a user input device:a processor configured to: receive image data corresponding to the object,render an image of the object from the image data on the display device,render a prob... | Koninklijke Philips N.V. | G06T019/00
US08812071 | B2 | 2014-08-19 | Transdermal analyte monitoring systems and methods for analyte detection | 2008-03-06 | Transdermal analyte monitoring systems (TAMS) having increased longevity and improved analyte detection are described herein. Kits for use with the TAMS and methods of using the TAMS and kits are also described. In a preferred embodiment, the TAMS includes a protective, semi-permeable membrane covering the surface of the hydrogel. The protective, semi-permeable membrane contacts with the skin of a user and prevents contamination or fouling of the hydrogel. Optionally, the hydrogel comprises one or more humectants and/or an immobilized enzyme. In another preferred embodiment, the TAMS contains at least one channel or pocket for increasing the amount of oxygen provided to the hydrogel. In one embodiment, a method for improving analyte detection by the TAMS is provided. For example, after the skin porosity is increased by an appropriate pretreatment, a skin preparation wipe is applied to the treated skin area and then the TAMS is applied to the treated area. | 1. A transdermal analyte monitoring system comprising: a sensor assembly, wherein the sensor assembly comprises a hydrogel and a sensor body containing a plurality of electrodes, wherein the sensor body is in fluid communication with the hydrogel, wherein the hydrogel comprises a humectant and an enzyme, and wherein the humectant is in an effective amount to increase the performance longevity of the transdermal analyte monitoring system, as indicated by decreased signal drift, compared to the same system in the absence of the humectant, andwherein the humectant is sodium pyrrolidone carboxylic acid (NaPCA). 2. The transdermal analyte monitoring system of claim 1, wherein the hydrogel comprises a polymer selected from the group consisting of polyethylene glycol diacrylate (PEGDA), agarose, polyethylene glycol diacrylate/polyethyleneimine (PEGDA-PEI), polyethylene glycol diacrylate-n-vinyl pyrrolidone (PEGDA-NVP), acrylate-polyethylene glycol-N-hydroxy succinimide (A-PEG-N), and blends and copolymers thereof. 3. The transdermal analyte monitoring system of claim 1, wherein the enzyme is oxidase enzyme. 4. The transdermal analyte monitoring system of claim 1, wherein the enzyme is covalently immobilized in the hydrogel. 5. The transdermal analyte monitoring system of claim 4, wherein the enzyme is covalently immobilized in the hydrogel using an A-PEG-N. 6. The transdermal analyte monitoring system of claim 1, wherein the sensor assembly comprises at least one channel or pocket for providing oxygen to the hydrogel. 7. The transdermal analyte monitoring system of claim 1, wherein the enzyme is immobilized in the hydrogel via non-covalent immobilization. 8. The transdermal analyte monitoring system of claim 1, further comprising a semi-permeable membrane on the surface of the hydrogel, wherein the membrane is in fluid communication with the hydrogel. 9. The transdermal analyte monitoring system of claim 8, wherein the hydrogel and the semi-permeable membrane form an interpenetrating polymer network. 10. A method for increasing analyte detection by a transdermal analyte monitoring system comprising: (a) treating a region of skin of the user to increase permeability; and(b) applying to the treated region of skin a transdermal analyte monitoring system comprisinga sensor assembly, wherein the sensor assembly comprises a hydrogel and a sensor body containing a plurality of electrodes, wherein the sensor body is in fluid communication with the hydrogel wherein the hydrogel comprises a humectant and an enzyme, and wherein the humectant is in an effective amount to increase the performance longevity of the transdermal analyte monitoring system, as indicated by decreased signal drift, compared to the same system in the absence of the humectant, andwherein the humectant is sodium pyrrolidone carboxylic acid (NaPCA). 11. The method of claim 10, wherein the region of skin is treated in step (a) by a method selected from the group consisting of tape stri... | Echo Therapeutics, Inc. | A61B005/05 A61B005/00 A61B005/145 A61B005/1486
US08892188 | B2 | 2014-11-18 | Identifying abnormal tissue in images of computed tomography | 2006-02-03 | An imaging method for identifying abnormal tissue in the lung is provided, comprising the recording of slice images of the lung by means of X-ray radiation, recording of blood vessels, differentiation of blood vessels and abnormal tissue, segmentation of the abnormal tissue and display of the segmented abnormal tissue on an output device. In addition, a computer tomograph for identifying abnormal tissue in the lung is provided, having a radiation source for recording slice images of the lung and blood vessels by means of X-ray radiation, a computer unit for differentiating the blood vessels from the abnormal tissue and for segmenting the abnormal tissue, as well as an output device for displaying the segmented abnormal tissue. Furthermore, a computer program is provided for controlling a computer tomograph for an identification of abnormal tissue in the lung by means of a radiation source, designed to record slice images of the lung and blood vessels by means of X-ray radiation, to differentiate the blood vessels from abnormal tissue, to segment the abnormal tissue and to control an output device for displaying the abnormal tissue. | 1. An imaging method for identifying abnormal tissue in a lung, comprising: recording slice images of the lung by means of X-ray radiation to produce a slice image of the lung, where the slice image of the lung includes at least one elongate blood vessel and at least one abnormal tissue;determining diameters of the at least one blood vessel and of the at least one abnormal tissue in the slice images by calculating a starting point corresponding to a first end of the at least one blood vessel, commencing from the starting point a wave front of constant speed, and measuring distances at the wave front between opposing walls of the blood vessel and the at least one abnormal tissue, wherein the measured distances provide direct approximations of the diameters, wherein the wave front propagates from the first end of the at least one blood vessel along a longitudinal axis of the at least one blood vessel to a second end of the at least one blood vessel, wherein the wave front is halted at boundaries of the at least one blood vessel;evaluating the diameters of the at least one blood vessel and of the at least one abnormal tissue at the different locations;differentiating the at least one blood vessel and the at least one abnormal tissue based on the evaluated diameters;segmenting the at least one abnormal tissue; anddisplaying the segmented abnormal tissue on an output device. 2. The imaging method as claimed in claim 1, wherein the wave front divides into multiple wave fronts at each branching point of the at least one blood vessel such that a wave front propagates along each branch. 3. The imaging method as claimed in claim 1, where differentiating the abnormal tissue is further done on a basis of a path of the wave front to the boundary of the blood vessels starting from a branching point of the blood vessels. 4. The imaging method as claimed in claim 1, where displaying the segmented abnormal tissue on the output device comprises displaying a lung cavity with the abnormal tissue and without the blood vessels on the output device. 5. The imaging method as claimed in claim 4, further comprising: determining the boundaries of the blood vessels with the lung cavity based on gray-scale values of recording data. 6. The imaging method as claimed in claim 1, where segmenting the abnormal tissue comprises segmenting the slice image of the lung into pulmonary blood vessels and lung cavity. 7. A computer tomograph for identifying abnormal tissue in a lung, comprising: a radiation source configured to record slice images of the lung and blood vessels having a longitudinal axis by means of X-ray radiation,a computer unit configured to determine diameters of the blood vessels and of the abnormal tissue in the slice images, to calculate starting points in the blood vessels corresponding to a first end of each blood vessel, commence from the starting points of a wave front of constant speed, and measure distances at the wave front between opposing walls of th... | Koninklijke Philips N.V. | A61B006/03 G06T007/00 G06K009/34 A61B006/00
US08560044 | B2 | 2013-10-15 | Garment accessory with electrocardiogram sensors | 2007-05-16 | A garment accessory includes a member having first and second ends, with the length of the member between the first and second ends being less than the circumference of a subject that the member is configured for and a pair of fastener mechanisms disposed in proximity to the first and second ends of the member, the fastener mechanisms configured to attach the member to an article of clothing worn by a subject. The garment accessory also includes at least a pair of sensors supported by the member with the sensors being at least one of ECG sensors, motion sensors, body temperature sensors and impedance plethysmography sensors. | 1. A garment accessory comprises: a member having first and second ends and a length between the first and second ends, with the length of the member between the first and second ends being approximately half of the circumference of the chest of a subject that the member is configured to be positioned on;a pair of fastener mechanisms disposed in proximity to the first and second ends of the member, the fastener mechanisms configured to attach the member to an article of clothing at portions of the article of clothing that are at the sides of the subject, when the article of clothing is worn by the subject; andat least a pair of sensors supported at the pair of fastener mechanisms of the member, the sensors having sensor membranes of the sensors configured to rest against the skin of the subject;the fastener mechanisms comprising tabs comprised of a stiff material of a sufficient stiffness and having fold portions to attach to the article of clothing, with the fold portions disposed to fold away from the skin of the subject. 2. The garment accessory of claim 1, wherein the member further has a fastener mechanism disposed at the center of the member. 3. The garment accessory of claim 2 wherein the fastener mechanisms at the first and second ends of the member and the center of the member are configured to couple the member to a brassiere. 4. The garment accessory of claim 3 wherein the member is comprised of a thin, firm, flexible band of material, the garment accessory further comprising: a set of conductors carried in the flexible band of material; anda circuit arrangement electrically coupled to the pair of sensors by the set of conductors, the circuit arrangement carried by the member. 5. The garment accessory of claim 4 wherein the circuit arrangement includes a wireless transmitter. 6. The garment accessory of claim 1 wherein the pair of sensors are removable from the member. 7. The garment accessory of claim 1 wherein the fastener mechanisms further comprise at least one of hooks, clips, elastic band, hook and loop fasteners or snaps. 8. The garment accessory of claim 1 wherein the tabs of the fastener mechanisms maintain a hook shape when bent to form hook mechanisms that hook over portions of the article of clothing. 9. The garment accessory of claim 1 wherein the tabs of the fastener mechanisms fold to form clips disposed on the bottom outside edge of the garment accessory, configured such that when the garment accessory is underneath the article of clothing that encircles the torso of the subject, the clips attach to the bottom edge of the article of clothing. 10. The garment accessory of claim 1 wherein the sensors are at least one of ECG sensors, motion sensors, body temperature sensors or impedance plethysmography sensors. 11. The garment accessory of claim 1 wherein the article of clothing that the garment accessory is configured to attach to is a brassiere. 12. The garment accessory of claim 1 wherein the garment accessor... | Medicomp, Inc. | A61B005/0408 A61B005/0416
US10456067 | B2 | 2019-10-29 | Blood collection assembly | 2005-07-20 | A blood collection assembly includes a hub with a distal end, a proximal end, a hub outer surface and an internal opening extending therethrough. The assembly also includes a holder housing defining a receiving chamber and having a rearward end adapted to receive a sample collection tube within the chamber and a forward end including a receiving port extending into the chamber, where the receiving port receives a portion of the proximal end of the hub therein. An internal opening of the hub accommodates a puncturing element at the proximal end thereof for providing passage of a fluid therethrough, and the puncturing element contacts the sample collection tube. The hub and the holder housing include an interengaging structure extending substantially perimetrically therearound for axially locking the hub with the holder housing. | 1. A blood collection assembly, comprising: a holder housing defining a receiving chamber and including a rearward end adapted to receive a sample collection tube within the chamber, a forward end including a receiving port having an interior surface, and a receiving structure formed within the interior surface of the receiving port, the receiving structure comprising a radial receiving portion comprising an annular groove extending perimetrically within the interior surface of the receiving port and dividing the interior surface into a first portion and a second portion and a longitudinal receiving portion; anda hub comprising a hub outer surface extending between a distal end and a proximal end, the hub further comprising an internal opening extending therethrough and a puncturing element at the proximal end thereof;wherein the hub outer surface comprises a radial locking component and an anti-rotation member, the radial locking component comprising an annular protrusion extending perimetrically about the hub outer surface and dividing the outer surface of the hub into a first portion and a second portion and the anti-rotation member protruding from and extending longitudinally along the hub outer surface perpendicular to the radial locking component,wherein a first portion of the anti-rotation member extends from the first portion of the outer surface of the hub on one side of the annular protrusion in a longitudinal direction towards the distal end of the hub and a second portion of the anti-rotation member extends from the second portion of the outer surface of the hub on an opposite side of the annular protrusion in a longitudinal direction towards the proximal end of the hub, the annular protrusion intersecting the anti-rotation member,wherein upon receipt of the anti-rotation member within the receiving port, the anti-rotation member and the longitudinal receiving portion are adapted to align the hub and the holder housing in a predetermined position with respect to one another and are adapted to prevent rotational movement of the hub within the receiving port;wherein, with the hub received within the receiving port of the holder housing, the radial locking component lockably engages the radial receiving portion and the anti-rotation member engages the longitudinal receiving portion thereby providing a snap-fit engagement for axially locking the hub to the holder housing with the puncturing element of the hub extending through the receiving port and into the chamber of the holder housing, andwherein the first portion of the interior surface of the receiving port contacts the first portion of the outer surface of the hub and the second portion of the interior surface of the receiving port contacts the second portion of the outer surface of the hub. 2. The assembly of claim 1, further comprising a needle cannula including an internal lumen mounted through the internal opening of the hub, the needle cannula having a forward intravenous end ... | Becton, Dickinson and Company | A61B005/15 A61B005/154
US09283288 | B2 | 2016-03-15 | Methods of using optical agents | 2006-10-11 | In certain aspects, the invention relates to processes for using renally excretable optical agents to detect one or more tissues of the renal system of a surgical patient. In certain aspects, the invention relates to a kit including a biocompatible composition containing one or more optical agents and instructions for using the optical agent(s) in a process of the present invention. | 1. A process for using a fluorescent agent in an abdominal or pelvic surgical procedure to reduce the risk of accidental injury to a ureter of a patient undergoing an abdominal or pelvic surgical procedure, the process comprising: delivering a fluorescent agent to the ureter of a patient undergoing an abdominal or pelvic surgical procedure,wherein the fluorescent agent is delivered intravenously, enterally, intraperitoneally, transdermally, or via inhalation to the patient,wherein at least a fraction of the fluorescent agent is excreted through the renal system such that at least a fraction of the fluorescent agent is present in the patient's urine,wherein the fluorescent agent is chosen from phenylxanthenes, phenothiazines, phenoselenazines, cyanines, indocyanines, squaraines, dipyrrolo pyrimidones, anthraquinones, tetracenes, quinolines, pyrazines, acridines, acridones, phenanthridines, azo dyes, rhodamines, phenoxazines, azulenes, azaazulenes, triphenyl methane dyes, indoles, benzoindoles, indocarbocyanines, benzoindocarbocyanines, derivatives having the general structure of 4,4-difluoro-4-bora-3a,4a-diaza-s-indacene, and conjugates thereof and derivatives thereof,irradiating a ureter of the patient's renal system with non-ionizing radiation to cause the fluorescent agent to fluoresce within the ureter; andoptically detecting the fluorescent agent in the irradiated ureter to distinguish the ureter from surrounding tissues and reduce the risk of accidental injury to the ureter. 2. The process of claim 1, wherein the fluorescent agent is selected from phenylxanthenes, phenothiazines, phenoselenazines, cyanines, indocyanines, squaraines, dipyrrolo pyrimidones, anthraquinones, tetracenes, quinolines, pyrazines, acridines, acridones, phenanthridines, azo dyes, rhodamines, phenoxazines, azulenes, azaazulenes, triphenyl methane dyes, indoles, benzoindoles, indocarbocyanines, benzoindocarbocyanines, conjugates thereof and derivatives thereof. 3. The process of claim 1, wherein the fluorescent agent is selected from fluorescein, and 3,6-diaminopyrazine-2,5-dicarboxylic acid. 4. The process of claim 1, wherein the non-ionizing radiation is selected from visible radiation, ultraviolet radiation, infrared radiation, and combinations thereof. 5. The process of claim 1, wherein the fluorescent agent is detected using at least one of unaided eye, camera, charged coupled device, photomultiplier tube, avalanche diode, and photodiodes. 6. The process of claim 1, wherein the abdominal or pelvic surgical procedure is selected from total or partial hysterectomy, oophorectomy, tubal ligation, surgical removal of ovarian cysts, anterior repair of the vaginal wall, caesarean section, repair of a pelvic prolapse, pelvic mass resection, removal of a fallopian tube, adnexectomy, removal of an ectopic pregnancy, vasectomy, prostatectomy, hernia repair surgery, colectomy, cholecystectomy, appendectomy, hepatobiliary surgery, splenectomy, distal or total pancreatectom... | MediBeacon, Inc. | A61B005/00 A61K049/00
US20080300499 | A1 | 2008-12-04 | Portable Apnea and Cardiac Monitor | 2008-05-19 | This invention relates to non-invasive monitoring devices, in monitoring minute movements made by the body to report the presence or absence of respiration and normal heart rate. The movements being measured are the pulsations made by the normal operation of both the heart and lungs. This is accomplished by converting the body movements into electrical signals with an adapted polymorph-piezoelectric transducer. The electrical signal that is detected is a mixed signal, a signal composed of a lower frequency (breathing contraction) and the higher frequency (heart pulse). The signals are detected by separating the body motions by an adaptive filter that has a break frequency at twice the frequency of the larger body movement signal (breathing contraction).;The separate body movement signals operate with associated logic circuitry to allow each signal to be independently measured, recorded, and acted upon to determine if the individual (patient) is undergoing a life threatening experience. Such an event by either signal will trigger an audio and visual display alarm which is attached on or near the patient. The alarm event trigger is also connected to a low voltage detector circuit to indicate that the supply voltage has dropped below an acceptable level. | 1. A portable motion detector for monitoring the heart rate and respirational activity of a patient, which comprises; polymorph-piezoelectric transducer means adapted for converting body expansion during the heart pumping and breathing process into a mixed electrical signal, 2. In a digital signal processing system, a method of computing a motion decision value from an electromechanical transducer which comprises the following steps: (A) A signal de-mixing method of splitting the bodily motion signal into a high pass signal (heart rate) and a low pass signal (breathing/respiration);(B) a logic or computational package;(C) a means of temporarily recording the information obtained by said computational package;(D) a means of transmitting the information recorded to long term storage/tracking/data manipulation means, and(E) an audio and visual warning device,(F) a self contained power source;(G) a charging means for the self contained power source; 3. The method according to claim 2, which further separates the electrical signal by using low-pass/high-pass filtering de-multiplexation of the composite signal into separate signals by an adaptive filter, a lower frequency (breathing/respiration contraction) and a higher frequency (heart pulse). The adaptive filter that has a break frequency at twice the frequency of the larger body movement signal (breathing/respiration contraction). 4. The method according to claim 2, wherein the step of forming the motion detection signal comprises computing where the separate body movement signals operate with associated logic circuitry to allow each signal to be independently measured, recorded, and acted upon to determine if the individual (patient) is undergoing a life threatening experience. 5. The method according to claim 2, wherein the step of forming the motion detection signal comprises separating each bodily motion signal where: (A) The counting circuit coupled to the output of said signal de-mixing means for producing an output signal representative of said heart pumping signal being not within an except-able limit over a predetermined counting period,(B) counting circuit coupled to the output of said signal mixing means for producing an output signal representative of the absence of a pumping or chest movement breathing signal over a predetermined counting period,(C) an oscillator circuit coupled to the input of said counting circuit for providing a signal for counting, 6. The method according to claim 4, wherein the output of said counting circuits for producing a signal indicating that heart rate is not within an acceptable limit or that respiration has ceased will trigger an audio and visual display alarm which is attached on or near the patient. 7. The method according to claim 2, wherein the low voltage detector circuit means coupled to the input of said audio output means for producing an output indicative of supply voltage dropping below a predetermined level, switch means coupled to the out... | - | A61B005/0205
US20080281550 | A1 | 2008-11-13 | SYSTEMS AND METHODS FOR CHARACTERIZING BALANCE FUNCTION | 2008-05-12 | The present invention relates to systems and methods for characterizing balance function. In particular, the present invention provides systems and methods for monitoring balance function (e.g., in a single individual and/or plurality of individuals), generating one or more databases comprising balance function data, processing and/or analyzing databases comprising balance function data, and characterizing balance function (e.g., in a single individual and/or plurality of individuals (e.g., utilizing databases comprising balance function data)). Systems and methods of the present invention find use in, among other things, research, diagnostic and therapeutic applications. | 1. A method for objectively scoring and/or assessing balance function in a subject comprising: a) providing a subject;b) instructing the subject to perform one or more physical tests;c) monitoring motion and/or movement of one or more locations present on said subject during performance of the physical tests, wherein monitoring motion and/or movement comprises use of a balance function data collection and analysis system; andd) correlating the subject's motion and/or movement to one or more balance function footprints to provide the subject a score and/or assessment of the subject's balance function or a diagnosis of a subject based on their balance function or balance function data, compared to a balance function footprint. 2. The method of claim 1, wherein said balance function data collection and analysis system is configured to detect, transmit, store, process and assess balance function data. 3. The method of claim 2, wherein detecting and transmitting balance function data comprises use of one or more motion sensors. 4. The method of claim 2, wherein storing balance function data comprises placement of balance function data on storage means. 5. The method of claim 2, wherein processing and assessing balance function data comprises use of a computer processor running software configured to analyze and assess the balance function data. 6. The method of claim 1, wherein said subject is instructed to perform one or more tasks present within one or more tests, wherein said one or more tests are selected from the group consisting of dynamic gait index test, functional gait assessment test, Berg balance test, timed up and go test, and the dynamic task test. 7. A method for generating a balance function footprint comprising: a) providing a plurality of subjects, wherein each subject of the plurality of subjects is identified as having a certain, common etiology;b) a balance function data collection and analysis system, wherein said system is configured to detect, transmit, store, process and assess balance function data, wherein detecting and transmitting balance function data comprises use of one or more motion sensors; wherein storing balance function data comprises placement of balance function data on storage means; and processing and assessing balance function data comprises use of a computer processor running software configured to analyze and assess the balance function data;c) instructing each subject of the plurality of subjects to perform an identical series of tasks,d) using the balance function data collection and analysis system to i) acquire motion and/or movement data from the subjects during the subject's performance of the task; andii) standardize the balance function data acquired to generate a balance function footprint. 8. A system comprising: a) a database comprising one or more balance function footprints;b) one or more sensors. 9. The system of claim 8, wherein said sensors are configured to transfer and/or transduce ... | WICAB, INC. | G06F015/00 A61B005/00
US20080275310 | A1 | 2008-11-06 | Apparatus for Monitoring Sweating State by Adjusting Temperature | 2006-04-06 | Disclosed is an apparatus for monitoring a sweating state by adjusting skin and core temperatures, in which infrared ray heating means for raising the body temperature and cameras for photographing the sweating states of the upper and the lower part of the body of a person to be diagnosed are provided in a chamber so that the sweating states of the respective parts in accordance with the increase in the body temperature are photographed and are stored in a database and the sweating patterns are analyzed to simply and correctly diagnose the focus and the aching part of the person to be diagnosed. The apparatus includes a chamber in which a person to be diagnosed lies to be diagnosed, a temperature controller provided in the chamber to maintain the temperature of the chamber, body heating means provided in the chamber to emit infrared rays to the body and to raise the body temperature, body temperature sensors for detecting the core and the skin temperature of the body, cameras provided in the chamber to photograph the sweating pattern caused by an increase in the body temperature, a controller for sensing the body temperature of the person to be diagnosed and the temperature of the previously set temperature, and a computer for receiving the sweating pattern of the person to be diagnosed that is transmitted from the cameras to store, mange, and monitor the sweating pattern and for analyzing data on the body temperature of the person to be diagnosed and on the environments of the chamber that are received from the controller to control the controller to maintain the optimal environments. chamber to control the body heating means at a previously set temperature, and a computer for receiving the sweating pattern of the person to be diagnosed that is transmitted from the cameras to store, mange, and monitor the sweating pattern and for analyzing data on the body temperature of the person to be diagnosed and on the environments of the chamber that are received from the controller to control the controller to maintain the optimal environments. | 1. An apparatus for monitoring a sweating state by adjusting a temperature comprising: a chamber in which a person to be diagnosed lies to be diagnosed;body heating means provided in the chamber to emit infrared rays to the body and to raise the body temperature;body temperature sensors for detecting the core surface and the skin temperature of the body;cameras provided in the chamber to photograph the sweating pattern caused by an increase in the body temperature;a controller for sensing the body temperature of the person to be diagnosed and the temperature of the chamber to control the body heating means at a previously set temperature; anda computer for receiving the sweating pattern of the person to be diagnosed that is transmitted from the cameras to store, mange manage, and monitor the sweating pattern and for analyzing data on the body temperature of the person to be diagnosed and on the environments of the chamber that are received from the controller to control the controller to maintain the optimal environments. 2. The apparatus as claimed in claim 1, further comprising: a temperature controller provided in the chamber to maintain the temperature of the chamber in accordance with the control of the controller; anda humidity controller provided in the chamber to maintain the humidity of the chamber to be in a predetermined range in accordance with the control of the controller. 3. The apparatus as claimed in claim 2, wherein the temperature controller comprises: a chamber temperature sensor for sensing the temperature of the chamber to output the sensed temperature to the controller; anda heater that operates in accordance with the control of the controller to heat the chamber. 4. The apparatus as claimed in claim 2, wherein the humidity controller comprises: a humidity sensor for sensing the humidity of the chamber to output the sensed humidity to the controller; anda humidifier that operates in accordance with the control of the controller to supply moisture to the inside of the chamber. 5. The apparatus as claimed in claim 1, wherein the body heating means are infrared ray heating devices provided on the internal top surface of the chamber 6. The apparatus as claimed in claim 1, wherein the cameras are a plurality of cameras for photographing the sweating states of the upper part of the body and the lower part of the body of the person to be diagnosed. 7. The apparatus as claimed in claim 1, wherein the body temperature sensors comprise: a core temperature sensor inserted into the mouth of the person to be diagnosed to measure the core temperature of the body; andskin temperature sensors attached to the legs, the body, and the face of the person to be diagnosed to measure the skin temperatures of the body. 8. The apparatus as claimed in claim 1, wherein a speaker and a mike are provided in the chamber so that the person to be diagnosed and a doctor have a conversation. 9. The apparatus as claimed in claim 1, wherein windows th... | - | A61B005/00
US20080275308 | A1 | 2008-11-06 | Premium-Based Civilly-Catastrophic Event Threat Assessment | 2006-10-09 | One accepts (101) consideration-based private civil security subscriptions with respect to providing the civilly-catastrophic event threat assessment. One gathers (102) information regarding specific civilly-catastrophic event causation agents. One may optionally gather (103) information regarding an authorized beneficiary. One then uses (104) the gathered information, at least in part, to formulate a civilly-catastrophic event threat assessment. One then provides (105) a civilly-catastrophic event threat assessment. One may optionally provide (106) the authorized beneficiary with communication equipment. One may optionally provide (107) survival instructions to the authorized beneficiary. One may optionally prompt (108) anticipatory steps by the authorized beneficiary. One may optionally provide (109) access to at least one life-sustaining resource. One may optionally provide (110) corresponding services and/or information. | 1. A method comprising: accepting consideration-based private civil security subscriptions from subscribers with respect to providing the civilly-catastrophic event threat assessments to an authorized beneficiary;gathering information regarding specific civilly-catastrophic event causation agents;using the information, at least in part, to formulate a civilly-catastrophic event threat assessment;providing to the authorized beneficiary of one of the consideration-based private civil security subscriptions the civilly-catastrophic event threat assessment related to any of a plurality of categorically different civilly-catastrophic events having at least one of: occurred; andpotentially occurring;such that an authorized beneficiary of a consideration-based private civil security subscription is able to reliably and predictably receive specific information regarding potential civilly-catastrophic events. 2. The method of claim 1 wherein an entity providing the civilly-catastrophic event threat assessment also provides civilly-catastrophic event-based access to at least one life-sustaining resource. 3. The method of claim 2 wherein the at least one life-sustaining resource comprises at least one of: hydration;nourishment;a life-sustaining shelter;transport away from an area of substantially sudden civil upheaval;environmentally-borne threat abatement;compressed or stored air; anda rescue service to come to an authorized beneficiary as corresponds to one of the subscriptions and move the authorized beneficiary away from a location of substantially sudden civil upheaval. 4. The method of claim 1 wherein gathering information regarding specific civilly-catastrophic event causation agents comprises gathering information from at least one of: a public resource;a non-public resource. 5. The method of claim 4 wherein the non-public resource comprises an authorized beneficiary input. 6. The method of claim 4 wherein the non-public resource comprises at least one monitoring instrument pre-positioned in a specific geographic area. 7. The method of claim 6 wherein the at least one monitoring instrument is a local sensor configured and arranged to detect at least one of the following: ionizing radiation;a chemical agent;a biological agent;a seismic event;a weather condition;a sonic event;a thermal event;a civil disturbance;a biological disturbance. 8. The method of claim 1 wherein using the information, at least in part, to formulate a civilly-catastrophic event threat assessment comprises using the information to formulate a civilly-catastrophic event threat assessment that is specifically targeted for individuals who comprise at least one of: individuals within a particular geographic area;individuals of a particular cultural affiliation;individuals of a particular religious affiliation;individuals of a particular economic affiliation;individuals of a particular ethnic affiliation;individuals of a particular tribal affiliation;individuals of a particular... | - | A61B005/00
US20080281259 | A1 | 2008-11-13 | Automatron Blood Pressure Manager for Nitroglycerine | 2007-05-07 | A means of tightly controlling blood pressure in the acute care setting is elusive. Presently, the bedside nurse manages the rate of the intravenous blood pressure medication being infused. The invention seeks to allow continuous polling of the blood pressure, through telemetry monitoring, and adjust the rate of the intravenous antihypertensive nitroglycerin. The method algorithm allows for a set point of either mean arterial pressure (MAP) or systolic blood pressure (SBP) to be input. The nitroglycerin drip rate is then adjusted to maintain the set parameter range. | 1. A method to control a patient's blood pressure through the use of parenteral (IV) nitroglycerin via a closed-loop process. 2. The method of claim 1 uses the patients measured blood pressure from either a non-invasive blood pressure cuff (NIBP) or a catheter placed in the patient's artery (arterial line). 3. The method of claim 1 intakes the patient's central venous pressure (CVP), if available, and partially basis the dose adjustment rate of nitroglycerin on the collected value. 4. The method of claim 1 intakes the patient's blood pressure measurement and compares it against a set or target blood pressure. Based on the difference between the patient's blood pressure and the target blood pressure, the dose adjustment rate is partially determined. 5. The method of claim 1 is based on the dose adjustment rate determined by claim 3 and 4. | Mr. Tyler Nicholas Carruth AUTOMATRON, LLC | A61B005/02
US20080242952 | A1 | 2008-10-02 | Effective response protocols for health monitoring or the like | 2007-05-29 | Systems, methods, computer program products, and media are described for obtaining data from occupational or leisure intercommunication at least between a device and a user and signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication; or for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user and applying one or more data extraction criteria to the health-status-indicative data surreptitiously captured from the interaction between the device and the user. | 1. An apparatus comprising: one or more physical media configured to bear a device-detectable implementation of a method including at leastreceiving an indication of an anomalous device-interactive performance of a specific individual; andselecting one or more diagnostic instructions at least partly based on the anomalous device-interactive performance of the specific individual. 2-6. (canceled) 7. The apparatus of claim 1 in which a portion of the one or more physical media comprises: one or more processors configured to perform one or more of optical image scanning or auditory pattern scanning upon the device-detectable implementation. 8-20. (canceled) 21. An apparatus comprising: one or more physical media bearing a device-detectable output indicating an occurrence ofreceiving an indication of an anomalous device-interactive performance of a specific individual; andselecting one or more diagnostic instructions at least partly based on the anomalous device-interactive performance of the specific individual. 22. The apparatus of claim 21 in which the one or more physical media comprise: an antenna-containing semiconductor chip. 23. (canceled) 24. The apparatus of claim 21 in which the one or more physical media comprise: a portable module including at least an auditory interface configured to be operated while the portable module is held or worn. 25. The apparatus of claim 21 in which a portion of the one or more physical media comprises: one or more of an image projection module or a touch screen. 26. The apparatus of claim 21 in which the one or more physical media include at least one of a repeater, a communication satellite, or another active module configured to accept first and second portions of the device-detectable output at first and second respective times. 27. The apparatus of claim 21 in which a portion of the one or more physical media comprises: one or more processors configured to perform one or more of text scanning or auditory pattern scanning upon the device-detectable output. 28. The apparatus of claim 21 in which a portion of the one or more physical media comprises: one or more processors configured to perform pattern recognition upon the device-detectable output. 29. The apparatus of claim 21 in which a portion of the one or more physical media comprises: circuitry for using an encryption constraint in at least some of the device-detectable output. 30. The apparatus of claim 21 in which at least one of the one or more physical media comprises: one or more signal-bearing media bearing at least one special-purpose instruction sequence as a portion of the device-detectable output. 31. The apparatus of claim 21 in which at least one of the one or more physical media comprises: one or more signal-bearing media bearing at least one signal from an implementation having at least circuitry for receiving an indication of an anomalous device-interactive performance of a specific individual. 32. (canceled) 33. The appa... | Searete LLC, a limited liablity corporation of the State of Delaware | A61B005/00
US20080268428 | A1 | 2008-10-30 | Chromosome 5 Genetic Variants Related to Dyslexia | 2004-11-12 | An isolated polynucleotide or genetic material from human Chromosome 5 that indicates the presence of dyslexia or a predisposition to develop dyslexia in the individual from whom which the sample was obtained. A method of diagnosing dyslexia or a predisposition to develop dyslexia. | 1. An isolated polynucleotide comprising at least about 17 consecutive nucleotides of SEQ ID NO:1 including residue 2285, where residue 2286 has an A to C substitution; or comprising at least about 17 consecutive nucleotides of SEQ ID NO:1 including residue 3281, where residue 3282 has a T to G substitution; or comprising at least about 25 consecutive nucleotides of SEQ ID NO:1 including residue 2285, where residue 2286 has an A to C substitution; or comprising at least about 25 consecutive nucleotides of SEQ ID NO:1 including residue 3281, where residue 3282 has a T to G substitution; or comprising at least about 40 consecutive nucleotides of SEQ ID NO:1 including residue 2285, where residue 2286 has an A to C substitution; or comprising at least about 40 consecutive nucleotides of SEQ ID NO:1 including residue 3281, where residue 3282 has a T to G substitution. 2. Isolated genetic material from human Chromosome 5 of an individual that indicates the presence of dyslexia or a predisposition to develop dyslexia in the individual from whom the material was obtained, the material comprising a sufficient portion of SEQ ID NO:1 comprising (Haplotype #1) an A to T substitution at residue 879 and a G to A substitution at residue 2613; or comprising (Haplotype #2) an A to C substitution at residue 424, a C to A substitution at residue 554, a C to T substitution at residue 1346, an A to C substitution at residue 2286, a G to A substitution at residue 2314 and a G to A substitution at residue 2613; or comprising (Haplotype #3) a G to A substitution at residue 1145 and a G to A substitution at residue 2613; or (Haplotype #4) comprising an A to C substitution at residue 424, a C to A substitution at residue 554, a C to T substitution at residue 1346, a G to A substitution at residue 2314, a G to A substitution at residue 2613 and a T to G substitution at residue 3282; or comprising (Haplotype #5) an A to C substitution at residue 424, a C to A substitution at residue 554, an A to T substitution at residue 879, a C to T substitution at residue 1346, a G to A substitution at residue 2314, a G to A substitution at residue 2613 and a T to G substitution at residue 3282; or comprising (Haplotype #6) an A to T substitution at residue 879; or comprising (Haplotype #7) an A to C substitution at residue 2286 and a G to A substitution at residue 2613; where except for these substitutions, residue 424 is A, residue 554 is C, residue 879 is A, residue 985 is C, residue 1145 is G, residue 1346 is C, residue 2275 is A, residue 2286 is A, residue 2314 is G, residue 2453 is C, residue 2613 is G, residue 3282 is T. 3. A method of diagnosing dyslexia or a predisposition to develop dyslexia, the method comprising: a) providing a sample from an individual containing genetic material from Chromosome 5; andb) analyzing the genetic material for the presence of one or more than one of an isolated polynucleotide comprising at least about 17 consecutive nucleotides of SEQ ID NO:1 i... | - | C12Q001/68 C07H021/04 G01N033/53 A61B005/00
US07433732 | B1 | 2008-10-07 | Real-time brain monitoring system | 2005-02-25 | An exemplary method for monitoring neonate brain function includes processing electroencephalographic data of the neonate to calculate at least one measure, the at least one measure including a non-linear measure, comparing the calculated at least one measure against at least one baseline threshold, wherein the at least one baseline threshold corresponds to a neonate age range, and outputting a result of the comparison. An exemplary system for monitoring neonate brain function includes a computing device for processing electroencephalographic data of the neonate to calculate at least one measure, the at least one measure including a non-linear measure, and for comparing the calculated at least one measure against at least one baseline threshold, wherein the at least one baseline threshold corresponds to a neonate age range, and a display for displaying a result of the comparison. | 1. A method for monitoring neonate brain function, comprising: processing electroencephalographic data of the neonate to calculate a plurality of measures, at least one measure including a non-linear measure, wherein the at least one measure is selected from a group consisting of signal energy, approximate entropy (ApEn), short-term maximum Lyapunov exponent (STLmax), cross-ApEn, coherence, mutual information, pattern-match ApEn, pattern-match regularity statistics, and time-frequency distributions;comparing the calculated at least one measure against at least one baseline threshold, wherein the at least one baseline threshold corresponds to a neonate age range;outputting a result of the comparison;generating an illness index of the neonate based on the comparison result; anddisplaying the illness index, wherein the illness index is a weighted combination of comparison results corresponding to the plurality of measures. 2. The method of claim 1, comprising displaying the result, wherein the result indicates health of the neonate. 3. The method of claim 1, wherein the result indicates whether the neonate's brain is functioning normally. 4. The method of claim 1, wherein the generating comprises: comparing a mean value of each of the plurality of measures for each of multiple channels of the electroencephalographic data over a time interval, against a corresponding normal confidence range;estimating a significance of each of the measures; andbased on the mean values and corresponding significances of the measures, combining the measures to form the illness index. 5. The method of claim 4, wherein the normal confidence range varies with age or age group of the neonate. 6. The method of claim 1, wherein the generating comprises: calculating a value of a weighted function of the plurality of measures; andestimating a significance of the calculated value based on a comparison of the value with an overall normal confidence range, wherein the estimated significance forms the illness index. 7. The method of claim 1, wherein the illness index corresponds to a probability that one of the plurality of measures for a channel of the encephalographic data has a probability distribution that is significantly different from a corresponding measure for a group of normal neonates. 8. The method of claim 7, wherein the probability distribution is significantly different when p&lt;0.05. 9. The method of claim 1, comprising triggering an alarm when the illness index crosses a predetermined threshold. 10. A system for monitoring neonate brain function, comprising: a computing device for processing electroencephalographic data of the neonate to calculate a plurality of measures, at least one measure including a non-linear measure, and for comparing the calculated at least one measure against at least one baseline threshold, wherein the at least one baseline threshold corresponds to a neonate age range; anda display for displaying a result of the comparison,wher... | University of Florida Research Foundation, Inc. | A61B005/00
US07461972 | B2 | 2008-12-09 | One point calibration integrated temperature sensor for wireless radio frequency applications | 2006-02-03 | A low power analog linear temperature sensor integrateable with digital and/or analog circuits in CMOS fabrication processes. The sensor is accurately calibrateable at a single temperature that yields a linear relationship between measured differential voltage and temperature over a wide variety of temperature ranges. The sensor provides stable voltage and current references that are essential for wireless sensor platforms. There are many applications where sensors require stable voltage/current references and the physics of the sensor's transduction mechanisms are themselves temperature dependent. Wireless platforms such as, but not limited to, passive RFID tags with the addition of on- or off-chip sensors provide a low cost solution for a variety of low cost sensor applications. | 1. An analog linear temperature sensor which provides an output, said sensor comprising: integrated temperature sensing elements configured to produce zero differential voltage at absolute zero temperature, wherein only a single measurement at a non-zero absolute temperature is required to calibrate the output over a range of temperatures;said sensor being integrated into a computational device and configured to measure environmental temperature; andsaid computational device being a cellular telephone, wherein said cellular telephone is configured to read temperature data from said sensor and data from at least one other sensor integrated into a radio frequency identifier (RFID) tag wherein said tag includes a wireless transceiver, said tag being attached to an object. 2. An analog linear temperature sensor which provides an output, said sensor comprising: integrated temperature sensing elements configured to produce zero differential voltage at absolute zero temperature, wherein only a single measurement at a non-zero absolute temperature is required to calibrate the output over a range of temperatures;said sensor being integrated into a computational device and configured to measure environmental temperature;said computational device being a cellular telephone, wherein said cellular telephone is configured to read temperature data from said sensor and data from at least one other sensor integrated into a radio frequency identifier (RFID) tag wherein said tag includes a wireless transceiver, said tag being attached to an object; andwherein said tag is configured as a passive diagnostic temperature skin patch sensor. | Altivera L.L.C. | G01K007/25 G01K015/00 A61B005/01 G08C019/00
US20080269640 | A1 | 2008-10-30 | Appliance for Recording Diagnostic Values in the Body | 2006-10-26 | In an appliance for recording diagnostic values in the body, with at least one sensor element that is assigned to a housing, the housing is intended to be able to be implanted with the at least one sensor element. | 1-19. (canceled) 20. An appliance for recording diagnostic values in the body, comprising at least one sensor element which has an associated housing, wherein the housing has the at least one sensor element implanted in it. 21. An appliance for recording diagnostic values in the body, comprising at least one sensor element which has an associated housing, wherein, in order to break open and/or to tear connective tissue with an implanted housing, the housing and/or the sensor element and/or a further element which is associated with the sensor element is designed to have a variable volume, extent and/or a variable surface area. 22. A housing for recording diagnostic values in the body comprising at least one sensor element which has an associated housing, wherein, in order to break open and/or to tear connective tissue with an implanted housing and with at least one sensor element, the at least one sensor element and/or an additional element can be moved with respect to the housing. 23. The appliance as claimed in at least one of claims 20 to 22, wherein, before recording a diagnostic value of the body, in particular of the blood, the volume, extent and/or surface area of the housing can be varied. 24. The appliance as claimed in claim 23, wherein the at least one sensor element and/or the at least one element is associated with the housing so that it can be moved, in particularly such that it can be moved linearly and/or can be pivoted. 25. The appliance as claimed in claim 23, wherein the at least one sensor element and/or at least one element can be moved with respect to the connective tissue with respect to the housing by means of at least one actuator immediately before recording of diagnostic values of the body, in particular of the blood, and the connective tissue can be torn open or torn in order to allow bodily fluid to come close to the area of the sensor element in order to determine the glucose or blood-sugar level. 26. The appliance as claimed in claim 23, wherein the housing is formed from at least two housing parts which are positioned coaxially one inside the other. 27. The appliance as claimed in claim 26, wherein at least one housing part has at least one associated sensor element for recording diagnostic values of the body, in particular of the blood. 28. The appliance as claimed in claim 26, wherein the housing parts which are positioned coaxially one inside the other can be moved linearly with respect to one another and/or can be rotated radially with respect to one another by means of at least one actuator, in particular a linear actuator. 29. The appliance as claimed in claim 26, wherein the housing parts are formed from a biocompatible material, or are coated with this material. 30. The appliance as claimed in claim 26, wherein a motor unit and/or a gearbox unit and/or a control unit with a spindle connected to them or it are or is provided in one housing part, with the second housing part engaging with the spind... | WITTENSTEIN AG | A61B005/00
US20080281176 | A1 | 2008-11-13 | METHOD OF MAKING AN ELECTROCHEMICAL SENSOR | 2007-10-30 | A process for the manufacture of small sensors with reproducible surfaces, including electrochemical sensors. One process includes forming channels in the surface of a substrate and disposing a conductive material in the channels to form an electrode. The conductive material can also be formed on the substrate by other impact and non-impact methods. In a preferred embodiment, the method includes cutting the substrate to form a sensor having a connector portion and a transcutaneous portion, the two portions having edges that define one continuous straight line. | 1. A sensor for the in vitro determination of the presence and/or level of glucose in a sample, the sensor comprising: a substrate; anda code on the substrate comprising sensor information,wherein the sensor comprises a working electrode, at least one other electrode, and a sample chamber to determine the presence and/or level of glucose in a sample applied to the sensor. 2. The sensor of claim 1, wherein the sensor information comprises production batch information. 3. The sensor of claim 1, wherein the sensor information comprises sensor chemistry information. 4. The sensor of claim 1, wherein the sensor information comprises calibration information. | THERASENSE, INC. | A61B005/1473
US20080281210 | A1 | 2008-11-13 | Arterial pressure sensing device | 2008-01-25 | An implantable medical device including a tubular housing defining a passage between a proximal end and a distal end of the housing. The passage providing fluid communication through the housing. A sensing unit is positioned within the passage and coupled to the housing. The sensing unit is configured to sense at least one of a physical, chemical, and physiological parameter within the passage. | 1. An implantable medical device comprising: a tubular housing defining a passage between a proximal end and a distal end of said housing, said passage providing fluid communication through said housing; anda sensing unit positioned within said passage and coupled to said housing, said sensing unit configured to sense at least one of a physical, chemical, and physiological parameter within said passage. 2. An arterial pressure sensor device in accordance with claim 1 wherein said housing further comprises: a distal retainer coupled to said housing at a midsection of said housing; anda proximal retainer initially coupled to said housing at said proximal end, with said housing positioned within a tissue, said proximal end contacts a skin surface and said distal retainer contacts a vessel wall to retain said housing positioned within the tissue. 3. An implantable medical device in accordance with claim 1 wherein said housing is affixed within a tissue using a biocompatible adhesive including one of an acrylic-based adhesive, an epoxy-based adhesive, a polyurethane-based adhesive, and a silicon-based adhesive. 4. An implantable medical device in accordance with claim 1 wherein said sensing unit comprises at least one sensor coupled to said housing. 5. An implantable medical device in accordance with claim 4 wherein said at least one sensor comprises one of a pressure sensor, an optical sensor, a biochemical sensor, a protein sensor, a motion sensor, an accelerometer, a gyroscope, a temperature sensor, a chemical sensor, a pH sensor, and a genetic sensor. 6. An implantable medical device in accordance with claim 4 wherein said at least one sensor is fabricated using a microelectromechanical systems (MEMS) technology. 7. An implantable medical device in accordance with claim 1 further comprising a void formed in an inner surface of said housing, said void configured to receive a corresponding sensor of said at least one sensor. 8. An implantable medical device in accordance with claim 7 wherein said sensor is coupled within said corresponding void to said inner surface using one of an acrylic-based adhesive, an epoxy-based adhesive, a polyurethane-based adhesive, and a silicon-based adhesive. 9. An implantable medical device in accordance with claim 7 wherein said at least one sensor is coupled to said housing using one of a chemical bonding process, a heat bonding process, a soldering process, a suturing process, and an outer packaging material. 10. An implantable medical device in accordance with claim 1 wherein said sensing unit is configured to facilitate obtaining data for at least one of arterial pressure analysis, temperature analysis, blood chemical analysis, blood osmolar analysis, and cellular count analysis, said sensing unit further configured to generate and transmit measurement data wirelessly to an external receiver. 11. An implantable medical device in accordance with claim 1 wherein said sensing unit further comprises a subst... | - | A61B005/0215
US07768261 | B2 | 2010-08-03 | Apparatus and method for image alignment for combined positron emission tomography (PET) and magnetic resonance imaging (MRI) scanner | 2008-04-25 | A phantom and method are provided for co-registering a magnetic resonance image and a nuclear medical image. The phantom includes a first housing defining a first chamber configured to receive a magnetic resonance material upon which magnetic resonance imaging can be performed in order to produce the magnetic resonance image. The phantom also includes three or more second housings configured to be attached to the first housing, where the second housings each define a second chamber configured to receive a radioactive material upon which nuclear imaging can be performed in order to produce the nuclear medical image and upon which the magnetic imaging can be performed in order to produce the magnetic resonance image. The first chamber has a volumetric capacity that is larger than a volumetric capacity of each second chamber. | 1. A phantom for co-registering a magnetic resonance image and a nuclear medical image, said phantom comprising: a first housing defining a first chamber configured to receive a magnetic resonance material upon which magnetic resonance imaging can be performed in order to produce the magnetic resonance image; andthree or more second housings configured to be attached to said first housing, wherein said second housings each define a second chamber configured to receive a radioactive material upon which nuclear imaging can be performed in order to produce the nuclear medical image and upon which the magnetic imaging can be performed in order to produce the magnetic resonance image,wherein said first chamber has a volumetric capacity that is larger than a volumetric capacity of each second chamber. 2. The phantom according to claim 1, wherein: said first housing has a first centroid;said second housings each define a respective second centroid; andsaid second housings are each configured to be attached to said first housing such that each second centroid is located at a same distance from said first centroid. 3. The phantom according to claim 2, wherein said second housings are each configured to be attached to said first housing such that each second centroid is located along a separate three-dimensional axis centered at said first centroid. 4. The phantom according to claim 1, wherein: said first housing is spherical; andsaid second housings are each spherical. 5. The phantom according to claim 1, wherein: said first chamber contains at least 200 ml of a silicone containing hydrogen as the magnetic resonance material; andeach second housing contains at least 0.5 ml of a radioactive silicon germanium mixture as the radioactive material. 6. The phantom according to claim 5, wherein said radioactive silicon germanium mixture in each second chamber has a specific activity of at least 0.3 mCi. 7. The phantom according to claim 1, wherein each second housing is detachably attached to said first housing. 8. The phantom according to claim 7, wherein: each second housing is mounted to a rod portion;said first housing has three or more tubular portions, said tubular portions each being configured to receive a respective rod portion of each second housing; andsaid tubular portions and respective rod portions being configured to threadedly engage with each other. 9. The phantom according to claim 1, wherein: said first housing has a first opening with a first valve that is configured to regulate a flow of the magnetic resonance material into and out of said first chamber; andsaid second housings each have a second opening with a second valve that is configured to regulate a flow of the radioactive material into and out of a respective second chamber. 10. A method of co-registering a magnetic resonance image and a nuclear medical image, said method comprising: placing a phantom in a patient scanning space, the phantom including a first housing definin... | Siemens Medical Solutions USA, Inc. | G01V003/00
US07460894 | B2 | 2008-12-02 | Biological optical measuring instrument | 2002-07-19 | A biomedical optical measurement apparatus comprising a light source unit for generating an inspection light containing multiple lights modulated at different frequencies, a light-receiving unit for receiving the light generated at said light source unit and passing through an object to be examined and for outputting the electric signals with the intensity corresponding to the received inspection light, and a detection means for detecting a signal with the same frequency of the reference signal in the output from said light-receiving unit. The detection means comprises an analog-digital conversion means for outputting digitized data by converting an input signal to a digital signal, a storage means for storing digitized data of multiple reference signals, a digital multiplication means for multiplying digitized data of input signals outputted from the analog-digital converting means by the digitized data of the reference signals read out from the storage means and for outputting the product of multiplication, and a digital band-limitation means for taking out DC data from the output from the digital multiplication means. Reference signal generating circuits of a number equal to that of frequencies of detected signals, which has been necessary for the conventional instrument, can be replaced by a single memory means and the configuration of the instrument can be simplified. Changes in frequency can be easily coped with by only re-writing the data of the storing means. | 1. A biomedical optical measurement apparatus comprising: a light source means for generating an inspection light containing multiple lights modulated at a given number of different frequencies,a light-receiving means for receiving the light generated at said light source means and passed through an object to be examined and for outputting synthesized signals made of a given number of electric signals each having a different frequency and an intensity corresponding to the received light, anda detection means for detecting a signal with the same frequency as a frequency of a reference signal having a frequency corresponding to frequencies of modulated multiple lights in the synthesized signals, whereinsaid detection means comprises:an analog-digital conversion means for outputting digitized synthesized signal data by converting said synthesized signals having different frequencies to a digital signal,a storage means for storing digitized data of a number of reference signals, where the number of reference signals is equal to the number of the synthesized signals output from the light-receiving means,a delaying means for delaying said reference signals data read out from the storage means based on phase differences between said synthesized signals data output from said analog-digital conversion means and said reference signals data read out from the storage means,a digital multiplication means for multiplying said synthesized signals data output from said analog-digital conversion means by said reference signals data delayed by the delaying means and for outputting the product of multiplication, anda digital band-limitation means for taking out DC data from the output from said digital multiplication means. 2. The biomedical optical measurement apparatus of claim 1, wherein said digital multiplication means and said digital band limitation means are composed by a digital signal processor. 3. The biomedical optical measurement apparatus of claim 1, further comprising a function generating means for inputting the digitized data of said reference signals and for generating a function that becomes “0 ” or “close to 0 ” near a level changing point of the digitized data. 4. The biomedical optical measurement apparatus of claim 1, wherein a digital signal processor including said digital multiplying means and said digital band-limitation means is provided for each modulation frequency. 5. The biomedical optical measurement apparatus of claim 4, wherein the digital multiplying means further comprises: a first digital multiplication means for multiplying digitized data of input signals outputted from said analog-digital conversion means by the digitized data of the reference signals read out from said storage means and for outputting the product of multiplication, anda second digital multiplication means for multiplying the output from said first digital multiplication means by a low-pass filter coefficient and for outputting the product of multiplicati... | Hitachi Medical Corporation | A61B005/1455
US07440797 | B2 | 2008-10-21 | Abdominal impedance based body composition meter | 2007-08-01 | An abdominal impedance based body composition meter is provided for estimating a body composition indicator with high accuracy. In one embodiment, the body composition meter measures pressing forces while load detection sections 3a-3d are pressed against the abdomen and abdominal impedances are measured by electrodes 2a-2d. Body composition indicators are estimated based on the measured pressing forces and the abdominal impedances. | 1. An abdominal impedance based body composition meter comprising: abdominal impedance measuring means;pressing force measuring means; andbody composition estimating means;wherein the abdominal impedance measuring means comprises a plurality of electrodes and a support which keeps the electrodes pressed against the abdomen and measures an abdominal impedance when the electrodes are in contact with the abdomen,wherein the pressing force measuring means measures a pressing force when the electrodes are pressed against the abdomen by the support,wherein the body composition estimating means estimates a body composition indicator based on the abdominal impedance measured by the abdominal impedance measuring means and the pressing force measured by the pressing force measuring means,wherein the electrodes comprise a pair of current-passing electrodes and a pair of measuring electrodes,wherein the pressing force measuring means comprises load detection sections which detect loads produced when the electrodes are pressed against the abdomen and determines a pressing force applied to the current-passing electrodes and a pressing force applied to the measuring electrodes based on the loads detected by the load detection sections,wherein the abdominal impedance measuring means determines an abdominal impedance based on a potential difference which is detected between the measuring electrodes when a low-frequency current is passed between the current-passing electrodes and determines an abdominal impedance based on a potential difference which is detected between the measuring electrodes when a high-frequency current is passed between the current-passing electrodes, and determines an additional abdominal impedance,wherein the body composition estimating means computes and estimates an abdominal total fat area as the body composition indicator by substituting the abdominal impedance determined by the abdominal impedance measuring means by passing the low-frequency current into Zlow, the additional abdominal impedance determined by the abdominal impedance measuring means by passing the high-frequency current into Zhigh, the pressing force applied to the current-passing electrodes which has been determined by the pressing force measuring means into SIavg, the pressing force applied to the measuring electrodes which has been determined by the pressing force measuring means into SVavg, and constants which provide a correlation with an abdominal total fat area determined by a standard technique into a, b, c and d, in an arithmetic expression represented by abdominal total fat area=a×Zlow+b×Zhigh+c×SIavg+d×SVavg. 2. An abdominal impedance based body composition meter comprising: abdominal impedance measuring means;pressing force measuring means; andbody composition estimating means;wherein the abdominal impedance measuring means comprises a plurality of electrodes and a support which keeps the electrodes pressed against the abdomen and measures an abdominal imped... | Tanita Corporation | A61B005/05
US08010183 | B2 | 2011-08-30 | Detecting system of position and posture of capsule medical device | 2008-03-25 | A detecting system of the position and the posture of a capsule medical device includes a main body of the capsule medical device that is inserted in a living body, a coil in a capsule that is arranged to the main body of the capsule medical device and forms a resonant circuit, a magnetic field generating device that is arranged around the living body and generates the Alternating magnetic field for generating an induced magnetic field in the coil in the capsule, and a plurality of magnetic field detecting devices that detect the strength of the induced magnetic field generated by the coil in the capsule by the magnetic field generated by the magnetic field generating device. | 1. A detecting system of the position and the posture of a capsule medical device, comprising a posture detecting system for detecting the posture of a main body of the capsule medical device: the posture detecting system comprising: a coil in the capsule provided to the main body of the capsule medical device;a magnetic field generating device for generating an alternating magnetic field for causing the coil in the capsule to generate an induced electromotive force;a magnetic field detecting device for detecting the strength of a magnetic field from the coil in the capsule;a calculation processing unit for calculating the direction and the position of the coil in the capsule in response to a detection value from the magnetic field detecting device; anda control device for calculating an angle formed by the direction of the coil in the capsule calculated by the calculation processing unit and the direction of the alternating magnetic field generated by the magnetic field generating device, andcontrolling the magnetic field generating device such that the calculated angle is not greater than a set value. 2. The detecting system of the position and the posture of a capsule medical device according to claim 1, wherein the control device controls such that the direction of the magnetic field generated by the magnetic field generating device matches the direction of the coil in the capsule when the angle formed by the direction of the coil in the capsule calculated by the calculation processing unit and the direction of the alternating magnetic field generated by the magnetic field generating device is greater than the set value. 3. The detecting system of the position and the posture of a capsule medical device according to claim 1, wherein the control device controls such that the direction of the magnetic field generated by the magnetic field generating device matches the direction of the coil in the capsule. 4. The detecting system of the position and the posture of a capsule medical device according to claim 1, wherein the magnetic field generating device generates the magnetic field in three perpendicular axial directions. 5. The detecting system of the position and the posture of a capsule medical device according to claim 1, wherein the magnetic field generating device includes at least one or more sets of facing coils. 6. The detecting system of the position and the posture of a capsule medical device according to claim 1, wherein the magnetic field generating device includes two or three sets of facing coils arranged to be orthogonal to one another. 7. The detecting system of the position and the posture of a capsule medical device according to claim 6, wherein the magnetic field detecting device includes two or three detecting coil arrays, with each detecting coil array corresponding to one of the two or three sets of facing coils comprising the magnetic field generating device. 8. The detecting system of the position and the posture... | Olympus Corporation | A61B005/05
US08267863 | B2 | 2012-09-18 | Procedure and system for monitoring a physiological parameter within an internal organ of a living body | 2008-04-29 | A system and surgical procedure for monitoring physiological parameters within an internal organ of a living body. The procedure entails making a first incision to expose the organ and a second incision through an external wall of the organ and into an internal cavity. A sensing unit is placed in the second incision such that a proximal end thereof remains outside the organ. The unit includes a sensing device having a sensing element for sensing the physiological parameter within the organ, and an anchor to which the sensing device is secured. The unit occludes the second incision and a distal end of the unit does not extend more than one centimeter into the cavity. The anchor is then secured to the wall of the organ, after which the first incision is closed and a readout device telemetrically communicates with the sensing device to obtain a reading of the physiological parameter. | 1. A surgical procedure comprising: making a first incision in a living body to expose at least a portion of an internal organ of the living body, the organ having an internal cavity within which an internal physiological parameter is desired to be sensed;making a second incision through an external wall of the organ and into the internal cavity;placing a sensing unit in the second incision such that a proximal end of the sensing unit at a first extremity thereof is outside the organ and a distal end of the sensing unit at an oppositely-disposed second extremity thereof is within the internal cavity of the organ, the sensing unit consisting essentially of a sensing device having a sensing element adapted to sense the physiological parameter within the organ and an anchor to which the sensing device is secured, the anchor comprising first and second portions, the first portion of the anchor comprising a tubular housing in which the sensing device is housed and an open end adjacent the distal end of the sensing unit through which the sensing device is received into the tubular housing and through which the sensing element is exposed, either the first portion of the anchor or the sensing device defining the distal end of the sensing unit, the second portion of the anchor projecting radially from and surrounding the first portion, defining the proximal end of the sensing unit, and comprising means for securing the anchor to an external wall of the organ by engaging a proximal surface of the external wall disposed nearer the proximal end of the sensing unit and not engaging a distal surface of the external wall disposed nearer the distal end of the sensing unit, the sensing unit being placed in the second incision so that the second portion of the anchor, the securing means thereof, and the proximal end of the sensing unit defined thereby are outside the organ, and so that only the first portion of the anchor and the distal end of the sensing unit extend entirely through the external wall, the distal end of the sensing unit does not extend more than one centimeter into the internal cavity of the organ, and the first portion of the anchor occludes the incision; andsecuring the anchor to the external wall of the organ such that the sensing device is secured within the second incision by only the anchor and the second incision is occluded by only the first portion of the anchor;closing the first incision; and thentelemetrically communicating with the sensing device to obtain a reading of the physiological parameter using a readout device located outside the living body. 2. The surgical procedure according to claim 1, wherein the organ is chosen from the group consisting of a heart, blood vessel, meninges, kidney, lung, bladder, and abdominal wall. 3. The surgical procedure according to claim 1, wherein the sensing device has a distal end, the sensing element is disposed at the distal end of the sensing device, and the distal end of the sensing device i... | Integrated Sensing Systems, Inc. | A61B005/07
US08597188 | B2 | 2013-12-03 | Health management devices and methods | 2008-06-20 | Methods and devices and systems including a data collection module for receiving and storing analyte data over a predetermined time period from a subject, a user interface unit coupled to the data collection module for providing one or more indication related to the analyte data, a control unit coupled to the data collection module and the user interface unit to control, at least in part the operation of the data collection module and the user interface unit, a communication module coupled to the control unit for communicating one or more signals associated with the analyte data to a remote location, where the user interface unit is configured to operate in a prospective analysis mode including substantially real time output of the analyte level received by the data collection module, and a retrospective analysis mode including limited output of information to the subject during the predetermined time period, and further where the communication module is configured to communicate with the remote location after the analyte data is received and stored in the data collection module over the predetermined time period, are provided. | 1. An analyte monitoring device, comprising: a data collection module for receiving and storing analyte data over a predetermined time period;a user interface unit coupled to the data collection module for providing one or more indication related to the analyte data;a control unit coupled to the data collection module and the user interface unit, and configured to control, at least in part an operation of the data collection module and the user interface unit; anda communication module coupled to the control unit for communicating one or more signals associated with the analyte data to a remote location; wherein the user interface unit is configured to operate in a prospective analysis mode including substantially real time output of an analyte level received by the data collection module, and a retrospective analysis mode including limited output of information during the predetermined time period;wherein the communication module is configured to communicate with the remote location after the analyte data is received and stored in the data collection module over the predetermined time period; andwherein the data collection module is operatively coupled to a transcutaneously positioned analyte sensor in fluid contact with an interstitial fluid under a skin layer, the analyte sensor generating signals corresponding to a monitored analyte level over the predetermined time period, wherein when the user interface unit is operating in the retrospective analysis mode, the user interface unit is configured to not display any information related to the received analyte data during the predetermined time period. 2. The device of claim 1 wherein the communication module is configured to communicate with the remote location using one or more of a USB cable connection, a serial cable connection, a radio frequency (RF) communication protocol, an infrared communication protocol, a Bluetooth communication protocol, or an 802.11x communication protocol. 3. The device of claim 1 wherein the limited output of information during the retrospective analysis mode includes output to the user interface unit of one or more of the analyte monitoring device operational status information, a time of day information, a user profile information, or an elapsed duration of the predetermined time period. 4. The device of claim 1 wherein the data collection module includes one or more of a data storage device or a memory device. 5. The device of claim 1 wherein the data collection module is configured to delete the stored analyte data after transferring the analyte data to the remote location. 6. The device of claim 1 wherein the remote location includes a data processing terminal. 7. The device of claim 1 wherein during the prospective analysis mode, the user interface unit is configured to visually output real time information related to the received analyte data. 8. The device of claim 7 wherein the visual output includes one or more of a graphical output, a numerical o... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/05 A61B005/145 A61B005/1473 A61B005/1486
US08649855 | B2 | 2014-02-11 | Burst suppression monitor for induced coma | 2007-10-01 | Methods and systems for monitoring subjects, analyzing EEG data, and maintaining subjects at desired sedation states. | 1. A method of monitoring a sedated subject using a logic processor in a monitoring system comprising: receiving digital data representing said subject's EEG signal at said processor;using said processor to determine one or more burst onsets;using said processor to determine one or more suppression onsets;using said processor to determine one or more interburst or burst suppression intervals by comparing one or more burst onsets to one or more previous suppression onsets;using said processor to determine a true interburst interval or true mean burst suppression (BS) interval for a specified time length with statistical confidence wherein said mean burst suppression interval indicates time between burst onsets and previous suppression onsets, regardless of burst duration;using said mean burst suppression interval to inform therapy for said subject to maintain said desired state of sedation;calculating an absolute amplitude difference across said EEG signal, with an absolute amplitude difference history interval;smoothing said absolute amplitude difference to minimizes effects of signal variation on onset detection; anddetecting burst onset when an absolute smoothed amplitude difference exceeds a median absolute amplitude times a specified multiplication factor. 2. The method of claim 1 further wherein: said using said mean burst suppression interval to inform therapy comprising presenting data indicating said mean burst suppression interval with statistical confidence to a user, thereby providing the user information to guide adjusting therapy to a subject to maintain the desired state of sedation. 3. The method of claim 1 further wherein: said using said mean burst suppression interval to indicate therapy comprising using a processor controlled medical device to adjust administration of one or more therapies using said mean burst suppression (BS) interval. 4. The method of claim 1 further wherein: said subject is being maintained in an induced coma and said method is used by an automated system to prevent over-sedation or under-sedation. 5. The method of claim 1 further comprising: using said processor to perform statistical analysis to correct for noise variability, other isoelectric variability, or other electric disturbances in said EEG. 6. The method of claim 1 further comprising: presenting data to a user indicating a true interburst interval with statistical accuracy. 7. The method of claim 1 further wherein: said digital data representing said subject's EEG is a stream of digital samples of amplitude values at different data capture time points;and determining burst or suppression onset further by a method comprising:computing changes in amplitude substantially at each data capture point;using said processor to perform a smoothing operation at each data capture point; andusing said processor to perform threshold differentiation at each data capture point. 8. The method of claim 7 further comprising: using said processor to determine... | The Regents of the University of California Wavestate, Inc. | A61B005/04
US08821587 | B2 | 2014-09-02 | Apparatus for control of a prosthetic | 2008-02-06 | A prosthetic device control apparatus includes at least one sensor worn by a user. The sensor(s) determines a user's movement. A control module is in communication with the sensor(s). The control module communicates movement information to a prosthetic. A method for controlling a prosthetic device includes sensing a user's movement, communicating the movement through a control module to a prosthetic device; and controlling the movement of a prosthetic device. | 1. A prosthetic device control apparatus comprising: at least one inertial sensor adapted to be worn by a user, the at least one inertial sensor detecting body movements that are indicative of a desired movement of a prosthetic device; anda control module in communication with the at least one sensor and receiving at least one signal based on the detected body movements therefrom, the control module having a plurality of user selectable control modes, each user selectable control mode defining a different motion for the prosthetic device, the plurality of user selectable control modes including at least a bulk control mode for moving the prosthetic device into a desired position and a finesse control mode for manipulating a prosthetic hand of the prosthetic device once the prosthetic device is in the desired position;wherein the control module commands a first plurality of prosthetic actuators to control shoulder movement, elbow movement and humeral rotation of the prosthetic device based on the detected body movements indicated by the at least one signal from the at least one inertial sensor to achieve bulk movement of the prosthetic device when the bulk control mode has been selected, thereby moving the prosthetic device into the desired position through a combination of shoulder movement, elbow movement and humeral rotation of the prosthetic device; andwherein the control module commands a second plurality of prosthetic actuators that is different than the first plurality of actuators to control wrist movement and hand movement of the prosthetic device based on a signal from the same at least one inertial sensor used to control the first plurality of actuators to achieve finesse movement of the prosthetic device when the finesse control mode has been selected, thereby manipulating the prosthetic hand through a combination of wrist movement and hand movement of the prosthetic device;wherein the at least one inertial sensor detects at least one of x, y or z plane positioning. 2. The prosthetic device control apparatus according to claim 1, wherein the control module is in wireless communication with the at least one inertial sensor. 3. The prosthetic device control apparatus according to claim 1, wherein the control module is in wireless communication with the prosthetic device. 4. The prosthetic device control apparatus according to claim 1, wherein the control module controls movement of the prosthetic device to be substantially proportional to the body movement detected by the at least one sensor. 5. The prosthetic device control apparatus according to claim 1 wherein the at least one inertial sensor is an acceleration gyro. 6. The prosthetic device control apparatus according to claim 1 wherein the at least one inertial sensor is an accelerometer. 7. The prosthetic device control apparatus according to claim 1 wherein the at least one inertial sensor wirelessly communicates information to said control module. 8. A prosthetic device co... | DEKA Products Limited Partnership | A61F002/70 A61B005/103 A61F002/68 A61F002/54 A61F002/72 A61F002/74 A61F002/76 A61F002/78 A61B005/0488 A61B005/00
US08038615 | B2 | 2011-10-18 | Inferring wellness from physiological conditions data | 2007-05-22 | A method for inferring the physiological condition an individual using a physiological monitoring system including providing a set of personal data regarding the individual, the personal data including wellness parameter data or semantic data or both, wherein the wellness parameters correspond to known physiological attributes for the individual; acquiring physiological data of the individual and associated capture condition data, during multiple capture events; calculating wellness parameter data for the individual, using the physiological data and the capture condition data; analyzing the wellness parameter data to identify changes in the wellness parameter data; and inferring the physiological condition of the individual from the changes in the wellness parameter data in view of the semantic data. | 1. A method for automatically inferring one or more physiological conditions of an individual using a physiological monitoring system comprising: (a) operating a physiological monitoring system, including an inference engine, an image processor, and an image capture device to automatically assess physiological conditions of a subject individual from a captured image of the subject individual, where the physiological monitoring system is operated in a fixed capture area and the image capture device automatically acquires color images of a plurality of individuals, where the subject individual resides in a family unit of individuals that live together in a common residence;(b) automatically analyzing the captured images using the image processor to identify an individual as a subject individual being monitored by the system;(c) automatically analyzing the captured images using the image processor to derive both physiological data of the individual and associated image capture condition data during multiple image capture events, wherein the image capture condition data includes ambient lighting condition data;(d) automatically associating a set of personal data regarding the subject individual to the subject individual identified by the system, the personal data including wellness parameter data or semantic data or both, wherein the wellness parameters correspond to known physiological attributes for the individual;(e) automatically analyzing the physiological data or image capture condition data using the image processor to determine acceptability, where acceptability of the capture condition data is determined relative to various criteria, including those related to ambient light conditions;(f) calculating wellness parameter data for the subject individual, using the acceptable physiological data and the acceptable image capture condition data;(g) analyzing the wellness parameter data to identify changes in the wellness parameter data; and(h) automatically inferring physiological condition of the subject individual, using the inference engine, from the changes in the wellness parameter data, in view of the semantic data or the image capture condition data associated with the subject individual, and also providing comparative inferences of the physiological condition of the subject individual relative to wellness parameter data, physiological conditions, including family health history, or semantic data of other individuals in the family unit. 2. The method according to claim 1 wherein the method for inferring further includes automatic testing for false positives with respect to the physiological condition of the subject individual, so as to determine whether inferred changes in one or more physiological conditions are attributable to image capture conditions, or changes thereof, rather than actual physiological changes. 3. The method according to claim 2 wherein the test for false positives includes evaluating changes in image capture condition... | Eastman Kodak Company | A61B005/00
US20080269617 | A1 | 2008-10-30 | Absorption and Scattering Map Reconstruction For Optical Fluorescence Tomography | 2006-10-24 | Optical fluorescence tomography is a highly sensitive method to image contrast agents in the body. However, current reconstruction methods suffer from a high complexity or even instability. According to an exemplary embodiment of the present invention, a method for absorption/scattering map reconstruction for optical fluorescence tomography may be provided, which uses a spectral model. This may provide for a subsequent fluorescence reconstruction with improved image quality. | 1. An optical fluorescence tomography apparatus (100) for examining an object of interest (101), the optical fluorescence tomography apparatus (100) comprising: a detector unit (107) adapted for detecting first light transmitted through the object of interest (101) and second light emitted by a contrast agent inside the object of interest (101), resulting in detection data; anda reconstruction unit (108) adapted for performing a fluorescence reconstruction on the basis of the detection data and a spectral model, resulting in reconstruction data comprising a spatial distribution of the contrast agent inside the object of interest (101). 2. The optical fluorescence tomography apparatus of claim 1, wherein the spectral model comprises a first concentration of a contrast agent. 3. The optical fluorescence tomography apparatus of claim 1, wherein the spectral model comprises an absorption model and a scattering model; wherein the reconstruction unit (108) is further adapted for: determining a spatial distribution of absorption at a fluorescence wavelength on the basis of the absorption model; anddetermining a spatial distribution of scattering at the fluorescence wavelength on the basis of the scattering model. 4. The optical fluorescence tomography apparatus of claim 3, wherein the absorption model readsμa(λ,x⇀)=∑i=1NCi(x⇀)ɛ(i,λ)wherein λ is a wavelength, N is a number of chromophores in the model, Ci is a concentration of the chromophore i, and ε(i, λ) is an absorption of chromophore i at wavelength λ. 5. The optical fluorescence tomography apparatus of claim 3, wherein the scattering model reads μ′s(λ, {right arrow over (x)})=A({right arrow over (x)})λB({right arrow over (x)}) wherein A is a scatter amplitude and B is a scatter power. 6. The optical fluorescence tomography apparatus of claim 2, wherein the spectral model comprises a second concentration of oxygenated haemoglobin, a third concentration of deoxygenated haemoglobin, and a fourth concentration of water. 7. The optical fluorescence tomography apparatus of claim 1, wherein the detector unit (107) comprises detecting elements; wherein each detecting element is adapted for detecting the transmitted or emitted light in an spectrally resolved manner. 8. The optical fluorescence tomography apparatus of claim 1, further comprising an excitation source (102) adapted for emitting electromagnetic radiation to the object of interest (101). 9. The optical fluorescence tomography apparatus of claim 1, wherein the excitation source (102) and the detector unit (107) are adapted for moving around the object of interest (101). 10. An image processing device for examination of an object of interest (101), the image processing device comprising: a memory for storing detection data of the object of interest (101), the detection data comprising detected first light transmitted through the object of interest (101) and detected second light emitted by a contrast agent inside the object of inter... | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61B005/00
US20080243007 | A1 | 2008-10-02 | Pulmonary Artery Pressure Signals And Methods of Using | 2007-03-28 | Embodiments of the invention are related to methods and systems for using a pulmonary artery pressure signal to detect and/or monitor physiological parameters, physiological status, and aspects of disorders and diseases, amongst other things. In an embodiment, the invention includes a method for detecting pulmonary symptoms of a disorder. In an embodiment, the invention includes a method for detecting a pathological change to a tissue, structure, or fluid volume in or around the lung. In an embodiment, the invention includes a method for detecting a disorder affecting airflow. Other aspects and embodiments are provided herein. | 1. A method for detecting pulmonary symptoms of a disorder comprising: chronically implanting a pulmonary artery pressure sensor;obtaining a pulmonary artery pressure signal from the pulmonary artery pressure sensor; andmonitoring the pulmonary artery pressure signal to identify a change in the signal over a baseline value. 2. The method of claim 1, the disorder associated with pathological pulmonary structural changes. 3. The method of claim 1, the change including an increase in the pulmonary artery pressure signal exceeding a threshold amount. 4. The method of claim 1, the change including a decrease in the pulmonary artery pressure signal exceeding a threshold amount. 5. The method of claim 1, the change persisting for a period of time exceeding a threshold amount. 6. The method of claim 5, the threshold amount greater than about one minute. 7. The method of claim 1, further comprising converting the pulmonary artery pressure signal into a respiration signal. 8. The method of claim 7, further comprising monitoring the respiration signal for changes in respiration rate. 9. The method of claim 7, further comprising monitoring the respiration signal for changes in tidal volume. 10. The method of claim 1, further comprising monitoring the respiration signal for changes consistent with a condition selected from the group consisting of pulmonary edema, pulmonary embolism, pleural effusion, pulmonary arteriovenous malformation (PAVM), indicative of combined obstructive pulmonary disease (COPD), emphysema, and asthma. 11. A method for detecting a pathological change to a tissue, structure, or fluid volume in or around the lung, the method comprising: establishing a baseline pulmonary artery pressure signal with a pressure sensor; andmonitoring the pulmonary artery pressure signal to identify a change in the pulmonary artery pressure signal compared to the baseline signal. 12. A method for detecting a disorder affecting airflow comprising: chronically implanting a pulmonary artery pressure sensor;obtaining a pulmonary artery pressure signal from the pressure sensor; andmonitoring the pulmonary artery pressure signal to identify a respiration pattern consistent with the disorder. 13. The method of claim 12, the disorder selected from the group consisting of snoring, apnea, Cheyne-Stokes syndrome, hypopnea, hyperpnea, tachypnea, and dyspnea. 14. The method of claim 12, the disorder comprising a central sleep apnea. 15. The method of claim 12, the disorder comprising an obstructive sleep apnea. 16. The method of claim 12, the respiration pattern characterized by a plurality of apneas. 17. The method of claim 12, the respiration pattern characterized by a plurality of hypopneas. 18. The method of claim 12, the respiration pattern characterized by a plurality of hyperpneas. 19. The method of claim 12, the respiration pattern characterized by a five or more obstructed breathing events per hour, the obstructed breathing event selected from ... | CARDIAC PACEMAKERS, INC. | A61B005/08 A61B005/00
US20080312559 | A1 | 2008-12-18 | Method and Apparatus for Performing Gastric Bypass Surgery | 2008-06-11 | A specially configured esophageal-gastric tube can be inserted into a patient's esophagus with the tip of the tube in the region of an end-to-end anaotomosis connection. The joined members are occluded upstream and downstream of the anastomosis connection. Fluid is pumped into the occluded section through the esophageal-gastric tube until a desired pressure level is attained. The anastomosis connection then is checked for leaks. The invention also includes a bougie having an end portion made of a material suitable for resisting electro-cauterization temperatures. | 1. Surgical apparatus suitable for performing gastric bypass surgery, comprising: a first tube made of a flexible material that defines a first passageway, the first tube having a tip end, a central portion, and an entrance/exit end, the tip end being provided with openings that establish fluid communication with the first passageway;a first syringe that can be attached to the entrance/exit end;an expansible balloon disposed adjacent the openings in the tip end and surrounding the first tube, the balloon being positioned intermediate the openings in the tip end and the central portion;a second passageway extending along at least a portion of the length of the tube, the second passageway establishing fluid communication with the balloon;a branch tube projecting from the side of the first tube in the region of the entrance/exit end, the branch tube being in fluid communication with the second passageway;a shutoff valve disposed in the branch tube; anda second syringe that can be attached to the shutoff valve. 2. The surgical apparatus of claim 1, wherein the tube includes a wall, and the second passageway is formed as part of the wall. 3. The surgical apparatus of claim 1, further comprising a second tube that is disposed within the first passageway and which defines the second passageway. 4. The surgical apparatus of claim 1, wherein the first tube is made of a material selected from the group consisting of latex, vinyl, and silicone. 5. The surgical apparatus of claim 1, further comprising a plurality of radiographic markers disposed at intervals along the length of the central portion. 6. The surgical apparatus of claim 5, wherein the radiographic markers are positioned approximately 10 cm from each other. 7. The surgical apparatus of claim 1, wherein the entrance/exit end is tapered, and further comprising an adapter having first and second ends, the first end being tapered for insertion into the entrance/exit end and the second end defining a luer slip connector to receive the first syringe. 8. The surgical apparatus of claim 1, wherein the connection between the shutoff valve and the second syringe is made by a luer lock connector. 9. Apparatus especially adapted for use in gastric bypass surgery, comprising a bougie having an elongate body portion and an enlarged tip, the tip being made of a material resistant to temperatures encountered during electro-cauterization. 10. The apparatus of claim 9, where the body portion and the tip are made of high-temperature silicone. 11. The apparatus of claim 9, wherein the tip is coated with a temperature-resisting substance such as TEFLON fluoropolymer. 12. A method of testing the integrity of an end-to-end anastomosis connection in a patient in which first and second conduits are joined, comprising the steps of: occluding the first conduit upstream of the anastomosis connection;occluding the second conduit downstream of the anastomosis connection;pumping fluid into the occluded section;incre... | - | A61B005/103 A61M029/02
US20080275337 | A1 | 2008-11-06 | BREATHING DETECTION APPARATUS AND METHOD | 2008-04-25 | A device is disclosed for remote monitoring of breathing movements of a patient, comprising a sender transmitting a microwave signal toward the patient, a receiver arranged to receive a signal reflected by the patient, and a processor. The processor mixes the received signal with the transmitted signal and with a signal in quadrature with the transmitted signal to determine a breathing frequency of the patient. In some embodiments, two signals at different frequencies are emitted toward the patient and processed in order to compensate for breathing patterns that are not uniquely resolvable at a single frequency. | 1. A device for remote monitoring of breathing movements of a patient, comprising: a transmitter transmitting a microwave signal toward the patient;a receiver arranged to receive a signal reflected by the patient; anda processor, comprising: a first mixer arranged to mix the transmitted signal and the received signal to provide a first signal component;a second mixer arranged to mix one of the received signal or the transmitted signal with a signal in quadrature with the other of the transmitted signal or the received signal to provide a second signal component; andan output device arranged to provide a signal representative of the patient's breathing movements based on the first and second signal components. 2. The device of claim 1 wherein the second mixer is operable to mix the received signal with a signal in quadrature with the transmitted signal to provide the second signal component. 3. The device according to claim 1 wherein the microwave signal has a frequency in the range of 10.6 GHz-30 GHz. 4. The device of claim 3 wherein the microwave signal has a frequency in the range of 20 GHz-26 GHz. 5. The device of claim 1 wherein the microwave signal is a non-modulated signal. 6. The device of claim 1 wherein the microwave signal is a continuous wave signal. 7. The device of claim 1 wherein the output device is arranged to provide a signal resulting from the addition of the first signal component to the second signal component. 8. The device of claim 1 wherein the transmitted signal comprises two microwave signals of different frequency and the system comprises a processor with a first mixer, a second mixer and an output device for each microwave signal. 9. The device of claim 8 wherein the output device is arranged to combine signals from the first and second mixers to provide a signal representative of the patients breathing movements. 10. The device of claim 8 wherein the microwave signals have a frequency in the range of 10.6 GHz-30 GHz. 11. The device of claim 8 wherein the microwave signals have a frequency in the range of 20 GHz-26 GHz. 12. A device for remote monitoring of breathing movements of a patient, comprising: a transmitter transmitting at least two electromagnetic signals toward the patient, the at least two electromagnetic signals having respective frequencies that are different from each other;a receiver arranged to receive a signal reflected by the patient; anda processor operable to mix the transmitted signal and the received signal to provide a first signal component, and to mix one of the received signal or the transmitted signal with a signal in quadrature with the other of the transmitted signal or the received signal to provide a second signal component; the processor further being operable to provide a signal representative of the patient's breathing movements based on the first and second signal components. 13. The device of claim 12 wherein the processor is operable to mix the received signal with a sig... | - | A61B005/055
US20080294071 | A1 | 2008-11-27 | Assay Method | 2005-05-10 | An assay for the assessment of drug side effects particularly the side effect of dizziness comprising the steps of: providing a first control animal located on abeam; inducing the control animal to traverse the beam; recording the number of footslips made by the animal during the traversal; providing a second test animal located on the beam or duplicate beam; inducing the test animal to traverse the beam; recording the number of footslips made by the animal during the traversal and determining whether there is an increase, decrease or no change in the number of footslips made by the test animal in comparison to the control animal. | 1. A method for determining the degree to which an animal experiences dizziness comprising the following steps: a) providing a first control animal located on a beam,b) inducing the control animal to traverse the beam,c) recording the number of footslips made by the animal during the traversal,d) providing a second test animal located on the beam or duplicate beam,e) inducing the test animal to traverse the beam,f) recording the number of footslips made by the animal during the traversal,g) determining whether there is an increase, decrease or no change in the number of footslips made by the test animal in comparison to the control animal. 2. A method for assaying a compound for the effect of producing dizziness in animal comprising the following steps: a) providing a control animal located on a beam,b) inducing the control animal to traverse the beam,c) recording the number of footslips made by the animal during the traversal,d) providing a test animal which has been dosed with a test compound,e) providing the dosed test animal located on the beam or duplicate beam,f) inducing the test animal to traverse the beamg) recording the number of footslips made by the animal during the traversal,h) determining whether there is an increase, decrease or no change in the number of footslips made by the test animal in comparison to the control animal. 3. The method according to claim 1 further comprising the steps of a) performing a second different test designed to measure the degree of locomotor activity for the test animal and control animalb) determining whether there is an increase decrease or no change in the degree of locomotor activity measured for the test animal in comparison to the control animal. 4. The method according to claim 3 wherein the measured locomotor activity is vertical and/or horizontal locomotion. 5. The method according to claim 3 wherein the second test is the locomotor activity test. 6. The method according to claim 1 wherein the control animal and the test animal are the same animal. 7. The method according to claim 1 wherein more than one control and/or test animal are used. 8. The method according to claim 2 further comprising the steps of a) performing a second different test designed to measure the degree of locomotor activity for the test animal and control animalb) determining whether there is an increase decrease or no change in the degree of locomotor activity measured for the test animal in comparison to the control animal. 9. The method according to claim 8 wherein the measured locomotor activity is vertical and/or horizontal locomotion. 10. The method according to claim 8 wherein the second test is the locomotor activity test. 11. The method according to claim 2 wherein the control animal and the test animal are the same animal. 12. The method according to claim 2 wherein more than one control and/or test animal are used. | - | A61B005/11
US20080269583 | A1 | 2008-10-30 | Detection and Monitoring of Stress Events During Sleep | 2006-02-07 | A computer-implemented method for monitoring a patient (22) includes receiving physiological signals from the patient during sleep and processing at least one of the signals to detect a spontaneous stress event. One or more of the signals following the stress event are analyzed so as to evaluate a stress response of the patient. | 1. A computer-implemented method for monitoring a patient, comprising: receiving a signal associated with respiration of the patient during sleep;processing the signal to detect a pattern of periodic breathing;extracting from the signal a shape characteristic of the pattern; andclassifying an etiology of the periodic breathing responsively to the shape characteristic. 2. The method according to claim 1, wherein classifying the etiology comprises determining the periodic breathing to be predominantly central or obstructive in origin. 3. The method according to claim 2, wherein extracting the shape characteristic comprises computing a parameter that is characteristic of a symmetry of the pattern. 4. The method according to claim 3, wherein determining the periodic breathing comprises classifying the periodic breathing as predominantly central in origin if the pattern is symmetrical. 5. The method according to claim 1, wherein receiving the signal comprises receiving an indication of least one of a flow through an airway of the patient, a movement of an abdomen or thorax of the patient, a heart rate of the patient, a respiration rate of the patient, a blood flow of the patient, a blood oxygen saturation level of the patient, and a pulse transit time (PTT) of the patient. 6. The method according to claim 1, wherein receiving the signal comprises receiving a photoplethysmographic signal from an organ of the patient. 7. The method according to claim 6, wherein receiving the signal comprises fastening a belt holding a reflective photoplethysmographic sensor around the patient, and measuring both a respiratory movement and a blood oxygen saturation level of the patient using the belt and photoplethysmographic sensor. 8. The method according to claim 1, wherein the pattern comprises a hyperventilation interval and a hypoventilation interval, and wherein extracting the shape characteristic comprises fitting first and second waveforms to the hyperventilation and hypoventilation intervals, respectively. 9. The method according to claim 8, wherein the first and second waveforms comprise analytical functions. 10. The method according to claim 9, wherein the analytical functions comprise polynomial waveforms. 11. The method according to claim 9, wherein the analytical functions comprise tilted parabolic waveforms. 12. The method according to claim 8, and comprising extracting features from the signal responsively to the first and second waveforms, and selecting a treatment to administer to the patient based on the extracted features. 13. A computer-implemented method for monitoring a patient, comprising: receiving physiological signals from the patient during sleep;processing at least one of the signals to detect a spontaneous stress event; andanalyzing one or more of the signals following the stress event so as to evaluate a stress response of the patient. 14. The method according to claim 13, wherein receiving the physiological signals comprises ... | WIDEMED LTD. | A61B005/087 A61B005/11 A61B005/02 A61B005/145
US20080294070 | A1 | 2008-11-27 | JUGULAR VENOUS PRESSURE GAUGE | 2007-05-27 | The present invention is directed to an apparatus for metering the jugular venous pressure of a patient, and a method thereof. The apparatus comprises: a slider which comprises a scale for indicating the venous pressure of a patient; a pointer connected to the sliding member of the slider perpendicularly to the scale, for pointing on the filling level of the jugular venous pressure of the patient. The jugular venous pressure is indicated by positioning the patient at a certain angle, such as 30 to 45 degrees with reference to the floor, placing the tip at the sternal-angle of the patient while the slider is held vertically, sliding the pointer until pointing on the top of the filling level of the jugular vein of the patient. | 1. An apparatus for metering the jugular venous pressure of a patient, the apparatus comprising: a slider comprising a scale indicating the jugular venous pressure of a patient; anda pointer, connected to a sliding member of said slider perpendicularly to said scale, for pointing on the filling level of the jugular vein of said patient. 2. An apparatus according to claim 1, wherein said scale indicates the distance between each point of said scale to a tip of said slider. 3. An apparatus according to claim 1, wherein said scale indicates normal jugular venous pressure of a patient. 4. An apparatus according to claim 1, wherein said scale indicates abnormal jugular venous pressure of a patient. 5. An apparatus according to claim 1, wherein said pointer comprises a bar. 6. An apparatus according to claim 1, further comprising a spirit level, for indicating verticality of said scale. 7. An apparatus according to claim 1, further comprising a spirit level, for indicating horizontality of said pointer. 8. An apparatus according to claim 1, further comprising a dowel and a corresponding depression, for holding said pointer perpendicularly to said scale while rendering said metering. 9. An apparatus according to claim 1, further comprising a dowel and corresponding depression, for holding said pointer parallel to said scale while said apparatus is idle. 10. An apparatus according to claim 1, wherein said pointer is foldable. 11. An apparatus according to claim 10, further comprising a slit corresponding to said pointer, said slit being used as housing for said pointer when folded. 12. An apparatus according to claim 1, further comprising a clip, for attaching said apparatus to a fabric. 13. An apparatus according to claim 1, further comprising a push button for pushing said pointer to protrude from the body of said slider. 14. An apparatus according to claim 1, wherein the length of said first bar is of pocket size. 15. A method for metering the jugular venous pressure of a patient, the method comprising the steps of: providing an apparatus comprising a first bar perpendicular to a second bar, said first bar comprising a scale indicating the distance between a tip thereof, said second bar slideable along said first bar;positioning said patient at an angle with reference to the floor; andplacing said tip at the sternal-angle of said patient while said first bar is held vertically, sliding said pointer over said first bar until reaching the top of the filling level of the jugular vein of said patient. 16. A method according to claim 15, wherein said angle is substantially 30 degrees. 17. A method according to claim 15, wherein said angle is substantially 45 degrees. 18. A method according to claim 15, wherein said apparatus further comprise a spirit level for indicating the verticality of said apparatus. 19. An apparatus for metering the jugular venous pressure of a patient, the apparatus comprising: a ruler the scale of which is adapted... | - | A61B005/103 G01C005/00
US20080275355 | A1 | 2008-11-06 | Method For Diagnosing An Infectioin Condition | 2007-05-04 | A method and an apparatus for determining an infection condition in an organism by measuring the level of nitrous oxide present in a gas sample taken from the organism. In one embodiment the nitrous oxide content of a gas sample is measured to diagnose systemic inflammatory response in a living organism. In an alternative embodiment, the nitrous oxide level of a gas sample taken from a living organism may be compared the with an expected nitrous oxide level for a healthy organism or with a prior measured nitrous oxide level of the living organism to diagnose the presence or absence of an infection in the living organism. A method and apparatus for determining response to a course of therapy is provided. The method and apparatus compares the nitrous oxide levels of a living organism before and after the administration of a therapy to the living organism to determine a response to the therapy. | 1. A method for diagnosing systemic inflammatory response in a living organism, comprising: collecting a first gas sample from the living organism;measuring the nitrous oxide content of the first gas sample to acquire a first measured nitrous oxide value; andcomparing first measured nitrous oxide value to a nitrous oxide reference value typical for a healthy and similar living organism to determine the presence of systemic inflammatory response. 2. A method for diagnosing sepsis comprising a method according to claim 1. 3. The method of claim 1 wherein the gas sample is either nasally or orally expired. 4. The method of claim 3 wherein the nitrous oxide is endogenously present in the gas sample. 5. The method of claim 1 wherein measuring the nitrous oxide content of the gas sample comprises illuminating the gas sample. 6. The method of claim 5 wherein illuminating the gas sample further comprises passing a light beam from a spectrometer light source through the gas sample. 7. The method of claim 1 further comprising: collecting a second gas sample from the living organism;measuring the nitrous oxide content of the second gas sample to acquire a second measured nitrous oxide value; andcomparing the second measured nitrous oxide value to the first measured nitrous oxide value to determine the presence of systemic inflammatory response in the living organism. 8. The method of claim 7 further comprising administering a therapy to treat the living organism before collecting the second gas sample. 9. The method of claim 1 further comprising administering a treatment to the living organism whereby the systemic inflammatory response is inhibited. 10. The method of claim 9 further comprising: collecting a second gas sample from the living organism;measuring the nitrous oxide content of the second gas sample to acquire a second measured nitrous oxide value; andcomparing the second measured nitrous oxide value to the first measured nitrous oxide value to determine an effectiveness of the dosage of medication based on a reduced level of nitrous oxide in the second gas sample. 11. The method of claim 1 wherein the living organism comprises a primate. 12. The method of claim 1 wherein measuring the nitrous oxide content of the first gas sample may comprise measuring the content of a nitrous oxide isotope. 13. The method of claim 1 wherein measuring the nitrous oxide content of the first gas sample comprises illuminating the first gas sample with an infrared light. 14. The method of claim 1 wherein the first gas sample further comprises a reference gas and wherein the method further comprises: measuring the reference gas content of the first gas sample to acquire a measured reference gas value; anddetermining a ratio of the first measured nitrous oxide content value to the measured reference gas value to determine the presence or absence of systemic inflammatory response. 15. The method of claim 14 wherein the reference gas comprises carbon dioxid... | Ekips Technologies, Inc. | A61B005/00 G01N021/00 G01N033/497
US20080249383 | A1 | 2008-10-09 | ANALYTE SENSING DEVICE HAVING ONE OR MORE SENSING ELECTRODES | 2008-04-04 | Embodiments of the present invention provide an analyte sensing device having one or more sensing electrodes. In embodiments, each sensing electrode may serve as an anode, a cathode, or a combination of an anode and cathode. In embodiments, there may be any suitable number of electrodes. Related analyte sensor insertion aids and methods of using the disclosed embodiments are also provided. | 1. An analyte sensing device, comprising: a main body configured to reside on skin of an individual when in use, the main body having one or more electrical components; andan analyte sensing electrode extending substantially perpendicularly from and electrically coupled to the main body, the analyte sensing electrode configured for insertion into the skin of the individual. 2. The analyte sensing device of claim 1, wherein the analyte sensing electrode is an anode. 3. The analyte sensing device of claim 1, wherein a skin-contact surface of the main body is a cathode. 4. The analyte sensing device of claim 3, wherein the skin-contact surface of the main body comprises surface protrusions which comprise the cathode. 5. The analyte sensing device of claim 1, wherein the analyte sensing electrode is a combination anode and cathode. 6. The analyte sensing device of claim 1, wherein the analyte sensing electrode extends from the main body by approximately 4-6 mm. 7. The analyte sensing device of claim 1, wherein the main body is a flexible substrate. 8. The analyte sensing device of claim 1, further comprising patterned adhesive on a skin-contact surface of the main body. 9. The analyte sensing device of claim 1, further comprising adhesive on a skin-contact surface of the main body, wherein the adhesive contains silver/silver chloride. 10. An analyte sensing device, comprising: a main body configured to reside on skin of an individual when in use, the main body having one or more electrical components; anda plurality of analyte sensing electrodes extending from and electrically coupled to the main body, the analyte sensing electrodes configured for insertion into the skin of the individual. 11. The analyte sensing device of claim 10, wherein the plurality of analyte sensing electrodes comprise at least one anode separate from at least one cathode. 12. The analyte sensing device of claim 10, wherein the plurality of analyte sensing electrodes comprise a plurality of anodes and a skin-contact surface of the main body is a cathode. 13. The analyte sensing device of claim 10, wherein the plurality of analyte sensing electrodes extend substantially perpendicularly from the main body. 14. The analyte sensing device of claim 10, wherein the plurality of analyte sensing electrodes are configured to detecting multiple analytes. 15. A method for measuring analyte in an individual, comprising: inserting a portion of an analyte sensing device into skin of an individual, the analyte sensing device comprising a main body configured to reside on the skin of the individual when in use, the main body having one or more electrical components; anda plurality of analyte sensing electrodes extending from and electrically coupled to the main body, the analyte sensing electrodes configured for insertion into the skin of the individual; andmeasuring an amount of analyte detected directly or indirectly by one or more of the plurality of analyte sensing electrode... | ISENSE CORPORATION | A61B005/1468
US07468045 | B2 | 2008-12-23 | Titanium molybdenum alloy guidewire | 2001-01-12 | A guidewire for medical use such as in vascular and nonvascular systems. The guidewire made from a titanium molybdenum alloy wire with a composition of approximately 78% titanium 11.5% molybdenum 6% zirconium and 4.5% tin by weight such that it is softer than stainless steel guidewires and stiffer than NiTi alloy guidewires. The distal end of the guidewire is of a smaller diameter and softer than the proximal end and fitted with a coil for springiness such that the distal end will bend when encountering curves in the body passageways. The distal tip may be heat treated for a gradient of softness with the distal tip being the softest. The distal end may also be tapered to provide an additional gradient of softness. A distal tip on the distal end of the guidewire protects the wall of the passageway from being punctured as the guidewire travels through the passageway. The resulting guidewire has properties between those of stainless steel guidewires and NiTi alloy guidewires for better torsion and stiffness characteristics. | 1. An intravascular guidewire adapted for insertion into the vascular system of a patient during the course of a catheterization procedure, comprising: a titanium molybdenum alloy wire having approximately 78% titanium, 11.5% molybdenum, 6% zirconium and 4.5% tin by weight, the wire having a diameter in a range of from 0.005 inch and 0.040 inch over a predetermined length dimension thereof, said wire having a proximal end portion and a distal end portion that is tapered to a lesser diameter than the diameter of the proximal end portion and that terminates in a rounded distal tip. 2. The guidewire as in claim 1 having, the distal end portion with a helical coil attached, and where the coil touches a distal tip of the guidewire, the coil providing springiness proximate the distal tip inhibiting kinking of the coil. 3. The guidewire as in claim 1 having, a rounded distal tip member on the end of the distal end portion of the wire to prevent the distal end of the wire from penetrating tissue in the wall of a body lumen upon passage of the guidewire through the body lumen. 4. The guidewire as in claim 1 wherein, the wire has a lubricious polymer coating. 5. The guidewire as in claim 1 wherein, the wire has a hydrophilic coating. 6. An intravascular guidewire adapted for insertion into the vascular system of a patient during the course of a catheterization procedure comprising a titanium molybdenum alloy wire having approximately between about 75% and about 83% titanium, between about 8% and about 14% molybdenum, between about 4% and about 8% zirconium and between about 2% and about 6% tin by weight, the wire having a diameter in a range of from 0.005 inch and 0.040 inch over a predetermined length dimension thereof, said wire having a proximal end portion and a distal end portion where the distal end portion is tapered to a lesser diameter than the diameter of the proximal end portion and terminates in a rounded distal tip. 7. The guidewire as in claim 6 having coil attached to a distal tip member such that the coil provides springiness at the distal tip portion to prevent kinking of the coil. 8. The guidewire as in claim 6 having, a distal tip member on the distal end portion to prevent the distal end of the wire from penetrating tissue in the wall of said body passageway. 9. The guidewire as in claim 6 wherein, the wire has a lubricious polymer coating thereon. 10. The guidewire as in claim 6 wherein, the guidewire has a hydrophilic coating thereon. | Minnesota Medical Development, Inc. | A61B005/00 A61M025/00
US07469159 | B2 | 2008-12-23 | Method for measuring neurovascular uncoupling in fMRI | 2005-05-26 | Neurovascular uncoupling that causes fMRI data to underestimate viable neuronal activity in the brain is detected and indicated to assist in pre-surgical planning. Tasks associated with the cortical region in question are performed while the fMRI measurement of brain activity is measured and the same tasks are performed while task performance is measured directly. Neurovascular uncoupling is detected and indicated by comparing the results of fMRI measurement and direct measurement. | 1. A method for detecting neurovascular uncoupling in the brain of a subject, the steps comprising: a) acquiring fMRI data from the subject with a magnetic resonance imaging system that measures brain activity in response to at least one of a selected stimulus applied to the subject and a selected task performed by the subject;b) transforming the fMRI data to a first map;c) measuring a behavioral response of the subject to the at least one selected stimulus and selected task performed by the subject;d) creating a second map based on the measured behavioral response in step c);e) comparing the first map created in step b) with the second map created in step d); andf) indicating a difference between the first map and the second map as neurovascular uncoupling. 2. The method as recited in claim 1 in which step a) measures the subject's field of vision and step b) includes producing a functional field map from the acquired fMRI data that indicates the locations in the subject's visual field that correspond to locations of indicated brain activity. 3. The method as recited in claim 2 in which step c) includes measuring the subject's field of view and step d) includes producing a visual field map that indicates the locations in the subject's visual field which the subject can actually see. 4. The method as recited in claim 1 in which step f) includes: f)i) displaying an anatomical image of the subject's brain;f)ii) converting the difference regions between the two maps relating to neurovascular uncoupling to corresponding regions in the image of the brain of i); andf)iii) displaying on the anatomical image regions where neurovascular uncoupling is indicated. 5. The method as recited in claim 1 which includes: g) acquiring MRI data from the subject and reconstructing therefrom an anatomical image of the subject's brain;h) displaying regions of brain activity on the anatomical image indicated by the acquired fMRI data; andstep f) includes displaying regions of neurovascular uncoupling on the anatomical image. 6. The method as recited in claim 1 in which step a) measures the subject's response to at least one of active movement and passive stimulation at a plurality of locations on the subject's body; and step b) includes: producing a functional field map from the acquired fMRI data that indicates the locations on the subject's body that correspond to locations of indicated brain activity. 7. The method as recited in claim 6 in which step c) includes measuring the subject's response to the at least one of active movement and the passive stimulation and step d) producing a sensory map that indicates the locations on the subject's body which are responsive. 8. The method as recited in claim 1 in which step b) includes producing a functional field map from the acquired fMRI data that may be used in step e) to compare response of the subject to direct measurements. 9. The method as recited in claim 1 in which step f) includes indicating neurovascular ... | The MCW Research Foundation, Inc. | A61B005/05 G06K009/00
US07445601 | B2 | 2008-11-04 | Non-invasive device and method for the diagnosis of pulmonary vascular occlusions | 2004-04-01 | The invention involves a device and method for ascertaining the functioning of the respiratory system and determining the cause of abnormal respiratory function. The device comprises an apparatus for measuring the flow of air from a patient's breadth and determining the concentration of gases contained therein via real-time spectrometry. From this data, a processor computes the ratio of detected gases relative to each other and as a function of expired volume and plots the calculations on a display screen. Based on various plots, a physician can more easily estimate the probability of a number of pulmonary diseases and afflictions, such pulmonary embolism, emphysema, congestive heart failure, infection, and related problems. | 1. A system for non-invasively diagnosing abnormal respiratory function, comprising: a patient breathing tube;a flow meter connected to said tube;a spectrometer interconnected to and in fluid communication with said tube, wherein said spectrometer is adapted to detect the concentration of gases present in said tube; anddata processing means interconnected to said flow meter and said spectrometer, wherein said data processing means is programmed to calculate and display at least one gas concentration ratio selected from the group consisting of the ratio of NO relative to the concentration of CO, the ratio of CO2 to O2 relative to the ratio of NO to CO, the ratio of NO to CO relative to expired volume, the ratio of CO2 to O2 relative to the ratio of CO2 to O2 relative to expired volume, and the ratio of CO2 to O2 simultaneously with a plot of NO relative to expired volume. 2. The system of claim 1, wherein said gas concentration ratio is the concentration of CO2 relative to the concentration of O2. 3. The system of claim 1, wherein said gas concentration ratio is the concentration of NO relative to the concentration of CO. 4. The system of claim 1, wherein said gas concentration ratio is the ratio of CO2 to O2 relative to the ratio of NO to CO. 5. The system of claim 1, wherein said gas concentration ratio is the ratio of NO to CO relative to expired volume. 6. The system of claim 1, wherein said gas concentration ratio is the ratio of CO2 to O2 relative to NO. 7. The system of claim 1, wherein said gas concentration ratio is the ratio of CO2 to O2 simultaneously with a plot of NO relative to expired volume. 8. The system of claim 1, wherein said spectrometer comprises a laser diode spectrometer remotely interconnected to and in fluid communication with said tube via a port in said patient breathing tube. 9. The system of claim 8, further comprising a vacuum pump interconnected to said spectrometer and said port. 10. The system of claim 9, wherein said vacuum pump operates at a rate of between about 10 to 100 milliliters per minute. 11. The system of claim 8, wherein said laser diode spectrometer is adapted to measure analyte concentrations of NO, CO, CO2, and O2. 12. The system of claim 8, wherein said laser diode spectrometer simultaneously measures CO2, O2, NO and CO. 13. The system of claim 8, wherein said laser spectrometer alternates between the measurement of any combination of CO2, O2, NO, and CO at a rate of at least twenty times per second for each such molecule. 14. The system of claim 8, wherein said laser spectrometer comprises a monochromatic light source from at least one tunable diode laser operating in a ring-down cavity mode with two or more mirror to increase apparent path length. 15. The system of claim 8, wherein said tube further includes a coupling for attachment to standard ventilator circuits. 16. The system of claim 8, wherein said tube further includes a coupling for attachment to a patient mouthpiece. 17.... | Charlotte-Mecklenburg Hospital | A61B005/02
US07451646 | B2 | 2008-11-18 | Device and method for resonant high-speed microscopic impedance probe | 2006-07-28 | A resonant high-speed microscopic impedance probe useful for small scale impedance measurements and/or cell and particle counting. | 1. A device capable of measuring impedance variances in a small region comprising: a probe electrode separated by one or more non-conducting regions from one or more shield electrodes;a detecting end of said probe electrode that can be placed in close proximity to said small region;an inductor L operatively connected to said probe electrode and to a transmission conductor connecting to measurement electronics;a signal generator for generating a radiofrequency signal; andimpedance measurement electronics able to measure a reflected signal and thereby determine impedance variances at said functional end. 2. The device of claim 1 further comprising: a varactor diode that acts as a controllable capacitance Cv operatively connected to said probe electrode, said inductor and to a transmission conductor connecting to measurement electronics. 3. The device of claim 2 further wherein: a stray capacitance between said probe and said one or more shield electrodes C resonates with said inductor L and said varactor diode forming a tunable LC resonant circuit with an approximate resonance frequency f0=(½)(1/L(C+CV))1/2 wherein said resonance frequency can be tuned by adjusting the effective capacitance Cv. 4. The device of claim 2 further comprising: a second inductor L2 operatively connected to said varactor and said inductor L and a DC source that allows separate DC tuning of the varactor diode (indicated as a variable capacitor CV) and RF reflectance;wherein inductor L2 allows a DC bias circuit to be isolated from an RF part of the device. 5. The device of claim 1 further wherein: said probe electrode, said detecting end, and said inductor L are configured in a planar geometry. 6. The device of claim 2 further wherein: said probe electrode, said detecting end, said inductor L, and said varactor are configured in a planar geometry. 7. The device of claim 1 further wherein: a stray capacitance between said probe and said one or more shield electrodes C resonates with said inductor L, forming an LC resonant circuit with an approximate resonance frequency f0=(½π)(1/LC)1/2. 8. The device of claim 1 further wherein: said inductor L includes a small series resistance r, thereby generating a measured reflectance loss with a peak centered at the resonance frequency f0 such that if a sample is placed near the detecting end of the probe, so that it couples to the electromagnetic fields emanating from the probe end, the effective capacitance and resistance of the sample impedance changes the reflectance loss. 9. The device of claim 8 further wherein: a capacitive part of the sample impedance changes the resonance frequency f0; anda resistive part of the sample impedance changes the magnitude of the reflection loss. 10. A method for measuring impedance variations in a small volume comprising: placing a microfabricated probe with a detecting end in proximity to said volume; andusing a tank circuit to couple said microfabricated probe to radiofrequency circuitry;... | The Regents of the University of California | G01N019/00 G01F023/26 A61B005/05
US07804299 | B2 | 2010-09-28 | Diffusion weighted preparatory sequence for magnetic resonance imaging pulse sequence | 2007-05-04 | A preparatory pulse sequence is applied prior to an imaging pulse sequence during a diffusion-weighted MRI scan. The preparatory pulse sequence diffusion weights the longitudinal magnetization using a gradient waveform that is first moment nulled to reduce image artifacts caused by patient motion. | 1. A method for producing a diffusion-weighted image with a magnetic resonance imaging (MRI) system, the steps comprising: a) performing a preparation pulse sequence with the MRI system that diffusion weights longitudinal magnetization in a subject, the preparation pulse sequence including:a)i) applying a radio frequency (RF) excitation pulse to produce transverse spin magnetization;a)ii) applying a first moment nulled diffusion weighting gradient waveform to diffusion weight the transverse spin magnetization;a)iii) applying an RF excitation pulse to produce diffusion-weighted longitudinal spin magnetization from the diffusion weighted transverse magnetization; anda)iv) applying an RF refocusing pulse such that the diffusion weighted transverse magnetization is refocused when the RF excitation pulse is produced in step a)iii);b) performing an imaging pulse sequence following step a);c) repeating steps a) and b) as necessary to acquire a set of data from which an image can be reconstructed; andd) reconstructing a diffusion-weighted image from the acquired set of data. 2. The method as recited in claim 1 in which the diffusion-weighted gradient waveform is also second moment nulled. 3. The method as recited in claim 1 in which a pair of RF refocusing pulses are applied during step a)iv). | Wisconsin Alumni Research Foundation | G01V003/00
US07440786 | B2 | 2008-10-21 | Method and apparatus for presentation of noninvasive glucose concentration information | 2004-12-14 | The invention relates generally to the extraction and/or presentation of glucose concentrations estimated as a function of time into a format that facilitates conveyance of the underlying information. More particularly, glucose concentration histories are presented in terms of risk of behavior, in a video format, and/or in an audio format. The reduction of data into video format, selected by time period, cluster, or glucose concentration response, into an animated or video presentation allows diagnosis and treatment information to be more readily determined and used. Alternatively, glucose concentrations are output through a voice synthesizer or an earcon. These information presentations are useful for both the medical professional and the end user. The information presentation is preferably used with a noninvasive, implantable, semi-continuous, and/or continuous analyte analyzer, such as a glucose concentration analyzer. | 1. An apparatus for analyte concentration display, comprising: a near-infrared noninvasive glucose concentration analyzer;a display module integrated into said analyzer; and means for video graphic display of historical analyte concentrations in a time compressed format on said display module. 2. The apparatus of claim 1, wherein said historical analyte concentrations represent at least five glucose concentrations. 3. The apparatus of claim 2, said historical analyte concentrations further comprising a time period, wherein said video graphic display represents glucose concentrations within said time period. 4. The apparatus of claim 3, wherein said time period comprises a period of any of about: a period of a glucose excursion;a day;a week;a month; andseveral months. 5. The apparatus of claim 4, wherein said video graphic display comprises any of: a slider graph;a treadmill graph; anda tachometer graph. 6. The apparatus of claim 5, wherein said video graphic display comprises a first animated graphical representation of at least one of said historical glucose concentrations. 7. The apparatus of claim 6, wherein said first animated graphical representation represents any of: a current time stamped glucose concentration; andan individual glucose concentration, wherein said individual glucose concentration comprises one of said historical analyte concentrations. 8. The apparatus of claim 7, wherein said video graphic display comprises a second animated graphical representation of any of: a maximum achieved glucose concentration;a minimum achieved glucose concentration;a targeted glucose concentration;a targeted maximum glucose concentration; anda targeted minimum glucose concentration. 9. The apparatus of claim 3, further comprising a user interface. 10. The apparatus of claim 9, wherein said time period is controlled through said user interface. 11. The apparatus of claim 1, wherein said display module comprises a color display screen. 12. The apparatus of claim 11, said historical analyte concentrations further comprising at least three physiological glucose concentration ranges, wherein said video graphic display represents said historical glucose concentrations for each of said ranges in different colors. 13. The apparatus of claim 12, wherein said physiological glucose concentration ranges comprise at least two of: a range of hyperglycemic glucose concentrations;a range of normal glucose concentrations;a range of hypoglycemic glucose concentrations; anda targeted glucose concentration range. 14. The apparatus of claim 11, wherein said display module comprises an input touch screen. 15. The apparatus of claim 1, wherein said analyzer comprises: a base module in a first container;a communication bundle having a first end and a second end, wherein said first end connects to said base module; anda sample module in a second container separated from said first container, wherein said second end of said communication bundle connects to s... | Sensys Medical, Inc. | A61B005/1455
US07462151 | B2 | 2008-12-09 | Method and apparatus for facilitating physiological coherence and autonomic balance | 2006-09-27 | Method and apparatus for determining the state of entrainment between biological systems which exhibit oscillatory behavior such as heart rhythms, respiration, blood pressure waves and low frequency brain waves based on a determination of heart rate variability (HRV) and an evaluation of the power spectrum thereof. Entrainment reflects a harmonious balance between the two branches of the autonomic nervous system within the body. This internal state of heightened physiological efficiency enhances health and promotes optimal performance. According to one embodiment a method is used to determine the entrainment level based on an entrainment parameter related to HRV, The method first determines the power distribution spectrum (PSD) and then calculates an entrainment parameter (EP), which is a measure of the power distribution in the HRV spectrum. High EP values occur when this power is concentrated within a relatively narrow range of frequencies, and lower values when the power is distributed over a broader range of frequencies. In one embodiment, an apparatus is provided for monitoring the heart beat and presenting this information via a personal computer, handheld device, or other processing means. | 1. A method comprising: sampling a physiological measure of a subject;deriving a power spectrum distribution (PSD) from the physiological measure;tracking dynamically a highest PSD peak having a PSD peak amplitude and frequency which vary over time; andproviding an indication that said subject is in a state of physiological coherence when the highest PSD peak exhibits at least a predetermined relative amplitude to a total energy above and below the highest PSD peak. 2. The method of claim 1, wherein the providing the indication further comprises providing the indication that said subject is in the state of physiological coherence when the highest PSD peak is in a frequency range of between 0.03125 Hertz and 0.234 Hertz. 3. The method of claim 1, wherein the state of physiological coherence is further characterized by an increased synchronization between two or more oscillatory systems of said subject. 4. The method of claim 1, wherein the state of physiological coherence is further characterized by a state of relaxation. 5. The method of claim 1, wherein sampling the physiological measure comprises sampling the physiological measure where said physiological measure comprises respiratory sinus arrhythmia, heart rate variability, respiration patterns, blood pressure rhythms and a factor derived from ECG-R wave amplitudes. 6. The method of claim 1 wherein said state of physiological coherence is further characterized by increased positive emotions in said subject. 7. The method of claim 1, wherein the state is further characterized by a waveform having a sine-wave shape, and wherein the waveform is one of a respiratory sinus arrhythmia wave and a heart rate variability wave. 8. The method of claim 1 further comprising providing a second indication when the highest PSD peak exhibits less than the predetermined relative amplitude, wherein the second indication is representative that the subject is in a more stressful state. 9. The method of claim 1, wherein sampling further comprises analyzing the physiological measure in one of a frequency domain, a time domain, a period analysis and a template match. 10. The method of claim 1, wherein said predetermined relative amplitude is an amplitude ratio of the PSD peak amplitude to a plurality of amplitudes for other PSD peaks corresponding to the total energy above and below the highest PSD peak. 11. The method of claim 10, wherein the highest PSD peak occurs within a search window and said amplitude ratio is computed by i) computing a first ratio of said PSD peak amplitude to a set of lower frequency peaks located below the search window, ii) computing a second ratio of said PSD peak amplitude to a set of higher frequency peaks located above the search window, and iii) multiplying said first ratio and said second ratio. 12. An apparatus comprising: sampling means adapted to sample a physiological measure of a subject; and,information output means;a processor coupled to the sampling means and the ... | Quantum Intech, Inc. | A61B005/00
US07438691 | B2 | 2008-10-21 | Method and device for predicting the fertile phase of women | 2002-07-29 | A method and device to predict ovulation in a female human by measuring changes in the concentration of a number of ions in eccrine sweat is disclosed. The concentration, or changes in concentration, of one or more ions are determined throughout the day and analyzed against predetermined patterns in order to predict ovulation one to five days in advance. This permits the user to more accurately determine commencement of the fertile phase, which for female humans is generally considered to be about four days prior to ovulation to one day after ovulation. The concentration of the ions measured include sodium (Na+), potassium (K+), ammonium (NH4+), calcium (Ca2+), chloride (Cl−) and nitrate (NO3−). To further increase the accuracy of the reading, a large number of readings can be obtained throughout a day and statistically analyzed to determine the change over time. In addition, the concentration of two or more ions can be obtained to increase accuracy. Ratiometric measurements between two or more ions can be determined to increase accuracy and account for ion accumulation on the skin. Ratiometric measurements between ammonium (NH4+) and calcium (Ca2+) have been found to provide more distinct patterns because the concentration of these two ions change in opposite directions during the relevant period preceding ovulation. | 1. A device for determining a fertile phase of a female human comprising: (a) a sensor for sensing concentrations of at least two ions in the eccrine sweat of the female, said at least two ions comprising at least a first ion and a second ion different from the first ion, said sensor comprising a first electrode for sensing the concentration of the first ion and a second electrode for sensing the concentration of the second ion, and generating output signals indicative of concentrations of the at least two ions in the eccrine sweat;(b) a processor for controlling the sensor to sense the concentrations of at least two ions in the eccrine sweat substantially simultaneously and at least on a daily basis; andwherein the processor monitors the output signals from the sensor to identify a distinct change in the concentration of one of the at least two ions following an inversion during a single menstrual cycle of the female which indicative of the female human being in the fertile phase. 2. A device as claimed in claim 1, wherein the at least two ions sensed in the eccrine sweat are selected from the group consisting of sodium (Na+), potassium (K+), ammonium (NH4+), calcium (Ca2+), chloride (Cl−) and nitrate (NO3−) of the eccrine sweat. 3. A device as claimed in claim 2, wherein to identify a distinct change in the concentration of one of the least two ions following an inversion, the processor monitors the output signals to identify a surge of 25% followed by a drop of 40% in the concentration of at least one of the first ion and the second ion. 4. The device as claimed in claim 1, wherein the processor monitors a ratio of the concentration of the first ion to the second ion to identify a distinct change in the concentration of the at least two ions following an inversion indicating the female human is in the fertile phase. 5. The device as claimed in claim 4, wherein the first ion is selected from the group consisting of potassium (K+), nitrate (NO3−), ammonium (NH4+) and calcium (Ca2+), and, a second ion is selected from the group consisting of sodium (Na+) and chloride (Cl−). 6. The device as claimed in claim 4, wherein the first ion is ammonium (NH4+) and the second ion is calcium (Ca2+). 7. A device as claimed in claim 4, wherein the device further comprises a display for displaying characters indicating the female human is in the fertile phase. 8. A device as claimed in claim 4, further comprising a fastener for fastening the device to the female subject such that the sensor contacts the skin of the female at least six hours each day; and wherein the processor controls the sensor to sense the concentration of the at least two ions between eight to eighteen times each day to monitor a daily average of the concentrations. 9. A device as claimed in claim 4, wherein the fertile state of ovulation is predicted to occur within six days following the inversion. 10. The device as claimed in claim 4, wherein the first ion is selected from the gr... | HealthWatchSystems, Inc. | A61B010/00 A61B005/05
US07437188 | B2 | 2008-10-14 | Method for reducing artifacts in magnetic resonance images acquired with continuous table motion | 2005-02-18 | MRA data is acquired from a large, longitudinal region of interest by translating the patient through the more limited longitudinal FOV of the MRI system as a three-dimensional MRA data set are acquired. Patient table movement is controlled to track a bolus of contrast agent as it passes through the region of interest. A seamless image of the entire region of interest is reconstructed after correcting the phase of acquired MRA data to reduce the signal falloff at abutting longitudinal FOVs. Phase corrections are determined from the central DC views acquired during the scan. | 1. A method for producing an image of a subject over an extended longitudinal field of view (FOVtot) which is larger than a longitudinal field of view (FOV) of a magnetic resonance imaging (MRI) system having a defined longitudinal field of view (FOV), the steps comprising: a) moving the subject through the MRI system such that the extended longitudinal field of view (FOVtot) passes through the defined longitudinal field of view (FOV);b) continuously acquiring NMR data from the subject as the subject is moved through the longitudinal FOV during a scan by repeatedly performing an imaging pulse sequence which acquires NMR data comprising a view of the subject;c) Fourier transforming each acquired view;d) storing each transformed view in a data matrix;e) calculating phase correction values from central DC views acquired during the scan that have no phase encoding applied;f) phase correcting transformed views using the calculated phase correction values; andg) reconstructing an image from phase corrected and transformed views stored in the data matrix. 2. The method as recited in claim 1 in which the MRI system has a table, and step a) is performed by: i) placing the subject on the table; andii) moving the table. 3. The method as recited in claim 2 in which the table is moved continuously while performing step a). 4. The method as recited in claim 2 in which the table is moved at different velocities while performing step a). 5. The method as recited in claim 2 which includes: injecting the subject with a contrast agent; andin which the table is moved at a velocity which tracks the contrast agent as it moves through the extended longitudinal field of view (FOVtot). 6. The method as recited in claim 2 in which step d) includes adjusting the location in the data matrix in which the view is stored as a function of the table location at the time the view is acquired in step b). 7. The method as recited in claim 1 in which step b) is performed by repeating a view order in which one of said central DC views is acquired at the beginning of the view order and step f) is performed by correcting subsequent views acquired during the view order with phase correction values calculated in step e) using the one central view acquired at the beginning of the view order. 8. The method as recited in claim 7 in which the view order is elliptical centric. 9. The method as recited in claim 1 in which step f) is performed before performing step d). 10. The method as recited in claim 1 in which the imaging pulse sequence employed in step b) acquires the NMR data while producing a readout gradient directed along the direction of subject movement. 11. The method as recited in claim 1 in which step e) is performed by calculating the phase of the central DC view. 12. The method as recited in claim 11 in which step f) is performed by subtracting the phase of the central DC view from the corresponding phase of the transformed view. 13. A magnetic resonance imaging (MR... | Mayo Foundation for Medical Education and Research | A61B005/05
US07448799 | B2 | 2008-11-11 | Dual use thermometer | 2006-11-15 | A thermometer for measuring body temperature in two different regions of a patient's body is provided. The thermometer includes a body and a probe fastened to the body. The probe has a tip portion disposed away from the body, and a temperature sensor is disposed on the tip of the probe. A supporting arm is pivotably attached to the body so that it can pivot between a closed position and an open position. In the closed position, the arm nests with the body so that the probe may be used to measure oral or rectal temperature. In the open position, the arm is positioned away from the body to provide a retaining surface for positioning the probe to measure axillary temperature. | 1. A thermometer for measuring body temperature, comprising: a body;a probe having a base portion fastened to the body and a tip portion disposed away from the body;a sensor for detecting a living being's temperature disposed on the probe; andan arm movably attached to the body so that it can pivot between a first position and a second position, wherein in the first position the arm is positioned away from the probe and in the second position the arm is positioned adjacent to the probe. 2. A thermometer according to claim 1, wherein the body comprises a curved body that forms a seat for accommodating the arm. 3. A thermometer according to claim 1, further comprising a switch activated by the arm to change the mode of the thermometer. 4. A thermometer according to claim 1, further comprising a switch to activate a function of the thermometer. 5. A thermometer according to claim 1, wherein in the second position, the arm forms a retaining surface for supporting the probe to measure axillary temperature. 6. A thermometer according to claim 5, wherein in the first position, the length of the probe is adapted to measure rectal temperature. 7. A thermometer according to claim 5, wherein in the first position, the length of the probe is adapted to measure oral temperature. 8. A thermometer for measuring body temperature, comprising: a body;a probe carried by the body, the probe having a tip portion disposed away from the body;a sensor for detecting a living being's temperature disposed on the probe; anda supporting arm pivotably attached to the body so that it can pivot between a closed position and an open position, wherein in the closed position the arm is positioned away from the probe and in the open position the arm is positioned away from the body to provide a retaining surface for positioning the probe. 9. A thermometer according to claim 8, wherein the body comprises a curved body that forms a seat for accommodating the supporting arm. 10. A thermometer according to claim 8, further comprising a switch activated by the supporting arm to change the mode of the thermometer. 11. A thermometer according to claim 8, further comprising a switch to activate a function of the thermometer. 12. A thermometer according to claim 8, wherein in the open position, the supporting arm forms a retaining surface for supporting the probe to measure axillary temperature. 13. A thermometer according to claim 12, wherein in the closed position, the length of the probe is adapted to measure rectal temperature. 14. A thermometer according to claim 12, wherein in the closed position, the length of the probe is adapted to measure oral temperature. 15. A thermometer for measuring body temperature, comprising: a probe having a base portion and a tip portion;a sensor for detecting temperature disposed on the tip of the probe; andmeans for providing a retaining surface for supporting the probe in an axillary region of a living being's body. 16. A thermometer ac... | Becton, Dickinson and Company | G01K001/14 G01K001/08 A61B005/01
US07431701 | B2 | 2008-10-07 | Method and system for diagnosis of lower urinary tract dysregulation and related conditions | 2001-09-13 | The invention disclosed herein provides a system and method for diagnosing lower urinary system conditions by receiving urodynamic data and subjective symptomatic impressions concerning a patient. The system and method then assigns weighted point values to the urodynamic data, for example, urethal pressure profiles due to sphincter contraction conditions such as urgency (2) and stress incontinence (4), and the subjective symptomatic impressions, quantifies norms for lower urinary system function, as a whole or in part, and quantifies recognized pathologies of lower urinary system conditions as degrees of departure from the norms. The system and method then determines a lower urinary system condition diagnosis of the patient, as a function of the degrees of departure from the norms, by invoking a set of rules to compare the weighted point values assigned and the quantified norms. The invention then reports the lower urinary system condition diagnosis, treatment recommendation, or further diagnostic suggestions. | 1. A method for use in a computer system of diagnosing lower urinary system conditions of a patient, the method comprising: receiving urodynamic data concerning the patient;receiving subjective symptomatic impressions concerning the patient;assigning at least one weighted point value to at least one of the urodynamic data and the subjective symptomatic impressions;quantifying at least one norm for at least one part of lower urinary system function as at least one numeric score;quantifying at least one recognized pathology of lower urinary system conditions as a degree of departure from the at least one norm;determining a lower urinary system condition diagnosis of the patient, as a function of the degree of departure from the at least one norm, by invoking a set of rules to compare the at least one weighted point value assigned to the at least one numeric score quantified; andreporting the determined lower urinary system condition diagnosis. 2. The method as described in claim 1 wherein the urodynamic data comprises a urethral pressure profile of the patient. 3. The method as described in claim 2 wherein the urodynamic data further comprises sphincter dynamic data of the patient of a type of at least one the following: the internal anal sphincter, the external anal sphincter, the internal urethral sphincter, and the external urethral sphincter. 4. The method as described in claim 1 wherein the urodynamic data comprises data generated from at least one of the following diagnostic tools: radiology, ultrasound, motor sensory evoked response activity, electromyographic recordings, magnetic resonance tomography, computed tomography, and nuclear imaging. 5. The method as described in claim 1 wherein the urodynamic data comprises: urethra characteristics;bladder characteristics; andurinary system flow rates of the patient. 6. The method as described in claim 5 wherein the urethra characteristics and the bladder characteristics are measured in qualities of at least one of tone, reflex excitability, sensation, and structural change. 7. The method as described in claim 1 wherein the step of receiving urodynamic data further comprises: recognizing graphical urodynamic patterns; andconverting the graphical urodynamic patterns into the urodynamic data. 8. The method as described in claim 1 further comprising: receiving a plurality of urodynamic datum and subjective symptomatic impressions concerning a plurality of patients, respectively;assigning a plurality of weighted point values to the plurality of urodynamic datum and subjective symptomatic impressions, respectively; andwherein the step of quantifying the at least one norm further comprises: determining the at least one numeric score as a functional average of the plurality of weighted point values; andwherein the step of quantifying the at least one recognized pathology further comprises: determining the degree of departure from the at least one norm as a functional average of the plurality of we... | - | A61B005/103 A61B005/117 A61B005/00
US07442172 | B2 | 2008-10-28 | Ligament tension gauge | 2002-08-29 | An apparatus and method for assessing tension of a ligament in a patient's body with less reliance on a subjective feel of a human examiner. | 1. An apparatus of assessing tension of a ligament at a joint in a patient's body, the apparatus comprising: a base having a surface abuttable with a portion of the patient's body;a tubular housing mounted to the base, the tubular housing defining an opening transverse to the surface;a probe having an elongated shaft with a first end, a second end, and a longitudinal axis between the first and second ends, the probe being mounted for translation within the opening of the tubular housing transverse to the surface and along the longitudinal axis of the shaft, the probe extending from the base such that the first end is axially engage able with the ligament; andmeans for translating the probe axially within the opening of the tubular housing toward the ligament to apply a force to the ligament by pressing the first end axially against the ligament. 2. The apparatus of claim 1, wherein the means for translating the probe comprises a force applicator coupled to the probe, the force applicator including a first end extending from the second end of the probe, and a resilient member positioned between the second end of the probe and the first end of the force applicator, the force applicator being operable to compress the resilient member and transmit an axial force to the second end of the probe along the shaft axis. 3. The apparatus of claim 2, wherein the force applicator comprises a hollow cap having a closed end corresponding to the first end of the force applicator and an open end, the open end and engaging the second end of the probe in sliding relationship, the cap containing the resilient member between the closed end and the second end of the probe such that the cap is sidable on the probe from the second end of the probe toward the first end of the probe a predetermined distance until the resilient member is fully compressed. 4. The apparatus of claim 3, wherein the hollow cap includes at least one of a transparent portion and a slotted portion to permit observing the probe shaft. 5. The apparatus of claim 4, wherein the resilient member includes a coiled spring. 6. The apparatus of claim 2, wherein the tubular housing comprises a tubular member mounted for axial translation relative to the base, the tubular member defining an axial bore, the probe shaft being mounted in the axial bore of the tubular member for axial translation relative to the tubular member, the force applicator being operable to cause the probe shaft to translate within the tubular member until the force applicator abuts the tubular member. 7. The apparatus of claim 1, wherein the base is configured to simultaneously abut a distal femoral site and a proximal tibial site. 8. The apparatus of claim 1, wherein the probe includes a bendable portion. 9. The apparatus of claim 6, further comprising a locking mechanism engageable with the tubular member to lock the tubular member's position relative to the base. 10. The apparatus of claim 9, wherein the locking mechanism ... | Zimmer, Inc. | A61B005/00
US07470239 | B1 | 2008-12-30 | High performance coil wire | 2000-06-20 | A high performance coil over-core guide wire. The guide wire incorporates a nickel-titanium core with a stainless steel coil to provide a wire with improved kink resistance and good pushability. | 1. A guide wire comprising: an elongate core defining a proximal end and a distal end;a continuous, unitary coil exhibiting an outer diameter and an inner diameter surrounding a portion of the core, with a portion of the coil extending beyond the distal end of the core by a plurality of turns of the coil, the inner diameter of the coil defining an internal area inside the coil along the portion of the coil that extends beyond the core; anda polymer contacting the distal end of the core and entirely filling the internal area, without extending beyond a distal end of the coil; anda metal tip extending from a distal end of the polymer and contacting a distal surface of the coil. 2. A guide wire comprising: an elongate core composed of a nickel-titanium alloy including a proximal end and a distal end;a continuous, unitary coil exhibiting an outer diameter and an inner diameter surrounding a portion of the core, the inner diameter of the coil defining an internal area inside the coil;a first polymer contacting the core and at least partially filling the internal area, without extending beyond a distal end of the coil; anda second polymer defining a tip and extending from a distal end of the first polymer and contacting a distal surface of the coil. 3. The guide wire of claim 2, wherein a portion of the coil extends beyond the distal end of the core by a plurality of turns of the coil. 4. The guide wire of claim 3, wherein the first polymer entirely fills the internal area along the portion of the coil that extends beyond the core. 5. The guide wire of claim 2, wherein the core is not tapered along its distal portion. 6. A guide wire comprising: an elongate core composed of a nickel-titanium alloy including a proximal end and a distal end;a continuous, unitary coil exhibiting an outer diameter and an inner diameter, composed of a second material, surrounding a portion of the core, the inner diameter of the coil defining an internal area inside the coil;a polymer contacting the core and at least partially filling the internal area, without extending beyond a distal end of the coil;a tip extending from a distal end of the polymer and contacting a distal surface of the coil; andwherein the core is not tapered along its distal portion. 7. The guide wire of claim 6, wherein a portion of the coil extends beyond the distal end of the core by a plurality of turns of the coil. 8. The guide wire of claim 7, wherein the polymer entirely fills the internal area along the portion of the coil that extends beyond the core. 9. The guide wire of claim 7, wherein the portion of the coil extending beyond the distal end of the core has a pitch different from a remaining portion of the coil. 10. A guide wire comprising: an elongate core defining a proximal end and a distal end;a continuous, unitary coil having a circular cross-section, exhibiting an outer diameter and an inner diameter surrounding a portion of the core, with a portion of the coil extending beyond ... | Scimed Life Systems, Inc. | A61B005/00 A61M025/00
US07627369 | B2 | 2009-12-01 | QT-interval measurement in the electrocardiogram | 2003-10-10 | A method and apparatus for measuring the QT interval of an electrocardiogram (ECG) signal is provided wherein the end of the T wave is identified from ECG data. The end of the T wave is defined as the first time of intersection at which an upright T wave of a first set of derived ECG signal data intersects an inverted T wave of a second set of derived ECG signal data. The intersection of the two sets of ECG data is along an isoelectric line within the trough after the positive T wave peak when the superimposed isoelectric baselines from the upright and inverted ECG signals demonstrate the best least squares fit. | 1. A method of measuring a QT interval of an electrocardiogram (ECG) signal wherein an end of a T wave is identified from ECG data, the method comprising the steps of: (a) acquiring ECG signal data from a signal acquisition apparatus or from a data storage;(b) deriving a first set of ECG signal data from the acquired ECG signal data using an electronic system;(c) inverting the first set of ECG signal data to derive a second set of ECG signal data using the electronic system, said second set of ECG signal data comprising an inverted set of said first set of ECG signal data; and(d) determining the end of the T wave using the electronic system by reference to timing of at least one intersection at which an upright T wave of said first set of derived ECG signal data intersects an inverted T wave of a said second set of derived ECG signal data, the two sets of ECG data being superimposed so as to maximize data fit over a segment of ECG signal after a peak of the positive T wave. 2. The method as claimed in claim 1 wherein the data fit is maximised by a least squares calculation. 3. The method as claimed in claim 1, wherein: the steps of deriving a first set of ECG signal data from the acquired ECG signal data comprise deriving a first set of reduced noise ECG signal data from the acquired ECG signal data;the inverting step comprises inverting the first set of reduced noise ECG signal data to derive an inverted set of reduced noise ECG signal data; andthe method further comprises the steps of:(e) identifying a portion of each set of ECG signal data corresponding to said segment;(f) calculating an offset such as to fit the first set of data to the inverted set of data over said segment;(g) detecting at least one intersection between the first set of data and the inverted set of data by reference to said offset; and(h) determining an end of said QT interval by reference to timing of the detected intersection. 4. The method as claimed in claim 3 wherein in step (h) the end of the QT interval is determined by a first point intersection. 5. The method as claimed in claim 4 wherein the end of the T wave is defined at the first point of intersection in said segment, provided there is at least one other point of intersection after a predetermined interval. 6. The method as claimed in claim 5, further comprising the steps of: determining the noise content in a segment of the ECG signal deemed to be an isoelectric baseline segment; andvarying said interval according to noise content in said segment. 7. The method as claimed in claim 3 wherein: the signal acquisition apparatus comprises multiple ECG signal acquisition leads;the ECG signal data comprise an ensemble of ECG signals which have been acquired from the ensemble of multiple acquisition leads; andstep (b) comprises calculating a median signal for each time from the ensemble of ECG signals for each lead to reduce low frequency baseline noise. 8. The method as claimed in claim 7 wherein the step (b) ... | PSI Heartsignals (Global) Limited | A61B005/0452
US08141556 | B2 | 2012-03-27 | Metallization with tailorable coefficient of thermal expansion | 2007-04-27 | The present invention is directed to an interconnect for an implantable medical device. The interconnect includes a pad and a first layer introduced over the pad. At least one of the pad or the first layer comprise a negative coefficient of thermal expansion (CTE) material. | 1. An interconnect for an implantable medical device comprising: a pad; anda first layer introduced over the pad, wherein one of the pad and the first layer comprise a negative coefficient of thermal expansion material comprising niobium pentoxide (Nb2O5);wherein one of the pad and the first layer comprises a range of 25 to 75 of Nb and 75 to 25 of Nb2O5 in a Nb/Nb2O5 ratio. 2. The interconnect of claim 1, wherein one of the pad and the first layer comprises 27/73 Nb/Nb2O5 ratio. 3. The interconnect of claim 1, wherein the coefficient of thermal expansion ranges from 0.5 to 8.0. 4. A joint in an implantable medical device comprising: a substrate;a conductor;an interconnect coupled to the substrate and to the conductor, the interconnect comprising: a pad; anda first layer introduced over the pad, wherein one of the pad and the first layer comprises a negative coefficient of thermal expansion material;wherein one of the pad and the first layer comprises a range of 25 to 75 of Nb and 75 to 25 of Nb2O5 in a Nb/Nb2O5 ratio. 5. The joint in an implantable medical device of claim 4, wherein the substrate comprises one of silicon, silicon dioxide, borosilicate glass, silicon nitride, silicon nitride, silicon carbide, and zinc oxide. 6. A joint in an implantable medical device comprising: a first element;a second element;an interconnect coupled to the first and the second elements, the interconnect comprising: a pad; anda first layer introduced over the pad, wherein one of the pad and the first layer comprises a negative coefficient of thermal expansion material;wherein one of the pad and the first layer comprises a range of 25 to 75 of Nb and 75 to 25 of Nb2O5 in a Nb/Nb2O5ratio. 7. The joint in an implantable medical device of claim 6, wherein the first element comprises a substrate. 8. The joint in an implantable medical device of claim 7, wherein the first element comprises one of silicon, silicon dioxide, borosilicate glass, silicon nitride, silicon nitride, silicon carbide, and zinc oxide. 9. The joint in an implantable medical device of claim 6, wherein the coefficient of thermal expansion ranges from 0.5 to 8.0. 10. A joint in an implantable medical device comprising: a first element;a second element;an interconnect coupled to the first and the second elements, the interconnect comprising: a pad; anda first layer introduced over the pad, wherein one of the pad and the first layer comprises a negative coefficient of thermal expansion material;wherein the joint is used to connect a sensor and a conductor; andwherein the sensor being one of a pressure sensor, an optical sensor, a biochemical sensor, a protein sensor, a motion sensor, a temperature sensor, a chemical sensor, and a genetic sensor. | Medtronic, Inc. | A61B019/00
US07979109 | B2 | 2011-07-12 | Needle driver for magnetic resonance elastography | 2007-10-12 | The present invention is directed toward an acoustic, piezoelectric, electric, electro-mechanical or pneumatically driven surface drum driver, in mechanical engagement with a biopsy or acupuncture needle device and a method for its use for diagnosis of small e.g. 100 microns, tumors via the production of magnetic resonance elastographic images (MRE), without artifact production, in a magnetic resonance imaging (MRI) machine. In a second embodiment, the invention is directed toward an acoustically, pneumatically, piezoelectrically, electrically and/or electro-mechanically driven acupuncture needle, useful for simulating manual single-needle acupuncture treatments via a non-manually manipulated acupuncture needle; and further to a device and process for determination, using twin pneumatically driven surface drivers, of organ stiffness, e.g. brain stiffness, which can be quantified so as to be useful in elucidating and quantifying brain cognitive state, e.g. normal, mild cognitive impairment (MCI) or Alzheimer's dementia (AD). | 1. A process to establish a standard range of brain viscosity and elasticity in normal subjects, Mild Cognitive Impairment (MCI) patients and Alzheimer's Disease (AD) patients comprising: positioning a patient within a magnetic resonance imaging (MRI) device;providing a pair of pneumatic drum drivers useful within said MRI device, each driven synchronously by a single wave generator;positioning said pair of pneumatic drum drivers in contact with the patient's skull, thereby simultaneously generating a shear wave, and a corresponding interference wave effective to neutralize attenuation of said shear wave, each wave propagating in the brain;elucidating brain stiffness data as a function of shear wavelength; andestablishing a standard range of brain viscosity and elasticity in normal subjects, MCI patients and AD patients;and correlating changes in brain viscosity and elasticity between normal, MCI and AD patients. 2. A process for generating magnetic resonance elastographs of cancerous tumors comprising: positioning a patient within a magnetic resonance imaging (MRI) device;providing a first acoustic, piezoelectric, electric, electro-mechanical or pneumatically driven surface drum driver, useful within said MRI device, in mechanical engagement with a biopsy or acupuncture needle;providing a second acoustic, piezoelectric, electric, electro-mechanical or pneumatically driven surface drum driver in contact with said patient, in an area adjacent said acupuncture or biopsy needle;inserting said biopsy or acupuncture needle within a tissue suspected of containing a cancerous tumor;providing wave generating means for synchronous movement of said biopsy or acupuncture needle within said tissue, wherein both spherical and plane waves are generated, thereby enhancing specificity and sensitivity; andgenerating a magnetic resonance elastograph. 3. The process of claim 2, further including the act of conducting a biopsy by use of the biopsy needle. | Lawrence Group Medical Device Trust | A61B005/05
US08285363 | B2 | 2012-10-09 | Surgical tracker and implantable marker for use as part of a surgical navigation system | 2008-07-17 | A marker for use with a surgical navigation system for tracking the tissue with which the marker is associated. The overall dimensions of the mark is such that it is shaped to be disposed below the skin of the patient in which the marker is implanted. The marker has a stem or a spike that allows the marker to be implanted in hard tissue such as bone. A head, containing navigation system components is located above the stem or spike. Once the procedure is completed the marker is removed from the patient by breaking the head free from the stem Alternatively the head may be rotated to cause the whole of the marker to break free from the bone in which the marker is implanted. | 1. A surgical tracker, said tracker comprising: a housing, said housing having a surface area of 180 cm2 or less and a maximum depth of 8 cm so that said housing can be placed against skin that extends over body tissue which is to be tracked;at least one first navigation system component mounted to said housing for receiving or transmitting signals from or to a static navigation unit remote from said housing so that the position of said housing relative to the static navigation unit can be determined wherein, the signals exchanged by said at least one first navigation system component are signals other than electromagnetic or radiofrequency signals and there is no wired connection by that at least one first navigation system component and the static navigation unit;at least one second navigation system component mounted to said housing for transmitting electromagnetic or radio frequency signals to or receiving electromagnetic or radio frequency signals from the body tissue adjacent said housing;a marker separate from said housing, said marker dimensioned to be implanted subcutaneously, said marker including: a stem dimensioned to be embedded in the body tissue so as to hold said marker in a static state in the body tissue; and at least one marker navigation component that receives electromagnetic or radio frequency signals from or transmits electromagnetic or radio frequency signals to said second navigation system component in said housing so that, based on exchanged signals, the position of said marker relative to said housing can be determined wherein, said at least one second navigation system component and said at least one marker navigation component are spaced apart a maximum distance of 25 cm so that the said at least one of said at least one second navigation system component or said at least one marker navigation component that transmits signals transmits signals at a power level of 1.67 Watts or less;a cable that extends from said housing into the body tissue to said marker over which signals are transmitted between said housing and said marker at least one navigation system component;a data transmitter attached to said housing and connected to said at least one second navigation system component or, through said cable, to said at least one marker navigation system component, said data transmitter configured to wirelessly transmit to a processor remote from said housing signals based on the transmission of signals between said at least one second navigation system component and said at least one marker navigation component so that the processor, which also receives the exchanges signals between the static navigation unit and said first navigation system component, based on the exchanged signals between the static navigation unit and said at least one first navigation system component and the exchanged signals between said at least one second navigation system component and said at least one marker navigation component, can determine th... | Stryker Corporation | A61B005/05
US08068650 | B2 | 2011-11-29 | Lesion quantification and tracking using multiple modalities | 2008-03-27 | A method for lesion detection includes acquiring pre-therapy medical image data from a first modality. Post-therapy medical image data is acquired from a second modality. A transformation matrix for transforming from an image space of the first modality to an image space of the second modality is calculated. A volume of interest is defined from the medical image data of the first modality. The volume of interest includes one or more lesions. The volume of interest is automatically copied to the medical image data of the second modality using the calculated transformation matrix. Treatment is directed to the lesion using the medical image data of the second modality including the copied volume of interest data. | 1. A method for lesion detection, comprising: acquiring medical image data from a first modality;acquiring medical image data from a second modality;calculating a transformation matrix for transforming from an image space of the first modality to an image space of the second modality;defining a volume of interest from the medical image data of the first modality, wherein the volume of interest includes one or more lesions wherein defining a volume of interest includes:defining the volume of interest based on one or more threshold values;receiving user input including an adjustment to the one or more threshold values; andredefining the volume of interest based on the adjusted one or more threshold values; andautomatically copying the volume of interest to the medical image data of the second modality using the calculated transformation matrix. 2. The method of claim 1, additionally including directing treatment based on structural information provided by the medical image data of the second modality and the volume of interest copied thereto. 3. The method of claim 2, wherein the treatment includes radiotherapy. 4. The method of claim 1, wherein the first modality is a functional modality and the second modality is a structural modality. 5. The method of claim 1, wherein the first modality is a PET or a SPECT and the second modality is a CT or MR. 6. The method of claim 1, wherein the volume of interest is defined from manual segmentation performed on one or more image slices of the medical image of the first modality. 7. The method of claim 1, wherein the volume of interest is defined from a placement of a predetermined shape within the medical image of the first modality. 8. The method of claim 1, additionally including creating a radiotherapy object for directing radiotherapy by marking the volume of interest on the image data first modality, wherein the first modality is a PET. 9. The method of claim 1, wherein redefining the volume of interest based on the adjusted threshold values is performed in real time by referring to a predetermined hash table. 10. A method for lesion tracking, comprising: acquiring medical image data from a first modality at a first point in time;defining a volume of interest from the medical image data of the first modality, wherein the volume of interest includes a lesion wherein defining a volume of interest includes:defining the volume of interest based on one or more threshold values:receiving user input including an adjustment to the one or more threshold values; andredefining the volume of interest based on the adjusted one or more threshold values;calculating one or more original quantified characteristics for the lesion based on the defined volume of interest;acquiring medical image data from a second modality at a second point in time;transforming the volume of interest to a second modality space and copying the transformed volume of interest to the medical image data of the second modality;calculatin... | Siemens Information Systems, Ltd. Siemens Medical Solutions USA, Inc. | G06K009/00 A61B005/05
US08140158 | B2 | 2012-03-20 | Heart monitoring device and a system to detect vibrations concerning the status of the heart | 2005-06-16 | An implantable heart monitoring system has a control circuit that operates an implanted vibrator to emit a vibration signal that interacts with tissue in vivo. A vibration sensor detects vibrations after interaction with the tissue, and supplies a detection signal to the control circuit. The control circuit analyzes the vibrations in the detected signal relative to the vibration signal, and derives information concerning at least one mechanical property of the heart therefrom, such as stiffness and/or thickness of at least a part of the heart. | 1. An implantable heart monitoring system comprising: a housing configured for implantation in a subject;a vibrator configured for implantation in the subject and operable to emit a vibration signal from a first in vivo site into cardiac tissue exhibiting a tissue transfer function representing physical properties of the cardiac tissue in the subject;a vibration sensor configured for implantation at a second in vivo site in the subject, said vibration sensor configured to detect said vibration signal after said vibration signal has propagated through said tissue between said first in vivo site and said second in vivo site and is further configured to emit a detection signal that represents the detected vibrations; anda control circuit in said housing in communication with said vibrator and said vibration sensor, said control circuit being configured to execute a procedure comprising actuating said vibrator to emit said vibration signal and receiving said detection signal from said vibration sensor, and to derive based on the detected vibrations in relation to the vibration signal said tissue transfer function, and to derive from said tissue transfer function, information describing at least one medical property of the heart of the subject, selected from the group consisting of stiffness of a heart portion of the heart and a thickness of a portion of the heart. 2. A system as claimed in claim 1 wherein said vibrator is located outside of said housing and is separately implantable in the subject. 3. A system as claimed in claim 2 wherein said vibration sensor is built into said housing. 4. A system as claimed in claim 1 wherein said vibrator is built into said housing. 5. A system as claimed in claim 1 wherein said vibration sensor is located outside of said housing and is separately implantable in the subject. 6. A system as claimed in claim 1 wherein said control circuit, in said procedure, derives a mechanical transfer function representing mechanical properties of in vivo tissue between said vibrator and said vibration sensor, as said information. 7. A system as claimed in claim 1 comprising a memory in communication with said control circuit, and wherein said control circuit is operable in time cycles corresponding to heart cycles of the heart, and wherein said control circuit executes said procedure a plurality of times, with a plurality of time cycles between each execution, and stores a result from each execution of the procedure in the memory, as respective stored results, and derives said information based on a change between successive stored results. 8. A system as claimed in claim 7 wherein said control circuit executes said procedure in a same portion of each time cycle in which the procedure is executed. 9. A system as claimed in claim 1 wherein said control circuit operates said vibrator to emit respective vibration signals having at least two different frequencies, and wherein said vibration sensor discriminates between sai... | St. Jude Medical AB | A61N001/365
US08050729 | B2 | 2011-11-01 | Devices for continuous measurement of glucose in body fluid | 2005-06-28 | A device for monitoring a constituent of body fluid continuously or repeatedly over a time period includes a body fluid sampling system and an optical sensor arrangement. The sampling system has a pumping arrangement which pumps body fluid from a cannula to a fluid sink. The optical sensor determines a concentration of at least one constituent of the body fluid. The pumping arrangement preferably ejects drops of body fluid from a fine nozzle. According to certain embodiments, optical sensor arrangements ensure that the optical paths do not pass through surfaces wetted by the fluid to be examined. | 1. A device for monitoring, continuously or repeatedly over a time period, a concentration of an analyte extracted from within a body, the device comprising: (a) a fluid transfer system adapted to generate of flow of a fluid containing at least one analyte, the flow passing from the body to a sensing location; and(b) an optical sensor arrangement associated with said fluid transfer system and configured for determining the concentration of the at least one analyte at said sensing location,wherein said optical sensor arrangement includes an illumination source for illuminating at least a portion of the fluid along an illuminating optical path and a sensor element sensitive to at least one wavelength of radiation received from said portion of the fluid along a sensing optical path, and wherein said optical sensor arrangement is deployed such that at least one of said illuminating optical path and said sensing optical path does not pass through any surface which is wetted by the fluid. 2. The device of claim 1, wherein said optical sensor arrangement is deployed such that at least one of said illuminating optical path and said sensing optical path does not pass through any container wall. 3. The device of claim 1, wherein said fluid transfer system includes at least one cyclically actuated piezoelectric transducer deployed to generate a net flow of the fluid. 4. The device of claim 1, wherein said fluid transfer system includes a nozzle arrangement for ejecting drops of the fluid and a pulse generator for generating a pressure pulse within the fluid so as to eject a drop from said nozzle arrangement. 5. The device of claim 4, further comprising a pressure regulating valve for maintaining a pressure above atmospheric pressure in a fluid sink deployed for receiving the ejected drops. 6. The device of claim 4, wherein said nozzle arrangement is deployed so as to eject a drop from said nozzle arrangement towards a capillary uptake nozzle. 7. The device of claim 4, wherein said nozzle arrangement is deployed so as to eject a drop from said nozzle arrangement towards an absorbent medium. 8. The device of claim 4, wherein said nozzle arrangement is deployed so as to eject a drop through a bubble trap arrangement. 9. The device of claim 4, wherein said optical sensor arrangement is deployed for determining the concentration of the analyte in the fluid making up said ejected drop while said ejected drop is in flight. | G-Sense Ltd. | A61B005/1455
US08761851 | B2 | 2014-06-24 | Indicators for a spectrophotometric system | 2006-12-06 | A near-infrared spectrophotometric system (e.g., a cerebral oximeter) includes a sensor portion and a monitor portion. The monitor portion includes a processor that runs an algorithm which utilizes the amount of detected light to determine the value of the oxygen concentration (e.g., the absolute level of oxygen concentration). The monitor portion also includes a visual display that displays the determined oxygen concentration values in various formats. The monitor portion may also include an audible device (e.g., a speaker), that provides audible indications of the determined oxygen concentration values. Various visual indicators may include, for example, color-coded graphs of the determined oxygenation values to alert the system user, for example, whether one hemisphere of the brain, or one or more regions of the brain, is in danger of adverse and potentially permanent damage. Also, data may be pre-processed by selecting the most clinically concerning sensor value (e.g., the sensor with the lowest value), and displaying only that sensor value and its identification on the display screen. Alternatively, an average value of multiple sensor measurements may be displayed. This reduces screen clutter and increases the speed of interpretation by the system user. Also, all sensor values may be averaged, and the average value displayed. The determined oxygenation values may also be provided in an audible format. | 1. A method for non-invasively determining the blood oxygen saturation level within a subject's tissue, using a spectrophotometric system that includes a processor, a visual display, and a plurality of sensors, each sensor having at least one light source operable to transmit a light signal into the subject's tissue, and at least one light detector operable to sense the light signal after passage through the subject's tissue and to produce signal data representative of the sensed light signal, the method comprising the steps of: positioning the sensors relative to the subject to sense separate regions of the subject;sensing the subject's tissue and determining an absolute oxygen saturation level for each region using sensed signal data from each region;displaying the oxygen saturation level for a selected region, or an average of the oxygen saturation levels of all of the regions, on the visual display;comparing the absolute oxygen saturation level of each region to a predetermined value that is based on established data, and which predetermined value is an absolute value independent of oxygen saturation values determined from the subject's tissue, and determining if the absolute oxygen saturation level of each region represents an adverse condition relative to the predetermined value;replacing the oxygen saturation value of the selected region or average value then currently displayed with the absolute oxygen saturation value of the region having the most adverse oxygen saturation value relative to the other regions. 2. The method of claim 1, wherein the predetermined value represents a minimum oxygen saturation value, and the adverse condition represents an oxygen saturation value below the minimum oxygen saturation value. 3. The method of claim 1, wherein the step of displaying the oxygen saturation level includes displaying one or more colors, where each color is associated with a level of oxygen saturation. 4. The method of claim 3, wherein a first color represents a first range of oxygen saturation values, a second color represents a second range of oxygen saturation values, and a third color represents a third range of oxygen saturation values, wherein the first, second, and third ranges are independent of one another, and each color is distinguishable from the other colors. 5. The method of claim 1, further comprising the step of adjusting the predetermined value in relation to one or more physiological parameters of the subject. 6. The method of claim 5, wherein the physiological parameters include the subject's temperature. 7. The method of claim 1, further comprising the step of providing an audible indicator of the oxygen saturation level for the region being displayed. 8. The method of claim 7, wherein the audible indicator changes as the oxygen saturation level changes. 9. The method of claim 7, wherein the audible indicator changes in pitch or frequency. 10. The method of claim 1, further comprising the step of providing a... | CAS Medical Systems, Inc. | A61B005/00
US07899520 | B2 | 2011-03-01 | Medical device for monitoring biological signals | 2008-05-21 | A medical device having a sensor for sampling a biological signal, the biological signal representing a signal waveform and forming a waveform vector composed of the biological signal samples, and a memory for storing a least two threshold vectors composed of boundary samples representing at least two boundaries related to the biological signal defining subspaces for the biological signal samples. One threshold vector is an upper threshold vector composed of upper boundary samples and the other threshold vector is a lower threshold vector composed of lower boundary samples. An evaluation unit connected to the sensor determines a similarity index (ASCI) by comparing each of the biological signal samples of the waveform vector to corresponding boundary samples of the threshold vectors, thus determining to which subspace each biological signal sample belongs to and creating a trichotomized signal vector, and calculating the signed correlation product of two trichotomized signal vectors. | 1. A medical device comprising: a. at least one sensor for sampling a biological signal,b. an evaluation unit connected to the sensor, wherein the evaluation unit receives samples of the biological signal and forms a waveform vector composed of the biological signal samples, andc. a memory for storing at least two threshold vectors composed of boundary samples, the threshold vectors representing at least two boundaries related to the biological signal samples, the boundaries defining subspaces for the biological signal samples, wherein one threshold vector is an upper threshold vector composed of upper boundary samples and the other threshold vector is a lower threshold vector composed of lower boundary samples, wherein each lower boundary sample is less than or equal to the corresponding upper boundary sample,wherein the evaluation unit determines a similarity index (ASCI) by:(1) comparing each of the biological signal samples of the waveform vector to corresponding boundary samples of the threshold vectors,(2) creating trichotomized signal vectors dependent on the subspaces to which the biological signal samples correspond, and(3) calculating the signed correlation product of two trichotomized signal vectors. 2. The medical device of claim 1 wherein the evaluation unit: a. determines a positive threshold vector and a negative threshold vector which together define an adaptive threshold zone (ABZ);b. determines a similarity index (ASCI) by: (1) trichotomizing two waveform vectors based on the threshold vectors;(2) calculating the signed correlation product between each sample pair of the two waveform vectors;(3) calculating a signed correlation product between the two waveform vectors by summing the signed correlation products of all sample pairs; and(4) normalizing the signed correlation product between the two waveform vectors by the total count of biological signal sample pairs. 3. The medical device of claim 2 wherein the evaluation unit trichotomizes a waveform vector by: a. comparing each biological signal sample of the waveform vector to the threshold vectors;b. setting the biological signal sample's trichotomized value to +1 if the biological signal sample is greater than the corresponding upper boundary sample;c. setting the biological signal sample's trichotomized value to −1 if the biological signal sample is less than the corresponding lower boundary sample;d. setting the biological signal sample's trichotomized value to 0 if the biological signal sample is less than or equal to the corresponding upper boundary sample, but greater than or equal to the corresponding lower boundary sample. 4. The medical device of claim 2 wherein the evaluation unit determines the signed correlation product of two trichotomized waveform vectors by: a. taking one biological signal sample from a first vector and taking a corresponding biological signal sample from the second waveform vector, thus forming a biological signal sample pair,b. setting the ... | BIOTRONIK CRM Patent AG | A61B005/04
US08934961 | B2 | 2015-01-13 | Trackable diagnostic scope apparatus and methods of use | 2008-05-19 | A surgical procedure is provided. The procedure acquires a plurality of points on or near a bone abnormality and registers them with a surgical navigation system. The navigation system uses the acquired points to make an implant having a surface adapted to contact the bone, and particularly an implant having a portion whose shape substantially matches that of the bone abnormality. | 1. A method of performing a surgical procedure, comprising: providing a tracking system and a diagnostic scope having a reference array attached thereto, the reference array having a reflective marker that is trackable by the tracking system;inserting a portion of the diagnostic scope into a body cavity through an incision site;identifying an abnormality on a bone within the body cavity with the diagnostic scope, wherein the diagnostic scope includes a grippable handle having a button actuatable by a hand gripping the handle;acquiring and registering a plurality of points on or near the abnormality with the diagnostic scope by detecting a position of the reflective marker attached to the diagnostic scope in space during each point acquisition, wherein the position of the reflective marker is detected in response to actuating the button on the grippable handle;capturing a real-time image of the bone with the diagnostic scope as the button on the grippable handle is actuated; andusing the acquired plurality of points to make an implant having a portion whose shape substantially matches that of the bone abnormality. 2. The method of claim 1, wherein the acquisition of the plurality of points on or near the abnormality with the diagnostic scope comprises collecting and analyzing dimensional data of the abnormality. 3. The method of claim 2, wherein the making of the implant having a portion whose shape substantially matches that of the bone abnormality comprises custom manufacturing the implant based on the collected and analyzed dimensional data of the abnormality. 4. The method of claim 1, wherein the making of the implant having a portion whose shape substantially matches that of the bone abnormality comprises accessing a computer database containing one or more reference implant models. 5. The method of claim 1, further comprising generating a representative model of the bone abnormality from the acquired points. 6. The method of claim 5, wherein the representative model of the bone abnormality comprises a three-dimensional and manipulatable image of the bone abnormality, the image showing the acquired points. 7. The method of claim 1, further comprising inserting a fiducial marker into the bone at or near the abnormality so that images taken in a first procedure can be registered with images taken in a later procedure. 8. The method of claim 1, further comprising installing the implant onto the bone. 9. The method of claim 8, wherein the implant is installed onto the bone by using the tracking system. 10. The method of claim 1, wherein the bone abnormality is on a femur. 11. The method of claim 1, wherein the bone abnormality is on a tibia. | Biomet Manufacturing, LLC | A61B005/05 G06K009/00 A61B017/56 A61B019/00 A61B005/06 A61B001/04 A61B001/317
US08489183 | B2 | 2013-07-16 | Biometric apparatus | 2008-05-15 | The invention provides a biometric apparatus which is able to determine the orientation of a measured person irrespective of the orientation of the measured person on the biometric apparatus or irrespective of the orientation of arrangement of the biometric apparatus and hence flow an electric current constantly through the same route and measure the voltage generated in the route, thereby providing accurate and reliable biometric data. The invention provides a biometric apparatus includes a plurality of weight measuring units provided on a casing for allowing a measured person to get on the upper surface thereof; a comparative determination unit for comparing measured values measured by the weight measuring units respectively and determining the orientation of the body of the measured person on the casing on the basis of the result of comparison; and a biometric data acquiring unit for acquiring a biometric data of the measured person on the basis of the result of determination by the comparative determination unit. | 1. A biometric apparatus comprising: a casing for allowing a person to get on a surface thereof;a plurality of weight measuring units provided in the casing;a biometric data acquiring unit including a plurality of electrode members arranged on the casing; anda comparative determination unit that determines an orientation of a foot of the person by comparing a value measured by at least one of the plurality of weight measuring units with a value measured by at least another one of the plurality of weight measuring units, and sets at least one of the plurality of electrode members as a measuring electrode and at least another one of the plurality of electrode members as a conducting electrode based on the determined orientation of the foot. 2. The biometric apparatus of claim 1, wherein the conducting electrode applies an electric current to a foot of the person and the measuring electrode detects a voltage from a foot of the person. 3. The biometric apparatus of claim 1, further comprising a changeover unit that connects the at least one of the plurality of electrode members set as the measuring electrode to a voltage detecting circuit and that connects the at least another one of the plurality of electrode members set as the conducting electrode to a current generating circuit. 4. The biometric apparatus of claim 1, wherein the comparative determination unit determines that a heel of the person is placed on a side where at least one of the plurality of weight measuring units measures a value larger than a value measured by at least another one of the plurality of weight measuring units. 5. The biometric apparatus of claim 4, wherein the comparative determination unit sets at least one of the plurality of electrode members, located on the side where the heel of the person is determined to be placed, as the measuring electrode. 6. The biometric apparatus of claim 1, wherein the comparative determination unit sets at least one of the plurality of electrode members, located on a side where at least one of the plurality of weight measuring units measures a value larger than a value measured by at least another one of the plurality of weight measuring units, as the measuring electrode. 7. The biometric apparatus of claim 1, wherein the comparative determination unit sets at least one of the plurality of electrode members, located on a side where a heel of the person is determined to be placed, as the measuring electrode. 8. The biometric apparatus of claim 1, wherein the comparative determination unit sets at least one of the plurality of electrode members, located on a side where a toe of the person is determined to be placed, as the conducting electrode. 9. The biometric apparatus of claim 1, wherein the comparative determination unit sets the measuring electrode and the conducting electrode after the person gets on the casing. 10. The biometric apparatus of claim 1, wherein at least one of the plurality of weight measuring units includes a ... | Tanita Corporation | A61B005/05 A61B005/103 A61B005/117
US20080278332 | A1 | 2008-11-13 | ANALYTE MONITORING SYSTEM AND METHODS | 2008-05-08 | Methods and systems for providing data communication in medical systems are disclosed. | 1. An analyte monitoring system, comprising: a data processing unit including a communication module; anda control unit in wireless communication with the data processing unit, the control unit configured to receive a predefined close proximity command associated with the wireless communication status of the data processing unit to the data processing unit; wherein in response to the received predefined close proximity command, the data processing unit changes the operational status of the communication module. 2. The system of claim 1 wherein the predefined close proximity command includes one of an RF switch on command or an RF switch off command. 3. The system of claim 1 wherein the data processing unit modifies the operational status of the communication module in response to the received predefined close proximity command. 4. The system of claim 3 wherein the operational status includes one of an active status and an inactive status. 5. The system of claim 4 wherein when the communication module is in the inactive status, the data processing unit does not transmit any data to the control unit. 6. The system of claim 1 wherein the control unit includes an output section to output information associated with the predefined close proximity command. 7. The system of claim 6 wherein the output information includes one or more of a visual display output, an audible output, or a vibratory output. 8. The system of claim 6 wherein the output information includes an icon. 9. A method, comprising: determining a communication status;receiving a communication related close proximity command; andmodifying the communication status based on the received command. 10. The method of claim 9 wherein the communication status includes one of an RF communication module in active mode, or the RF communication module in inactive mode. 11. The method of claim 10 wherein modifying the communication status includes switching from the active mode to the inactive mode of the RF communication module, or from the inactive mode to the active mode of the RF communication module. 12. The method of claim 9 wherein modifying the communication status includes disabling wireless data transmission based on the received command. 13. The method of claim 12 wherein the wireless data transmission is disabled for a predetermined time period. 14. The method of claim 12 further including: receiving a further communication related close proximity command; andfurther modifying the communication status based on the received further command. 15. The method of claim 14 wherein further modifying includes activating the wireless data transmission based on the received further command. 16. The method of claim 10 including transmitting data associated with monitored analyte level of a user. 17. The method of claim 16 including wirelessly transmitting the data associated with the monitored analyte level for further processing. 18. The method of claim 10 including entering inactive... | Abbott Diabetes Care, Inc. | G08B023/00 A61B005/05
US20080300473 | A1 | 2008-12-04 | Stabilized Multi-Wavelength Laser System for Non-Invasive Spectrophotometric Monitoring | 2006-12-14 | A spectroscopy system that may be used for spectrophotometric oxygenation monitoring of tissue includes a monitor portion and a sensor portion. The sensor portion generally includes a light source and one or more light detectors. The sensor portion may attach to a human to sense light signals from the light source that have traversed biological tissue, the light signals ultimately being used by the system to determine biological tissue blood hemoglobin oxygenation levels. The monitor portion generally includes a processor and a visual display. A laser beam combiner may couple a plurality of laser diode output light signals into one optical fiber. To stabilize the output of each of the laser diodes, an optical fiber light stabilizer is coupled to the combined laser diode output. The optical fiber light stabilizer redistributes the modes in the optical fiber such that the higher-order modes are filled until an equilibrium mode distribution is established. A light sensor may also provide feedback with respect to the laser diode output, which allows for compensation of any laser diode light output instability independently of optical fiber related instabilities. | 1. A spectroscopic system that monitors oxygenation levels in biological tissue, comprising: a light source having a light output;a predetermined length of optical fiber coupled to the light output of the light source, where the light output of the light source propagates through the optical fiber;an optical fiber light stabilizer provided within the optical fiber, where the light output of the light source propagates through the optical fiber light stabilizer, and where the light stabilizer redistributes modes of the light output from the light source until an equilibrium mode distribution is established in the light output propagating within the optical fiber; anda sensor coupled at an output of the light stabilizer, where the sensor provides the light output with the equilibrium mode distribution at an output of the spectroscopic system. 2. The spectroscopic system of claim 1, where the light source comprises a laser diode. 3. The spectroscopic system of claim 1, where the light source provides the light output at one or more predetermined wavelengths. 4. The spectroscopic system of claim 1, where the light source provides the light output at one or more predetermined wavelengths of light. 5. The spectroscopic system of claim 1, where the light source comprises a plurality of laser diodes. 6. The spectroscopic system of claim 1, where the light source comprises a plurality of laser diodes, each of the laser diodes providing a laser beam output at a predetermined wavelength, and where the spectroscopic system further comprises a laser beam combiner that combines together the laser beam outputs and provides the combined laser beam output to the optical fiber. 7. The spectroscopic system of claim 1, where the light stabilizer comprises a spool and a predetermined length of the optical fiber wrapped around the spool. 8. The spectroscopic system of claim 1, where the light stabilizer comprises a spool and a predetermined length of the optical fiber wrapped around the spool, where the spool has a radius that is at least approximately equal to a long term bend radius of the optical fiber wrapped around the spool. 9. The spectroscopic system of claim 1, further comprising a light detector that detects the light output from the light source. 10. The spectroscopic system of claim 1, further comprising a light detector that detects the light output from the light source and provides a detected light signal indicative thereof, and a signal processor that is responsive to the detected light signal to compensate for instability relating to the light output from the light source. 11. The spectroscopic system of claim 1, further comprising a light detector that detects the light output from the light source and provides a reference input intensity used for quantitative measurement methods in spectrophotometric applications. 12. The spectroscopic system of claim 11, whereas said light detector is a component of the light source. 13. The spectroscop... | CAS Medical Systems , Inc. | A61B005/1455
US20080306444 | A1 | 2008-12-11 | INTEGRATED MEDICAMENT DELIVERY DEVICE FOR USE WITH CONTINUOUS ANALYTE SENSOR | 2008-06-05 | An integrated system for the monitoring and treating diabetes is provided, including an integrated receiver/hand-held medicament injection pen, including electronics, for use with a continuous glucose sensor. In some embodiments, the receiver is configured to receive continuous glucose sensor data, to calculate a medicament therapy (e.g., via the integrated system electronics) and to automatically set a bolus dose of the integrated hand-held medicament injection pen, whereby the user can manually inject the bolus dose of medicament into the host. In some embodiments, the integrated receiver and hand-held medicament injection pen are integrally formed, while in other embodiments they are detachably connected and communicated via mutually engaging electrical contacts and/or via wireless communication. | 1. A medicament delivery device for monitoring and treating at least one of a plurality of hosts, the medicament delivery device comprising: a communication module configured to interrogate a continuous glucose sensor and to receive sensor data therefrom, wherein the sensor data comprises a signal associated with an analyte concentration of a host;a processor module configured to process the sensor data and calculate a medicament therapy, wherein the processor module comprises programming for calculating the medicament therapy based at least in part on the sensor data; anda hand-held injector pen configured and arranged to deliver a medicament to the host, based at least in part on the medicament therapy. 2. The medicament delivery device of claim 1, further comprising a user interface configured and arranged for at least one of input of at least some medical information and display of at least some medical information, wherein medical information comprises at least one of host information, received sensor data, processed sensor data, the calculated medicament therapy, a delivered medicament therapy, an instruction, an alert, an alarm, and a failsafe. 3. The medicament delivery device of claim 2, wherein the user interface is detachably connected to the hand-held injector pen. 4. The medicament delivery device of claim 2, wherein host information comprises at least one of a host information, type of medicament to be delivered, a glucose target, predicted hypoglycemia, predicted hypoglycemia, a therapy protocol, an alert, and an alarm. 5. The medicament delivery device of claim 1, wherein the processor module is further configured for validation of the medicament therapy. 6. The medicament delivery device of claim 1, wherein the medicament therapy comprises at least one of a type of medicament to be delivered, an amount of medicament to be delivered and a time of delivery. 7. The medicament delivery device of claim 1, wherein the communication module is further configured to communicate treatment information to a central monitor, wherein the treatment information comprises at least one of host information, sensor data, the medicament therapy, and delivered medicament information. 8. The medicament delivery device of claim 1, wherein the communication module is configured for wireless communication. 9. The medicament delivery device of claim 8, wherein the wireless communication is selected from the group consisting of RF communication, IR communication, Bluetooth communication, and inductive coupling. 10. The medicament delivery device of claim 1, wherein the communication module and the medicament delivery device are integrally formed. 11. The medicament delivery device of claim 1, wherein the communication module and the medicament delivery device are detachably connected. 12. The medicament delivery device of claim 1, wherein the injector pen is configured for aseptic medicament delivery to a plurality of hosts. 13. The medicament d... | DexCom, Inc. | A61M005/00 A61B005/145
US20080287811 | A1 | 2008-11-20 | Method for Assessing The Functional Condition Of Cardiovascular System | 2005-04-14 | The invention relates to medicine, namely: to cardiology, and may be used for assessment of functional condition of the human cardiovascular system (CVS) and the character of its control by the autonomic nervous system and other regulatory systems of the homeostasis. A method of non-invasive examination of the human CVS was developed, the method enabling to continuously, during a necessary period of time and quite simply with the aid of a computer and a piezoceramic tranducer (FIG. 3), record differential sphygmograms (FIG. 4) and by these sphygmograms using the method of determining the “coding” points to perform express-analysis simultaneously of two main pulse characteristics: a) rhythmicity and b) pulse oscillation of the arterial pressure (AP). The automatic disposition of the “coding” points in the averaged graph of the cardiocycle and their additional visual correction (FIG. 5) guarantee precision of determining the amplitude-temporal parameters at each recognised normal pulsation of a selected pulsogram fragment. By this fragment, the cardiac rhythm and all the amplitude-temporal cardiohemodynamic parameters will be measured and analysed, the parameters characterising the left ventricle myocardium contractile capacity as well as the resilient-elastic properties of the arterial bed vessel walls. For this purpose, the conventional units of the computer “digitizing” will be calibrated and transformed into accepted units of the blood AP measurement (mm Hg) and then, by means of integrating by respective areas of the cardiocycle graphs, the values of the blood AP pulse increment will be determined for different stages of the cardiac cycle. The continuous monitoring of the pulsogram parameter changes provides fulfillment of spectral analysis of the cardiac rhythm variability as well as of the selected cardiohemodynamic parameters. By results of the statistical and spectral analyses of the measured parameters' variability, the functional condition and the character of the subject's CVS vegetative regulation will be assessed by comparing the obtained values with the average statistical numerical values of these same parameters established for the CVS of groups of people who were selected as control subjects. The results may be used for resolving the problems of differential diagnosis of the cardiovascular diseases under clinical conditions, for individual examination of patients as well as for performing an operative medical checkup of health condition in various groups of population. | 1. A method for pulsometric assessment of the functional state and the character of autonomic regulation of the human cardiovascular system wherein, from the subject under study in relative resting and under conditions of performing a load test, a pulsogram is recorded in a non-invasive way by the method of differential sphygmography with the aid of a respective transducer, the pulsogram being studied using the method of the “coding points”, characterized in that the pulsogram analogue signal picked off from the transducer is transformed into a digital signal continuously recorded and analysed with the aid of a computer, with parallel periodic measurement of the blood arterial pressure (AP) using a sphygmomanometer; then a pulsogram fragment at least 2-minute long is selected and used for constructing a graph of the averaged cardiocycle and determining the “coding points”; the principle of the “coding points” disposition is transferred to each recognised normal pulsation of the selected fragment; and then, by the fragment, the temporal parameters characterising the heart rhythm and its variability are measured; by the same pulsogram fragment, the average value of the pulse arterial pressure (PAP) in conventional digital units is determined by means of integrating a respective part of the differential sphygmogram curve; and then, comparing this value with the average PAP value measured at the same period with the sphygmomanometer, the calibration coefficient of the AP proportionality is determined; and, taking the calibration coefficient into account, the conventional digital units are re-computed into conventional mm Hg, and on their basis the values of the blood arterial pressure increment are computed for different stages of the cardiac cycle, these values then being used for determining the amplitude-temporal cardiohemodynamic parameters characterising the left ventricle myocardium contractile capacity, namely: normalised pulse arterial pressure−PAPn;-cardiohemodynamicindex=CHDI=ΔAPAaccelΔAPAdecel,where ΔAPAaccel and ΔAPAdecel are, respectively, accelerated and decelerated anacrotic increments of the arterial pressure; the average−VAPAaccel and maximum−VmaxAPA velocities of the anacrotic increment of the AP in the phase of accelerated expulsion of the blood from the left ventricle;and the resilient-elastic properties of the arterial bed vessel walls, such as:-rigidityindexoftheaortawallsRIA=ΔAPRSPAPn·100%,where ΔAPRS is the arterial pressure increment due to the pulse pressure wave parried from the periphery during the systole period, and-thearterialwallstoneindex-TIA=ΔAPDaccelΔAPAaccel,where ΔAPDaccel is the accelerated dicrotic increment of the blood arterial pressure during the starting period of diastole;after which statistical processing of all the parameters is performed and spectral analysis of the heart rhythm variability and the selected amplitude-temporal cardiohemodynamic parameters are fulfilled, th... | - | A61B005/02
US20080262378 | A1 | 2008-10-23 | IMPLANTABLE THERAPY DELIVERY SYSTEM HAVING MULTIPLE TEMPERATURE SENSORS | 2007-04-19 | A system includes an implantable medical device and an implantable therapy delivery element. The device includes a housing, electronics disposed in the housing, and a first temperature sensor operably coupled to the electronics and located in proximity to the housing. The element has a proximal end portion configured to be operably couplable to the device and a distal end portion configured to be placed in a patient at a location removed from the device. A second temperature sensor located on the element between the proximal end and the distal end. The second temperature sensor is operably couplable to the electronics. The electronics are configured to compare a value associated with the first temperature sensed by the first temperature sensor to a value associated with the second temperature sensed by the second temperature sensor. The system may be used for monitoring an infection in proximity to the device. | 1. A system comprising: (i) an implantable medical device comprising: a housing;electronics disposed in the housing, the electronics configured to control delivery of therapy; anda first temperature sensor operably coupled to the electronics and located in proximity to the housing; and(ii) an implantable therapy delivery element comprising: a proximal end portion configured to be operably couplable to the implantable medical device, the proximal end portion having a proximal end;a distal end portion configured to be placed in a patient at a location removed from the implantable medical device, the distal end portion having a distal end; anda second temperature sensor located on the element between the proximal end and the distal end, the second temperature sensor being operably couplable to the electronics,wherein the electronics are further configured to compare a value associated with the first temperature sensed by the first temperature sensor to a value associated with the second temperature sensed by the second temperature sensor. 2. The system of claim 1, wherein the implantable medical device further comprises a first electrical contact operably coupled to the electronics, and wherein the implantable therapy delivery element comprises a second electrical contact configured to contact the first electrical contact when the implantable therapy element is coupled to the implantable medical device, the second electrical contact being electrically coupled to the second temperature sensor. 3. The system of claim 1, wherein the electronics comprise an alert generating module. 4. The system of claim 3, wherein the electronics comprise a computer readable medium that when implemented cause the device to determine whether a value associated with the first temperature crosses a threshold; determine whether the value associated with the first temperature is greater than a comparable value associated with the second temperature by a predetermined amount; and activate the alert generating module if the value associated with the first temperature crosses the threshold value and is greater than the comparable value associated with the second temperature by the predetermined amount. 5. The system of claim 1, wherein the implantable medical device is an implantable electrical signal generator and wherein the therapy delivery element is an implantable medical lead. 6. The system of claim 1, wherein the implantable medical device is an implantable electrical signal generator and wherein the therapy delivery element is a lead extension. 7. The system of claim 1, wherein the implantable medical device is an implantable infusion device and wherein the therapy delivery element is an implantable catheter. 8. A method for monitoring an infection in proximity to a therapy delivering medical device implanted in a patient, the therapy delivering medical device being part of a system comprising an associated therapy element for delivering the therapy from the the... | MEDTRONIC, INC. | A61B005/01
US20080269629 | A1 | 2008-10-30 | Multimodal therapeutic and feedback system | 2007-04-25 | The generation of alpha waves in a mammalian subject may be stimulated through a process of measuring a wave-from produced by the electrical activity of the subject's brain, analyzing the wave-form to determine changes in its frequency, and delivering audio and visual stimulation to the subject, each at a rate selected to vary the frequency of the wave-form until any alpha state is achieved. The rate may be selected by the subject, by a therapist, or may be varied automatically through a feedback mechanism. | 1. A method of treating a mammalian subject, comprising the steps of: measuring a wave-form produced by the electrical activity of the subject's brain;analyzing said wave-form to determine changes in its frequency; anddelivering to the subject audio and visual stimulation, each at a rate selected to vary said frequency of said wave-form to a desired frequency of said wave-form. 2. The method of claim 1, wherein said step of delivering audio and visual stimulation includes the step of delivering pulses of light to one of the subject's eyes. 3. The method of claim 2, wherein said step of delivering pulses of light comprises the step of delivering pulses of light separately to the right and left visual fields of one of the subject's eyes. 4. The method of claim 1, wherein said step of delivering audio and visual stimulation includes the step of delivering pulses of sound to one of the subject's ears. 5. The method of claim 1, wherein said step of delivering audio and visual stimulation to the subject includes the steps of delivering pulses of light to one of the subject's eyes and delivering pulses of sound to one of the subject's ears. 6. The method of claim 5, wherein said steps of delivering pulses of light and delivering pulses of sound are performed simultaneously. 7. The method of claim 6, wherein said steps of delivering pulses of light and delivering pulses of sound are performed so that said pulses of light are synchronized with said pulses of sound. 8. The method of claim 1, further including the step of selecting said rate to correspond to a rate known to induce alpha wave-forms in the subject. 9. The method of claim 8, wherein said method includes the step of monitoring said changes in frequency, and said selecting step is performed in response to said changes in frequency of the wave form. 10. The method of claim 9, wherein said selecting step is performed by the subject. 11. The method of claim 9, wherein said selecting step is performed by a person other than the subject. 12. The method of claim 9, wherein said selecting step is performed automatically by a feedback mechanism. 13. A system for treating a mammalian subject, comprising: a brain wave measuring device for measuring a wave-form representative of the electrical activity of the subject's brain;an analyzer for determining changes in the frequency of said wave-form;a light-producing element for providing pulses of light to the right visual field of the subject at a first rate and to the left visual field of the subject at a second rate;a sound transducer for providing pulses of audible sounds to one of the subject's ears at a third rate; anda control system for varying said first, second and third rates in response to said changes in the frequency of said wave-forms. 14. The system of claim 13, wherein said measuring device and said analyzer are included within an electroencephalograph having an electrode. 15. The system of claim 13, wherein said control system c... | - | A61B005/0482
US07433731 | B2 | 2008-10-07 | Biological signal detection system | 2004-12-27 | A biological signal detection system includes a pair of electrodes, adapted to be attached on a living body; a transmitter, adapted to be attached on the living body, and including an electric circuit operable to process the biological signal and to telemeter the processed signal; and a biological signal inputting apparatus. The system may also include a receiver-transmitter operable to receive the processed signal telemetered from the transmitter and to telemeter the processed signal to the biological signal inputting apparatus by way of a wide area network; and a storage for storing the processed signal telemetered from the transmitter as biological signal data and for inputting the biological signal data to the biological signal inputting apparatus. Also, the system may include a storage for storing the processed signal telemetered from the transmitter as biological signal data and for inputting the biological signal data to the biological signal inputting apparatus. | 1. A biological signal detection system, comprising: a plurality of first electrodes for detecting a biological signal;a plurality of second electrodes for detecting the biological signal;a first support, supporting the first electrodes and adapted to be disposed on a living body so that a first one of the first electrode is located on a first position of the living body and a second one of the first electrodes is located on a second position of the living body;a second support, supporting a first one of the second electrodes and adapted to be disposed on the living body so that the first one of the second electrodes is located on a third position of the living body;a third support, supporting a second one of the second electrodes and adapted to be disposed on the living body so that the second one of the second electrodes is located on a forth position of the living body;a forth support, supporting a third one of the second electrodes and adapted to be disposed on the living body so that the third one of the second electrodes is located on a fifth position of the living bodya fifth support, supporting a forth one of the second electrodes and adapted to be disposed on the living body so that the forth one of the second electrodes is located on a sixth position of the living body;a transmitter, adapted to be disposed on the living body, comprising an electric circuit operable to process the biological signal and to telemeter the processed signal;a first connection section for electrically connecting the plurality of first electrodes to the transmitter; anda second connection section for electrically connecting the plurality of second electrodes to the transmitter;a biological signal inputting apparatus;a receiver-transmitter, adapted to be disposed on the living body, and operable to receive the processed signal telemetered from the transmitter and to telemeter the processed signal to the biological signal inputting apparatus using a wide area network; anda storage, adapted to be removably attached to the receiver-transmitter for storing the processed signal telemetered from the transmitter as biological signal data,wherein the first support, the second support, the third support, the fourth support and the fifth support are independent from each other so that relative distances there-between are variable;when the first one of the first electrodes, the second one of the first electrodes, the first one of the second electrodes, the second one of the second electrodes, the third one of the second electrodes, and the forth one of the second electrodes are respectively located on the first position, the second position, the third position, the forth position, the fifth position and the sixth position, at least one CM5 lead is established between the first position and the third position, at least one NASA lead is established between the second position and the fourth position, and at least one CC5 lead is established between the fifth position and the... | Nihon Kohden Corporation | A61B005/0408
US20080300495 | A1 | 2008-12-04 | Method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure | 2007-05-29 | A method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure uses a pressure sensor and an acceleration sensor to measure a voltage-time variation waveform of blood pressure detected by an acceleration sensor. An interference variation waveform is subtracted from the measured voltage-time variation waveform to calculate the correct blood pressure value. During measurement, the acceleration sensor detects and determines whether the vibration interference falls within a predetermined allowable range; if yes, a warning signal is issued to remind users to measure blood pressure again; if no, the acceleration sensor detects an interference variation waveform. The inference variation waveform obtained by a conversion function is subtracted from the voltage-time variation waveform when the blood pressure measurement is affected by external vibration interferences, and the original voltage-time variation waveform not affected by vibration interferences is restored, so as to measure a correct blood pressure value. | 1. A method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure, and the method utilizes a pressure sensor, an acceleration sensor and a warning unit in the sphygmomanometer to measure a correct blood pressure value when the measurement is taken under vibrations, and the method comprising the steps of: Step S01: powering on the sphygmomanometer to execute the function of measuring blood pressure;Step S02: detecting whether or not a vibration interference detected by the acceleration sensor of the sphygmomanometer falls within an allowable range, and executing Step S03 if the detected vibration interference falls within the allowable range, or else executing Step S04;Step S03: issuing a warning signal to remind a user to execute Step S02 again;Step S04: detecting voltage-time variation waveforms of the pressure sensor and the acceleration sensor respectively, and executing Step S05;Step S05: calculating an interference waveform variation from the voltage-time variation waveform detected by acceleration sensor under the vibration interference according to a conversion function, and executing Step S06;Step S06: subtracting the interference waveform variation from the detected voltage-time variation waveform of the blood pressure and pulse having the vibration interference, and executing Step S07; andStep S07: calculating the correct blood pressure value from the voltage-time variation waveform of the blood pressure and pulse minus the interference waveform variation. 2. The method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure as recited in claim 1, wherein the pressure sensor detects the voltage-time variation waveform by a pressure difference of the blood pressure and pulse. 3. The method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure as recited in claim 1, wherein the acceleration sensor produces a voltage variation by spatial position variation of different angles for detecting a voltage-time variation waveform, when a vibration interference occurs. 4. The method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure as recited in claim 3, wherein the acceleration sensor detects a two-dimensional space variation. 5. The method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure as recited in claim 3, wherein the acceleration sensor detects a three-dimensional space variation. 6. The method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure as recited in claim 1, wherein the warning unit is an LCD, a LED screen, a lamplight or a speaker. 7. The method of eliminating vibration interferences of a sphygmomanometer for measuring blood pressure as recited in claim 6, wherein the warning signal issued by the warning unit is voice, music, sound, sentence of different colors, word, lamplight or figure. | HEALTH & LIFE CO., LTD | A61B005/00
US20080242977 | A1 | 2008-10-02 | SYSTEMS, METHODS AND APPARATUS FOR LONGITUDINAL/TEMPORAL ANALYSIS OF PLAQUE LESIONS | 2007-03-30 | Systems, methods and apparatus are provided through which in some embodiments detection of a change in characteristics of plaque in a longitudinal exam is automated for the purpose of assessing change in disease due to therapy, patient behavior modifications or follow-up. In some embodiments, diagnosis and treatment of arterial lesions includes obtaining a plurality of sets of computed-tomography images of at least one arterial plaque lesion, wherein each set of computed-tomography images are acquired at a different time, then storing the computed-tomography images in a database and analyzing arterial plaque variations in the sets of computed-tomography images for changes in at least one parameter. | 1. A computer-accessible medium having executable instructions to support s diagnosis and treatment of arterial lesions, the executable instructions capable of directing a processor to perform: accessing a plurality of patient images that were acquired longitudinally; andanalyzing arterial plaque variations in the plurality of images for changes selected from a group of changes consisting of at least one change in size, at least one change in composition, at least one change in characteristics and at least one change in location. 2. The computer-accessible medium of claim 1, wherein the executable instructions capable of directing the processor to perform the analyzing further comprise executable instructions capable of directing the processor to perform: bookmarking each lesion; andcomparing longitudinally each bookmarked lesion. 3. The computer-accessible medium of claim 1, wherein the executable instructions capable of directing the processor to perform the analyzing further comprise executable instructions capable of directing the processor to perform: linking the bookmarked lesions on a vessel-by-vessel basis. 4. The computer-accessible medium of claim 1, wherein the executable instructions capable of directing the processor to perform the analyzing further comprise executable instructions capable of directing the processor to perform: registering vessels to a common reference or standard reference; andcomparing the current vessel with the corresponding vessel in a previous image. 5. The computer-accessible medium of claim 1, wherein longitudinally further comprises: across multiple studies. 6. The computer-accessible medium of claim 1, wherein longitudinally further comprises: over a long-term time frame. 7. The computer-accessible medium of claim 1, wherein the executable instructions capable of directing the processor to perform the acquiring further comprise executable instructions capable of directing the processor to perform: acquiring through computed-tomography. 8. The computer-accessible medium of claim 1, wherein the executable instructions capable of directing the processor to perform the acquiring further comprise executable instructions capable of directing the processor to perform: acquiring through magnetic resonance. 9. A method to support diagnosis and treatment of arterial lesions, the method comprising: obtaining a plurality of sets of computed-tomography images of at least one arterial plaque lesion, wherein each set of computed-tomography images are acquired at a different time;storing the computed-tomography images in a database; andanalyzing arterial plaque variations in the sets of computed-tomography images for changes in at least one parameter. 10. The method of claim 9, the analyzing further comprising: bookmarking each lesion; andcomparing longitudinally each bookmarked lesion. 11. The method of claim 9, the analyzing further comprising: registering images and locations in the images; anddetermining chan... | GENERAL ELECTRIC COMPANY | A61B005/05
US20080319505 | A1 | 2008-12-25 | Integrated Transcranial Current Stimulation and Electroencephalography Device | 2008-05-09 | In an integrated transcranial current stimulation and electroencephalography device, data obtained by EEG electrodes is received, amplified, converted, and then processed by a microcontroller that extracts frequency information from the sampled data and produces signals in response to the extracted EEG data. These signals are converted to create a software-definable alternating voltage used to control a current-source that connects to stimulation electrodes. The device may further download recorded EEG data from the microcontroller to a computer for further analysis. Using the device, EEG signals are detected, then received, processed, and analyzed by a microcontroller to identify the patient state. Based on the patient state and the desired protocol, a type and amount of current stimulation to apply to the patient is determined and a control signal is sent from the microcontroller to the current source in order to trigger the transcranial current stimulation of the patient. | 1. An integrated transcranial current stimulation and electroencephalography system, comprising: at least one electrode;a current source for providing stimulation current to the electrode; anda microcontroller adapted to: receive electroencephalography data from the electrode;analyze the received data to determine a stimulation current to be delivered by the electrode; andsend a signal to the current source to induce the electrode to deliver the determined stimulation current. 2. The device of claim 1, further comprising at least a second electrode, wherein the detection and stimulation functions are performed by different electrodes. 3. The device of claim 1, further comprising at least one amplifier for amplifying signals transmitted between the electrode and the microcontroller. 4. The device of claim 2, further comprising at least one amplifier for amplifying signals transmitted between at least one electrode and the microcontroller. 5. The device of claim 1, further comprising a connection to a computer for downloading the electroencephalography data for additional analysis. 6. The device of claim 1, further comprising a connection to a computer for downloading the electroencephalography data for additional analysis. 7. An integrated transcranial current stimulation and electroencephalography device, comprising: input for receiving electroencephalography data signals;microcontroller, connected to the input, comprising: electroencephalography data signal processor;electroencephalography data signal analyzer; andcurrent stimulus control signal generator; andoutput, connected to the microcontroller, for transmitting current stimulus control signals to a current source. 8. The device of claim 7, further comprising an interface for downloading electroencephalography data to a computer. 9. A method for integrated transcranial current stimulation and electroencephalography, comprising: detecting electroencephalography signals from patient; andin a single microcontroller, performing the steps of: receiving and processing the electroencephalography signals with a microcontroller;analyzing the processed signals to identify a patient state;based on the identified patient state, determining the type and amount of transcranial current stimulation to apply to patient; andsending a control signal to a current source to trigger the determined transcranial current stimulation of the patient. 10. The method of claim 9, further comprising the step of downloading the received and processed electroencephalography signals to a computer for additional analysis. 11. The method of claim 9, wherein the step of determining the type and amount of current stimulation to apply selects between precomputed control signals depending on the identified patient state. 12. The method of claim 9, wherein the step of determining the type and amount of current stimulation to apply adaptively computes the control signal depending on the identified patient state. 13. The... | Massachusetts Institute of Technology | A61N001/04 A61B005/0476
US07440790 | B2 | 2008-10-21 | Method and apparatus for monitoring cardiac activity | 2005-04-26 | A method and apparatus for monitoring cardiac activity starting from a series of images of the heart made by X-rays. Attenuation of the X-rays is sampled on an area subject to repetitive thickenings of at least one wall of the heart, and a repetitive attenuation signal is interpreted as a signal representative of the heart activity. | 1. A method for monitoring cardiac activity comprising: providing a series of images of the heart by imaging detected attenuation of radiation;for more than one of the series of images, sampling the attenuation of the radiation on an area subject to repetitive thickenings of at least one wall of the heart;calculating a weighted linear combination of two curvatures fit to each sampled attenuation; andanalyzing the weighted linear combinations over time to identify a repetitive attenuation signal as a signal representative of the cardiac activity based on a wall thickness of the at least one wall of the heart. 2. The method according to claim 1 wherein the sampling area has an approximately circular periphery approximately corresponding to the dimensions of the heart. 3. The method according to claim 2 wherein the circular periphery surrounds the heart. 4. The method according to claim 3 wherein the sampling area surrounds an area excluded from sampling. 5. The method according to claim 4 wherein the excluded area is an area with regulated intensity. 6. The method according to claim 2 wherein the sampling area surrounds an area excluded from sampling. 7. The method according to claim 6 wherein the excluded area is an area with regulated intensity. 8. The method according to claim 1 wherein the sampling area surrounds an area excluded from sampling. 9. The method according to claim 8 wherein the excluded area is an area with regulated intensity. 10. The method according to claim 1 wherein the sampling area comprises at least one segment that intersects a variable range projection of a wall thickness of the at least one wall of the heart. 11. The method according to claim 10 wherein the segment is oriented at 30° from a longitudinal axis of the heart. 12. The method according to claim 11 wherein the segment is oriented at 60° from a longitudinal axis of the heart. 13. The method according to claim 12 comprising: providing an attenuation profile on a path of segment(s); andidentifying one or more parameters in an analytic curve of a preset type with a shape conforming to the profile. 14. The method according to claim 13 comprising: identifying the parameters of two parabolas, one with a shape complying with an external silhouette of the profile, the other with the upper cavity of the profile; andproviding the signal representative of the cardiac activity based on a weighted combination of the parameters of the parabolas, wherein the curvature of at least one of the two parabolas varies as a function of an internal thickness of the heart. 15. The method according to claim 11 comprising: providing an attenuation profile on a path of segment(s); andidentifying one or more parameters in an analytic curve of a preset type with a shape conforming to the profile. 16. The method according to claim 15 comprising: identifying the parameters of two parabolas, one with a shape complying with an external silhouette of the profile, the other with the... | General Electric Company | A61B005/05 A61B006/00 A61B005/02 G06K009/00 G01N023/04
US07468038 | B2 | 2008-12-23 | Non-invasive electronic method and apparatus for measuring blood pressure | 2005-12-20 | Provided is a non-invasive electronic method and apparatus for measuring blood pressure, in which the data is processed with an algorithm of non-linear fitting for recovering trend envelope of an oscillating PW by a firmware, so that the trend envelope of PW amplitude can be recovered accurately, and the obtained average pressure shows higher degree of agreement with the real situation in clinic. In this apparatus, a 3-way solenoid valve (20) is arranged between a cuff (10) and a first pressure sensor (30), wherein the common port of the 3-way solenoid valve (20) is connected with the first pressure sensor (30), the normally open port thereof connected with the cuff (10), and the normally closed port thereof connected with the air atmosphere during zeroing process. Moreover, an independent timing circuit (45) is provided additionally. In the measurement method of the present invention, the zeroing is performed with the help of the 3-way solenoid valve (20). Once a state of overtime or overpressure occurs, an air bump (50) is immediately closed, and a quick-deflation solenoid valve (60) is opened so that the pressure in the cuff (10) is released to ensure safe application. | 1. A non-invasive electronic method for measuring blood pressure, which includes the following steps: a. inflating a cuff communicated with a first pressure sensor which has been zeroed, and then stopping inflation after a cuff pressure reaches a preset value;b. deflating the cuff at a preset speed and detecting whether a pulse signal occurs;c. after the pulse signal is detected, processing data with an algorithm for recovering a trend envelope of an oscillating pulse wave, and displaying the measured blood pressure;wherein during the step a, the first pressure sensor is zeroed in a state of communicating with atmosphere; andduring the step c, the algorithm for recovering the trend envelope of the oscillating pulse wave is a nonlinear fitting algorithm, wherein data processing during the step c includes the following steps:C1: sampling during the deflation process, and thereby obtaining a cuff pressure X and a pressure pulse wave amplitude Y of each sampling point;C2: taking a cuff pressure X0 and a pressure pulse wave amplitude Y0 of a first sampling point among the sampling points as initial values, and selecting a pressure step Δ;C3: selecting at least three adjacent sampling points while sequentially moving point-by-point on the basis of the sampling points, and then performing data processing as follows for the sampling points until the last sampling point;piecewise performing trend term fitting by constructing a multiple-order curve with said at least three adjacent sampling points, generating a cuff pressure Xm and a pressure pulse wave amplitude Ym of an interpolation point between two adjacent sampling points of said at least three adjacent sampling points and then saving Xm and Ym, wherein Xm is within a cuff pressure interval of said two adjacent sampling points, Xm=X0−mΔ, while each of n and m is a natural number which increases by one during such fitting process;C4: performing smooth processing by taking said pressure pulse wave amplitude of each saved interpolation point as an intermediate value, averaging the amplitude of each interpolation point with amplitudes of at least two adjacent interpolation points while moving point-by-point, and thereby obtain mg the smoothed amplitude of each interpolation point and then saving it; andC5: searching the maximum of pulse wave amplitude among the saved and smoothed amplitudes, constructing a multiple-order curve with the data of the point corresponding to the maximum of pulse wave amplitude and at least one point on either side thereof, and calculating the maximum of pulse wave amplitude of the multiple-order curve as the extreme value of trend envelope curve of the oscillating pulse wave, wherein a cuff pressure corresponding to the extreme value is an average pressure. 2. The non-invasive electronic method for measuring blood pressure according to claim 1, wherein a 3-way solenoid valve is used so that the first pressure sensor is communicated with atmosphere during the zeroing process,... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/02
US07798979 | B2 | 2010-09-21 | Method and apparatus for penetrating tissue | 2008-04-14 | These and other objects of the present invention are achieved in a body fluid sampling system for use on a tissue site that includes an electrically powered drive force generator. A penetrating member is operatively coupled to the force generator. The force generator moves the member along a path out of a housing having a penetrating member exit, into the tissue site, stops in the tissue site, and withdraws out of the tissue site. An analyte detecting member is positioned to receive fluid from a wound created by the penetrating member. The detection member is configured to determine a concentration of an analyte in the fluid using a sample of less than 1 μL of the fluid. | 1. A body fluid sampling system for use on a tissue site, the system comprising: an electrically powered drive force generator, wherein the drive force generator is selected from a voice coil drive force generator and a rotary voice coil drive force generator;a penetrating member operatively coupled to said force generator, said force generator moving said penetrating member along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;a sample chamber with capacity of less than 1 μL positioned to receive fluid from a wound created by said penetrating member; andan analyte detecting member configured to determine a concentration of an analyte in the fluid using a sample of less than 1 μL of the fluid in the sample chamber. 2. The system of claim 1, wherein the driver is a voice coil drive force generator. 3. The system of claim 1, wherein the driver is a rotary voice coil drive force generator. 4. The system of claim 1, further comprising: a tissue stabilizer device coupled to the housing. 5. The system of claim 4, wherein the tissue stabilizer device is configured to enhance fluid flow from said target tissue. 6. The system of claim 4, wherein the tissue stabilizer device creates a stretching of said skin surface. 7. The system of claim 4, wherein the tissue stabilizer device includes a plurality of protrusions. 8. The system of claim 4, wherein the tissue stabilizer device applies a vacuum to said target tissue. 9. The system of claim 4, wherein the tissue stabilizer device is configured to apply a force to said target tissue and cause the target tissue to press in an inward direction relative to the housing member. 10. The system of claim 4, wherein the tissue stabilizing member applies a stimulation to said target tissue. 11. The system of claim 1 further comprising a vacuum source to provide a low pressure environment to draw fluid from a wound created by the penetrating member in said tissue. | Pelikan Technologies, Inc. | A61B005/00 A61B017/14 A61B017/32 B65D081/00
US07444756 | B2 | 2008-11-04 | Surgical depth instrument | 2007-01-23 | A depth gauge for measuring the depth of a hole in a bone is disclosed having a digital readout for providing the measurement to a surgeon. The depth gauge has a probe with a tip inserted into and positioned proximate to the depth the hole, and has a reference member adjustably positioned relative the probe and against the bone proximate the hole. A measuring device is provided with the depth gauge for measuring the relative distance between the tip and the reference member, and the measuring device has a large, digital display for providing the relative distance. | 1. An electronic instrument for measuring the depth of a bore, the instrument comprising: a generally cylindrical sleeve defining a lengthwise central axis;a probe which is disposed for reciprocating movement within the sleeve along the lengthwise central axis, the probe having a radially extending protrusion near a tip of the probe;an electronic sensor that generates an electrical signal relating to the distance between a tip of the sleeve and the protrusion of the probe; anda display generally centered on and oriented along the lengthwise central axis, which receives a signal from the sensor and displays information relating to the distance between the tip of the sleeve and the protrusion of the probe. 2. The electronic instrument of claim 1 wherein the display is an electronic digital display. 3. The electronic instrument of claim 2 wherein the electronic digital. 4. The electronic instrument of claim 2 wherein the electronic digital display displays a numerical representation of the distance between the tip of the sleeve and the protrusion of the probe. 5. The electronic instrument of claim 2 wherein the electronic digital display displays information relating to a fastener to be secured within the bore. 6. The electronic instrument of claim 1 wherein the diameter of the probe at the radially extending protrusion is smaller than the diameter of the bore. | Eidosmed LLC | A61B005/107
US07455660 | B2 | 2008-11-25 | Locking guidewire straightener | 2004-10-07 | A guidewire straightener assembly (50,200,300,400) including a guidewire tube (56,250,450) with a tubular passage therethrough having a first diameter and a distal end, optionally including a collar (58,260,462) having a second diameter larger than the first diameter. The assembly also includes a guidewire straightener (10,210,410) having a distal end (12,212,412), a proximal end (14,214,414) having an exterior surface with a diameter sized to fit within the guidewire tube, and a longitudinal passageway (18,218,418) therethrough sized to accept the guidewire (52,202,402). The proximal end of the guidewire straightener further includes a biasing member (20,220,420) with at least one finger (26,216,416) therealong. The assembly further includes a guidewire (52,202,402) disposed within the tubular passage and the longitudinal passageway(18,218,418). When the guidewire straightener (10,210,410) is disposed within the guidewire tube in a first position with its biasing member in the tubular passage, the at least one finger (26,216,416) is biased by the tubular passage into frictional engagement with the guidewire; when disposed in a second position, the at least one finger is disposed away from frictional engagement with the guidewire. | 1. A guidewire straightener assembly comprising: a guidewire tube including: a tubular passage sized to accept a guidewire and having a distal end;and wherein the guidewire tube includes a collar at its distal end, wherein the collar comprises a longitudinal opening having a diameter larger than the diameter of a guidewire straightener proximal end,the guidewire straightener including: a distal end;a proximal end having an exterior surface sized to fit within the guidewire tube distal end wherein the proximal end of the guidewire straightener further includes a biasing member with at least one finger extending proximal from the proximal end to enter the guidewire tube distal end; anda longitudinal passageway having a diameter sized to accept the guidewire, the passageway extending between the proximal end and the distal end; anda guidewire disposed within the tubular passage and the longitudinal passageway; such that, when the guidewire straightener proximal end is disposed within the guidewire tube in a first position in the tubular passage, the at least one finger is biased by the tubular passage into frictional engagement with the guidewire, and, when the guidewire straightener proximal end is disposed in a second, more distal position with respect to the guidewire tube, the at least one finger is disposed within the guidewire tube distal end, and away from frictional engagement with the guidewire. 2. The guidewire straightener assembly according to claim 1, wherein the at least one finger comprises a plurality of fingers. 3. The guidewire straightener assembly according to claim 1, wherein the distal end of the guidewire straightener has a generally conical shape. 4. The guidewire straightener assembly according to claim 1, wherein, when the guidewire straightener is in its second, more distal position, its biasing member is spaced distally of the guidewire tube distal end. 5. The guidewire straightener assembly according to claim 1, wherein, when the guidewire straightener is in its second, more distal position, its biasing member is disposed within a recess of the collar sufficiently large to enable the biasing member to be essentially free of biasing engagement with an interior surface thereof. 6. A method of inserting a guidewire into a patient comprising: providing the guidewire straightener assembly of claim 1;distally disposing the guidewire straightener from the distal end of the tube such that the at least one finger is translated from the tubular passage to the collar such that the at least one finger is biased away from the guidewire;inserting the proximal end of the guidewire straightener into an introducer device; andadvancing the guidewire distally through the introducer device and into the patient. 7. The method according to claim 6, further comprising, after inserting the guidewire into the patient a desired length, removing the guidewire straightener and the guidewire tube from the guidewire by sliding the guidewire str... | Medical Components, Inc. | A61M005/00 A61M031/00 A61M025/00 A61B005/00
US07438689 | B2 | 2008-10-21 | Method for arbitrary two-dimensional scaling of phonocardiographic signals | 2003-10-09 | A heart signal comprising both first and second heart sounds and murmurs occurring due to various heart conditions is modeled by a sum of sinusoids automatically selected to represent the heart sound without a noise component. In an analysis step the signal parameters are measured, in an interpolation and transformation part Ω(t) and A(t) are interpolated from one window position to the next window position, and a synthesis part reconstructs the transformed heart signal. The transformations are transformations in either time scale (with unchanged frequency content) or frequency scale (with unchanged time progression), or both. Furthermore a dynamic transformation of an arbitrary heart signal to a fixed number of beats per minute is made possible. | 1. A method for arbitrary time and/or frequency scaling by analysis, classification and subsequent re-synthesis of phonocardiographic signals obtained from a transducer and subjected to digital spectral analysis, characterised in that the signals are converted on a running basis to a sinusoidal model representation, that a time and/or frequency axis scaling is defined and used to control the amplitudes and phases of the sinusoids, which are subsequently added to create a time and/or frequency scaled representation of the phonocardiographic signal and comprising the further steps of: obtaining the frequency content of the signal by applying a Short Time Fourier Transform with overlapping segments, a Discrete Fourier Transform being performed on each segment,performing a frequency peak search on each segment by consecutive removal of the spectral components having the highest energy content, identifying each peak by its frequency valuerepeating the peak search until a maximum number of peaks have been identified or until the energy content of the last peak is below a preset minimum,establishing a segment-by-segment map of spectral peaks, said peaks forming a track over timeoptionally subjecting the frequency values of the spectral peaks to a multiplicationdefining a synthesis frame [frame m] of time, based on said segmentsoptionally subjecting each frame to a multiplication of the time scaleadjusting the phase of sine generators centered on the frequencies of the tracksadjusting the amplitudes of said sine generatorssumming the outputs of all sine generators active at any one instant for a given frame length T. 2. A method according to claim 1, for scaling a phonocardiographic signal on the frequency axis by a desired factor, characterised in that the spectral peaks are multiplied by a factor q. 3. A method according to claim 1, for scaling a phonocardiographic signal on the time axis by a desired factor, characterised in that the frame length is multiplied by a factor p. 4. A method according to claim 1, for autoscaling a phonocardiographic signal on the time axis, characterised in that the scaling factor p is set such that the frame length [T] multiplying factor is equal to the heart rate divided by 60. 5. A method according to claim 1, or 3, characterised in that the number of sine generators is equal to or below 50. 6. A method for arbitrary time and/or frequency scaling by analysis, classification and subsequent re-synthesis of phonocardiographic signals obtained from a transducer and subjected to digital spectral analysis, characterised in that the signals are converted on a running basis to a sinusoidal model representation, that a time and/or frequency axis scaling is defined and used to control the amplitudes and phases of the sinusoids, which are subsequently added to create a time and/or frequency scaled representation of the phonocardiographic signal and including the following steps: obtaining the frequency content of the signal b... | Bang & Olufsen Medicom A/S | A61B005/02
US07465275 | B2 | 2008-12-16 | Method and apparatus for measuring functional residual capacity (FRC) and cardiac output(CO) | 2002-05-29 | The present invention relates to a method and an apparatus for measuring Functional Residual Capacity (CO). Further, it relates to a method and an apparatus for measuring cardiac output (CO). The method comprises the steps of determining the amount of oxygen and carbon dioxide being inhaled and exhaled during at least two breathing cycles, identifying at least one difference between said breathing cycles, and estimating the FRC and/or CO based on the determined uptake of oxygen and excretion of carbon dioxide and said identified difference. | 1. A method for estimation of functional residual capacity (FRC) in a person, comprising: establishing a breathing pattern;determining the amount of oxygen and carbon dioxide being inhaled and exhaled during a first breathing cycle comprised by said breathing pattern;accomplishing a first change in said breathing pattern of known magnitude;determining the amount of oxygen and carbon dioxide being inhaled and exhaled during a second breathing cycle;estimating the FRC based on the determined amounts of oxygen and carbon dioxide in said first and second breathing cycles. 2. The method of claim 1, wherein said first change of breathing pattern is selected from apnea, a breath volume larger than the breath volume of said first breathing cycle, a breath volume smaller than the breath volume of said first breathing cycle, a duration difference between the different phases in said first and second breathing cycles, a difference in dead space volume. 3. The method of claim 1, wherein said second breathing cycle is followed by a second change in breathing pattern of known magnitude compensatory to said first change in breathing pattern of known magnitude to re-establish a steady state allowing a new estimation of FRC with minimal delay. 4. The method of claim 3, wherein said first change in breathing pattern is an increase of the time period between exhalation and inhalation or a breath with substantially decreased breath volume and said second change in breathing pattern is a breath with substantially increased breath volume. 5. The method of claim 3, wherein said first change in breathing pattern is a breath with substantially increased breath volume and said second change in breathing pattern is an increase of the time period between inhalation and exhalation or a breath with substantially decreased breath volume. 6. The method of claim 2, comprising an apnea period of from 5 to 25 seconds. 7. The method of claim 6, comprising an apnea period of from 8 to 15 seconds. 8. The method of claim 1, wherein the amount of oxygen and carbon dioxide being inhaled and exhaled is determined by measuring the end-tidal concentration of oxygen and carbon dioxide and the inspirational concentration of oxygen. 9. The method of claim 2, wherein said difference in dead space volume is increased during a limited period of time. 10. The method of claim 9, wherein said period of time is short enough to avoid re-circulation of carbon dioxide in mixed venous blood. 11. The method of claim 1, comprising the determination of oxygen uptake and carbon dioxide excretion during said first breathing cycle based on the determined amount of oxygen and carbon dioxide being inhaled and exhaled. 12. The method of claim 1, wherein the estimation of FRC is repeated for continuous monitoring of FRC. 13. The method of claim 1, wherein the estimation of FRC comprises several of said determinations so as to provide an average FRC estimation. 14. The method of claim 1, wherein said ... | GE Healthcare Finland Oy | A61B005/02 G01N001/22 G01N031/00 G01N033/497
US07808391 | B2 | 2010-10-05 | Remote caregiver support system | 2008-06-04 | The current invention provides a remote caregiver support system which allows caregivers to provide informal care to care receivers. The systems of the current invention comprise a caregiver system and a care receiver system operating on personal computers. Each system is linked by a communications network. Preferably, the communications network is a high speed internet connection. Additionally, the current invention provides methods for delivering informal care to a care receiver remotely located from the caregiver. In the methods of the current invention, the caregiver accesses a remote caregiver support system and interacts with the care receiver providing text reminders and other informal care necessary to ensure the continued health and wellbeing of the care receiver. | 1. A non-medical system comprising: a caregiver system;a care receiver system to facilitate remote informal care for at least one care receiver; and,a communication network linking the caregiver system and the care receiver system; wherein the care receiver system is operable in an unattended mode to enable operation of the care receiver system without any input on the part of the care receiver whereby the care receiver provides no input to and interacts with the care receiver system. 2. The system of claim 1 wherein the care receiver system is operable in an attended mode whereby the care receiver provides an input to and interacts with the care receiver system. 3. The system of claim 2 wherein the input enables the care receiver to acknowledge a reminder. 4. The system of claim 2 wherein the input enables the care receiver to accept, reject and initiate a videoconference. 5. The system of claim 1 further comprising: a care receiver client enabling interaction between the caregiver system and the care receiver system; and a caregiver client for accessing and configuring the care receiver system. 6. The system of claim 5 wherein the caregiver client accesses and configures a graphical user interface of the care receiver system. 7. The system of claim 5 wherein the caregiver client enables a caregiver to establish a videoconference with the care receiver system, the videoconference establishable in an interactive mode or an observational mode. 8. The system of claim 5 wherein the caregiver client enables the creation, editing and display of a slideshow on the care receiver system. 9. The system of claim 5 wherein the caregiver client enables the creation, editing and display of a reminder on the care receiver system. 10. The system of claim 9 wherein the display of the reminder is accompanied by an audible tone to draw attention to the reminder. 11. The system of claim 9 wherein the reminder is accompanied by a recording of the caregiver's voice. 12. The system of claim 9 wherein the reminder is accompanied by a visual cue. 13. The system of claim 9 wherein the care receiver system further comprises a database storing the reminder and any acknowledgement to the reminder generated by the care receiver system. 14. The system of claim 5 further comprising a plurality of global settings on the care receiver system adjustable by the caregiver client. 15. The system of claim 5 further comprising a plurality of multimedia settings dynamically adjustable by the caregiver client. 16. The system of claim 5 further comprising a client server monitor for monitoring operation of the care receiver client and automatically restarting the care receiver client if the care receiver client is detected to be unresponsive or operating outside of predetermined parameters. 17. The system of claim 1 further comprising a sensor for initiating a signal to be transmitted to an environment and event monitoring system, the signal analyzed against predetermined... | - | G08B023/00 A61B005/00
US07457652 | B2 | 2008-11-25 | Method and circuit for indicating quality and accuracy of physiological measurements | 2003-11-12 | Sensors and monitors for a physiological monitoring system having capability to indicate an accuracy of an estimated physiological condition. The sensor senses at least one physiological characteristic of a patient and is connectable to a monitor that estimates the physiological condition from signals detected by the sensor. The sensor includes a detector for detecting the signals from the patient which are indicative of the physiological characteristic. The sensor is associated with a memory configured to store data that defines at least one sensor signal specification boundary for the detected signals. The boundary is indicative of a quality of the signals and an accuracy of the physiological characteristic estimated from the signals by the monitor. The sensor further includes means for providing access to the memory to allow transmission of the data that defines the at least one sensor boundary to the monitor. | 1. A system comprising a sensor and a monitor for sensing at least one physiological characteristic of a patient, the sensor being connectable to a monitor that estimates a physiological characteristic from signals detected by the sensor, the sensor comprising: a detector for detecting the signals from the patient which are indicative of the physiological characteristic;a memory connected with the sensor, said memory being physically located on one of a sensor body, sensor cable, sensor connecting plug or a sensor adapter module, and said memory being configured to store data defining at least one sensor signal specification boundary for the detected signals, the sensor signal specification boundary being indicative of a quality of the signals and an accuracy of the physiological characteristic estimated from the signals by the monitor, wherein the sensor signal specification boundary includes limits for an AC modulation component and DC component of the signals; andmeans for providing access to the memory to allow transmission of the data defining the at least one sensor boundary to the monitor;wherein the monitor determines to display or not display the estimate of the physiological characteristic based on the signals and their relationship relative to a plurality of sensor signal specification boundaries and to a plurality of monitor boundaries preprogrammed into the monitor. 2. The system of claim 1, wherein the monitor computes calculated values, having AC and DC components, from the signals, wherein the sensor signal specification boundary constitutes limits on the AC and DC components of the calculated values, and wherein the AC and DC components are dependent on either a physiological status of the patient, sensor type, or sensor location. 3. The system of claim 1, wherein the signals detected from the patient include first and second sets of signals derived from detected light scattered from the patient, the light having fist and second wavelengths, the signals derived from detected light each having an AC modulation component and a DC component, and the sensor signal specification boundary including limits on the AC and DC components. 4. The system of claim 3, wherein the signals derived from detected light are indicative of an arterial oxygen saturation of the patient. 5. The system of claim 1, wherein the memory comprises a digital memory configured to store a digital representation of the at least one sensor signal specification boundary, and wherein the physiological characteristic is arterial oxygen saturation. 6. A monitor for providing an indication of an accuracy of an estimated physiological condition of a patient, the monitor being connectable to a sensor that detects signals indicative of at least one physiological characteristic of the patient, the monitor comprising: a receiving circuit configured to receive the signals indicative of the at least one physiological characteristic from the sensor and data defining at lea... | Mallinckrodt Inc. | A61B005/1455
US07463920 | B2 | 2008-12-09 | Rapid MR visualization of interventional devices | 2005-03-25 | A method and apparatus for visualizing an interventional device in a body using a magnetic resonance (MR) imager. In operation an MR pulse sequence is applied to the body so as to develop MR signals, which are processed so as to acquire one or more images. An operator of the imager uses a hands-free device to control parameters of the applied pulse sequence so as to cause the images to show an intensity distortion artifact in the area of the interventional device that is increased as compared to intensity distortion artifact in conventional MR images; and, after an image distortion which is expected to be produced by the interventional device is seen, the operator changes the applied pulse sequence parameters, in a “hands-free” manner, so that said one or more acquired images show an intensity distortion artifact in the area of the interventional device that is reduced as compared to said increased intensity distortion. | 1. A method for identifying and enhancing visualization of an interventional device using a magnetic resonance (MR) imaging system, comprising the following steps: employing a pulse sequence in MR imaging a body portion of a patient and deriving response MR signals for use in providing MR images of said body portion;processing the MR signals to provide MR images of said body portion, in response to first parameters of the applied pulse sequence predetermined for acquisition of a first gradient echo image, using a first amplitude readout gradient and first receiver bandwidth, said first gradient echo image including a first intensity distortion artifact in the area of an interventional device andsecond parameters of the applied pulse sequence predetermined for acquisition of a second gradient echo image, using a second amplitude readout gradient and second receiver bandwidth, said second gradient echo image including a second intensity distortion artifact in the area of an interventional device, said first parameters providing a relatively low amplitude readout gradient and relatively low receiver bandwidth and increased spatial image intensity distortion artifact compared to said second parameters and are selected to facilitate locating said interventional device relative to an image slab; andselecting either said first or second parameters in response to user command. 2. The method of claim 1, including using a hands-free device for controlling selection of said first and second parameters to selectively acquire gradient echo images using a relatively long echo time to show said increased spatial image intensity distortion artifact, and selectively acquire gradient echo images using a relatively short echo time to show a reduced intensity distortion artifact. 3. The method of claim 1, including using a hands-free device for controlling selection of said first and second parameters to selectively acquire gradient echo images using a relatively low amplitude readout gradient, relatively low receiver bandwidth, and relatively long echo time to show said increased spatial image intensity distortion artifact, and selectively acquire gradient echo images using a relatively high amplitude readout gradient, relatively high receiver bandwidth, and relatively short echo time to show a reduced intensity distortion artifact. 4. The method of claim 1, including using a hands-free device for controlling selection of said first and second parameters selectively acquire gradient echo images using a relatively low amplitude readout gradient, relatively low receiver bandwidth, and relatively long echo time a person who guides the interventional device in the body to show said increased spatial image intensity distortion artifact, and selectively acquire spin echo images so as to show a reduced intensity distortion artifact. 5. The method of claim 1, including the step of, in response to user command, alternating between acquisition of a sequence of images wit... | Siemens Medical Solutions USA, Inc. | A61B005/00 G01V003/00
US07470420 | B2 | 2008-12-30 | Optical determination of glucose utilizing boronic acid adducts | 2005-12-07 | The present invention concerns an improved optical method and optical sensing device for determining the levels of polyhydroxyl-substituted organic molecules in vitro and/or in vivo in aqueous media. The range of detection is between about 400 and 800 nm. In particular, a sensory devise is implemented in a mammal to determine sugar levels. Specifically, a dye is combined with a conjugated nitrogen-containing heterocyclic aromatic boronic acid-substituted bis-onium compound in the presence of a sugar, such as fructose or glucose. The viologens are preferred as the aromatic conjugated nitrogen-containing boronic acid substituted compounds. The method is useful to determine sugar levels in a human being. | 1. A device for optically determining an analyte concentration, which device comprises: an analyte permeable component;a fluorophore associated with the analyte permeable component and configured to absorb light at a first wavelength and emit light at a second wavelength, wherein the fluorophore comprises a substituted pyrene sulfonate;a quencher associated with the analyte permeable component and configured to modify the light emitted by the fluorophore by an amount related to the analyte concentration wherein the quencher comprises a boronic acid substituted viologen;a light source; anda detector. 2. The device of claim 1, wherein the analyte permeable component comprises a polymer matrix. 3. The device of claim 2, wherein the polymer matrix is a hydrogel. 4. The device of claim 3, wherein the hydrogel is formed by polymerization of hydrophilic monomers or by cross-linking hydrophilic polymers. 5. The device of claim 4, wherein the hydrophilic monomers are selected from the group consisting of 2-hydroxyethyl-methacrylate, polyethylene glycol methacrylate, methacrylic acid, hydroxyethyl acrylate, N-vinyl pyrrolidone, acrylamide, N,N′-dimethyl acrylamide, methacryloylaminopropyl trimethylammonium chloride, diallyl dimethyl ammonium chloride, sodium sulfopropyl methacrylate, optionally having cross-linkers selected from ethylene dimethacrylate, polyethyleneglycol dimethacrylate (PEGDMA), methylene-bis-acrylamide and trimethylolpropane triacrylate, and combinations thereof. 6. The device of claim 2, wherein the polymer matrix is insoluble in water. 7. The device of claim 6, wherein the water-insoluble polymer matrix is prepared from monomers selected from the group consisting of HPTS(Lys-MA)3, HPTS-MA, HPTS-CO2-MA, APTS-BuMA, and APTS-DegMA. 8. The device of claim 7 wherein the water-insoluble polymer matrix comprises copolymers 2-hydroxyethylmethacrylate (HEMA) and polyethylene glycol dimethacrylate or N,N′-dimethylacrylamide and methylene-bis-acrylamide. 9. The device of claim 1, wherein the analyte permeable component comprises a membrane, which confines the fluorophore and the quencher. 10. The device of claim 1, wherein the substituted pyrene sulfonate is a pyranine derivative selected from the structure:wherein R1, R2, and R3 are each —NHR4, R4 is —CH2—CH2(—O—CH2—CH2)n—X1;wherein X1 is —OH, OCH3—CO2H, —CONH2, —SO3H, or —NH2; and n is between about 70 and 10,000. 11. The device of claim 1, wherein the boronic acid substituted viologen further comprises an aromatic boronic acid moiety. 12. The device of claim 11, wherein the quencher is prepared from a precursor selected from the group consisting of:wherein X is bromide or chloride. 13. The device of claim 1, wherein the viologen is prepared from a precursor having a structure selected from the group consisting of: 14. The device of claim 1, wherein the quencher is configured to bind an amount of the analyte. 15. The device of claim 13, wherein the quencher is configured to reduce... | The Regents of the University of California | A61B010/00 A61B005/00 A61B008/00
US07440788 | B2 | 2008-10-21 | Oral health measurement clamping probe, system, and method | 2004-08-26 | Devices, systems, and methods determine the health of oral objects by providing objective measurements using a detachable probe body. The detachable probe body may isolate reusable system components (including an electromagnetic signal detection, signal transmission, energy generation, and or energy transmitting components) from the oral cavity, optionally by encasing at least a portion of one or more of these components in a sheath or the like. A window of the probe body maintains sterile isolation and transmits electromagnetic energy to and or signals from the oral object. Accuracy can be enhanced by a clamp or other structure for engaging a surface of the oral object so as to maintain a fixed alignment between the signal receiver and the oral object. | 1. A detachable probe body for use with an oral health measurement system for oral health measurements of an oral object within a mouth of a patient, the mouth having oral surfaces, the health measurement system having reusable components including an electromagnetic signal receiver coupled to a processor for evaluating a health-related characteristic of the mouth from an electromagnetic signal, the probe body comprising: a window capable of transmitting the electromagnetic signal from the oral object to the signal receiver;a casing extending from the window, the casing and window together defining an opening for receiving the signal receiver so as to provide isolation of the reusable components from the mouth; andat least one probe surface oriented to engage one or more oral surfaces so as to maintain alignment between the signal receiver and the oral object. 2. The probe body of claim 1, wherein the at least one probe surface comprises a first probe surface region and a second probe surface region, the probe receiving the oral object with the first and second probe surface regions. 3. The probe body of claim 2, wherein the probe body comprises a first probe body portion having the first probe surface region and a second probe body portion having the second probe surface region, the second body portion movable relative to the first body portion. 4. The probe body of claim 3, further comprising a biasing structure urging the first probe surface region toward the second probe surface region when the oral object is at a desired position there between so as to accommodate a range of differing oral object sizes. 5. The probe body of claim 4, wherein the first portion slidingly receives the second portion, the biasing structure comprising a spring urging the second region toward the first region when the oral object is received there between. 6. The probe body of claim 3, further comprising a ratchet coupling the first probe body portion to the second probe body portion, the ratchet oriented to allow movement of the first probe surface region toward the second probe surface region and to inhibit movement of the first probe surface region away from the second probe surface region. 7. The probe body of claim 3, wherein at least one of the first body portion, the second body portion, and an intermediate portion coupling the first and second body portions comprises a deformable material. 8. The probe body of claim 7, wherein the deformable material resiliently biases the first probe surface region toward the second probe surface region. 9. The probe body of claim 7, wherein the deformable material defines the first probe surface region, the deformable material being sufficiently soft to deform against the oral object when a dentist manually applies pressure. 10. The probe body of claim 5, wherein the first probe body portion is configured to move toward the second body portion when a pressure is applied with an orientation lateral relative to the... | Kelvyn Enterprises, Inc. | A61B005/1455
US08287531 | B2 | 2012-10-16 | Catheter with ablation needle and mapping assembly | 2008-06-17 | The invention is directed to a catheter that creates enhanced lesions using a needle electrode and can simultaneously map electrical activity at a plurality of points using an enhanced mapping assembly. The catheter comprises an elongated catheter body having at least one lumen extending longitudinally therethrough. A needle control handle is provided at the proximal end of the catheter body. A needle electrode assembly extends through the catheter body and needle control handle and has a proximal end attached to the needle control handle and a distal end within the distal end of the catheter body. A mapping assembly is mounted at the distal end of the catheter body and comprises at least two flexible spines. Each spine has a proximal end attached at the distal end of the catheter body and a free distal end. Each spine carries at least one electrode. The distal end of the needle electrode assembly is extendable past the proximal end of the mapping assembly upon manipulation of the needle control handle. | 1. An injection catheter comprising: an elongated catheter body having at least one lumen extending longitudinally therethrough;a needle control handle at a proximal end of the catheter body;a needle electrode assembly extending through the catheter body and needle control handle and having a proximal end attached to the needle control handle and a distal end within a distal end of the catheter body;a mapping assembly, having proximal and distal ends, mounted at the distal end of the catheter body and comprising at least two flexible spines, each spine having a proximal end attached at the distal end of the catheter body and a free distal end, wherein each spine carries at least one electrode; andat least one location sensor;wherein the distal end of the needle electrode assembly is extendable past the proximal end of the mapping assembly upon manipulation of the needle control handle. 2. The catheter of claim 1, wherein each spine comprises a non-conductive covering having a support arm that has shape memory disposed therein. 3. The catheter of claim 2, wherein each support arm comprises Nitinol. 4. The catheter of claim 1, wherein the mapping assembly is moveable between an expanded arrangement, in which each spine extends radially outward from the catheter body, and a collapsed arrangement, in which each spine is disposed generally along a longitudinal axis of the catheter body. 5. The catheter of claim 4, wherein, when the mapping assembly is in its expanded arrangement, each spine extends radially outwardly from the catheter body and forms an arced shape. 6. The catheter of claim 1, wherein the mapping assembly comprises at least five spines. 7. The catheter of claim 1, further comprising an electrode lead wire having a first end electrically connected to the needle electrode assembly and a second end electrically connected to a source of ablation energy. 8. The catheter of claim 1, wherein the needle electrode assembly comprises a distal tubing having proximal and distal ends and comprising an electrically conductive material and a proximal tubing comprising a material more flexible than the distal tubing and having a distal end attached, directly or indirectly, to the proximal end of the distal tubing. 9. The catheter of claim 8, wherein the distal tubing comprises metal and the proximal tubing comprises plastic. 10. The catheter of claim 8, wherein the distal tubing comprises Nitinol or stainless steel. 11. The catheter of claim 8, wherein the proximal tubing comprises polyimide or PEEK. 12. The catheter of claim 8, wherein the distal end of the distal tubing forms a beveled edge. 13. The catheter of claim 8, wherein the needle electrode assembly further comprises an intermediate tubing joining the proximal and distal tubings. 14. The catheter of claim 8, wherein the needle electrode assembly further comprises an outer plastic tubing fixedly attached, directly or indirectly, to the proximal and distal tubings so that the oute... | Biosense Webster, Inc. | A61B018/14 A61F002/00
US08041420 | B2 | 2011-10-18 | Apparatus for measuring tissue samples electrical impedance | 2006-06-02 | Apparatus for measuring the electrical impedance of a tissue sample comprises a probe and a sheath comprising an elongated tubular body having one closed end and one open end providing an internal cavity. The sheath is composed of a material which when contacted with a tissue sample is capable of providing a conductive path through the sheath between the electrodes and the tissue sample. The resistivity of the material forming the sheath when contacted with the tissue sample is greater than the resistivity of the tissue sample. The probe is received within the internal cavity of the sheath. The sheath is impervious to bacteria and viruses. | 1. An apparatus configured to measure the electrical impedance of a tissue sample, wherein the apparatus comprises: (a) an elongate probe having electrodes at one end, wherein the electrodes are configured to transmit an electrical signal between the apparatus and the tissue in contact with the apparatus; and(b) a sheath removably disposed on the elongate probe, the sheath comprising an elongate tubular body having a closed end and an open end and defining an internal cavity, in which the end of the probe on which the electrodes are provided is received within the cavity, and in which at least a portion of the sheath is formed from a material which is capable of providing a conductive path through the sheath between the electrodes and the tissue sample when the sheath is contacted with the tissue sample, and in which the resistivity of the material when the material is contacted with the tissue sample is greater than the resistivity of the tissue sample, and in which the material of the sheath is formed from a polymeric material which has a porous structure such that in use the polymeric material can be impregnated with body fluid to provide an ionically conductive path through the sheath between the electrodes and the tissue sample, and in which the porous material has a pore size effective to block the passage of the polio virus through the porous material, wherein the sheath is removable from the probe after use, such that the probe is configured to be covered with another sheath and used on another patient without requiring sterilization of the probe. 2. The apparatus as claimed in claim 1, in which the sheath is composed of one or more materials selected from cellulose acetate, polyethersulphone, polyamide, and cellulose. 3. The apparatus as claims in claim 1, which includes a quantity of an aqueous solution of a salt between the electrodes and the sheath. 4. The apparatus as claimed in claim 1, in which the material of the sheath at the closed end in the vicinity of the electrodes is different from the material in other parts of the sheath. 5. The apparatus as claimed in claim 1, wherein the ratio of the resistivity of the material of the sheath, when impregnated with body fluid, to the resistivity of the tissue sample is at least about 10. 6. An apparatus configured to measure the electrical impedance of a tissue sample, wherein the apparatus comprises: (a) an elongate probe having electrodes at one end, wherein the electrodes are configured to transmit an electrical signal between the apparatus and the tissue in contact with the apparatus; and(b) a sheath removably disposed on the elongate probe, the sheath comprising an elongate tubular body having a closed end and an open end and defining an internal cavity, in which the end of the probe on which the electrodes are provided is received within the cavity, and in which at least a portion of the sheath is formed from a material which is capable of providing a conductive path through t... | The University of Sheffield Sheffield Teaching Hospitals NHS Foundation Trust | A61B005/05
US08075509 | B2 | 2011-12-13 | Medical apparatus | 2005-08-25 | A medical apparatus (1) comprises a user interface (30), a number of sensors (16,27,21,22,29) and a control system (18) programmed for displaying on the user interface screen (31) a number of indicia (37), each relating to a corresponding parameter and comprising an indicium border (45) delimiting an indicium selection area (45a); the control system allows selection of an indicium for identifying the parameter a user intends to modify, allows setting of a new set value for the selected parameter; the control system also receives from the sensors (29) a measure of the actual value for the parameter and displays on the indicium selection area (45a) of the indicium both the actual value and the set value for the same parameter. The programming software for the control system is also disclosed. | 1. Medical apparatus comprising: at least a user interface for enabling setting of a plurality of parameters pertinent to operation of said apparatus or pertinent to a process to be performed by said apparatus, the user interface including at least a screen;a number of sensors for detecting actual values of said parameters;a control system for controlling operation of said medical apparatus, the control system being connected to said sensors, being responsive to actions by a user on said user interface, and being programmed for executing the following steps: displaying on said screen a number of indicia, each relating to a corresponding parameter and comprising an indicium border delimiting an indicium selection area;allowing selection of an indicium among said number for identifying the parameter a user intends to modify,allowing setting of a new set value for said parameter;receiving from at least one of the sensors a measure of the actual value for said parameter;displaying on the indicium selection area of the indicium both the actual value and the set value for the corresponding parameter,the control system being programmed to contemporaneously display in the indicium selection area of a same parameter the following regions: a first region displaying the set value a user has set for the parameter,a second region, distinct from the first region, displaying the current actual value for the same parameter,the first region and the second region occupying respective areas having different sizes on the indicium selection area. 2. Apparatus according to claim 1, wherein the control system is programmed to store in a memory the set value for each parameter and the measured value for the same parameter, the measured values being periodically updated and corresponding updated values displayed in the selection area of the corresponding indicium. 3. Apparatus according to claim 1, wherein the control system is programmed to display in the indicium selection area the actual value for the parameter and, on a different time instant, the set value a user has set for the parameter. 4. Apparatus according to claim 1, wherein the control system is programmed to: display in the indicium selection area the actual value for the parameter,detect if a set value of a parameter is under modification, and thendisplay in the indicium selection area the set value a user has set for the same parameter. 5. Apparatus according to claim 4, wherein the control system is programmed to detect if a set value of a parameter is under modification by at least detecting if the corresponding indicium has been selected. 6. Apparatus according to claim 4, wherein the control system is programmed to control the user interface not to display in the indicium selection area the actual value of the parameter during the set value modification and to re-display the actual value when the set value modification has been completed. 7. Apparatus according to claim 1, wherein the control sy... | Gambro Lundia AB | A61M037/00 C02F001/44
US09084553 | B2 | 2015-07-21 | Heart imaging with adaptive inversion time | 2008-04-02 | In a method for acquisition of magnetic resonance images of the heart, MR signals of the heart are acquired using an imaging sequence, wherein the magnetization is inverted by an RF inversion pulse before the acquisition of the MR signals; and of the heart activity is detected, and the point in time of the switching of the RF inversion pulse dependent on the detected heart activity. | 1. A method for acquiring a magnetic resonance image of a heart comprising: administering contrast agent to a patient and, while the contrast agent is effective in the heart of the patient, controlling a magnetic resonance apparatus from a control computer to acquire magnetic resonance signals from the heart by generating, and exposing the patient in the magnetic resonance apparatus to, a magnetic resonance imaging pulse sequence in successive cardiac cycles, including at least one pulse during which magnetic resonance signals from the heart are acquired, and including emitting a radio-frequency (RF) inversion pulse that inverts magnetization at least in the heart, at a point in time before said at least one pulse;while acquiring said magnetic resonance signals, detecting a cardiac signal over a plurality of successive cardiac cycles of the heart, each of said cardiac cycles comprising an R-spike, said successive cardiac cycles, and thus said R-spike, being subject to a changing rate during said acquisition of said magnetic resonance signals;providing said cardiac signal to said control computer and, in said control computer, for each respective cardiac cycle, among said successive cardiac cycles, in which said magnetic resonance data are acquired, determining a time span TF from an end of acquisition of said magnetic resonance signals of an immediately preceding cardiac cycle, and determining a time span TT from said R-spike of the respective cardiac cycle until a beginning of acquisition of magnetic resonance signals in the respective cardiac cycle, and calculating a time span TD after said R-spike in the respective cardiac cycle, dependent on TT and TF, thereby making TD dependent on said changing rate;in each of said cardiac cycles in which said magnetic resonance signals are acquired, emitting said RF inversion pulse at an end of TD calculated for that respective cardiac cycle and thereby beginning acquisition of said magnetic resonance signals in the respective cardiac cycle at a time TI, following the R-spike in the respective cardiac cycle in which said magnetic resonance signals are acquired, that because TD is dependent on said changing rate, is also dependent on said changing rate, and that minimizes a longitudinal magnetization of the myocardium of the heart; andreconstructing an image that includes the myocardium of the heart from the magnetic resonance signals acquired during the respective cardiac cycle, and making the reconstructed magnetic resonance image available in electronic form as a data file. 2. A method as claimed in claim 1 comprising employing, as said imaging sequence, a segmenting imaging sequence in which data are entered only into a portion of k-space during a cardiac cycle of the heart in the imaging sequence. 3. A method as claimed in claim 1 comprising employing a gradient echo-based imaging sequence as said imaging sequence, with a bipolar gradient as said at least one pulse, and wherein only N RF excitation p... | Siemens Aktiengesellschaft | A61B005/05 A61B005/055 G01R033/56 A61B005/0456 A61B005/00 G01R033/563 G01R033/565 G01R033/567
US09125548 | B2 | 2015-09-08 | Method and apparatus for providing data processing and control in a medical communication system | 2008-05-14 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A method comprising: transcutaneously positioning an in vivo analyte sensor in fluid contact with interstitial fluid under a skin layer, the analyte sensor configured to generate a plurality of signals corresponding to a monitored analyte level in the interstitial fluid over a predetermined time period;generating one or more data packets from the plurality of signals generated by the analyte sensor in fluid contact with the interstitial fluid;communicating the generated one or more data packets to a remote location;monitoring at the remote location, the one or more data packets for one or more predetermined conditions associated with the one or more data packets during the predetermined time period;categorizing the one or more data packets into at least one bin that corresponds to a rate of change of the analyte level monitored over the predetermined time period, wherein a position of a trend indicator is associated with the at least one bin;changing a characteristic of an output display of the trend indicator when an uncertainty parameter associated with time spacing between current and prior signals in the generated one or more data packets exceeds a predetermined threshold level, the uncertainty parameter being dynamically modified based on the time spacing between the current and prior signals in the one or more data packets; andoutputting a notification when the one or more predetermined conditions associated with the one or more data packets is detected, wherein the one or more predetermined conditions are associated with an adverse data condition associated with the analyte sensor. 2. The method of claim 1, wherein the one or more predetermined conditions includes an out of range temperature value associated with the detected one or more data packets, a persistent low detected one or more data packets, or an error condition of the analyte sensor. 3. The method of claim 2, wherein the persistent low detected one or more data packets includes at least two consecutive signals from the detected one or more data packets that are below a predetermined signal level. 4. The method of claim 3, wherein the predetermined signal level is associated with a failed analyte sensor state. 5. The method of claim 2, wherein the error condition of the analyte sensor includes unacceptable data quality of the detected one or more data packets. 6. The method of claim 1, wherein the notification is output when the same one or more predetermined conditions are detected more than once during the predetermined time period. 7. The method of claim 1, wherein outputting the notification includes one or more of displaying the generated output notification, audibly outputting the generated output notification or vibratorily outputting the generated output notification. 8. The method of claim 1, wherein the predetermined time period is less than approximately one hour from the positioning of the analyte sensor in continuous contact with the interstitial fluid un... | Abbott Diabetes Care Inc. | A61B005/145 A61B005/00 A61B005/1473 A61B005/1495
US08282568 | B2 | 2012-10-09 | Method for estimating cardiac pumping capacity | 2008-04-30 | A system and a method for estimating changes in cardiac pumping capacity in response to an intervention from data recorded by an accelerometer positioned on an outer surface of the heart. Interventions such as intravenous administration of nitroprusside, betablocker, epinephrine (adrenaline), or changes in the intravascular volume such as by fluid filling, are considered. It is shown that it is possible to directly correlate changes in cardiac pumping capacity to changes in parameters or graphical representation derived from signals characteristic to the acceleration an outer surface of the heart. In particular, it is shown how the effect of the interventions can be directly read from these parameters or graphical representations, and can thereby assist in decisions regarding the treatment of a cardiac surgery patient. The invention is preferably used in a post-surgery monitoring unit for surveillance of cardiac surgery patients. | 1. A system for assisting a decision relating to intravenous administration of medicaments to, or changes in the intravascular volume of, a cardiac surgery patient, the system comprising an accelerometer to be positioned on, in or just below an outer surface part of the myocardium in the supply region of the one or more of left anterior descending artery (LAD), circumflex artery (CX) or right coronary artery (RCA) on which the cardiac surgery has been performed, for registering an acceleration signal characteristic to accelerations of said outer surface part(s) as a function of time in three different directions; anda signal processing unit comprising:a recording module configured to record the registered acceleration signal from the accelerometer;a data processing module configured to generate a parameter or a graphical representation based on the obtained acceleration signal, or based on an integrated signal representing velocity or position, the parameter or the graphical representation correlating with cardiac pumping capacity of the patient's heart, the data processing module being configured to generate a parametric curve being a function of time in at least one of:a three dimensional graph with axes corresponding to the acceleration, or integrated signals representing relative velocities or positions, in the three different directions, ora graph with a first axis corresponding to acceleration, velocity, or position, and a second axis corresponding to correlated pressure from synchronically registered blood pressure;a storage medium configured to store of one or more baselines for the parameter or graphical representation generated over a period of time including or following intravenous administration of medicaments or changes in the intravascular volume; anda comparison module configured to compare the parameter or graphical representation generated after intravenous administration of medicaments or changes in the intravascular volume with a stored baseline to graphically present a development over time of the parameter or graphical representation, this being useful for evaluating the effect of the intravenous administration of medicaments or change in the intravascular volume and whether this should be continued, halted, or reversed,wherein the parametric curve being a function of time is based on an amplitude of the accelerometer signal or an amplitude of an integrated signal representing velocity or position. 2. The system according to claim 1, wherein the data processing software is further configured to correlate the acceleration signal, or integrated signals representing relative velocities or positions, with electrocardiogram and/or blood pressure to assign segments of the parametric curve to stages of the cardiac cycle and determine the length of one or more of said segments as a parameter or graphical representation which correlates with cardiac pumping capacity. 3. The system according to claim 1, wherein the accelerometer is ... | Bio-Medisinsk Innovasjon AS | A61B005/0402
US08721544 | B2 | 2014-05-13 | System for in-vivo measurement of an analyte concentration | 2008-03-20 | The analyte concentration, such as glucose, in a human or animal body is measured with an implantable sensor that generates measurement signals. The measurement signals are compressed through statistical techniques to produced compressed measurement data that can is easier to process and communicate. A base station carries the implantable sensor along with a signal processor, memory, and a transmitter. A display device is also disclosed that can receive the compressed measurement data from the base station for further processing and display. | 1. A system for in-vivo measurement of blood glucose concentration in a human, comprising: a base station comprising, a base power source to supply power to components of the base station;an implantable sensor for generating measurement signals that are correlated to a blood glucose concentration to be measured, the measurement signals being raw data recorded over first time intervals;a signal processor that executes instructions for processing the measurement signals to produce compressed measurement data for second time intervals which are at least ten-times longer than the first time intervals, the measurement data being compressed at least ten-times compared to the measurement signals,base memory for storing the compressed measurement data;a transmitter for wireless transmission of the compressed measurement data; and,a display device comprising, a display device power source that has greater capacity than the base power source to supply power to components of the display device;a user input for inputting data into the display device;a receiver for receiving the wireless transmission of the compressed measurement data from the base unit,a microprocessor for processing the compressed measurement data to determine the blood glucose concentration;a display for displaying blood glucose concentration. 2. The system as in claim 1 wherein the measurement signals are compressed through the selected sampling of the measurement signals. 3. The system as in claim 1 wherein the measurement signals are compressed through statistical techniques. 4. The system as in claim 3 wherein the statistical techniques select more than one measurement signal and determine the median of the more than one measurement signal. 5. The system as in claim 4 wherein the median of the more than one measurement signal is calculated by determining a slope of a line connecting pairs of measurement signals and calculating a median value of each slope of the line connecting each pair of measurement signals. 6. The system as in claim 5 wherein the median of the more than one measurement signal for each first time interval is calculated from the median value of each slope of the line connecting each pair of measurement signals. 7. The system as in claim 1 wherein display device further comprises a blood glucose meter. 8. The system as in claim 1, further comprising, a pre-amplifier for amplifying the measurement signals. 9. A base station for in-vivo measurement of blood glucose concentration in a human, comprising: a base power source to supply power to components of the base station;an implantable sensor for generating measurement signals that are correlated to a blood glucose concentration to be measured, the measurement signals being raw data recorded over first time intervals;a signal processor that executes instructions for processing the measurement signals to produce compressed measurement data for second time intervals which are at least ten-times longer than the fir... | Roche Diagnostics Operations, Inc. | A61B005/145 G06G007/60
US08734347 | B2 | 2014-05-27 | Analytical method and investigation system | 2008-05-30 | Metabolism diseases are evaluated by using a sensor to measure concentration of a medically significant analyte in a human or animal body fluid. The measurement data are correlated with the concentration of a medically significant analyte in body fluid over a time period of at least eight hours. An analytical facility analyzes time intervals within the time period to determine a stability parameter that characterizes the analyte concentration dynamics of change. The analytical facility further analyzes the stability parameters to determine disease-related particularities of metabolism. | 1. A system for investigation of the metabolism of a human or animal with regard to a medically significant analyte, comprising: a sensor for determining measuring data that are correlated with a concentration of a medically significant analyte in a human or animal body fluid, andan analytical facility for analyzing a series of the measuring data that have been determined by the sensor and are correlated with the concentration of the medically significant analyte in the human or animal body fluid for time points that are distributed over a period of time of at least eight hours, the analytical facility being free from need to obtain a reference value of the medically significant analyte to calibrate the measuring data;whereby the analytical facility is adapted such that in operation, it selects from the period of time multiple time intervals, which each of the time intervals extend over at least one hour and the analytical facility determines a stability parameter that characterizes the analyte concentration dynamics of change in the time interval for each of the time intervals by analyzing measuring data that are from the time interval, and the analytical facility analyzes the stability parameter and assigns the stability parameter to one of a plurality of predetermined parameter ranges in order to determine disease-related particularities of metabolism. 2. The system as in claim 1 wherein the analytical facility calculates a first derivative over time of the measuring data in a calculation step in order to calculate the stability parameter. 3. The system as in claim 2 wherein the analytical facility calculates a standard deviation of the first derivative over time in a further calculation step. 4. The system as in claim 1 wherein the analytical facility performs a Fourier transformation of a derivative of the measuring data over time to calculate the stability parameter. 5. The system as in claim 1 wherein the analytical facility, performs a linear transformation when operating on the measuring data. 6. The system as in claim 5 wherein the analytical facility selects the linear transformation for each individual time interval such that the mean of the measuring data of a corresponding time interval corresponds to a pre-determined value. 7. The system as in claim 6 wherein the analytical facility selects the linear transformation for each individual time interval such that the standard deviation of the measuring data of the corresponding individual time interval corresponds to a pre-determined value. 8. The system as in claim 1 wherein the analytical facility, is a hand-held device that comprises an output unit, an input unit, and a memory. 9. The system as in claim 1 wherein the analyte is glucose. 10. The system as in claim 9 wherein the analytical facility assigns therapeutic recommendations of a dosage of insulin doses to disease-related particularities of glucose metabolism that have been determined by analysis of the stability par... | Roche Diagnostics Operations, Inc. | A61B005/145 G06G007/60
US09662059 | B2 | 2017-05-30 | Imaging agents for monitoring changes of dopamine receptors and methods of using thereof | 2008-04-25 | The present invention is related generally to a method for screening subjects to determine those subjects more likely to develop diabetes by quantization of insulin producing cells. The present invention is also related to the diagnosis of diabetes and related to monitor disease progression or treatment efficacy of candidate drugs. | 1. A non-invasive method of monitoring changes in beta cell mass in a subject comprising: detecting a first time point level of Dopamine receptor expression in endocrine pancreatic islets in the pancreas of said subject,detecting a second time point level of Dopamine receptor expression in endocrine pancreatic islets in the pancreas of said subject, andmonitoring change in beta cell mass by comparing the first time point level of Dopamine receptor expression with the second time point level of Dopamine receptor expression,wherein a reduction of Dopamine receptor expression in endocrine pancreatic islets in the pancreas of said subject is indicative of reduction in beta cell mass, and wherein each of said detecting the first time point level and second time point level comprises an assay comprising the steps of: a. administering 18F-fallypride capable of binding with a high level of specificity to a D2-like receptor into said subject in the presence of haloperidol;b. measuring an imaging signal of 18F-fallypride in the pancreas of said subject in the presence of haloperidol;c. administering 18F-fallypride into said subject in the absence of haloperidol;d. measuring an imaging signal of 18F-fallypride in the pancreas of said subject in the absence of haloperidol;e. detecting the level of Dopamine receptor expression in endocrine pancreatic islets by subtracting the imaging signal in the presence of haloperidol from the imaging signal in the absence of haloperidol;and wherein the reduction in beta cell mass is predictive of being at risk or predisposed to developing clinical symptoms of diabetes in the subject. 2. The non-invasive method of claim 1, wherein said imaging signals of said 18F-fallypride is measured by positron emission tomography (PET) or positron emission tomography-computerized tomography (PET/CT) or positron emission tomography-magnetic resonance imaging (PET/MRI). 3. The method of claim 1, wherein said subject is at preclinical stage of diabetes. 4. The method of claim 3, wherein said diabetes is Type 1 or Type II diabetes. 5. The method of claim 1, wherein said subject is at clinical stage of diabetes. 6. The method of claim 5, wherein said diabetes is Type I or Type II diabetes. 7. The method of claim 1 wherein said subject is at preclinical stage or clinical stage of a diabetes-related metabolic syndrome. 8. The method of claim 1 wherein the administrations of 18F-fallypride are administered with glucose. 9. The method of claim 1 wherein an oral glucose test is conducted simultaneously. 10. The non-invasive method of claim 1, wherein said Dopamine receptor is D2 receptor. | The Regents of the University of California | A61B006/00 A61B005/00 A61B005/145 A61B006/03 A61B005/055
US20080319498 | A1 | 2008-12-25 | Apparatus and Method for Detecting Diastolic Heart Failure | 2004-08-31 | In a method and implantable medical apparatus for detecting diastolic heart failure (DHF), and a pacemaker embodying such an apparatus, movement of the valve plane of the heart is measured and analyzed to identify a slowing of the movement of the valve plane as an indication of a DHF state of the heart. A signal indicative of this DHF state is emitted and, in the pacemaker, is used to control the administration of a pacing pulse therapy to the heart. | 1-29. (canceled) 30. An implantable medical apparatus for detecting diastolic heart failure of the heart of a patient, comprising: a movement of the valve plane of the heart that emits an electrical signal representing a measure of movement of the valve plane of a heart;a movement analyzer, supplied with said electrical signal, that analyzes the movement of the valve plane represented by said electrical signal; andsaid movement analyzer comprising a determination unit that identifies slowing of said movement of said valve plane and that, upon identifying said slowing, emits a signal indicating a DHF state of the heart of the patient. 31. An apparatus as claimed in claim 30 wherein said movement analyzer comprises a comparator that compares said measurement of said movement of said valve plane with at least one predetermined reference value. 32. An apparatus as claimed in claim 30 wherein said movement analyzer identifies a time range comprising early diastole, before an atrial contraction, of said heart and analyzes said measurement of said movement of said valve plane in said time range. 33. An apparatus as claimed in claim 32 wherein said movement analyzer comprises a detector that detects a T-wave of the heart, and wherein said movement analyzer analyzes said measurement of said movement of the valve plane at a time substantially coinciding with said T-wave. 34. An apparatus as claimed in claim 30 comprising an activity sensor that detects an activity condition of the patient and that emits a further electrical signal to said movement analyzer representing said activity condition, and wherein said movement analyzer analyzes said measurement of said movement of the valve plane when said further electrical signal indicates a resting condition of the patient. 35. An apparatus as claimed in claim 30 wherein said movement of said valve plane exhibits a velocity, and wherein said movement measuring unit measures said velocity. 36. An apparatus as claimed in claim 30 wherein said movement of said valve plane exhibits an acceleration, and wherein said movement measurement unit measures said acceleration. 37. An apparatus as claimed in claim 36 wherein said movement measurement unit is an accelerometer configured for placement on or close to said valve plane. 38. An apparatus as claimed in claim 37 wherein said movement analyzer comprises a calculating unit that calculates a velocity of said valve plane by summing or integrating respective values representing said acceleration measured by said movement measuring unit. 39. An apparatus as claimed in claim 30 wherein said movement measurement unit measures a pressure correlating movement of said valve plane, at a location selected from the group consisting of an inner wall of the coronary sinus, an inner wall of the great cardiac vein, and an inner wall of a coronary vein. 40. An apparatus as claimed in claim 39 wherein said movement measurement unit is a pressure sensor having a circumferenti... | - | A61N001/365 A61B005/02
US20080269672 | A1 | 2008-10-30 | Analyte Monitoring Device and Methods of Use | 2008-05-01 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. A sensor control unit comprising: a housing adapted for placement on skin and adapted to receive a portion of an electrochemical sensor having a plurality of contact pads;a plurality of conductive contacts disposed on the housing and configured for coupling to the plurality of contact pads on the sensor; anda transmitter disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. 2-120. (canceled) 121. A diabetes management system comprising: an insulin delivery device including:means for sending and receiving RF signals;means for dispensing insulin in accordance with dispensing instructions received from a remote controller;a glucose sensor assembly including:means for generating sensor data corresponding to the user's glucose level;means for storing the sensor data;means for receiving a sensor data download request;means for transmitting the sensor data from the data storage means in the sensor to data storage means in another device in response to a sensor data download request;a remote controller comprising:means for receiving input from the user;means for delivering output to the user;means for sending dispensing instructions to the delivery device by RF signals;means for sending sensor data download requests to the sensor assembly by RF signals;means for receiving sensor data from the sensor assembly by RF signals;means for storing the sensor data. 122. The system of claim 121, wherein the sensor assembly comprises means for sending an alarm signal to the delivery device in response to the sensor assembly sensing a predetermined glucose condition in the user; andthe delivery device comprises means for receiving an alarm signal from the sensor assembly and means for generating an alarm signal that may be discerned by the user. 123. The system of claim 122, wherein the user discernable delivery device alarm signal is at least one of an audible, visual or tactile signal. 124. The system of claim 123, wherein the sensor assembly comprises means for generating a second alarm signal to the delivery device in response to the sensor assembly sensing a second predetermined glucose condition in the user. 125. The system of claim 124, wherein the delivery device comprises means for generating a first alarm signal in response to the first sensor assembly alarm signal and a second alarm signal in response to the second sensor assembly alarm signal. 126. The system of claim 121, wherein the sensor assembly senses the user's glucose substantially continuously and the sensor assembly memory is sufficient to store at least 12 hours of substantially continuous sensor data. 127. The system of claim 121, wherein the means for sending sensor data download requests to the sensor assembly is activated by user input on the remote controller. 128. The system of claim 121, wherein the sensor assembly comprises software for predicting whether the user's glucose level will exceed or fall belo... | Abbott Diabetes Care, Inc. | A61B005/055 A61M031/00
US20080294036 | A1 | 2008-11-27 | Surgical Metal Detection Apparatus and Methods | 2008-04-22 | A surgical metal detection apparatus. The apparatus comprises an elongate, tubular cannula and a removable element, which may be a trocar, releasably engaged within the cannula. A metal detection coil is positioned within the removable element or, in some embodiments, within the cannula. The metal detection coil is connected to a metal detection circuit which is adapted to detect the presence of a metallic element proximate to the end of the removable element or cannula. When a metallic element has been detected, the removable element may be removed from the cannula to permit the passage of a removal tool, such as a screwdriver, through the cannula without removing or changing the position of the cannula. | 1. A surgical metal detection apparatus, comprising: an elongate, generally tubular cannula having a first end, a second end opposite the first end, and a passageway between the first and second ends that is open at both ends;an elongate removable element having first and second ends corresponding to the first and second ends of the cannula, the removable element being sized and adapted (1) to be removably retained in an engaged position substantially within the passageway of the cannula such that the second end of the removable element is proximate to the second end of the cannula, and (2) to be removed from the cannula to allow for the passage of a removal tool through the cannula;a detection coil within the cannula or the removable element; anda detection circuit connected to the detection coil, the detection circuit and detection coil being constructed and adapted to detect the presence of a metallic element proximate to the second end of the cannula or the second end of the removable element when the removable element is in its engaged position substantially within the passageway of the cannula. 2. The surgical metal detection apparatus of claim 1, wherein the cannula is substantially rigid. 3. The surgical metal detection apparatus of claim 1, wherein the cannula and the removable element are formed of biocompatible materials. 4. The surgical metal detection apparatus of claim 3, wherein the biocompatible materials are plastics. 5. The surgical metal detection apparatus of claim 1, wherein the detection circuit is a beat frequency oscillation metal detection circuit. 6. The surgical metal detection apparatus of claim 1, wherein the removable element is a trocar. 7. The surgical metal detection apparatus of claim 1, wherein the cannula and the removable element have cooperating engaging structures constructed and arranged to retain the removable element in the engaged position. 8. The surgical metal detection apparatus of claim 7, wherein the cooperating engaging structures comprise a set of threads on an interior surface of the cannula and a cooperating set of threads on an exterior surface of the removable element. 9. The surgical metal detection apparatus of claim 1, wherein the detection coil is within the removable element proximate to the second end of the removable element. 10. The surgical metal detection apparatus of claim 9, wherein the detection coil includes a ferrite core. 11. The surgical metal detection apparatus of claim 1, wherein the detection coil is within the cannula proximate to the second end of the cannula. 12. The surgical metal detection apparatus of claim 11, wherein the removable element includes a ferrite element proximate to its second end positioned such that when the removable element is in the engaged position, the ferrite element acts as a ferrite core for the detection coil within the cannula. 13. The surgical metal detection apparatus of claim 1, the detection circuit further comprising one or ... | DEVICE EVOLUTIONS, LLC | A61B005/00
US20080262321 | A1 | 2008-10-23 | Early Detection of Harmful Agents: Method, System and Kit | 2005-08-04 | A portable system for detecting agents present in a sample is disclosed. The system comprises a sensing device, having a substrate formed with a plurality of reaction chambers and a plurality of channels interconnecting at least a portion of the plurality of reaction chambers, wherein at least a portion of the plurality of reaction chambers comprises a sensor, capable of generating a detectable signal when exposed to the agents. The system further comprises a detector, which receives signals from the sensing device and provides an image of sensors generating the optical signals. The portable system is connected to a communication network via a communication unit. | 1. A portable system for detecting agents present in a sample, the system comprising: a sensing device having a substrate formed with a plurality of reaction chambers and a plurality of channels interconnecting at least a portion of said plurality of reaction chambers, wherein at least a portion of said plurality of reaction chambers comprises a sensor, capable of generating a detectable signal when exposed to the agents;a detector capable of receiving signals from said sensing device and providing an image of sensors generating said detectable signals; anda communication unit, for connecting the portable system to a communication network. 2. A kit for detecting agents present in a sample, the kit comprising: (a) a portable detection system, comprising:a sensing device having a substrate formed with a plurality of reaction chambers and a plurality of channels interconnecting at least a portion of said plurality of reaction chambers, wherein at least a portion of said plurality of reaction chambers comprises a sensor, capable of generating a detectable signal when exposed to the agents;a detector capable of receiving signals from said sensing device and providing an image of sensors generating said optical signals; anda communication unit, for connecting the portable system to a communication network; and(b) a sampling device being connectable to said detection system and capable of sampling fluids from the environment and delivering said fluids to said sensing device. 3. The kit of claim 2 wherein said sampling device comprises a plurality of containers and treating elements for treating fluids in said plurality of containers prior to said transfer of said fluids to said sensing device. 4. The kit of claim 2, further comprising a medical injection device, for injecting at least one medicament to a mammal. 5. The kit of claim 4, wherein said injection device comprises: an injection needle;a plurality of syringes, each having a plunger slidably disposed therein;a driving mechanism having a plurality of drive members, each drive member of said plurality of drive members being operable to engage a plunger of a respective syringe of said plurality of syringes; anda redirector for redirecting medicaments, released by at least one syringe of said plurality of syringes, into said injection needle. 6. The kit of claim 2, further comprising a storage unit for storing said sensing device, said sampling device and/or said fluids, and maintaining a predetermined temperature thereof. 7. The kit of claim 6, wherein said storage unit comprises: at least one cavity having a plurality of temperature control units and configured to receive a plurality of objects in a manner such that a plurality of thermal communications are established between said plurality of temperature control units and said plurality of objects, each of said plurality of objects being selected from the group consisting of said sensing device, said sampling device and said fluids; anda c... | Ramot At Tel Aviv University Ltd. Yissum Research Development Company of the Hebrew University of Jerusalem | G01N033/00 C12Q001/00 A61M005/20 B01L003/00 G08B021/12 A61B005/00 G01N021/76 C12M001/34
US20080262330 | A1 | 2008-10-23 | Analyte Sensors and Compositions for Use Therein | 2006-06-30 | The present invention relates to analyte sensors and compositions, membranes for use in analyte sensors. For example, an analyte sensor is provided that comprise an electrode surface comprising an enzyme; an a biocompatible analyte permeable composition comprising a nitric oxide generating agent. | 1. An analyte sensor suitable for implantation in a patient comprising: an electrode surface comprising an enzyme; anda biocompatible analyte permeable composition comprising a nitric oxide generating agent, wherein said composition generates nitric oxide when in contact with nitrosothiols. 2. The analyte sensor of claim 1, wherein said enzyme is an oxidase. 3. The analyte sensor of claim 1, wherein said composition further comprises a pharmaceutically effective agent. 4. The analyte sensor of claim 2, wherein the enzyme is glucose oxidase. 5. The analtye sensor of claim 1, wherein said composition substantially prevents biofouling of said sensor for three days or more after implantation. 6. The analyte sensor of claim 1, wherein said sensor is capable of continuously measuring blood glucose levels in a patient for 3 days or more. 7. The analyte sensor of claim 1, wherein said sensor is subcutaneously implantable. 8. The analyte sensor of claim 1, wherein said composition is disposed on said electrode surface. 9. The analyte sensor of claim 1, wherein said analyte permeable composition can substantially exclude at least one of: ascorbate, urate, cysteine or paracetamol. 10. The analyte sensor of claim 1, wherein said analyte permeable composition is substantially glucose permeable. 11. The analyte sensor of claim 1, wherein said nitric oxide generating agent is a metal ion binding moiety. 12. The analyte sensor of claim 11, wherein said metal ion binding moiety is a Nx-donor ligand where x is 2,3,4,5,6,7,8or9. 13. The medical device of claim 11, wherein said nitric oxide generating compound is a Sy-donor ligand where y is 2, 3, 4, 5, 6, 7, 8 or 9. 14. The analyte sensor of claim 11, wherein said nitric oxide generating agent is a metal-ligand complex comprising said metal ion binding moiety and a metal ion that is redox active. 15. The analyte sensor of claim 14, wherein said metal ligand complex is a metal-cyclen complex or a metal-cyclam complex. 16. The analyte sensor of claim 15, wherein said metal-cyclen complex is a copper-cyclen complex or a copper-cyclam complex. 17. The analyte sensor of claim 14, wherein said metal-ligand complex is a N4 donor type macrocycle. 18. The analyte sensor of claim 1, wherein the nitric oxide generating agent is represented by Formula I:wherein:R1 and R2 represent each independently: alkyl, alkenyl, alkynl, H, or any R1 and R2 together form an aryl or an heteroaryl;------- represents a bond with any bond order;M is a metal; and salts thereof. 19. The analyte sensor of claim 18, wherein M is metallic ion selected from the group consisting of Ca, Cu, Zn, Co, Mn, Al, Fe, V, Cr, and Ti. 20. The analyte sensor of claim 18, wherein the nitric oxide generating compound is a Cl, I, F, or Br salt of Formula I. 21. The analyte sensor of claim 18, wherein M is Cu(II). 22. The analyte sensor of claim 1, wherein said nitric oxide generating agent is selected from at least one of: Cu(II)-dibenzo[e,k]-2,... | - | A61B005/1486
US20080287817 | A1 | 2008-11-20 | APPARATUS FOR DETECTING HUMAN PHYSIOLOGICAL AND CONTEXTUAL INFORMATION | 2007-10-29 | The invention comprises an apparatus for determining the contextual or physiological status of the individual wearing the apparatus. The apparatus is designed to be consumable or disposable. In most embodiments the invention comprises an adhesive housing. In some embodiments, two different sensors are secured to the housing. The apparatus is in electronic communication with a processing unit that can derive analytical status data by using the data received from the two sensors. In some embodiments, the processing unit is programmed to control other devices, or is programmed to trigger an event. In still other embodiments, the apparatus is in electronic communication with a separate computing device, which may contain the processing unit. | 1. An apparatus for monitoring an individual's physiological or contextual status, said apparatus able to be disposed of after use, said apparatus comprising: a. a housing having a surface, said surface comprising an adhesive material, said adhesive material enabling said housing to be removably attached to selected portions of said individual's body;b. a first sensor secured to said housing, said first sensor capable of generating heart-related data of said individual;c. a second sensor secured to said housing, said second sensor capable of generating temperature-related data of said individual; andd. a processing unit in electronic communication with said first and second sensors, said processing unit programmed to use both of said heart-related data from said first sensor and said temperature-related data from said second sensor to derive analytical status data. | - | A61B005/02
US20080242947 | A1 | 2008-10-02 | Configuring software for effective health monitoring or the like | 2007-03-30 | A system, method, computer program product, and carrier are described for obtaining an indication of an activity status change of an application program and causing a movement of data distillation code relating to physiology-indicative data in response to the indication of the activity status change of the application program; or for obtaining (a) first clinical data acceptable to a clinical analysis module and (b) a pointer to the clinical analysis module, configuring one or more applications using the pointer to the clinical analysis module after receiving at least the first clinical data acceptable to the clinical analysis module, and using second clinical data acceptable to the clinical analysis module with at least one of the one or more applications configured using the pointer to the clinical analysis module after receiving at least the first clinical data acceptable to the clinical analysis module. | 1. A method comprising: obtaining an indication of an activity status change of an application program; andcausing a movement of data distillation code relating to physiology-indicative data in response to the indication of the activity status change of the application program. 2. The method of claim 1 in which obtaining an indication of an activity status change of an application program comprises: receiving information about an active process of the application program. 3-4. (canceled) 5. The method of claim 1 in which obtaining an indication of an activity status change of an application program comprises: configuring a module with information specific to the application program. 6. (canceled) 7. The method of claim 1 in which causing a movement of data distillation code relating to physiology-indicative data in response to the indication of the activity status change of the application program comprises: configuring a processor in response to one or more physiological expressions. 8-9. (canceled) 10. The method of claim 1 in which causing a movement of data distillation code relating to physiology-indicative data in response to the indication of the activity status change of the application program comprises: requesting the data distillation code relating to physiology-indicative data. 11. (canceled) 12. The method of claim 1 in which obtaining an indication of an activity status change of an application program comprises: obtaining an invocation addressing the application program. 13. The method of claim 1 in which causing a movement of data distillation code relating to physiology-indicative data in response to the indication of the activity status change of the application program comprises: receiving an indicator of an association between the data distillation code and a data type. 14. The method of claim 1 further comprising: executing the data distillation code at least upon physiology-indicative output from the application program. 15. The method of claim 1 further comprising: executing the data distillation code upon data about more than one person. 16. (canceled) 17. A system comprising: means for obtaining an indication of an activity status change of an application program; andmeans for causing a movement of data distillation code relating to physiology-indicative data in response to the indication of the activity status change of the application program. 18. (canceled) 19. The system of claim 17 in which the means for obtaining an indication of an activity status change of an application program comprises: means for receiving an output from the application program. 20. The system of claim 17 in which the means for obtaining an indication of an activity status change of an application program comprises: means for receiving an indicator of a user of the application program. 21. (canceled) 22. The system of claim 17 in which the means for causing a movement of data distillation code relating to physiology-indicativ... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00
US20080287821 | A1 | 2008-11-20 | Computational user-health testing | 2008-05-07 | Methods, apparatuses, computer program products, devices and systems are described that carry out accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application; accepting brain activity measurement data proximate to the interaction; associating the at least one user-health test function output with the brain activity measurement data; and presenting the at least one user-health test function output and the brain activity measurement data at least partly based on the associating the at least one user-health test function output with the brain activity measurement data. | 1. A method comprising: accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application;accepting brain activity measurement data proximate to the interaction;associating the at least one user-health test function output with the brain activity measurement data; andpresenting the at least one user-health test function output and the brain activity measurement data at least partly based on the associating the at least one user-health test function output with the brain activity measurement data. 2. The method of claim 1 wherein the accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application comprises: accepting at least one personal computing device output as the at least one user-health test function output. 3. The method of claim 1 wherein the accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application comprises: accepting at least one alertness or attention test function output at least partly based on an interaction between at least one user and at least one device-operable application. 4. The method of claim 1 wherein the accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application comprises: accepting at least one memory test function output at least partly based on an interaction between at least one user and at least one device-operable application. 5. The method of claim 1 wherein the accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application comprises: accepting at least one speech test function output at least partly based on an interaction between at least one user and at least one device-operable application. 6. The method of claim 1 wherein the accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application comprises: accepting at least one calculation test function output at least partly based on an interaction between at least one user and at least one device-operable application. 7. The method of claim 1 wherein the accepting at least one user-health test function output at least partly based on an interaction between at least one user and at least one device-operable application comprises: accepting at least one neglect or construction test function output at least partly based on an interaction between at least one user and at least one device-operable application. 8. The method of claim 1 wherein the accepting at least one user-health test function output at least partly based on an interaction betw... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/12 A61B005/00 A61B005/0476 A61B005/11
US20080248871 | A1 | 2008-10-09 | INTERFACE DEVICE | 2008-03-14 | The present invention relates generally relates to an interface device for operative connection between a computer and a one-dimensional, two-dimensional or three-dimensional motion input system and/or tactile force sensing input system. The interface device includes a first interface for operative connection to the input system and a second interface for operative connection to the computer; and a processor operatively connected to the first and second interfaces for receiving input signals from the motion tracking/tactile force input system. The device, or computer, configures the input signals to output signals for input to the computer for controlling the computer. The invention is particularly beneficial in a therapy/rehabilitation setting wherein people or patients can participate in almost any video game by configuring the input system to the video game through the interface device. | 1. An interface device for operative connection between a computer and a motion input system comprising: a first interface for operative connection to the motion input system and a second interface for operative connection to the computer; anda processor operatively connected to the first and second interfaces for receiving input signals from the motion input system and enabling configuration of the input signals to output signals for input to the computer for controlling the computer. 2. An interface device as in claim 1 wherein the motion input system is a one, two or three dimensional motion input system. 3. An interface device as in claim 1 wherein the motion input system is a tactile force input system. 4. An interface device as in claim 1 wherein the motion input system is combination of a tactile force and a one, two or three dimensional input system. 5. An interface device as in claim 1 wherein the processor includes means for dynamically selecting any one of or a combination of keyboard, mouse and joystick equivalent outputs. 6. An interface device as in claim 1 wherein the processor includes means for testing communication between the interface device and the computer. 7. An interface device as in claim 1 wherein the processor includes means for receiving input from a user to adjust the linear and angular motion of a sensor of the input system to a defined range of motion for input to the computer. 8. An interface device as in claim 7 wherein the processor includes means for selectively excluding selected linear or angular motion of a sensor based on the measured movement of the sensor by a user. 9. An interface device as in claim 7 wherein the processor includes means for selectively smoothing input signals from the sensor based on the measured movement of the sensor. 10. An interface device as in claim 7 wherein the processor includes means for selectively amplifying input signals from the sensor based on the measured movement of the sensor. 11. An interface device as in claim 1 wherein the processor includes means for testing input and output signals for simulating output signals to verify compatibility with the computer's software. 12. An interface device as in claim 1 wherein the processor includes means for providing output signals to a USB port. 13. An interface device as in claim 1 further comprising a keyboard input interface operatively connected to the processor for providing keyboard input to the processor. 14. An interface device as in claim 1 further comprising a keypad input interface operatively connected to the processor for providing keypad input to the processor. 15. An interface device as in claim 1 further comprising a display operatively connected to the processor for displaying graphic output to the user. 16. An interface device as in claim 1 wherein the processor includes firmware enabling a menu system for configuring the motion input system. 17. An interface as in claim 16 wherein the menu syste... | - | A61B005/103 A63F009/24
US20080264426 | A1 | 2008-10-30 | DEVICE TO TREAT SNORING AND OBSTRUCTIVE SLEEP APNEA IN ADULTS AND TO PREVENT INFANTS FROM SLEEPING NON-SUPINE | 2008-04-28 | This is a position sensing monitor that provides private sensory biofeedback to adults, in the form of auditory stimulus or tactile sensation (vibration), to promote non-supine sleep. The essential feature of this invention is to eliminate/reduce position-related snoring and obstructive sleep apnea and not disrupting sleep of bed partner. A variation of this concept is its use in infants to promote sleeping supine by parental notification of rolling from supine to side or prone positions, allow re-position to supine and thereby reducing the risk for SIDS. | 1. A small monitor that identifies supine and non-supine sleeping position. 2. A system as in claim 1 to treat snoring and obstructive sleep apnea in adults without disturbing bed-partner's sleep 3. A system as in claim 2, that is based upon biofeedback principles and provides an intermittent audible tone and/or tactile stimulation when an adult individual is supine position, and ceasing once move to a non-supine position. 4. A system as in claim 3 where the audible can be provided privately by means of a small bud earphone, either wireless or connected by fine wire to the monitor, to not to disturb the individual's bed partner. 5. A system as in claim 2 is worn on chest by means of a strap, adhesive material that attaches to the skin, or attaches to bed clothing 6. A system as in claim 2 that, by learning principles, can condition an individual from sleeping supine. 7. A system as in claim 2 that can store information regarding time spent in the supine and non-supine position. 8. A system as in claim 1 where is attached to an infants bed clothing. 9. A system as in claim 8 that emits an audible alarm to notify parents when an infant moves to a non-supine position. | - | A61F005/56 A61B005/11
US20080287830 | A1 | 2008-11-20 | Apparatus and Method for Portable Blood Cell Counting | 2006-12-01 | The invention relates to a novel self-contained apparatus and method for portable and rapid counting of blood cells in a sample drawn from a subject in need of determining his or her blood cell counts. The apparatus may be about the size of a credit card, and rapidly determines a subject's blood cell counts. | 1. A self-contained apparatus for counting blood cells comprising a substrate with an upper surface and a lower surface, the substrate comprising; (a) a microneedle pad, the microneedle pad being capable of drawing a blood sample from the digit of a human subject;(b) a pump capable of forcing the blood sample into a blood and reagent mixing bulb, wherein the blood sample is treated with reagents to yield a treated sample free of undesired blood components, and of desired viscosity and dilution;(c) blood cell filtration elements to filter white blood cells into a white blood cell counting channel and red blood cells into a red blood cell counting channel, the diameter of at least the distal portion of the channels being sufficient for only a single file of blood cells to occupy the counting channels, wherein the distal ends of the channels narrow further to prevent the further passage of blood cells; and(d) an array of counting sensors adjacent to the blood cell counting channels, the sensors being in communication with a logic chip adapted to determine blood cell counts from the data communicated by the sensors and to display the blood cell counts on cell count indicators on the upper surface of the substrate. 2. The apparatus of claim 1, further comprising at least one component selected from the group consisting of a validity indicator circuit, a clean sample indicator, a bladder seal, a bladder adhesion ring, a sample holding chamber, a flat cell battery, a solar cell, a blackflow check valve, an air accumulator bulb, air suction channels, an air suction draw tube, an air entry channel, and instructions printed or etched on the upper surface of the substrate. 3. The apparatus of claim 1, wherein the substrate comprises an upper substrate and a lower substrate, wherein the joinder of the upper and lower substrates forms the channels and chambers by complementary grooves in an interior surface of the upper and lower substrates. 4. The apparatus of claim 1, wherein the substrate comprises a plurality of layers laid down by thin layer deposition under computer aided design control. 5. The apparatus of claim 1, wherein the substrate comprises an upper substrate, a middle substrate, and a lower substrate joined together. 6. The apparatus of claim 1, wherein the pump is manually operated by the subject successively pressing and releasing the microneedle pad. 7. The apparatus of claim 1, wherein the blood sample drawn into the apparatus is of volume between about 50 μl and about 1 ml. 8. The apparatus of claim 1, wherein the blood cell counts are displayed on the upper surface of the substrate by LEDs. 9. A self-contained apparatus for counting red blood cells comprising a substrate with an upper surface and a lower surface, the substrate comprising; (a) a microneedle pad, the microneedle pad being capable of drawing a blood sample from the digit of a human subject;(b) a pump capable of forcing the blood sample into a blood and reagent mixing ... | - | A61B005/157
US07454242 | B2 | 2008-11-18 | Tissue sensing adaptive radar imaging for breast tumor detection | 2004-09-17 | A tissue-sensing adaptive radar method of detecting tumours in breast tissue uses microwave backscattering to detect tumours which have different electrical properties than healthy breast tissue. The method includes steps for reducing skin reflections and for constructing a three-dimensional image using synthetic focusing which shows the presence or absence of microwave reflecting tissues. | 1. A method of detecting the presence or absence of a tumour within a breast comprising skin and an interior volume, comprising the steps of: (a) illuminating the breast with microwaves from a plurality of locations and recording the reflections received at each location as a signal;(b) identifying a first skin reflection and a second skin reflection separated by a period of time and time-gating the signal by setting all data arriving before the first skin reflection and after the second skin reflection to zero;(c) creating a first estimate of reflections from the skin and subtracting said first skin reflections from each signal;(d) creating a second estimate of reflections from the skin for a single location from the signals received in at least two adjacent locations and subtracting the second skin reflections from each signal;(e) constructing a three-dimensional image of the interior volume from the signals showing the presence or absence of microwave reflecting tissues. 2. The method of claim 1 further comprising the step of adaptively estimating reflections from strongly reflecting non-tumour tissue within the interior volume and subtracting such reflections from the signals. 3. The method of claim 1 wherein the image construction step comprises focusing based on relative distance between the focal point and the antenna. 4. The method of claim 3 wherein the image reconstruction step is free of a path loss compensation step. 5. The method of claim 3 wherein the resulting image is filtered to emphasize locations closer to the centre of the breast, away from the antenna. 6. The method of claim 1 further comprising the step of co-registering a magnetic resonance image of the same breast and comparing features between the microwave radar image and the magnetic resonance image. 7. The method of claim 1 further comprising the steps of creating a first estimate of reflections from a non-tumour reflecting structure within the interior volume and subtracting said non-tumour reflections from each signal; and creating a second estimate of reflections from the non-tumour reflecting structure for a single location from the signals received in at least two adjacent locations and subtracting the second estimate from each signal. 8. The method of claim 1 wherein a person whose breast is being scanned lies face down while being scanned. 9. A method of detecting the presence or absence of a tumour within a breast comprising skin and an interior volume, comprising the steps of: (a) illuminating the breast with microwaves from a plurality of locations and recording the reflections received at each location as a signal;(b) identifying a first skin reflection and a second skin reflection separated by a period of time and time-gating the signal by setting all data arriving before the first skin reflection and after the second skin reflection to zero;(c) creating a first estimate of reflections from the skin and subtracting said first skin reflections from e... | - | A61B005/05
US07454243 | B2 | 2008-11-18 | Aptitude testing | 2004-08-18 | A method of assessing the cognitive aptitude of a subject to a predetermined task, the method including the steps of: (i) presenting to the subject a group of cognitive tasks; (ii) detecting brain response signals from the subject during presentation of the group of cognitive tasks; (iii) calculating SSVEP amplitude, phase and/or coherence responses from the brain response signals; and (iv) comparing the SSVEP responses to known SSVEP responses obtained from individuals with high and/or low aptitudes to the predetermined task in order to assess the subject's aptitude for the predetermined task. | 1. A method of assessing the cognitive aptitude of a subject to a predetermined task, the method including the steps of: (i) simultaneously presenting to the subject one of a group of cognitive tasks and a visual flicker;(ii) detecting brain response signals from the subject during presentation of said cognitive task and visual flicker;(iii) calculating amplitude, phase and/or coherence of SSVEP responses elicited by the visual flicker from said brain response signals; and(iv) comparing said SSVEP responses to known SSVEP responses obtained from individuals with high and/or low aptitudes to said predetermined task in order to assess the subject's aptitude for said predetermined task. 2. A method as claimed in claim 1 wherein including the step of presenting said group of cognitive tasks to said individuals in order to obtain said known SSVEP responses and storing said known SSVEP response in a database. 3. A method as claimed in claim 1 or 2 wherein the cognitive tasks are selected so that they place demands on the subject which are similar to demands experienced when carrying out the predetermined task. 4. A method as claimed in claim 3 wherein the cognitive tasks are selected so that they place one or more of the following demands on the subject: attention, analytical thinking, holistic thinking, verbal thinking, visuo-spatial thinking, working memory, recognition memory and identifying emotional expressions. 5. A method as claimed in claim 1 including the steps of: repeating the presentation of said cognitive tasks in order to calculate multiple SSVEP responses;statistically analysing said SSVEP responses in order to determine statistically significant changes in SSVEP amplitude, phase and/or coherence; andcomparing said statistically significant changes to said known SSVEP responses in order to assess the subject's aptitude for said predetermined task. 6. A method as claimed in claim 1 wherein the step of comparing said SSVEP responses to known SSVEP responses includes the step of assessing the subject's thinking style. 7. A method as claimed in claim 1 wherein steps (i), (ii) and (iii) are performed at a local site and wherein step (iv) is performed at a remote site. 8. A method as claimed in claim 7 including the step of maintaining a database of said known SSVEP responses at said remote site. 9. A method as claimed in claim 8 including the step of communicating the amplitude, phase and/or coherence SSVEP responses from the local site via the Internet to said remote site. 10. Apparatus for assessing the cognitive aptitude of a subject to a predetermined task, the apparatus including: (i) means for simultaneously presenting to the subject one of a group of cognitive tasks and a visual flicker;(ii) means for detecting brain response signals from the subject during presentation of said cognitive task and visual flicker;(iii) means for calculating amplitude, phase and/or coherence of SSVEP responses elicited by the visual flicker from ... | SSPT Pty Ltd. | A61B005/00
US07447345 | B2 | 2008-11-04 | System and method for generating PET-CT images | 2003-12-16 | The invention relates to a system and method for executing a scanning procedure comprising the steps of generating CT image data for a scan, generating CT attenuation correction data for the scan, acquiring a current frame of PET data for the scan in a 3D format, and simultaneously conducting the following steps: reconstructing at least a portion of a PET image for the current frame, including the step of overlapping a portion of the current frame with an adjacent frame, and acquiring at least a portion of a next frame of PET data. | 1. A method for executing a scanning procedure comprising the steps of: (a) generating CT image data for a scan;(b) generating CT attenuation correction data for the scan;(c) acquiring a current frame of PET data for the scan in a 3D format; and(d) simultaneously conducting the following steps: reconstructing at least a portion of a PET image for the current frame, including the step of overlapping a portion of the current frame with an adjacent frame, andacquiring at least a portion of a next frame of PET data. 2. The method of claim 1, further comprising repeating step (d) for at least one subsequent frame. 3. The method of claim 1, wherein the 3D format comprises a projection plane format. 4. The method of claim 3, wherein the step of reconstructing at least a portion of a PET image comprises applying a Fourier rebinning process to directly convert the PET data from the projection plane format to a sinogram format. 5. The method of claim 1, wherein the step of overlapping comprises computing a weight for each overlapping slice in the overlapping portion of the current frame and the adjacent frame. 6. The method of claim 1, wherein the adjacent frame comprises a previous frame, and the step of overlapping comprises retrieving a stored overlapping portion of the previous frame. 7. The method of claim 1, further comprising the step of storing an overlapping portion of the current frame to enable overlapping during reconstruction of the next frame. 8. The method of claim 1, further comprising the step of automatically configuring a scan protocol for the PET data based on a scan protocol used to acquire the image CT data. 9. The method of claim 1, further comprising the step of generating a fused PET-CT image with the CT image data and the PET image. 10. The method of claim 1, wherein the reconstructing step includes applying an attenuation correction to the PET image using the CT attenuation correction data. 11. A PET-CT system comprising: a CT detector;a PET detector;at least one memory; andat least one processor, which is programmed to control a generation of CT image data for the scan;control a generation of CT attenuation correction data for the scan;control an acquisition of a current frame of PET data for the scan in a 3D format; andcontrol a simultaneous execution of (a) a reconstruction of at least a portion of a PET image for the current frame, including overlapping a portion of the current frame with an adjacent frame, and (b) an acquisition of at least a portion of a next frame of PET data. 12. The PET- CT system of claim 11, wherein the at least one processor is programmed to repeat, for at least one subsequent frame, the step of controlling the simultaneous execution of (a) and (b). 13. The PET-CT system of claim 11, wherein the 3D format comprises a projection plane format. 14. The PET-CT system of claim 13, wherein the reconstruction of at least a portion of a PET image comprises application of a Fourier rebinning proce... | General Electric Company | G06K009/00 A61B005/05 G01T001/164
US07587945 | B2 | 2009-09-15 | Device for measuring tibio-femoral force amplitudes and force locations in total knee arthroplasty | 2008-08-19 | A probe used during a total knee arthroplasty for measuring forces and locations of their points of application and thereby moments includes two load sensitive plates t to be inserted in one joint-compartment of a knee joint each and each being provided with a top surface and a bottom surface. At least two load sensors may be situated on the top surfaces and/or the bottom surface of each load sensitive plate. | 1. A probe used during a total knee arthroplasty for measuring force amplitudes and force locations, comprising: first and second load sensitive plates connected to each other, each for insertion into a joint-compartment of a knee joint and each having a top surface and a bottom surface, wherein each of the first and second load sensitive plates includes at least three load sensors that are non-colinearly arranged in such a way as to allow measurement of the force amplitudes and force locations along two perpendicular axes. 2. A probe according to claim 1, wherein at least two of the load sensors are situated on the bottom surface of one of the first and second load sensitive plates. 3. A probe according to claim 1, wherein at least one of the load sensors is situated on the top surface of one of the first and second load sensitive plates, and at least another one of the load sensors is situated on the bottom surface of one of the first and second load sensitive plates. 4. A probe according to claim 1, further comprising a base plate configured to be in contact with the tibia. 5. A probe according to claim 1, further comprising a set of wedges. 6. A probe according to claim 1, wherein each load sensor comprises a bridge-shaped structure and at least one piezoresistive sensor. 7. A probe according to claim 1, wherein measured parameters are used as inputs of a computational biomedical model of the knee joint acting as an assistive expert system. 8. A probe according to claim 1, wherein wireless telemetry is used to transmit measurement signals emitted by the load sensors to a computer. 9. A probe according to claim 1, wherein each load sensor provides measurement data for a discrete point. 10. A probe according to claim 9, wherein the discrete point measurements of the load sensors are spaced apart relative to one another by a distance greater than 2 mm. 11. A probe according to claim 1, wherein the connection between the first and second load sensitive plates is via a flexible joining element. 12. A probe according to claim 11, wherein a gap is present between the connected first and second load sensitive plates. 13. A probe according to claim 11, wherein the connection between the first and second load sensitive plates is along a segment of adjoining inner lateral side walls of the first and second load sensitive plates. 14. A probe used during a total knee arthroplasty for measuring force amplitudes and force locations, comprising: first and second load sensitive plates separated from one another at a gap between adjoining side walls thereof, the first and second plates being adapted for insertion into a joint-compartment of a knee joint, wherein each of the first and second load sensitive plates includes at least two load sensors, the first and second plates being electrically connected to one another across the gap between the adjoining side walls. 15. A probe according to claim 14, wherein the at least two load sensors on the f... | Synthes USA, LLC | G01B005/30
US07457655 | B2 | 2008-11-25 | Motion correction of magnetic resonance images using moments of spatial projections | 2004-10-08 | Magnetic resonance images reconstructed from a radial/projection acquisition are corrected for motion corruption caused by in-plane translational and in-plane rotational motion of an imaged subject using only the projection data itself. The method is based on the consistency properties of the 0th, 1st, and 2nd order moments of the spatial domain projections. In-plane translational motion is corrected by shifting/aligning the spatial projections according to the center of mass of each projection, which is calculated using the 0th and 1st moments. In-plane rotational motion is accounted for by determining the rotational motion time record using the 2nd moment information. The determination of the rotational motion time record using the 2nd moments is made possible by acquiring the successive MR projections at a view angle spacing that is substantially 45° and thus achieves sufficient linear independence. The translation-corrected spatial projections are reconstructed using the actual projection view angles, which are sum of the intended projection view angles plus the detected in-plane rotational motion time record. The calculation may also be performed in real-time during data acquisition in order to reacquire projections found to be too inconsistent due to through plane motion or other changes that are not in-plane rigid body motions or to acquire projections determined to be missing from the data set due to rotational motion. | 1. A method for producing an image of an object with a magnetic resonance imaging (MRI) system, the steps comprising: a) acquiring a projection of NMR data by performing a pulse sequence with the MRI system which samples through the center of k-space at a selected view angle;b) repeating step a) a plurality of times at different view angles such that the difference in view angles of successive projection acquisitions is an angle spacing that achieves sufficient linear independence;c) performing a one-dimensional Fourier transformation of each projection to form a plurality of spatial projections;d) calculating the 0th moment of each spatial projection;e) calculating the first moment of each spatial projection;f) correcting spatial projections for in-plane translational motion of the object using the calculated 0th and first moment information;g) calculating the second moment of each translation-corrected spatial projection;h) determining a rotational motion time record based on the 2nd moments of the translation-corrected spatial projections; andi) calculating the actual projection view angles, which are perturbed by the in-plane rotational motion, using the detected rotational motion time record; andj) reconstructing an image using the translational-corrected spatial projections with a reconstruction method which incorporates the detected rotational motion time record to reconstruct the spatial projections based on their actual projection view angle. 2. The method as recited in claim 1 in which the pulse sequence is a two-dimensional projection reconstruction (2DPR) pulse sequence. 3. The method as recited in claim 2 in which the pulse sequence acquires a 3-D k-space data set as a plurality of 2-D slices. 4. The method as recited in claim 1 in which the projections acquired in step b) sample k-space substantially uniformly. 5. The method as recited in claim 1 in which step f) includes: i) calculating the center of mass of each spatial projection using the first moment information and the 0th moment information; andii) shifting the spatial projection by an amount determined by its center of mass. 6. The method as recited in claim 1 in which the angle spacing is substantially 45° between successive projection acquisitions. 7. The method as recited in claim 1 in which step h) includes: vi) producing a second moment trajectory from the calculated second moments of successively acquired projections; andvii) the in-plane rotation motion time record is calculated by performing an inversion in which the in-plane rotational motion time record is fit to the second moment trajectory. 8. The method as recited in claim 7 in which the inversion is a weighted least squares fit. 9. The method as recited in claim 8 which includes: calculating the mean 0th moment of all the acquired spatial projections;calculating the deviation of each spatial projection 0th moment from the mean; andusing the deviations in the weighted least squares fit step. 10. A metho... | Mayo Foundation For Medical Education and Research | A61B005/05
US07469383 | B2 | 2008-12-23 | Medical apparatus for diagnosis or treatment | 2004-11-04 | The invention relates to a medical apparatus for diagnosis and/or treatment using a procedure (20), with an operating console (11) to control the use of components of the procedure (20), the operating console (11) having an input device (22), an output device (27) and an assistance module (26) to assist the user, which can be adapted to users' varying levels of training. This enables the user interfaces of the operating console to be adapted to the varying levels of knowledge of the medical operating personnel, so that the user in question is offered only those commands, with corresponding additional information, which the user has for example already learned during training. | 1. A medical apparatus for executing a medical procedure, comprising: an application device for executing the medical procedure having a plurality of application components;an operating console for controlling the application components;an input device for inputting a control command;an output device for displaying a graphic user interface; andan assistance module adapted to support a user for operating the medical apparatus according to a level of competence of the user;a control function configured in conjunction with the assistance module to control a user's access rights, and thereby vary and limit the user's ability to operate the device with the application components, based on a profile of the user so that, given multiple levels of access, access by a user to operate the device with the application components is limited to a level consistent with information in the profile, such that different users may be provided with different levels of access to control or operate and wherein when access is made available to a user the apparatus is configured to a mode of operation corresponding to a designated level for that user. 2. The medical apparatus according to claim 1, wherein the medical procedure includes an element chosen from the group consisting of acquiring medical data related to bodily functions, diagnosis of a medical condition and treatment of a medical condition of a patient and wherein access by a user to operate the device with the application components is limited to a level consistent with information in the profile corresponding to a level of knowledge of the user, such that different users may be provided with different levels of access to operate the device based on level of knowledge of each user. 3. The medical apparatus according to claim 1, wherein the assistance module includes a plurality of support data sets adapted to be displayed on the output device corresponding to a plurality of levels of competence. 4. The medical apparatus according to claim 1, wherein the operating console includes a sensor device for verifying the user. 5. The Medical apparatus according to claim 1, wherein the operating console includes a checking device for checking an access authorization, a skill level, and the competence level of the user using a data set representing the user. 6. The medical apparatus according to claim 1, wherein the graphic user interface is adapted to change a user-related screen displayed on the graphic user interface based on a data set representing the user. 7. The medical apparatus according to claim 1, wherein the operating console includes a memory device connected to the input device, the memory device having a plurality of stored user profiles for dynamically adapting the graphic user interface according to a current user profile input via the input device. 8. The medical apparatus according to claim 1, wherein a functionality of an application software for executing the medical procedure includes a hie... | Siemens Aktiengesellschaft | G06F003/00 G06F003/048 H04K001/00 G05B015/00 A61B005/00
US07438894 | B2 | 2008-10-21 | Dyes for organ function monitoring | 2003-10-07 | Highly hydrophilic indole and benzoindole derivatives that absorb and fluoresce in the visible region of light are disclosed. These compounds are useful for physiological and organ function monitoring. Particularly, the molecules of the invention are useful for optical diagnosis of renal and cardiac diseases and for estimation of blood volume in vivo. | 1. A composition comprising a compound of formulawherein R46, R47, R48, R49, R50, R51, R52, R53, R54, R55, R56, R57, R58, Y6, and Z6 are independently selected from the group consisting of —H, C1-C10 alkoxyl, C1-C10 polyalkoxyalkyl, C1-C20 polyhydroxyalkyl, C5-C20 polyhydroxyaryl, saccharides, amino, C1-C10 aminoalkyl, cyano, nitro, halogen, hydrophilic peptides, arylpolysulfonates, C1-C10 alkyl, C5-C20 aryl, —SO3T, —CO2T, —OH, —(CH2)aSO3T, —(CH2)aOSO3T, —(CH2)aNHSO3T, —(CH2)aCO2(CH2)bSO3T, —(CH2)aOCO(CH2)bSO3T, —(CH2)aCONH(CH2)bSO3T, —(CH2)aNHOC(CH2)bSO3T, —(CH2)aNHCONH(CH2)bSO3T, —(CH2)aNHCSNH(CH2)bSO3T, —(CH2)aOCONH(CH2)bSO3T, —(CH2)aPO3HT, —(CH2)aPO3T2, —(CH2)aOPO3HT, —(CH2)aOPO3T2, —(CH2)aNHPO3HT, —(CH2)aNHPO3T2, —(CH2)aCO2(CH2)bPO3HT, —(CH2)aCO2(CH2)bPO3T2, —(CH2)aOCO(CH2)bPO3HT, —(CH2)aOCO(CH2)bPO3T2, —(CH2)aCONH(CH2)bPO3HT, —(CH2)aCONH(CH2)bPO3T2, —(CH2)aNHCO(CH2)bPO3HT, —(CH2)aNHCO(CH2)bPO3T2, —(CH2)aNHCONH(CH2)bPO3HT, —(CH2)aNHCONH(CH2)bPO3T2, —(CH2)aNHCSNH(CH2)bPO3HT, —(CH2)aNHCSNH(CH2)bPO3T2, —(CH2)aOCONH(CH2)bPO3HT, —(CH2)aOCONH(CH2)bPO3T2, —CH2(CH2—O—CH2)c—CH2—OH, —(CH2)d—CO2T, —CH2—(CH2—O—CH2)e—CH2—CO2T, —(CH2)f—NH2, —CH2—(CH2—O—CH2)g—CH2NH2, —(CH2)h—N(Ra)—(CH2)i—CO2T, and —(CH2)j—N(Rb)—CH2—(CH2—O—CH2)kCH2—CO2T, with the proviso that when either Y6 or Z6 is alkyl the other is not alkyl, and when either Y6 or Z6 is —(CH2)dCO2T the other is not —(CH2)dCO2T;W6 is —S—;X6 is —S—;V6 is a single bond or is selected from the group consisting of —O—, —S—, —Se—, and —NRa;D6 is a single or a double bond;A6, B6 and E6 may be the same or different and are selected from the group consisting of —O—, —S—, —Se—, —P—, —NRa, —CRcRd, CRc, alkyl, and —C═O;A6, B6, D6, and E6 may together form a 6 or 7 membered carbocyclic ring or a 6 or 7 membered heterocyclic ring optionally containing one or more oxygen, nitrogen, or sulfur atoms;a b, d, f, h, i, and j independently vary from 1-10;c, e, g, and k independently vary from 1-100;a6 and b6 vary from 0 to 5;Ra, Rb, Rc, and Rd are defined in the same manner as Y6;T is either H or a negative charge. 2. The composition of claim 1 further comprising a contrast agent. 3. The composition of claim 1 wherein the compound comprises a radioactive halogen. 4. The composition of claim 1 further comprising a radioactive metal ion or a paramagnetic metal ion. 5. The composition of claim 1 in a formulation selected from the group consisting of a liposome, a micelle, a microcapsule, a microparticle, and combinations thereof. 6. A method of optically imaging a mammal comprising administering to a mammal an effective amount of a compound of the formula of claim 1 and thereafter performing the optical physiological function monitoring imaging procedure. 7. The method of claim 6 wherein said procedure utilizes light of wavelength in the region of 350-1300 nm. 8. The method of claim 6 wherein said procedure comprises monitoring a blood clearance profile by fluorescence using light of wavelength in the region of 350 to 13... | Mallinckrodt, Inc. | A61B010/00 A61B005/00 A61B008/00
US07445604 | B2 | 2008-11-04 | Blood sampling kit and method of using same | 2004-10-04 | Kits and methods for closed-system sampling blood in patients are disclosed. The kit includes a four-port stopcock adapted to allow communication at least between two ports. The stopcock has a first port adapted to receive a syringe, a second port for communicating with an infusion solution reservoir, a third port for receiving a flush solution, and a fourth port for communication with a patient. The kit also includes at least one sampling site adapted to fluidically communicate with the fourth port and to be positioned between the four-port stopcock and the patient. At least one of the sampling sites is adapted to allow extraction of a fluid being communicated between the stopcock and the patient. | 1. A method of blood sampling, comprising: a) drawing a fluid through a catheter port of a four-port stopcock into a syringe engaged to an axial port of the stopcock, the catheter port being in fluid communication with a patient, a solution inlet port of the four-port stopcock being adapted to communicate with a reservoir of infusion solution to be directed to the patient through the catheter port, said solution inlet port not in fluid communication with said patient while said fluid is drawn through said catheter port, said drawing thereby causing blood from said patient to flow into said catheter port via said fluid communication and to reach a sampling site positioned between said catheter port and said patient;b) extracting a blood sample through said sampling site; andc) returning fluid from said syringe through said catheter port to said patient. 2. The method according to claim 1, further comprising: d) manipulating said stopcock to thereby establish communication between a saline port of said four-port stopcock and said syringe;e) drawing a flush fluid through said saline port of the four-port stopcock into said syringe, said saline port being in communication with a reservoir of flush fluid;f) manipulating said stopcock to thereby establish communication between said syringe and said catheter port;g) directing said flush fluid from said syringe through said catheter port; andh) manipulating said stopcock to thereby establish communication between said solution inlet port and said catheter port. 3. The method according to claim 1, wherein the fluid drawn in step a) includes the infusion solution. 4. The method according to claim 1, wherein step a) includes drawing sufficient fluid to ensure availability of blood at the sampling site. 5. The method according to claim 1, wherein the syringe is a sheathed syringe having a reservoir for holding a fluid. 6. The method according to claim 5, wherein said syringe includes a silicone sheath. 7. The method according to claim 1, wherein the solution inlet port is adapted to communicate with the infusion solution reservoir through a pump. 8. A method of blood sampling in a system including a four-port stopcock and at least two sampling sites, the stopcock having an axial port adapted to receive a syringe, a solution inlet port for communicating with a reservoir of infusion solution, a saline port for receiving a flush solution, and a catheter port for communication with a patient, said sampling sites being positioned serially between said catheter port and a patient, said method comprising: a) drawing a clearance fluid through a first sampling site, thereby causing blood to flow from said patient to a second sampling site, said drawing sufficient to provide access to blood at a second sampling site said second sampling site being positioned closer to the patient than said first sampling site;b) blocking flow between said first sampling site and said second sampling site; andc) drawing blood th... | Biosensors International USA | A61B005/00 B65D081/00 A61M005/28
US07468040 | B2 | 2008-12-23 | Methods and systems for implantably monitoring external breathing therapy | 2004-08-30 | An implantable device is used to monitor one or more conditions associated with an external breathing therapy delivered to the patient. The device may monitor therapy parameters including therapy effectiveness, impact of the therapy on the patient, therapy usage, compliance with a prescribed usage, therapy interactions, and/or other parameters. | 1. A method of monitoring respiration therapy delivered to a patient, comprising: sensing one or more conditions associated with patient-external breathing therapy; andimplantably monitoring the patient-external breathing therapy based on the one or more sensed conditions;wherein implantably monitoring the patient-external breathing therapy comprises implantably monitoring patient use of the patient-external breathing therapy; andwherein implantably monitoring the patient use of the patient-external breathing therapy comprises comparing the patient use of patient-external breathing therapy with a prescribed use of the patient-external breathing therapy. 2. The method of claim 1, wherein sensing the one or more conditions comprises sensing one or more conditions associated with positive airway pressure breathing therapy. 3. The method of claim 1, wherein sensing the one or more conditions comprises sensing one or more cardiovascular system conditions. 4. The method of claim 1, wherein sensing the one or more conditions comprises sensing one or more respiratory system conditions. 5. The method of claim 1, wherein sensing the one or more conditions comprises sensing a proximity of the patient to a location of a therapy device delivering the patient-external breathing therapy. 6. The method of claim 1, wherein implantably monitoring the patient-external breathing therapy comprises implantably monitoring an effectiveness of the patient-external breathing therapy. 7. The method of claim 6, wherein implantably monitoring the effectiveness of the patient-external breathing therapy comprises: detecting disordered breathing episodes; andmonitoring the effectiveness of the patient-external breathing therapy based on the detected disordered breathing episodes. 8. The method of claim 6, wherein implantably monitoring the effectiveness of the patient-external breathing therapy comprises determining night-to-night changes in the effectiveness of the patient-external breathing therapy. 9. The method of claim 1, wherein implantably monitoring the patient use of the patient-external breathing therapy comprises: analying respiratory waveform morphology; andmonitoring the patient use of the patient-external breathing therapy based on the respiratory waveform morphology analysis. 10. The method of claim 1, further comprising providing one or more alert signals associated with the patient use of the patient-external breathing therapy. 11. The method of claim 1, wherein implantably monitoring the patient-external breathing therapy comprises implantably monitoring an impact of the patient-external breathing therapy on the patient. 12. The method of claim 1, wherein implantably monitoring the patient-external breathing therapy comprises implantably monitoring one or more interactions between the patient-external breathing therapy and another therapy delivered to the patient. 13. The method of claim 1, further comprising adapting the patient-external breathing... | Cardiac Pacemakers, Inc. | A61B005/08
US07457653 | B2 | 2008-11-25 | Method of solving magnetoencephalographic and electroencephalographic inverse problems | 2002-05-31 | A method of solving magnetoencephalographic and electroencephalographic inverse problems provides better grouping of moments of equivalent current dipoles than conventional methods. Locations of dipoles are fixed based on information using fMRI or the like. The magnitudes and orientations of the dipole moments are adjusted to fit magnetic fields and/or electrical potential distributions produced by the dipoles to magnetoencephalographically and/or electroencephalographically measured data. The method also includes grouping the dipoles in two stages, using the correlation coefficient of the magnitudes of the dipole moments as a criterion. | 1. A method of tracking brain activity, comprising: acquiring an image of brain activity showing a brain activation volume by a method other than magnetoencephalography or electroencephalography;dividing the brain activation volume into subvolumes;placing equivalent current dipoles in the subvolumes, respectively;adjusting magnitudes and orientations of equivalent current dipole moments of the equivalent current dipoles such that magnetic fields produced by the equivalent current dipoles fit data obtained by magnetoencephalography;grouping the equivalent current dipoles based on a criterion that a correlation coefficient of adjusted magnitudes is greater than a threshold; andoutputting to one of a display and a printer an activation state of each of parts of the brain by using grouped equivalent current dipoles to reproduce time sequences of data obtained by the magnetoencephalography. 2. A method of tracking brain activity, comprising: acquiring an image of brain activity showing a brain activation volume by a method other than magnetoencephalography or electroencephalography;dividing the brain activation volume into subvolumes;placing equivalent current dipoles in the subvolumes, respectively;adjusting magnitudes and orientations of equivalent current dipole moments of the equivalent current dipoles such that magnetic fields produced by the equivalent current dipoles fit data obtained by magnetoencephalography;grouping the equivalent current dipoles in a first grouping step of summing the equivalent current dipoles provided in a distance such that the equivalent current dipoles influence one another to set a new dipole at a centroid position weighted by magnitude of the equivalent current dipole moments;grouping the moments of the new equivalent current dipole set in the first grouping step and the equivalent current dipoles not grouped in the first grouping step in a second grouping step based on a criterion that a correlation coefficient of adjusted magnitudes is greater than a threshold; andoutputting to one of a display and a printer an activation state of each of parts of the brain by using grouped equivalent current dipoles obtained by the second grouping step to reproduce time sequences of data obtained by the magnetoencephalography. 3. A method of tracking brain activity, comprising: acquiring an image of brain activity showing a brain activation volume by a method other than magnetoencephalography or electroencephalography;dividing the brain activation volume into subvolumes;placing equivalent current dipoles in the subvolumes, respectively;adjusting magnitudes and orientations of equivalent current dipole moments of the equivalent current dipoles such that electrical fields produced by the equivalent current dipoles fit data obtained by electroencephalography;grouping the equivalent current dipoles based on a criterion that a correlation coefficient of adjusted magnitudes is greater than a threshold; andoutputting to one of a displa... | National Institute of Information Communications Technology | A61B005/05
US08355550 | B2 | 2013-01-15 | Methods and apparatus for virtual coronary mapping | 2008-05-01 | A virtual map of vessels of interest in medical procedures, such as coronary angioplasty is created so that doses of contrasting agent given to a patient may be reduced. A position of a coronary guidewire is determined and locations of vessel boundaries are found. When the contrast agent has dissipated, virtual maps of the vessels are created as new images. The locations of the determined vessel boundaries are imported to a mapping system and an image obtained without using a contrast agent is modified based on the imported locations of vessel boundaries. This creates a virtual map of the vessels. | 1. A method for mapping vessels comprising: determining a plurality of positions of a coronary guidewire in a plurality of non-contrast enhanced fluoroscopic images;determining, in a plurality of contrast enhanced fluoroscopic images, a plurality of reference points of a vessel segment corresponding to the plurality of positions of the coronary guidewire in a plurality of contrast enhanced fluoroscopic images; andmodifying a non-contrast enhanced fluoroscopic image based at least in part on the determined plurality of reference points of a vessel segment. 2. The method of claim 1 further comprising: determining a reference system relative to the coronary guidewire;detecting a contrast agent contrasting the vessels; andextracting the locations of a plurality of vessel boundaries relative to the determined reference system based on the contrasted vessels. 3. The method of claim 2 wherein the reference system is a relative coordinate system wherein a first dimension of the relative coordinate system is a length along the coronary guidewire relative to a reference point and a second dimension of the relative coordinate system is a distance between the coronary guidewire and a vessel boundary. 4. The method of claim 3 wherein the reference point is a location of a catheter along the coronary guidewire. 5. The method of claim 1 further comprising: determining a position of a coronary guidewire in a first of the plurality of non-contrast enhanced fluoroscopic images by applying an image-based guidewire detector, and image-based guidewire tip detector, and an image-based catheter end detector. 6. The method of claim 5 further comprising: learning the image-based guidewire detector from a set of guidewire images using marginal space learning and a probabilistic boosting tree;learning the image-based guidewire tip detector from a set of guidewire tip images using a probabilistic boosting tree and steerable features; andlearning the image-based catheter end detector from a set of catheter end images using a probabilistic boosting tree. 7. The method of claim 1 further comprising: determining a position of the coronary guidewire by tracking a two-dimensional guidewire shape in a plurality of images. 8. The method of claim 7 further comprising: tracking the coronary guidewire with a robust estimation of translational motion of a two-dimensional guidewire shape; andtracking the coronary guidewire with a robust estimation of non-rigid deformation of a two-dimensional guidewire shape. 9. The method of claim 8 wherein the position of the coronary guidewire is further refined by integrating results of a guidewire tip detection and a catheter end detection. 10. The method of claim 1 further comprising: determining a centerline and a plurality of vessel boundaries of a coronary vessel segment by applying an image-based vessel detector learned from a set of vessel segment images using marginal space learning and a probabilistic boosting tree. 11. The method... | Siemens Aktiengesellschaft | G06K009/00
US08108025 | B2 | 2012-01-31 | Flexible array for use in navigated surgery | 2007-04-24 | A system can be used to determine a position of a structure in physical space. The system can include a flexible member, such as a drape, cloth, sheet, etc. that can include one or more trackable devices associated therewith. The trackable devices can be incorporated into the weave of a woven cloth or the matrix of a polymer member. The trackable devices can be used within a tracking system to determine the location of the trackable devices. | 1. A system to determine a contour of a surface of a patient, the system comprising: a flexible member having an edge, a top surface, and a bottom surface, wherein the top surface and the bottom surface are substantially co-extensive and have an area greater than the edge, wherein the flexible member is at least partially woven including a first fiber interlaced with a second fiber; anda tracking device including a first coil formed on the first fiber and a second coil formed on the second fiber, wherein the tracking device is operable to determine a distinct position information regarding the surface of the patient;wherein the flexible member is formable to a contour of the patient and is operable to assist in a registration of a patient space to an image space. 2. The system of claim 1, at least one of the first coil or the second coil is operable to receive an electromagnetic field, produce an electromagnetic field, or combinations thereof. 3. The flexible member system of claim 2, further comprising: a tracking system including a localizer;wherein the localizer includes at least one of a coil operable to receive an electromagnetic field, a coil operable to produce an electromagnetic field, or combinations thereof. 4. The system of claim 3, further comprising: a processor operable to determine a contour of the patient based upon a determined position of the tracking sensor in the member. 5. The system of claim 4, wherein the contour is at least one of a 2D contour, a 3D contour, or combinations thereof. 6. The system of claim 1, where the flexible member is at least one of substantially non-porous or a completely woven sheet. 7. The system of claim 1, wherein the flexible member is a substantially non-porous member. 8. The flexible member system of claim 1, wherein the first coil is substantially orthogonal to the second coil. 9. The system of claim 8, wherein the flexible member includes a plurality of the first coils and a plurality of the second coils to form a plurality of the tracking devices. 10. The system of claim 4, wherein the member includes a plurality of the tracking devices; wherein each of the plurality of the tracking devices is operable to provide distinct position information;wherein a plurality of the distinct position information can be analyzed with the processor to determine a contour of the patient. 11. The system of claim 10, further comprising: an imaging system operable to acquire image data of the patient;wherein the processor is operable to determine an image data contour;wherein the processor is operable to correlate the determined contour of the patient and the image data contour. 12. The system of claim 11, wherein the determined contour of the patient and the image data contour are each at least one of a 2D contour, a 3D contour, or combinations thereof. 13. The flexible member system of claim 1, wherein the member defines a passage and a passage cover; wherein the passage cover is operable to be mov... | Medtronic, Inc. | A61B005/05
US08092445 | B2 | 2012-01-10 | Fluid connection device, system and method for continuously taking fluid microsamples using this device | 2007-12-12 | The invention relates to a fluid connection device and to a continuous automated system and method for taking fluid microsamples using this device.;A device (20) according to the invention, which is intended to be connected to a first line (21) via a first opening and includes a second opening through which a second line (28) passes, comprises: a female fluid connector (22) that defines the first opening and has an internal fitting surface (23) terminating in a female end (24) into which this first line opens; anda male fluid connector (25) that defines the second opening, which is fitted into the female connector via its external surface (26) and terminates via its male radial end (27) inside the female connector.;According to the invention, this second line is formed from a flexible microtube pushed right through the male connector beyond the male end, the free end of the second line pressing in a sealed manner against the female end so as to minimize the dead volume between the first line and the male connector. | 1. A fluid connection device which is useable for microsamples to be taken or to be injected and which is connected to a first line via a first opening in this device, which includes a second opening through which a second line passes, the second line communicating with the first line in order to transfer said microsamples, the device comprising: said first and second lines;a female fluid connector that defines said first opening and has an internal fitting surface terminating in a female radial end into which this first line opens; anda male fluid connector that defines said second opening, which is fitted into the female connector via its external surface and terminates via its male radial end inside the female connector,wherein this second line is formed from a flexible microtube which is pushed through the male connector axially beyond said male end, the free end of this second line pressing in a sealed manner against said female end so as to minimize a dead volume between the first line and the male connector. 2. The fluid connection device as claimed in claim 1, wherein said internal fitting surface of the female connector is a conical surface converging on said female end with the same conicity as said external surface of the male connector, which converges on said male end. 3. The fluid connection device as claimed in claim 2, wherein said male and female connectors are both connectors of the “Luer” type, as defined by the ISO 594-1 standard of 1986, or else of the “Luer-lock” type, as defined by the ISO 594-2 standard of 1998. 4. The fluid connection device as claimed in claim 3, wherein said second line extends beyond said male end by an axial length at least equal to the minimum distance (dl) separating said respective ends of the two connectors when the male connector is pushed into the connection position in the female connector in accordance with one or other of said standards. 5. The fluid connection device as claimed in claim 1, wherein said second line is provided, around its cylindrical wall and near its free end, with a stiffening means capable of stiffening it inside said female connector. 6. The fluid connection device as claimed in claim 5, wherein said stiffening means is formed from a ring made of a material that has a stiffness at least equal to that of said second line and is capable of being fastened thereto, this ring being mounted so as to bear between said second line and said conical internal fitting surface of said female connector. 7. The fluid connection device as claimed in claim 1, wherein said first line a flexible microtube suitable for taking said liquid from an animal or injecting it thereinto. 8. The fluid connection device as claimed in claim 6, wherein said stiffness of the ring is greater than that of said second line, said ring comprising a polymeric resin. | Commissariat a l'Energie Atomique | A61M025/16
US07972272 | B2 | 2011-07-05 | Electrostatically driven image probe | 2008-01-22 | The present invention provides an imaging probe for imaging mammalian tissues and structures using high resolution imaging, including ultrasound and optical coherence tomography. The imaging probe uses electrostatic discharge for actuating a metalized cantilever which holds the end portion of imaging assembly which emits energy in the imaging probe device by holding the cantilever at a potential such that it is neither grounded or charged such that the only electrical path to ground is through a dissipative polymer forming part of the device which is enveloped by a wire or coil held at ground potential. A high voltage electrode attracts the metalized cantilever and the dissipative polymer is used to connect to the cantilever to ground causing the cantilever to move about thereby scanning the region in front of the probe. | 1. Electrostatically driven imaging probe, comprising: a) an elongate hollow catheter sheath having distal front and back sections and an elongate proximal section and having a diameter suitable to permit insertion of the elongate hollow catheter sheath into bodily lumens and cavities, the distal back section containing an electrically dissipative polymer sealed therein which is wrapped by a metal coil for connection to ground potential and a timing circuit;b) imaging means located in said distal front section for emitting energy and receiving energy reflected back from interior surfaces of said bodily lumens and cavities, said distal front section containing one of a medium and vacuum sealed therein which allows transmission of imaging energy emitted by the imaging means, said imaging means being connected to an imaging conduit which extends through a proximal end of the elongate hollow catheter sheath for connection to an image processing and display system, said imaging conduit being enveloped by metal and a portion of the metal enveloped imaging conduit in the distal front section forming a cantilever; andc) an elongate electrode located in the elongate hollow catheter sheath having an elongate uninsulated electrode section located in said front distal section, the elongate electrode being connected to a high voltage power supply, wherein in operation when a high voltage is applied to the elongate electrode the cantilever is electrically attracted to the elongate uninsulated electrode section and undergoes deflection towards the elongate uninsulated electrode section and upon contacting the elongate uninsulated electrode section the portion of the metal enveloped imaging conduit in the distal front section acquires an electrical charge from the elongate uninsulated electrode section thereby causing the metal enveloped imaging conduit in the distal front section to be repelled therefrom thereby causing the imaging means to scan the field of view. 2. The electrostatically driven imaging probe according to claim 1 wherein said imaging conduit is an optical fiber and wherein said imaging means includes optical radiation emitted from an end portion of said optical fiber located in said distal front section, said optical fiber being optically coupled to a source of optical radiation through said proximal end of the elongate hollow catheter sheath. 3. The electrostatically driven imaging probe according to claim 2 wherein said imaging means includes focusing optics located in a distal end portion of said distal front section for directing and focusing light emitted from said optical fiber and collecting light reflected back into said distal front section. 4. The electrostatically driven imaging probe according to claim 3 wherein said focusing optics includes a GRIN lens mounted in the distal end portion of said distal front section. 5. The electrostatically driven imaging probe according to claim 3 wherein said focusing optics includes a ball le... | University Health Network | A61B008/00 A61B005/00 G02B026/10
US08165844 | B2 | 2012-04-24 | Motion tracking system | 2007-05-15 | A system is provided for capturing motion of a moving object via a plurality of motion sensor modules placed on various body segments. The sensor modules capture both 3D position and 3D orientation data relating to their respective body segments, thereby gathering motion data having six degrees of freedom with respect to a coordinate system not fixed to the body. Each body sensor collects 3D inertial sensor data and, optionally, magnetic field data. In embodiments, either DSP circuitry within the sensor modules or an external computing device, processes the sensor data to arrive at orientation and position estimates by using an estimation algorithm, such as a Kalman filter or a particle filter. The processing includes biomechanical model constraints that allow flexibility in the joints and provide characteristics for various joint types. To improve estimation accuracy, the system may be integrated with various types of aiding sensors. | 1. A method for capturing motion of an object having a plurality of segments via a plurality of motion sensor modules located on the segments, each of two or more of the segments of the object having thereon at least one sensor module, each sensor module being adapted to capture 3D position data and 3D orientation data associated with its respective segment, the method comprising: collecting 3D inertial sensor data at each sensor module via at least one 3D accelerometer sensor and at least one 3D angular velocity sensor; andcalculating a 3D orientation and 3D position of each segment having a sensor module by integrating the 3D angular velocity data and doubly integrating the 3D accelerometer data in time while applying a joint constraint to correct the calculated segment positions and orientations such that the segments are constrained to be linked to one another via a joint having at least one joint characteristic property, wherein applying the joint constraint is executed consistent with the at least one joint characteristic property and includes allowing a predetermined degree of laxity in one or more joints. 2. The method according to claim 1, wherein applying a joint constraint to correct the calculated segment positions and orientations further includes providing for two or more distinct biomechanical characteristics corresponding to two or more respective distinct joint types. 3. The method according to claim 1, further comprising updating the 3D orientation and 3D position calculation with one or more aiding sensors selected from the group consisting of a magnetic sensor, a GPS receiver, an RF-based local positioning sensor, a barometer, a pressure sensor, and a camera. 4. The method according to claim 3, wherein the one or more aiding sensors include a camera and the camera is adapted to improve the object's 3D position and 3D orientation calculation by providing at least 2D positional data of the object. 5. The method according to claim 1, further comprising updating the 3D orientation and 3D position calculation by detecting an external contact of the object with predetermined surroundings of the object based on at least one of position, velocity, and acceleration of one or more segments. 6. The method according to claim 1 wherein the object's surroundings are characterized by one or more predetermined assumptions, the method further comprising updating the one or more predetermined assumptions via a simultaneous localization and mapping technique. 7. A system for capturing motion of an object having a plurality of segments, the system comprising: a plurality of motion sensor modules including at least one sensor module located on each of two or more segments, each sensor module being adapted to capture 3D inertial sensor data; andat least one sensor fusion circuit for estimating 3D position and 3D orientation of the plurality of segments based on the 3D inertial sensor data captured by the motion sensor modules, and being adapt... | Xsens Holding B.V. | A61B005/103
US07940975 | B2 | 2011-05-10 | Medical image diagnosing support method and apparatus, and image processing program configured to extract and correct lesion candidate region | 2005-07-28 | A medical image diagnosing support method is provided that includes a site region extracting step for obtaining a tomographic image which is picked up by a medical diagnostic imaging apparatus and extracting a predetermined site region from the obtained tomographic image, a first region extracting step for extracting a first lesion candidate region from the site region based on pixel values of the site region extracted in the predetermined site region extracting step, a second region extracting step for extracting a second lesion candidate region from the site region based on a distribution of the pixel values of the site region extracted in the predetermined site region extracting step, and a region correcting step for correcting the first lesion candidate region extracted in the first region extracting step by using the second lesion candidate region extracted in the second region extracting step. | 1. A medical image diagnosing support method, comprising: a site region extracting step of obtaining, by a medical image diagnosing apparatus, a tomographic image, and extracting a predetermined site region from the obtained tomographic image;a first region extracting step of extracting a first lesion candidate region from the site region in the tomographic image based on pixel values of the site region extracted in the predetermined site region extracting step;a second region extracting step of extracting a second lesion candidate region from the site region in the tomographic image based on a distribution of pixel values of the site region extracted in the predetermined site region extracting step; anda region correcting step of extracting as a pulmonary emphysema a corrected lesion candidate region where the first lesion candidate region overlaps with the second lesion candidate region. 2. The medical image diagnosing support method according to claim 1, wherein the distribution of the pixel values of the predetermined site region is calculated by using at least one of an area ratio, a mean ratio, and a standard deviation ratio of the predetermined site region and the first lesion candidate region. 3. The medical image diagnosing support method according to claim 1, wherein the distribution of the pixel values of the predetermined site region is calculated by dividing the extracted predetermined site region into at least two parts, and using at least one of an area ratio, a mean ratio, and a standard deviation ratio of the divided predetermined site region and the first lesion candidate region. 4. The medical image diagnosing support method according to claim 3, wherein the division of the predetermined site region includes a division into a number of parts, the number depending on the predetermined site region. 5. The medical image diagnosing support method according to claim 1, further comprising: a progression calculating step of calculating a progression of the first lesion candidate region which is corrected in the region correcting step; anda displaying step of changing a color of the corrected first lesion candidate region depending on the progression of the overlapped lesion candidate region calculated in the progression calculating step, and displaying the overlapped lesion candidate region having the changed color in association with the tomographic image. 6. The medical image diagnosing support method according to claim 5, wherein, in the displaying step, the tomographic image is displayed superimposed on the overlapped region. 7. The medical image diagnosing support method according to claim 5, wherein, in the displaying step, the overlapped region is displayed next to the tomographic image. 8. The medical image diagnosing support method according to claim 5, wherein the displaying step includes a selecting step of selecting a display or non-display mode of the overlapped region. 9. A medical image diagnosing support appara... | Hitachi Medical Corporation | G06K009/00
US08187202 | B2 | 2012-05-29 | Method and apparatus for acoustical outer ear characterization | 2006-09-13 | The present invention relates to a method and apparatus for characterizing acoustical properties of an outer ear 130, the method comprising the steps of: transmitting a first acoustic signal 125 towards the outer ear 130, receiving a second acoustic signal 150 from the outer ear 130, and characterizing acoustical properties 165 of the outer ear on the basis of the second acoustic signal 150. The method is characterized in that the first acoustic signal 125 comprises at least one of the following elements: music, and speech. The present invention further relates to a method and apparatus for enrolling, authenticating and identifying a person on the basis of acoustical properties of an outer ear 130. | 1. A method of characterizing acoustical properties of an outer ear, the method comprising: transmitting a first acoustic signal towards the outer ear, the first acoustic signal comprising at least one of music and speech,receiving a second acoustic signal from the outer ear,characterizing acoustical properties of the outer ear on the basis of the second acoustic signal; andadding spectral components to the at least one of music and speech to form the first acoustic signal so as to complement the audible spectrum of the at least one of music and speech. 2. The method of claim 1, wherein the method further comprises transmitting an acoustic noise signal towards the outer ear. 3. The method of claim 1, wherein the first acoustic signal is based on a representation of an audio signal, the representation comprising one of the following: an electromagnetic representation of an audio signal,an optical representation of an audio signal,an electrical representation of an audio signal, anda digital representation of an audio signal. 4. The method of claim 3, wherein the digital representation is a lossy compressed audio signal. 5. The method of claim 3, wherein the method further comprises processing the representation of the audio signal before converting the representation into an acoustic signal. 6. The method of claim 5, wherein the processing comprises the addition of noise to the representation of the audio signal. 7. The method of claim 5, wherein the processing comprises the addition of noise to the representation of the audio signal, the noise formed using a psycho-acoustic model of the human auditory system in order to reduce perceptual disturbance as a result of the added noise. 8. The method of claim 1, wherein the characterizing the acoustical properties of the outer ear involves a frequency domain analysis of the second acoustic signal. 9. The method of claim 8, wherein the characterizing the acoustical properties of the outer ear involves characterizing a transfer function of the acoustical properties of the outer ear. 10. The method of claim 1, further comprising enrolling the acoustical properties of an outer ear of a person being the object of the characterizing. 11. The method of claim 10, further comprising authenticating a person using previously enrolled acoustical properties of an outer ear. 12. The method of claim 10, further comprising identifying a person using previously enrolled acoustical properties of an outer ear. 13. An apparatus for characterizing acoustical properties of an outer ear, the apparatus comprising: a transmitter for transmitting a first acoustic signal towards the outer ear, the first acoustic signal comprising at least one of music and speech,an encoder for adding spectral components to the at least one of music and speech to form the first acoustic signal so as to complement the audible spectrum of the at least one of music and speech,a receiver for receiving a second acoustic signal from the out... | Koninklijke Philips Electronics N.V. | A61B005/00
US08131352 | B2 | 2012-03-06 | System and method for automatically adjusting detection thresholds in a feedback-controlled neurological event detector | 2007-06-20 | Methods and systems for detecting neurological events and approximating a target detection rate are disclosed. A target detection rate may be identified for the neurological event. Electrographic signals incident on a neurological event detector may be monitored. Each signal may be compared to a threshold value for a parameter. As the threshold value varies, it has a predictable effect on a detection rate of the neurological event. A rate at which the electrographic signals exceed the threshold value may be measured and compared to the target detection rate. The threshold value may be adjusted to minimize the difference between the measured detection rate and the target detection rate. | 1. A method for approximating a target detection rate for a neurological event with a neurological event detector intended for use with a human patient, comprising: identifying a target detection rate for the neurological event;monitoring electrographic signals incident on the neurological event detector;comparing, with a processor, each electrographic signal to a threshold value for a parameter the variation of which has a predictable effect on a detection rate of the neurological event;measuring a rate at which the electrographic signals exceed the threshold value to determine a measured detection rate;comparing the measured detection rate to the target detection rate; andautomatically adjusting the threshold value to minimize a difference between the measured detection rate and the target detection rate based on the parameter and the predictable effect that the variation of the parameter has on the rate of detection of the neurological event. 2. The method of claim 1, wherein comparing each electrographic signal to a threshold value for a parameter the variation of which has a predictable effect on a detection rate of the neurological event includes: comparing each electrographic signal to a threshold value for a line length parameter which corresponds to a percentage difference between accumulated line lengths for a first electrographic signal in a window of time and accumulated line lengths for a second electrographic signal in the window of time. 3. The method of claim 1, wherein comparing each electrographic signal to a threshold value for a parameter the variation of which has a predictable effect on a detection rate of the neurological event includes: comparing each electrographic signal to a threshold value for a line length parameter which corresponds to a percentage difference between accumulated line lengths for an electrographic signal in a first window of time and accumulated line lengths for the electrographic signal in a second window of time. 4. The method of claim 3, wherein the accumulated line lengths for the electrographic signal in a first window of time are normalized based on a duration of the first window of time and the accumulated line lengths for the electrographic signal in the second window of time are normalized based on a duration of the second window of time. 5. The method of claim 1, wherein comparing each electrographic signal to a threshold value for a parameter the variation of which has a predictable effect on a detection rate of the neurological event includes: comparing each electrographic signal to a threshold value for an area parameter which corresponds to the difference between an accumulated area under the curve for a first electrographic signal in a window of time and an accumulated area under the curve for a second electrographic signal in the window of time. 6. The method of claim 1, wherein comparing each electrographic signal to a threshold value for a parameter the variation of which has a ... | NeuroPace, Inc. | A61B005/04
US08388543 | B2 | 2013-03-05 | Photoplethysmography apparatus and method employing high resolution estimation of time-frequency spectra | 2007-05-16 | A photoplethysmography apparatus and method is provided for high resolution estimating of Time-Frequency Spectra (TFS) and associated amplitudes using Variable Frequency Complex Demodulation (VFCDM), in a two-step procedure using a Time-Varying Optimal Parameter Search (TVOPS) technique to obtain TFS, followed by VFCDM to obtain even greater TFS resolution and instantaneous amplitudes associated with only specific frequencies of interest, via the combined TVOPS and VFCDM. | 1. A digital filtering method for estimating Time-Frequency Spectra (TFS) and associated amplitudes of a photoplethysmography (PPG) signal using Variable Frequency Complex Demodulation (VFCDM), the method comprising: obtaining a first TFS of the PPG signal by performing, with a controller, a time-varying optimal parameter search;performing, with the controller, the VFCDM with the first TFS of the PPG signal to obtain at least one variable frequency complex demodulation signal having improved TFS resolution and instantaneous amplitudes associated with only specific frequencies of interest; andsimultaneously deriving, with the controller, a heart rate and a respiratory rate of a subject from the at least one variable frequency complex demodulation signal. 2. A photoplethysmography (PPG) apparatus for monitoring heart and respiratory rates, the apparatus comprising: a controller configured to: obtain a first Time-Frequency Spectra (TFS) of a PPG signal by performing a time-varying optimal parameter search,perform a Variable Frequency Complex Demodulation (VFCDM) with the first TFS of the PPG signal to obtain at least one VFCDM signal having improved TFS resolution and instantaneous amplitudes associated with only specific frequencies of interest, andsimultaneously derive a heart rate and a respiratory rate of a subject from the at least one VFCDM signal. 3. The apparatus of claim 2, wherein the apparatus is configured to provide improved diagnosis treatment of sleep apnea. 4. The apparatus of claim 2, wherein the apparatus is configured to provide an electrocardiograph signal. 5. The apparatus of claim 2, wherein the apparatus is configured to detect sudden infant death syndrome. 6. The apparatus of claim 2, wherein the apparatus is configured to detect hypovolemia. 7. The apparatus of claim 2, wherein the apparatus is configured to determines renal blood flow data. | The Research Foundation of State University of New York | A61B005/02
US08444595 | B2 | 2013-05-21 | Methods to pair a medical device and at least a remote controller for such medical device | 2007-06-15 | A disease management system, methods, and devices are shown and described. In one embodiment, the system includes an infusion pump and a remote controller with the ability to be paired to each other. A method to verify a wireless connection between an infusion pump and a remote controller is shown and described herein. In a further embodiment, a method to verify a wireless connection between an infusion pump and a remote controller is provided. In addition, a method of operating a diabetes management system is provided in which the system includes an infusion pump and at least a remote controller. | 1. A method to verify a wireless connection between an infusion pump and a remote controller, the method comprising: placing a remote controller in a wireless link with the infusion pump;providing identification information specific to the infusion pump to the remote controller;providing identification information specific to the remote controller on the infusion pump; andconfirming that the infusion pump's identification is received by the remote controller and that remote controller identification information is received by the infusion pump, in which the confirming includes displaying identification specific to the remote controller on a display of the infusion pump and identification information specific to the infusion pump on a display of the remote controller. 2. The method of claim 1, in which the identification information specific to the infusion pump is selected from a group consisting essentially of a serial number of the device; names; icons; avatars, speech identification, sounds, or combinations thereof. 3. The method of claim 1, in which the identification information specific to the controller is selected from a group consisting essentially of a serial number of the device; names; icons; avatars, speech identification, sounds, or combinations thereof. 4. The method of claim 1, further comprising connecting another remote controller to the infusion pump via a wireless link;providing identification information specific to the infusion pump to the another remote controller; andproviding identification information specific to the another remote controller on the infusion pump. 5. The method of claim 4, further comprising: ignoring any command from the remote controller upon confirmation of the another remote controller by the user. 6. The method of claim 1, in which the remote controller identification information comprise a user selected indicia. 7. The method of claim 6, in which the user-selected indicia is selected from a group consisting of a user selected text nomenclature, user selected icon, user selected audio, user selected visual indicia, and combinations thereof. | Animas Corporation | A61M031/00
US09622681 | B2 | 2017-04-18 | Method and system eliminating electromagnetic interference in a medical image, in real-time | 2006-01-30 | In a medical apparatus including a medical imaging system and a medical position and navigation system (MPS), the medical imaging system including an imaging transmitter, periodically emitting imaging radiation and an imaging detector, the medical position and navigation system including at least one MPS transmitter periodically transmitting MPS radiation and at least one MPS detector, the MPS radiation electromagnetically interfering with at least one mode of operation of the imaging detector, a device for eliminating interference to the imaging detector caused by positioning radiation, the device comprising a synchronizer, coupled with the medical imaging system and with the medical position and navigation system, synchronizing the imaging detector and each the at least one MPS transmitter, so that neither of the at least one MPS transmitter transmits during the at least one mode of operation of the imaging detector. | 1. In a medical apparatus including a medical imaging system and a medical position and navigation system (MPS), the medical imaging system including an imaging transmitter periodically emitting imaging radiation and a separate imaging detector having at least an imaging radiation acquisition mode of operation and a sampling mode of operation, the MPS including at least one MPS transmitter periodically transmitting MPS radiation and at least one MPS detector, with one of the at least one MPS transmitter and the at least one MPS detector being positionable within a body, said MPS radiation electromagnetically interfering with at least one of said modes of operation of said imaging detector, a device for eliminating interference to the imaging detector caused by positioning radiation, the device comprising: a synchronizer configured to be coupled with said medical imaging system and with said MPS, wherein said synchronizer is configured to synchronize said imaging detector and said at least one MPS transmitter, so that said at least one MPS transmitter does not transmit during said at least one of said modes of operation of said imaging detector, wherein said at least one of said modes of operation of said imaging detector is said imaging radiation acquisition mode. 2. The device according to claim 1, wherein said at least one of said modes of operation of said imaging detector is said sampling mode. 3. A combined imaging and positioning apparatus comprising: a medical imaging system, for obtaining a representation of the anatomy of a portion of a body, said medical imaging system including an imaging radiation transmitter for periodically transmitting imaging radiation and a separate imaging detector having at least an imaging radiation acquisition mode of operation and a sampling mode of operation;a medical position and navigation system (MPS), said MPS including at least one MPS transmitter for transmitting MPS radiation, said MPS radiation electromagnetically interfering with at least one of said modes of operation of said imaging detector, and at least one MPS detector for detecting MPS radiation, with one of said at least one MPS transmitter and said at least one MPS detector being positionable within the body; anda synchronizer coupled with said medical imaging system and with said MPS, wherein said synchronizer is configured to synchronize said imaging detector and said at least one MPS transmitter, so that said at least one MPS transmitter does not transmit during said at least one of said modes of operation of said imaging detector, wherein said at least one of said modes of operation of said imaging detector is said imaging radiation acquisition mode. 4. The system according to claim 3, wherein said at least one of said modes of operation of said imaging detector is said sampling mode. 5. The apparatus according to claim 3, further including a display unit coupled with said medical imaging system and said medical positioning system, ... | MediGuide Ltd. | A61B005/05 A61B005/06 A61B090/00 A61B034/20
US20080319286 | A1 | 2008-12-25 | Optical Probes for Non-Invasive Analyte Measurements | 2008-08-04 | An optical probe for non-invasively measuring an analyte property in a biological sample of a subject, comprises a plurality of illumination fibers that deliver source light from an optical probe input to a sample interface, a plurality of collection fibers that deliver light returned from the sample interface to an optical probe output, and wherein the illumination and collection fibers are oriented substantially perpendicular to the sample interface and the illumination and collection fibers are stacked in a plurality of linear rows to provide a stack of fibers arranged in a rectangular pattern. The optical probe is amenable to manufacturing on a scale consistent with a commercial product. | 1. An optical probe for non-invasively measuring an analyte property in a biological sample of a subject, comprising: a plurality of illumination fibers that deliver source light from an optical probe input to a sample interface,a plurality of collection fibers that deliver light returned from the sample interface to an optical probe output, andwherein the illumination and collection fibers are oriented substantially perpendicular to the sample interface and the illumination and collection fibers are stacked in a plurality of linear rows to provide a stack of fibers. 2. The optical probe of claim 1, wherein the stack of fibers forms a rectangle. 3. The optical probe of claim 2, wherein the stack of fibers forms a square. 4. The optical probe of claim 1, wherein the illumination and collection fibers comprise separate rows in the stack of fibers. 5. The optical probe of claim 4, wherein the illumination and collection fibers comprise alternating separate rows in the stack of fibers thereby providing a linear stack of fibers. 6. The optical probe of claim 1, wherein every other row in the stack of fibers consists of illumination fibers and the intervening rows comprise both illumination and collection fibers. 7. The optical probe of claim 6, wherein the intervening rows comprise alternating illumination and collection fibers such that each collection fiber has eight adjacent illumination fibers thereby providing a linear stack 8:1 of fibers. 8. The optical probe of claim 1, wherein each linear row comprises alternating illumination and collection fibers such that each collection fiber has four adjacent illumination fibers thereby providing an alternating linear stack of fibers. 9. The optical probe of claim 1, further comprising an optical homogenizer at the optical probe input to homogenize the source light at the input of the illumination fibers. 10. The optical probe of claim 1, further comprising an optical homogenizer at the optical probe output to homogenize the return light at the output of the collection fibers. 11. The optical probe of claim 10, further comprising an aperture at the output of the optical homogenizer to reduce the size of the optical probe output. 12. The optical probe of claim 1, wherein the numerical aperture of the illumination fibers is different than the numerical aperture of the collection fibers. 13. The optical probe of claim 1, wherein the illumination and collection fibers comprise a silica core and a cladding comprises fused silica, doped silica, Teflon, or a fluoropolymer. 14. The optical probe of claim 1, wherein the relative spacing, angle, numerical aperture, and placement of the illumination and collection fibers are arranged to achieve depth targeting. 15. The optical probe of claim 1, further comprising means to control the temperature of the sample interface. 16. The optical probe of claim 1, further comprising an index matching fluid at the optical interface between the sample and the samp... | - | A61B005/1455
US20080287775 | A1 | 2008-11-20 | ARRANGEMENT FOR AFFIXING A MAGNETIC RESONANCE RF COIL TO A PATIENT | 2008-05-15 | In an arrangement to affix RF coils (in particular local coils) on a patient in examinations with a magnetic resonance apparatus, using a horizontal board to accommodate the patient for a magnetic resonance examination to be implemented; with the RF coil that has coil electronics; a fastening strap is fashioned to affix the coil on the patient and the fastening strap has an connection cable integrated therewith. The coil electronics are supplied with energy via the integrated connection cable. | 1. An arrangement to affix an RF coil on a patient for a magnetic resonance examination using the RF coil, comprising: a board configured to receive a patient thereon in a prone position for implementing a magnetic resonance examination of the patient;an RF coil for use in said magnetic resonance examination, said coil having coil electronics physically associated therewith;a fastening strap connected to the coil to affix the coil on the patient on the board; anda connection cable connected to said coil electronics for supplying power to said coil electronics, said connection cable being integrated with said fastening strap. 2. An arrangement as claimed in claim 1 comprising a latch connection that connects said fastening strap to said board. 3. An arrangement as claimed in claim 2 wherein said latch connection comprises electrical contacts integrated therein. 4. An arrangement as claimed in claim 3 wherein said contacts are galvanic contacts. 5. An arrangement as claimed in claim 1 wherein said connection cable integrated with said fastening strap is connected with said RF coil via an electrical connection selected from the group consisting of a galvanic connection, inductive coupling elements and capacitive coupling elements. 6. An arrangement as claimed in claim 1 wherein said connection cable or said fastening strap comprises a signal coupler integrated therewith for transferring electrical signals between the RF coil and the connection cable. 7. An arrangement as claimed in claim 6 wherein said integrated signal coupler is a signal coupler selected from the group consisting of capacitive couplers and inductive couplers. 8. An arrangement as claimed in claim 7 wherein said RF coil has a housing comprising a coupler counterpart connected to the signal coupler to transfer signals between the RF coil and the signal coupler. 9. An arrangement as claimed in claim 8 comprising a fastening strap guide, and wherein said coupler counterpart is located below said fastening strap guide and is pressed by the fastening strap onto the signal coupler. 10. An arrangement as claimed in claim 1 wherein said RF coil is configured as a local coil. | - | A61B005/055
US20080306350 | A1 | 2008-12-11 | Methods for Implementing a Weight Loss Program | 2007-06-11 | A method for implementing a weight loss program is provided. The diet meal provider educates a health care professional about a diet meal plan or a dietary aid. The health care professional may recommend the plan to his or her patient who need to loose weight. Based on such recommendation, the patient will subscribe to receive the diet meal plan or the dietary which are supplied to the patient by the diet meal provider. | 1. A method for implementing a weight loss program, the method comprising: providing information about a diet meal plan and a dietary aid directly to a health care professional;receiving a first request for the diet meal plan or a dietary aid for a patient wherein the request is caused by the health care professional's recommendation after evaluating the patient;supplying the diet meal plan or the dietary aid to the patient; andbilling and receiving a payment from the health care professional for the supplied diet meal plan. 2. The method of claim 1 wherein the first request is received from either the patient or the health care professional. 3. The method of claim 1 wherein the dietary aid is selected from the group consisting of dietary drugs, fiber supplements, dietary supplements, and combinations thereof. 4. The method of claim 1 further comprising billing the patient for the supplied diet meal plan. 5. (canceled) 6. The method of claim 1, wherein the patient is supplied with the diet meal plan once a week. 7. The method of claim 1 wherein the first request from the patient comprises the patient's selection for individual breakfast, lunch, and dinner for each day. 8. The method of claim 1 further comprising receiving a second altering the first request. 9. The method of claim 1 wherein the diet meal plan is supplied to the patient's home, doctor's office, or both. 10. The method of claim 1 further comprising modifying the diet meal plan supplied to the patient based on the patient's progress. 11. The method of claim 1, wherein the patient suffers one or more health condition selected from the group consisting of hypertension, dyslipidemia, diabetes, coronary heart disease, stroke, gallbladder diseases, osteoarthritis, sleep apnea and respiratory problems, endometrial cancer, breast cancer, and colon cancer. 12. The method of claim 1, wherein the health care professional is selected from the group consisting of medical doctors, doctors of osteopathic medicine, nutritionists, dietitians, chiropractors, physical therapists, and veterinarians. 13. The method of claim 1 further comprising training the health care professional's assistant about different options within the diet meal plan wherein the patient is educated by the health care professional's assistant about different options within the meal plan prior to making the request for the diet meal plan. 14. A method for implementing a weight loss program, the method comprising: providing information about a diet meal plan and a dietary aid directly to a health care professional;receiving information about a patient from the health care professional's office comprising the patient's payment information and a first request for the diet meal plan or a dietary aid wherein the request is caused by the health care professional's recommendation after evaluating the patient;in response to receiving information about the patient, creating an account for the patient;in response to the reque... | - | A61B005/00
US20080269597 | A1 | 2008-10-30 | Gauze or the Like Having a Device for its Detection at a Distance | 2006-08-30 | A gauze or other material for surgical use comprising at least one presence detection device for remotely identifying said gauze or other material for surgical use. The presence detection device can be a radio-opaque element or can be activated by a radio-frequency or a magnetic signal. It can be provided with a chip, an antenna and a memory. | 1. A gauze or other material for surgical use, characterized in that it comprises at least one presence detection device for remotely identifying said gauze or other material for surgical use. 2. The gauze or the like according to claim 1, characterized in that it comprises more warp and weft webs wherein said at least one detection device is fastened to a web or between said webs by bonding or within a pocket obtained on said gauze or at the surface of said gauze by means of gripping elements. 3. The gauze or the like according to claim 1, characterized in that at least one detection device is substantially thread-like and is a part of the warp and/or weft of said webs. 4. The gauze or the like according to claim 1, characterized in that said presence detection device comprises at least one element that is capable of inducing an intensity variation of a magnetic field. 5. The gauze or the like according to claim 1, characterized in that said presence detection device comprises at least one integrated electronic circuit (CHIP) and at least one antenna for said integrated electronic circuit, said at least one antenna being capable of receiving and transmitting radio frequency waves. 6. The gauze or the like according to claim 5, characterized in that said presence detection device comprises a memory element capable of containing useful additional information to be communicated to the outside. 7. The gauze or the like according to claim 1, characterized in that it comprises at least one radiopaque element. | - | A61B005/06 A61F013/00
US20080269569 | A1 | 2008-10-30 | FOLLOW-UP SUPPORT SYSTEM FOR IMPLANTABLE MEDICAL DEVICES | 2007-04-30 | Individual implantable medical device (IMD) follow-up support unit supports follow-up session for IMD. IMD includes a data input for receiving IMD data which includes an IMD identification code identifying an individual IMD and IMD type, and IMD irregularity messages including data identifying type/date of irregularity, a database for storing IMD, a follow-up history dataset that includes date(s) of latest follow-up session. Includes a data processor that generates a list of irregularities taking place after previous follow-up session, calculate an individual IMD irregularity index for each type of irregularity based on IMD irregularity messages for an individual IMD via IMD identification code, calculate for type of irregularity an average irregularity index over the same irregularity over all IMD irregularity messages associated to the same type of IMD, compare calculated individual IMD irregularity index (IIII) with the calculated average irregularity index (ARI) and generate an indication if IIII significantly differs from the ARI. | 1. An IMD follow-up support unit for supporting a follow-up session for an individual implantable medical device (IMD), said IMD follow-up support unit comprising: a data input for receiving IMD data, said IMD data including an IMD identification code identifying an individual IMD and an IMD type, and IMD irregularity messages including data identifying a type of irregularity and a date of an irregularity;a database for storing said IMD data received via said data input;a follow-up history dataset comprising at least a date of a latest follow-up session for said individual IMD;a data processor adapted to: generate a list of irregularities which took place after a previous follow-up session;calculate an individual IMD irregularity index for each said type of irregularity based on said IMD irregularity messages associated with said individual IMD by way of said IMD identification code to form a calculated individual IMD irregularity index;calculate for each said type of irregularity an average irregularity index over same irregularity over all said IMD irregularity messages associated to said IMD type to form a calculated average irregularity index;compare said calculated individual IMD irregularity index with said calculated average irregularity index; and,generate an indication if said individual IMD irregularity index significantly differs from said average irregularity index. 2. The IMD follow-up support unit of claim 1, wherein said data processor is further adapted to calculate a list of hints upon generating said indication, wherein said list of hints is calculated on a basis of said irregularity messages stored in said data base and wherein said list of hints comprises at least one of: static suggestions which are always associated with an irregularity leading to said indication, and/ordynamic suggestions, wherein said data processor is adapted to calculate for each said dynamic suggestion an IMD setting index of device settings currently configured to treat said irregularity and to compare said IMD setting index to an average index over all device settings of IMDs of this type. 3. The IMD follow-up support unit of claim 1, wherein said IMD follow-up support unit is at least remotely connected to an IMD follow-up workstation comprising: a display for displaying said hints calculated by said data processor;input means for entering of commands and/or data by a user;a telemetry link adapted for wireless transmission of data to an IMD;wherein said IMD follow-up support unit is further adapted to generate display messages that includes a hint generated andresponse buttons that allow a user to respond to said hints; and,wherein said follow-up support unit is further adapted to respond to actuation of a response button and to eventually initiate reprogramming of said IMD according to said hint. 4. The IMD follow-up support unit of claim 3, wherein said response buttons can be actuated by means of a computer mouse. 5. The IMD follow-up support ... | - | A61B005/00
US20080287833 | A1 | 2008-11-20 | Method of evaluating gastroparesis using an ingestible capsule | 2007-09-06 | A method of evaluating gastroparesis comprising the steps of providing an ingestible capsule having a pH sensor and a pressure sensor, having a subject ingest the capsule, recording pH measurements from the pH sensor as a function of time as the capsule moves through at least a portion of the gastrointestinal tract of the subject, recording pressure measurements from the pressure sensor as a function of time as the capsule moves through at least a portion of the gastrointestinal tract of the subject, determining the capsule's location at a first position in the gastrointestinal tract, deriving a pressure pattern as a function of time and the pressure measurements, providing a reference pressure pattern, and analyzing the subject's pressure pattern relative to the reference pressure pattern to evaluate the subject with respect to gastroparesis. The location may be the junction between the stomach and the small bowel of the subject. The pressure pattern may be the number of contractions relative to a baseline over a given time interval or area under the curve of pressure measurements over a given time interval. | 1. A method of evaluating gastroparesis comprising the steps of: providing an ingestible capsule having a pH sensor and a pressure sensor;having a subject ingest said capsule;recording pH measurements from said pH sensor as a function of time as said capsule moves through at least a portion of the gastrointestinal tract of said subject;recording pressure measurements from said pressure sensor as a function of time as said capsule moves through at least a portion of said gastrointestinal tract of said subject;determining said capsules location at a position in said gastrointestinal tract;deriving a pressure pattern as a function of time and said pressure measurements;providing a reference pressure pattern; andanalyzing said subject's pressure pattern relative to said reference pressure pattern to evaluate said subject with respect to gastroparesis. 2. The method set forth in claim 1, wherein said position is the junction between the stomach and the small bowel of said subject. 3. The method set forth in claim 1, wherein said pressure pattern is the number of contractions relative to a baseline for a given time interval. 4. The method set forth in claim 1, wherein said pressure pattern is the area under the curve of said pressure measurements for a given time interval. 5. The method set forth in claim 1, wherein said pressure pattern is amplitude. 6. The method set forth in claim 1, wherein said step of evaluating said subject comprises diagnosing gastroparesis. 7. The method set forth in claim 3, wherein said step of analyzing said pressure pattern comprises the step of determining if said subject's pressure pattern is significantly lower than said reference pressure pattern, wherein said reference pressure pattern is derived from a healthy control group. 8. The method set forth in claim 7, wherein said subject's pressure pattern is at least ten percent less than said reference pressure pattern. 9. The method set forth in claim 1, wherein said time is the time said capsule resides in the stomach of said subject. 10. The method set forth in claim 1, wherein said time extends from about one hour before said capsules is determined to be at said position to about the time said capsule is determined to be at said position. 11. The method set forth in claim 1, wherein said step of analyzing said subject's pressure pattern relative to said reference pressure pattern to evaluate said subject with respect to gastroparesis comprises the step of comparing said subject's pressure pattern to said reference pressure pattern for three consecutive twenty minute intervals before and after said capsule is determined to be at said position. 12. The method set forth in claim 1, and further comprising the steps of determining the transit time between a first location and a second location, and evaluating said transit time relative to a reference transit time. 13. The method set forth in claim 12, wherein said first location is the point at which said capsu... | - | A61B005/00
US20080255442 | A1 | 2008-10-16 | REGISTRATION SYSTEM AND METHOD | 2007-10-08 | Methods and systems for registering the position of a body part of a patient with a computer aided surgery system are described. A computer implemented method includes determining the position of a plurality of points of the body part, fitting the points to a representation of the body part, determining a closeness of fit for each of the points and displaying an indication of the closeness of fit. The system includes a tracking system for determining the position of an instrument engageable with a plurality of points of the body part and a control system. The control system is configured to determine the position of the plurality of points of the body part, fit the points to a representation of the body part, determine a closeness of fit for each of the points and display an indication of the closeness of fit for at least one of the points. | 1. A computer implemented method for registering the position of a body part of a patient with a computer aided surgery system, comprising: determining the position of a plurality of points of the body part;fitting the plurality of points to a representation of the body part;determining a closeness of fit for each of the plurality of points; anddisplaying an indication of the closeness of fit for at least one of the plurality of points. 2. A method as claimed in claim 1, wherein displaying an indication of the closeness of fit includes displaying the indication at the determined position of the point overlaid on an image of the body part at a fitted position of the image of the body part. 3. A method as claimed in claim 1 or 2, where the indication is a visual indication. 4. A method as claimed in claim 3, wherein the visual indication is colour. 5. A method as claim in claim 4, wherein the colour varies depending on the closeness of fit. 6. A method as claimed in any preceding claim, wherein an indication of the closeness of fit is displayed for each of the plurality of points. 7. A method as claimed in any preceding claim and further comprising: re-determining the position of at least a one of the plurality of points; andre-fitting those points of the plurality of points whose position has not been re-determined and those points whose position has been re-determined. 8. A method as claimed in any preceding claim, wherein the plurality of points includes a set of landmark points and a set of surface points. 9. A method as claimed in claim 8, wherein the positions of the landmark points are used to determine the direction and/or the size and/or the orientation of the body part. 10. A method as claimed in claim 8 or 9, wherein the positions of the surface points are used to fit the positions of the plurality of points to the representation of the body part. 11. A method as claimed in any preceding claim, wherein the representation of the body part is a captured image of the body part or derived from a captured image of the body part. 12. A method as claimed in any of claims 1 to 10, wherein the representation of the body part is a virtual model of the body part or derived from a virtual model of the body part. 13. A method as claimed in any preceding claim and further comprising assessing the spatial distribution of the closeness of fit of the plurality of points to identify at least a one of the plurality of points causing an unsuccessful fit. 14. A method as claimed in any preceding claim, wherein one of the plurality of points is a landmark point, and wherein the closeness of fit of the landmark point is determined, and if the closeness of fit of the landmark point exceeds a threshold, then displaying a visual indication to recapture the landmark point. 15. A method as claimed in claim 14, and further comprising: re-determining the position of the landmark point; andre-fitting using the landmark point whose position has been re-de... | - | A61B005/00
US20080262346 | A1 | 2008-10-23 | Method And Apparatus For Blood Vessel Parameter Determinations | 2006-03-03 | An apparatus and method for the determination of the flow, the coronary reserve and relative coronary reserve of a specific coronary artery. The apparatus and method employ a three-dimensional model (50) providing the volume of a segment of the artery at a plurality of points in time, and disclose various options (56) for determining the arterial flow through the artery segment during one or more heart beat cycles or parts thereof. The determination of the coronary reserve (62) and relative coronary reserve (64) follow from the volume (54) and the flow through the artery. Alternatively, the coronary reserve is determined directly from a velocity profile relating to one or more fixed artery segments and the three-dimensional model. | 1. A method for determining the arterial reserve of a subject having a blood flow, said blood flow having a velocity below or equal to a maximal velocity value, the method comprising the following steps: receiving an at least one first model, said model representing an at least one substantially fixed segment of an at least one artery of the subject at a plurality of points in time associated with and during an at least one part of an at least one first heart beat cycle, said artery segment having a proximal cross section and a distal cross section;injecting contrast agent into the at least one artery of the subject, said subject being in a non-hyperemic state, said contrast agent is injected at an injection area having a proximal cross section and having a distance from the at least one substantially fixed segment;determining an at least one first parameter from an at least one first angiogram representing the substantially fixed artery segment, said first angiogram taken from a first projection angle; injecting the subject with substance that simulates hyperemia;receiving an at least one second model, representing the at least one substantially fixed segment of the at least one artery of the subject at a plurality of points in time associated with and during an at least one part of an at least one second heart beat cycle;injecting the contrast agent to the at least one artery of the subject, said subject being in a hyperemic state, said contrast agent is injected at the injection area;determining an at least one second parameter from an at least one second angiogram representing the substantially fixed segment, said second angiogram taken from a second projection angle; anddetermining the arterial reserve as the ratio between one of any of the at least one first parameter or a combination thereof and any of the at least one second parameter or a combination thereof. 2. The method of claim 1 wherein determining the first parameter comprises the steps of: determining from the at least one first angiogram, arterial segment output values of the blood flow within the at least one fixed segment of the at least one artery being in a non-hyperemic state, at times corresponding to the plurality of points in time associated with the at least one part of the at least one first heart beat cycle;summing the arterial segment output values determined during the at least one part of the at least one first heart beat cycle, said arterial segment being in a non-hyperemic state; anddividing the summed arterial segment output values by the duration of the at least one part of the at least one first heart beat cycle, yielding an average non-hyperemic artery segment flow output. 3. The method of claim 1 wherein determining the second parameter comprises the steps of: determining from the at least one second angiogram, arterial segment output values of the blood flow within the at least one fixed segment of the at least one artery being in a hyperemic state, at tim... | - | A61B005/0275 A61B006/00
US20080275356 | A1 | 2008-11-06 | RESPIRATORY SENSING BELT USING PIEZO FILM | 2007-05-03 | A respiratory belt sensing device to aid in monitoring breathing patterns of a user, particularly in diagnosis and treatment of persons with breathing disorders. The device includes a PVDF film member having first and second metalized major surfaces and a perimeter containing a plurality of spaced apart alignment features. It also includes a first tabbed lead and a second tabbed lead respectively attached to the first and second major surfaces of the PVDF film member by an electrically conductive adhesive. The resulting assembly is affixed to the surface of a body encircling belt member. Signals are generated when the PVDF film on the belt member is stressed due to stretching as a user breathes in and out. The generated signals can thereafter be used in diagnosis and analysis. | 1. A respiratory sensing belt comprising: (a) a film member having piezoelectric properties and first and second metalized major surfaces wherein the surfaces are of symmetric shape and contain a plurality of spaced apart alignment features around the surface perimeter;(b) a first tabbed lead and a second tabbed lead respectively attached to the first and second metalized major surfaces of the film member by an electrically conductive adhesive;(c) a label for concealing the film member; and(d) a body encircling elastic belt member wherein the film member is affixed to the surface of the belt member such that stretching of the belt member produces an electrical signal between the first and second tabbed leads. 2. The respiratory sensing belt as in claim 1 wherein the film member is polyvinylidene fluoride. 3. The respiratory sensing belt as in claim 2 wherein the film member has an elongated rectangular portion. 4. The respiratory sensing belt as in claim 3 wherein the film member has an enlarged head at both ends. 5. The respiratory sensing belt as in claim 1 wherein the film member has a dumbbell shaped perimeter. 6. The respiratory sensing belt as in claim 5 wherein the body encircling belt member is adapted to fit around the abdomen of a user. 7. The respiratory sensing belt as in claim 5 wherein the body encircling belt member is adapted to fit around the chest of a user. 8. The respiratory sensing belt as in claim 1 wherein the respiratory sensing belt is disposable. 9. The respiratory sensing belt as in claim 1 wherein adhesive foam members are placed around the film member and tabbed leads. 10. The respiratory sensing belt as in claim 9 wherein a urethane member forms a seal around the outside of the foam members. 11. The respiratory sensing belt as in claim 10 wherein the label, the foam members, and urethane layers are stitched to the belt member. 12. The respiratory sensing belt of claim 11 wherein the respiratory sensing belt is reuseable. 13. A respiratory sensing belt comprising: (a) a film member having piezoelectric properties and first and second metalized major surfaces;(b) a first tabbed lead and a second tabbed lead respectively attached to the first and second major surfaces of the film member by an electrically conductive adhesive;(c) a pair of foam members sandwiching the film member and tabbed leads therebetween;(d) a urethane member surrounding the foam members in a moisture resistant configuration;(e) a label concealing the film member; and(f) a body encircling elastic belt member wherein the assembly elements (a)-(e) is affixed to the surface of the belt member. 14. The respiratory sensing belt of claim 13 wherein the film member, the foam members, and urethane coating have corresponding features in their shape for alignment. 15. The respiratory sensing belt of claim 14 wherein said assembly is stitched to the belt. 16. The respiratory sensing belt of claim 15 wherein the respiratory sensing belt is reusab... | - | A61B005/08
US20080249380 | A1 | 2008-10-09 | Protection Mechanism for Spectroscopic Analysis of Biological Tissue | 2005-04-19 | The present invention provides a protection mechanism for a spectroscopic analysis system being adapted to determine a property of a biological structure in a volume of interest of a patient. The spectroscopic system preferably makes use of high power radiation, and provides a protection mechanism for preventing an accidental exposure of light sensitive tissue of a body. The invention provides a variety of approaches to detect whether a measurement head of the spectroscopic system is in a measurement position. The measurement position of the measurement head can effectively be determined by making use of e.g. a pressure sensor, a sensor measuring the electric resistance of the skin of the patient, or even by optical means analyzing the intensity or the spatial structure of a monitoring beam providing a visual image of the volume of interest. | 1. A spectroscopic system (100) for determining of a property of a biological structure in a volume of interest (152) of a patient, the spectroscopic system comprising: a measurement head (120) for directing an excitation beam (116) into the volume of interest,at least one detector element (104; 160; 170, 172) for determining if the measurement head is in a measurement position,means for switching (108) the excitation beam in response of an output from the at least one detector element. 2. The spectroscopic system (100) according to claim 1, wherein the at least one detector element is a pressure sensor (104) being adapted to generate an output being indicative of a contact pressure between the measurement head (120) and the surface of the skin (150). 3. The spectroscopic system (100) according to claim 1, wherein the at least one detector element comprising two electrodes (162, 164), the at least one detector element being adapted to generate an electrical signal being indicative of the electric resistance between the two electrodes. 4. The spectroscopic system (100) according to claim 1, further comprising monitoring means for directing a monitoring beam (124) into the volume of interest (152) and for collecting monitoring return radiation (118) from the volume of interest, the monitoring return radiation being indicative of a visual image of the volume of interest. 5. The spectroscopic system (100) according to claim 4, wherein the detector element (170) being adapted to generate an intensity signal being indicative of the intensity of the monitoring return radiation (118). 6. The spectroscopic system (100) according to claim 4, wherein the detector element is an image analysis unit (172) being adapted to recognize the biological structure in the visual image of the volume of interest (152) and being further adapted to generate a detection signal for the means for switching (108) the excitation beam in response of recognizing the biological structure in the visual image. 7. The spectroscopic system (100) according to claim 1, wherein the means for switching (108) of the excitation beam are adapted to release the excitation beam (116) into the volume of interest (152) only when the contact pressure is within a predefined range. 8. The spectroscopic system (100) according to claim 3, wherein the means for switching (108) of the excitation beam are further adapted to release the excitation beam (116) into the volume of interest (152) only when the resistance between the electrodes is within a predefined range. 9. The spectroscopic system (100) according to claim 5, wherein the means for switching (108) of the excitation beam (116) are further adapted to release the excitation beam into the volume of interest (152) only when the intensity signal is within a predefined range. 10. The spectroscopic system (100) according to claim 5, wherein the means for switching (108) of the excitation beam (116) are further adapted to release the excitat... | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61B005/145
US20080319314 | A1 | 2008-12-25 | INJECTION CATHETER | 2008-08-26 | An injection catheter is provided. The catheter comprises a catheter body comprising a flexible tubing having proximal and distal ends and at least one lumen therethrough. A tip section comprising a flexible tubing is mounted at the distal end of the catheter body. The tip section has a needle passage extending therethough. The needle passage has a proximal region having a proximal diameter and a distal region having a distal diameter less than the proximal diameter. A needle control handle is provided at the proximal end of the catheter body. An injection needle extends through the tip section, catheter body, and needle control handle and has a proximal end attached to the needle control handle and a distal end within the needle passage. The injection needle is longitudinally slidable so that its distal end can extend beyond the distal end of the tip section upon suitable manipulation of the needle control handle. | 1. A steerable cardiac drug injection catheter comprising: a catheter body having an outer wall, proximal and distal ends, and at least one lumen extending therethrough;a control handle fixedly attached to the proximal end of the catheter body;a tip section comprising flexible tubing having proximal and distal ends and at least one lumen extending therethrough, the proximal end of the tip section being fixedly attached to the distal end of the catheter body;an injection needle which is straight and smooth over its entire length extending through the lumen in the catheter body and the lumen in the tip section, said needle being slidable from a first position in which the needle is withdrawn into the tip electrode to a second position in which the needle extends out of the distal end of the tip section;a needle control knob mounted on the control handle and connected to the proximal end of the injection needle for sliding the injection needle from the first position to the second position; and,a deflection control mounted on the control handle for deflecting the catheter upon manipulation of the deflection control, said deflection control being longitudinally movable with respect to the control handle for deflecting the tip section of the injection catheter. 2. An injection catheter according to claim 1, further comprising an electromagnetic mapping sensor disposed in the tip section for producing electrical signals indicative of the position of the electromagnetic mapping sensor relative to the position of a reference electrode sensor. 3. An injection catheter according to claim 2, further comprising a sensor cable electrically attached to the electromagnetic mapping sensor and extending through the lumen in the tip section, through the lumen in the catheter body and into the control handle, and the sensor cable is electrically attached to a circuit board situated within the control handle. 4. An injection catheter according to claim 1, wherein the control handle comprises a first member fixedly attached to the proximal end of the catheter body and a second member that is movable relative to the first member. 5. An injection catheter according to claim 4, wherein the deflection control comprises a puller wire having a proximal end and a distal end, the puller wire extending from the control handle, through the lumen in the catheter body and being fixedly secured within the tip section, the proximal end of the puller wire being fixedly secured to the second member of the control handle, whereby manipulation of the first member of the control handle relative to the second member of the control handle moves the puller wire relative to the catheter body resulting in deflection of the tip section. 6. An injection catheter as defined in claim 5, wherein the first member of the control handle is movable from a first position to a second position relative to the second member of the control handle, and the puller wire is fixedly secured to the second... | BIOSENSE WEBSTER, INC. | A61M025/092 A61B005/055
US20080243019 | A1 | 2008-10-02 | Respiratory function measuring equipment and storage medium | 2007-12-19 | A respiratory function measuring device comprises: a three-dimensional measuring unit that measures a chest movement and an abdomen movement of a breathing animal; a first measuring unit that measures a time T1 where a rate of volume decrease of the abdomen is maximized in an expiration; a second measuring unit that measures a time T2 where a rate of volume decrease of the chest is maximized in the expiration; and a respiratory time difference outputting unit that computes and outputs a value Tde corresponding to T2-T1. This allows measuring respiratory function to diagnose an obstructive pulmonary disease, a restrictive pulmonary disease, and the like in a natural state, for a subject of a breathing animal, even if the subject does not have a sense of self-awareness. | 1. A respiratory function measuring device comprising: a three-dimensional measuring means that measures a chest movement and an abdomen movement of a breathing animal;a first measuring means that measures a time T1 where a rate of volume decrease of the abdomen is maximized in an expiration;a second measuring means that measures a time T2 where a rate of volume decrease of the chest is maximized in the expiration; anda respiratory time difference outputting means that computes and outputs a value Tde corresponding to T2-T1. 2. The respiratory function measuring device according to claim 1, wherein the respiratory time difference outputting means computes Tde in terms of multiple expirations and computes and outputs a value Av(Tde) corresponding to an average value thereof. 3. A respiratory function measuring device comprising: a three-dimensional measuring means that measures a chest movement and an abdomen movement of a breathing animal;a third measuring means that measures a time T3 where a rate of volume increase of the abdomen is maximized in an inspiration;a fourth measuring means that measures a time T4 where a rate of volume increase of the chest is maximized in the inspiration; anda respiratory time difference outputting means that computes and outputs a value Tdi corresponding to T4-T3. 4. The respiratory function measuring device according to claim 3, wherein the respiratory time difference outputting means computes Tdi in terms of multiple inspirations and computes and outputs a value Av(Tdi) corresponding to an average value thereof. 5. A respiratory function measuring device comprising: a three-dimensional measuring means that measures a body movement of a breathing animal;a fifth measuring means that measures an inspiration time Ti of a respiration;a sixth measuring means that measures an expiration time Te of the respiration; anda respiratory ratio outputting means that computes and outputs a value R corresponding to Ti/Te. 6. The respiratory function measuring device according to claim 5, wherein the respiratory ratio outputting means measures R in terms of multiple respirations and computes and outputs a value Av(R) corresponding to an average value thereof. 7. A respiratory function measuring device comprising: a three-dimensional measuring means that measures a body movement of a breathing animal; anda respiratory minute volume outputting means that outputs a value corresponding to a respiratory minute volume. 8. A computer-readable storage medium having a program recorded thereon where the program makes a computer as the respiratory function measuring device according to claim 1. 9. A computer-readable storage medium having a program recorded thereon where the program makes a computer as the respiratory function measuring device according to claim 3. 10. A computer-readable storage medium having a program recorded thereon where the program makes a computer as the respiratory function measuring device according to cla... | KEIO University | A61B005/08
US20080306369 | A1 | 2008-12-11 | DETACHABLE LATCHING MECHANISM | 2007-06-06 | A detachable latching mechanism is disclosed herein. The latching mechanism for latching a first apparatus to a second apparatus comprises: a detachable housing configured to be attached to the first apparatus. The housing includes a first and second housing member forming a channel there between. The latching mechanism includes a first latch supported by the first housing member and biased to extend into the channel in a first direction and a second latch supported by the second housing member and biased to extend into the channel in a second direction opposite from the first direction. Further the latching mechanism is provided with a latching member having a first end configured to be attached to the second apparatus and a second end having a protrusion configured to push against the first and second latches to move them apart when the latching member is inserted into the channel. The movements of the latches are perpendicular to the movement of the latching member. | 1. A latching mechanism for latching a first apparatus to a second apparatus, comprising: a detachable housing configured to be attached to the first apparatus, the housing including first and second housing members forming a channel there between;a first latch supported by the first housing member and biased to extend into the channel in a first direction;a second latch supported by the second housing member and biased to extend into the channel in a second direction opposite from the first direction; anda latching member having a first end configured to be attached to the second apparatus and a second end having a protrusion configured to push against the first and second latches to move them apart when the latching member is inserted into the channel, the movements of the latches being perpendicular to the movement of the latching member. 2. A latching mechanism as in claim 1, wherein the first and second latches each has a head, and the protrusion of the latching member is configured to push against the heads of the first and second latches to overcome their respective biases to move them apart and allow the insertion of the latching member. 3. A latching mechanism as in claim 2, wherein the first and second latches are biased to re-extend back into the channel after the latching member has been fully inserted into the channel, thereby latching the latching member in the channel to prevent its removal. 4. A latching mechanism as in claim 1, further comprising a first spring and a second spring, wherein the first spring biases the first latch into the channel in the first direction and the second spring biases the second latch into the channel in the second direction. 5. A positioner latching assembly comprising: a detachable housing having a plurality of housing members configured to define a cavity, the housing configured to attach to a support;a plurality of latches supported by the housing members and biased to extend into the cavity; anda latching member having a first end configured to attach to a carriage assembly and a second end having a protrusion configured to engage between the detachable housing and the plurality of latches, while latching;wherein the latches are biased to move away from the cavity while latching to allow the insertion of the latching member into the cavity and to re-extend back into the cavity after the latching member has been fully inserted into the cavity. 6. A latching assembly as in claim 5, wherein each latch has a head, the heads of the latches being pushed apart from the cavity by inserting the protrusion into the cavity. 7. A latching assembly as in claim 6, wherein each head has a first tapering surface and a second tapering surface, the first tapering surface being located in the direction of movement of the latching member for latching and the second tapering surface is located opposite to the first tapering surface. 8. A latching assembly as in claim 7,wherein the first tapering surface is mor... | GENERAL ELECTRIC COMPANY | G05G005/06 A61B005/00
US07460899 | B2 | 2008-12-02 | Apparatus and method for monitoring heart rate variability | 2005-02-25 | A wrist-worn or arm band worn heart rate variability monitor is provided. Heart rate variability (“HRV”) refers to the variability of the time interval between heartbeats and is a reflection of an individual's current health status. Over time, an individual may use the results of HRV tests to monitor either improvement or deterioration of specific health issues. Thus, one use of the HRV test is as a medical motivator. When an individual has a poor HRV result, it is an indicator that they should consult their physician and make appropriate changes where applicable to improve their health. If an individual's HRV results deviate significantly from their normal HRV, they may be motivated to consult their physician. In addition, the inventive monitor is capable of monitoring the stages of sleep by changes in the heart rate variability and can record the sleep (or rest) sessions with the resulting data accessible by the user or other interested parties. Alternate embodiments of the invention allow assistance in the diagnosis and monitoring of various cardiovascular and sleep breathing disorders and/or conditions. Other embodiments allow communication with internal devices such as defibrillators or drug delivery mechanisms. Still other embodiments analyze HRV data to assist the user in avoiding sleep. | 1. A wrist or arm worn heart rate variability monitor, comprising: at least one sensor embedded in a membrane for detecting heart rate signals generated within a body when placed in contact with the body's wrist, wherein the sensor is capable of monitoring signals selected from the group consisting of optical, infra-red, pressure, ultrasonic and electrical signals;a circuit that conditions the electrical signals and converts the analog signals to digital signal data;a heart rate variability signal processor that monitors and analyzes the digital signal data, calculates parameters comprising the mean digital signal value and at least one standard deviation of the digital signal data monitored and establishes a threshold level based on the calculated parameters evaluated when the user is in an awake state and just prior to real-time heart rate variability data monitoring, wherein the calculated threshold level is neither predetermined nor preset and does reflect the user's actual heart rate and heart rate variability over the defined time interval, and wherein the processor monitors the digital signal data and recognizes the crossing of the established threshold level by the digital signal data and may be programmed to establish a new threshold level into a subsequent sleep state that reflects the body's actual heart rate and heart rate variability within the subsequent sleep state and then performs at least one HRV test using the new threshold level while the body is within the subsequent sleep state and then continues to monitor the body's HRV for additional threshold crossings;an alarm that may be activated to alert the user of abnormal conditions or events or the crossing of a specified threshold level by the heart rate variability; anda memory capable of storing the established threshold levels and real time digital signal data. 2. The heart rate variability monitor of claim 1, wherein the at least one sensor comprises at least one piezoelectric sensor. 3. The heart rate variability monitor of claim 1, further wherein the at least one sensor comprises at least one micro electromechanical sensor. 4. The heart rate variability monitor of claim 1, further comprising the at least one sensor being selected from the group consisting of electrode sensors, optical sensors, cMUT sensors and piezoelectric sensors. 5. The heart rate variability monitor of claim 1, further comprising: a transceiver for two-way communication with an internal device, wherein the alarm may be activated upon communication of a significant event to the transceiver by the internal device. 6. The heart rate variability monitor of claim 5, wherein the internal device communicates diagnostic data to the heart rate variability monitor. 7. The heart rate variability monitor of claim 5, further comprising communicating to the internal device that the significant event is permitted to occur, not permitted to occur, or permitted to occur but only after a period of time has passed... | Quiescent, Inc. | A61B005/04 A61B005/02 G08B023/00
US07438466 | B2 | 2008-10-21 | Calibrating a thermal scanning system | 2006-12-22 | A method for recalibrating a bilateral thermal scanner comprising a first thermal sensor and a second thermal sensor can comprise acquiring a first set of thermal scanning data with the thermal sensors in their normal positions; optionally analyzing the first set of thermal scanning data is analyzed to determine if the data has a one-sided bias; acquiring a second set of thermal scanning data with the positions of the thermal sensors reversed; determining a calibration from the first and second sets of thermal scanning data; and optionally correcting at least one of the first and/or second set of thermal scanning data using the calibration. | 1. A method for calibrating a bilateral thermal scanner, the method comprising: acquiring, with the bilateral thermal scanner comprising a first thermal sensor and a second thermal sensor, a first set of bilateral temperature data using the first thermal sensor and the second thermal sensor;acquiring a second set of bilateral temperature data, wherein the positions of the first thermal sensor and second thermal sensor are substantially interchanged from their position when acquiring the first set of bilateral temperature data; andcomparing the first and second sets of bilateral temperature data to determine at least one calibration value for the bilateral thermal scanner. 2. The method of claim 1, wherein the acquiring steps involve rolling the bilateral thermal scanner over a surface. 3. The method of claim 1, wherein the acquiring steps involve placing the first and second thermal sensors, which are operable to determine temperature, against a surface. 4. The method of claim 1, additionally comprising activating a component of the bilateral thermal scanner to initiate data acquisition. 5. The method of claim 1, wherein the acquiring steps involve placing the first and second thermal sensors to collect thermal data from opposite sides of a patient's spine. 6. The method of claim 1, wherein the acquiring steps involve sensing the intensity of radiation from a surface. 7. The method of claim 1, additionally comprising displaying at least some of the acquired first and second sets of bilateral temperature data. 8. The method of claim 1, further comprising determining whether the first set of bilateral temperature data has a one-sided bias. 9. The method of claim 8, wherein determining whether the first set of bilateral temperature data has a one-sided bias comprises calculating a difference between an average of temperature data from the first thermal sensor and an average of temperature data from the second thermal sensor. 10. The method of claim 8 additionally comprising prompting a user to perform a calibration test if the one-sided bias is greater than a predetermined value. 11. The method of claim 1, wherein one of acquiring a first set of bilateral temperature data and acquiring a second set of bilateral temperature data comprises scanning up a patient's spine, andthe other of acquiring a first set of bilateral temperature data and acquiring a second set of bilateral temperature data comprises scanning down a patient's spine. 12. The method of claim 1, wherein the at least one calibration value comprises at least one of a gain and an offset. 13. The method of claim 1, further comprising applying the at least one calibration value to at least one of the first and second sets bilateral temperature data. 14. The method of claim 1, further comprising applying the at least one calibration value to the output of at least one of the first and second thermal sensors. 15. The method of claim 1, further comprising saving the at least one ... | - | G01K015/00 G01K013/00 G01K003/00 G01J005/00 A61B005/00
US07507208 | B2 | 2009-03-24 | Method and apparatus for continuous pulse contour cardiac output | 2008-05-21 | A system and method are provided for sensing cardiac electrogram (EGM) signals and ventricular pressure signals and for using the sensed EGM and sensed pressure signals for estimating stroke volume (SV). A measure of cardiac output can be computed from the estimated SV and a heart rate determined from the EGM signals. The sensed ventricular pressure signal and the sensed EGM signal are used to derive landmark points such as an estimated pulmonary diastolic pressure, a mean pulmonary artery pressure, a peak right ventricular pressure (RVP), and various time intervals used in computing an area or a pulse contour integral. The pulse contour integral is used to estimate SV. The estimated pulmonary diastolic pressure, mean pulmonary artery pressure and CO computed from the estimated SV can be used to compute a pulmonary vascular resistance. | 1. A method, comprising: sensing a cardiac electrogram (EGM) signal of the heart;sensing a right ventricular pressure (RVP) signal;computing an estimated stroke volume using the sensed EGM signal and the sensed RVP signal;wherein computing an estimated stroke volume comprises computing a RVP pulse contour integral by: identifying an integration start time corresponding to an estimated time of pulmonic valve opening;identifying an integration end time corresponding to an estimated time of pulmonic valve closure;computing the area under the right ventricular pressure signal between the integration start time and the integration end time; andwherein computing the RVP pulse contour comprises computing a plurality of correction areas used to adjust the area under the right ventricular pressure signal between the integration start time and the integration end time; andone of displaying and storing in a memory structure at least one of: the RVP pulse contour, the estimated stroke volume, the right ventricular pressure signal, the area. 2. A method according to claim 1 wherein sensing the right ventricular pressure signal further comprises: detecting an R-wave event from the sensed EGM signal;setting an analysis window for a predetermined interval of time following the R-wave event detection; andacquiring and storing the RVP signal during the analysis window. 3. A method according to claim 1 wherein computing an estimated stroke volume further comprises: deriving an estimated pulmonary artery diastolic pressure from the RVP signal;deriving a peak right ventricular pressure amplitude from the RVP signal; andcomputing a difference between the estimated pulmonary artery diastolic pressure and the peak right ventricular pressure amplitude. 4. A method according to claim 1 wherein the estimated time of pulmonic valve opening is a time corresponding to a maximum rate of pressure rise in the right ventricular pressure signal. 5. A method according to claim 1 wherein the estimated time of pulmonic valve closure is a time corresponding to a right ventricular pressure amplitude occurring on a descending portion of the right ventricular pressure signal and equaling a right ventricular pressure amplitude at the integration start time. 6. A method according to claim 1 wherein computing a plurality of correction areas comprises: deriving an estimated pulmonary artery diastolic pressure using the right ventricular pressure signal and the EGM signal, andcomputing the area below the estimated pulmonary artery diastolic pressure and between the integration start time and the integration end time. 7. A method according to claim 1 wherein computing a plurality of correction areas comprises: deriving a mean pulmonary artery pressure using the RVP signal and the EGM signal;identifying a corrected integration end time corresponding to a time that a RVP signal amplitude occurring on a descending portion of the right ventricular pressure signal equals the estimated mean pu... | Medtronic, Inc. | A61B005/00
US07467006 | B2 | 2008-12-16 | Magnetic resonance imaging system and contrast-enhanced angiography | 2004-10-05 | An MRI system capable of acquiring a high-quality blood vessel image even during a measuring period other than at an injected contrast medium concentration peak by always picking up an image under optimum conditions while following up with an injected contrast medium concentration that changes every moment in the body. A flip angle is changed according to a contrast medium concentration b(t) as shown by a curve (102). During a period (Da) where a contrast medium concentration b(t) gradually increases, a flip angle (FA) increases according to a contrast medium concentration b(t); and during a period (Db) where a contrast medium concentration b(t) gradually decrease, a flip angle is gradually decreased. If a flip angle is controlled so as to follow up with a contrast medium concentration b(t) and reach an Ernst's angle at which a single intensity is at a maximum, a high-quality blood vessel image can be acquired even during a measuring period other than at an injected contrast medium concentration peak. | 1. A magnetic resonance imaging system comprising: static magnetic field generating means for applying a static magnetic field to an object to be examined;gradient magnetic field generating means for applying a gradient magnetic field:RF magnetic field transmitting means for irradiating, to nuclear spins within said object, RF magnetic field pulses to cause nuclear magnetic resonance of the nuclear spins;an echo signal receiving means for detecting an echo signal emitted by the nuclear magnetic resonance;pulse sequence control means for controlling a pulse sequence including at least a repetition time TR as an imaging parameter and executed to receive the echo signal;signal processing means for reconstructing an image of a blood vessel by using the echo signal detected by said echo signal receiving means; anddisplay means for displaying said blood vessel image,wherein, during execution of said pulse sequence, said pulse sequence control means changes a value of the at least said repetition time TR in said pulse sequence depending on a concentration profile of a contrast agent, such that the repetition time has a shorter value during a higher concentration period than a value during a lower concentration period, which has been injected into said object, in said blood vessel. 2. The magnetic resonance imaging system according to claim 1, wherein said signal processing means estimates the concentration of the contrast agent based on information regarding changes in the concentration of the contrast agent, the information being obtained in advance, andsaid pulse sequence control means changes the value of said repetition time and takes said blood vessel image in accordance with a predicted value of the concentration of the contrast agent. 3. The magnetic resonance imaging system according to claim 2, further comprising an input unit for receiving an input to instruct start of the imaging of said blood vessel image, wherein said pulse control means successively takes monitoring images including said blood vessel,said display means successively displays said monitoring images, andsaid pulse sequence control means switches over the imaging from said monitoring images to said blood vessel image in accordance with the start instruction. 4. The magnetic resonance imaging system according to claim 1, further comprising contrast agent injecting means, said contrast agent being injected by said contrast agent injecting means. 5. The magnetic resonance imaging system according to claim 1, wherein said pulse sequence control means sets a first period and a second period depending on the concentration of the contrast agent, and said imaging parameter has different values during said first period and said second period, respectively. 6. The magnetic resonance imaging system according to claim 5, wherein said pulse sequence control means selects at least two imaging parameters including said repetition times, and a different imaging parameter is selected for ... | Hitachi Medical Corporation | A61B005/05 A61B006/00 G01V003/00
US07444179 | B2 | 2008-10-28 | Devices, systems and methods for characterization of ventricular fibrillation and for treatment of ventricular fibrillation | 2004-04-29 | A method of determining a state of ventricular fibrillation, includes: measuring the rhythm of the heart during ventricular fibrillation for a period of time; creating a lagged phase space reconstruction of the measured rhythm; determining a first value related to the rate of change of the leading edge of the phase space reconstruction over the period of time; and determining the state of ventricular fibrillation by relating the first value to the state of ventricular fibrillation. A defibrillation system includes at least one processor in communication with a sensor to measure heart rhythm and an applicator to apply a defibrillation pulse. The processor is adapted to create a lagged phase space reconstruction of ventricular fibrillation heart rhythm and to determine a first value related to the rate of change of the leading edge of the phase space reconstruction over a period of time. | 1. A method of determining a state of ventricular fibrillation, comprising; measuring the rhythm of the heart during ventricular fibrillation for a period of time;creating a lagged phase space reconstruction of the measured ventricular fibrillation heart rhythm using a constant time lag;determining a first value related to the rate of change of the leading edge of the lagged phase space reconstruction over the period of time; anddetermining the state of ventricular fibrillation based upon the determined first value to determine a treatment. 2. The method of claim 1 wherein the first value is average angular velocity for the period or angular area for the period. 3. The method of claim 2 wherein the average angular velocity for the period of time is the rate of rotation of the leading edge of the lagged phase space reconstruction about the center of mass of points at an average rate. 4. The method of claim 1 wherein the lagged phase space reconstruction is a two-dimensional reconstruction. 5. The method of claim 1 wherein the lagged phase space reconstruction is a three-dimensional reconstruction. 6. The method of claim 1 further comprising: determining a second value related to the fractal self-similarity dimension of the ventricular fibrillation heart rhythm for the period of time, the step of determining the state of fibrillation including the step of relating at least one of the first value and the second value to the state of fibrillation. 7. The method of claim 6 wherein the second value is a scaling exponent. 8. The method of claim 7 wherein the first value is average angular velocity. 9. The method of claim 8 wherein the determined state of ventricular fibrillation is associated with a probability of success of a mode of treatment of ventricular fibrillation. 10. The method of claim 9 wherein the mode of treatment is defibrillation shock. 11. The method of claim 10 wherein the probability of success of the defibrillation shock is associated with the scaling exponent and the angular velocity. 12. A method of determining a treatment for a patient experiencing ventricular fibrillation, comprising: measuring the rhythm of the heart during ventricular fibrillation for a period of time;creating a lagged phase space reconstruction of the measured ventricular fibrillation heart rhythm using a constant time lag;determining a first value related to the rate of change of the leading edge of the lagged phase space reconstruction over the period of time to determine the treatment; andrelating the first value to a treatment for the patient. 13. The method of claim 12 wherein the first value is average angular velocity for the period or angular area for the period. 14. The method of claim 13 wherein the average angular velocity is the rate of rotation of the leading edge of the lagged phase space reconstruction about the center of mass of points at an average rate. 15. The method of claim 12 wherein the lagged phase space reconstruction is ... | University of Pittsburgh-Of the Commonwealth System of Higher Education | A61B005/046
US07444866 | B2 | 2008-11-04 | System for rehabilitation training using muscle force data measured of a person both before impaired and when impaired | 2007-08-14 | A system for rehabilitation training is able to specify the muscle of a person in need of rehabilitation, and also to help the person to set a proper goal of training. To this end, the system includes a trainer device for measuring living data of the person and for training him or her, and a data server for keeping the living body data of the person measured by the trainer device. The system further includes a menu formulator. When the person is to undergo the process of rehabilitation, the menu formulator compares his or her living body data measured by the trainer device to his or her living body data measured when he or she was not impaired and kept in the data server to find out a difference therebetween. A menu for rehabilitation is formed based on the difference, so that rehabilitation effects will be enhanced. | 1. A system for rehabilitation training comprising: a trainer device for measuring living data of a person and for training the person;a data server for storing living body data of the person measured by said trainer device; anda menu formulator for forming a menu for rehabilitation of the person,said trainer device including a bi-articular arm device having a structure simulating an upper limb or a lower limb of a human being,said bi-articular arm device comprising at least one link on which the upper limb or lower limb of the human being is secured when the trainer device is used, and multiple actuators for generating a driving force for rotating said at least one link,said menu formulator comparing the living body data measured of the person, when undergoing a process of rehabilitation, by said trainer device to the living body data which was measured of the person, when not impaired, and is stored in said data server, to find out a difference between both of the living body data, thereby forming the menu based on the difference,said bi-articular arm device having the actuators generating forces at the distal end of said at least one link in multiple directions. 2. The system in accordance with claim 1, wherein the living body data is of force of a muscle, said menu formulator forming the menu with the muscle of which muscular force is weakened specified. 3. The system in accordance with claim 1, wherein said bi-articular arm device measuring muscular force of the upper limb or the lower limb of the person. 4. A system for rehabilitation training comprising: trainer means for measuring living data of a person and for training the person;data keeping means for keeping living body data of the person measured by said trainer means; andmenu forming means for forming a menu for rehabilitation of the person,said trainer device including a bi-articular arm device having a structure simulating an upper limb or a lower limb of a human being,said bi-articular arm device comprising at least one link on which the upper limb or lower limb of the human being is secured when the trainer device is used, and multiple actuators for generating a driving force for rotating said at least one link,said menu forming means comparing the living body data measured of the person, when undergoing a process of rehabilitation, by said trainer means with the living body data which was measured of the person, when not impaired, and is stored in said data keeping means, to find out a difference between both of the living body data, thereby forming the menu based on the difference,said bi-articular arm device having the actuators generating forces at the distal end of said at least one link in multiple directions. 5. The system in accordance with claim 1, wherein the bi-articular arm device comprises a first link and a second link respectively corresponding to a lower arm and an upper arm, where a root end of the second link is rotatably connected to a base and a root end ... | Oki Electric Industry Co., Ltd. | A61B005/22
US07672718 | B2 | 2010-03-02 | Thoracic impedance detection with blood resistivity compensation | 2008-06-16 | This document discusses, among other things, a cardiac rhythm management device or other implantable medical device that uses thoracic impedance to determine how much fluid is present in the thorax, such as for detecting or predicting congestive heart failure, pulmonary edema, pleural effusion, hypotension, or the like. The thoracic fluid amount determined from the thoracic impedance is compensated for changes in blood resistivity, which may result from changes in hematocrit level or other factors. The blood-resistivity-compensated thoracic fluid amount can be stored in the device or transmitted to an external device for storage or display. The blood-resistivity-compensated thoracic fluid amount can also be used to adjust a cardiac pacing, cardiac resynchronization, or other cardiac rhythm management or other therapy to the patient. This document also discusses applications of the devices and methods for predicting or indicating anemia. | 1. A method comprising: detecting a thoracic impedance signal of a thorax of a subject;detecting a blood resistivity signal by sensing between at least two intravascular or intracardiac electrodes disposed in the same chamber of the heart or the same blood vessel; andusing a signal processing circuit, extracting physiological information from the thoracic impedance signal, the extracting physiological information including using the blood resistivity signal to reduce or eliminate an effect of a change in blood resistivity on the physiological information. 2. The method of claim 1, wherein extracting physiological information includes extracting cardiac stroke information from the thoracic impedance signal. 3. The method of claim 1, wherein extracting physiological information includes extracting respiration information from the thoracic impedance signal. 4. The method of claim 1, wherein extracting physiological information includes extracting an abdominal fluid indication from the thoracic impedance signal. 5. The method of claim 1, comprising filtering the thoracic impedance signal to obtain a near-DC thoracic impedance signal, wherein the physiological information is generated at least in part by using the near-DC thoracic impedance signal. 6. The method of claim 5, wherein filtering the thoracic impedance signal includes accounting for posture. 7. The method of claim 1, wherein detecting the blood resistivity includes using first and second intravascular electrodes. 8. The method of claim 1, wherein detecting the blood resistivity signal includes measuring the blood resistivity at like times during at least one of a cardiac cycle and a respiratory cycle. 9. The method of claim 1, comprising: storing a history of the physiological information; anddetermining whether a physiological trend is present using the history of physiological information. 10. The method of claim 9, comprising determining whether heart failure decompensation is present using a physiological trend determination. 11. The method of claim 1, comprising transmitting the physiological information from an implantable medical device to an external device. 12. The method of claim 11, comprising displaying the physiological information on the external device. 13. A system comprising: an implantable medical device including: a thoracic impedance measurement circuit configured to provide a thoracic impedance signal;a blood resistivity measurement circuit configured to provide a blood resistivity signal sensed between at least two intravascular or intracardiac electrodes disposed in the same chamber of the heart or the same blood vessel; anda controller coupled to the blood resistivity measurement circuit, the controller configured to extract physiological information from the thoracic impedance signal, the controller configured to use the blood resistivity signal to reduce or eliminate an effect of a change in blood resistivity on the physiological information. 14. The ... | Cardiac Pacemakers, Inc. | A61B005/00
US07447536 | B2 | 2008-11-04 | System and method for measurement of local lung function using electron beam CT | 2003-10-09 | Certain embodiments include a method and system for measuring and imaging local lung function. The method includes triggering an image scan of at least a lung cross-section of a patient, scanning the lung cross-section during at least one of an inspiration and an expiration of air by the patient to obtain lung image data and measuring a lung function during at least one of an inspiration and an expiration of air by the patient to obtain lung function data. The method may also include performing a preview scan of the patient to identify the lung cross-section for imaging. Additionally, the lung image data and lung function data may be combined for use in diagnosis of the patient. The lung image data and lung function data may also be output. The method may further include processing the lung function data to generate a plot of lung attenuation versus time. | 1. A method for lung imaging, said method comprising: triggering a series of image scans, each of a lung cross-section of a patient;scanning the plurality of lung cross-sections during at least one of an inspiration and an expiration of air by the patient to obtain lung image data;measuring a local lung function during at least one of an inspiration and an expiration of air by the patient to obtain local lung function data;reconstructing local lung images from the plurality of lung cross-section scans to provide cross-sectional images of the lung; andcombining said lung image data from said local lung images and said local lung function data for use in diagnosis of the patient based on an area of a lung portion at different points in time. 2. The method of claim 1, further comprising performing a preview scan of the patient to identify the lung cross-section for imaging. 3. The method of claim 1, further comprising outputting said lung image data and said local lung function data. 4. The method of claim 1, further comprising processing said local lung function data to generate a plot of lung attenuation versus time or area versus time. 5. The method of claim 1, wherein said triggering step comprises triggering said image scan based on air flow in the patient's lung. 6. The method of claim 1, wherein said scanning step dynamically obtains said lung image data as at least one of air flow and air capacity in the lung are changing. 7. The method of claim 1, wherein said scanning step obtains said lung image data with at least one of air flow and air capacity in the lung held constant. 8. The method of claim 1, wherein said scanning step obtains lung image data during a sweep of at most 33 ms. 9. The method of claim 1, wherein said scanning step obtains lung image data during a sweep of at most 50 ms. 10. A diagnostic imaging system, said system comprising: a trigger triggering an imaging scan of said patient based on at least one of air flow in said lung, volume of said lung, movement of said lung, and density of said lung;an energy source generating an electron beam to strike a target;a target producing radiation in response to said electron beam, said radiation irradiating at least a portion of said lung of said patient;a detector receiving said radiation after said radiation irradiates said at least a portion of said lung of said patient during said imaging scan;a data acquisition system obtaining local lung image data from said detector based on said radiation; anda reconstruction system combining local lung function data and said lung image data for use in diagnosis of said patient,wherein said system scans a first portion of said lung at a first location and said reconstruction system reconstructs a first image of the lung at the first location, wherein said system scans a second portion of said lung at a second location and said reconstruction system reconstructs a second image of the lung at the second location such that extracted lo... | G.E. Medical Systems Global Technology Company, LLC | A61B005/05
US07459145 | B2 | 2008-12-02 | Multifunctional magnetic nanoparticle probes for intracellular molecular imaging and monitoring | 2003-10-27 | The present invention provides multifunctional magnetic nanoparticle probe compositions for molecular imaging and monitoring, comprising a nucleic acid or polypeptide probe, a delivery ligand, and a magnetic nanoparticle having a biocompatible coating thereon. The probe compositions may further comprise a fluorescent or luminescent resonance energy transfer moiety. Also provided are compositions comprising two or more such multifunctional magnetic nanoparticle probes for molecular imaging or monitoring. In particular, the nucleic acid or polypeptide probes bind to a target and generate an interaction observable with magnetic resonance imaging (MRI) or optical imaging. The invention thereby provides detectable signals for rapid, specific, and sensitive detection of nucleic acids, polypeptides, and interactions thereof in vivo. | 1. A nanoparticle probe composition for facilitating molecular imaging or monitoring, comprising: a detectable moiety comprising a magnetic nanoparticle having a biocompatible coating thereon;a targeting probe attached to the biocompatible coating, wherein the targeting probe is selected from the group consisting of a nucleic acid probe, an antibody, an antibody fragment, and an aptamer; andan intracellular delivery ligand attached to the biocompatible coating. 2. The composition of claim 1, wherein the targeting probe is a nucleic acid probe. 3. The composition of claim 2, wherein the nucleic acid probe hybridizes to a nucleic acid target sequence on a subject nucleic acid and forms a stem-loop structure when not bound to the nucleic acid target sequence. 4. The composition of claim 2, wherein the nucleic acid probe comprises a modification of the nucleic acid backbone for the increased stability of the nucleic acid as compared to a naturally occurring nucleic acid. 5. The composition of claim 1, wherein the targeting probe is an antibody or fragment thereof. 6. The composition of claim 1, wherein the targeting probe is an aptamer. 7. The composition of claim 1, wherein the delivery ligand comprises a protein transduction peptide selected from the group consisting of HIV-1 TAT, HSV VP22, and ANTP. 8. The composition of claim 1, wherein the intracellular delivery ligand is selected from the group consisting of the W/R peptide (SEQ ID NO:4), NLS* peptide (SEQ ID NO:5), the AlkCWK18 peptide (SEQ ID NO:6), DiCWK18 peptide (SEQ ID NO:7), the transportan peptide (SEQ ID NO:8), the DipaLytic peptide (SEQ ID NO:9), the K16RGD peptide (SEQ ID NO:10), the P1 peptide (SEQ ID NO:11), the P2 peptide (SEQ ID NO:12), the P3 peptide (SEQ ID NO:13), the P3a peptide (SEQ ID NO:14), the P9.3 peptide (SEQ ID NO:15), the Plae peptide (SEQ ID NO:16), the Kplae peptide (SEQ ID NO:17), the cKplae peptide (SEQ ID NO:18), the MGP peptide (SEQ ID NO:19), the HA2 peptide (SEQ ID NO:20), the LARL46 peptide (SEQ ID NO:21), the Hel-11-7 peptide (SEQ ID NO:22), the KK peptide (SEQ ID NO:23), the KWK peptide (SEQ ID NO:24), the RWR peptide (SEQ ID NO:25), and the Loligomer peptide (SEQ ID NO:26). 9. The composition of claim 1, wherein the intracellular delivery ligand facilitates receptor-mediated endocytosis of the composition. 10. The composition of claim 1, wherein the intracellular delivery ligand facilitates entry into a cell by permeabilizing the cell membrane. 11. The composition of claim 1, wherein the magnetic nanoparticle comprises a metal detectable by use of an MRI instrument that is selected from the group consisting of selected from the group consisting of iron, cobalt, zinc, cadmium, nickel, gadolinium, chromium, copper, manganese, terbium, europium, gold, silver, platinum, and alloys thereof. 12. The composition of claim 11, wherein the magnetic nanoparticle is selected from the group consisting of monocrystalline iron oxide nanoparticle (MION), chela... | Georgia Tech Research Corporation Emory University | A61B005/055
US07441343 | B2 | 2008-10-28 | Saddle fit system and method | 2007-08-14 | In an embodiment, a bicycle saddle fit system comprises first and second supports, a substrate coupled to the first and second supports, and a measurement gauge coupled to the substrate. The first and second supports are each configured to receive sit bones of an adult human. At least one of the first and second supports is selectively moveable to locate the sit bones and to define a sit bone width. The substrate is configured to facilitate the selective movability of the at least one of the first and second supports. The measurement gauge is configured to visually indicate the sit bone width. In another embodiment, a method of fitting a bicycle saddle to a human rider comprises locating the first and second sit bones of the rider, reading a sit bone width from the measurement gauge, and assessing a saddle characteristic based on the rider's sit-bone width. | 1. A bicycle saddle fit system comprising: a substrate of sufficient size to span the sit bones of an adult human;a support being coupled to the substrate and being fixed relative to the substrate for receiving at least one of the sit bones;a measurement gauge being positioned relative to the substrate for enabling a sit-bone width of the adult human to be measured;first and second indicators coupled to the substrate and positioned adjacent the measurement gauge, at least one of the first and second indicators being selectively moveable along the substrate such that the measurement gauge and the substrate are moveable relative to at least one of the first and second indicators to locate the sit bones and to define a sit bone width. 2. The bicycle saddle fit system of claim 1, wherein the first and second indicators further comprise respective first and second supports configured to receive a respective sit bone. 3. The bicycle saddle fit system of claim 2, wherein the first and second supports are fitting cups. 4. The bicycle saddle fit system of claim 3, wherein the substrate is a substantially flat layer extending beyond a periphery of each of the first and second supports. 5. The bicycle saddle fit system of claim 2, wherein the first and second supports are formed separately from the substrate. 6. The bicycle saddle fit system of claim 1, wherein the first and second indicators are each selectively moveable relative to the substrate. 7. The bicycle saddle fit system of claim 1, wherein the at least one of the first and second indicators is linearly slidable along the substrate. 8. The bicycle saddle fit system of claim 1, wherein the measurement gauge extends along the substrate. 9. The bicycle saddle fit system of claim 8, wherein the measurement gauge comprises a scale that is coupled to the substrate. 10. The bicycle saddle fit system of claim 9, wherein the scale extends at least intermediate the first and second indicators to provide a visual indication of the sit bone width. 11. The bicycle saddle fit system of claim 1, wherein the measurement gauge is fixed to the substrate. 12. The bicycle saddle fit system of claim 1, wherein the substrate comprises a slot extending longitudinally along the substrate. 13. The bicycle saddle fit system of claim 12, wherein at least one of the first and second indicators is selectively moveable along the slot of the substrate. 14. The bicycle saddle fit system of claim 12, wherein the slot of the substrate is a single, continuous slot. 15. The bicycle saddle fit system of claim 1, wherein the support is a fitting cup. 16. The bicycle saddle fit system of claim 1, wherein the support is a first support that receives a first sit bone, and wherein the system comprises a second support for receiving a second sit bone. 17. The bicycle saddle fit system of claim 16, wherein the first support comprises a first fitting cup and the second support comprises a second fitting cup. 18. The bicycle ... | Specialized Bicycle Components, Inc. | A61B005/103
US07715906 | B2 | 2010-05-11 | Method and apparatus for detecting noise in an implantable medical device | 2007-06-04 | An implantable medical device and associated method monitor a physiological signal for sensing physiological events and detecting a physiological condition in response to the sensed physiological events. The device senses a first event from the physiological signal, senses a noise signal in the physiological signal and senses a next event from the physiological signal wherein the first event and the next event define a signal interval. The signal interval is declared as a noisy interval in response to the sensed noise signal. | 1. An implantable medical device, comprising: sensing circuitry for receiving a cardiac signal, sensing cardiac depolarization signals and measuring depolarization intervals occurring between the sensed cardiac depolarization signals;an arrhythmia detection module for detecting arrhythmia events from the measured depolarization intervals; anda noise detection module for sensing a noise signal in the cardiac electrical signal and declaring a depolarization interval defined by a first sensed cardiac depolarization signal and a next sensed cardiac depolarization signal as a noisy interval in response to the noise signal being sensed during the depolarization interval;wherein the detection circuitry configured to inhibit detecting an arrhythmia event in response to the noisy depolarization interval. 2. The device of claim 1 wherein sensing the noise signal comprises sensing one or more signals exceeding a sensing threshold within a predetermined time interval of the first sensed cardiac depolarization signal. 3. The device of claim 2 further comprising threshold setting circuitry for setting the sensing threshold as an automatically adjusted threshold comprising a first threshold level and a second threshold level, the first and second threshold levels separated by a first linear decay rate and the second threshold level followed by a second linear decay rate different than the first linear decay rate. 4. The device of claim 1 wherein sensing the noise signal comprises sensing an over range signal. 5. The device of claim 4 wherein the sensing circuitry comprises an input amplifier and an analog-to-digital converter and the over range signal is a signal causing one of an input amplifier saturation, an analog-to-digital converter amplitude over range, and an analog-to-digital converter signal differential over range. 6. The device of claim 5 further comprising means for clearing the over range signal by shorting a differential feedback of the input amplifier to restore a zero input state of the input amplifier in response to the over range signal. 7. The device of claim 1 further comprising means for storing time markers of sensed signals and corresponding signal labels wherein the next sensed cardiac depolarization signal is assigned a unique label corresponding to a noisy interval. 8. The device of claim 1 wherein the arrhythmia detection module comprises an arrhythmia event interval counter that is incremented in response to a measured cardiac depolarization interval falling into a detection interval range, and wherein the arrhythmia detection module inhibits detecting an arrhythmia event by not incrementing the arrhythmia event interval counter in response to the noisy interval. 9. The device of claim 8 wherein the arrhythmia event interval counter corresponds to one of an arrhythmia onset detection interval counter and an arrhythmia termination detection interval counter. 10. The device of claim 8 wherein not incrementing the arrhythmia e... | Medtronic, Inc. | A61B005/04
US07431700 | B2 | 2008-10-07 | Body movement and respiration monitor | 2001-12-06 | A monitor which can detect respiration of a sleeping person without being affected by the attitude of the sleeping person or the indoor illumination light and can easily evaluate detected respiration quantitatively through image measurement. The monitor comprises means (1) for projecting a specified illumination pattern, means (5) for picking up the image of projected light continuously, means (7) for calculating inter-frame moving amount of the illumination pattern from the image of two frames acquired by the image pickup means at different times, means (8) for generating a moving amount waveform data comprising inter-frame moving amounts arranged in time series, and means (9) for detecting the movement of an object from the moving amount waveform data. | 1. A monitor comprising: lighting pattern projecting means for projecting a specified lighting pattern;image pickup means, whose optical axis is different from that of the lighting pattern projecting means, for picking up light of a projected wavelength continuously;moving amount calculating means for calculating an inter-frame moving amount of the lighting pattern from two frames of images acquired by the image pickup means at different times;moving amount waveform generating means for generating moving amount waveform data comprising inter-frame moving amounts arranged in time series; andwaveform detecting means for detecting a movement of an object to be picked up from the moving amount waveform data;wherein the lighting pattern projecting means projects the lighting pattern that has a plurality of lighting spots or slit rays; andthe moving amount calculating means calculates an inter-frame moving amount of each lighting spot in an axial direction connecting the lighting pattern projecting means and the image pickup means while providing a positive or negative sign in accordance with a moving direction, and calculates a sum of moving amounts of pixels corresponding to the plurality of lighting spots or slit rays. 2. The monitor as set forth in claim 1, wherein the waveform detecting means detects a body movement and respiratory movement of a sleeping person, the monitor further comprising: safety deciding means for deciding the safety of the sleeping person from the body movement and respiratory movement of the sleeping person; andsignal outputting means for, when the safety deciding means decides that the sleeping person is in a dangerous state, outputting a signal indicating it. 3. The monitor as set forth in claim 2, wherein the moving amount calculating means calculates an inter-frame moving amount in a partial region of the lighting pattern in an axial direction connecting the lighting pattern projecting means and the image pickup means, and determines a region in an image where an inter-frame moving amount in the partial region appears much greater than a predetermined amount or where an inter-frame movement in the partial region appears frequently beyond a predetermined amount as a region where the sleeping person exists, and the safety deciding means decides that the sleeping person is in a dangerous state when the region where the sleeping person exists moves in a shorter time than a predetermined time and moves frequently beyond a predetermined amount. 4. The monitor as set forth in claim 2, wherein the moving amount calculating means calculates an inter-frame moving amount in a partial region of the lighting pattern in an axial direction connecting the lighting pattern projecting means and the image pickup means, and determines a region in an image where an inter-frame moving amount in the partial region appears much greater than a predetermined amount or where an inter-frame movement in the partial region appears more frequently... | KEIO University Sumitomo Osaka Cement Co., Ltd. | A61B005/08 A61B005/103 A61B005/117
US07470238 | B2 | 2008-12-30 | Unit for piercing, and piercing device | 2003-07-01 | A lancing unit (U1) includes a lancing member (2), an auxiliary part (3) which is separate from the lancing member (2), and a supporter (1) detachably supporting these. Preferably, the lancing unit (U1) further includes a cap (29) which covers a needle (21) of the lancing member (2) and which is detachable from the lancing member (2), and the lancing member (2) is supported by the supporter (1) via the cap (29). | 1. A lancet supplying unit comprising a case having a closed bottom and a closable open top, a lancing member housed in the case, a sensor housed in the case and provided separately from the lancing member, and a supporter detachably supporting each of the lancing member and the sensor in the housing, wherein the case is temporarily attached to a lancing apparatus for supplying the lancing member and the sensor to the lancing apparatus,wherein the lancing member and the sensor remain attached to the lancing apparatus even after the case is detached from the lancing apparatus,wherein the lancing member includes a needle,wherein the lancet supplying unit further comprises a cap for covering the needle, the cap being detachable from the lancing member,wherein the lancing member is supported by the supporter via the cap, andwherein the cap is integrally formed on the supporter. 2. The lancet supplying unit according to claim 1, wherein the sensor is held by a sensor holder, and wherein, when lancing of a skin is performed by utilizing the lancing member, the sensor holder engages the lancing member to control lancing depth in the skin. 3. The lancet supplying unit according to claim 1, wherein the lancing member includes a body holding the needle, and wherein the cap is integrally formed on the body. 4. The lancet supplying unit according to claim 3, wherein a boundary portion between the cap and the body has a structure which causes a stress to be concentrated on the boundary portion more than on other portions of the cap and the body. 5. The lancet supplying unit according to claim 4, wherein the boundary portion has a constricted configuration. 6. The lancet supplying unit according to claim 1, wherein the supporter includes a portion for fitting to a part of the cap to hold the cap in a standing posture. 7. The lancet supplying unit according to claim 1, wherein the case includes a tubular portion for accommodating the cap together with the lancing member and the sensor. 8. The lancet supplying unit according to claim 1, further comprising a lid for closing the open top the case. 9. The lancet supplying unit according to claim 1, wherein a direction in which the sensor is detachable from the supporter corresponds to a direction in which the cap is detachable from the lancing member. 10. The lancet supplying unit according to claim 1, wherein the sensor is detachably supported by the cap. 11. A lancet supplying unit comprising a case having a closed bottom and a closable open top, a lancing member housed in the case, a sensor housed in the case and provided separately from the lancing member, and a supporter detachably supporting each of the lancing member and the sensor in the housing, wherein the case is temporarily attached to a lancing apparatus for supplying the lancing member and the sensor to the lancing apparatus,wherein the lancing member and the sensor remain attached to the lancing apparatus even after the case is detached fro... | Arkray, Inc. | A61B005/00 A61B017/32
US08340744 | B2 | 2012-12-25 | System, apparatus, and method for reproducible and comparable flow acquisitions | 2006-12-08 | A system, apparatus and method are provided for measuring and removing the influence of pulsatility on contrast agent flow in a region of interest of a vascular system of a patient. Once the change of blood speed over the cardiac cycle is known (pulsatility), this influence is removed from acquired image sequence for outcome control such that “quasi-stationary”, regular flow acquisition is passed on to subsequent visualization and analysis processes. A contrast agent injector is also provided that simultaneously measures and uses ECG to inject a known contrast agent at a fixed point over the cardiac cycle or such that a known amount of contrast agent will arrive at a known time at a region of interest in the vasculature of a patient, thus controlling one of the main unwanted variables in an acquisition of blood flow sequences. | 1. A method for visualization and analysis of an acquired time intensity sequence of images of a blood flow of the advance of a contrast agent through a region of interest in a vascular system of a patient, comprising the steps of: obtaining information on the pulsatile nature of the blood flow;modifying the acquired time intensity sequence of images using the obtained information to remove therefrom an influence of pulsatile flow, the modifying producing a modified sequence of images;determining a temporal offset between the acquired time intensity sequence of images and the modified sequence of images such that a concentration of contrast agent at a pre-selected point in time has maximal similarity between the two sequences of images;offsetting the modified sequence of images by the determined temporal offset; andproviding both the acquired time intensity sequence of images and the offset sequence of images to an image analysis system. 2. The method of claim 1, wherein where the information on the pulsatile nature of the blood flow is obtained by performing a method comprising: acquiring, with an imaging system, the time intensity sequence of images, wherein the imaging system is selected from the group consisting of interventional x-ray, ultrasound, and repeated acquisitions using angiography with computed tomography or magnetic resonance imaging, the acquiring comprising the substeps of: varying a concentration of the contrast agent over time at an injection site thereof;imaging the contrast agent downstream of the injection site at the region of interest to obtain an observed sequence thereof; andsimultaneously obtaining an electrocardiogram (ECG) signal from the patient;identifying a phase of a cardiac cycle of the patient from the ECG;obtaining a predicted sequence of the concentration of the contrast agent at the injection site from a contrast agent prediction model, wherein the contrast agent prediction model is based on a predetermined template as a function of the identified phase and the function further based on model parameters representing the transport of the mix of contrast agent and blood downstream of the injection site by each of convection, dispersion due to varying speeds over a vessel cross section, diffusion, and variable transport speed due to pulsatility;matching the observed sequence with the predicted sequence by adjusting the model parameters such that the error between the observed sequence and the predicted sequence is minimized; andextracting a degree of pulsatility from the adjusted model parameters. 3. The method of claim 2, wherein the template is a simplified template comprising a high average flow speed in systole and a low average flow speed in diastole. 4. The method of claim 3, wherein the acquired sequence of images comprises two sequences of images, the two sequences of images comprising a sequence of images acquired before an interventional step and a sequence of images acquired after an interventiona... | Koninklijke Philips Electronics N.V. | A61B006/00 G06K009/00 G06G007/58
US08062287 | B2 | 2011-11-22 | Device for controlled release of chemical molecules | 2006-12-12 | The invention relates to a device comprising an ingestible capsule which measures at least one body parameter and starts a drug release program upon predefined changes in the body parameter. | 1. A medical device for drug administration comprising: a) a device which measures at least one body parameter of a patient for at least one measuring cycle and at least one monitoring cycle,b) a central control unit programmed to generate a normalized curve for each of the at least one measured body parameter of the patient from measured data obtained in the at least one measuring cycle and to adjust the measured data obtained in the at least one monitoring cycle to normalize the measured data based on the normalized curve to generate normalized monitoring data, andc) an ingestible capsule which starts releasing the drug in response to a change of at least 25%/hour on a normalized scale being observed in the normalized monitoring data for a period of at least 1 hour. 2. The device according to claim 1, wherein the at least one body parameter includes at least one of body temperature, core body temperature, skin surface temperature, blood pressure, melatonin level, triacylglycerol level, and cortisol level. 3. The device according to claim 1, wherein the central control unit controls the ingestible capsule to start releasing the drug during the at least one monitoring cycle in response to a change of at least 40%/hour on the normalized scale being observed in the normalized monitoring data for a period of at least 3 hours. 4. The device according to claim 1, wherein the central control unit controls the ingestible capsule to start releasing the drug during the at least one monitoring cycle in response to the normalized monitoring data deviating from a minimum point of the normalized curve by a threshold value of at most 25%. 5. The device according to claim 1, wherein the central control unit controls the ingestible capsule to start releasing the drug during the at least one monitoring cycle in response to the normalized monitoring data deviating from a maximum point of the normalized curve by a threshold value of at most 25%. 6. The device according to claim 1, wherein the central control unit controls the ingestible capsule to delay releasing of the drug for at most 2 measuring and/or monitoring cycles. 7. The device according to claim 1, wherein the central control unit controls the ingestible capsule to delay releasing of the drug for between 1 and 24 hours. 8. The device according to claim 1, wherein the central control unit controls the ingestible capsule to stop releasing the drug after one of: at most 30 cycles andwhen the normalized monitoring data return to between at least −60% and at most 60% of the preselected point of the normalized curve. 9. The device according to claim 1, wherein the drug is also released in response to the normalized monitoring data deviating from a preselected point on the normalized curve by more than a threshold value of at most 50% on a normalized scale. 10. A method for the controlled release of drugs from a capsule, comprising: a) measuring at least one body parameter of a patient for at least 3 m... | Koninklijke Philips Electronics N V | A61K009/22
US07869866 | B2 | 2011-01-11 | Device for the monitoring of physiologic variables through measurement of body electrical impedance | 2005-09-06 | A device for the measurement of body impedance includes a current injecting device for injecting a test current through a first set of electrodes to the body of a patient, and a voltage reading device for reading voltage variations between two electrodes belonging to a second set of electrodes using a coherent demodulator. | 1. A device for the measurement of body impedance comprising: means for injecting a test current through a first set of electrodes to the patient's body, including: at least one oscillator and at least one voltage—current converter and means to read voltage variations between two electrodes belonging to a second set of electrodes including: a demodulator; at least one filter; at least one differential amplifier with two outputs having opposite phase and at least one low pass filter, wherein said at least one oscillator is a sinewave oscillator and wherein said demodulator is a coherent demodulator further including at least one comparator and at least one switch, and wherein the output of said sinewave oscillator is connected to the input of said voltage-current converter, the output of said voltage-current converter is connected to said first set of electrodes, said second set of electrodes is connected to the input of said at least one filter, the output of said at least one filter is connected to the input of said at least one differential amplifier, the first output of said at least one differential amplifier is connected to a first input of said at least one switch, the second output of said at least one differential amplifier is connected to a second input of said at least one switch and to the input of said at least one comparator, the output of said at least one comparator is connected to the control input of said at least one switch, the output of said at least one switch is connected to the input of said at least one low pass filter. 2. A device for the measurement of body impedance comprising: means for injecting a test current through a first set of electrodes to the patient's body, including: at least one oscillator and at least one voltage—current converter and means to read voltage variations between two electrodes belonging to a second set of electrodes including: a demodulator; at least one filter; at least one differential amplifier with two outputs having opposite phase and at least one low pass filter, wherein said at least one oscillator is a sinewave oscillator and wherein said demodulator is a coherent demodulator further including at least one comparator and at least one switch, and wherein the output of said sinewave oscillator is connected to the input of said voltage-current converter and to the input of said at least one comparator, the output of said voltage-current converter is connected to said first set of electrodes, said second set of electrodes is connected to the input of said at least one filter, the output of said at least one filter is connected to the input of said at least one differential amplifier, the first output of said at least one differential amplifier is connected to a first input of said at least one switch, the second output of said at least one differential amplifier is connected to a second input of said at least one switch, the output of said at least one comparator is connected to the contr... | Smartex S.r.l. | A61B005/00
US08594779 | B2 | 2013-11-26 | Seizure prediction | 2007-04-30 | Seizure prediction systems and methods include measuring impedance and a potential within a brain of a patient to determine whether the brain is in a state indicative of a possibility of seizure. In some embodiments, at least one of the measured impedance or the measured potential may be used as a primary indication of the brain state indicative of a possibility of seizure. In one embodiment, if one of the measured impedance or the measured potential indicates a seizure, the other measurement (impedance or potential) may be used to validate whether the brain is in the state indicative of the possibility of seizure. | 1. A method comprising: receiving, with a processor, impedance measurements indicative of an impedance of a brain of a patient, wherein the impedance of the brain is indicative of a physiological state of the brain;receiving, with the processor, potential measurements indicative of a potential of the brain;selecting, with the processor, a frequency component of a waveform of the impedance measurements from a plurality of frequency components of the waveform of the impedance measurements; anddetermining, with the processor, whether the brain is in a state indicative of a possibility of a future seizure based on the potential measurements and the selected frequency component of the waveform of the impedance measurements. 2. The method of claim 1, further comprising measuring, with a sensing module, the potential of the brain to generate the potential measurements by at least measuring the potential via at least one of electroencephalography (EEG) or electroencephalography (ECoG). 3. The method of claim 1, further comprising measuring, with a sensing module, the potential of the brain to generate the potential measurements by at least measuring the potential of the brain at a different time than measuring the impedance of the brain to generate the impedance measurements. 4. The method of claim 1, further comprising measuring, with a sensing module, the impedance of the brain of the patient to generate the impedance measurements and measuring the potential of the brain to generate the potential measurements by at least alternating between measuring the impedance of the brain and measuring the potential of the brain. 5. The method of claim 1, wherein determining whether the brain is in the state indicative of a possibility of a future seizure based on the potential measurements and the selected frequency component of the waveform of the impedance measurements comprises: determining whether the selected frequency component of the waveform of the impedance measurements indicates the brain is in the state indicative of the possibility of a future seizure; andvalidating that the brain is in the state indicative of the possibility of a future seizure based on the potential measurements. 6. The method of claim 5, wherein validating that the brain is in the state indicative of the possibility of a future seizure based on the potential measurements comprises determining whether the potential measurements indicate the brain is in the state indicative of the possibility of a future seizure. 7. The method of claim 6, further comprising storing an indicator if the potential measurements do not indicate the brain is in the state indicative of the possibility of a future seizure. 8. The method of claim 1, wherein determining whether the brain is in the state indicative of a possibility of a future seizure based on the potential measurements and the selected frequency component of the waveform of the impedance measurements comprises: determining whether the b... | Medtronic, Inc. | A61B005/04 A61B005/00
US08469899 | B2 | 2013-06-25 | Management of gastro-intestinal disorders | 2008-03-03 | The present invention relates to the field of methods and apparatus for the determination of various conditions of gastric and gastro-intestinal malfunction, especially those performed by means of breath tests. | 1. A method for using a breath test for the determination of gastro-intestinal disorders in a subject, the method comprising: detecting, using a first measurement of a breath test, the generation of a predefined gas in said subject, after the subject was administered a meal comprising at least a first marker material and a second marker material, wherein said first marker material being such that it is not generally absorbed in the subject's stomach, and releases said predefined gas in the presence of intestinal bacteria, and wherein said second marker material being such that it indicates a location of said meal within the subject's gastro-intestinal tract,determining, using a second measurement of said breath test, the position within the subject's gastro-intestinal tract at which said predefined gas is generated, using said second marker material, andevaluating said gastro-intestinal condition of said subject from the outcome of said breath test. 2. The method according to claim 1, wherein the method further comprises: detecting a by-product of said second marker material, using said breath test, anddetermining the position of said predefined gas generation in the gastro-intestinal tract of said subject, by the temporal relationship between the appearance of said predefined gas and of said by-product of said marker material in said subject's breath. 3. The method according to claim 2, wherein said second marker material is labeled with a carbon isotope, and said by-product is isotonically labeled carbon dioxide. 4. The method according to claim 2, wherein the method further comprises determining whether said subject is suffering from one or both of a sugar intolerance and a bacterial overgrowth, using a relative time and quantity of detection of said predefined gas and labeled by-products of said second marker material, wherein said second material is an isotonically labeled material generally absorbed in the small intestine. 5. The method according to claim 4, wherein the method further comprises indicating that said subject is suffering a bacterial overgrowth, using the detection of a small quantity of said predetermined gas, characteristic of a small part of said first material in the presence of bacteria, occurring concurrently with the detection of said labeled by-products of said second marker material. 6. The method according to claim 4, wherein the method further comprises indicating that said subject is suffering from said sugar intolerance, based on the detection of said predefined gas later than the detection of said labeled by-products of said second marker material. 7. The method according to claim 4, wherein the method further comprises indicating that said subject is suffering said sugar intolerance and a bacterial overgrowth, based on the detection of a large quantity of said predefined gas, characteristic of the majority of said first material in the presence of bacteria, occurring essentially concurrently with the detec... | Exalenz Bioscience Ltd. | A61B005/00 G01N033/497
US10534129 | B2 | 2020-01-14 | System and method providing intracoronary laser speckle imaging for the detection of vulnerable plaque | 2008-03-28 | Apparatus and method according to an exemplary embodiment of the present invention can be provided for analyzing tissue. For example, the apparatus can include at least one first arrangement configured to illuminate at least one anatomical structure with at least one of at least one electro-magnetic radiation. The apparatus can also include at least one second arrangement that may include at least two wave-guiding arrangements associated with one another that are configured to receive a further electro-magnetic radiation reflected from the tissue and transmit at least one speckle pattern associated with the further electro-magnetic radiation. The wave-guiding arrangements may be structured so as to reduce crosstalk therebetween. | 1. An apparatus for intracoronary laser speckle imaging, comprising: a light-delivery optical fiber configured to illuminate a coronary wall with laser light that scatters from the coronary wall to form laser speckle patterns;a light-collecting fiber bundle configured to transmit laser speckle patterns formed by the scattered light, wherein fibers in the light-collecting fiber bundle are separated by at least 3 times a width of at least one fiber core so as to reduce crosstalk between fibers down to no greater than 10 percent of collected light; anda detector configured to detect the laser speckle patterns transmitted by the light-collecting fiber bundle;wherein the apparatus is configured to diminish effects of intervening blood on detection of laser speckle patterns by allowing no more than 200 μm of blood to flow between the light-collecting fiber bundle and the coronary wall. 2. The apparatus of claim 1 wherein fibers in the light-collecting fiber bundle include a primary acid-resistant cladding and a secondary acid-soluble cladding that bonds individual fibers. 3. The apparatus of claim 1 wherein the light-collecting fiber bundle fiber bundle has a partial core size no greater than 0.36, wherein the partial core size is core area divided by individual fiber area. 4. The apparatus of claim 1 wherein the light-collecting fiber bundle is leached. 5. The apparatus of claim 1 wherein the crosstalk is measurable by illuminating one fiber in the light-collecting fiber bundle and detecting light leakage in fibers adjacent to the illuminated fiber. 6. The apparatus of claim 1 further including a catheter and a motorized, computer-controlled translation stage for translating the catheter along a coronary segment. 7. A method for speckle imaging, comprising: using a light-delivery optical fiber to illuminate an anatomical structure with coherent light that scatters from the anatomical structure to form a speckle pattern;using a light-collecting fiber bundle to collect scattered light and transmit the speckle pattern formed by the scattered light, wherein fibers in the light-collecting fiber bundle are separated by at least 3 times a width of at least one fiber core so as to reduce crosstalk between fibers; andto diminish effects of intervening blood on detection of the laser speckle pattern, allowing no more than 200 μm of whole blood to flow between the light-collecting fiber bundle and the anatomical structure. 8. The method of claim 7 wherein the light-collecting fiber bundle reduces a measurable crosstalk down to no greater than 10 percent of light collected by the light-collecting fiber bundle. 9. The method of claim 8 wherein the light-collecting fiber bundle includes cladding, and wherein the crosstalk is measurable by illuminating one fiber in the light-collecting fiber bundle and detecting light leakage in the cladding and fibers adjacent to the illuminated fiber. 10. The method of claim 8 wherein a radius of curvature of fibers in th... | The General Hospital Corporation | G02B006/04 G02B006/06 A61B001/00 A61B005/00 G02B006/32
US20080312556 | A1 | 2008-12-18 | Device for Measuring Data Relating to Urine Production of a Patient | 2005-06-14 | The invention relates to a device for measuring data relating to urine production of a patient for medical application. The device is provided with a collecting device (1) for collecting the urine of the patient and a processor (14) for connecting to the collecting device. The processor is provided with means (12) for measuring the data and memory means (13) for storing the measurement data. The urine collecting device (1) is arranged for placing in a toilet bowl. According to the invention the collecting device is intended as a disposable article, whereby the hygiene of the patient is guaranteed. | 1. Device for measuring data relating to urine production of a patient for medical application, which device is provided with: a collecting device (1) for collecting the urine of the patient;means (12) for measuring the data; anda processor (14) provided with memory means (13) for storing the measurement data,characterized in that the urine collecting device (1) is arranged for mounting on a toilet bowl. 2. Device as claimed in claim 1, wherein the urine collecting device (1) is funnel-shaped and provided with means (4, 4A, 5, 5A, 6) for mounting in the toilet bowl. 3. Device as claimed in claim 1 or 2, wherein the urine collecting device (1) takes an at least partially double-walled (2,3) form and is provided with a liquid lock (7, 8), and wherein the measuring means are arranged to measure air displacement as a result of the urine production. 4. Device as claimed in any of the preceding claims, wherein the measuring means (12) are incorporated in the processor (14). 5. Device as claimed in any of the preceding claims, wherein the measuring means comprise a flow sensor (12) for measuring data relating to the urine flow from the patient. 6. Device as claimed in claim 4 or 5, wherein the processor (14) can be mounted releasably on the collecting device (1). 7. Device as claimed in claim 6, wherein the collecting device (1) is provided with a cover (18) for at least partially receiving the processor (14). 8. Device as claimed in any of the preceding claims, further comprising an interface (16) for connection to the processor (14) for displaying the measurement data. 9. Device as claimed in claim 8, wherein the processor (14) is further provided with a transmitter (15) for wireless transmission of the measurement data to the interface (16), this interface being provided with a receiver (17) for receiving the measurement data. 10. Device as claimed in claim 8 or 9, wherein the interface (16) comprises a game computer, such as a Gameboy®. | BEST MEDICAL B.V. | A61B005/20 G01F019/00
US20080281235 | A1 | 2008-11-13 | Angular Displacement Sensor for Joints And Associated System and Methods | 2008-05-16 | A system for monitoring joint position following introduction of a joint prosthesis in a patient includes a first angular movement sensor positioned adjacent a first side of a bodily joint of a patient and a second angular movement sensor positioned adjacent a second, opposite side of the bodily joint. A receiver can receive data from the angular movement sensors. A processor in signal communication with the receiver has a comparator for determining whether the received data from the sensors indicate that movement of the bodily joint is less than a predetermined limit. An indicator in signal communication with the comparator is provided for warning the patient if the joint movement is determined by the comparator to be above the predetermined limit. The received data can be tracked over time. | 1. A system for monitoring joint position following introduction of a joint prosthesis in a patient, the system comprising: a first angular movement sensor positionable adjacent a first side of a bodily joint of a patient;a second angular movement sensor positionable adjacent a second side of the bodily joint of the patient, the second side on an opposite side of the bodily joint from the first side;a receiver adapted to receive data from the first and the second angular movement sensors;a processor in signal communication with the receiver having comparison means resident thereon for determining whether the received data from the first and the second angular movement sensors indicate that movement of the bodily joint is less than a predetermined limit; andan indicator in signal communication with the comparison means forwarning the patient if the joint movement is determined by the comparison means to be above the predetermined limit. 2. The system recited in claim 1, wherein the prosthesis introduction comprises a hip replacement, the hip first side comprises a location on a pelvis of the patient, and the hip second side comprises a location on the hip prosthesis. 3. The system recited in claim 1, wherein the first and the second angular movement sensor each comprise a battery-powered gyroscope adapted for wireless communication with the receiver. 4. The system recited in claim 1, further comprising an input device adapted for signal communication with the processor, the input device comprising means for receiving the predetermined angular movement limit from a user and for transmitting the received predetermined angular movement limit to the processor. 5. The system recited in claim 4, wherein the input device comprises a handheld input device comprising means for entering angular limits of abduction/adduction, flexion/extension, and internal/external rotation. 6. The system recited in claim 5, wherein the entering means comprises a numerical keypad. 7. The system recited in claim 1, wherein the indicator comprises one of a vibrator affixable to a body of the patient and a means for issuing an audio signal. 8. The system recited in claim 1, wherein the indicator comprises means for issuing a warning signal to the patient, and wherein an intensity of an issued warning signal is related to a difference between the predetermined limit and the received data from the first and the second angular movement sensors. 9. The system recited in claim 1, wherein the processor further has a timing element, and further comprising a memory device in signal communication with the processor, for storing a time history of the received data from the first and the second angular movement sensors. 10. The system recited in claim 9, further comprising a reading device for downloading the stored time history from the memory device thereinto, for use in tracking the movement of the bodily joint over time. 11. The system recited in claim 1, wherein the first ... | - | A61B005/117
US20080319291 | A1 | 2008-12-25 | Blood Testing Apparatus Having a Rotatable Cartridge with Multiple Lancing Elements and Testing Means | 2008-04-23 | A blood testing apparatus includes a device for withdrawing blood, a membrane-type testing element, an evaluating device, and a display device forming a complete system which can be manipulated as a single piece of apparatus. Multiple testing elements can be inserted into the apparatus and brought successively to a work position for carrying out several measurements. The blood withdrawing device includes multiple pricking elements. One pricking element is pushed through one testing element and pricks the surface of the skin of a user. The pricking position is disposed so that blood withdrawn from the surface of the skin can impinge upon a testing element. | 1. A blood testing apparatus, comprising: a test membera laser source configured to produce a wound from which blood flows;electronics for analysis, anda display, wherein the test member, laser source, electronics and display form a glucose monitoring system that is integrated in a single apparatus, 2. The apparatus of claim 1, wherein the test member includes an opening from which a beam from the laser can pass through. 3. A blood testing apparatus, comprising: a plurality of test members;a laser source configured to produce a wound from which blood flows;electronics for analysis, anda display, wherein the plurality of test members, laser source, electronics and display form a glucose monitoring system that is integrated in a single apparatus, 4. The apparatus of claim 2, wherein the plurality of test members are arranged radially around an axis of rotation of a rotatable, disposable device having a circumferential wall section that runs cylindrically on an outer periphery, each of the test member and the at least one lancing element having a longitudinal axis that is substantially perpendicular relative to the axis of rotation, and wherein the test members are disposed with surfaces normal in a radial direction with respect to a center of the rotatable, disposable device 5. The apparatus of claim 2, wherein each of a test member includes an opening from which a beam from the laser can pass through. 6. A method for testing an analyte in a blood sample, comprising: providing a blood testing apparatus that includes a test member, a laser source, electronics for analysis and a display that are integrated in a single apparatus;producing a wound with the laser;receiving a blood sample from the wound of the test member; andusing the electronics to determine a concentration of glucose in the blood sample. 7. The method of claim 6, wherein a plurality of test members are provided. | - | A61B005/055
US20080300664 | A1 | 2008-12-04 | METHOD AND APPARATUS FOR PLACING A CORONARY SINUS/CARDIAC VEIN PACING AND DEFIBRILLATION LEAD WITH ADJUSTABLE ELECTRODE SPACING | 2008-05-06 | A medical electrical lead, which may be useful in coupling an implantable medical device, is comprised of a first and second lead. The first lead has a first electrode coupled adjacent a distal end portion thereof. The distal end portion of the first lead is anchorable in the coronary sinus of a patient. The second lead is coupled with and moveable along the first lead. The second lead has a second electrode located thereon wherein the position of the first and second electrodes may be varied relative to one another by movement of the second lead along the first lead. A rubber tip holds the relative position of each electrode. | 1. A medical electrical lead, comprising: a first lead having a first electrode coupled adjacent a distal end portion thereof, said distal end portion of said first lead being anchorable in the coronary sinus of a patient; anda second lead coupled with and moveable along said first lead, said second lead having a second electrode located thereon wherein the position of the first and second electrodes may be varied relative to one another by movement of said second lead along said first lead,wherein said first lead includes a core of electrically conductive material surrounded by a layer of insulative material,wherein said first lead includes an expandable helical coil coupled to its distal end and a fixed helix for attachment to a heart tissue. 2. A medical electrical lead, as set forth in claim 1, wherein said second lead is an over-the-wire lead. 3. A medical electrical lead, as set forth in claim 1, wherein said second lead includes an elongated flexible body having first and second end portions and a passage extending therebetween, and wherein said first lead passes through said passage. 4. A medical electrical lead, as set forth in claim 3, wherein the elongated flexible body comprises a flexible tube. 5. A medical electrical lead, as set forth in claim 3, wherein the elongated flexible body comprises a polyurethane tube. 6. A medical electrical lead, as set forth in claim 3, wherein said second lead includes an anode formed of electrically conductive material coupled to a distal end portion of the elongated flexible body, and wherein said anode includes a bore in communication with said passage, and wherein said first lead passes through said bore. 7. A medical electrical lead, as set forth in claim 6, wherein said anode includes a seal coupled thereto and in sliding contact with said first electrode to resist fluid egress into said bore and said passage. 8. A medical electrical lead, as set forth in claim 7 wherein said seal holds said first and second electrodes in relative position during implant. 9. A medical electrical lead, as set forth in claim 6, wherein said anode is tapered along a longitudinal axis of the elongated flexible body and the anode. 10. A medical electrical lead, as set forth in claim 6 wherein the anode further comprises an annular lip and wherein the elongated flexible body extends over the annular lip. 11. A medical electrical lead, as set forth in claim 1, wherein the first and second leads are adapted to be coupled together at their proximal end portions to prevent relative longitudinal movement therebetween. 12. A medical electrical lead as set forth in claim 1, wherein the first and second leads are adapted to electrically couple said first lead to a connector pin of said second lead. 13. A medical electrical lead, comprising: a first lead having a first electrode coupled adjacent a distal end portion thereof, said distal end portion of said first lead being anchorable in the coronary sinus of a pati... | - | A61N001/05 A61B005/042
US20080275320 | A1 | 2008-11-06 | METHOD AND DEVICE FOR MEASUREMENTS IN BLOOD | 2008-05-30 | We present an optical probe arrangement that surrounds blood in a receptacle. It comprises LED's and light detector arranged to overcome the variations when the receptacle is translucent medical tubing and the like. Also, a signal processing algorithm is used to average signals from a plurality of light detectors, to further enhance results when measuring hematocrit. The invention makes it possible to add the feature of hematocrit measurement into dialysis system without major alterations to the dialysis machine or transport tubing. | 1-3. (canceled) 4. Sensor device for the measuring of the density of a fluid flowing through a transparent tubing, said device comprising an elongate frame sized to fit around the tubing, said frame carrying a light beam emitter facing the tubing, and one or several sensor(s) also facing the vessel and so arranged that its or their sense sectors do not intersect the beam of the light source in the vessel or tubing. 5. Sensor device according to claim 4, wherein the locations of the light source and sensor(s) respectively are separated lengthwise of the frame. 6. Sensor device according to claim 4, wherein two sensors are arranged with their sensing directions perpendicular to the light beam. 7. An optical probe arrangement that surrounds a fluid flowing through an elongate transparent tubing, said optical probe arrangement comprising a frame sized to surround said receptacle in part, said frame supporting at least two sets of light emitters and light detectors, each set comprising one light emitter and at least one detector, each set arranged to transilluminate the fluid at a preferred angle between said light emitter and said light detector—or detectors—of each set, where said angle is at least sufficient to avert direct light from said light emitter to said light detector, for the detection of constituents in said fluid. 8. An optical probe arrangement according to claim 7, comprising four sets of light emitters and two or three light detectors in each set, wherein a light detector represents a detector incorporated in an adjacent set. 9. An optical probe arrangement according to claim 7, wherein the light emitters are arranged as an array to encircle said elongated tubing at longitudinally one location around said tubing's circumference, and the light detectors are arranged to encircle the tubing at a different circumferential location. 10. An optical probe arrangement according to claim 7, wherein a second array of light detectors are longitudinally located at a third location around said tubing's circumference, and the light detectors are arranged to encircle the tubing at that circumferential location. 11. A method for processing signals from sensor devices as claimed in claim 4, including an amplifier for amplifying signals from the sensor devices, which comprises applying a signal processing algorithm on the signals from said sensor devices to calculate density. 12. A method for processing signals from light detectors as claimed in claim 11, which comprises applying a signal processing algorithm on the signals from said light detectors, to detect said constituents. 13. A method according to claim 12, which comprises applying a multi variable analysis of signals from all light detectors engaged in the signaling process. 14. A sensor device as claimed in claim 4, wherein a third array of light sensors is longitudinally located at a fourth location around said tubing's circumference, and the light sensors are arranged to encircle ... | - | A61B005/1455 G01N033/49
US20080306358 | A1 | 2008-12-11 | Body Insertable System, Receiving Apparatus, and Body Insertable Apparatus | 2005-09-13 | A receiving apparatus 3 which is an element of a body insertable system is configured with a reciving unit 6 including receiving antennas 8a to 8d and a reception processing device 9, and a postion detecting unit 7 including transmitting antennas 10a to 10d, a first linear magnetic field generator 11a, a second linear magnetic field generator 11b, a diffusr magnetic field generator 12, and a processing device 13, and the receiving unit 6 and the postion detecting unit 7 are formed seperately and independently of each other. Therefore, when the body insertable system is used for acquisition of intra-subject information which is acquired by the capsule endoscope 2 and position detection of the capsule endoscope 2, both the receiving unit 6 and the position detecting unit 7 are used, whereas when the body insertable system is used only for the acquisition of the intra-subject information, it is possible to use only the receiving unit 6, whereby a burden on the subject at a time of use is restricted to a minimum degree according to the purpose of use while increase in operational cost is suppressed. | 1. A body insertable system comprising: a body insertable apparatus that is introduced inside a subject, acquires intra-subject information as information concerning the subject, and transmits radio signals including the acquired intra-subject information; anda receiving apparatus that performs reception processing of the radio signals transmitted by the body insertable apparatus, whereinthe body insertable apparatus includes: an intra-subject information acquiring unit that acquires the intra-subject information;a magnetic field sensor that detects a magnetic field in a region where the body insertable apparatus is located; anda radio transmitting unit that transmits radio signals including at least the intra-subject information, andthe receiving apparatus includes: a receiving unit that includes at least a receiving antenna which receives the radio signals transmitted by the body insertable apparatus and a receiving circuit which performs reception processing on the radio signals received by the receiving antenna; anda position detecting unit that includes a magnetic field generator which generates a predetermined magnetic field for position detection in a region where the body insertable apparatus can be present, and a position calculator that calculates a position of the body insertable apparatus based on a result of detection of the magnetic field for position detection acquired by the magnetic field sensor, the position detecting unit being formed separately and independently of the receiving unit. 2. The body insertable system according to claim 1, wherein the radio transmitting unit transmits radio signals including a result of detection acquired by the magnetic field sensor in addition to the intra-subject information, andthe position detecting unit acquires the result of detection acquired by the magnetic field sensor via the receiving unit. 3. The body insertable system according to claim 1, wherein the position detecting unit is arranged in a fixed state relative to the subject at a time of use, andthe receiving unit is arranged in a movable state relative to the subject at a time of use. 4. A receiving apparatus performing reception processing of a radio signal transmitted from a predetermined detection target, comprising: a receiving unit that includes at least a receiving antenna which receives the radio signal transmitted from the detection target, and a receiving circuit which performs reception processing on the radio signal received by the receiving antenna; anda position detecting unit that includes a magnetic field generator which generates a predetermined magnetic field for position detection in a region where the detection target can be present, and a position calculator which calculates a position of the detection target based on a result of detection of the magnetic field for position detection in the region where the detection target can be present, the position detecting unit being formed separately and independently... | - | A61B001/00 A61B005/05
US20080262344 | A1 | 2008-10-23 | RELATIVE VALUE SUMMARY PERFUSION MAP | 2007-04-23 | Disclosed is a system and method for diagnosing a perfusion condition in the human brain. The process used includes measuring parameter values at a location of interest in one hemisphere and determining relative differences between those measured values against values measured at a substantially symmetrically opposite location in the other hemisphere. These relative differences are used to generate a composite output which is displayed using different colors. Each color is reflective of a different tissue state, e.g., ischemic core, penumbra, etc. If the substantially symmetrically opposite location is abnormal, and thus, is not a good reference, an alternative reference location which is not substantially symmetrically opposite the location of interest is chosen, and a relative difference map generated. Alternatively, an absolute value map can be generated. | 1. A process for diagnosing a perfusion condition in the brain of a patient, comprising: selecting an anatomical location of interest in a first hemisphere of a brain;identifying a substantially symmetrically opposite location in an other hemisphere of said brain;measuring a first set of values for plurality of parameters at said location of interest;measuring a second set of values for said plurality of parameters at said substantially symmetrically opposite location in said other hemisphere;determining relative differences between said first and second sets of values; andusing said relative differences to generate a composite output which is reflective of said condition at said location of interest. 2. The process of claim 1, comprising: providing a device to display said output as an image;using a plurality of colors to differentiate between different tissue states in said image. 3. The process of claim 1, comprising: adapting said image to be a perfusion map. 4. The process of claim 1 wherein said process of using said relative differences to generate a composite output step comprises: classifying tissues into different perfusion states utilizing threshold values applied to each of said plurality of parameters. 5. The process of claim 1, comprising: determining whether tissues at both of said location of interest and said substantially symmetrically opposite locations are abnormal; anddirecting the process to an alternative method of diagnosing said perfusion condition if both of said location of interest and said substantially symmetrically opposite location include tissues which are abnormal. 6. The method of claim 5 wherein said alternative method comprises: selecting an alternative reference location which is not substantially symmetrically opposite said location of interest, said alternative reference location including tissues in a normal condition in one of said first and other hemispheres;measuring said plurality of parameter to determine an alternative set of values at said alternative reference location;determining alternative relative differences between said first set of values and said alternative set of values; andusing said alternative relative differences to generate a composite output which is reflective of said condition at said location of interest. 7. The method of claim 5 comprising: generating an alternative summary map using absolute values. 8. The method of claim 1 comprising: providing the user with an ability to view output values relating to each of said parameters individually rather than in composite form. 9. The method of claim 1 comprising: including MTT in said plurality of parameters for the purpose of determining the existence of ischemia; andincluding CBV in said plurality of parameters for the purpose of determining the existence of infarct. 10. The method of claim 9 comprising: including CBF in said plurality of parameters for the purpose of determining the existence of ischemia along with said MT... | - | A61B006/00 A61B005/05
US20080255440 | A1 | 2008-10-16 | Method of Forming a Sterilized Sensor Package and a Sterilized Sensor Package | 2005-07-13 | A sensor package is provided comprising an implantable sensor having an electrode area and an electric contact area, the package comprising a shielding packaging enclosing at least the electrode area of the sensor while exposing the electric contact area. The sensor package fascilitates handling of the implantable sensor and provides for the possibility of sterilising the electric contact area leaving the electrode area protected. | 1. A method of forming a sterilised sensor package for an electrochemical sensor, said method comprising the steps of: providing a sterilised sensor having an electrode area and an electric contact area,forming a shielding packaging that is impermeable to micro-organisms in such a manner that at least the electrode area is enclosed in the shielding packaging, the electric contact area extending outside the shielding packaging. 2. A method according to claim 1, wherein at least the electrode are of the sensor is sterilised by radiation. 3. A method according to claim 1, wherein the shielding packaging is impermeable to gas. 4. A method according to claim 1, and wherein the sensor comprises electronics, and wherein the method further comprising the step of connecting the electronics to the electric contact area succeeding sterilisation of the electrode area of the sensor. 5. A method according to claim 1, wherein the electronics are sterilised and enclosed by a further shielding packaging. 6. A method according to claim 3, wherein the electronics are sterilised by means of a gas. 7. A method according to claim 6, wherein the gas is ethylene oxide (EtO). 8. A method of sterilising a series of sensors according to claim 1, wherein the test samples of the sterilised sensors are extracted; that calculation of calibration parameters is performed by measurement; and that the part of the sensors that projects outside the shielding packaging is provided with an information carrier that contains the calibration parameters. 9. A method according to claim 8, wherein passive circuits are used as information carrier. 10. A method according to claim 8, wherein active electronic circuits are used as information carrier. 11. A sensor package comprising an implantable sensor having electrodes and electric coupling areas, which sensor is kept temporarily sterile by means of a shielding packaging or housing, wherein a base packaging is provided that shields at least the electrodes; and that the electric coupling areas are situated outside the base packaging. 12. A sensor package according to claim 11, wherein an additional packaging is provided configured for completely or partially shielding those parts of the sensor that project outside the base packaging. 13. A sensor package according to claim 12, wherein the additional packaging encloses the base packaging. 14. A sensor package according to claim 11, wherein the base packaging contains means for transdermal insertion of the electrodes of the sensor. 15. A sensor package according to claim 12, wherein the base packaging and the additional packaging adjoin each other along an endless rim comprising a tearing line or a rupture zone. 16. A sensor package according to claim 12, wherein at least a part of the base packaging is constituted by a form-stable housing. 17. A sensor package according to claim 14, wherein the means for transdermal insertion is a needle which is supported by the form-stable ho... | Novo Nordisk A/S | A61B005/00 B65B055/08
US20080243022 | A1 | 2008-10-02 | SEIZURE PREDICTION USING BRAIN SIGNAL TELEMETRY | 2007-03-30 | An ambulatory intrinsic brain signal processor circuit is coupled to a plurality of electrodes. The signal processor circuit can include a digital multiplexer circuit coupled to the electrodes to multiplex brain signal data from different electrodes together into a multiplexed data stream. An ambulatory transceiver circuit wirelessly communicates information to and from a remote transceiver. A controller circuit permits a user to control which of the electrodes contribute data, a data resolution, and whether the data includes one or both of neural action or local field potential data. Seizure prediction components and methods are also described. | 1. A method comprising: receiving a Normal template providing an indication of correlation of intrinsic brain potentials during at least one non-seizure time period of a subject, wherein the non-seizure time period excludes a seizure time period of a seizure, and wherein the non-seizure time period excludes at least a first specified time period preceding the seizure;receiving a Non-Normal template providing an indication of correlation of the brain potentials during at least one pre-seizure time period or seizure time period of the subject, wherein the pre-seizure time period is less than or equal to a second specified time period before the seizure, and wherein the seizure occurs during the seizure time period;monitoring intrinsic brain potentials using at least two different locations of a brain of the subject and forming an indication of correlation of the brain potentials at the at least two different locations during a sampling time period; andpredicting an upcoming seizure at least in part by comparing the indication of correlation of the brain potentials obtained during the sampling time period to each of the Normal and Non-Normal templates. 2. The method of claim 1, comprising: receiving a seizure occurrence input to establish a time of at least one known seizure of a subject;monitoring brain potentials using at least two different locations of the brain of the subject; andforming the Normal and Non-Normal templates using information from the monitoring and the time of the at least one known seizure of the subject. 3. The method of claim 1, wherein the intrinsic brain potentials are local field potentials. 4. The method of claim 1, wherein the intrinsic brain potentials are intrinsic neuronal action potentials. 5. The method of claim 4, wherein the monitoring intrinsic brain potentials comprises: acquiring and digitizing neuronal action potential signals at separate locations of different electrodes;communicating information about the digitized action potential signals to an ambulatory transmitter circuit located at the subject; andtransmitting information about the digitized action potential signals to at least one of a local or remote user-interface device. 6. The method of claim 4, wherein the monitoring intrinsic brain potentials comprises monitoring single-unit activity (SUA) of individual neurons. 7. The method of claim 4, wherein the monitoring intrinsic brain potentials comprises monitoring multi-unit activity (MUA) of a set of nearby individual neurons. 8. The method of claim 1, wherein the monitoring includes counting a number of neuronal signal energy indications that exceed a specified threshold value. 9. The method of claim 1, wherein the monitoring includes integrating a neuronal signal over time. 10. The method of claim 1, wherein the first specified time period is at least one hour. 11. The method of claim 1, wherein the second specified time period is less than or equal to one hour. 12. The method of claim 1, ... | - | A61B005/04
US20080242954 | A1 | 2008-10-02 | IMAGE PRESENTATION APPARATUS, IMAGE PRESENTATION METHOD AND PROGRAM | 2008-03-27 | An image presentation apparatus is provided that includes a display section that sequentially displays a plurality of images, a biosignal acquisition section that acquires a signal value of a biosignal indicating a biological response of an observer watching the image displayed on the display section, and a display controlling section that controls a timing at which an image subsequent to the image being displayed on the display section is displayed on the display section based on the signal value acquired by the biosignal acquisition section. | 1. An image presentation apparatus that presents an image, comprising: a display section that sequentially displays a plurality of images;a biosignal acquisition section that acquires a signal value of a biosignal indicating a biological response of an observer watching the image displayed on the display section; anda display controlling section that controls a timing at which an image subsequent to the image being displayed on the display section is displayed on the display section based on the signal value acquired by the biosignal acquisition section. 2. The image presentation apparatus according to claim 1, further comprising: a comparison value storage section that stores a predetermined comparison value of a biosignal; anda biosignal value comparison section that compares the signal value acquired by the biosignal acquisition section to the comparison value, wherein the display controlling section controls the timing at which the subsequent image is displayed on the display section based on a comparison result by the biosignal value comparison section. 3. The image presentation apparatus according to claim 2, wherein the display controlling section displays the subsequent image on the display section in a case where the signal value acquired by the biosignal acquisition section reaches the comparison value. 4. The image presentation apparatus according to claim 2, wherein: the comparison value storage section stores a first comparison value of a biosignal and a second comparison value of a biosignal, which is greater than the first comparison value; andthe display controlling section controls the timing at which the subsequent image is displayed on the display section to keep the signal value acquired by the biosignal acquisition section between the first comparison value and the second comparison value. 5. The image presentation apparatus according to claim 4, wherein the display controlling section displays the subsequent image on the display section in a case where the signal value acquired by the biosignal acquisition section is outside of a range between the first comparison value and the second comparison value. 6. The image presentation apparatus according to claim 5, wherein the display controlling section displays the subsequent image on the display section in a case where the signal value acquired by the biosignal acquisition section is outside of the range between the first comparison value and the second comparison value for a predetermined length of time. 7. The image presentation apparatus according to claim 2, wherein: the comparison value storage section stores a comparison value for every period of time that passes since the image begins to be displayed on the display section; andthe biosignal value comparison section compares the signal value acquired by the biosignal acquisition section to the comparison value corresponding to the amount of time that has passed since the image began to be displayed on the display se... | FUJIFILM Corporation | A61B005/00
US20080287783 | A1 | 2008-11-20 | SYSTEM AND METHOD OF TRACKING DELIVERY OF AN IMAGING PROBE | 2008-04-25 | A system and method to track delivery of a surgical instrument through an imaged subject is provided. The system comprises a controller and an imaging system including an imaging probe in communication with the controller. The imaging probe includes a transducer array operable to acquire image data through a range of motion about a longitudinal axis and in a direction of image acquisition with the imaging probe stationary. The system also includes a tracking system to track a position of the imaging probe relative to a second object tracked by the tracking system, and a display illustrative of a direction of image acquisition of the imaging probe relative to an illustration of a position of the second object. | 1. A system to track delivery of a surgical instrument through an imaged subject, comprising: a controller; andan imaging system including an imaging probe in communication with the controller, the imaging probe having a transducer array operable to acquire image data through a range of motion about a longitudinal axis and in a direction of image acquisition with the imaging probe stationary;a tracking system to track a position of the imaging probe relative to a second object tracked by the tracking system;a display illustrative of a direction of image acquisition of the imaging probe relative to an illustration of a position of the second object. 2. The system of claim 1, wherein the imaging probe includes a tracking element attached at a transducer array, the transducer array operable to move the range of motion about the longitudinal axis and to acquire ultrasound image data of the imaged subject. 3. The system of claim 2, wherein the imaging probe includes a marker attached at a transducer array, the transducer array operable to move the range of motion about the longitudinal axis and to acquire ultrasound image data of the imaged subject, and wherein the controller is operable to detect an illustration of the marker in the acquired imaged data. 4. The system of claim 1, the display further comprising an illustration of a path of delivery of the second object to a target site. 5. The system of claim 1, wherein the second object is an ablation catheter, and further comprising a steering system operable to move delivery of both the imaging probe and the ablation catheter through the imaged subject. 6. The system of claim 5, wherein the controller is operable to receive an instructions via an input device representative of a selection between a manual steering mode and an automatic steering mode. 7. The system of claim 6, wherein in the automatic steering mode, the steering system automatically moves the imaging probe so that the direction of image acquisition follows movement of a second object traveling through the imaged subject. 8. The system of claim 7, wherein the second object is an ablation catheter, wherein in the manual steering mode and the controller receives instructions to align the direction of image acquisition relative to one of the target site, and wherein the steering system moves both the imaging probe and the ablation catheter in the direction of image acquisition through the imaged subject. 9. The system of claim 1, wherein a transducer array of the imaging probe attaches to both a tracking element of the tracking system and a marker detectable in acquired image data of the imaging probe, and wherein the controller receives instructions representative of selection of tracking of movement of at least one of the tracking element and the marker. 10. A method of tracking delivery of an imaging probe through an imaged subject, the method comprising the steps of: rotating a transducer array about a longitudinal axis of ... | GENERAL ELECTRIC COMPANY | A61B005/05
US20080306401 | A1 | 2008-12-11 | Muscular volume balance evaluation apparatus, a muscular volume balance valuation method | 2008-06-10 | A muscular volume balance measuring apparatus that measures the balance of the muscular volume in a specific part is provided. At least two sets of current applying, devices for inserting a specific part of the body and applying current to living bodies, or a voltmeter, which measures voltage, are provided. A bioelectricity impedance measuring device which measures bioelectricity impedance of a measured region from applied current and measured voltage is also provided. A muscular volume calculating device which computes muscular volume from measured bioelectricity impedance is also provided. Further, balance of muscular volume of the inserted part is calculated from each set of the computed muscular volume. In addition, there is a feature of being able to expand and contract for the longitudinal width, which is the distance which inserts a specific part, and breadth, which is the distance between current applying device or between the voltmeter. | 1. A muscular volume balance evaluation apparatus, which comprises: at least two sets of current applying devices and voltmeter are symmetrically located to a specific part of the body;a bioelectricity impedance measuring device measures bioelectricity impedance of a measured region from current applied by the current applying devices and from voltage measured by the voltmeter;a muscular volume calculating device computes muscular volume from the measured bioelectricity impedance; andthe muscular volume balance evaluation apparatus determines for balance of muscular volume in the specific part by the computed set of the muscular volume. 2. The muscular volume balance evaluation apparatus according to claim 1, wherein longitudinal width, which is an inter-electrode distance symmetrically located to the specific part, andbreadth, which is the distance between the electrode pairs,are able to expand and contract. 3. The muscular volume balance evaluation apparatus according to claim 1, wherein further comprises: a longitudinal width measurement device measures the longitudinal width;a breadth measurement device measures the breadth; anda volume calculating device computes volume for the muscle to measure by the longitudinal width and the breadth. 4. A muscular volume balance evaluation apparatus according to claim 1, wherein the measurement of balance of the muscular volume is for measuring balance of muscular volume about anterior surface and posterior surface of leg. 5. A muscular volume balance evaluation apparatus according to claim 1, wherein further comprises evaluation data in a memory device to evaluate about balance of muscular volume, and computes a desired value of balance of the muscular volume. 6. A muscular volume balance evaluation apparatus according to claim 1, wherein each of the muscular volume for the computed set is displayed separately. 7. A muscular volume balance valuation method, comprising: computing muscular volume from a value of bioelectricity impedance measured in a specific part of the body; andcalculating balance of muscular volume of the specific part measured by the computed set of muscular volume. | Tanita Corporation | A61B005/053
US20080281179 | A1 | 2008-11-13 | ANALYTE MONITORING SYSTEM AND METHODS | 2008-05-08 | Methods and systems for providing data communication in medical systems are disclosed. | 1. An analyte sensor, comprising: a substrate;a plurality of electrodes provided on the substrate, at least a portion of one of the plurality of electrodes positioned in fluid contact with an analyte of a user; anda conductive trace provided on the substrate and coupled to one of the plurality of electrodes. 2. The sensor of claim 1 wherein the conductive trace comprises carbon. 3. The sensor of claim 1 wherein the conductive trace is connected to a ground terminal. 4. The sensor of claim 1 wherein the plurality of electrodes includes a working electrode, a reference electrode and a counter electrode. 5. The sensor of claim 4 wherein the conductive trace is connected to the counter electrode. 6. The sensor of claim 1 wherein the plurality of electrodes are positioned in a stacked configuration. 7. The sensor of claim 1 wherein the conductive trace is configured to establish electrical contact with a power supply to provide electrical signal to one or more of the plurality of electrodes. 8. The sensor of claim 1 wherein the conductive trace is connected to a ground terminal. 9. The sensor of claim 1 wherein the conductive trace is connected to a guard trace. 10. A system for powering a data processing device, comprising: an analyte sensor, including: a substrate;a plurality of electrodes provided on the substrate, at least a portion of one of the plurality of electrodes positioned in fluid contact with an analyte of a user; anda conductive trace provided on the substrate and coupled to one of the plurality of electrodes; anda data processing device including a contact point for electrically connecting to one of the plurality of electrodes, the data processing device further including a power supply wherein when the contact point is in electrical connection with the one of the plurality of electrodes, the power supply is configured to transition the data processing device from a low power state to an active power state. 11. The system of claim 10 wherein the power supply includes a battery. 12. The system of claim 10 wherein the one of the plurality of electrodes includes a counter electrode of the analyte sensor. 13. The system of claim 10 wherein the conductive trace and the data processing device are coupled to a ground terminal. 14. The system of claim 10 wherein the sensor includes a guard trace disposed on the substrate, and further, wherein the data processing device includes a guard contact point for electrically coupling to the guard trace. 15. The system of claim 10 wherein the data processing device includes a data communication unit to transmit one or more signals related to the monitored analyte level received from the analyte sensor. 16. The system of claim 15 wherein the data communication unit includes a close proximity receiver for receiving one or more close proximity commands. 17. The system of claim 10 wherein the analyte sensor includes a glucose sensor. 18. The system of claim 10 wherein the analyte sensor is ... | Abbott Diabetes Care, Inc. | A61B005/145 G06F001/00
US20080249429 | A1 | 2008-10-09 | Spirometers | 2008-03-28 | A portable hand-held spirometer apparatus is described. A main casing has a pair of opposed arms extending to one side thereof. A flow head component unit has a channel through which air may be blown, with formations in the channel imparting a rotary movement to air blown through the channel. A free-spinning vaned member is mounted for rotation about an axis substantially coaxial with that of the channel. The flow head component unit has two open-ended recesses extending to either side of the vaned member and dimensioned to receive the opposed arms when the flow head component unit is fitted on to the main casing. This construction enables the successive use of detachable flow head components, each may be used once and then a batch of which may be sterilized before re-use. | 1. A portable hand-held spirometer apparatus comprising a first casing housing electronics adapted to measure and display one or more lung function parameters and a detachable flow head unit fitted thereto, the flow head unit incorporating a flow head component including a channel through which air may be blown, formations in the channel imparting a rotary movement to air blown through the channel and a free-spinning vaned member mounted for rotation about an axis substantially coaxial with that of the channel, wherein the first casing has a pair of arms extending to one side thereof, and the flow head unit comprises a second casing incorporating two open-ended recesses extending from a side of the second casing to a pair of positions located either side of the flow head component and dimensioned to receive the pair of arms when the first casing and the second casing are fitted together, with the vaned member of the flow head component lying between the pair of arms. 2. The spirometer apparatus of claim 1, wherein a part of the first casing incorporating the flow head unit has, at an inlet end of the channel, a flange or ring for engagement with a disposable spirometer mouthpiece containing an anti-bacterial filter. 3. The spirometer apparatus of claim 1, wherein the pair of arms carry optical means enabling rotation of the vane member to be observed and processed by circuitry in the first casing. 4. The spirometer apparatus of claim 3, wherein the first casing houses a liquid-crystal display to display results deriving from processing by said circuitry. 5. The spirometer apparatus of claim 1, wherein the detachable flow head unit is manufactured from molded pieces which snap fit together to produce said unit which can be fitted on to the pair of arms (20, 21) and removed therefrom after use. 6. The spirometer apparatus of claim 1, wherein the detachable flow head unit includes a pair of members, each of said members containing a portion of a flow passage and each of said members having a central hub into which the free-spinning vaned member is fitted. 7. The spirometer apparatus of claim 6, wherein the vaned member has a shaft, one end of which is fitted into each said central hub on assembling the pair of members of the flow head unit together. 8. The spirometer apparatus of claim 6, further including a transparent cylindrical wall section in a casing part adjacent the vaned member to enable motion of the vaned member to be observed and measurements obtained. 9. The spirometer apparatus of claim 6, wherein one or both of the pair of members of the detachable flow head unit has a set of vanes therein constructed and arranged to impart a peripheral rotational movement to air blown through the channel. | Vitalograph (Ireland) Limited | A61B005/08
US20080287813 | A1 | 2008-11-20 | Blood Pressure Monitoring Device and Methods for Making and for Using Such a Device | 2004-03-03 | Disclosed are a system and a method for noninvasively and continuously monitoring blood pressure. Also disclosed is a method for making such a device. The system includes a semiconductor chip comprising a transducer array of individual pressure or force sensors and associated circuitry providing control signals to and/or processing signals from these sensors, all of the above integrated in the chip. Also disclosed is a specific sensor structure provided on said chip. The invention further encompasses a system for measuring and/or tracking the blood pressure waveform and for combining the latter with related blood values like the heartbeat, derived from the above or other measuring devices. | 1. A transducer for continuously and noninvasively measuring blood pressure in a blood vessel, having integrated into a single semiconductor chip: an array of transducer or sensor elements,means for reading out data from said sensor elements, andmeans for converting said data into a transmittable form. 2. The blood pressure transducer according to claim 1, further including interface means for transmitting said data from said chip to an external computer. 3. The blood pressure transducer according to claim 1, wherein the semiconductor chip is a CMOS chip. 4. The blood pressure transducer according to claim 1, wherein each sensor element comprises a fluid-filled capacitive sensor having a flexible electrode or membrane and a rigid electrode and a fluid gap connected to an opening. 5. The blood pressure transducer according to claim 1, wherein each sensor element comprises a resistive sensor having strain sensitive resistors on a flexible structure of cross-linked beams a flexible protective membrane and a fluid gap connected to openings. 6. The blood pressure transducer according to claim 1, wherein the sensor elements are arranged in a square array. 7. The blood pressure transducer according to claim 6, wherein the array comprises at least 2×2 sensor elements, preferably 4×4 sensor elements. 8. The blood pressure transducer according to claim 4, wherein the array comprises 2×2 sensor elements arranged with adjacent openings located in the center of said array. 9. The blood pressure transducer according to claim 7, wherein the array of sensor elements is placed close to one end of the semiconductor chip. 10. The blood pressure transducer according to claim 1, wherein the semiconductor chip is part of a sensing device which further includes a power source. 11. The blood pressure transducer according to claim 2, wherein the interface means for transmitting the data from said chip to an external computer is a wireless transmission means. 12. A system for continuously and noninvasively measuring and monitoring blood pressure in a blood vessel, comprising a sensing device including a single semiconductor chip having integrated: an array of sensor elements overlying said blood vessel,means for reading out data from said sensor elements,means for converting said data, andmeans for interfacing with external evaluation means. 13. The measuring and monitoring system according to claim 12, further including a power source on the sensing device and wireless means for interfacing with the external evaluation means. 14. A method for making a transducer for continuously and noninvasively measuring and monitoring blood pressure in a blood vessel, the method including the steps of: fabricating on a single semiconductor chip with conventional semiconductor technology, preferably CMOS technology,an array of sensor elements,means for converting, andmeans for transmitting said data to external evaluation means. 15. A method for using a transducer accordi... | EIDGENOSSISCHE TECHNISCHE HOCHSCHULE ZURICH UNIVERSITAT ZURICH | A61B005/02 A61B005/04
US07445767 | B2 | 2008-11-04 | In-vivo diagnostic method by near infrared radiation | 2006-03-22 | This invention relates to an in-vivo diagnostic method based on near infrared radiation (NIR radiation) that uses water-soluble dyes and their biomolecule adducts, each having specific photophysical and pharmaco-chemical properties, as a contrast medium for fluorescence and transillumination diagnostics in the NIR range, to new dyes and pharmaceuticals containing such dyes. | 1. An in-vivo diagnostic near infrared radiation method comprising: (a) administering a compound of the general formula I BL−(F−Wn)m (I)wherein L represents a number from 1 to 6n represents a number from 0 to 10, andm represents a number from 1 to 100;B is a biological detecting unit having a molecular weight of up to 30,000 that bonds to specific cell populations or selectively to receptors, or accumulates in tissues or tumours, or generally stays in the blood, or is a macromolecule that bonds non-selectively;F represents a cyanine dye of the general formula IIawherein r represents the numbers 0, 1 or 2, wherein, for r=2, the respective L6 and L7 groups that occur in duplicate may be the same or different, L1 to L7 are the same or different, each independently representing a CH or CR group, whereR is a halogen atom, a hydroxy, carboxy, acetoxy, amino, nitro, cyano or sulfonic acid group or an alkyl, alkenyl, hydroxyalkyl, carboxyalkyl, alkoxy, alkoxycarbonyl, sulfoalkyl, alkylamino, dialkylamino or halogenalkyl group containing up to 6 carbon atoms, aryl, alkylaryl, hydroxyaryl, carboxyaryl, sulfoaryl, arylamino, diarylamino, nitroaryl or halogenaryl group containing up to 9 carbon atoms, or where R represents a bond that bonds to another group R and forms a 4−to 6-member ring together with the interspersed groups L1 to L7, or two of L1 to L7 that are linked via a —CO— group,R3 to R12 are the same or different, each independently representing a hydrogen atom, a group B or W, or an alkyl or alkenyl group containing up to 6 carbon atoms or an aryl or aralkyl group optionally carrying an additional group W, or to each pair of adjacent groups R3 to are annealed, with due regard for the interspersed C atoms, 5−to 6−member rings that may be saturated, unsaturated or aromatic, and that optionally carry an additional group R, andX and Y are the same or different, each independently representing an O, S, Se or Te or a —C(CH3)2—, —CH═CH—or —CR13R14- group, with the proviso that X and Y are never S at the same time,where R13 and R14 independently represent a hydrogen atom, a group B or W, or an alkyl or alkenyl group containing up to 6 carbon atoms or an aryl or aralkyl group containing up to 9 carbon atoms, the alkyl, alkenyl, aryl or aralkyl group optionally carrying an additional group W,W represents a hydrophilic group that improves water-solubility, with n-octanol-water distribution coefficient of the compound according to formula I being less than or equal to 2.0 for 1=0; as well as a physiologically tolerable salt, W being selected from the group consisting of (i) a carboxy or sulfonic acid group; (ii) a carboxyalkyl group; and (iii) an alkoxycarbonyl or alkoxyoxoalkyl group containing up to 12 carbon atoms, wherein W takes the position of (i) both R4 and R8 and/or (ii) both R6and R10 and/or (iii) both R11 and R12; and(b) irradiating the compound with light in the near infrared range; and(c) recording the radiation produced by the compound. 2. A m... | Institut fur Diagnostikforschung GmbH an der Freien Universitat Berlin | A61B010/00 A61B005/00 A61B008/00
US07435220 | B2 | 2008-10-14 | Monitoring, system for cardiac surgical operations with cardiopulmonary bypass | 2006-05-11 | A monitoring system for cardiac operations with cardiopulmonary bypass comprising: a processor operatively connected to a heart-lung machine; a pump flow detecting device connected to a pump of the heart-lung machine to continuously measure the pump flow value and send it to the processor; a hematocrit reading device inserted inside the arterial or venous line of the heart-lung machine to continuously measure the blood hematocrit value and to send it to the processor; a data input device to allow the operator to manually input data regarding the arterial oxygen saturation and the arterial oxygen tension; computing means integrated in the processor to compute the oxygen delivery value on the basis of the measured pump flow, the measured hematocrit value, the preset value of arterial oxygen saturation, and the preset value of arterial oxygen tension; and a display connected to the processor to display in real-time the computed oxygen delivery value. | 1. A monitoring system for cardiac operations with cardiopulmonary bypass comprising: a processor operatively connected to a heart-lung machine;a pump flow detecting device connected to a pump of the heart-lung machine to continuously measure the pump flow value (Qp) and send it to the processor;a hematocrit reading device inserted inside the arterial or venous line of the heart-lung machine to continuously measure the blood hematocrit (HCT) value and to send it to the processor;a data input device to allow the operator to manually input data regarding the arterial oxygen saturation (Sat(a)) and the arterial oxygen tension (PaO2);computing means integrated in the processor to compute the oxygen delivery (DO2i) value on the basis of the measured pump flow (Qp), the measured hematocrit (HCT) value, the preset value of arterial oxygen saturation (Sat(a)), and the preset value of arterial oxygen tension (PaO2); anda display connected to the processor to display in real-time the computed oxygen delivery (DO2i) value. 2. A monitoring system according to the claim 1, further comprising: comparing means to compare the above mentioned oxygen delivery (DO2i) value with a threshold value of oxygen delivery (DO2icrit); andan alarm which is triggered when the oxygen delivery (DO2i) value falls below the threshold value of oxygen delivery (DO2icrit). 3. A monitoring system according to the claim 2, wherein the threshold value of oxygen delivery (DO2icrit) is preset by the operator at a value of about 270 ml/min/m2. 4. A monitoring system according to claim 2, wherein the data input device allows the operator to manually input data regarding the weight and height of the patient so that the computing means is able to compute the body surface area (BSA) of the patient. 5. A monitoring system according to claim 1, further comprising a temperature detecting device able to continuously measure the body temperature (T) of the patient and to send the temperature values to the processor, to be subsequently displayed by the display. 6. A monitoring system according to claim 5, further comprising an oxygen delivery threshold computing means, receiving as input the temperature (T) value of the patient as detected. 7. A monitoring system according to claim 1, wherein the processor includes a hemoglobin value computing means able to calculate the hemoglobin (Hb) value from the detected hematocrit (HCT) value, and to send it to the oxygen delivery computing means, to be subsequently displayed in the display. 8. A monitoring system according to claim 1, wherein the processor includes a computing means able to calculate the indexed pump flow (Qpi) from the pump flow (Qp) and the value of body surface area (BSA) of the patient. 9. A monitoring system according to claim 1, wherein the pump flow detecting device includes a Doppler reading cell. | Sorin Group Italia S.r.l. | A61B005/02 A61B005/00 A61M016/00 A62B007/00
US07441342 | B2 | 2008-10-28 | System and method for sizing a heart for treating congestive heart failure | 2005-04-27 | A method for use in reconstructing the left ventricle of the heart of a patient from a diseased state to an appropriate reconstructed state is disclosed. The method includes the steps of determining the body surface area of the patient; determining a characteristic of the left ventricle which is the same in the diseased state as in the reconstructed state; determining an appropriate sizer based on the determined body surface area of the patient and the characteristic of the left ventricle which is the same in the diseased state as in the reconstructed state; identifying akinetic tissue within a heart chamber wall; making an incision through the akinetic tissue in the chamber wall; inserting said appropriate sizer into the chamber through the incision; removing the sizer; and, closing the incision. | 1. A method for use in reconstructing the left ventricle of the heart of a patient from a diseased state to an appropriate reconstructed state, the method comprising the steps of: determining the body surface area of the patient;determining a characteristic of the left ventricle which is the same in the diseased state as in the reconstructed state;determining an appropriate sizer based on the determined body surface area of the patient and the characteristic of the left ventricle which is the same in the diseased state as in the reconstructed state;identifying akinetic tissue within a heart chamber wall;making an incision through the akinetic tissue in the chamber wall;inserting said appropriate sizer into the chamber through the incision;removing the sizer; and,closing the incision. 2. A method according to claim 1 wherein the body surface area of the patient is determined based on the height of the patient and the weight of the patient. 3. A method according to claim 1 wherein said step of determining the appropriate sizer further comprises: determining a first size of the sizer based on body surface area of the patientdetermining a second size of the sizer based on the characteristic of the left ventricle which is the same in the diseased state as in the reformed state; and,determining the volume of the appropriate sizer by selecting the lesser of the first size and the second size. 4. A method for use in reconstructing the left ventricle of the heart of a patient from a diseased state to an appropriate reconstructed state, the method comprising the steps of: determining the body surface area of the patient;determining the short axis left ventricle diameter at papillary muscle tips of the heart of the patient;determining an appropriate sizer based on the determined body surface area of the patient and the short axis left ventricle diameter at papillary muscle tips of the heart of the patient;identifying akinetic tissue within a heart chamber wall;making an incision through the akinetic tissue in the chamber wall;inserting said appropriate sizer into the chamber through the incision;removing the sizer; and,closing the incision. 5. A method according to claim 4 wherein the body surface area of the patient is determined based on the height of the patient and the weight of the patient. 6. A method according to claim 4 wherein said step of determining the appropriate sizer further comprises: determining a first size of the sizer based on body surface area of the patientdetermining a second size, V2, of the sizer based on short axis left ventricle diameter at papillary muscle tips of the heart of the patient; and,determining the volume of the appropriate sizer by selecting the lesser of the first size and the second size. 7. A method according to claim 6 wherein said step of determining a second size, V2, of the sizer based on short axis left ventricle diameter at papillary muscle tips, d, of the heart of the patient comprises calculating V2 a... | CHF Technologies, Inc. | A61B005/103
US07444071 | B2 | 2008-10-28 | Method for diagnosing disease from tongue image | 2005-04-28 | A digital image processing method for a centralized multi-mode medical tongue image inspection system, comprising the steps of: forming at least one service unit (300) for performing tongue record (220) construction; forming a processing unit (320) for tongue record inspection (308); and providing communication links (330) between the service units and the processing unit. | 1. A method for diagnosing diseases by inspecting images of an individual's tongue comprising: forming a first image of said tongue, a color checker, and a portion of said individual's face;forming a second image of said tongue and said portion of said individual's face;performing a color comparison of said portion of said individual's face in said second image to said portion of said individual's face in said first image;matching the colors in said second image to the colors in said first image based on said color comparison; anddiagnosing at least one disease based on changes in said first tongue image and said second tongue image. 2. The method as in claim 1 wherein a color calibration is performed on said first image prior to said color comparison step. 3. The method of claim 1 wherein said diagnosis is aided by comparison to reference images of other individuals tongues. 4. A method for diagnosing diseases by inspecting images of an individual's tongue comprising: forming a first image of said tongue and a portion of said individual's face;forming a second image of said tongue and said portion of said individual's face;performing a color comparison of said portion of said individual's face in said second image to said portion of said individual's face in said first image;matching the colors in said second image to the colors in said first image based on said color comparison; anddiagnosing at least one disease based on changes in said first tongue image and said second tongue image. 5. The method of claim 4 comprising the additional steps of: forming a third image of said tongue, said portion of said individual's face, and a color checker;performing a color calibration on said third image;performing a color comparison of said portion of said individual's face in said third image to said portion of said individual's face in said first and second image; andmatching the colors in said first and second image to the colors in said third image based on said color comparison. 6. The method of claim 4 comprising the additional steps of: transmitting said first image and said second image to a process unit; andperforming said diagnosis step at said process unit. 7. The method of claim 1 further comprising the steps of: retrieving the patient's tongue history records from an image storage facility;reviewing and classifying the first or second image of said tongue and the retrieved history records;transmitting, displaying, or storing the classification, or tongue records. 8. The method of claim 7 wherein the step of classifying the acquired and retrieved tongue images using an inspection mode further comprises the steps of: collecting a plurality of sample tongue images from healthy and unhealthy individuals;extracting relevant color information from said plurality of sample tongue images in a transformed color space;extracting relevant non-color information from said plurality of sample tongue images;training classifiers with the extracted colo... | Carestream Health, Inc. | G03B041/00 G03B043/00 A61B005/00
US07452338 | B2 | 2008-11-18 | Intracoelomic mobile body, and capsule-type ultrasonic endoscope | 2004-11-04 | An intracoelomic mobile body comprising: a capsule body which has a body to be rotated and can be introduced into the human body; and a rotation drive unit which has a rotation shaft for freely rotating the body to be rotated back and forth. The rotation drive unit is arranged in the capsule body in such a way that a first line obtained by linking the positions of center of gravity of cross-sections formed up to a limit of the formation when the capsule body is cut along a plane perpendicular to the rotation shaft does not overlap a second line obtained by extending the rotation shaft as the same line. | 1. An intracoelomic mobile body comprising: a swallowable capsule body which has an internal body to be rotated, the swallowable capsule body can be introduced into the human body and having first and second ends offset in a longitudinal direction; anda rotation drive unit which has a rotation shaft for freely rotating the internal body to be rotated back and forth, the rotation drive unit being arranged in the swallowable capsule body between the first and second ends of the swallowable capsule body in such a way that a first line obtained by linking the positions of center of gravity of cross-sections formed up to limit of the formation when the capsule body is cut along a plane perpendicular to the rotation shaft does not overlap a second line obtained by extending the rotation shaft as the same line. 2. The intracoelomic mobile body according to claim 1, wherein the internal body to be rotated is an information acquisition unit for acquiring biological information. 3. The intracoelomic mobile body according to claim 2, wherein the information acquisition unit is an ultrasonic transducer unit for transmitting and receiving ultrasonic waves. 4. The intracoelomic mobile body according to claim 3, wherein the first line and the second line do not overlap at all. 5. The intracoelomic mobile body according to claim 2, wherein the information acquisition unit is an image pickup unit for picking up optical images. 6. The intracoelomic mobile body according to claim 3, wherein the first line is a straight line. 7. The intracoelomic mobile body according to claim 6, wherein the first line and the second line are not parallel. 8. The intracoelomic mobile body according to claim 6, wherein the first line and the second line are substantially parallel. 9. The intracoelomic mobile body according to claim 8, wherein the distance between the first line and the second line is less than about 10 mm. 10. The intracoelomic mobile body according to claim 3, wherein the first line is formed from a combination of a straight line and a curved line. 11. The intracoelomic mobile body according to claim 10, wherein the first line has at least two straight lines and at least one of the straight lines overlaps the second line. 12. The intracoelomic mobile body according to claim 3, wherein the rotation drive unit is arranged in the capsule body in such a way that the first line and the second line form an angle. 13. The intracoelomic mobile body according to claim 3, wherein the rotation drive unit is arranged in the capsule body in such a way that the second line is eccentric from the first line. 14. A capsule-type ultrasonic endoscope comprising: a swallowable capsule body which has an internal body to be rotated, the swallowable capsule body can be introduced into the human body and having first and second ends offset in a longitudinal direction; anda rotation drive unit which has a rotation shaft for freely rotating the internal body to be rotated back a... | Olympus Corporation | A61B005/117
US07463921 | B2 | 2008-12-09 | Method and system for analyzing an electrocardiographic signal | 2002-08-13 | A method for analyzing an ECG signal includes obtaining a measurement of an area based repolarization interval from at least one beat in the electrocardiogram signal and detecting an altered ventricular repolarization based on the obtained measured area based repolarization interval. | 1. A method for analyzing an effect of a pharmacological agent on an electrocardiogram (ECG) signal, comprising: obtaining a first ECG signal having at least one beat, the beat having a first repolarization interval represented by a portion of the first ECG signal from a first T-wave starting point to a first T-wave ending point;quantifying an area under the first ECG signal during the first repolarization interval;quantifying a first time to reach at least one percentage of the area under the first ECG signal during the first repolarization interval;administering the pharmacological agent;after administering the pharmacological agent, obtaining a second ECG signal having at least one beat, the beat having a second repolarization interval represented by a portion of the second ECG signal from a second T-wave starting point to a second T-wave ending point;quantifying an area under the second ECG signal during the second repolarization interval;quantifying a second time to reach the at least one percentage of the area under the second ECG signal during the second repolarization interval;detecting an altered ventricular repolarization based on the quantified first and second times. 2. The method of claim 1, wherein: the first T-wave starting point comprises a point during the at least one beat of the first ECG signal at a first time after a first R peak location of the first ECG signal; andthe second T-wave starting point comprises a point during the at least one beat of the second ECG signal at a second time after a second R peak location of the second ECG signal. 3. The method of claim 2, wherein the first time after the first R peak location is approximately equal to the second time after the second R peak location. 4. The method of claim 2, wherein the first time after the first R peak location is approximately equal to 100 msec. 5. The method of claim 2, wherein the first time after the first R peak location is selected to avoid QRS complex components. 6. The method of claim 1, wherein: the first T-wave ending point comprises a point during the at least one beat of the first ECG signal where a slope of the first repolarization interval is equal to or passes below a first threshold; andthe second T-wave ending point comprises a point during the at least one beat of the second ECG signal where a slope of the second repolarization interval is equal to or passes below a second threshold. 7. The method of claim 6, wherein the first threshold is substantially equal to the second threshold. 8. The method of claim 1, wherein: the area under the first repolarization interval comprises a total area under the first repolarization interval; andthe area under the second repolarization interval comprises a total area under the second repolarization interval. 9. The method of claim 1, wherein at least one of the area under the first repolarization interval and the area under the second repolarization interval are determined using integration. 10. The... | University of Rochester | A61B005/0402
US07466850 | B2 | 2008-12-16 | Computerized method and apparatus for iterative calculation of the general linear model using only substantially orthogonal model functions | 2006-03-10 | In a method and apparatus for incremental calculation of a general linear model given intermittent correlation of the model functions, data composed of a number of data sets with a number of different random samples are processed. For each independent random sample contained in the data set, a dependency on an order quantity is compared with the dependency on the order quantity in model functions contained in a model matrix G using the general linear model in order to check the occurrence of specific characteristics in the dependency of the random sample on the order quantity. The calculation of the general linear model ensues incrementally from data set to data set. Before the calculations for the data of a data set, a check is made as to whether the model functions contained in the model matrix G exhibit orthogonal portions to a sufficient degree in a segment (represented by the data set) of the dependency on the order quantity for the calculations. Given a degree of parallelism that is too high, one or more model functions of the model matrix G that interfere with the orthogonality are masked out to a predeterminable degree in the calculations for this data set, in order to achieve the sufficient degree of orthogonal portions. An incremental calculation of the general linear model with corresponding output of intermediate results is thereby enabled, even given the presence of intermittent correlating model functions in the model matrix. | 1. A computerized method for processing data comprising a plurality of data sets respectively containing a plurality of different random samples, said method comprising the steps of: for each individual random sample contained in a data set, automatically electronically checking for occurrence of specific characteristics of a dependency of that random sample on an order quantity by comparing, in a predetermined sequence of said data sets, the dependency of that random sample on said order quantity with a dependency on said order quantity of model functions contained in a model matrix, using the general linear model, to obtain an intermediate comparison result;before said checking of each individual random sample, automatically electronically determining whether said model functions in said model matrix exhibit, within a time segment represented by the data set containing the random sample to be checked, a degree of orthogonality sufficient to cause said model matrix to represent a determinable equation system, and masking out, to a predetermined degree, model functions that interfere with said sufficient degree of orthogonality from contributing to the comparison; andstoring each intermediate comparison result as it occurs and making the stored intermediate comparison results, and a final result comprising the respective intermediate comparison results for all of said data sets, electronically available. 2. A method as claimed in claim 1 comprising masking out said model functions that interfere with said sufficient degree of orthogonality by setting the interfering model functions to have a lower degree of contribution to said comparison than model functions that do not interfere with said sufficient degree of orthogonality. 3. A method as claimed in claim 2 comprising setting said interfering model functions to have a lower predetermined degree of contribution to said comparison by augmenting a diagonal of a quadratic matrix M, generated from said comparisons, at respective positions associated with the interfering model functions, with M=GT·G, wherein G is said model matrix. 4. A method as claimed in claim 1 comprising masking out said interfering model functions by setting said predetermined degree of contribution as a function of a degree of orthogonal portions of said interfering model function with another model function with which the interfering model function exhibits a highest degree of parallelism. 5. A method as claimed in claim 1 wherein the step of checking said sufficient degree of orthogonality comprises automatically electronically calculating a correlation function between respective model functions. 6. A method as claimed in claim 5 comprising comparing said correlation coefficient to a predetermined value and determining that the correlated model functions have said sufficient degree of orthogonality if said correlation coefficient is below said predetermined value. 7. A method as claimed in claim 5 comprising automatic... | Siemens Aktiengesellschaft | G06K009/00 A61B005/05
US08527028 | B2 | 2013-09-03 | Harness with sensors | 2007-05-16 | A garment for ambulatory, physiological monitoring of a patient includes a belt, having first and second end portion with closures at the end portions to wrap around a user's chest, a strap having a first end coupled to a portion of the belt with the strap having a second end, a pair of shoulder strap portions each shoulder strap portion having a first end coupled together at the second end of the strap and a second end, and a back portion that joins the second ends of the pair of shoulder strap portions, with at least one of the belt, strap portions and back portion having an accommodation for carrying a sensor. Other embodiments are described. | 1. A garment comprises: a belt, having first and second end portions with closures at the end portions to wrap around a user's chest, and fastening mechanism disposed between the end portions;a Y-shaped chest band of unitary construction, the chest band having a front strap portion having a substantially constant width and having a length extending from a first end that is configured to couple to the fastening mechanism on the belt to a second end where the front strap portion integrally joins with first ends of a pair of shoulder strap portions of the Y-shaped chest band, each shoulder strap portion further having a second end and a substantially constant width; anda back portion, the back portion having a broad region that is substantially greater in width in comparison with the width of the shoulder strap portions and less than a width of the user's back and configured to rest against the back of the user, with upper portions of the broad region joining the second ends of the pair of shoulder strap portions, and with a lower portion of the back portion integrally joined with the belt, with at least one of the belt, shoulder strap portions and back portion having an accommodation for carrying a sensor. 2. The garment of claim 1 wherein the pair of shoulder strap portions of the Y-shaped chest band are configured to rest against a portion of the user's torso where the user's neck meets the user's shoulders. 3. The garment of claim 1 wherein the second end of the front strap portion and the first ends of the pair of shoulder strap portions are coupled together by stacked fastener mechanisms. 4. The garment of claim 1 wherein the back portion is unitary in construction with the pair of shoulder strap portions. 5. The garment of claim 1 wherein the back portion is coupled to the pair of shoulder strap portions by a fastener. 6. The garment of claim 1 wherein the back portion, the front strap portion, and the pair of shoulder strap portions are comprised of synthetic rubbers based on polychloroprene material. 7. The garment of claim 1 wherein the accommodation for the sensor is disposed in the middle of the back portion of the garment, and the garment further comprises second and third accommodations disposed in each of the shoulder strap portions and a fourth accommodation disposed in the belt of the garment. 8. The garment of claim 1 wherein the accommodation is a pouch provided in the garment. 9. The garment of claim 1 wherein the accommodation is a snap that mates with a corresponding snap on the sensor. 10. The garment of claim 1 wherein the accommodation is provided in one of the belt, the pair of shoulder strap portions and the back portion and the garment further comprises: at least one conductor to deliver a signal from the sensor to a processing circuit. 11. The garment of claim 3 wherein the fastener mechanisms are selected from the group consisting of snaps and hook and loop fasteners. 12. The garment of claim 1 wherein the ac... | Medicomp, Inc. | A61B005/0408 A61B005/0416
US08423109 | B2 | 2013-04-16 | Method for enhancing pulse oximery calculations in the presence of correlated artifacts | 2008-06-20 | A method for determining a physiological parameter in the presence of correlated artifact, including obtaining two digital waveforms, x and y, the waveforms being representative of the absorption of two wavelengths of electromagnetic energy received from a blood-perfused tissue, and where each of the waveforms has a component corresponding to a plethysmographic waveform and a component corresponding to the correlated artifact; calculating several weighted difference waveforms of the form x−R*y, where R is a multiplier, by varying R over a range; evaluating the several weighted difference waveforms using a shape characteristic of the weighted difference waveform; identifying a weighted difference waveform most closely representative of and one most different from the plethysmographic waveform; determining a pleth-based physiological parameter using the waveform most closely representative of the plethysmographic waveform; determining at least one artifact-based physiological parameter using the waveform most different from the plethysmographic waveform; and rejecting other possible candidate values for the pleth-based physiological parameter using the artifact-based physiological parameter. | 1. A method comprising: receiving two waveforms corresponding to two different wavelengths of light from a patient;combining the two waveforms to form a plurality of weighted difference waveforms, wherein the plurality of weighted difference waveforms vary from one another by a value of a multiplier;identifying one of the weighted difference waveforms from the plurality of weighted difference waveforms using a characteristic of one or more of the plurality of weighted difference waveforms; andcalculating a physiological parameter based at least in part on the identified weighted difference waveform. 2. The method of claim 1, wherein the physiological parameter comprises a pulse rate. 3. The method of claim 1, wherein the physiological parameter comprises an oxygen saturation value. 4. The method of claim 1, wherein identifying one of the weighted difference waveforms from the plurality of weighted difference waveforms comprises identifying the waveform that most closely resembles a plethysmographic waveform. 5. The method of claim 1, wherein identifying one of the weighted difference waveforms from the plurality of weighted difference waveforms comprises identifying the waveform that least closely resembles a plethysmographic waveform. 6. The method of claim 1, comprising identifying another one of the weighted difference waveforms from the plurality of weighted difference waveforms using another characteristic of one or more of the plurality of weighted difference waveforms, and wherein calculating the physiological parameter is based at least in part on the another identified weighted difference waveform. 7. The method of claim 1, wherein the characteristic of one or more of the plurality of weighted difference waveforms comprises at least one of a measure of complexity of one or more of the plurality of weighted difference waveforms or a measure of skew of one or more of the plurality of weighted difference waveforms. 8. A pulse oximeter, comprising: a processor capable of: receiving two waveforms corresponding to different wavelengths of light from a patient;combining the two received waveforms to form a plurality of weighted difference waveforms, wherein the plurality of weighted difference waveforms vary from one another by a value of a multiplier;identifying one of the weighted difference waveforms from the plurality of weighted difference waveforms using a characteristic of one or more of the plurality of weighted difference waveforms; andcalculating a physiological parameter based at least in part on the identified weighted difference waveform. 9. The oximeter of claim 8, wherein identifying one of the weighted difference waveforms from the plurality of weighted difference waveforms comprises identifying the waveform that most closely resembles a plethysmographic waveform. 10. The oximeter of claim 8, wherein identifying one of the weighted difference waveforms comprises identifying the waveform that least closely resembles a pl... | Covidien LP | A61B005/1455
US08195274 | B2 | 2012-06-05 | Mapping of vascular perfusion territories | 2008-04-28 | Techniques, systems and computer program products are disclosed for mapping of vascular perfusion territories by placing one or more blood vessels of the vascular perfusion territories in a tag condition and others in a control condition by applying a train of pseudo-continuous radio frequency tagging pulses. In addition, an encoding scheme is applied to fully invert or relax the blood vessels of the vascular perfusion territories. Also, a tagging efficiency is measured for each blood vessel based on the applied encoding scheme. Further, the vascular perfusion territories are separated by using the measured tagging efficiency in a decoding process. | 1. A method for enabling mapping vascular perfusion territories comprising: applying a train of pseudo-continuous radio frequency tagging pulses to place two or more blood vessels of vascular perfusion territories in a tag condition and others in a control condition;applying a differential encoding scheme, by differentially tagging the two or more blood vessels, to fully invert or relax the blood vessels of the vascular perfusion territories;measuring a tagging efficiency for each blood vessel based on the applied encoding scheme, wherein the tagging efficiency for each blood vessel is proportional to a difference in magnetization in control and tag conditions; andseparating the vascular perfusion territories by using the measured tagging efficiency in a decoding process. 2. The method as in claim 1, wherein selecting and said differentially tagging the two or more of the blood vessels comprises using pseudo-continuous arterial spin labeling. 3. The method as in claim 2, further comprising applying pulses of magnetic field gradients across a tagging plane between the train of pseudo-continuous radio frequency pulses. 4. The method as in claim 2, further comprising differentially encoding two or more of the blood vessels within the tagging plane by using a modified pseudo-continuous arterial spin labeling. 5. The method as in claim 1, further comprising applying a single labeling gradient waveform in a direction of blood flow with non-zero mean for the tag and control conditions. 6. The method as in claim 5, further comprising applying additional pulses of magnetic field gradients perpendicular to the single labeling gradient waveform to generate phase shifts between the blood vessels. 7. The method as in claim 1, wherein applying the encoding scheme comprises using a Hadamard encoding scheme. 8. The method as in claim 1, wherein measuring the tagging efficiency comprises measuring the tagging efficiency on a per-vessel basis to improve the decoding process. 9. The method as in claim 1, further comprising quantitatively measuring a perfusion of each vascular perfusion territory. 10. The method as in claim 1, further comprising measuring and quantifying a contribution of two or more of the blood vessels to the perfusion of each voxel. 11. The method as in claim 1, wherein applying the encoding scheme comprises optimizing a signal-to-noise ratio. 12. The method as in claim 1, further comprising mapping the measured tagging efficiencies to blood vessel coordinates. 13. The method as in claim 1, wherein measuring the tagging efficiencies comprises measuring the tagging efficiencies by clustering. 14. A magnetic resonance imaging (MRI) system, comprising: an MRI imaging module to acquire MRI image; anda controller which controls the MRI imaging module and configured to perform operations comprising: applying a train of pseudo-continuous radio frequency tagging pulses to place two or more blood vessels of vascular perfusion territories in a... | The Regents of the University of California | A61B005/05 G01V003/00
US08298148 | B2 | 2012-10-30 | Integrated heart monitoring device and method of using same | 2008-05-12 | A device for monitoring the heart of a patient including a housing, a computing device, an optical sensor adapted to provide signals to the computing device indicative of a distance from the optical sensor to a vessel carrying blood, as well a diameter of the vessel, a Doppler sensor adapted to provide signals to the computing device indicative of a velocity of the blood through the vessel, and an ECG sensor adapted to provide signals to the computing device indicative of a plurality of electrical stimuli that cause the heart to pump. The computing device uses signals from the optical sensor, the Doppler sensor, and the ECG sensor to compute parameters including oxygen saturation of the blood, blood flow, blood pressure, heart rate, and cardiac output. | 1. A device for monitoring the heart of a patient, including: a housing;a computing device mounted within the housing;an optical sensor mounted within the housing and adapted to provide signals to the computing device sufficient to compute a distance from the optical sensor to a vessel canying blood, as well a diameter of the vessel;a Doppler sensor mounted within the housing and adapted to direct ultrasonic waves based on the distance and diameter, and to provide signals to the computing device indicative of a velocity of the blood through the vessel; andan ECG sensor mounted within the housing and adapted to provide signals to the computing device indicative of a plurality of electrical stimuli that cause the heart to pump;the computing device using signals from the optical sensor, the Doppler sensor, and the ECG sensor to compute parameters including oxygen saturation of the blood, blood flow, blood pressure, heart rate, and cardiac output. 2. The device of claim 1, further including a temperature sensor mounted within the housing and adapted to provide signals to the computing device indicative of a temperature of the patient. 3. The device of claim 1, further including a communication device mounted within the housing, the communication device being coupled to the computing device and configured to transmit information relating to the parameters. 4. The device of claim 1, wherein the housing is configured for subcutaneous implantation. 5. The device of claim 1, wherein timing of the operation of the Doppler sensor is based on the signals from the ECG sensor. 6. The device of claim 1, wherein the optical sensor includes an emitter array and a detector array. 7. The device of claim 1, further including a connector mounted to the housing, the connector being adapted to couple to ECG leads. 8. The device of claim 7, wherein the connector is configured to enable one of programming of the computing device and downloading information relating to the parameters. 9. A monitoring device for measuring parameters indicative of behavior of the heart of a patient, the device including: a plurality of sensors mounted within an implantable housing, the plurality of sensors including an optical sensor for measuring oxygen saturation of blood flowing through a vessel, the optical sensor further configured to provide signals sufficient compute a distance from the optical sensor to a vessel carrying blood and a diameter of the vessel,a Doppler sensor configured to direct ultrasonic waves to the vessel based on the distance and diameter and to receive corresponding frequency-shifted signals for measuring a velocity of the blood,an ECG sensor for measuring electrical activity of the heart, anda temperature sensor for measuring a temperature of the patient;a communication device mounted within the housing, the communication device being configured to wirelessly transmit information relating to the measured parameters; anda computing device that executes a ... | Cardio Art Technologies Ltd | A61B008/00 A61B005/02
US08275445 | B2 | 2012-09-25 | System and method for determining the position of an instrument | 2008-04-22 | The invention relates to a system and a method for a system featuring an electromagnetic position detection system and an instrument of which the position is able to be determined in at least one electromagnetic field created with the electromagnetic position detection system, with a device distorting the electromagnetic field being located at least partly in the electromagnetic field or a device distorting the electromagnetic field being able to be brought at least partly into the electromagnetic field, and with the position of the instrument being determined taking into account the position of the device or of a part of the device in the electromagnet field and the distortion of the electromagnetic field associated with it. | 1. A system for determining a position of an instrument in a procedure, comprising: an electromagnetic position detection system comprising a transmitter that creates an electromagnetic field;an x-ray device having a C-arm, an x-ray source, and an x-ray detector that distorts the electromagnetic field;a position sensor of the electromagnetic position detection system arranged on the C-arm that determines a position of the x-ray device in the electromagnetic field; anda further position detection system for determining positions of the C-arm, the x-ray source, and the x-ray detector, wherein the further position detection system comprises: optical markers arranged on the C-arm, the x-ray source, and the x-ray detector, anda camera that records the positions of the optical markers,wherein the electromagnetic position detection system is configured to determine the position of the instrument based on the position of the x-ray device in the electromagnetic field and the distortion of the electromagnetic field. 2. The system as claimed in claim 1, wherein the instrument comprises a sensor. 3. The system as claimed in claim 1, wherein the instrument is a catheter. 4. The system as claimed in claim 1, wherein the x-ray device and the electromagnetic position detection system are linked to each other via an interface for transmitting data of the position of the x-ray device to the electromagnetic position detection system. 5. The system as claimed in claim 1, wherein the distortion of the electromagnetic field is determined analytically or numerically by a field calculation based on the position of the x-ray device, material of the device, or dimensions of the x-ray device. 6. The system as claimed in claim 1, wherein the instrument is located at least partly in the electromagnetic field or able to be brought at least partly into the electromagnetic field. 7. The system as claimed in claim 1, wherein the positions of the C-arm, the x-ray source, and the x-ray detector are used in determining the position of the instrument. 8. The system as claimed in claim 1, wherein the electromagnetic position detection system and the x-ray device are calibrated for the distortion of the electromagnetic field. 9. The system as claimed in claim 8, wherein resulting displacements are determined for the calibration. 10. The system as claimed in claim 8, wherein data of the calibration is interpolated for determining the position of the instrument. 11. A method for determining a position of an instrument in a procedure, comprising: creating an electromagnetic field by an electromagnetic position detection system comprising a transmitter;distorting the electromagnetic field by an x-ray device having a C-arm, an x-ray source, and an x-ray detector;determining a position of the x-ray device in the electromagnetic field by a position sensor of the electromagnetic position detection system arranged on the C-arm;determining the position of the instrument based on the p... | Siemens Aktiengesellschaft | A61B005/05
US07957795 | B2 | 2011-06-07 | Abdominal impedance measurement apparatus | 2008-03-26 | An abdominal impedance measurement apparatus includes a plurality of electrodes for measuring an abdominal impedance of a human subject, and includes an electrode supporting member for supporting the electrodes in such a manner that the electrodes protrude from the electrode supporting member. The electrode supporting member includes a plurality of segments aligned in a direction, the electrodes being respectively mounted on different segments, neighboring segments being connected rotatably with each other. The electrode supporting member further includes rotation-angle restricting parts for restricting relative rotation-angles between neighboring segments. | 1. An abdominal impedance measurement apparatus comprising: a plurality of electrodes for measuring an abdominal impedance of a human subject; andan electrode supporting member for supporting the electrodes in such a manner that the electrodes protrude from the electrode supporting member, the electrode supporting member comprising a plurality of segments aligned in a direction, each single electrode being mounted on a segment that is different from every other segment on which another electrode is mounted, such that no two electrodes are mounted on the same segment, neighboring segments being connected rotatably with each other, the electrode supporting member comprising rotation-angle restricting parts for restricting relative rotation-angles between neighboring segments,wherein the plurality of electrodes comprise two current-supplying electrodes adapted to contact an abdomen of the human subject for applying an electrical current to the abdomen of the human subject and two voltage-measurement electrodes adapted to be brought into contact with two parts of the abdomen of the human subject for measuring a voltage between the parts with which the voltage-measurement electrodes are in contact, the current-supplying electrodes being mounted on two of the segments of the electrode supporting member, the voltage-measurement electrodes being mounted on two of the segments of the electrode supporting member that are different from the segments on which the current-supplying electrodes are mounted,wherein at least one segment on which no electrode is mounted exists between the segment on which one of the current-supplying electrodes is mounted and the segment on which one of the voltage-measurement electrodes is mounted, andwherein the rotation-angle restricting parts of the electrode supporting member are provided such that the rotation-angle restricting parts located nearer a center of the electrode supporting member restrict a relative rotation-angle between neighboring segments within a larger range and such that the rotation-angle restricting parts located farther from the center of the electrode supporting member restrict a relative rotation-angle between neighboring segments within a smaller range. 2. The abdominal impedance measurement apparatus of claim 1, wherein segments each on which no electrode is mounted exist at all positions between neighboring segments on which the electrodes are mounted. 3. The abdominal impedance measurement apparatus of claim 1, wherein neighboring segments are rotatably jointed by pins, and the rotation-angle restricting parts are adapted to restrict relative rotation-angles about the pins between neighboring segments. 4. The abdominal impedance measurement apparatus of claim 1, wherein each of the electrodes comprises a cylindrical part protruding from the electrode supporting member and a convex curved surface adapted to be brought into contact with the abdomen of the human subject, the convex curved surface ... | Tanita Corporation | A61B005/05
US08608655 | B2 | 2013-12-17 | Sleep evaluation device | 2008-06-06 | A sleep evaluation device 1 is provided with a respiration detection unit 7 that detects changes in respiration and outputs respiration signals. CPU 6 generates six primary parameters indicating sleep states based on the respiration signals and four secondary parameters based on the primary parameters. The secondary parameters are linearly independent of one another. Evaluation unit 20 computes a sleep score indicating the degree of the quality of sleep by multiplying each of the four secondary parameters by second coefficients and totalizing results of the multiplications. Results of the evaluations are displayed on display unit 4. | 1. A sleep evaluation system, comprising: a biometric signal unit for measuring a state of a human subject, and for outputting a biometric signal; anda central processing unit coupled to the biometric signal unit, comprising:a primary parameter computation unit for generating, based on the biometric signal, n number of primary parameters indicating a state of sleep, n being a natural number that is at least two;a secondary parameter computation unit for generating m number of secondary parameters by multiplying each of the n number of primary parameters by an eigenvector computed based on a correlation matrix, with the eigenvector being a first coefficient, and totalizing results of the multiplications, m being a natural number satisfying n >m and a correlation coefficient of any two of the secondary parameters being smaller than a correlation coefficient of any two of the primary parameters; andan evaluation unit for computing a sleep index indicating a quality of sleep based on a result of computation obtained by multiplying each of the secondary parameters by a second coefficient and totalizing results of the multiplications. 2. The system of claim 1, wherein the second coefficient is a factor loading obtained by multiplying the first coefficient by the square root of an eigenvalue. 3. The system of claim 1, wherein the primary parameter computation unit comprises a program code portion for generating , based on the biometric signals, as the n number of primary parameters, at least three of sleep latency, sleep efficiency, mid-arousal number, deep sleep latency, deep sleep time, movement number, total bed time, out-of-bed latency, sleep time, total sleep time, mid-arousal period, REM sleep latency, light sleep period, REM sleep period, sleep stage change number, light sleep occurrence number, REM sleep occurrence number, deep sleep occurrence number, REM sleep duration, REM sleep interval, REM sleep cycle, sleep cycle, the ratio of light sleep in a first and a second half of sleep, the ratio of REM sleep in the first and the second half of the sleep, and the ratio of deep sleep in the first half and the second half of the sleep. 4. The system of claim 3, wherein the primary parameter computation unit comprises a program code portion for generating, as the n number of primary parameters, the sleep latency, the mid-arousal number, the deep sleep latency, the deep sleep time, and the movement number, and wherein the secondary parameter computation unit comprises a program code portion for selecting, as the m number of secondary parameters, four secondary parameters. 5. The system of claim 1, further comprising an inputter for receiving an input of characteristics data of the human subject being examined, wherein the primary parameter computation unit stores a plurality of the first coefficients each corresponding to each of plural groups of populations categorized in accordance with the characteristics of human subjects and computes the prima... | Tanita Corporation | A61B005/00
US08565504 | B2 | 2013-10-22 | Ultrasonic image processing apparatus and ultrasonic image processing method | 2008-04-25 | Three slices, i.e., Basal, Mid, and Apical slices, which correspond to clinically useful ASE segmentation are designated, and the positions of the three slices are tracked through at least one cardiac cycle by performing three-dimensional speckle tracking in the remaining time phases. Three C-mode projection images concerning the tracked positions are reconstructed. In addition, arbitrary myocardial wall motion parameters at the tracked positions are computed and displayed upon being superimposed on C-mode images or projected/displayed on a polar map. As a C-mode projection image method, one of the following techniques can be used detecting and projecting only movement components perpendicular to slices determined in an initial time phase, detecting and projecting average movement components of the respective walls, and tracking and projecting each myocardial position. The obtained C-mode images are simultaneously displayed together with markers indicating the positions of long-axis images and C-mode images. | 1. An ultrasonic image processing apparatus comprising: a storage unit which stores volume data acquired for each time phase concerning periodic motion of an object to be examined which periodically moves by scanning the object with ultrasonic waves;a setting unit which sets an arbitrary slice in volume data corresponding to a predetermined time phase;a movement vector processing unit which determines a movement amount of the object between two different time phases;a tracking unit which moves a plane corresponding to the arbitrary slice in said each volume data corresponding to remaining time phases concerning the periodic motion by executing tracking processing of tracking a temporal change in a position of the arbitrary slice in the predetermined time phase, in the tracking processing, a position of each slice being determined based on a plurality of motion vectors at points on a slice corresponding to a preceding time phase;an image generating unit which generates a first ultrasonic image in said each time phase on the basis of data corresponding to the plane corresponding to the arbitrary slice in said each time phase; anda display unit which displays the first ultrasonic image for each time phase,wherein the tracking unit computes a movement vector concerning the plane corresponding to the arbitrary slice in the predetermined time phase, and executes the tracking processing by using the movement vector in the predetermined time phase and the arbitrary slice in the predetermined time phase. 2. An apparatus according to claim 1, wherein the tracking unit sets an area to which a plane corresponding to the arbitrary slice in an immediately preceding or succeeding time phase is moved by using the movement vector in the immediately preceding or succeeding time phase as a plane corresponding to the arbitrary slice in the time phase, for each of remaining time phases ti (where i is an integer satisfying 1≦i≦n and n is an integer satisfying n≧2) concerning the periodic motion when the predetermined time phase is represented by tj (i≠j), calculates a movement vector concerning the plane corresponding to the arbitrary slice in the time phase, and moves an arbitrary slice in an immediately preceding or succeeding time phase by using the arbitrary slice in the time phase and the movement vector concerning the arbitrary slice in the time phase. 3. An apparatus according to claim 1, wherein the tracking unit calculates a three-dimensional movement vector at each position on the arbitrary slice for each time phase, and calculates an average of projection components of a three-dimensional movement vector at said each position with respect to a normal line to the arbitrary slice as a movement vector concerning the arbitrary slice for each time phase. 4. An apparatus according to claim 1, wherein the tracking unit calculates a three-dimensional movement vector at each position on the arbitrary slice for each time phase, and calculates an average of a three... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporationq | G06K009/20 A61B008/00
US09788750 | B2 | 2017-10-17 | Seizure prediction | 2007-04-30 | Seizure prediction systems and methods include measuring impedance within a brain of a patient to determine whether the brain is in a state indicative of a possibility of an onset of a seizure. In some embodiments, the measured impedance is compared to a predetermined threshold in order to determine whether the brain is in a state indicative of a possibility of a seizure. In other embodiments, a trend of the impedance measurements is correlated to a template. In other embodiments, a frequency component of a waveform of the impedance measurement amplitudes over time is correlated to frequency components of a template waveform. Upon detecting a state in which a seizure is likely to occur, a seizure indicator may be generated, which, in some embodiments, may be used to activate therapy delivery to the patient or, in other embodiments, activate an alarm. | 1. A method comprising: receiving, with a processor, impedance measurements indicative of an impedance of a brain of a patient, wherein the impedance of the brain is indicative of a physiological state of the brain;determining, with the processor, a trend in the impedance measurements over time;determining, with the processor, that the brain is in a state indicative of a possibility of a future seizure based on the trend; andcontrolling, with the processor, delivery of therapy to the patient based on determining that the brain is in the state indicative of the possibility of the future seizure. 2. The method of claim 1, further comprising measuring the impedance of the brain of the patient with an impedance sensing module, wherein measuring the impedance of the brain of the patient comprises delivering a stimulation current in a range of about 500 nanoamps to about 10 microamps across tissue in one or more regions of the brain of the patient. 3. The method of claim 1, further comprising, with the processor, comparing an amplitude of at least one of the impedance measurements to a predetermined threshold and determining that the brain is in the state indicative of the possibility of the future seizure based on the comparison. 4. The method of claim 3, wherein comparing the amplitude of the at least one of the impedance measurements to the predetermined threshold value comprises comparing an average amplitude of a plurality of the impedance measurements over a period of time to the predetermined threshold value. 5. The method of claim 1, further comprising: with the processor, comparing the trend to a template; andbased on the comparison of the trend to the template, determining the trend substantially correlates to the template, wherein determining that the brain is in the state indicative of the possibility of the future seizure comprises determining that the brain is in the state indicative of the possibility of a future seizure based on determining the trend substantially correlates to the template. 6. The method of claim 5, wherein determining the trend substantially correlates to the template comprises determining a first frequency component of the trend substantially correlates to a second frequency component of the template. 7. The method of claim 1, wherein the trend indicates a rate of change of the impedance measurements over time. 8. The method of claim 1, wherein controlling delivery of therapy comprises at least one of controlling delivery of a pharmaceutical agent or controlling delivery of electrical stimulation to tissue of the patient. 9. The method of claim 1, further comprising, with the processor, controlling the therapy to inhibit an onset of the future seizure. 10. The method of claim 1, further comprising, with the processor, controlling the therapy to reduce severity or duration of the future seizure. 11. The method of claim 1, wherein the state indicative of the possibility of a future seizure is a first state, t... | Medtronic, Inc. | A61B005/04 A61B005/053 A61B005/0478 A61B005/00 A61B005/024 A61B005/0402 A61B005/0488 A61N001/08 A61N001/36
US09597029 | B2 | 2017-03-21 | System and method for remotely evaluating patient compliance status | 2007-06-19 | A system and method for remotely evaluating patient compliance status is presented. Compliance to a course of treatment is periodically checked through qualitative inquiry of a patient under remote management. Physiometry of the patient, that indicates emotional state, is remotely monitored contemporaneously with the qualitative state. Patient status is evaluated, including at least one of corroborating the compliance and identifying an emotional imbalance, based upon the indicated emotional state of the patient. | 1. A system for remotely evaluating compliance of a patient to a prescribed course of treatment, comprising: an implantable medical device (IMD) implanted in the patient, the IMD comprising one or more sensors;a query module configured to check compliance to the course of treatment of a patient under remote management on a periodic and regular basis,the query module comprising a communicator configured to provide an interactive user interface for performing a qualitative inquiry of patient compliance and emotional state,the communicator further configured to receive responses to the qualitative inquiry;wherein the one or more sensors are configured to remotely monitor health related physiometry of the patient contemporaneously with the qualitative inquiry;a data storage device configured to store physiometry and the responses to the qualitative inquiry; andan evaluation module configured to evaluate patient compliance based upon a comparison of the physiometry to baseline physiometric data and configured to assess contemporaneously received responses to the qualitative inquiry to determine an emotional state and to corroborate patient compliance with the determined emotional state,wherein the evaluation module is further configured to identify emotional imbalance based upon the comparison of the physiometry to baseline physiometric data and the assessment of the contemporaneously received responses to the qualitative inquiry, andwherein the evaluation module is further configured to output information regarding at least one of corroborating compliance, identifying an emotional imbalance, or a combination thereof, which is related to patient compliance toward the prescribed course of treatment. 2. A system according to claim 1, wherein the data storage device further comprises: stored factors bearing on patient compliance selected from the group consisting of medication, dietary restrictions, lifestyle modifications; and combinations thereof. 3. A system according to claim 1, wherein the evaluation module further comprises: a relative scale for determining emotional stability along a continuum or discrete set of emotional levels. 4. A system according to claim 1, wherein the interactive user interface comprises one or more of: a strictly formatted qualitative inquiry, comprising questions and answers;an intermediately formatted qualitative inquiry, comprising one or more of visual analog scales, five factor wellness scales, comfort scales, pain scales, stress scales, and spiritual wellness scales;a loosely formatted qualitative inquiry, comprising one or more of written wellness journals, spoken wellness journals, present-time day logs, and time-to-time day logs; or a combination thereof. 5. A system according to claim 1, wherein the evaluation module performs the evaluation following one or more of each response to the qualitative inquiry, at an end of a section of the qualitative inquiry, and upon completion of the qualitative inquiry. 6. A... | Cardiac Pacemakers, Inc. | A61B005/00 G06F019/00 A61N001/00 A61B005/16 A61N001/372
US09259164 | B2 | 2016-02-16 | Controller for neuromuscular testing | 2006-08-21 | A peripheral nervous system (PNS) neuromuscular disorder testing system incorporates a testing apparatus that is affixed to a patients skin, positioning a stimulus transducer (e.g., electrical contact(s), movable sharp point, etc.) into a nerve pathway of interest with an electrode and reference electrodes appropriately positioned relative thereto to measure the electrical signal produced. A wireless connection to a control box relays the data associated with the time of the stimulus and the sensed response to a diagnostic system that analyzes the waveform per selectable testing protocols with user tailorable view capabilities and data dissemination communication paths. | 1. A kit operable to perform both evoked response potential (ERP) and peripheral nervous system (PNS) testing, the kit comprising: a PNS testing device positionable upon a subject's skin having a signal electrode, a reference electrode and a tactile stimulus producer;an ERP testing device configured to perform an ERP test;a control unit device operably configured to communicate a plurality of testing control administration parameters to the testing devices, wherein the control unit device is operable to perform both a peripheral nervous system nerve conduction test with the PNS testing device and an evoked response potential test with the ERP testing device, wherein the control unit device comprises a first port that is selectively connectable to a selected one of the PNS testing device or the ERP testing device;a remote storage device, wherein the remote storage device is configured to store classifiers associated with a plurality of neuromuscular disorders;a computer processor device in communication with the remote storage device, wherein the computer processor device is remote from the control unit device, wherein the computer processor device is configured to receive PNS test results and ERP test results from the control unit device, wherein the computer processor device is configured to compare the PNS test results or ERP test results from the control unit device against the classifiers stored on the remote storage device to correlate the PNS test results or ERP test results with one or more of the neuromuscular disorders, wherein the computer processor device is further configured to generate additional classifiers based at least in part on the PNS test results or ERP test results, the additional classifiers being associated with one or more neuromuscular disorders; anda second port in communication with the control unit device, wherein the a second port is configured to:(i) provide remote registration of the control unit device with the computer processor device via a communication network,(ii) provide communication of one or more PNS testing protocols or ERP testing protocols from the computer processor device to the control unit device via the communication network, and(iii) provide communication of the PNS test results or ERP test results from the control unit device to the computer processor device via the communication network,wherein the testing control administration parameters communicated from the control unit device to the testing devices are provided in accordance with the one or more PNS testing protocols or ERP testing protocols. 2. The kit of claim 1, wherein the tactile stimulus producer comprises an electrically actuated projecting member to jab the skin. 3. The kit of claim 1, wherein the tactile stimulus producer comprises a pair of selectively electrified contacts. 4. The kit of claim 1, wherein the plurality of testing control administration parameters are selected from a group consisting of a stimuli start time, a ... | Neuronetrix, Inc. | A61B005/05 A61B005/04 A61B005/00 A61B005/11
US20080287768 | A1 | 2008-11-20 | NECKLACE TYPE OF DETECTOR FOR ELECTROCARDIOGRAPHIC AND TEMPERATURE INFORMATION | 2007-08-01 | The present invention provides an non-adhering electrocardiographic and temperature signal detector comprising: (a) an electrocardiographic detecting working electrode, an electrocardiographic reference electrode and a temperature sensor connected together by a conduction chain; (b) a signal processing unit, for processing those signals obtained from said electrode or sensor, (c) a wireless transceiver for sending and receiving processed signal of electrocardiographic to the receiver at a far end unit or receiving data from the tar end unit; and (d) a power supply supplying power to the operation of the detector. The detector improves the uncomfortable of conventional adhering detector and increase the willing of a use due to the aesthetic, outlook of the detector and convenient to detect. Accordingly, the electrocardiographic and body temperature gas and opportunity to apply home care device. | 1. A non-stuck detector for electrocardiographic and temperature information measurement comprising: (a) a working electrode, a reference electrode and a temperature sensor connected by a chain; (b) a signal processing unit for processing those signals got from said electrode/sensor, (c) a wireless transceiver for sending/receiving messages to/from a far end; and (d) a power supply. 2. The detector according to claim 1, wherein the working electrode and the reference electrode are constructed into a two-electrode system, the former detects a surface voltage of a use's skin and the latter provides a reference voltage, thus an electrocardiographic signal is got by comparing the surface voltage and the reference voltage. 3. The detector according to claim 2, wherein the working electrode comprises an induction coil to detect the surface voltage of the user's skin. 4. The detector according to claim 1, wherein the reference electrode is applied on detection and amplification of electrocardiographic signals. 5. The detector according to claim 1, wherein the working electrode and the reference electrode are placed separately on the chain part and a pendant part linked to said chain. 6. The detector according to claim 5, wherein the electrode placed on the chain part is the chain itself or is a conductive plate on the chain. 7. The detector according to claim 5, wherein the electrode placed on the pendant is a conductive outer shell of said pendant or is an induction coil underneath the shell. 8. The detector according to claim 7, wherein the conductive outer shell is made by metal or plastic having electrical conductivity. 9. The detector according to claim 8, wherein the plastic is conductive polymer, or is polymer processed by electroplating or coating with conductive material. 10. The detector according to claim 1, wherein the temperature sensor detects body temperature through resistance thermometer, thermistor, thermocouple, or any non-invasive electric-thermometer. 11. The detector according to claim 1, wherein the signal processing unit comprising: (a) a signal amplifier; (b) an electronic filter; and (c) an analog-to-digital converter (ADC). 12. The detector according to claim 11, wherein the electronic filter is selected from the group consisting of low-pass filter, high-pass filter, band-pass filter, band-stop filter and all-pass filter. 13. The detector according to claim 1, wherein the wireless transceiver transports signals through radio wave interface or through infrared interface. 14. The detector according to claim 13, wherein the radio wave interface is IEEE 802 standard, RFID, GSM, PHS or CDMA. 15. The detector according to claim 1, wherein the power supply is battery. 16. The detector according to claim 1, wherein the signal processing unit, the wireless transceiver and the power supply are placed in the pendant. 17. The detector according to claim 1, which further comprises a data storage unit. 18. The detector acco... | NATIONAL YANG-MING UNIVERSITY | A61B005/0408 A61B005/01
US20080243017 | A1 | 2008-10-02 | BREATHING SOUND ANALYSIS FOR ESTIMATION OF AIRLOW RATE | 2007-03-28 | Apparatus for use detection of apnea includes a microphone mounted in the ear of the patient for detecting breathing sounds and a second external microphone together with an oximetric sensor. A transmitter at the patient compresses and transmits the signals to a remote location where there is provided a detector module for receiving and analyzing the signals to extract data relating to the breathing. The detector uses the entropy or range of the signal to generate an estimate of air flow while extracting extraneous snoring and heart sounds and to analyze the estimate of air flow using Otsu's threshold to detect periods of apnea and/or hypopnea. A display provides data of the detected apnea/hypopnea episodes and related information for a clinician. | 1. Apparatus for use in analysis of breathing of a patient comprising: a microphone arranged to be located on the patient for detecting breathing sounds;a detector module for receiving and analyzing the signals to extract data relating to the breathing;the detector module being arranged to analyze the signals to generate an estimate of air flow;and a display for displaying for a clinician the estimated air flow rate relative time;wherein the detector module is arranged to calculate a function representing the range of the signal or the entropy of the signal providing an estimate of air flow during breathing. 2. The apparatus according to claim 1 wherein the detector module is arranged to cancel heart sounds from the function. 3. The apparatus according to claim 1 wherein the function using entropy is defined by the following formula.H(p)=∑i=1Nptlogpi,where p, is the probability distribution function of the ith event. 4. The apparatus according to claim 1 wherein the function is the range of signal which is defined as the log of the difference between minimum and maximum of the signal within each short window (i.e. 100 ms) of data. 5. The apparatus according to claim 1 wherein the extraneous sounds are removed by the detector module prior to flow estimation. 6. The apparatus according to claim 1 wherein the display is arranged to display the estimated air flow versus time in any desired time length being chosen by the user. 7. The apparatus according to claim 1 wherein the display is capable of zoom-in and zoom-out functions in the same window. 8. The apparatus according to claim 16 wherein the display is capable of playing the breathing sounds in any data window. 9. The apparatus according to claim 1 wherein the microphone is arranged to be located in the ear of the patient. 10. The apparatus according to claim 9 wherein the microphone in the ear includes a transmitter arranged for wireless transmission to a receiver. 11. The apparatus according to claim 1 wherein the apparatus is arranged such that inspiration and expiration are monitored by an initial calibration wherein the patient is instructed to initialize the system by taking a deep breath, hold it, start up of the monitoring system, then exhale and continue breathing. 12. The apparatus according to claim 1 wherein an estimate of flow rate is calibrated using a look-up table of previously measured flow-sound relationship data that is sorted based on characteristics of the subjects. 13. The apparatus according to claim 1 wherein the characteristics in the look-up table include BMI, gender, height, neck circumference, and smoking history of the subject. 14. The apparatus according to claim 1 wherein the detector module is arranged to cancel heart sounds. 15. Apparatus for use in use in analysis of breathing of a patient during sleep comprising: a microphone arranged to be located on the patient for detecting breathing sounds;a detector module for receiving and analyzing th... | - | A61B005/08
US20080275328 | A1 | 2008-11-06 | SENSING BODY FUNCTIONS | 2007-12-27 | A body function sensor (1, 2, 3) for detecting breathing or heart beat or the like has an RF source (5) driving a transmit antenna (20) to create a low frequency (LF) electromagnetic detection field at a location to be occupied by an animal body. A receive antenna (21) is arranged to drive an analyser means (23, 13, 15,) to produce a signal representing the difference between the transmitted and detected signals detect a change in the field caused by the breathing or heart beat of an animal body in the location. A control means (16) controls the frequency of the RF source (5) to ensure resonance, and also controls the analyser means (23, 13), a display monitor (17) and an alarm. A memory (18) records movements of the subject, breathing patterns, pulse rate and posture of the body. | 1. A body function sensor (1, 2, 3) arranged to detect a body function of an animal body by detecting any phase difference between electromagnetic energy to which the animal body is exposed and electromagnetic energy received by an analyser means, characterised in that the sensor is arranged to create a low frequency (LF) electromagnetic field at a location where the animal body is to be positioned and to detect any phase changes in this electromagnetic field, a control means to adjust the frequency of the electromagnetic field until it is resonant, and the analyser means is arranged to assess minor phase changes in the electromagnetic field as an indication of the body function. 2. A body function sensor as claimed in claim 1, characterised in that the analyser means is arranged to determine the presence or movement of an animal body at the location by detecting a change in resonant frequency caused by the presence or movement of the animal body. 3. A body function sensor as claimed in claim 1, characterised in that the analyser means is arranged to determine at least one body function of the animal body by detecting changes in the phase of the electromagnetic field caused by the movement of the animal body at the location. 4. A body function sensor as claimed in claim 1, characterised in that a variable frequency waveform generator has its frequency controlled by the control means and is arranged to transmit its output through a transmit means to an antenna for transmitting the low frequency (LF) electromagnetic field at the location. 5. A body function sensor as claimed in claim 4, characterised in that the antenna includes a single conductor positioned at the location and also serving to transmit a received signal to an analyser means. 6. A body function sensor as claimed in claim 4, characterised in that a receive antenna is provided to receive the electromagnetic field and to convey it to the analysis means. 7. A body function sensor as claimed in claim 4, characterised in that the antenna (4) and the receive antenna include separate conductors positioned at the location. 8. A body function sensor as claimed in claim 7, characterised in that the separate conductors are intercalated. 9. A body function sensor as claimed in claim 4, characterised in that the variable frequency waveform generator includes a source of LF radio-frequency radiation. 10. A body function sensor as claimed in claim 4, characterised in that the variable frequency waveform generator is connected to the transmit antenna. 11. A body function sensor as claimed in claim 1, characterised in that the analyser means includes an Analogue/Digital converter to receive a signal representative of the electromagnetic field, and a Fast Fourier Transformer, or its electronic equivalent, to determine the difference in phase between the transmitted LF signal and a received LF signal changed by said body function. 12. A body function sensor as claimed in claim 11, characteri... | - | A61B005/05
US20080275359 | A1 | 2008-11-06 | Mobile in vivo brain scan and analysis system | 2008-07-03 | Described is a mobile in vivo brain scan and analysis system. The system includes a data collection subsystem and a data analysis subsystem. The data collection subsystem is formed to collect brain data that is reflective of firing neurons in a mobile subject and transmit the brain data to the data analysis subsystem. The data analysis subsystem is configured to generate and display a three-dimensional image that depicts a location the firing neurons. The data analysis subsystem is also configured to compare the brain data against a library of brain data to detect an anomaly in the brain data, and notify a user of any detected anomaly in the brain data. | 1. A mobile brain scan and analysis system, comprising: a data analysis subsystem, the data analysis subsystem including one or more processors that are configured to receive brain data from a remote data collection subsystem, the brain data being reflective of firing neurons in a mobile subject, and wherein the data analysis subsystem is further configured to generate and display a three-dimensional image that depicts a location of the firing neurons. 2. A mobile brain scan and analysis system as set forth in claim 1, wherein the data analysis subsystem is further configured to: compare the brain data against a library of brain data to detect an anomaly in the brain data, the anomaly being indicative of an abnormal brain function; andnotify a user of any detected anomaly in the brain data. 3. A mobile brain scan and analysis system as set forth in claim 2, further comprising a data collection subsystem, the data collection subsystem being formed to collect brain data and transmit the brain data to the data analysis subsystem. 4. A mobile brain scan and analysis system as set forth in claim 3, wherein the data collection subsystem is further formed to transmit the brain data wirelessly to the data analysis subsystem. 5. A mobile brain scan and analysis system as set forth in claim 4, wherein data collection subsystem further comprises a helmet with an array of electrodes, with each of the electrodes being formed to detect electroencephalograph (EEG) data from a user. 6. A mobile brain scan and analysis system as set forth in claim 5, wherein the data analysis subsystem further comprises: a receiver system, the receiver system being configured to receive the transmitted brain data from the data collection subsystem;a data processing system, the data processing system having a relational database management system (RDBMS) controller for connecting with and operating an RDBMS having a library of brain data, and further being configured to receive the brain data from the receiver system and compare the brain data to the RDBMS to detect an anomaly in the brain data. 7. A mobile brain scan and analysis system as set forth in claim 6, wherein the data analysis subsystem is further configured to compare a detected anomaly in the brain data with an RDBMS to generate a diagnosis of the detected anomaly. 8. A mobile brain scan and analysis system as set forth in claim 7, wherein the data analysis subsystem is further configured to compare the three-dimensional image with a RDBMS having a library of three-dimensional images to detect an anomaly in the brain data. 9. A mobile brain scan and analysis system as set forth in claim 1, further comprising a data collection subsystem, the data collection subsystem being formed to collect brain data and transmit the brain data to the data analysis subsystem. 10. A mobile brain scan and analysis system as set forth in claim 9, wherein the data collection subsystem is further formed to transmit the brain data ... | - | A61B005/0476
US20080242950 | A1 | 2008-10-02 | Computational user-health testing | 2007-05-24 | Methods, apparatuses, computer program products, devices and systems are described that carry out detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection; mapping the user data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection to at least one user-health test function set; and selecting at least one user-health test function in response to the at least one user-health test function set. | 1. A method comprising: detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection;mapping the user data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection to at least one user-health test function set; andselecting at least one user-health test function in response to the at least one user-health test function set. 2. The method of claim 1 wherein the detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: detecting user input data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection. 3. The method of claim 1 wherein the detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: detecting passive user data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection. 4. The method of claim 1 wherein the detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: detecting user reaction time data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection. 5. The method of claim 1 wherein the detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: detecting user speech or voice data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection. 6. The method of claim 1 wherein the detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: detecting user hearing data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection. 7. The method of claim 1 wherein the detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: detecting user body movement, pupil movement, or eye movement data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection. 8. (canceled) 9. The method of claim 1 wherein the detecting user data from an interaction bet... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00
US07470233 | B2 | 2008-12-30 | Method and apparatus for muscle function measurement | 2005-01-26 | The present invention provides an apparatus and method for monitoring muscle function based on an index derived from a pressure or force signal. The muscle function index is derived from an instantaneous muscle stiffness ratio computed as the ratio of the first time derivative of the pressure or force waveform to the corresponding instantaneous pressure or force. The instantaneous stiffness ratio, Ė/E(t), is in units of 1/sec and relates to the rate of strong bond formation and will be influenced by calcium handling properties of the muscle fibers and the intracellular calcium concentration. As such, an index derived from Ė/E(t) provides a measure of the inotropic status of the muscle. | 1. A method, comprising: sensing a pressure signal associated with a muscular contraction event via a chronically-implantable pressure sensor coupled to a chronically implanted medical device;computing an instantaneous stiffness ratio waveform as a ratio of a first time derivative of the pressure signal to an instantaneous pressure metric within said IMD; andderiving an index of muscle function from the instantaneous stiffness ratio waveform within said IMD,wherein the computing and deriving steps of said method are carried out within via operative circuitry wholly disposed within said IMD; andwherein the index of muscle function comprises a peak of the stiffness ratio waveform. 2. A method according to claim 1, wherein the pressure signal comprises a cardiac chamber pressure signal. 3. A method according to claim 1, wherein the muscular contraction event comprises at least part of one of a myocardial depolarization sequence and a myocardial repolarization event. 4. A method according to claim 1, wherein the pressure signal is obtained using an intracardiac pressure transducer. 5. A method according to claim 4, wherein the intracardiac pressure transducer is adapted to be disposed within a portion of a right ventricular chamber. 6. A method, comprising: sensing a force signal associated with a muscular contraction event, wherein the force signal is provided by a chronically-implantable sensor which is coupled to a chronically implanted medical device (IMD);computing an instantaneous stiffness ratio waveform as a ratio of a first time derivative of the force signal to an instantaneous force metric within the IMD; andderiving an index of muscle function from the stiffness ratio waveform within the IMDwherein the computing and deriving steps are performed via circuitry disposed within the. IMD; andwherein the index of muscle function comprises a peak of the stiffness ratio waveform. 7. A method according to claim 6, wherein the force signal comprises a cardiac chamber force signal. 8. A method according to claim 6, wherein the muscular contraction event comprises at least part of one of a myocardial depolarization sequence and a myocardial repolarization event. 9. A method according to claim 6, wherein the force signal is obtained using an intracardiac force transducer. 10. A method according to claim 9, wherein the intracardiac force transducer is adapted to be disposed within a portion of a right ventricular chamber. 11. A chronically implantable medical device (IMD), comprising: means for sensing a pressure signal associated with a muscular contraction event wherein the pressure signal is provided by a chronically-implantable sensor and said sensor couples to the (IMD);means within the IMD for computing an instantaneous stiffness ratio waveform as the ratio of the first time derivative of the pressure signal to the instantaneous pressure; andmeans within the IMD for deriving an index of muscle function as a feature of the instantaneous ... | Medtronic, Inc. | A61B005/02
US07458167 | B2 | 2008-12-02 | Arm angle indicator | 2004-12-30 | An angle measurement device is provided for determining the version angle for the prosthetic humeral head during shoulder replacement surgery. The device has a first collar for contacting the forearm and has a first coupling element in the form of a rod aligned with the ulna. The device has a second collar for contacting the upper arm when the elbow is flexed to 90°. The second collar has a second coupling element coupled thereto extending in a direction generally perpendicular to the humerus. The second coupling element is pivotally coupled to the first coupling element for rotation about a pivot point axis. An angle indicator having angle indicia is located at said pivot point and fixed to one of said first or second arm contacting collars for indicating the angle between the ulna and the humerus during external rotation of the forearm. | 1. An arm angle measurement device for determining the version angle of the humeral head during shoulder surgery when the forearm is externally rotated with the elbow flexed to 90° comprising: a first element for contacting the forearm having a first coupling element aligned with the ulna;a second element for contacting the upper arm, said second element having a second coupling element aligned with to the humerus, said second coupling element pivotally coupled to said first coupling element for rotation about a pivot point; andan angle indicator having angle indicia located at said pivot point and fixed to one of said first or second elements for indicating the external rotation angle between said first and second coupling elements, wherein the external rotation angle is indicative of the version angle of the humeral head. 2. The angle measurement device as set forth in claim 1 wherein said angle indicator has a face with angle indicia from ± 40° centered about a position wherein the axis of the ulna is in a plane parallel to the sagittal plane. 3. The arm angle measurement device as set forth in claim 2 further comprising means for locking the first and second coupling element at a selected angle of said angle indicator. 4. The arm angle measurement device as set forth in claim 1 further comprising means for locking the first and second coupling element at a selected angle of said angle indicator. 5. The arm angle measurement device as set forth in claim 1 wherein the first and second elements are connected to the forearm and upper arm by collars which have open portions which are open to the posterior. 6. The arm angle measurement device as set forth in claim 5 wherein said open portions are oriented in generally perpendicular directions. 7. The arm angle measurement device as set forth in claim 5 wherein the collars are part circular with sectors open to the posterior for receiving the forearm and upper arm. 8. The arm angle measurement device as set forth in claim 7 wherein the collars are adapted to receive straps for respectively encompassing posterior portions of the forearm and upper arm. 9. The arm angle measurement device as set forth in claim 1 wherein the angle indicator is fixed to the second element. 10. The arm angle measurement device as set forth in claim 1 wherein the first and second elements have arm receiving portions oriented in generally perpendicular directions. 11. The arm angle indicator as set forth in claim 1 wherein said first and second coupling elements are rods extending from said first and second elements. 12. An angle indicator for displaying the external rotation angle between a 90° flexed forearm and upper arm with respect to the sagittal plane comprising: a first collar shaped to receive the forearm;a second collar shaped to receive the upper arm pivotally coupled to said first collar at a pivot point;a dial adjacent said pivot point having angle indicia thereon for displaying the angle between sai... | Howmedica Osteonics Corp. | A61B005/103
US07438694 | B2 | 2008-10-21 | Lancing device | 2006-03-07 | A lancing device is used with a lancet for lancing body tissue to result in a wound for bleeding. An improved lancing device has a priming knob having a slanted priming notch rotatably connected to the lancing device about an axis of rotation generally parallel to a central axis of the lancing device and a strike path of a lancet carrier. Rotation, translation, or both rotation and translation of the priming knob allows a user to prime the lancing device. | 1. A lancing device, for use with a lancet for lancing body tissue to result in a wound for bleeding, the lancing device comprising: (1) a generally annular housing having a central axis and an internal channel extending along the central_axis terminating at a priming end of the housing and at an opposed lancing end of the housing;(2) a priming knob rotatably and translationally connected about the housing to permit rotation about an axis of rotation generally parallel to the central axis of the housing and translation in a direction parallel to the axis of rotation, wherein the priming knob comprises a slanted priming notch for priming interaction with a priming protrusion of a lancet carrier,(3) a lancet carrier translationally mounted within the internal channel of the housing for carrying a lancet along a strike path that is generally parallel with the central axis, the strike path of the lancet carrier starting from a primed position toward the priming end of the housing wherein the lancet is shielded by the housing and ending at a lanced position wherein a tissue penetrating portion of the lancet extends outwardly from the lancing end of the housing, the lancet carrier comprising a priming protrusion for priming interaction with the slanted priming notch of the priming knob; and(4) a lancet carrier driver disposed within the internal channel of the housing, the lancet carrier driver being operatively engaged with the lancet carrier when driving the lancet carrier along the strike path from the primed position to the lanced position; wherein when the slanted priming notch is in priming interaction with the priming protrusion of the lancet carrier: (i) rotation of the priming knob is by itself sufficient to prime the lancet driver and places the lancet carrier into a primed position; (ii) movement of the priming knob along an axis generally parallel to the central axis of the housing is by itself sufficient to prime the lancet driver and places the lancet carrier into a primed position and (iii) a combination of both rotation of the priming knob and movement of the priming knob along an axis generally parallel to the central axis of the housing, prunes the lancet driver and places the lancet carrier into the primed position. 2. The device of claim 1, wherein the priming knob is rotatably connected to the housing at the priming end of the housing. 3. The device of claim 1, wherein the slanted priming notch is symmetrical such that when the priming notch is in priming interaction with the priming protrusion of the lancet carrier, clockwise rotation, counterclockwise rotation, or both clockwise and counterclockwise rotation of the priming knob primes the lancet driver and places the lancet carrier into the primed position. 4. The device of claim 1, wherein the slanted priming notch comprises a series of priming teeth and the priming protrusion comprises a priming tooth, and wherein when the slanted priming notch is in priming interaction with... | Agamatrix, Inc. | A61B005/00 A61B017/32
USRE040587 | E1 | 2008-11-25 | Antenna for magnetic resonance imaging and method of use | 2000-08-10 | A medical appliance for use in magnetic resonance imaging procedures. A guidewire for vascular procedures is formed by a coaxial cable acting as antenna in a magnetic resonance imaging system. | 1. A medical appliance comprising an elongated signal-receiving antenna for detecting and providing magnetic resonance response signals, the antennas adapted to be inserted into the body during magnetic resonance imaging procedures and for providing the response signals used for calculating a position of the medical appliance in the body, wherein the antenna comprises an open wire length including first and second conductor means having proximal ends adapted and arranged for interconnection to a receiver to couple the detected resonance response signals to the receiver, spaced-part distal ends, and at least a first insulator means for physically separating and electrically insulating adjacent portions of the first and second conductor means, the distal ends of the first and second conductor means and the at least first insulator means adapted and arranged for exposure to a field of electromagnetic energy during a magnetic resonance procedure to couple electromagnetic energy from the field into the antenna and detect and provide the magnetic resonance response signals to the proximal ends of the conductor means. 2. A medical appliance according to claim 1, wherein the open wire length antenna is formed of a coaxial cable including the first and second conductors in a coaxial arrangement. 3. A medical appliance according to claim 1, wherein the open wire length antenna is formed of a cable having the first conductor enclosed in the first insulator, the first insulator surrounded by the second conductor and the second conductor encased in a second insulator, and wherein said first conductor and second conductor have the same length. 4. A medical appliance according to claim 1, wherein the open wire length antenna is formed of a cable having the first conductor enclosed in the first insulator, the first insulator surrounded by the second conductor, and the second conductor encased in a second insulator, and wherein said first conductor and second conductor have unlike lengths. 5. A medical appliance according to claim 1, wherein the open wire length antenna is made of the first conductor, the first insulator includes a first insulating coating applied on said first conductor, the second conductor includes a conducting coating surrounding said first insulating coating, and the antenna further includes a second insulating coating applied on said conducting coating, and wherein said first conductor and conducting coating have the same length. 6. A medical appliance according to claim 1, wherein the open wire length antenna is made of the first conductor, the first insulator includes a first insulating coating applied on said first conductor, the second conductor includes a conducting coating surrounding said first insulating coating, and the antenna further includes a second insulating coating applied on said conducting coating, and wherein said first conductor and conducting coating have unlike lengths. 7. A medical appliance according to claim 1... | Schneider (Europe) A.G. | A61B005/055
US20080272918 | A1 | 2008-11-06 | MEDICAL DEVICE AND METHOD FOR PREVENTING FALLS | 2007-05-04 | A medical device is beneficial to protect a patient in a care facility from a fall by remotely alerting a caregiver of patient's unsafe mobility before the fall using a motion sensor having an alarm. The medical device uses an operating module with a wireless transmitter integrated into the motion sensor for broadcasting the alarm as an alert signal to notify the caregiver of patient's unsafe motion. A wireless receiver with the caregiver receives the alert signal from the wireless transmitter. The alert signal identifies the motion sensor that was activated so the caregiver can respond to prevent the patient from falling. | 1. A medical device to protect a patient in a care facility from a fall by remotely alerting a caregiver of patient's motion before the fall comprising: a motion sensor;an operating module and a wireless transmitter in communication with the motion sensor for broadcasting an alert signal to notify the caregiver of patient movement by a particular patient;a wireless receiver with the caregiver for receiving the alert signal from the wireless transmitter; andthe alert signal for identifying the motion sensor that was activated so the caregiver can respond to prevent the patient from falling. 2. The medical device of claim 1 wherein the wireless receiver has a reset button so the caregiver who responds first can send a reset signal from the wireless receiver to the motion sensor. 3. The medical device of claim 1 wherein the motion sensor is a pressure sensor positioned beneath where the patient lies or sits. 4. The medical device of claim 1 wherein the motion sensor is positioned to detect patient movement before the fall. 5. The medical device of claim 1 wherein the alert signal from the wireless transmitter displays on the wireless receiver the motion sensor activated, patient name, patient number, room number, and/or patient location. 6. The medical device of claim 1 wherein the wireless receiver sends a programming signal to the motion sensor to thereby program the operating module. 7. The medical device of claim 1 wherein the wireless receiver is a pager, a handheld or a cellular device. 8. A medical device to protect a patient in a care facility from a fall by remotely alerting a caregiver of patient's motion before the fall comprising: a motion sensor having a silent alarm;a wireless transmitter powered by and integrated into the motion sensor;an operating module adapted to convert the alarm to an alert signal broadcast from the wireless transmitter within the care facility;a wireless receiver with the caregiver for receiving and displaying the alert signal from the wireless transmitter;a reset button on the wireless receiver so the caregiver can send a reset signal from the wireless receiver to the operating module to reset the motion sensor; andthe alert signal for identifying the motion sensor that was activated so the caregiver can respond to prevent the patient from falling. 9. The medical device of claim 8 wherein the wireless receiver signals the caregiver with audible tones, visual cues, and/or vibration pulses when the motion sensor is activated. 10. The medical device of claim 8 wherein the wireless receiver is programmed to identify alert signals from multiple motion sensors. 11. The medical device of claim 8 wherein the motion sensor is a pressure sensor positioned to monitor unsafe motion of the patient. 12. The medical device of claim 8 wherein the wireless receiver has a visual display to interface with and program the operating module. 13. The medical device of claim 8 wherein the wireless receiver sends a programm... | - | G08B023/00 A61B005/12
US20080294029 | A1 | 2008-11-27 | System to Measure Blood Coagulation Related Parameters | 2006-11-24 | System to measure blood coagulation related parameters comprising one first channel (6) adapted to contain a blood sample; said system containing, at least partially, an expandable material which is able to increase its volume when activated by an exciting source; said system furthermore comprising several excitable regions distributed close to said first channel (6), in such a way that, when one of said excitable region is activated, said expandable material increases to such an extend that the channel cross section is reduced to a location situated at or near to said excitable region. | 1. System to measure blood coagulation related parameters comprising one first channel adapted to contain a blood sample; said system containing, at least partially, an expandable material which is able to increase its volume when activated by an exciting source; said system furthermore comprising several excitable regions distributed close to said first channel, in such a way that, when one of said excitable region is activated, said expandable material increases to such an extend that the channel cross section is reduced at a location situated at or near to said excitable region. 2. System according to claim 1 furthermore comprising at least a second channel, said second channel being parallel to and communication at one of its end with said first channel. 3. System according to claim 1 wherein said material is adapted to be activated by heat, light, pH variation or by a chemical, electromagnetic or radioactive element as exciting factor. 4. System according to claim 1 wherein said material is a mixture comprising Expancel®. 5. System according to claim 1 in combination with an exciting source, such as a laser, which is distinct from said system. 6. System according to claim 1 furthermore comprising an exciting source such as an electrical resistance. 7. System according to claim 1 wherein said channel is defined by two horizontal and two lateral walls. 8. System according to claim 7 wherein at least one of said horizontal walls is made of said expandable material. 9. System according to claim 1 wherein said channel cross section is adapted to be reduced along a vertical direction. 10. System according to claim 7 wherein both lateral walls are made of said expandable material. 11. System according to claim 10 wherein said channel cross section is adapted to be reduced along a horizontal direction. 12. System according to claim 1 wherein the said material is expandable and is a mixture comprising PDMS or a biocompatible material. 13. System according to claim 1 wherein said channel(s) comprise(s) at least a region of narrowing. 14. System according to claim 1, said system having a multilayered structure comprising successively a substrate layer, an expandable layer, a rigid layer and a cover layer, said substrate layer incorporating said exciting source and said channel(s) being defined in said rigid layer. 15. System according to claim 14 wherein said substrate layer is made of PCB, said expandable layer is made of PDMSXB et said rigid layer is made of PDMS. 16. System according to claim 1 furthermore comprising sensing means. 17. System according to claim 16 wherein said sensing means are diodes, optical density or impedance measuring means. 18. System according to claim 16 wherein said sensing means are adapted to locate where the blood stops. 19. System according to claim 16 wherein said sensing means are adapted to detect blood displacement. 20. Assembly comprising a system as defined in claim 1 in combination with at lea... | - | A61B005/00
US20080262372 | A1 | 2008-10-23 | Medical Device Adapted to the Monitoring of Muscle Behaviour in Patients | 2005-06-03 | The present invention is related to a medical device for monitoring limb muscle behaviour in a patient, said limb corresponding to a leg, an arm, or the neck, said medical device comprising: a moving unit (1) adapted to receive a first part of a limb, said moving unit (1) presenting at least one degree of freedom (11) according to which the moving unit (1) may move;a haptic device (6) adapted to receive a second part of the limb and to measure in the form of an analogical signal the mechanical response of said second part to a movement of the moving unit (1);a set of electromyographic electrodes (4) able to measure in the form of an analogical signal the electrical activities in the muscle(s) of the limb;an analogical signal acquisition unit (9) able to acquire the analogical signals from the electromyographic electrodes and the haptic device (6) and converting said analogical signals into digital signals;a controller (8) able to process said digital signals and to control the moving unit (1). | 1. A medical device for monitoring limb muscle behaviour in a patient, said limb corresponding to a leg, an arm or the neck, said medical device comprising: a moving unit adapted to receive a first part of a limb corresponding to a second part of a limb, said moving unit presenting at least one degree of freedom according to which the moving unit may move;a haptic device adapted to receive a second part of the limb corresponding to the second part of the limb respectively, and to measure in the form of an analogical signal the mechanical response of said second part to a movement of the moving unit;a set of electromyographic electrodes able to measure in the form of an analogical signal the electrical activities in the muscle(s) of the limb;an analogical signal acquisition unit able to acquire the analogical signals from the electromyographic electrodes and the haptic device and converting said analogical signals into digital signals; anda controller able to process said digital signals and to control the moving unit. 2. The medical device according to claim 1, wherein the moving device is driven by a motor under the control of the controller. 3. The medical device according to claim 1, wherein the electromyographic electrodes are needle electrodes and/or surface electrodes. 4. The medical device according to claim 1, which comprises four electromyographic needle electrodes and two electromyographic surface active electrodes. 5. The medical device according to claim 1, wherein the electromyographic electrodes are microelectrodes. 6. The medical device according to claim 1, wherein the moving unit presents at least one rotation axis around which the moving unit is able to rotate relatively to the haptic device with a rotation angle (α) comprised between about +45° and about −45° and a defined maximal rotation velocity. 7. The medical device according to claim 6, wherein the maximum rotation velocity of the moving unit is about 300°/sec. 8. The medical device according to claim 6, wherein the moving unit (1) has a position accuracy of about 0.1 mm. 9. The medical device according to claim 1, which further comprises an amplification system for amplifying the analogical signals measured by the electrodes. 10. The medical device according to claim 1, wherein the controller maintains in direction and in intensity the driving force of the motor and thereby in direction and in intensity the motion of the moving unit. | - | A61B005/0488
US20080294015 | A1 | 2008-11-27 | Administration system | 2008-05-22 | An administration system proposes necessary exercise content in accordance with dining selections and a physical condition of a user. A portable terminal obtains and stores physical condition information and information on calories ingested upon dining. The portable terminal selects exercise content candidates based on the calorie ingestion information and the physical condition information and displays the selected exercise content candidates. A user can be informed before the start of exercise of necessary exercise contents corresponding to selected dining in consideration of physical condition. | 1. An administration system comprising: first acquisition means for acquiring calorie information from a food provision facility terminal in a dining facility;second acquisition means for acquiring physical condition information associated with a physical condition of a user;necessary exercise content selection means for selecting a necessary exercise content based on the calorie information acquired by the first acquisition means while considering the physical condition information acquired by the second acquisition means; andoutput means for outputting the exercise content selected by the necessary exercise content selection means. 2. The administration system according to claim 1, wherein the first acquisition means acquires a menu and the calorie information for each food of the menu from the food provision facility terminal, and wherein the necessary exercise content selection means selects the exercise content based on the calorie information corresponding to a food selected by selection means from the menu. 3. The administration system according to claim 1, wherein the physical condition information acquired by the second acquisition means includes one of a body temperature, a heart rate, a respiration rate, blood pressure, a perspiration amount, a body weight, a basal metabolic rate, a muscle volume and a body fat percentage. 4. The administration system according to claim 1, wherein the output means includes a display means having an image display function. 5. The administration system according to claim 1, wherein the output means outputs reservation information based on the exercise content to an exercise facility reservation management device. 6. The administration system according to claim 1, wherein the first acquisition means is incorporated in a portable terminal, and wherein the second acquisition means, the necessary exercise content selection means and the output means are mounted on a vehicle. 7. The administration system according to claim 1, wherein the first acquisition means, the necessary exercise content selection means and the output means are incorporated in a portable terminal, and wherein the second acquisition means is mounted on a vehicle. 8. The administration system according to claim 6, wherein the second acquisition means acquires the physical condition information from a driver sensor provided in the vehicle. 9. The administration system according to claim 7, wherein the second acquisition means acquires the physical condition information from a driver sensor provided in the vehicle. 10. A vehicle mounted terminal for coupling to a system for administering health related information and activities of a user of a vehicle, the vehicle mounted terminal comprising: a first interface capable of a near field communication for connecting, when within a predetermined near field range, with a portable terminal;a second interface capable of a vehicle network communication for connecting with an electronic contr... | DENSO CORPORATION | A61B005/00
US07431696 | B1 | 2008-10-07 | Monitoring device, method and system | 2007-12-10 | A monitoring device (20) for monitoring the vital signs of a user is disclosed herein. The monitoring device (20) is preferably an article (25) having an optical sensor (30) and a circuitry assembly (35). The optical sensor (30) preferably comprises a photodetector (130) and a plurality of light emitting diodes (135). The monitoring device (20) preferably provides for the display of the following information about the user: pulse rate; blood oxygenation levels; calories expended by the user of a pre-set time period; target zones of activity; time; or distance traveled. | 1. A method of monitoring a user's vital signs, the method comprising: deactivating a processor for a deactivation period of 128 to 640 microseconds;activating an optical sensor during the deactivation period to generate a signal corresponding to the flow of blood through at least one artery of the user, the optical sensor comprising a photodetector and a plurality of light-emitting diodes;reactivating the processor to process the signal to obtain an average pulse value of a first set of time periods;processing the signal to obtain an average pulse value of a second set of time periods, the second set of time periods less than the first set of time periods;subtracting the average pulse value of the second set of time periods from the average pulse value of the first set of time periods to generate a first filtered pulse data value;processing the first filtered pulse data value to obtain a first heart rate value;using the first heart rate value to filter the subsequent average pulse values generated from the digital signal to obtain a real-time heart rate value of the user wherein the real-time heart rate value is based on 60 to 180 sample readings and each of the sample readings is based on 4 to 15 sub-sample readings; andprocessing the real-time heart rate data of the user for analysis of real-time calories expended by the user and for real-time display of a plurality of the user's vital signs. | Impact Sports Technologies, Inc. | A61B005/00
US07454241 | B2 | 2008-11-18 | Method for cardiac magnetic resonance imaging | 2004-08-09 | The invention relates to a method for cardiac magnetic resonance imaging. The method comprises the steps of monitoring the ECG of a patient and detecting the occurrence of R-waves (R1, R2, R3) in the ECG, acquiring a plurality of phase encoded MR signals (a1-3, b1-3, c1-3, d1-3, e1-3) per cardiac cycle by subjecting the patient to a sequence of RF pulses and magnetic field gradient pulses, and reconstructing an MR image from at least a part of the MR signals (a1-3, b1-3, C1-3, d1-3, e1-3). For providing an improved cardiac MRI method which enables the acquisition and reconstruction of high quality MR images of the heart in cases in which the heart cycle of the examined patient changes over time or is irregular because of arrythmia, the invention proposes that the ECG of the patient is monitored continuously during acquisition of MR signals, wherein MR signals (d1 to c1, b2 to a2, c3 to b3) used for image reconstruction are retrospectively selected depending on the time interval between successive R-waves (R1, R2, R3). | 1. A method for cardiac magnetic resonance imaging, the method comprising the following steps: A) monitoring the ECG of a patient and detecting the occurrence of R-waves in the ECG;B) acquiring a plurality of phase encoded MR signals per cardiac cycle by subjecting the patient to a sequence of RF pulses and magnetic field gradient pulses;C) reconstructing an MR image from at least a part of the MR signals acquired in step B);wherein the ECG of the patient is monitored continuously during acquisition of MR signals in step B), wherein MR signals used for image reconstruction in step C) are retrospectively selected depending on the time interval between successive R-waves;wherein the MR signals are acquired over a plurality of cardiac cycles;wherein the acquisition of MR signals for each cardiac cycle begins after a delay period in an early diastolic phase of the cardiac cycle after detection of an R-wave and continues until the next R-wave is detected;wherein the delay period is adapted in accordance with the actually detected duration of the interval between successive R-waves. 2. The method for cardiac magnetic resonance imaging according to claim 1, wherein MR signals acquired during the diastolic phase are retrospectively selected from a series of MR signal sets acquired with different delay periods. 3. The method for cardiac magnetic resonance imaging according to claim 1, wherein a phase encoding scheme is applied in step B) such that the phase encoding values of the acquired MR signals are repeated cyclically. 4. The method for cardiac magnetic resonance imaging according to claim 1, wherein the sequence of RF pulses and magnetic field gradient pulses is selected such that a steady state acquisition scheme is applied. 5. The method for cardiac magnetic resonance imaging according to claim 4, wherein the sequence of RF pulses and magnetic field gradient pulses is a balanced turbo field echo sequence. 6. A device for magnetic resonance imaging of a body of a patient placed in a stationary and substantially homogeneous main magnetic field, the device comprising means for establishing the main magnetic field, means for generating magnetic field gradients superimposed upon the main magnetic field, means for radiating RF pulses towards the body, control means for controlling the generation of the magnetic field gradients and the RF pulses, means for receiving and sampling MR signals generated by sequences of RF pulses and magnetic field gradient pulses, reconstruction means for forming an image from the signal samples, and means for monitoring the ECG of the patient, wherein the control means is configured to execute an imaging procedure according to a method of claim 1. 7. A digital medium programmed with a program code, the program code being configured to control a computer processor to perform the imaging procedure according to a method of claim 1 to be carried out on a magnetic resonance device. 8. A cardiac MR imaging method which adj... | Koninklijke Philips Electronics N.V. | A61B005/05
US07738946 | B2 | 2010-06-15 | System, method and apparatus for the detection of patient-borne fluorescing materials | 2007-10-03 | System, method and apparatus wherein a probe employing non-imagining optics is utilized in conjunction with a fluorescing (e.g., nanocrystal) tracer at the body of a patient. Excitation components within the probe working end are utilized to excite the nanocrystals to fluoresce at wavelengths in the near infrared region, such fluorescent energy is homogenized by interacting with involved tissue to provide a uniform fluorescing intensity over the surface of a photo-detector. Initialization and background determination procedures are described along with a technique for determining statistically significant levels of fluorescing activity. | 1. A method for the surgical treatment of patients afflicted with neoplastic tissue, which comprises the steps: (a) administering to a patient an effective amount of a locator incorporating a fluorescing tracer and which specifically binds a marker produced by or associated with neoplastic tissue;(b) permitting time to elapse following step (a) for said locator to preferentially concentrate at any said marker and for unbound locator to be cleared so as to increase the ratio of fluorescing-based radiation from specifically bound locator to fluorescing-based radiation representing background in said patient;(c) after said time has elapsed in step (b) surgically accessing an operative field of the patient;(d) providing a hand manipulative probe having forwardly disposed excitation components energizable to cause any said fluorescing tracer to fluoresce at a detection wavelength or wavelengths and a forwardly disposed photo-detector configured for response to non-imaged said detection wavelengths so as to develop detection outputs corresponding with fluorescing radiation intensity;(e) using said probe, determining and storing the fluorescing radiation intensity at said background;(f) based upon the fluorescing radiation intensity at said background, determining a statistically significant fluorescing radiation intensity value whereat a perceptible cue is generated;(g) removing tumor burden from the patient using the probe where necessary to locate and differentiate neoplastic tissue; and(h) using the probe, determine and remove lymph tissue sites exhibiting detection outputs. 2. The method of claim 1 in which: step (a) administers said locator wherein a fluorescing nanocrystal is configured to fluoresce in the near infrared region of the spectrum. 3. A probe configured for detecting fluorescing tracers at the body, the tracers being responsive to wavelength defined excitation energy to fluoresce at one or more detection wavelengths including the near-infra-red region of the spectrum, comprising: a working portion having a transparent window with a forward transmission surface generally contactable with tissue of the body;a photo-detector disposed generally centrally and rearwardly of said window and configured for response to non-imaged fluorescing-based radiation;a generally cylindrical excitation mount within said working portion having a forward edge symmetrically disposed in spaced adjacency from said photo-detector;an array of excitation diodes supported at said mount forward edge in adjacency with said window, energizable under a modulation scheme to project photon excitation energy at an excitation wavelength or wavelengths through said window;a longpass filter located intermediate said photo-detector and said window configured to substantially block excitation and ambient illumination;an integrating pre-amplifier located rearwardly of said photo-detector and electrically coupled therewith to amplify a value associated with said non-imaged f... | Actis, Ltd. | A61B006/00 A61B008/00
US07435222 | B2 | 2008-10-14 | Tele-diagnostic device | 2006-09-25 | A system, probe, and method for collecting a plurality of diagnostic information and transmitting the diagnostic information to a remote location. The system comprises a member contoured to at least a portion of a person's hand, and an interface unit in electrical communication with the member. The interface unit is capable of transmitting information to a remote location. The member comprises at least eight sensors. | 1. A system for collecting diagnostic information and transmitting the diagnostic information to a remote location, the system comprising: a member comprising a palm portion, a wrist portion and a plurality of phalange portions comprising an index finger phalange portion and a middle finger phalange portion, with the index finger phalange portion being at least as long as the middle finger phalange portion of the member, the member comprising an EKG diagnostic device comprising at least eight EKG sensors located on the member on at least two of the palm portion, the wrist portion, and at least one of the phalange portions; andan interface unit in communication with the member, wherein the interface unit is capable of transmitting information to a remote location. 2. The system of claim 1 wherein the EKG sensors are located on the palm portion, the wrist portion, and at least one of the phalange portions of the member. 3. The system of claim 1 wherein the EKG diagnostic device comprises at least 10 sensors. 4. The system of claim 3 wherein the EKG diagnostic device comprises 11 sensors. 5. The system of claim 1 wherein the index finger phalange portion is longer than the middle finger phalange portion of the member. 6. The system of claim 1 wherein at least four of the EKG sensors are located on the index finger phalange portion of the member. 7. The system of claim 1 wherein at least five of the EKG sensors are located on the index finger phalange portion of the member. 8. The system of claim 1 wherein the plurality of phalange portions comprise a thumb portion, with at least one of the EKG sensors being located on the thumb portion of the member. 9. The system of claim 1 wherein the plurality of phalange portions comprise a pinky finger portion, with at least one of the EKG sensors being located on the pinky finger portion of the member. 10. The system of claim 1 wherein at least one of the EKG sensors is located on a palmer surface of the palm portion of the member. 11. The system of claim 1 wherein the member has a shape that corresponds to at least a substantial portion of a person's hand such that the member is capable of being worn on a person's hand. 12. The system of claim 11 wherein the member has a portion shaped to contour to a person's palm. 13. The system of claim 11 wherein the member has a portion shaped to contour to a person's finger. 14. The system of claim 13 wherein the member has a portion shaped to contour to a person's palm. 15. The system of claim 11 wherein the member comprises a plurality of diagnostic devices. 16. The system of claim 15 wherein the plurality of diagnostic devices includes the EKG diagnostic device, a blood pressure and pulse diagnostic device, and a temperature device. 17. The system of claim 16 wherein the plurality of diagnostic devices further includes a percent O2 diagnostic device. 18. The system of claim 17 wherein the plurality of diagnostic devices further includes an auscultati... | Ineedmd.com, Inc. | A61B005/02 A61B005/0205 A61B005/024 A61B005/0402 A61B005/01
US07463919 | B2 | 2008-12-09 | Cardiac diagnostics using wall motion and perfusion cardiac MRI imaging and systems for cardiac diagnostics | 2003-07-29 | Cardiac information of a patient is displayed by obtaining a plurality of MRI cine loops of the heart of the patient at a plurality of heart rates, the plurality of cine loops including both wall motion cine loops and at least one perfusion cine loops and simultaneously displaying both the wall motion cine loops and the at least one perfusion cine loop. | 1. A method of displaying cardiac information of a patient having a cardiac cycle, comprising: obtaining a plurality of MRI cine loops of the heart of the patient at a plurality of heart rates, the plurality of cine loops including cine loops including frames of wall motion images at different heart beat rates and at least one cine loop including frames of perfusion images of at least one cardiac location;synchronizing images in the wall motion cine loops so that heart motion at the different heart beat rates correspond to the same portion of the cardiac cycle and adjusting the at least one perfusion cine loop to have the same number of frames as the wall motion cine loops; andsimultaneously displaying both wall motion cine loops and the at least one perfusion cine loop. 2. The method of claim 1, further comprising adjusting a number of frames in ones of the displayed wall motion cine loops and/or the perfusion cine loop so that the displayed wall motion cine loops and the perfusion cine loop have the same number of frames, wherein each displayed corresponding image of wall motion in the cine loops is displayed at approximately the same percent of time within a heart beat. 3. The method of claim 1, wherein obtaining a plurality of MRI cine loops comprises acquiring a plurality of MRI cine loops while a stress test is administered to the patient. 4. The method of claim 1, further comprising evaluating the displayed MRI cine loops to determine a presence or absence of coronary artery disease based on the displayed cine loops. 5. A method of displaying cardiac information of a patient, comprising: obtaining a plurality of MRI cine loops of the heart of the patient at a plurality of heart rates, the plurality of cine loops including cine loops including frames of wall motion images and at least one cine loop including frames of perfusion images of at least one cardiac location;synchronizing the MRI wall motion cine loops to concurrently display images corresponding to the same portion of a cardiac cycle of the patient; andsimultaneously displaying both the wall motion cine loops and the at least one perfusion cine loop,wherein simultaneously displaying comprises simultaneously displaying a plurality of cine loops for differing locations associated with the heart of the patient for a single dosage of a stress inducing agent. 6. A method of displaying cardiac information of a patient, comprising: obtaining a plurality of MRI cine loops of the heart of the patient at a plurality of heart rates, the plurality of cine loops including cine loops including frames of wall motion images and at least one cine loop including frames of perfusion images of at least one cardiac location;synchronizing the MRI cine loops to concurrently display images corresponding to the same portion of a cardiac cycle of the patient; andsimultaneously displaying both wall motion cine loops and the at least one perfusion cine loop,wherein simultaneously displaying comprises si... | Wake Forest University Health Sciences | A61B005/05
US07847565 | B2 | 2010-12-07 | System for electrical impedance tomography and method thereof | 2006-07-28 | A system for electrical impedance tomography and method thereof are disclosed, by which electrical characteristics within a measurement target can be precisely detected. The present invention includes the steps of injection a current to a measurement target via at least one electrode pair selected form a plurality of electrodes (250) attached to the measurement target, detecting voltage of a surface of the measurement target using a plurality of voltmeters (260) connected to the electrodes that are not selected, respectively, adjusting gains of the voltmeters according to maximum values of the detected voltages, respectively, amplifying the detected voltages using the gain-adjusted voltmeters, respectively, and imaging an internal part of the measurement target based on the amplified voltages. | 1. A method of electrical impedance tomography, injecting the at least one comprising the steps of: injecting a current to a measurement target via at least one electrode pair selected from a plurality of electrodes attached to the measurement target;detecting voltage of a surface of the measurement target using a plurality of voltmeters connected to the electrodes that are not selected, respectively;removing noise included in the detected voltages based on slopes of the detected voltages;adjusting gains of the voltmeters according to maximum values of the detected voltages, respectively;amplifying the detected voltages using the gain-adjusted voltmeters, respectively; andimaging an internal part of the measurement target based on the amplified voltages. 2. The method of claim 1, the step of injecting the current to the measurement target via the at least one selected electrode pair, comprising the steps of selecting the electrode pair and a frequency;generating a voltage signal according to the selected frequency;converting the voltage signal to a current; andinjecting the current into the measurement target via the selected electrode pair. 3. The method of claim 1, the step of injecting the current to the measurement target via the at least one selected electrode pair, comprising the steps of converting a voltage signal to two currents differing in phase;calibrating the two currents to have a same amplitude and frequency; andinjecting the two currents into the measurement target via the selected electrode pair, respectively. 4. The method of claim 1, wherein if the slope of the detected voltage exceeds a threshold, a voltage of a section corresponding to the slope exceeding the threshold is replaced by a preset voltage value. 5. The method of claim 1, the step of adjusting the gains of the voltmeters according to the maximum values of the detected voltages, respectively, comprising the step of adjusting the gains of the voltmeters to have the maximum values of the detected voltages equal to or greater than 90% of maximum outputs of the voltmeters, respectively. 6. The method of claim 1, the step of imaging the internal part of the measurement target based on the amplified voltages, comprising the steps of: adjusting sizes of the amplified voltages according to gain values of the voltmeters, respectively; andimaging the internal part of the measurement target based on the adjusted voltage sizes. 7. A system for electrical impedance tomography, comprising: a plurality of electrodes attached to a measurement target for current injection and voltage detection;a current source supplying a current to a plurality of the electrodes;a board including a plurality of switches to selectively provide the current to at least one selected electrode pair from the current source;a plurality of voltmeters detecting voltages of a surface of the measurement target via the electrodes not selected, the voltmeters removing noises included in the detected voltag... | University-Industry Cooperation Group of Kyunghee University | G01R027/08
US07440929 | B2 | 2008-10-21 | Multimodal authorization method, system and device | 2006-05-26 | The invention includes methods, systems and devices for authenticating transactions. A method may begin by enrolling at least two biometric specimens to a database. A false acceptance ratio (“FAR”) is determined for each of the specimens, and authorization options are identified. A cost value is calculated for at least one of the options to provide a calculated cost value (“CCV”). The CCV may be a function of the FAR(s) of the specimen(s) corresponding to the option. An acceptable cost value range (“ACV range”) may be identified, and compared to the CCV. If it is determined that the CCV is in the ACV range, then the option is selected. If the CCV is in the ACV range, then a set of biometric samples is provided (the “sample set”). The sample set is compared to the biometric specimens, and it is determined whether the biometric samples match the biometric specimens. If the biometric samples match the biometric specimens, then the transaction is authorized. | 1. A method of authorizing a transaction, comprising: enrolling at least two biometric specimens, a first one of the specimens being a first type, and a second one of the specimens being a second type;determining a false acceptance ratio (“FAR”) for each of the specimens, wherein at least one of the false acceptance ratios is determined using the following equationFAR=PType⁢⁢I=P⁡(x∈RAu¦Im)=∫RAu⁢f⁡(x¦Im)⁢ⅆx;identifying authorization options, each option requiring a match to one or more of the biometric specimens;calculating a cost value for at least one of the options to provide a calculated cost value (“CCV”), wherein the CCV is a function of the FAR(s) of the specimen(s) corresponding to the option;identifying an acceptable cost value range (“ACV range”);comparing the CCV to the ACV range;determining whether the CCV is in the ACV range;selecting the option if the CCV is in the ACV range;providing a set of biometric sample(s) (the “sample set”), the sample set having biometric sample(s) of the same type(s) as those corresponding to the selected option;comparing the biometric sample(s) to the biometric specimen(s);determining whether the biometric sample(s) match the biometric specimen(s); andif the biometric sample(s) match the biometric specimen(s), then authorizing the transaction. 2. The method of claim 1, wherein the types are selected from fingerprint, facial image, retinal image, iris scan, hand geometry scan, voice print, signature, signature gait, keystroke tempo, blood vessel pattern, palm print, skin composition spectrum, lip shape and ear shape. 3. The method of claim 1, wherein a first type corresponds to a first body part and a second type corresponds to a second body part. 4. The method of claim 1, wherein the biometric specimens comprise a set, and the options comprise permutations of the set. 5. The method of claim 1, wherein at least one of the cost values is calculated using the following equationCost=c1+c2⁢FAR1+c3⁢∏i=12⁢FARi+…+cN+1⁢∏i=1N⁢FARi. 6. The method of claim 1, wherein the ACV range is provided by an administrator. 7. A computer readable memory device having stored thereon instructions that are executable by a computer to cause the computer to: enroll at least two biometric specimens, a first one of the specimens being a first type, and a second one of the specimens being a second type;determine a false acceptance ratio (“FAR”) for each of the specimens, wherein the instructions cause the computer to determine at least one of the false acceptance ratios using the following equationFAR=PType⁢⁢I=P⁡(x∈RAu¦Im)=∫RAu⁢f⁡(x¦Im)⁢ⅆx;identify authorization options, each option requiring a match to one or more of the biometric specimens;calculate a cost value for at least one of the options to provide a calculated cost value (“CCV”), wherein the CCV is a function of the FAR(s) of the specimen(s) corresponding to the option;identify an acceptable cost value range (“ACV range”);compare the CCV to the ACV range;determine whether... | Ultra-Scan Corporation | G07C009/00 G06K009/00 A61B005/00
US08175677 | B2 | 2012-05-08 | MRI-guided medical interventional systems and methods | 2008-06-06 | An MRI-guided interventional system includes a frame with a cooperating targeting cannula. The frame is configured to be secured to the body of a patient, and is configured to translate and rotate such that the targeting cannula can be positioned to a desired intrabody trajectory. The frame may include one or more MRI-visible fiducial markers that allow frame location/orientation to be determined within an MRI image. A plurality of user-activatable actuators are configured to translate and rotate the frame relative to the body of a patient so as to position the targeting cannula to a desired intrabody trajectory. The targeting cannula includes an axially-extending guide bore therethrough that is configured to guide placement of an interventional device in vivo. Various instrumentation and equipment, such as stimulation leads, ablation catheters, injection catheters, etc., can be inserted through the targeting cannula to execute diagnostic and/or surgical procedures. | 1. An MRI-guided interventional system, comprising: a frame that supports a cooperating targeting cannula that is movable along a Z-direction, wherein the frame is configured to translate and rotate such that the targeting cannula can be positioned to a desired intrabody trajectory, and wherein the targeting cannula includes a guide bore therethrough that is configured to guide placement of an interventional device in vivo, wherein the frame comprises: a base having a patient access aperture formed therein, wherein the base is configured to be secured to the body of a patient;a yoke movably mounted to the base and rotatable about a roll axis; anda platform movably mounted to the yoke and rotatable about a pitch axis,wherein an X-Y support table is movably mounted to the platform and is configured to translate in both an X-direction and Y-direction relative to the Z-direction; anda plurality of user-activatable actuators operably connected to the frame that are configured to translate and rotate the frame relative to the body of the patient so as to position the targeting cannula to a desired intrabody trajectory. 2. The system of claim 1, wherein the frame comprises: a roll actuator operably connected to the yoke and configured to rotate the yoke about the roll axis;a pitch actuator operably connected to the platform and configured to rotate the platform about the pitch axis;an X-direction actuator operably connected to the platform and configured to move the X-Y support table in the X-direction; anda Y-direction actuator operably connected to the platform and configured to move the X-Y support table in the Y-direction. 3. The system of claim 2, wherein the roll actuator, pitch actuator, X-direction actuator, and Y-direction actuator each have respective free ends which extend outwardly from the frame along a common direction. 4. The system of claim 1, wherein the base comprises a plurality of locations for attaching the base to a body of a patient via fasteners. 5. The system of claim 4, wherein each location comprises a plurality of adjacent apertures configured to receive a fastener therethrough. 6. The system of claim 1, wherein the base is configured to be secured to the skull of a patient about a burr hole formed therein, and wherein the targeting cannula guide bore is configured to guide intra-brain placement of a device in vivo. 7. The system of claim 6, wherein the base comprises a plurality of locations for attaching the base to a skull of a patient via fasteners. 8. The system of claim 1, wherein the yoke comprises a pair of spaced apart arcuate arms, and wherein the platform engages and moves along the arcuate arms when rotated about the pitch axis. 9. The system of claim 8, wherein at least one of the yoke arcuate arms comprises a thread pattern formed in a surface thereof, and wherein the pitch actuator comprises a rotatable worm with teeth configured to engage the thread pattern, and wherein rotation of the worm causes the ... | MRI Interventions, Inc. | A61B005/055
US08011508 | B2 | 2011-09-06 | Packaged systems for implanting markers in a patient and methods for manufacturing and using such systems | 2008-06-27 | Packaged systems for implanting a marker in a patient and methods for manufacturing and using such systems. In one embodiment, a packaged system comprises an introducer having a cannula and a stylet configured to be received in the cannula, a marker in the cannula, and a package having a sterile compartment. The marker can have a casing configured to be implanted in a patient and a resonating circuit in the casing. The resonating circuit can comprise a coil configured to wirelessly transmit a target signal in response to a wirelessly transmitted excitation signal. The introducer is contained within the sterile compartment. In another embodiment, the marker is not loaded in the introducer within the compartment of the package. | 1. A packaged system for implanting a marker in a patient, comprising: an introducer having a cannula and a stylet configured to be received in the cannula;a marker in the cannula, the marker having a casing configured to be implanted in a patient and a resonating circuit in the casing, the resonating circuit comprising a coil configured to wirelessly transmit a target signal in response to a wirelessly transmitted excitation signal; anda package having a sterile compartment, wherein the introducer is contained within the sterile compartment. 2. The packaged system of claim 1 wherein the marker comprises a first marker and the system further comprises a second marker in the cannula, and wherein the second marker comprises another casing configured to be implanted in the patient and another resonating circuit that wirelessly transmits another target signal in response to another wirelessly transmitted excitation signal. 3. The packaged system of claim 1 wherein: the marker comprises a first marker, the casing comprises a first casing, and the resonating circuit comprises a first resonating circuit configured to wirelessly transmit a first target signal at a first frequency in response to a first excitation signal; andthe system further comprises a second marker having a second casing and a second resonating circuit in the second casing, the second resonating circuit being configured to wirelessly transmit a second signal at a second frequency in response to a second excitation signal, wherein the first frequency is different than the second frequency. 4. The packaged system of claim 1 wherein the marker is the sole marker in the cannula. 5. The packaged system of claim 1 wherein: the introducer comprises a first introducer and the marker comprises a first marker, the casing comprises a first casing, and the resonating circuit comprises a first resonating circuit configured to produce a wirelessly transmitted first target signal in response to a first excitation signal; andthe system further comprising a second introducer in the package and a second marker in the second introducer, the second marker comprising a second casing configured to be implanted in the patient and a second resonating circuit in the second casing, and the second resonating circuit being configured to wirelessly transmit a second target signal in response to a second excitation signal. 6. The packaged system of claim 5 wherein the package comprises a tray and a cover sealed to the tray, the tray having a first introducer retainer in which the first introducer is positioned and a second introducer retainer in which the second introducer is positioned. 7. The packaged system of claim 1 wherein: the introducer comprises a first introducer and the marker comprises a first marker, the casing comprises a first casing, and the resonating circuit comprises a first resonating circuit configured to produce a wirelessly transmitted first target signal at a first frequency in respon... | Calypso Medical Technologies, Inc. | B65D083/10 A61M005/00 A61B017/34
US07945312 | B2 | 2011-05-17 | Multisensor probe for tissue identification | 2008-08-22 | A multisensor probe for continuous real-time tissue identification. The multisensor probe includes a tissue penetrating needle, a plurality of sensors useful in characterizing tissue and a position sensor to measure the depth of the needle into the tissue being diagnosed. The sensors include but are not limited to an optical scattering and absorption spectroscopy sensor, an optical coherence domain reflectometry sensor, an electrical impedance sensor, a temperature sensor, a pO2 sensor, a chemical sensor and other sensors useful in identifying tissue. The sensors may take the form of a plurality of optical fibers extending through said needle. A retractable sheath may be disposed around the distal section of the needle to protect the needle when not in use. The sheath retracts when the probe is inserted into tissue and the position of the sheath relative to the needle may be measured to determine the needle's depth. Systems and methods for tissue identification are also provided. | 1. A multisensor probe for tissue identification, the probe connectable to a controller via a cable, the probe comprising: a handle;an elongate body having a distal tip and a lumen extending through the elongate body to the distal tip;a slideable sheath disposed over a distal section of the elongate body and being retractable from the elongate body toward the handle;a resilient member configured to bias the sheath toward the distal tip of the elongate body; anda plurality of optical fibers extending from the controller, through the cable, through the lumen, to the distal tip of the elongate body when the probe is connected to the controller, wherein at least two of the plurality of optical fibers are optical scattering and absorption spectroscopy fiber optics and wherein at least one of the plurality of optical fibers is an optical coherence domain reflectometry fiber optic, and wherein a position sensor is incorporated into the probe, the position sensor adapted to measure a depth to which the elongate body is inserted into tissue, the position sensor configured to read a position of the sheath relative to the elongate body. 2. The multisensor probe of claim 1 wherein the position sensor is a sensor selected from the group consisting of a resistive sensor and a linear optical encoder. 3. The multisensor probe of claim 1 having a conducting element extending through the lumen. 4. The multisensory probe of claim 1, further comprising a marking switch to identify a location in the tissue the distal section of the elongate body is inserted into the tissue. 5. The multisensory probe of claim 1, wherein the resilient member is a spring. 6. A tissue detection system comprising: a probe, the probe comprising a multisensor needle comprising a plurality of optical fibers, the probe further comprising: a handle;a slideable sheath disposed over a distal section of the needle and being retractable from the needle toward the handle;a resilient member configured to bias the sheath over the distal section of the needle;a position sensor adapted to sense depth of the needle inserted into tissue using the sheath wherein the sensor reads a position of the sheath relative to the needle; anda controller configured to deliver and collect light through the plurality of optical fibers wherein at least one of the plurality of optical fibers is utilized as an optical coherence domain reflectometry sensor configured for measuring backscattered and reflected light as a function of depth into the tissue and wherein at least one the optical fibers is utilized for optical scattering and absorption spectroscopy. 7. The system of claim 6 further comprising at least one electrode for sensing electrical information about the tissue. 8. The system of claim 6 wherein the sheath retracts when the needle is inserted into the tissue. 9. The system of claim 6, wherein at least two of the plurality of optical fibers are utilized for optical scattering and absorption spectroscopy... | Bioluminate, Inc. | A61B006/00 A61B005/00
US07928850 | B2 | 2011-04-19 | Analyte monitoring system and methods | 2008-05-08 | Methods and systems for providing data communication in medical systems are disclosed. | 1. An analyte monitoring device, comprising: a counter; anda data processing section coupled to the counter, wherein the data processing section is configured to: establish a communication range between the analyte monitoring device and a remote device when the remote device is placed within a predetermined distance from the analyte monitoring device;determine when a transcutaneously positioned analyte sensor is initialized;increment a count stored in the counter based on the determination;associate the incremented count with the transcutaneously positioned analyte sensor; andtransmit, using a communication module, the incremented count to the remote device when the remote device is placed within the established communication range, wherein the incremented count is used to verify the identity of the new transcutaneously positioned analyte sensor. 2. The device of claim 1 wherein the counter includes a nonvolatile memory unit. 3. A control device for communicating with an analyte monitoring unit, comprising: a data communication section configured to establish a communication range with the analyte monitoring unit when the control unit is placed within a predetermined distance from the analyte monitoring unit and further configured to receive data, including a count information associated with a transcutaneously positioned analyte sensor, when the control device is placed within the established communication range; anda processing unit coupled to the data communication section to process the received data, the processing unit further configured to compare the count information in the received data with a stored count to verify the identity of the transcutaneously positioned analyte sensor, and when it is determined that the received count information is different than the stored count, the processing unit is configured to generate one or more signals related to the status of the transcutaneously positioned analyte sensor including the identity of the transcutaneously positioned analyte sensor. 4. The device of claim 3 wherein the generated one or more signals include a command to disable the data communication section. 5. The device of claim 3 wherein the generated one or more signals is output to a user. 6. The device of claim 5 including a display unit and wherein the generated one or more signals is displayed on the display unit. 7. The device of claim 5 wherein the one or more signals output includes one or more of a visual indicator, an audible indicator or a vibratory indicator. 8. A method, comprising: establishing a communication range between an analyte monitoring device and a remote device when the remote device is placed within a predetermined distance from the analyte monitoring device;detecting an initialization of a transcutaneously positioned analyte sensor;associating a count with the transcutaneously positioned analyte sensor;transmitting the count associated with the transcutaneously positioned analyte sensor to the remote... | Abbott Diabetes Care Inc. | G08B023/00 A61B005/05
US09063203 | B2 | 2015-06-23 | Method for determining at least one pet parameter | 2008-05-21 | A method is disclosed for determining at least one PET parameter for a PET examination in a combined magnetic resonance/PET scanner designed for isocentric measurement. In at least one embodiment, at least one acquisition time for the PET examination is determined as a PET parameter from magnetic resonance data, recorded for determining patient-specific examination parameters for a magnetic resonance examination, and/or from at least one patient-specific information value derived therefrom, and is displayed and/or used for controlling the scanner. | 1. A method for controlling a position emission tomography (PET) examination in a combined magnetic resonance/PET scanner designed for isocentric measurement, the method comprising: recording magnetic resonance data with the combined magnetic resonance/PET scanner;deriving patient-specific information from the magnetic resonance data, wherein the patient specific information includes a specific absorption rate;determining PET parameters from the specific absorption rate, wherein the PET parameters include at least one acquisition time for the PET examination; andcontrolling the combined magnetic resonance/PET scanner to perform the PET examination with the determined PET parameters. 2. The method as claimed in claim 1, wherein the patient-specific information further includes at least one of a volume, a mass of the patient and a mass of a part of the patient located in the examination area. 3. The method as claimed in claim 1, wherein the PET parameters further include at least one PET parameter, which is determined as a function of a target volume of interest. 4. The method as claimed in claim 3, wherein the patient-specific information further includes the target volume of interest. 5. The method as claimed in claim 3, wherein at least one of the magnetic resonance data and the patient-specific information are linked to the PET parameters via a look-up table. 6. The method as claimed in claim 3, wherein at least one of the magnetic resonance data and the patient-specific information are linked to the PET parameters via a formula. 7. The method as claimed in claim 1, wherein at least one of the magnetic resonance data and the patient-specific information are linked to the PET parameters via a look-up table. 8. The method as claimed in claim 1, wherein at least one of the magnetic resonance data and the patient-specific information are linked to the PET parameters via a formula. 9. The method as claimed in claim 1, wherein the magnetic resonance data is recorded by a fast full-body measurement by way of magnetic resonance. 10. The method as claimed in claim 9, wherein only a coronal or sagittal section is recorded as the magnetic resonance data. 11. The method as claimed in claim 9, wherein three-dimensional magnetic resonance data are recorded. 12. A medical examination device, comprising: a combined magnetic resonance/PET scanner with a control device, the control device being configured to carry out the method as claimed in claim 1. 13. A tangible computer readable medium including program segments for, when executed on a computer device, causing the computer device to implement the method of claim 1. 14. A medical examination device, comprising: a combined magnetic resonance/positron emission tomography (PET) scanner including, means for recording magnetic resonance data with the combined magnetic resonance/positron emission tomography (PET) scanner;means for deriving patient-specific information from the magnetic resonance data,... | SIEMENS AKTIENGESELLSCHAFT | A61B005/055 G01R033/48 A61B006/03 A61B006/00 G01T001/16 A61B006/02 G01T001/161 G01R033/54
US08359084 | B2 | 2013-01-22 | Containment tubes and imaging systems employing same | 2008-03-27 | An isolation system with imaging or radiation therapy capability is disclosed. At least one containment barrier (14, 15, 16, 17) defines an isolation region (10). An imaging or therapy system (20) is disposed outside of the isolation region. The containment barrier includes a substantially hollow tubular extension (24, 42, 44, 124, 224, 324) protruding away from the isolation region (10). The substantially hollow tubular extension surrounds an interior volume (26) that is in fluid communication with the isolation region and is in fluid isolation from the imaging or therapy system. The substantially hollow tubular extension is made at least partially of a material providing operative communication between the imaging or therapy system and the interior volume of the substantially hollow tubular extension. | 1. An imaging or treatment method comprising: surrounding a human subject with a sealed container while the human subject is disposed in an isolation region;transporting the sealed container from the isolation region to an imaging or treatment apparatus with the human subject disposed in and surrounded by the sealed container; andperforming imaging or treatment of the human subject surrounded by the sealed container using the imaging or treatment apparatus;wherein the sealed container is sized and shaped for insertion into a cylindrical bore of the imaging or treatment apparatus and the transporting includes conformably sliding the sealed container into the cylindrical bore of the imaging or treatment apparatus. 2. The imaging or treatment method as set forth in claim 1, wherein the transporting further comprises: after the surrounding, decontaminating an exterior of the sealed container. 3. The imaging or treatment method as set forth in claim 1, further comprising: disposing the human subject in the container;sealing the container using a sealing portion configured to seal the container with a human subject inside; andafter the sealing, decontaminating the container. 4. The imaging or treatment method as set forth in claim 1, wherein the transporting includes transferring the sealed container with the human subject disposed in and surrounded by the sealed container out of the isolation region via one or more airlocks. 5. An imaging or treatment method comprising: surrounding a human subject with a sealed container while the subject is disposed in an isolation region;transporting the human subject surrounded by the sealed container from the isolation region to an imaging or treatment apparatus;performing imaging or treatment of the human subject surrounded by the sealed container using the imaging or treatment apparatus; andproviding breathable air to the human subject surrounded by the sealed container by providing a self-contained breathing apparatus within the sealed container. 6. An imaging or treatment method comprising: extending a flexible sheath or containment bag away from a wall separating an isolation region from an imaging or therapy system into an extended configuration, the flexible sheath or containment bag in the extended configuration defining an interior that overlaps an examination region of the imaging or therapy system and is isolated from the imaging or therapy system by the flexible sheath or containment bag, the interior being in fluid communication with the isolation region via an opening in the wall sized to admit a subject from the isolation region into the interior overlapping the examination region, wherein the flexible sheath or containment bag is elastic and the extending comprises biasing the elastic flexible sheath or containment bag against the elasticity into the extended configuration and retaining the elastic flexible sheath or containment bag in the extended configuration using suspenders or anchors dis... | Koninklijke PHilips Electronics N.V. | A61B005/055
US08696586 | B2 | 2014-04-15 | Method and device for measuring blood flow in the bone after a fracture | 2008-05-05 | A method and device for measuring blood flow in the bone after a fracture, wherein a hole is drilled into the bone. The method includes the steps of inserting the metal sleeve of an osteoscope into the hole; inserting the optics of an osteoscope into the metal sleeve and focused on the cavity created by the drill; filling the device and the cavity with a physiological solution such that the pressure of the solution is above the systolic blood pressure; decreasing the pressure of the solution while observing the cavity through the osteoscope, and recording the pressure in the system at the commencement of bleeding. | 1. A method for determining circulation of blood flow in a bone after a femoral neck fracture, comprising: a) drilling a hole into the bone with a drill;b) inserting a metal sleeve of an osteoscope into the hole;c) inserting optics of the osteoscope into the metal sleeve, the optics being focused on a cavity created by the drill;d) filling the device and the cavity with a physiological solution such that a pressure of the solution is above a systolic blood pressure of a patient;e) decreasing the pressure of the solution while observing the cavity through the osteoscope;f) recording a system pressure when bleeding commences; andg) determining acceptable circulation when a difference between the systolic pressure and the pressure recorded in step f) is less than 60 mm Hg, and unacceptable circulation when a difference in the pressure is more than 60 mm Hg. 2. The method as claimed in claim 1, wherein the metal sleeve is inserted into the hole to leave free the wall of a small part of the cavity. 3. The method as claimed in claim 2, wherein a diameter of the small part of the cavity is drilled to be smaller, than an internal diameter of the sleeve. 4. The method as claimed in claim 2, wherein a surface of the small part of the cavity is at least 35 mm2. 5. The method as claimed in claim 2, wherein a surface of the small part of the cavity is at least 30 mm2. 6. The method as claimed in claim 1, wherein the optics is arranged in the sleeve to be about 2 mm shorter, than the sleeve. 7. The method as claimed in claim 1, wherein a second drill is applied through the hole of the sleeve thereby removing fragments of the bone. 8. The method as claimed in claim 1, wherein the hole is flushed out with physiological solution, after step b). 9. The method as claimed in claim 8, wherein the drill is provided with an external diameter approximately the same size as an external diameter of the sleeve. 10. The method as claimed in claim 9, wherein the drill is provided with an end section having smaller diameter, than the internal diameter of the sleeve. 11. A device for determining blood flow in a bone after a femoral neck fracture, comprising: a metal sleeve of an osteoscope to be inserted into a bore drilled in the fractured bone;an optics of an osteoscope to be inserted into said metal sleeve and connected to a visual display of the osteoscope;a reservoir with physiological solution connected to the metal sleeve;means for controlling the pressure of the solution;a manometer for monitoring the pressure of the solution; anda manometer for measuring the blood pressure of the patient,wherein the device is configured for determining acceptable circulation when a difference between a systolic pressure of a patient and a system pressure recorded when bleeding commences is less than 60 mm Hg, and unacceptable circulation when a difference in the pressure is more than 60 mm Hg. 12. The device as claimed in claim 11, wherein the metal sleeve has a flow in con... | - | A61B001/317 A61B005/026
US08396534 | B2 | 2013-03-12 | Intravital-information imaging apparatus | 2008-05-30 | An intravital-information imaging apparatus includes light sources for irradiating a living body with light, and acoustic-wave detectors that detect acoustic waves generated from a light absorber having absorbed part of energy of the light from the light sources, the light absorber existing in the living body. Furthermore, the intravital-information imaging apparatus includes signal processors that calculate information representing a distribution of optical characteristic values of the living body using sound pressures of the acoustic waves generated from the light absorber. | 1. An intravital-information imaging apparatus that performs imaging using an acoustic wave generated by irradiating a subject with light, the intravital-information imaging apparatus comprising: a light source configured to irradiate the subject with light;an acoustic-wave detector configured to detect an acoustic wave generated from a light absorber that has absorbed a part of energy of the light from said light source, the light absorber existing inside the subject; anda signal processor configured to calculate an average effective attenuation coefficient of the subject using relative-position information representing a relative position between the light absorber and an irradiation region of the subject and sound pressures of the acoustic wave, the relative-position information and the sound pressures being obtained by analyzing electric signals corresponding to the acoustic wave detected by said acoustic-wave detector. 2. The intravital-information imaging apparatus according to claim 1, wherein said light source includes two or more light sources, and said signal processor is further configured to calculate the average effective attenuation coefficient using a first distance, which is a distance measured from a first one of said light sources, and a second distance, which is a distance measured from a second one of said light sources, said first and second light sources being provided at mutually different positions. 3. The intravital-information imaging apparatus according to claim 1, further comprising an optical-path converter configured to convert an optical path of the light from said light source, wherein said signal processor is further configured to calculate the average effective attenuation coefficient using a first distance, which is a distance measured from a first position, and a second distance, which is a distance measured from a second position, the second position being varied from the first position by said optical-path converter. 4. The intravital-information imaging apparatus according to claim 1, wherein said light source includes two or more light sources, and said signal processor is further configured to calculate the average effective attenuation coefficient using differences between sound pressures obtained by simultaneous irradiation with light from n light sources and sound pressures obtained by simultaneous irradiation with light from n−1 light sources. 5. The intravital-information imaging apparatus according to claim 1, wherein said signal processor is further configured to calculate an average effective attenuation coefficient of the living body using a straight line identified on the basis of coordinate values represented by the relative-position information and logarithms of the sound pressures of the acoustic wave generated from the light absorber. 6. The intravital-information imaging apparatus according to claim 5, wherein the straight line is identified using the least square method on the basis of... | Canon Kabushiki Kaisha | A61B006/00
US08818497 | B2 | 2014-08-26 | Biological signal processing unit, wireless memory, biological signal processing system, and control system of device to be controlled | 2005-07-12 | It is an object to provide a means for easily operating various devices to be controlled (i.e. a word processor or a car) using a brain wave signal of a patient with Alzheimer's disease or a psychological disorder, and for supporting the active performance of a user by employing wireless communication with a memory (wireless memory) capable of communicating without wires in which a command on the human movement is memorized. A present biological signal processing unit has an electrode for detecting a biological signal (an electric signal) from a living body, or for transmitting an electric signal into a living body, an interface, and an antenna which can communicate with an external device (i.e. wireless memory, reader/writer); therefore, convenience of the user is improved in the case where a user utilizes an electronic device or the like. In addition, the active performance of a user can be supported. | 1. A biological signal processing unit comprising: an electrode for at least one of detecting a biological signal from an inside of a living body and transmitting an electric signal to the inside of the living body;a plastic substrate having flexibility;an antenna over the plastic substrate having flexibility, wherein the antenna is configured to wirelessly communicate with a wireless memory configured to be attached to the living body, which is located outside the biological signal processing unit;an interface comprising a rectification circuit provided between the electrode and the antenna and over the plastic substrate having flexibility; anda processor over the plastic substrate having flexibility, wherein the processor comprises a transistor electrically connected to the antenna, the transistor comprising: a first insulating film;a first electrode over the first insulating film;a second insulating film over the first insulating film and the first electrode;a semiconductor layer over the second insulating film, so as to overlap with the first electrode, wherein the semiconductor layer comprises a channel region, a first impurity region, and a second impurity region between the channel region and the first impurity region, and wherein a concentration of an impurity element in the second impurity region is lower than a concentration of the impurity element in the first impurity region;a third insulating film over the semiconductor layer;a second electrode over the third insulating film, so as to overlap with the semiconductor layer;a third electrode over the second electrode, wherein a width of the third electrode is shorter than a width of the second electrode in a channel length direction;a fourth insulating film over the third insulating film, second electrode and the third electrode; anda fourth electrode over the fourth insulating film, the fourth electrode electrically connected to the first impurity region through a contact hole in the third insulating film and the fourth insulating film,wherein at least the antenna, the rectification circuit, and the processor are provided over the plastic substrate having flexibility with an adhesive agent interposed therebetween. 2. The biological signal processing unit according to claim 1, wherein the plastic substrate comprises one selected from the group consisting of polyethylene terephthalate, polycarbonate, polyarylate, and polyethersulfone. 3. The biological signal processing unit according to claim 1, wherein the second impurity region is overlapped with the second electrode and is not overlapped with the third electrode. 4. A biological signal processing unit comprising: an electrode for transmitting an electric signal including information to an inside of a living body;a plastic substrate having flexibility;an antenna over the plastic substrate having flexibility, the antenna configured to wirelessly communicate with a reader/writer located outside of the living body;an interface compris... | Semiconductor Energy Laboratory Co., Ltd. | A61B005/04 A61F004/00 A61B005/0478 G06F003/01 A61B005/00
US20080255431 | A1 | 2008-10-16 | Early Detection and Warning of Harmful Agents | 2005-08-04 | A method of detecting the presence of harmful agents in the environment is disclosed. The method comprises: providing a plurality of portable agent detectors, placing the detectors on vectors in the environment, connecting the detectors over a communication network, and, upon detection of an agent by a first detector, signaling detectors identified as being nearby to carry out corresponding agent detection tests, thereby to confirm or localize an initial agent detection. | 1. A method of detecting the presence of harmful agents in the environment comprising: providing a plurality of portable agent detectors,placing said detectors on vectors in said environment,connecting said detectors over a communication network,upon detection of an agent by a first detector, signaling detectors identified as being nearby to carry out corresponding agent detection tests,thereby to confirm or localize an initial agent detection. 2. The method of claim 1, further comprising monitoring said detectors at a central location via said communication network. 3. The method of claim 2, wherein said signaling said nearby detectors comprises remotely activating said nearby detectors. 4. A method of detecting the presence of harmful agents in the environment comprising: providing a plurality of portable agent detectors able to generate a detection event in the presence of any of said harmful agents,placing said detectors on vectors in said environment,connecting said detectors over a communication network,at a central location monitoring said detectors to identify clustering of said detection events, said clustering being to confirm or localize an initial agent detection. 5. The method of claim 4, further comprising, upon a detection event generated by a first detector, signaling detectors identified as being nearby to said first detector, to carry out corresponding agent detection tests. 6. The method of claim 2 further comprising, upon a detection event generated by a first detector, alerting vectors identified as being nearby to said first detector. 7. A method of detecting the presence of harmful agents, comprising: providing a plurality of portable agent detectors able to generate a detection event in the presence of any of said harmful agents,associating said detectors with cellular communication units to allow central monitoring of said detectors, andplacing said detectors and cellular communication units on vectors in said environment. 8. The method of claim 7, further comprising associating said detectors with positioning units, and monitoring locations of individual detectors using a positioning network. 9. The method of claim 2 further comprising obtaining atmospheric conditions at a location of said first detector and using said atmospheric condition to predict a propagation path of said agent. 10. A portable agent detector for releasing into an environment to detect harmful agents released into said environment, the agent detector comprising: a removable sensing cassette comprising sensing wetware; anda processing unit for processing signals from said sensing cassette. 11. The portable agent detector of claim 10, further comprising an activation unit for activating or selecting an operational mode of the portable agent detector. 12. The portable agent detector of claim 11, further comprising a user interface, wherein said activation unit is controllable by said user interface. 13. The portable agent detector of claim ... | Ramot At Tel Aviv Univeristy Ltd. | A61B005/00 G08B021/00
US20080300502 | A1 | 2008-12-04 | Method Of Correlating Internal Tissue Movement | 2005-06-23 | Internal tissue movement is correlated with the external body surface movement by tracking external surface movement at a plurality of surface point (12), simultaneously imaging internal tissue movement and correlating the external and internal movement using partial least squares regression to obtain a correlation model. In subsequent techniques, internal tissue movement can be predicted from measured external surface movement using the correlation model. | 1. A method of correlating internal tissue movement with external surface movement comprising tracking surface movement at a plurality of surface points, simultaneously obtaining tissue movement and correlating the surface movement to the tissue movement by a regression technique to obtain a correlation model. 2. A method as claimed in claim 1 in which the regression technique is partial least squares regression (PLSR). 3. A method as claimed in claim 1 in which the internal tissue movement comprises cardiac movement. 4. A method as claimed in claim 1 in which the surface movement is respiratory induced. 5. A method as claimed in claim 1 in which the surface movement is tracked using a tension jacket. 6. A method as claimed in claim 1 in which the surface movement is tracked remotely. 7. A method as claimed in claim 6 in which the surface movement is tracked optically. 8. A method as claimed in claim 1 in which the surface movement is tracked by measuring strain or curvature or both. 9. A method as claimed in claim 1 in which tissue movement is obtained using at least one of 3D magnetic resonance, Positron Emission Tomography, Computer Tomography or 3D Echo Cardiography. 10. A method of modeling internal tissue movement from surface movement comprising tracking surface movement at a plurality of surface points and modeling correlated internal tissue movement therefrom using a correlation model. 11. (canceled) 12. An internal tissue movement correlation apparatus comprising an external surface movement tracker reference having a plurality of trackable points. 13. An apparatus as claimed in claim 12 in which the reference comprises a garment on which the trackable points are provided. 14. An apparatus as claimed in claim 13 in which the trackable points comprise strain sensors. 15. An apparatus as claimed in claim 12 in which the trackable points are remotely trackable. 16. An apparatus as claimed in claim 12 in which the trackable points are optically trackable. 17. An apparatus as claimed in claim 12 in which the garment includes one or more optical fibres for sensing surface movement. 18. An apparatus as claimed in claim 17, the optical fibres being sensitised to detect bending thereof. 19. An apparatus as claimed in claim 18, the optical fibres including a longitudinal portion with a segment of its cross sectional profile being removed. 20. An apparatus as claimed in claim 12 further comprising a tracker apparatus arranged to track movement of the trackable points. 21. An apparatus as claimed in claim 12 further comprising an apparatus arranged to obtain movement of internal tissue. 22. An apparatus as claimed in claim 21 further comprising a processor arranged to process the obtained internal tissue movement and tracked external movement and derive a correlation model. 23. (canceled) 24. A method of correcting an internal tissue movement image using a related external surface movement comprising obtaining an imaged repre... | IMPERIAL INNOVATIONS LIMITED | A61B005/113
US20080262320 | A1 | 2008-10-23 | System for Monitoring a Physical Parameter of a Subject | 2006-06-28 | A system for monitoring a physical parameter of a subject includes a sensor, a transmitter, a repeater, and a base station. The sensor is located within the body of the subject and is operable to sense a physical parameter of the subject. The transmitter is electrically connected to the sensor, and is operable to transmit a spread spectrum encoded signal using a digital spreading code, the encoded signal carrying information indicative of the sensed physical parameter and being transmitted to a local region outside the body. The repeater is located in the local region, is movable with the subject, and is operable to receive the transmitted encoded signal and retransmit it to a remote region outside the body. The base station is located in the remote region and is operable to receive and decode the retransmitted encoded signal. | 1. A system for monitoring a physical parameter of a subject, the system comprising: a sensor associated with the body of the subject and operable to sense a physical parameter of the subject;a transmitter electrically connected to the sensor and operable to transmit a spread spectrum encoded signal using a digital spreading code, the encoded signal carrying information indicative of the sensed physical parameter and being transmitted to a local region outside the body;a repeater located in the local region and operable to receive the transmitted encoded signal and retransmit it to a remote region outside the body; anda base station located in the remote region and operable to receive and decode the retransmitted encoded signal, and remotely monitor the physical parameter of the subject. 2. The system of claim 1, wherein the sensor is one selected from the group including a temperature sensor, a heart rate sensor, an accelerometer, a pressure sensor, a pH sensor, a pO2 sensor, and an imaging sensor. 3. The system of claim 1, wherein the transmitter is a radio frequency transmitter. 4. The system of claim 1, wherein the transmitter is a transponder that transmits the encoded signal in response to an interrogation signal. 5. The system of claim 4, wherein the repeater generates the interrogation signal for the transponder. 6. The system of claim 5, wherein the repeater generates the interrogation signal for the transponder upon a request from the base station. 7. The system of claim 4, wherein the base station generates the interrogation signal for the transponder. 8. The system of claim 1, wherein the base station is operable to monitor a physical parameter of each of a plurality of subjects wherein each subject is associated with an individual transmitter having a corresponding identification code. 9. The system of claim 1, wherein the sensor and the transmitter are located within a swallowable capsule. 10. The system of claim 1, wherein power to the transmitter can be controlled by an activation signal provided by one of the repeater and the base station. 11. A system for monitoring a physical parameter of a subject, the system comprising: a sensor located within the body of the subject and operable to sense a physical parameter of the subject;a transmitter electrically connected to the sensor and operable to transmit a signal to a local region outside the body, the transmitted signal including information indicative of the sensed physical parameter;a repeater located in the local region, movable with the subject, and operable to receive the transmitted signal and retransmit it to a remote region outside the body; anda base station located in the remote region and operable to receive the retransmitted signal, and remotely monitor the physical parameter of the subject. 12. The system of claim 11, wherein the sensor is one selected from the group including a temperature sensor, a heart rate sensor, an accelerometer, a pressure sensor, a... | - | A61B005/00
US07833165 | B2 | 2010-11-16 | System for monitoring neural shunt function and associated methods | 2006-08-15 | A system and method for evaluating neural shunt functionality is provided. Accordingly, in one aspect a method for monitoring neural shunt functionality may include measuring a first intracranial pressure pulse inside the shunt, measuring a second intracranial pressure pulse outside the shunt, and comparing pulsatile characteristics from the first measurement to the second measurement in order to determine shunt functionality. | 1. A method for monitoring neural shunt functionality, comprising: measuring a first intracranial pressure pulse inside the shunt using a first sensor;measuring a second intracranial pressure pulse outside the shunt using a second sensor; andcomparing pulsatile characteristics from the first measurement to the second measurement using a computational device in order to determine shunt functionality. 2. The method of claim 1, wherein measuring inside the shunt further includes inserting the first sensor into the shunt to measure the first intracranial pressure pulse. 3. The method of claim 2, wherein inserting the first sensor into the shunt further includes inserting the first sensor into a structure that is fluidically coupled to the shunt. 4. The method of claim 1, wherein measuring outside the shunt further includes measuring the second intracranial pressure pulse from a location outside the cranium. 5. The method of claim 4, wherein measuring from a location outside the cranium includes measuring the second intracranial pressure pulse from tympanic membrane displacements. 6. The method of claim 4, wherein measuring from a location outside the cranium includes measuring the second intracranial pressure pulse from a supraorbital artery. 7. The method of claim 1, wherein the pulsatile characteristics include pulse amplitude. 8. The method of claim 1, wherein the pulsatile characteristics include pulse risetime or pulse falltime. 9. The method of claim 1, wherein comparing pulsatile characteristics further includes determining a correlation measurement between the first intracranial pressure pulse and the second intracranial pressure pulse. 10. The method of claim 1, further comprising clearing an occluded portion of the neural shunt. | - | A61B005/00 A61B005/02
US07435223 | B2 | 2008-10-14 | Identifying infants at risk for sudden infant death syndrome | 2007-03-27 | A method for identifying infants at risk for SIDS includes applying electrodes to an infant, receiving electrical signals from the electrodes, analyzing the received electrical signals to measure alternans of a heart of the infant, and identifying whether the infant is at risk for SIDS.;A system for identifying infants at risk for SIDS includes an input unit configured to receive electrical signals from electrodes applied to an infant, a processor connected to the input unit and configured to process the received electrical signals to measure alternans of a heart of the infant, and a comparator configured to compare the measured alternans with alternans in a population of infants. | 1. A system for identifying infants at risk for SIDS, comprising: an input unit configured to receive electrical signals from electrodes applied to an infant;a processor connected to the input unit and configured to process the received electrical signals to measure alternans of a heart of the infant and configured to compare the measured alternans with data associated with alternans in a population of infants and identify, based on the results of the comparison, whether an infant is at risk for SIDS. 2. The system of claim 1 wherein: the processor is configured to process the received electrical signals to measure T-wave alternans of the heart of the infant and is configured to compare the measured T-wave alternans with data associated with T-wave alternans in a population of infants. 3. The system of claim 1 wherein the processor is further configured to process the received electrical signals to measure a QT interval of the heart of the infant. 4. The system of claim 3 wherein the processor is configured to compare the measured QT interval with data associated with QT intervals in a population of infants. 5. The system of claim 1 wherein the received electrical signals comprise electrocardiogram signals. 6. The system of claim 1 further comprising an analog-to-digital converter configured to sample the received electrical signals at a frequency less than or equal to twice a frequency corresponding to alternans to generate sample points. 7. The system of claim 6 wherein the processor is configured to compute an alternating energy at each of the sample points for a series of beats. 8. The system of claim 7 wherein the processor comprises: a matrix constructor configured to construct sample point matrices, having rows and columns, from the generated sample points;an adder configured to sum the alternating energy over a set of sample points to generate a total alternating energy; anda divider configured to normalize the total alternating energy with respect to an energy of an average waveform. 9. The system of claim 6, wherein the processor is configured to create a processed signal having an asymmetric spectrum, and to process the processed signal to generate an indication of alternans in the received electrical signals. 10. The system of claim 9 wherein the processor is configured to create the processed signal as an analytical signal. 11. The system of claim 10 wherein the processor is configured to create the processed signal as an analytical signal by generating a frequency domain representation of the received electrical signals, modifying the frequency domain representation to remove components corresponding to negative frequencies, and generating the analytical signal as a time domain representation of the modified frequency domain representation. 12. The system of claim 1 wherein: the input unit comprises circuitry configured to receive an electrocardiogram signal; andthe system further comprises an electrode connected to the i... | Cambridge Heart, Inc. | A61B005/0452
US20080249376 | A1 | 2008-10-09 | Distributed Patient Monitoring System | 2008-04-04 | A distributed patient monitoring system for visually monitoring patients and patient parameters using portable processing devices in different remote locations includes a monitoring processor. The monitoring processor is responsive to user initiated commands from multiple different portable processing devices in different remote locations and includes an input processor and a data processor. The input processor acquires vital sign parameters and associated video data representative of multiple sequences of video images of corresponding multiple different patients. The data processor processes the vital sign parameters and associated video data to provide processed first video data representing an image sequence including a composite image including a first area showing live video of a selected first patient and a second area presenting vital sign parameters of the selected first patient. The data processor also processes the vital sign parameters and associated video data to provide processed second video data representing an image sequence including a composite image including a first area showing live video of a selected second patient and a second area presenting vital sign parameters of the selected second patient. A communication network has bandwidth sufficient to communicate the processed first video data and second video data to first and second portable processing devices respectively of the multiple different portable processing devices in different remote locations in response to commands received from the first and second portable processing devices respectively. | 1. A distributed patient monitoring system for visually monitoring patients and patient parameters using a plurality of portable processing devices in different remote locations, comprising: a monitoring processor responsive to user initiated commands from a plurality of different portable processing devices in different remote locations and including, an input processor for acquiring vital sign parameters and associated video and audio data and for controlling the selected viewing of said data representative of plurality of sequences of video images of a corresponding plurality of different patients; anda data processor for processing said vital sign parameters and associated video data to provide, processed first video data representing an image sequence including a composite image including a first area showing live video of a selected first patient and a second area presenting vital sign parameters of said selected first patient andprocessed second video data representing an image sequence including a composite image including a first area showing live video of a selected second patient and a second area presenting vital sign parameters of said selected second patient; anda communication network of bandwidth sufficient to communicate said processed first video data and second video data to first and second portable processing devices respectively of said plurality of different portable processing devices in different remote locations in response to commands received from said first and second portable processing devices respectively. 2. A system according to claim 1, wherein said data processor processes said vital sign parameters and associated video data to provide processed first and second video data by encoding said video data with a compression function. 3. A system according to claim 2, wherein said compression function is compatible with at least one of, (a) MPEG-4, (b) MPEG-2 and (c) DIVX. 4. A system according to claim 1, wherein said input processor acquires audio data of said plurality of different patients,said data processor processes said audio data to provide processed audio data andsaid communication network is of bandwidth sufficient to communicate said processed first video data and second video data and audio data to said first and second portable processing devices respectively. 5. A system according to claim 4, wherein said input processor acquires said audio data from a plurality of different microphones in patient rooms associated with said plurality of different patients. 6. A system according to claim 4, wherein said data processor processes said audio data to provide processed audio data by encoding said audio data with a compression function. 7. A system according to claim 1, wherein said input processor acquires said vital sign parameters from patient monitoring devices attached to said plurality of different patients. 8. A system according to claim 1, wherein said input processor acquires said associated v... | Siemens Medical Solutions USA, Inc. | A61B005/00
US20080319351 | A1 | 2008-12-25 | Spine tilt monitor with biofeedback | 2008-06-20 | The present invention relates to a system and method for continually monitoring the spine of a user via device containing electromechanical transducer, processor, delay, transmitter, and a feedback component. Through the device and method of employing, the user will be able to learn and sustain a proper posture, thereby avoiding future back pain. | 1. A spine monitoring device for monitoring the lumbar region and cervical region of the spine, comprising an electromechanical transducer;a signal processor;a time switch;a transmitter; anda feedback means. 2. The spine monitoring device of claim 1, wherein said electromechanical transducer is a sensor selected from the group consisting of accelerometer, barometer, or a speedometer. 3. The spine monitoring device of claim 1, wherein said electromechanical transducer measures pressure, force, acceleration, or velocity. 4. The spine monitoring device of claim 1, further comprising a communication means. 5. The spine monitoring device in claim 1, further comprising a memory storage. 6. The spine monitoring device in claim 1, wherein said delay is based on time, current flow, or energy capacitance. 7. The spine monitoring device in claim 1, wherein said feedback means is an audible device, a sensory device, or a visual device. 8. The spine monitoring device in claim 7, wherein said audible device is an alarm or buzzer. 9. The spine monitoring device in claim 7, wherein said sensory device is an electrostimulator. 10. The spine monitoring device in claim 7, wherein said visual device is a monitor. 11. The spine monitoring device in claim 1, wherein said device is mounted on a base that is body-mountable. | The Hong Kong Polytechnic University | A61B005/103
US20080275360 | A1 | 2008-11-06 | Arrangement and Method for Assessing the Motility of a Generally Tubular Anatomical Organ | 2005-05-04 | The invention relates to an arrangement for assessing the motility of a generally tubular anatomical organ (1), said arrangement comprising a longitudinally extending catheter (6) being adapted for introducing into said organ (1). The invention furthermore comprises an electrode arrangement (7) arranged along the longitudinal extension of said catheter (6), a detection unit (8) for measuring the electric potential (V(d)) at least partly along the length of said organ (1) and an evaluation unit (13) for detecting a generally step-like change (19) in the measured electric potential and determining the distance between a position along said catheter (6) corresponding to said step-like change (19) and a predetermined fixed point (16) to which the catheter (6) can be attached. The invention also relates to a method for such motility assessment. | 1. Arrangement for assessing the motility of a generally tubular anatomical organ, comprising: a longitudinally extending catheter being adapted for introducing into said organ and being adapted to be attached to a predetermined fixed point when assessing motilityan electrode arrangement arranged along a longitudinal extension of said catheter,a detection unit for measuring the electric potential (V(d)) at least partly along a length of said organ, andan evaluation unit for detecting a generally step-like change in the measured electric potential and determining the distance between a position along said catheter corresponding to said step-like change and the predetermined fixed. 2. The arrangement according to claim 1, wherein the detection unit is to measure said electric potential (V(d)) in organs having different electrical properties due to different mucosal morphology in said organs. 3. The arrangement according claim 1 wherein said electrode arrangement is arranged to straddle a transition zone, thereby measuring the electric potential on both sides of said step-like change to detect a position of the transition zone in relation to said fixed catheter. 4. The arrangement according to claim 1, wherein said electrode arrangement comprises a plurality of electrodes mounted longitudinally spaced apart along said catheter. 5. The arrangement according to claim 4, wherein said electrodes include a plurality of separated lumens extending along the interior of said catheter, each of said lumens having an opening facing the inside of said organ. 6. The arrangement according to claim 4, wherein said electrodes include a plurality of electrically conductive contacts arranged in a spaced-apart configuration along said catheter. 7. The arrangement according to a reference electrode is connected to said detection unit. 8. The arrangement according to claim 7, wherein said reference electrode is adapted for arranging subcutaneously in, or in contact with, an individual for which an assessment of said motility is to be carried out. 9. The arrangement according to claim 7, one of said electrodes also constitutes a reference electrode. 10. The arrangement according to claim 1, wherein the evaluation unit is adapted for determining the degree of contraction of the esophagus of a human or animal. 11. The arrangement according to claim 1, wherein said catheter also comprises means for side-hole manometry in said organ, for assessing circumferential muscular contractions of said organ. 12. The arrangement according to claim 1, wherein said catheter also comprises sensor means of at least one of ph-metry or impedance measurements along said catheter. 13. The arrangement according to claim 1, wherein said catheter is provided with a sensor for indicating the geometrical configuration of said catheter, and said sensor comprises a measuring unit for cooperating with said catheter and being adapted for assessing the degree of contraction of an anatomical... | - | A61B005/0488
US20080294021 | A1 | 2008-11-27 | System and Method for the Management or Control of Cardiovascular Related Diseases, Such as Hypertension | 2006-09-26 | A system and a method for the management or control of cardiovascular related diseases, in particular of hypertension, basically comprises a blood pressure measuring device and means for determining an activity status of the autonomic nervous system. In particular, this status is determined on the basis of heart rate variability or on the basis of a composition ratio of granulocytes and lymphocytes in leukocytes of the individual. | 1. A system for the management or control of cardiovascular related diseases, in particular of hypertension of an individual, comprising a blood pressure measuring device for measuring the blood pressure of the individualmeans for determining an autonomic nervous system activity status of said individual. 2. A system according to claim 1, the system further comprising a heart rate variability determination means for determination of the heart rate variability (HRV) of said individual. 3. A system according to claim 1, the system further comprising means for determining or entering a composition ratio of granulocytes and lymphocytes in leukocytes of the individual. 4. A system according to claim 2, further comprising a calculating means for determining an autonomic nervous system activity status based on said heart rate variability (HRV). 5. A system according to claim 1, wherein said system comprises a display for display of said autonomic nervous system activity status. 6. A system according to claim 1, wherein said blood pressure measuring device and said heart rate variability (HRV) determination means are integrated within the same physical device. 7. A system according to claim 2, wherein said heart rate variability determination means is adapted to determine heart rate variability (HRV) on the basis of blood pressure pulse measurements. 8. A system according to claim 2, wherein the heart rate variability determination means include an ECG measuring means and wherein said heart rate variability means is adapted to determine heart rate variability (HRV) on the basis of electrocardiographic information of said individual. 9. A system according to claim 1, wherein said system comprises a user input interface adapted for entry of information relating to at least one of risk factors of said individual to suffer from diseases, in particular from cardiovascular related diseases and/or driving factors for such disease. 10. A system according to claim 9, wherein said user input interface is adapted for entry of information selected from the group of sex, age, smoking, total cholesterol level, family history, target organ damage, diabetes, associated clinical conditions and renin concentration or other biochemical information. 11. A system according to claim 1, wherein said blood pressure measuring device is designed for measurement of blood pressure according to a predetermined or predeterminable clinically validated measurement schedule. 12. A system according to claim 1, further comprising a decision making means coupled or adapted to be coupled to said blood pressure measuring device and to said autonomic nervous system activity status determination means and wherein said decision making means comprises data processing means for determining at least one of a risk level for the patient to suffer from a disease and/or driving factors for such disease and/or appropriate measures for treatment of the individual in view of said disease on th... | CONGENER WELLNESS CORP. | A61B005/021 A61B005/02
US20080280276 | A1 | 2008-11-13 | Virtual reality tools and techniques for measuring cognitive ability and cognitive impairment | 2007-07-30 | Techniques and tools for measuring cognitive ability and/or detecting cognitive impairment or decline. For example, techniques and tools are described that can be used to diagnose or test susceptibility to cognitive impairments in children or in elderly people (such as cognitive impairments associated with Alzheimer's Disease). Techniques and tools are described that can be used to evaluate treatment effects and/or measure cognitive decline over time. | 1. A method comprising: receiving input from a user as the user interacts with software presenting a virtual reality environment that includes a first-person, three-dimensional graphical rendering of the virtual reality environment;measuring performance of the user in the virtual reality environment based at least in part upon the received input; andusing the measured performance of the user in the virtual reality environment in analysis of cognitive status;wherein the measured performance includes one or more of distance traversed in the virtual reality environment, time elapsed before reaching a target in the virtual reality environment, percentage of successful trials, time spent in a target area of the virtual reality environment, and velocity of movement in the virtual reality environment. 2. The method of claim 1, wherein the virtual reality environment is organized in a coordinate space, and wherein the measuring comprises tracking position of the user over time in the coordinate space. 3. The method of claim 2, wherein a file includes results of the tracking as time-stamped coordinate locations. 4. The method of claim 1, wherein the virtual reality environment includes multiple areas, and wherein the measured performance further includes pattern of movement between the multiple areas. 5. The method of claim 1, wherein the virtual reality environment includes multiple areas, and wherein the measured performance further includes pattern of time spent in respective areas of the multiple areas. 6. The method of claim 1, wherein the distance traversed is total distance or cumulative distance. 7. The method of claim 1, further comprising: presenting the user with one or more visual or audible cues as to location of the target in the virtual reality environment, wherein the performance of the user is measured in conjunction with presentation of the cues, and wherein the performance of the user is subsequently measured without presentation of the cues to the user. 8. The method of claim 1, further comprising: measuring neural activity of the user with a magnetic resonance imaging tool as the user interacts with the software presenting the virtual reality environment, wherein the analysis of cognitive status is based at least in part on the measured neural activity. 9. The method of claim 1, further comprising: measuring neural activity of the user with a functional magnetic resonance imaging tool as the user interacts with the software presenting the virtual reality environment, wherein the analysis of cognitive status is based at least in part on the measured neural activity. 10. The method of claim 1, wherein a personal computer system presents the virtual reality environment. 11. The method of claim 10, wherein a server computer system provides the virtual reality environment to the personal computer system over a network connection. 12. The method of claim 1, wherein the virtual reality environment is graphically rendered to the us... | Oregon Health & Science University and Oregon Research Institute | A61B005/00
US07435228 | B2 | 2008-10-14 | High fidelity hearing restoration | 2003-07-18 | Method for accurately measuring hearing loss includes the steps of selecting a series of audio tones within the normal range of hearing (502) and then measuring a relative sensitivity of a test subject with respect to the ability to hear each of the audio tones, exclusive of the effects of tinnitus. (504, 506, 508, 510, 512) The relative sensitivity of the test subject to hear the tones can be measured by determining (510) for each tone an intensity necessary for the test subject to hear the tones at a subjectively equal loudness level which is selected to exceed a perceived level of noise attributable to tinnitus for the test subject. | 1. A method for accurately measuring hearing loss, comprising the steps of: selecting a series of audio tones within the normal range of hearing;measuring a relative sensitivity of a test subject with respect to the ability to hear each of said audio tones, exclusive of the effects of tinnitus, said measuring step including selecting a plurality of audio tones, and determining for each said audio tone an intensity necessary for a test subject to hear said audio tone at a subjectively equal loudness level; anddetermining a difference between said intensity measured for each of said audio tones and an intensity predicted by a standard loudness contour for each of said audio tones;wherein said subjectively equal loudness level exceeds a noise level attributable to said tinnitus at a frequency of each said audio tone. 2. A method for accurately measuring hearing loss, comprising the steps of: selecting a series of audio tones within the normal range of hearing;measuring a relative sensitivity of a test subject with respect to the ability to hear each of said audio tones, exclusive of the effects of tinnitus said measuring step including determining for each said audio tone an intensity necessary for said test subject to hear said audio tones at a subjectively equal loudness level which exceeds a noise level attributable to said tinnitus at a frequency of each said audio tone; anddetermining a difference between said intensity measured for each of said audio tones and an intensity predicted by a standard loudness contour for each of said audio tones. 3. The method according to claim 2 further comprising the step of selecting said standard loudness contour to be at least one of a Fletcher-Munson Loudness Contour and a functional equivalent of a Fletcher-Munson Loudness Contour. 4. The method according to claim 1 further comprising the step of measuring a noise level attributable to tinnitus. 5. The method according to claim 1 further comprising the step of configuring at least one gain setting of a hearing aid to compensate for said hearing loss determined in said measuring step. 6. A method for setting a frequency dependent audio gain of a hearing aid device for a person suffering from tinnitus, comprising the steps of: measuring a test subject's loss of hearing attributable exclusively to dispersion in the hearing channel; andsetting for each of a plurality of frequency bands of said hearing aid device an audio gain level to compensate exclusively for said dispersion loss;wherein said measuring step comprises selecting a plurality of audio tones, determining for each said audio tone an intensity necessary for a test subject to hear said audio tone at a subjectively equal loudness level which exceeds a noise level attributable to tinnitus at a frequency of each said audio tone, and determining a difference between said intensity measured for each of said audio tones and an intensity predicted by a standard loudness contour for each of said audio ... | Harris Corporation | A61B005/00 A61B005/12 H04R029/00 H04R025/00 H04R003/02 A61N001/00
US07464595 | B2 | 2008-12-16 | Portable audiometer enclosed within a patient response mechanism housing | 2008-02-08 | A portable audiometer for performing hearing testing enclosed within a patient response mechanism housing is disclosed. The audiometer may include a printed circuit board containing circuitry, plugs or other connectors for wired or wireless headphones and bone conductor, a battery charging unit, and one or more buttons for indicating response to pure tone and other stimuli, including, but not limited to noise, speech, and visual and audio instructions and graphical representations of objects or concepts used in a comprehensive hearing test. The instrument can be connected via a wired or wireless interface to a variety of controlling devices, including PC's, Personal Digital Assistants (PDA's), cellular phones, smartphones, including hybrid phone and PDA, and other devices such as media players, gaming systems, personal audio players, and other devices that can operatively communicate with the portable device. | 1. A portable audiometer device, comprising: a portable enclosure that is adapted for a user to hold in the user's hand, be worn around the user's neck, or be placed on a table-top remote from a microprocessor-based controlling or monitoring device;a patient response mechanism in said portable enclosure for signaling responses to sound stimuli;electronic audiometer circuitry in said portable enclosure that generates or delivers electrical signals representing the sound stimuli and records one or more patient responses;an acoustic transducer that is coupled to the audiometer circuitry for receiving the electrical signals and producing the sound stimuli; anda communications interface in said portable enclosure for communicatively coupling the audiometer circuitry with said microprocessor-based controlling or monitoring device. 2. The audiometer device of claim 1, wherein the patient response mechanism includes at least one of a button, switch, dial, or other actuator. 3. The audiometer device of claim 1, wherein the patient response mechanism also enables navigation through or response to information presented during a testing process. 4. The audiometer device of claim 1, wherein the acoustic transducer also produces other aurally perceptible information. 5. The audiometer device of claim 1, wherein the transducer interface enables the acoustic transducer to be removably connected to the enclosure. 6. The audiometer device of claim 1, wherein the acoustic transducer includes an earphone. 7. The audiometer device of claim 1, wherein the acoustic transducer includes an insertion earphone. 8. The audiometer device of claim 1, wherein the acoustic transducer includes a bone vibrator. 9. The audiometer device of claim 1, further comprising a power supply port for connecting a power supply to the audiometer circuitry. 10. The audiometer device of claim 1, wherein the microprocessor-based device is at least one of a personal computer, a personal digital assistant, a phone, a custom or hybrid device, a media player, a gaming system, a personal audio player, or another device that can operatively communicate with the audiometer circuitry. 11. The audiometer device of claim 1, wherein the microprocessor-based device is adapted to manage, store, or retrieve data about a test subject or a test in such a way that an operation and general appearance of a respective user interface on each of a plurality of such microprocessor-based devices is similar in look, feel, and general operational form, independent of screen size, orientation, or other attributes that may differ from one such microprocessor-based device to another. 12. The audiometer device of claim 1, wherein the microprocessor-based device is adapted to monitor or control a plurality of tests performed simultaneously by a respective plurality of test subjects. 13. The audiometer device of claim 1, wherein the audiometer device comprises a patient-friendly cover over at least a portion of the... | Otovation, LLC | A61B005/00 H04R029/00
US07455406 | B2 | 2008-11-25 | Ophthalmic ultrasonic diagnostic apparatus | 2007-09-28 | To provide an ophthalmic ultrasonic diagnostic apparatus capable of obtaining measurement result with high reliability without lengthy measurement. The apparatus which measures constituent parts of an eye through A-mode includes an ultrasonic probe for A-mode measurement including a transducer emitting an ultrasonic wave into the eye and receiving reflected echoes from tissues of the parts, calculation means for judging whether signals of the echoes are appropriate and calculating measurement values of the parts based on the signals, a memory storing the calculated values, grouping means for setting grouping ranges having predetermined ranges of permissible dispersion with reference to the measurement values and determining that either the grouping range previously set or the grouping range newly set based on the obtained measurement value remains, and central value calculating means counting the numbers of the measurement values falling within the grouping ranges and calculating a central value based on the measurement values. | 1. An ophthalmic ultrasonic diagnostic apparatus which measures constituent parts of an examinee's eye through A-mode, the apparatus comprising: an ultrasonic probe for A-mode measurement including a transducer which emits an ultrasonic wave into the examinee's eye, and receives echoes reflected respectively from tissues of the constituent parts;calculation means for judging whether signals of the received reflected echoes are appropriate by analyzing the reflected echo signals and calculating measurement values of the constituent parts based on the reflected echo signals which are judged as appropriate;a memory which stores the calculated measurement values;grouping means for setting grouping ranges each having a predetermined permissible dispersion range with reference to the measurement value, and determining that either one of; the grouping range which is previously set; andthe grouping range which is newly set based on the obtained measurement value; andcentral value calculating means for counting the number of the measurement values which fall within the respective grouping ranges, and calculating a central value based on the measurement values which fall within the grouping range such that the number of the measurement values falling within the grouping range reaches a predetermined number. 2. The ophthalmic ultrasonic diagnostic apparatus according to claim 1, wherein the grouping means sets the grouping ranges with reference to the measurement value which is obtained first. 3. The ophthalmic ultrasonic diagnostic apparatus according to claim 1, wherein the grouping means sets the grouping ranges in steps of a predetermined width in a range including the lowest and highest measurement values stored in the memory. 4. The ophthalmic ultrasonic diagnostic apparatus according to claim 1, wherein the grouping means sets the grouping ranges with reference to the measurement values at every obtainment of the measurement values. 5. An ophthalmic ultrasonic diagnostic apparatus which measures constituent parts of an examinee's eye through A-mode, the apparatus comprising: an ultrasonic probe for A-mode measurement including a transducer which emits an ultrasonic wave into the examinee's eye, and receives echoes reflected respectively from tissues of the constituent parts;a calculation unit which judges whether signals of the received reflected echoes are appropriate by analyzing the reflected echo signals and calculates measurement values of the constituent parts based on the reflected echo signals which are judged as appropriate;a memory which stores the calculated measurement values;a grouping unit which sets grouping ranges each having a predetermined permissible dispersion range with reference to the measurement value, and determines that either one of; the grouping range which is previously set; andthe grouping range which is newly set based on the obtained measurement value; anda central value calculating unit which counts the number of ... | Nidek Co., Ltd. | A61B003/10 A61B005/05
US07455646 | B2 | 2008-11-25 | Polymer coated guide wire | 2007-06-26 | A guidewire or section thereof, that has a core member or the like with a plurality of contiguous tapered segments having taper angles that are configured to produce a linear change in stiffness over a longitudinal portion of the device. The device may also have a core section with a continuously changing taper angle to produce a curvilinear profile that is configured to produce a linear change in stiffness of the core over a longitudinal portion of the device. An embodiment has a plurality of radiopaque elements that may be intermittent, continuous or in the form of a helical ribbon for scaled measurement of intracorporeal structure under flouroscopic imaging. Another embodiment has at least one layer of polymer over the distal end of the device. | 1. A guidewire comprising: an elongate core member having a proximal section and a distal section with a longitudinal portion which tapers distally to a reduced transverse dimension; wherein the taper creates a substantially linear change in stiffness of the core member over a length thereof; anda tubular polymer member disposed about the distal section of the elongate core member having a plurality of longitudinal segments with at least one of the longitudinal segments being radiopaque and at least one of the segments being radiolucent. 2. The guidewire of claim 1 wherein the longitudinal portion has a curvilinear taper. 3. The guidewire of claim 1 wherein the longitudinal portion has a length of about 10 to about 25 cm. 4. The guidewire of claim 1 further comprising a polymer layer disposed about the distal section of the elongate core member. 5. The guidewire of claim 4 wherein the polymer layer is selected from a group consisting of polyurethanes, polyamide, copolymers of polyurethane and copolymers of polyamide. 6. A guidewire comprising: an elongate core member having a proximal section and a distal section with a longitudinal portion which tapers distally to a reduced transverse dimension wherein the taper creates a substantially linear change in stiffness of the core member over a length thereof; anda flexible body disposed about the distal section of the core member comprising a polymer layer disposed about a radiopaque layer. 7. The guidewire of claim 6 wherein the flexible body disposed about the distal section of the core member comprises a first polymer layer disposed about the distal section, a radiopaque layer disposed about the first polymer layer and a second polymer layer disposed about the radiopaque layer. 8. The guidewire of claim 6 wherein said radiopaque layer is continuous in an axial direction. 9. The guidewire of claim 8 wherein the radiopaque layer is shorter in an axial direction than the flexible body. 10. The guidewire of claim 6 wherein the radiopaque layer is intermittent in an axial direction. 11. The guidewire of claim 10 wherein the intermittent radiopaque layer is comprised of radiopaque bands spaced apart a predetermined axial distance. 12. The guidewire of claim 11 wherein the radiopaque bands are spaced from about 0.2 to about 2 cm apart in an axial direction. 13. The guidewire of claim 11 wherein the radiopaque bands are spaced about 0.8 to about 1.2 cm apart in an axial direction. 14. The guidewire of claim 11 wherein the radiopaque bands are about 0.5 to about 5 mm in width. 15. The guidewire of claim 11 wherein the bands are about 1 to 2 mm in width. 16. The guidewire of claim 6 wherein the radiopaque layer is comprised of a radiopaque helical ribbon coil in an axial direction. 17. The guidewire of claim 16 wherein the radiopaque helical ribbon coil material is about 0.0005 to about 0.0040 inch in thickness. 18. The guidewire of claim 16 wherein the radiopaque helical ribbon coil materia... | Advanced Cardiovascular Systems, Inc. | A61B005/00 A61M025/00
US07766862 | B2 | 2010-08-03 | Baseline acquisition for infection monitoring | 2008-01-31 | A method includes obtaining a baseline value for an indicator of infection in proximity to an implantable medical device and monitoring the indicator of infection in proximity to the device. The method further includes determining whether the monitored indicator is indicative of infection in proximity to the device. The determination as to whether the monitored indicator is indicative of infection includes comparing a value associated with the monitored indicator to the baseline value. | 1. A method performed by an implantable medical device, comprising: obtaining via a sensor an initial baseline value for an indicator of infection in proximity to the implantable medical device; monitoring the indicator of infection in proximity to the device via the sensor; and determining whether the monitored indicator is indicative of infection, wherein determining whether the monitored indicator is indicative of infection comprises comparing a value associated with the monitored indicator to the initial baseline value;receiving instructions from an external device to obtain an additional baseline value for an indicator of infection in proximity to the implantable medical device; andobtaining the additional baseline value via the sensor operably coupled to the device following receipt of the instructions; monitoring the indicator of infection in proximity to the device via the sensor; and determining whether the monitored indicator is indicative of infection in proximity to the device, wherein determining whether the monitored indicator is indicative of infection comprises comparing a value associated with the monitored indicator to the additional baseline value. 2. The method of claim 1, wherein the instructions received from the external device instruct the implantable device to obtain the additional baseline value at a predetermined time following implantation of the device. 3. The method of claim 1, wherein the indicator of infection is temperature. 4. The method of claim 1, further comprising providing an alert if the monitored indicator is determined to be indicative of infection. 5. A non-transitory computer readable medium containing instructions that when implemented cause an implantable medical device to: obtain via a sensor an initial baseline value for an indicator of infection in proximity to the implantable medical device; monitor the indicator of infection in proximity to the device via the sensor; and determine whether the monitored indicator is indicative of infection, wherein the determination comprises comparing a value associated with the monitored indicator to the initial baseline value;receive instructions from an external device to obtain an additional baseline value for an indicator of infection in proximity to the implantable medical device; andobtain the additional baseline value via the sensor for an indicator of infection in proximity to the device; monitor the indicator of infection in proximity to the device; and determine whether the monitored indicator is indicative of infection in proximity to the device, wherein the determination comprises comparing a value associated with the monitored indicator to the additional baseline value. 6. An implantable medical device comprising: a sensor capable of detecting an indicator of infection;electronics operably coupled to the sensor and configured to monitor an infection in proximity to the device; anda computer readable medium according to claim 5 readable and exec... | Medtronic, Inc. | A61M031/00
US07470231 | B2 | 2008-12-30 | Fatigue degree measurement device, fatigue detection device and computer program to be used therein | 2004-10-22 | A fatigue degree measurement device includes a living body signal peak value detecting step to detect the peak values of respective cycles of the original waveform of the living body signal data; a power value calculating step to calculate the difference between a peak value on the upper limit side and a peak value on the lower limit side for every prescribed time period from respective peak values obtained by the living body signal peak value detecting step to set the difference as a power value; and a power value inclination calculating step to determine the inclination of the power value, to calculate an integral value by absolute value treatment of the time series signals of the inclination of the power values to determine the integral value as the degree of fatigue. | 1. A fatigue degree measurement device, comprising: a living body signal peak value detecting means for detecting the peak value in each cycle of an original waveform of the living body signal data collected by a living body signal measurement device;a power value calculating means for calculating the difference between a peak value on the upper limit side and a peak value on the lower limit side for every prescribed time period from each peak value obtained by said living body signal peak value detecting means and for setting the difference as the power value;a power value inclination calculating means for determining an inclination of said power values to the time base during the prescribed time period by slide calculating the prescribed times at a prescribed overlap rate for the prescribed time period;a fatigue degree calculating means for calculating an integral value by absolute value treatment of time base signal of power value inclination obtained from the slide calculation by said power value inclination calculating means to determine the obtained integral value as the degree of fatigue;a maximum Lyapunov index calculating means for calculating the maximum Lyapunov index by chaos analyzing said living body signal data;a maximum Lyapunov index peak value detecting means for detecting the peak value in each cycle of a time series change waveform of the calculated maximum Lyapunov index;a maximum Lyapunov index inclination calculating means for determining an inclination of each peak value of the maximum Lyapunov indexes obtained by the maximum Lyapunov index peak value detecting means to the time base during the prescribed time period by slide calculating the prescribed times at a prescribed overlap rate for the prescribed time period, in addition to said inclination of the power value; anda comparing and determining means for determining as the generating point of a fatigue signal when the inclination of the power value obtained by slide calculating using said power value inclination calculating means and the maximum Lyapunov index obtained by slide calculating using the maximum Lyapunov index inclination calculating means stably showing the phase difference of substantially 180° among time series signals,wherein said comparing and determining means includes a fatigue state determining means for determining the state of fatigue based on the inclinations of power value and the maximum Lyapunov index appearing in time series, andwherein said fatigue state determining means includes a means to perform frequency analysis of the change in the inclinations of power value and the maximum Lyapunov index appeared in time series, and determines a central fatigue predominant state when the power spectrum of the inclination of the maximum Lyapunov index is large, and a peripheral fatigue predominant state when the power spectrum of the inclination of the power value is large. 2. The fatigue degree measurement device according to claim 1, wherein said... | Delta Tooling Co., Ltd. | A61B005/00
US07465272 | B2 | 2008-12-16 | Vehicle with on-board dieters' self-acquiring overweight preventive system and method | 2004-08-20 | A vehicle has an onboard dieters self acquiring overweight prevention system which they can utilize a weighing device connected to the seat of the driver or passenger and a further weighing device upon which the feet may be placed. The vehicle computer can be programmed with a weight-trend analyzer routine predicting overweight or obesity condition and capable of automatically warning the person and a healthcare worker of the condition. | 1. A motor vehicle comprising: an on-board vehicle microcomputer;a memory and a display connected to said microcomputer;wherein said vehicle microcomputer forecasts an overweight or obesity condition in at least one person selected from at least one of said vehicle driver and at least one passenger in said vehicle;a first on-board weighing apparatus for a driver;a second on-board weighing apparatus for a passenger, wherein the passenger is not the driver;wherein at least one of said first and second weighing apparatuses weighs said driver and/or said at least one passenger occupying said vehicle and provides a weight of said driver and/or said at least one passenger to said on-board vehicle microcomputer;said on-board vehicle microcomputer being programmed with a weight trend analyzer routine predicting a developing overweight or obesity condition of said driver and/or said at least one passenger; andsaid on-board vehicle microcomputer automatically warning at least one of said driver and/or said at least one passenger and/or a health-care professional of a developing overweight or obesity condition. 2. The motor vehicle defined in claim 1 wherein said on-board weighing apparatus has at least two weighing platforms including a weighing platform for a body of said person and an additional weighing platform for the feet of the person, said apparatus comprising determining a true weight from said platforms by processing acquired data over an interval between said person engaging said platforms and a point at which said person engages a seat belt whereby said true weights reflect a measurement during which hands of the person are not resting upon a surface. 3. The motor vehicle defined in claim 1 wherein said microcomputer is further programmed with: a weight trial analyzer routine evaluating the conditions of a dieter's weight loss trial plan when overweight of said dieter is forecast or exists;a trial period assignment routine evaluating a time period of said dieter's weight loss trial plan;an overweight warning routine responsive to a negative result of said dieter's weight loss trial plan and sending overweight or obese warnings at a driver's or a passenger's discretion;a driver weighing routine responsive to weight and position of said driver in said vehicle;a passenger weighing routine responsive to weight and position of said passenger in said vehicle;a transposition method of an individual weighing routine to provide a convenient and correct weighing of said individual in said vehicle;a system initialization routine responsive to accessing of said vehicle by a vehicle operator and powering of said weighing apparatus;a system activation routine for determining whether the vehicle is in a state for employing said display by said driver;an introduction routine establishing interconnections between said dieter and self-acquiring overweight preventive system in said vehicle;a weight identifier routine for identifying a reason for any changes in ... | - | A61B005/00
US20080269644 | A1 | 2008-10-30 | Precision Athletic Aptitude and Performance Data Analysis System | 2008-04-25 | Systems and methods provide collection and analysis of athletic and other human performance and related environmental data. The systems include performance and environmental measurement hardware and data collection and analysis software. Data can be collected in various ways including one of several standardized precision athletic tests. Certification of performance data is provided if environmental data and performance measurement hardware status satisfy preset limits. Performance data can be normalized based on non-standard environmental conditions. The systems and methods establish baseline indications of athletic ability, predict athletic potential for specific sports and team positions, compares individuals to others and to norms, tracks athletic progress, specifies tailored training programs for athletic improvement for specific sports and team positions, provides visual data and analysis, and aids training, physical therapy, and rehabilitation. | 1. A system for collecting human performance data, comprising: a first performance sensor;a locator device adapted for determining an actual location of the first performance sensor and for providing a location signal;preset location parameters;a data processor adapted to compare the location and the preset parameters and to output at least one of a plurality of signals based on the comparisons. 2. The system of claim 1, further comprising: a first signal provided by the data processor upon the actual location satisfying the preset location parameters; andan indicator adapted to receive and respond to the first signals. 3. The system of claim 1, wherein the preset location includes a preset displacement and the actual location includes a relative displacement between the first performance sensor and a reference point. 4. The system of claim 3, wherein the plurality of signals includes: a first signal provided by the data processor upon the relative displacement satisfying the preset displacement parameters;a second signal provided by the data processor upon the relative displacement being less than the preset displacement parameters;a third signal provided by the data processor upon the relative displacement being greater than the preset displacement parameters; andan indicator, the indicator adapted to display a different indication for each of the first, second, and third signals. 5. The system of claim 1, further comprising: an environmental sensor adapted to provide an environmental measurement;preset environmental parameters associated with the environmental measurement; anda certification signal provided by the data processor upon the environmental measurement satisfying the preset environmental parameters and the actual location satisfying the present location parameters. 6. The system of claim 1, further comprising: a second performance sensor, andwherein the preset and actual location include a displacement between the first and second performance sensors. 7. The system of claim 1, wherein the locator device includes a distance measuring laser. 8. The system of claim 1, wherein the locator device includes a GPS receiver. 9. The system of claim 1, further comprising a video camera for capturing image data, and wherein the data processor is further adapted to time associate the image data and human performance data. 10. The system of claim 1, further comprising a geographically remotely located database for storing the human performance data. 11. The system of claim 1, further comprising: a first data structure including data fields for storing human performance data associated with a first person and the first performance sensor; anda second data structure including data fields for storing human performance data associated with at least a second person and spanning a first period of development; andwherein the data processor is further adapted to determine a predicted human performance of the first person over a second period of... | - | A61B005/22
US07682355 | B2 | 2010-03-23 | Refined infection monitoring | 2007-04-19 | A method for monitoring infection in proximity to an implanted medical device includes monitoring an indicator of infection in proximity to the device and determining whether a value associated the indicator of infection crosses a first threshold indicative of infection. The method further includes detecting an event associated with the device. The event is capable of affecting the indicator of infection. The method also includes determining whether the value associated with the indicator of infection crosses a second threshold indicative of infection if the event is detected. The method may further include issuing an alert if (i) the first threshold is crossed and the event is not detected, or (ii) the second threshold is crossed and the event is detected. | 1. A method for monitoring an infection in proximity to an implanted medical device, comprising: monitoring, via the implantable medical device, an indicator of infection in proximity to the device;determining whether a value associated with the indicator of infection crosses a first threshold indicative of infection;determining if an event associated with the device that is capable of affecting the indicator of infection has occurred, wherein determining whether the event associated with the device has occurred, comprises detecting a recharging event;determining whether the value associated with the indicator of infection crosses a second threshold indicative of infection if the event has occurred; andissuing an alert if: (i) the first threshold is crossed and the event has not occurred, or(ii) the second threshold is crossed and the event has occurred. 2. The method of claim 1 wherein monitoring the indicator of infection comprises monitoring temperature. 3. The method of claim 2, wherein the second threshold includes a higher temperature parameter than the first threshold. 4. A computer readable medium containing instructions that when implemented cause an implantable medical device to (i) monitor an indicator of infection in proximity to the device; (ii) determine whether a value associated with the indicator of infection crosses a first threshold indicative of infection; (iii) determine if an event associated with the device has occurred, the event capable of affecting the indicator of infection, wherein the event is a recharging event; (iv) determine whether the value associated with the indicator of infection crossed a second threshold indicative of infection if the event has occurred; and (v) provide an alert either if the first threshold is crossed and the event has not occurred or if the second threshold is crossed and the event has occurred. 5. An implantable medical device comprising: electronics capable of executing computer readable instructions and capable of detecting an event associated with the device, the event capable of affecting an indicator of infection;a sensor operably coupled to the electronics, the sensor capable of detecting the indicator of infection;an alarm capable of providing a sensory cue; anda computer readable medium containing instructions that when implemented by the electronics cause the device to (i) determine whether a value associated with the indicator of infection crosses a first threshold indicative of infection, wherein the event is a recharging event; (ii) determine whether the value associated with the indicator of infection crosses a second threshold indicative of infection if the event is detected; and (iii) activate the alarm either if the first threshold is crossed and the event is not detected or if the second threshold is crossed and the event is detected. 6. The device of claim 5, wherein the sensor is a temperature sensor. | Medtronic, Inc. | A61K009/22
US07467010 | B2 | 2008-12-16 | Stress-at-work evaluating device and method | 2004-09-16 | A stress-at-work evaluating device evaluates a stress of a subject at work by measuring activities of one of the right and the left masseter muscles, the work including exercise of arms or feet of the subject. The device evaluates the stress from a normalized intensity of the myoelectric potential one of the right and the left masseter muscles. The normalized intensity is calculated by dividing the intensity of the measured myoelectric potential of the masseter muscles by the level of the external force acting on the head of the subject at work. | 1. A stress-at-work evaluating method for evaluating stress of a subject at work based on measuring the activities of one of the right and the left masseter muscles, the work including exercise of the subject's arms or feet, wherein the stress-at-work evaluating method comprises: a myoelectric potential measurement step of sensing and amplifying a myoelectric potential of one of the right and the left masseter muscles, the potential being generated through the activities of one of the right and the left masseter muscles at work;a processing step of processing time-series data of the measured myoelectric potential of the one of the right and the left masseter muscles and calculating the intensity information on myoelectric potential of the one of the right and the left masseter muscles;a judging step of judging the level of an external force acting on the subject's head, the external force measured simultaneously with the myoelectric potential; andan evaluating step of evaluating the level of stress of the subject at work, wherein the level of stress is obtained by dividing the calculated intensity of the myoelectric potential by the judged level of the external force. 2. The stress-at-work evaluating method according to claim 1, wherein the judging step of judging the level of the external force uses measured external force data of a measurement time region, which is identical to the measurement time region used in the myoelectric potential measurement step. 3. A stress-at-work evaluating device for evaluating stress of a subject at work based on measuring the activities of one of the right and the left masseter muscles, the work including exercise of the subject's arms or feet, wherein the stress-at-work evaluating device comprises: a sensor unit for sensing a myoelectric potential of one of the right and the left masseter muscles, the myoelectric potential being generated through the activities of one of the right and the left masseter muscles at work;an amplifier for amplifying the myoelectric potential sensed by the sensor unit;a myoelectric-potential data processing part for processing time-series data of the amplified myoelectric potential of the one of the right and the left masseter muscles, to thereby calculate the intensity of the myoelectric potential of the one of the right and the left masseter muscles;an external-force level judging part for judging the level of an external force acting on a portion of the head of the subject at work, the external force measured simultaneously with the myoelectric potential; andan evaluating part for evaluating the level of stress of the subject at work, wherein the level of stress is obtained by dividing the calculated intensity of the myoelectric potential by the judged level of external force. 4. The stress-at-work evaluating device according to claim 3, wherein the external-force level judging part judges the level of the external force acting on the portion of the subject's head based on the... | Yokohama Rubber Co., Ltd. | A61B005/04
US07452334 | B2 | 2008-11-18 | Antenna stent device for wireless, intraluminal monitoring | 2004-09-13 | Disclosed herein is a stent device useful for maintaining the patency of a lumen while monitoring an intraluminal characteristic. The device includes a structure having a set of extendible bands that are capable of plastic deformation to form a scaffolding having an inductance, and further includes a capacitance coupled to the set of extendible bands and responsive to the intraluminal characteristic. The capacitance and the inductance form a tank circuit after the plastic deformation of the set of extendible bands to enable wireless transmission of an indication of the intraluminal characteristic. | 1. A stent device, comprising: a structure comprising a set of extendible bands that are capable of plastic deformation to form a scaffolding having an inductance, wherein the set of extendible bands has a non-helical configuration prior to the plastic deformation and wherein a portion of the non-helical configuration forms the scaffolding upon plastic deformation; anda capacitance electrically coupled to the set of extendible bands such that the capacitance and the inductance form a tank circuit after the plastic deformation of the set of extendible bands. 2. The stent device of claim 1, wherein the structure comprises a longitudinal beam coupled to the set of extendible bands. 3. The stent device of claim 2, wherein the longitudinal beam comprises a side beam having a reduced-width link between adjacent bands of the set of extendible bands such that the reduced-width link breaks a connection between the adjacent bands upon plastic deformation of the structure. 4. The stent device of claim 2, wherein the longitudinal beam and the set of extendible bands are directly connected by a single link. 5. The stent device of claim 2, wherein the structure comprises a further set of extendible bands such that the scaffolding comprises a further inductance, the stent device further comprising a further capacitance coupled to the further set of extendible bands such that the further capacitance and the further inductance form a further tank circuit after plastic deformation of the further set of extendible bands. 6. The stent device of claim 5, wherein the longitudinal beam is a common electrical node for the first-named inductance and the further named inductance and the first-named capacitance and the further named capacitance. 7. The stent device of claim 6, wherein the first-named capacitance and the further capacitance comprise first and second capacitive sensors, respectively, and wherein the first and second capacitive sensors are disposed at opposite ends of the longitudinal beam. 8. The stent device of claim 7, wherein the longitudinal beam has a length such that the opposite ends of the longitudinal beam are positioned beyond where an expandable deployment device is disposed during the plastic deformation of the structure. 9. The stent device of claim 8, wherein the structure comprises first and second platforms upon which the first and second capacitive sensors are mounted, respectively. 10. The stent device of claim 1, wherein the set of extendible bands comprises a non-conductive support and a conductive trace disposed on the non-conductive support. 11. The stent device of claim 1, wherein the scaffolding comprises a helix portion and a ring portion such that the ring portion provides structural support at an end of the helix portion. 12. The stent device of claim 1, wherein the non-helical configuration is planar such that each extendible band of the set of extendible bands comprises a plurality of involute, planar sections prior to... | The Regents of The University of Michigan | A61B005/02 A61F002/06
US07434325 | B2 | 2008-10-14 | Systems and methods for determining optimal retractor length in minimally invasive procedures | 2004-07-26 | A measurement instrument is provided with indicia extending therealong. The measurement instrument indicia is correlated with an indicator of a reference instrument. The measurement instrument is positionable adjacent the entry location of a portal into a patient when the reference instrument is in the portal. The indicator of the reference instrument is correlated to a target location in the patient, and a location of the indicator along the indicia provides an indication of the optimal length for a retractor to be positioned in the portal. | 1. A system for determining an optimal length for a retractor for use in a minimally invasive procedure with a patient, comprising: a reference instrument extending between a distal end and a proximal end, said reference instrument including a predetermined length between said distal end and an indicator thereon, said distal end being positionable in a portal adjacent to a target location in the patient with said indicator located proximally of an entry location of the portal; anda measurement instrument extending between a distal end and a proximal end, said measurement instrument including indicia extending therealong, wherein when said reference instrument is in the portal with said distal end thereof adjacent the target location and said measurement instrument is positioned along the reference instrument with said distal end of said measurement instrument adjacent the entry location, a location of said indicator of said reference instrument along said indicia of said measurement instrument provides an indication of the optimal length of the retractor for positioning in the portal. 2. The system of claim 1, wherein said indicia includes a numerical scale increasing in value from said proximal end toward said distal end of said measurement instrument. 3. The system of claim 1, wherein said indicia provides a measurement of 0 a portion of said predetermined length of said reference instrument, said portion extending distally from the entry location to said distal end of said reference instrument when said reference instrument is positioned in the portal. 4. The system of claim 1, wherein said indicia is located on an external surface of said measurement instrument. 5. The system of claim 1, wherein said measurement instrument is a first dilator and said reference instrument is a second dilator, said second dilator being positionable about said first dilator when in said portal, said first dilator being removable with said second dilator in the portal and being positionable along the second dilator with said distal end adjacent the entry location of the portal. 6. The system of claim 1, wherein said measurement instrument includes a passage sized for positioning about at least a proximal portion of said reference instrument. 7. The system of claim 6, wherein said measurement instrument includes a window, said window extending along said measurement instrument between said distal and proximal ends of said measurement instrument. 8. The system of claim 7, wherein said indicia extends along said window. 9. The system of claim 1, wherein said reference instrument includes a number of dilators positionable one over the other to sequentially dilate tissue along the portal. 10. The system of claim 9, wherein each of said number of dilators includes a predetermined length between a distal end and an indicator thereof, and said measurement instrument is selected from a set of measurement instruments each with indicia correlated to said predetermi... | Warsaw Orthopedic, Inc. | A61B005/107 A61B001/313 G01B005/18
US08111888 | B2 | 2012-02-07 | Method for automatically selecting a display mode for an image data record of an organ to be examined | 2008-05-20 | At least one example embodiment discloses a method for automatically selecting a display mode for an image data record is provided. In the at least one example embodiment, spatially resolved values of a parameter are determined in a first image data record of the organ to be examined, with the parameter allowing an evaluation of a functionality of the organ to be examined. Thereupon a deviation of the determined values of the parameter from a tolerance range is localized. The result of the localization is saved, and a display mode of a second image data record is selected on the basis of the result of localization, with the second image data record being a 3D image data record. The second image data record is shown in the selected display mode. | 1. A method for automatically selecting a display mode for an image data record of an organ to be examined, comprising: displaying a first image data record of the organ to be examined;site-specific determining values of a parameter in the first image data record of the organ to be examined, the parameter allowing an evaluation of a functionality of the organ to be examined;localizing a deviation of the determined values of the parameter from a tolerance range;saving a result of the localization;selecting a display mode of a second image data record based on the result of the localization of the deviation, the second image data record being a 3D image data record; anddisplaying the second image data record in the selected display mode. 2. The method as claimed in claim 1, wherein the localization of the deviation takes place aided by a subdivision of the organ to be examined into segments. 3. The method as claimed in claim 2, wherein a display mode is assigned to each segment. 4. The method as claimed in claim 1, wherein the display mode prescribes the viewing angle. 5. The method as claimed in claim 1, wherein the result of the localization is marked in a schematic representation of the organ to be examined when the result of the localization is being saved. 6. The method as claimed in claim 1, wherein the first image data record is a 2D image data record or a 3D image data record or a series of 2D or 3D image data records. 7. The method as claimed in claim 1, wherein the heart is the organ to be examined. 8. The method as claimed in claim 1, wherein the second image data record displays the coronary arteries. 9. The method as claimed in claim 1, wherein the parameter for evaluating the functionality of the organ to be examined is a measure of contrast agent enhancement or washout. 10. The method as claimed in claim 1, wherein the parameter for evaluating the functionality of the organ to be examined is a measure of a perfusion value. 11. The method as claimed in claim 1, wherein the parameter for evaluating the functionality of the organ to be examined is a measure of a contractility value. 12. The method as claimed in claim 1, wherein the parameter for evaluating the function of the organ to be examined is a measure of a thickness of a part of the organ. 13. The method as claimed in claim 1, wherein the parameter for evaluating the function of the organ to be examined is a measure of a motility value. 14. A non-transitory computer readable medium including program segments for, when executed on a computer device, causing the computer device to implement the method of claim 1. 15. A method for automatically selecting a display mode for an image data record of an organ to be examined, comprising: displaying a first image data record of the organ to be examined;determining values of a parameter in the first image data record of the organ to be examined, the parameter allowing an evaluation of a functionality of the organ to be exami... | Siemens Aktiengesellschaft | G06K009/00
US08206309 | B2 | 2012-06-26 | Blood pressure monitoring apparatus and method | 2008-04-04 | The blood pressure monitoring apparatus and method are disclosed that can monitor a blood pressure of a subject using an electrocardiogram signal, a pulse wave signal and a body characteristic information of the subject, wherein the electrocardiogram signal and the pulse wave signal of the subject are monitored to remove a noise signal generated from monitoring of the pulse wave signal, allowing monitoring a precise blood pressure of the subject, and calculating the pulse wave analysis information using the monitored pulse wave signal, and using the electrocardiogram signal and the pulse wave signal to calculate a pulse transit time (PPT), and plugging a calculated pulse wave propagation time, pulse wave analysis information and body characteristic information of the subject into a predetermined regression equation to monitor the blood pressure. | 1. A blood pressure monitoring apparatus comprising: an electrocardiogram monitoring unit configured to monitor an electrocardiogram signal of a subject;a photoplethysmography (PPG) monitoring unit configured to emit an optical signal of a predetermined frequency to a subject's body, detect an optical signal that has reflected from or penetrated the subject's body, and deemphasize noise within the detected optical signal to generate a pulse wave signal;a body characteristic information input unit configured to enable input of body characteristic information of the subject;a micro computer configured to calculate pulse wave analysis information using the pulse wave signal, to calculate a pulse wave propagation time using the pulse wave signal and the electrocardiogram signal, and to calculate a blood pressure of the subject using the calculated pulse wave analysis information, the pulse wave propagation time, and the body characteristic information inputted by the body characteristic information input unit, wherein the micro computer is configured to calculate the blood pressure by inputting the pulse wave propagation time, the pulse wave analysis information and the body characteristic information into the following regression equation:Blood⁢⁢pressure⁢⁢(P)=k1×pulse⁢⁢wave⁢⁢propagation⁢⁢time⁢+∑i⁢⁢k2⁢i×pulse⁢⁢wave⁢⁢analysis⁢⁢information⁢+∑j⁢⁢k3⁢j×⁢body⁢⁢characteristic⁢⁢information,where k1, K2i, and K3jdefine coefficient values respectively corresponding to the pulse wave propagation time, the pulse wave analysis information, and the body characteristic information monitored from the regression analysis, which are predetermined values that are experimentally obtained; anda display unit displaying the blood pressure calculated by the micro computer. 2. The apparatus as claimed in claim 1, wherein the electrocardiogram monitoring unit comprises: electrocardiogram monitoring electrodes configured to monitor an electric signal of the body by being brought into contact with the body; andan electrocardiogram monitoring circuit configured to generate an electrocardiogram signal using the electric signal monitored by the electrocardiogram monitoring electrodes and output the generated electrocardiogram signal to the micro computer. 3. The apparatus as claimed in claim 1, wherein the PPG monitoring unit comprises: a PPG sensor configured to generate an optical signal of a predetermined frequency, emit the generated optical signal to the body of the subject, receive an optical signal reflected from the body, and convert the received optical signal to an electric signal; anda PPG monitoring circuit configured to discern a noise signal and a signal of the predetermined frequency from the electric signal converted by the PPG sensor, generate the pulse wave signal using the discerned signal of the predetermined frequency, and input the generated pulse wave signal to the micro computer. 4. The apparatus as claimed in claim 3, wherein the PPG sensor comprises: a... | LG Electronics Inc. | A61B005/021 A61B005/0402 A61B005/00 A61B005/01
US07927285 | B2 | 2011-04-19 | Nasal mask for carbon dioxide gas measurement | 2008-06-30 | A nasal mask is adapted to attached on a face of a subject so as to cover a nose of the subject. A shell has a contact edge adapted to be brought into close contact with the face of the subject, thereby forming a sealed space therein. An attachment portion is disposed in the shell and has an inner space. The attachment portion is configured to detachably engage with a sensor disposed outside the shell and operable to measure concentration of carbon dioxide gas expired from the nostrils. A guide member is disposed inside the shell and forming a passage adapted to introduce expired gas from the nostrils into the inner space. The shell is formed with an inlet adapted to introduce external gas into the sealed space. The expired gas is guarded by the guide member from the external gas introduced through the inlet. The attachment portion is formed with an outlet adapted to discharge the expired gas from the inner space to the sealed space. | 1. A nasal mask, adapted to be attached to a face of a subject so as to cover a nose of the subject, comprising: a shell, having a contact edge adapted to be brought into close contact with the face of the subject, thereby forming a sealed space therein;an attachment portion, disposed in the shell and having an inner space, the attachment portion being configured to detachably engage with a sensor disposed outside the shell and operable to measure concentration of carbon dioxide gas expired from the nostrils; anda guide member, disposed inside the shell and forming a passage adapted to introduce expired gas from nostrils into the inner space,a cup member integrally provided with the attachment portion and adapted to cover the nose and to block the external gas introduced through the inlet wherein:the shell is formed with an inlet adapted to introduce external gas into the sealed space;the expired gas is guarded by the guide member from the external gas introduced through the inlet; andthe attachment portion is formed with an outlet adapted to discharge the expired gas from the inner space to the sealed space. 2. The nasal mask as set forth in claim 1, wherein: the guide member is a nasal tube member adapted to be placed within the nostrils so as to communicate the nostrils and the inner space. 3. The nasal mask as set forth in claim 1, wherein: the guide member is a cup member integrally provided with the attachment portion and adapted to cover the nose; andan inner space of the cup member is communicated with the inner space of the attachment portion. 4. The nasal mask as set forth in claim 1, wherein: the shell is formed with a recessed part adapted to receive the sensor, and the attachment portion is projected into the recessed part. 5. The nasal mask as set forth in claim 4, wherein: the attachment portion is monolithically formed with the shell. 6. The nasal mask as set forth in claim 4, wherein: the attachment portion is detachably fitted with an aperture formed in the recessed part. 7. The nasal mask as set forth in claim 1, wherein: the nasal mask is configured such that the expired gas travels first from the nostrils and then through the inner space and then to the sealed space. | Nihon Kohden Corporation | A61B005/08 A62B018/02
US08036826 | B2 | 2011-10-11 | Sports sensor | 2008-05-16 | A data logger for a monitoring sports which includes an accelerometer, a gyro sensor to sense angular displacement, a GPS unit to sense position and velocity, a magnetometer to sense direction of movement, a heart rate monitor, and a controller programmed to manipulate the data and provide a display of the heart rate, speed, and other sport parameters. The data can be stored or transmitted to a remote computer for use by the coach. The device is useful in football codes, athletics, swimming, snow sports and cycling. | 1. A data acquisition system for use in sporting events which incorporates a. at least one inertial sensor measuring angular acceleration in at least one dimensionb. at least one accelerometer to derive acceleration and velocity data in 3 dimensionsc. a microcontroller with a clock to measure the angular acceleration and accelerometer datad. a power supplye. optional communication means for transmission of angular acceleration and accelerometer data from the microcontroller to a computer devicef. said computer device and/or said micro controller being programmed to use the angular acceleration and accelerometer data to provide accurate output of parameters such as velocity, acceleration, changes of direction and distance traveled. 2. A data acquisition system as claimed in claim 1 which also includes a GPS sensor and velocity is derived from the global position sensor and the accelerometer and angular change data is sampled to obtain movement characteristics of the sport being monitored. 3. A data acquisition system as claimed in claim 2 which also includes a magnetometer which combined with GPS data can provide directional data. 4. A data acquisition system as claimed in claim 1 which also includes a physiological sensor. 5. A data acquisition system as claimed in claim 4 in the physiological sensor is a heart rate monitor. 6. A data acquisition system as claimed in claim 1 for a cyclist which also includes a pedal cadence sensor. 7. A data acquisition system for use in sporting events which incorporates a) at least one accelerometer to derive acceleration and velocity data in 3 dimensionsb) at least one magnetometer sensing the direction of magnetic northc) a Global Positioning Sensor (GPS)d) a microcontroller with a clock to measure the magnetometer, GPS and accelerometer datae) a power supplyf) optional communication means for transmission of magnetometer, GPS and accelerometer data from the microcontroller to a computer deviceg) said computer device and/or said micro controller being programmed to use the magnetometer, GPS and accelerometer data to provide accurate output of parameters such as velocity, acceleration, changes of direction and distance traveled. 8. A data acquisition system as claimed in claim 7 which also includes at least one inertial sensor measuring angular acceleration in at least one dimension. 9. A data acquisition system as claimed in claim 7 which also includes a physiological sensor. 10. A data acquisition system as claimed in claim 9 in the physiological sensor is a heart rate monitor. | MNT Innovations Pty Ltd | G01C021/00
US08043227 | B2 | 2011-10-25 | Non-invasive system and method for measuring skin hydration of a subject | 2006-12-26 | A non-invasive system and method for measuring skin hydration of a subject includes a thermistor used in transient mode to obtain a measurement of the thermal conductivity of the skin of the subject and a processor for determining a skin hydration value from the thermal conductivity measurement. The system for measuring skin hydration further includes a non-invasive system for detecting blood analyte concentration, such as glucose, with a spectroscopic device having e.g., an infrared source which generates infrared beam and detector for detecting transmitted radiation through portion (e.g., finger) of a subject. The system may also include skin hydrator which moisturizes the skin and is connected in a control loop to the system for measuring skin hydration. The system for detecting blood analyte concentration may include a photoacoustic device or a metabolic heat conformation device. | 1. A non-invasive system for measuring hydration of at least one of a skin and exterior mucosa and detecting a blood analyte concentration of a subject comprising: a thermal conductivity measuring apparatus for obtaining a measurement of a thermal conductivity of the at least one of the skin and the exterior mucosa of the subject; anda processor configured to determine a hydration value from the thermal conductivity measurement and to determine the blood analyte concentration, wherein the processor is further configured to postpone determination of the blood analyte concentration until the hydration value is within a predetermined range,further comprising a hydrating device for providing a hydrating agent to the least one of the skin and the exterior mucosa and a hydration controller connected to the hydrating device so the hydration controller maintains the hydration value at a predetermined level within the predetermined range. 2. The non-invasive system according to claim 1, wherein the thermal conductivity measuring apparatus comprises: a thermistor in thermal connection with the at least one of a skin and exterior mucosa of the subject in use: anda current supply for supplying current to the thermistor sufficient to maintain the resistance or temperature of the thermistor at a predetermined level,wherein the thermal conductivity measuring apparatus obtains a measurement of the thermal conductivity in dependence upon the current supplied to the thermistor. 3. The non-invasive system according to claim 1, wherein the system for detecting the analyte blood concentration comprises one or more of: a spectroscopic device;a photoacoustic device; anda device for determining blood glucose concentration based on the metabolic heat conformation method. 4. The non-invasive system according to claim 1, wherein the analyte concentration is glucose concentration. 5. The non-invasive system of claim 1, wherein the processor is further configured to continuously compare the determined hydration value with the predefined range until the determined hydration value is within the predefined range. 6. The non-invasive system of claim 1, wherein the processor is further configured to determine the blood analyte concentration from a radiation detected by a detector, wherein the radiation is emitted by a source and transmitted through a portion of the subject. 7. The non-invasive system of claim 1, wherein the processor is further configured to determine the blood analyte concentration from a pressure wave detected by an acoustic sensor in response to a pulsed radiation emitted by source to cause a temperature rise that generates the pressure wave. 8. A method for measuring hydration of at least one of a skin and exterior mucosa and a blood analyte concentration of a subject, the method comprising the acts of: non-invasively measuring a thermal conductivity of the at least one of the skin and the exterior mucosa;determining a hydration level of the at least o... | Konklijke Philips Electronics N.V. | A61B005/00
US08221721 | B2 | 2012-07-17 | Polycyclo dyes and use thereof | 2008-02-08 | The invention relates to a family of fluorescent compounds of Formula (I). The compounds can be chemically linked to biomolecules, such as proteins, nucleic acids, and therapeutic small molecules. The compounds can be used for imaging in a variety of medical, biological and diagnostic applications, and are particularly useful for the in vivo imaging of regions of interest within a mammal.; | 1. A compound represented by Formula I:or a salt thereof, wherein: X1 and X2 each independently are selected from the group consisting of C(CH2K1)(CH2K2), O, S, and Se, wherein K1 and K2 each independently are selected from the group consisting of H and C1-C20 alkyl; wherein K1 and K2 together can optionally form part of a cyclic ring;R1 and R2 each independently are selected from the group consisting of H, C1-20 alkyl, alkylaryl, and —R9T1, wherein R9 is independently selected from the group consisting of C1-C20 alkyl, aryl, alkylaryl, and —(CH2—O—CH2)yCH2, wherein y is an integer selected from 1 to 100, andT1 is a linker moiety capable of forming a covalent bond with a biomolecule;X3 is selected from the group consisting of H, halogen, -T2R10, -T2R9T1, and —R9T1, whereinT2 is selected from the group consisting of S, O, and NR′,R10 is selected from the group consisting of H, C1-20 alkyl, aryl, arylalkyl;X4 and X5 each independently are selected from the group consisting of (CK1K2)i; NR′, —NR′R9T1, O, and S, wherein i is an integer selected from 1 to 7, andR′ is selected from the group consisting of H and alkyl;W1 and W2 each represent nonmetal atoms as part of an aryl moiety;n1 and n2 each independently are selected from 0 to 4;R3, R4, R5 and R6 each independently are selected from the group consisting of H, halogen, carboxylate, carboxylic acid, carboxylic ester, amine, amide, sulfonamide, hydroxyl, O-alkyl, S-alkyl, alkyl, sulfonic acid, sulfonate, a phosphoryl, SO2NR7-Q-CHR8—[(CH2)m—Y—(CH2)p—(O)k]h(CH2)dT1, and —R9T1, wherein R7 independently is selected from the group consisting of C1-C20 alkyl, aryl, alkylaryl,R8 independently is selected from the group consisting of H, C1-C20 alkyl, aryl, alkylaryl and arylalkyl;R7 and R8, when taken in combination, optionally form a 4-, 5-, 6- or 7-member, optionally substituted, saturated or unsaturated ring,Q independently is selected from the group consisting of a bond, carbonyl, C1-C6 alkyl, and C1-C6 cycloalkyl,Y independently is selected from the group consisting of a bond, —SO2NR1—, —O—, —COO—, and —CONR′—,h is an integer selected from 0 to 70,k is 0 or 1,d is an integer from 0 to 12,m is an integer from 0 to 12, andp is an integer from 0 to 12. 2. The compound of claim 1, wherein the molecule has an absorption and emission wavelength in the range from about 500 nm to about 1100 nm. 3. The compound of claim 2, wherein the molecule has an absorption and emission wavelength in the range from about 600 nm to about 900 nm. 4. The compound of claim 1, wherein X3 is selected from the group consisting of: O—(CH2)j-T1, S—(CH2)j-T1, O-Ph-(CH2)j-T1, S-Ph-(CH2)j-T1, wherein j is an integer from 0 to 6; andPh is phenyl. 5. The compound of claim 4, wherein T1 is selected from the group consisting of —NH2, —OH, —SH, —SO3H, carboxyl, —COCl, —(CO)O(CO)R13, —CONHNH2,which is substituted or unsubstituted,which is substituted or unsubstituted,which is substituted or unsubstituted,azido, —NCS, —CHO, —COCH2I,phthal... | VisEn Medical, Inc. | A61B005/00 C07D263/62 C07D495/02
US08447380 | B2 | 2013-05-21 | Method for diagnosis of functional lung illnesses | 2004-05-17 | In a method for diagnosing functional lung illnesses, image exposures of the lungs are obtained at various phase points in time of the respiration of a subject, such as at maximum inhalation and maximum expiration, and the image exposures are segmented and at least two of the image exposures are compared on a segment-by-segment basis to identify a change in tissue density between the compared segments, as an indicator of lung functioning. | 1. A method for diagnosing functional lung illnesses comprising the steps of: operating a medical imaging apparatus to acquire, at each of a plurality of different respiration phase points in time in a succession during respiration of a subject, an image exposure of the lungs of the subject; andin a processor supplied with the image exposures, automatically electronically dividing each of said image exposures into a plurality of segments and automatically electronically comparing corresponding segments in at least two of said image exposures to automatically identify a change in tissue density of tissue in the lungs between the compared segments, and making an indication of said change available in electronic form at an output of said processor. 2. A method as claimed in claim 1 comprising generating said image exposures of the lungs of the subject at respective respiration phase points in time representing maximum inhalation and maximum expiration. 3. A method as claimed in claim 1 comprising acquiring said image exposures with a magnetic resonance apparatus as said medical imaging apparatus, and automatically determining a change in proton density between the compared segments as said change of said tissue density. 4. A method as claimed in claim 1 comprising, from said change in tissue density, automatically electronically calculating lung function data selected from the group consisting of ventilation, tidal volume, residual volume, and reserve capacity. 5. A method as claimed in claim 1 wherein the step of automatically electronically segmenting each of said image exposures comprises dividing each of said image exposures into a plurality of macroscopic areas, and wherein the step of comparing corresponding segments comprises comparing corresponding macroscopic areas in at least two of said image exposures by determining a signal representing air in the tissue contained in the compared macroscopic areas. 6. A method as claimed in claim 5 comprising determining corresponding macroscopic areas to be compared in said at least two of said image exposures by segmenting anatomical features in said at least two of said image exposures, associating at least one macroscopic area in each of said at least two image exposures with the segmented anatomical features therein, and localizing respective macroscopic areas in said at least two of said image exposures by determining displacement of said anatomical features in said at least two of said image exposures. 7. A method as claimed in claim 5 comprising identifying displacement of adjacent anatomical features in said at least two of said image exposures, and determining at least one of expansion or compression of said macroscopic area from displacement of said adjacent anatomical features. 8. A method as claimed in claim 5 comprising automatically electronically generating a color-coded image of the tissue density in the respective macroscopic areas, and superimposing said color-coded image on an... | Siemens Aktiengesellschaft | A61B005/08
US08371299 | B2 | 2013-02-12 | Ventilator aerosol delivery | 2008-04-10 | The delivery of a medicament to a mechanically ventilated patient may include the steps of delivering a respiratory maneuver to such a patient, monitoring the physiological effects of the respiratory maneuver on such patient, automatically establishing an adjustment to the respiratory maneuver in response to the monitored physiological effects, and delivering the adjusted respiratory maneuver to such a patient. | 1. A method for delivering a medicament during a single treatment session to a mechanically ventilated patient receiving a flow of breathing gas, the method being implemented in a respiratory therapy system, the method comprising: delivering a first respiratory maneuver to a patient during a single treatment session, wherein the single treatment session comprises multiple deliveries of respiratory maneuvers to the patient, wherein the first respiratory maneuver includes:(i) performing a temporary adjustment of a ventilation parameter pertaining to the flow of breathing gas for the patient, wherein the temporary adjustment is applied during a first time period within the first respiratory maneuver during the single treatment session, and wherein the temporary adjustment alters the flow of breathing gas delivered to the patient;(ii) activating a drug delivery device structured to aerosolize the medicament during a second time period within the first respiratory maneuver during the single treatment session, wherein the second time period overlaps with the first time period and has a duration that is less than a duration of the first time period; and(iii) delivering the aerosolized medicament to the patient via the flow of gas;measuring signals corresponding to physiological effects of the first respiratory maneuver on the patient, wherein the signals are measured during the single treatment session;determining a modification of the temporary adjustment of the ventilation parameter for a second respiratory maneuver included in the single treatment session, wherein the modification is determined based on the measured signals corresponding to the physiological effects of the first respiratory maneuver; anddelivering the second respiratory maneuver of the single treatment session to the patient using the modified temporary adjustment of the ventilation parameter, wherein the single treatment session further comprises a designated time period and a prescribed dose of medicament, wherein the prescribed dose of medicament is designated to be delivered during the designated time period using the multiple deliveries of respiratory maneuvers, and wherein the designated time period ranges from 5 to 30 minutes. 2. The method of claim 1, wherein determining the modification of the temporary adjustment of the ventilation parameter further includes determining the first time period and the second time period of the second respiratory maneuver. 3. The method of claim 1, wherein the ventilation parameter is a normal tidal volume, wherein the temporary adjustment of the normal tidal volume comprises increasing the normal tidal volume during the first respiratory maneuver, and wherein the modification of the temporary adjustment of the normal tidal volume used during the second respiratory maneuver comprises a modified normal tidal volume that is different than the increased normal tidal volume during the first respiratory maneuver. 4. The method of claim 1, wherei... | Respironics Respiratory Drug Delivery | A62B007/00
US08163003 | B2 | 2012-04-24 | Active blood vessel sleeve methods and systems | 2007-09-17 | Methods, apparatuses, computer program products, devices and systems are described that include accepting one or more blood vessel sleeve dimensions based on blood vessel data from an individual; and making a rapid-prototyped blood vessel sleeve at least partly based on the one or more blood vessel sleeve dimensions. | 1. A system comprising: circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data associated with at least one blood vessel;circuitry configured for determining at least one irregularity at least partly by identifying at least one change over time in the blood flow data associated with at least one blood vessel; andcircuitry configured for signaling, in response to determining the at least one irregularity, at least one actuation mechanism of the at least one blood vessel sleeve to release at least one therapeutic agent from at least one reservoir associated with the at least one blood vessel sleeve. 2. The system of claim 1, wherein the circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data associated with at least one blood vessel comprises: circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve, and in response to at least one request, blood flow data associated with at least one blood vessel. 3. The system of claim 1, wherein the circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data associated with at least one blood vessel comprises: circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve condition data associated with at least one blood vessel. 4. The system of claim 1, wherein the circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data associated with at least one blood vessel comprises: circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve rhythm data associated with at least one blood vessel. 5. The system of claim 1, wherein the circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data associated with at least one blood vessel comprises: circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve pressure data associated with at least one blood vessel. 6. The system of claim 1, wherein the circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data associated with at least one blood vessel comprises: circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve chemical data associated with at least one blood vessel. 7. The system of claim 1, wherein the circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data associated with at least one blood vessel comprises: circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve movement data associated with at least one blood vessel. 8. The system of claim 1, wherein the circuitry configured for receiving from at least one sensor of at least one blood vessel sleeve blood flow data as... | The Invention Science Fund I, LLC | A61F002/06 G06Q010/00 G06Q050/00 A61M029/00
US08560038 | B2 | 2013-10-15 | Method and apparatus for providing data processing and control in a medical communication system | 2008-05-14 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A method, comprising: receiving, using one or more processors, a signal associated with a monitored analyte level from an in vivo analyte sensor;retrieving, using the one or more processors, a predetermined number of stored signals associated with the monitored analyte level;processing, using the one or more processors, the received signal and the retrieved predetermined number of prior analyte level signals to determine a tolerance parameter;comparing, using the one or more processors, the tolerance parameter to a predetermined tolerance range corresponding to a dynamically varying uncertainty parameter determined based upon a number of data available for determining a rate of change of the monitored analyte level; anddetermining, using the one or more processors, a glucose trend information based on the received signal and the retrieved predetermined number of stored signals when the tolerance parameter is within the predetermined tolerance range such that no more than a predetermined number of the retrieved predetermined number of stored signals is missing, wherein the tolerance parameter is related to a temporal spacing of the received signal and the retrieved stored signals. 2. The method of claim 1 wherein the glucose trend information is determined based on analyte sensor sensitivity. 3. The method of claim 1 wherein the predetermined tolerance range includes a predetermined number of temporally spaced signals related to the monitored analyte level necessary to determine the glucose trend information. 4. The method of claim 1 wherein each of the predetermined number of retrieved stored signals associated with the monitored analyte level and the received signal associated with the monitored analyte level are temporally separated by approximately one minute or less. 5. The method of claim 1 wherein each of the predetermined number of retrieved stored signals associated with the monitored analyte level and the received signal associated with the monitored analyte level are temporally separated by approximately more than one minute. 6. The method of claim 1 wherein the glucose trend information includes a projected alarm associated with one or more of the rate of change of the monitored analyte level, or a direction of a projected analyte level. 7. The method of claim 1 including outputting the glucose trend information. 8. The method of claim 7 wherein the glucose trend information includes one or more of a visual output, an audible output, or a vibratory output. 9. An analyte monitoring device, comprising: a communication unit to receive an analyte level related signal from an in vivo analyte sensor;a data processing unit coupled to the communication unit; anda memory storing instructions which, when executed by the data processing unit, causes the data processing unit to retrieve a predetermined number of stored signals associated with a monitored analyte level, process the received analyte level related signal and the retrieved ... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/145 A61B005/05
US08718944 | B2 | 2014-05-06 | Patient-specific image-based computational modeling and techniques for human heart surgery optimization | 2008-05-22 | A method for determining cardiac status comprises for a given patient, constructing a patient-specific, three-dimensional, computational model of the patient's heart; and executing the constructed computational model, said executing generating a quantitative analysis of cardiac function. A method of performing cardiac surgeries comprises: a) assessing surgical options based on a patient-specific, three-dimensional, computational model of a patient's heart; and b) performing surgery based on one or more of the surgical options. A computer system comprises: a) a data source containing data of a patient's heart; b) a modeler coupled to receive data from the data source, the modeler generating a patient-specific, three-dimensional, computational model of the heart based on the heart data; and c) a processor routine for computationally providing information about a certain cardiac function using the three-dimensional heart model and for applying computational, quantitative analysis of the cardiac function, wherein the quantitative analysis of the cardiac function provides an assessment for surgical options, optimizing surgical techniques, or predicting outcomes. | 1. A method for determining cardiac status, comprising: for a given patient,a) using a computer-based modeler, constructing a patient-specific, three-dimensional, computational model of the patient's heart, the model including fluid-structure interactions and right ventricle, left ventricle, and patch and scar tissues of the patient's heart, wherein the fluid-structure interactions include deformation of the patient's heart and are represented by a system of differential equations for solid mechanics and fluid mechanics;by a processor, executing the constructed computational model, said executing generating a quantitative analysis of cardiac function, including assessment of ventricle volume based on a patient-specific simulation which simulates heart expansion and contraction by solving the differential equations, wherein the patient-specific simulation is based on ventricle pressure measured in the patient and includes mechanical stress/strain, flow velocity, shear stress and pressure; andvalidating by the processor the model with pre-operative patient-specific data, including the measured ventricle pressure, prior to assessing surgical options to establish a validated model useable for surgical design,the processor outputting to a user results of the generated quantitative analysis. 2. The method of claim 1, wherein the step of constructing comprises basing the model on patient-specific data including one or more of the data from ventricle morphology, cardiac motion, blood flow, material properties and pressure and volume conditions. 3. The method of claim 2, further comprising providing one or more of the patient-specific data before constructing the model. 4. The method of claim 3, wherein the step of providing one or more of the patient-specific data comprises acquiring the one or more patient-specific data using cardiac magnetic resonance. 5. The method of claim 1, further comprising constructing a three-dimensional, computational model of a healthy heart to serve as baseline control. 6. The method of claim 1, wherein the model further includes right ventricle/left ventricle/patch/scar interactions. 7. The method of claim 1, wherein the model includes valve mechanics. 8. The method of claim 7, wherein the valve mechanics comprise one or both of flow and solid stress/strain results. 9. The method of claim 1, wherein the model includes pulmonary regurgitation. 10. The method of claim 1, wherein the model includes fiber orientation and single-, double-, or multiple-layer anisotropic models. 11. The method of claim 10, wherein the anisotropic models comprise one or both of passive and active anisotropic models. 12. The method of claim 1, wherein the model includes an active contraction model. 13. The method of claim 12, wherein the step of constructing comprises introducing an external force field which is tied to fiber structure and orientations, thereby adding the active contraction model. 14. The method of claim 1, further comp... | Worcester Polytechnic Institute | G06F019/00 G06K009/00 G06K009/20 G06K009/36 A61B005/02
US09265451 | B2 | 2016-02-23 | Method and apparatus for determining spasticity | 2006-03-31 | There is provided a system and a method for quantitative measurement of spasticity in a patient. It has been found that stretch reflex measurements, that are quantitatively indicative of spasticity, can be obtained by recording an EMG signal while the limb is being moved at a variety of angular velocities. The method advantageously allows the clinician to perform the test at the bedside by eliminating the need for cumbersome mechanical components for moving the limb while providing quantitative measurements. | 1. A method of patient care to reduce spasticity and/or to improve motor control of a limb, the method comprising: a) providing measurement apparatus for measuring a joint angle and EMG activity in said limb;b) determining a threshold EMG activity value in said limb;c) determining a zero angle defining an initial position of said limb;d) imparting a movement to said limb by a clinician from said initial to a final position while measuring an EMG activity value in muscles of said limb and angles and velocities at which said limb is moved;e) recording angle relative to said zero angle and velocity as a data point at which said measured EMG value crosses said threshold EMG activity value to determine a threshold angle at said velocity;f) repeating said movement at a range of velocities until a set of data points are acquired and recorded to provide a relationship between velocities and threshold angles;g) algorithmically validating said data points based on said relationship between velocities and threshold angles;h) repeating one of step b), c), d), e), f) or combination thereof if said data points are outside predetermined quality criteria to obtain a set of valid data points;i) computing a stretch reflex threshold value and a sensitivity of said threshold value to velocities value using said valid data points;j) outputting said stretch reflex threshold value and said sensitivity of said threshold angles to velocities value to provide a quantitative measure of spasticity in said limb; andtreating a patient with a physical, pharmacological or surgical treatment based on said quantitative measure. |  | A61B005/04 A61B005/11 A61B005/0488 A61B005/00
US20080312515 | A1 | 2008-12-18 | System and method for providing cardiovascular disorder diagnosis services | 2007-12-07 | System and method provide an on-line high-performance diagnosis service for cardiovascular disorders. A client requests a high-performance diagnosis by transmitting real electrocardiographic treatment data and magnetocardiographic treatment data of a human body being a treatment object and virtual heart simulation parameters to a medical service server. The medical service server, in response to the diagnosis request, analyzes the real electrocardiographic treatment data to generate an electrocardiographic analysis result, and performs a virtual heart simulation using the simulation parameters to generate a pseudo electrocardiogram and magnetocardiogram. Further, the medical service server determines a disease state of the human body on the basis of the electrocardiographic analysis result, the magnetocardiographic treatment data and the pseudo electrocardiogram and magnetocardiogram, and generates definitive diagnosis data through comparison among the real magnetocardiographic treatment data, the electrocardiographic analysis result, the disease state, and a diagnosis criteria. The definitive diagnosis data is provided to the client. | 1. A diagnosis system for providing cardiovascular disorder diagnosis services through a network, comprising: a client group having one or more clients, each of which transmits real electrocardiographic treatment data and magnetocardiographic treatment data of a human body being a treatment object along with a cardiovascular disorder diagnosis request through the network, receives definitive diagnosis data as a reply to the cardiovascular disorder diagnosis request through the network; anda medical service server for analyzing the real electrocardiographic treatment data received through the network from the client in accordance with a task schedule utilizing available resource information, determining a disease state of the human body on the basis of the electrocardiographic analysis result, the real magnetocardiographic treatment data, and pseudo electrocardiogram and magnetocardiogram obtained through a virtual heart simulation, creating definitive diagnosis data on cardiovascular disorders of the human body on the basis of the real magnetocardiographic treatment data, the electrocardiographic analysis result and the determined disease state, and transmitting the created definitive diagnosis data through the network to the client. 2. The diagnosis system of claim 1, wherein each of the clients comprises: an electrocardiographic analysis block for analyzing an electrocardiographic signal in the treatment data to generate the real electrocardiographic treatment data;a magnetocardiographic analysis block for analyzing a magnetocardiographic signal in the treatment data to generate the real magnetocardiographic treatment data;a control block for transmitting the retrieved real electrocardiographic treatment data and the magnetocardiographic treatment data and virtual heart simulation parameters provided thereto along with the cardiovascular disorder diagnosis request to the medical service server, and receiving the definitive diagnosis data to be delivered to the client from the medical service server; anda diagnosis data storage block for storing the definitive diagnosis data. 3. A method of providing cardiovascular disorder diagnosis services through a network, comprising: requesting, by a client, a high-performance diagnosis on cardiovascular disorders by transmitting real electrocardiographic treatment data and magnetocardiographic treatment data of a human body being a treatment object and virtual heart simulation parameters through the network to a medical service server;analyzing, by the medical service server, in response to the high-performance diagnosis request, the real electrocardiographic treatment data to generate an electrocardiographic analysis result;performing, by the medical service server, a virtual heart simulation using the simulation parameters to generate a pseudo electrocardiogram and magnetocardiogram;determining, by the medical service server, a disease state of the human body on the basis of the electrocardiographic a... | RESEARCH AND INDUSTRIAL COOPERATION GROUP | A61B005/00
US20080312547 | A1 | 2008-12-18 | Cough Detecting Apparatus and Cough Detecting Method | 2006-09-13 | A cough detecting apparatus capable of high-precision couching detection is provided. This cough detecting apparatus includes a voice measuring section for detecting the voice of a subject, a body motion measuring section for detecting the body motion of the subject, and a cough detecting section for detecting coughing, based on the voice detected by the voice measuring section and the body motion detected by the body motion measuring section. | 1. A cough detecting apparatus comprising: a voice measuring section for detecting a voice of a subject to obtain voice information;a body motion measuring section for detecting a body motion of the subject to obtain body motion information; anda cough detecting section for detecting coughing, based on the voice information and the body motion information. 2. The cough detecting apparatus of claim 1, wherein the cough detecting section comprises; a first determining section for determining if coughing has occurred or not, based on the voice information; anda second determining section for determining if coughing has occurred or not, based on the body motion information;wherein, when both the first determining section and the second determining section determine that coughing has occurred, it is determined that the cough detecting section has detected coughing. 3. The cough detecting apparatus of claim 1wherein the cough detecting section detects coughing by computation based on the voice information and the body motion information. 4. The cough detecting apparatus of claim 1, wherein the body motion measuring section includes an accelerometer. 5. The cough detecting apparatus of claim 1, wherein the body motion measuring section includes a body motion sensor to be mounted on an abdominal region or a chest region. 6. A cough detecting method comprising steps of: measuring a voice of a subject by a voice measuring section to obtain voice information;measuring a body motion of the subject by a body motion measuring section to obtain body-motion information; anddetecting coughing by inputting the voice information and the body motion information into a cough detecting section. 7. The cough detecting method of claim 6, wherein the coughing detecting step comprises steps of: first determining if coughing has occurred or not, based on the voice information;second determining if coughing has occurred or not, based on the body motion information;wherein, when in both the first determining step and the second determining step it is determined that coughing has occurred, it is determined that coughing has been detected. | - | A61B005/11 A61B005/08
US20080262374 | A1 | 2008-10-23 | EVENT TRIGGERED INFECTION MONITORING | 2007-04-19 | Implantable medical devices include a sensor module capable of detecting an indicator of infection in proximity to the device. The sensor module is activated following a detection of an event associated with the device, such as, for example, refill of an implantable infusion device or implantation. The sensor module is deactivated two or more hours following activation. Methods include detecting a first event associated with an implanted medical device; activating a sensor module of the implanted medical device at or after the detection of the first event; deactivating the sensor module two hours or more after the sensor module is activated; and determining whether information regarding the indicator of infection is indicative of an infection. | 1. A method comprising: detecting a first event associated with an implanted medical device;activating a sensor module of the implanted medical device at or after the detection of the first event, the sensor module capable of obtaining and transmitting information regarding an indicator of infection;deactivating the sensor module two hours or more after the sensor module is activated; anddetermining whether the information regarding the indicator of infection is indicative of an infection. 2. The method of claim 1, wherein detecting the first event comprises detecting an event associated with risk of infection in proximity to the implanted medical device. 3. The method of claim 1, wherein detecting the first event comprises detecting an event capable of interfering with determining whether the information regarding the indicator of infection is indicative of an infection. 4. The method of claim 1, wherein detecting the first event comprises receiving instructions from a programmer. 5. The method of claim 1, wherein detecting the first event comprises detecting device activation following implantation. 6. The method of claim 1, wherein the implanted device is a refillable infusion device and wherein detecting the first event comprises detecting an event associated with refilling the device. 7. The method of claim 1, wherein detecting the first event comprises detecting an end of a recharge event. 8. The method of claim 7, wherein activating the sensor module comprises activating the sensor module ten minutes or more after the detection of the end of the recharge event. 9. The method of claim 1, wherein activating the sensor module comprises providing power to a sensor from a power source of the implanted device. 10. The method of claim 1, wherein activating the sensor module comprises activating a circuit to store information provided by a sensor in memory of the implanted device. 11. The method of claim 1, wherein activating the sensor module comprises activating the module such that the module provides sensed information in discrete time intervals. 12. The method of claim 1, wherein deactivating the sensor module two hours or more after the sensor module is activated comprises deactivating the sensor module between two hours and 90 days after the sensor module is activated. 13. The method of claim 1, wherein determining whether information regarding the indicator of infection is indicative of an infection is determined in the implantable medical device. 14. The method of claim 1, further comprising sending the information regarding the indicator of infection to a second device, wherein determining whether information regarding the indicator of infection is indicative of an infection is determined in the second device. 15. The method of claim 1, further comprising: detecting a second even associated with the implanted medical device; and reactivating the sensor module at or after the detection of the second event. 16. The method of... | MEDTRONIC, INC. | A61B005/053 A61B005/01
US20080312660 | A1 | 2008-12-18 | DEVICES AND METHODS FOR MEASURING THE SPACE AROUND A NERVE ROOT | 2008-06-16 | Described herein are method, systems and devices for measuring the region adjacent to or around a nerve root, such as the space within an intervertebral foramen before, during and/or after a spinal decompression procedure. Measurement devices may be advanced by pulling on them using a guidewire passing through the intervertebral foramen and out of the subject. The measurement device may include sounds for determining one or more dimensions of the space around a nerve root within an intervertebral space, lateral recess or central canal. Various embodiments of sounds, including calibrated, inflatable, expandable, moldable, and tapered sounds (or combinations of these) are described. | 1. A method of measuring the size of a compliant region adjacent to a patient's nerve root, the method comprising: advancing a guidewire from a first position outside of the patient's body, through an intervertebral foramen, and out of the patient's body at a second position;advancing a measurement device adjacent to a portion of the nerve root, wherein the measurement device is coupled to the guidewire; andestimating a size of the compliant region adjacent to the nerve root, based on the advancement of the measurement device. 2. The method of claim 1, wherein advancing the measurement device comprises pulling it into the intervertebral foramen and/or the lateral recess behind the guidewire, wherein the measurement device is coupled to the proximal end of the guidewire. 3. The method of claim 1, wherein multiple measurement devices are provided, each of a different diameter, and wherein estimating the size of the compliant region adjacent to the nerve root comprises determining a largest of the devices that can pass adjacent to the nerve root. 4. The method of claim 1, further comprising expanding an expandable region of the measurement device. 5. The method of claim 4, further comprising passing fluid into the expandable region of the measurement device to expand the region; andestimating the size based on an amount of fluid that can be passed into the expandable portion. 6. The method of claim 1, wherein estimating the size comprises estimating a cross sectional area of the intervertebral foramen. 7. The method of claim 1, wherein estimating the size comprises estimating a volume of the intervertebral foramen. 8. The method of claim 1, further comprising applying neural stimulation from the measurement device and monitoring for EMG signals. 9. The method of claim 1, further comprising molding a moldable region of the measurement device within the compliant region adjacent to the nerve root and withdrawing the molded region. 10. The method of claim 1, wherein guidewire is percutaneously advanced. 11. A method of measuring the size of a compliant region adjacent to a patient's nerve root as part of a spinal decompression procedure, the method comprising: advancing a guidewire from a first position outside of the patient's body, through an intervertebral foramen, and out of the patient's body at a second position;pulling the measurement device at least partially into the intervertebral foramen, wherein the measurement device is coupled to the proximal portion of the guidewire;expanding a portion of the measurement device; andestimating a size of the compliant region adjacent to the nerve root, based on the expansion of the measurement device. 12. The method of claim 11, wherein the step of expanding a portion of the measurement device comprises expanding a portion of the measurement device within the intervertebral foramen. 13. The method of claim 11, wherein expanding the portion of the measurement device comprises passing a fluid int... | Baxano, Inc. | A61B017/58 A61B019/00 A61B005/00
US20080317206 | A1 | 2008-12-25 | MEDICAL EXAMINATION SYSTEM CONTROL APPARATUS AND CONTROL METHOD THEREFOR | 2008-04-22 | A medical examination system causes a display unit to display an object for shifting to examination processing together with patient information corresponding to identification information of a patient received by a receiving unit if the patient information of the patient is input via an operation unit when the receiving unit receives the identification information of the patient from an external device. The medical examination system shifts to the examination processing without causing the display unit to display the object if the patient information is not input via the operation unit when the receiving unit receives the patient information. | 1. A medical examination system comprising: a receiving unit configured to receive identification information of a patient from an external device;an operation unit configured to input patient information about the patient; anda processing unit configured to cause a display unit to display an object for shifting to examination processing together with the patient information corresponding to the identification information of the patient received from the receiving unit if the patient information is input via the operation unit, and to execute the examination processing without causing the display unit to display the object if no patient information is input via the operation unit, when the receiving unit receives the identification information of the patient. 2. The medical examination system according to claim 1, wherein the processing unit causes the display unit to display the patient information corresponding to the identification information of the patient received by the receiving unit without causing the display unit to display the object if no patient information is input via the operation unit when the receiving unit receives the identification information of the patient. 3. The medical examination system according to claim 2, wherein the processing unit sets time that the display unit displays the patient information corresponding to the identification information of the patient received by the receiving unit without causing the display unit to display the object. 4. A medical examination system comprising: a receiving unit configured to receive identification information of a patient from an external device;an operation unit configured to input patient information about the patient; anda processing unit configured to register as examination reservation information the patient information corresponding to the identification information of the patient received by the receiving unit if the patient information is input via the operation unit when the receiving unit receives the identification information of the patient. 5. The medical examination system according to claim 4, wherein the processing unit causes the display unit to display the patient information corresponding to the identification information of the patient received by the receiving unit if the patient information is input via the operation unit when the receiving unit receives the identification information of the patient. 6. A method for a medical examination system, the method comprising: causing a display unit to display an object for shifting to examination processing together with patient information corresponding to identification information of a patient received by a receiving unit if the patient information is input via an operation unit when the receiving unit receives the identification information of the patient from an external device; andexecuting examination processing without causing the display unit to display the object if no patient information is in... | CANON KABUSHIKI KAISHA | H05G001/64 A61B005/00
US20080312513 | A1 | 2008-12-18 | Neurosurgical Candidate Selection Tool | 2006-03-21 | A system and method for neurosurgery candidacy assessment includes multiple data sources, wherein results of tests from at least some of the multiple data sources are integrated. A neurosurgery candidacy assessment report including a recommendation regarding candidacy for neurosurgery is provided based on the integrated results. The multiple data sources may include cognitive tests, a background data source, a medical data source, an anxiety/depression data source, and a motor skills data source. The medical data source may include a FLASQ-PD questionnaire. | 1-29. (canceled) 30. A computerized system for evaluating candidacy of a subject for neurosurgery, the system comprising: a cognitive testing data source, including at least one cognitive test for testing at least one cognitive domain of the subject, said at least one cognitive test providing cognitive data for said at least one cognitive domain;a medical data source, said medical data source for collecting data relating to indications and contraindications for candidacy for neurosurgery, said medical data source providing medical data;a processor for integrating said cognitive data and said medical data, said processor comprising a classification algorithm based on expert knowledge about the relevance of said cognitive data and said medical data to the candidacy for neurosurgery, said processor providing an output value for the candidacy of the subject for neurosurgery, said output value based on said classification algorithm; anda reporting module in communication with said processor and configured to display said output value. 31. The system of claim 30, further comprising an additional data source, said additional data source comprising at least one of: a background data source, an anxiety data source, a depression data source, and a motor skills data source, the additional data source configured to provide additional data for feeding into said classification algorithm to integrate said additional data with said collected cognitive data and said collected medical data. 32. The system of claim 30, wherein said medical data source is a neurosurgical candidacy questionnaire. 33. The system of claim 32, wherein said neurosurgical candidacy questionnaire comprises questions relating to at least one of: a diagnosis of idiopathic Parkinson's disease, a diagnosis of non-idiopathic Parkinson's disease, neurological function, and a history of medication. 34. The system of claim 30, wherein said at least one cognitive test is a battery of computerized tests for testing various cognitive domains. 35. The system of claim 30, wherein said classification algorithm is at least one of: a decision tree algorithm and a flow chart algorithm. 36. The system of claim 30, wherein said classification algorithm includes exclusion thresholds, wherein if individual data points within said cognitive data or said medical data exceed said exclusion thresholds, said output value is a recommendation for non-candidacy. 37. The system of claim 30, wherein said output value is one of: a recommendation for candidacy, a recommendation for non-candidacy, a probable recommendation for candidacy, a probable recommendation for non-candidacy and an inconclusive result. 38. A method of providing a recommendation for candidacy of a subject for neurosurgery, the method comprising: providing a computerized cognitive test to said subject;collecting cognitive data, for said subject, from said computerized cognitive test;providing a medical data source to the subject, the medical d... | - | A61B005/00 G06N005/02
US20080306363 | A1 | 2008-12-11 | Specialized Human Servo Device And Process For Tissue Modulation Of Human Fingerprints | 2006-01-06 | Apparatus and methods for noninvasive spectroscopic measurement of an analyte in a subject that have been optimized for producing uniform and repeatable tissue modulation across test subjects and for the same test subject on different occasions are provided. The apparatus comprises an ergonomically shaped grip that substantially conforms to a subject's hand; a surface for placement of at least one of the subject's fingertips upon grasping the grip; and an optically transparent aperture, or a plurality of apertures, disposed within the surface. A modification to the surface of the apparatus adjacent to the aperture that is detectable via the tactile sense of the subject can be added to provide tactile feedback to the subject to guide correct placement of the fingertip over the aperture. The apparatus and methods can also incorporate feedback methods to guide and optimize placement and conditions of the fingertip to further improve accuracy of measurements. | 1. An apparatus for noninvasive spectroscopic measurement of an analyte in a subject comprising: (a) an ergonomically shaped grip that substantially conforms to a subject's hand;(b) a surface for placement of at least one of the subject's fingertips upon grasping the grip; and(c) an optically transparent aperture disposed within the surface. 2. The apparatus of claim 1, further comprising a modification to the surface adjacent to the aperture, wherein the modification is detectable via the tactile sense of the subject. 3. The apparatus of claim 2, wherein the modification comprises at least one raised nub, bump and/or ridge on the surface of the apparatus. 4. The apparatus of claim 1, further comprising one or more additional apertures. 5. The apparatus of claim 1, wherein the ergonomically shaped grip comprises two or more ridges that define recesses, whereby the recesses conform to the subject's fingers upon grasping the grip. 6. The apparatus of claim 1, wherein the grip is shaped to facilitate placement of the volar side of the subject's thumb tip on a portion of the grip that opposes the surface for placement of the subject's fingertips, whereby the subject moves the fingertips and opposing thumb tip towards each other upon grasping the grip. 7. The apparatus of claim 1, further comprising a sensor in the surface of the grip adjacent to the aperture, wherein the sensor detects proper positioning of a fingertip over the aperture. 8. The apparatus of claim 7, wherein the sensor comprises an electrode array. 9. The apparatus of claim 8, wherein the electrode array comprises a plurality of electrodes, and each electrode detects 60 Hz electrical activity. 10. The apparatus of claim 7, wherein the sensor detects pressure applied by a fingertip. 11. The apparatus of claim 7, wherein the sensor detects moisture. 12. The apparatus of claim 11, wherein moisture is detected via resistance, capacitance or impedance between the electrodes. 13. The apparatus of claim 1, further comprising a feedback loop that transmits a detectable signal that corresponds to information detected by the sensor. 14. The apparatus of claim 13, wherein the detectable signal is transmitted to a processor. 15. The apparatus of claim 13, wherein the detectable signal is transmitted to the subject. 16. The apparatus of claim 1, further comprising means for measuring temperature of the surface. 17. The apparatus of claim 1, further comprising means for altering the temperature of the surface. 18. The apparatus of claim 1, further comprising a spectroscopic measurement system that directs light through the aperture toward the subject's fingertip and detects spectra emitted from the subject's fingertip through the aperture. 19. A method for noninvasive spectroscopic measurement of an analyte in a subject comprising: (a) contacting the subject's hand with an apparatus of claim 1;(b) positioning a fingertip of the subject over the aperture;(c) directing light through... | Lightouch Medical, Inc. | A61B005/1455
US20080242949 | A1 | 2008-10-02 | Computational user-health testing | 2007-05-15 | Methods, apparatuses, computer program products, devices and systems are described that carry out implementing in at least one device at least one user-health test function that is structurally distinct from at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device; obtaining user data in response to an interaction between a user and the at least one application; and presenting an output of the at least one user-health test function at least partly based on the user data. | 1. A method comprising: implementing in at least one device at least one user-health test function that is structurally distinct from at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device;obtaining user data in response to an interaction between a user and the at least one application; andpresenting an output of the at least one user-health test function at least partly based on the user data. 2. (canceled) 3. The method of claim 1 wherein the implementing in at least one device at least one user-health test function that is structurally distinct from at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device comprises: implementing in the at least one device at least one alertness or attention test function that is structurally distinct from the at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device. 4. The method of claim 1 wherein the implementing in at least one device at least one user-health test function that is structurally distinct from at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device comprises: implementing in the at least one device at least one memory test function that is structurally distinct from the at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device. 5. The method of claim 1 wherein the implementing in at least one device at least one user-health test function that is structurally distinct from at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device comprises: implementing in the at least one device at least one speech test function that is structurally distinct from the at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device. 6. The method of claim 1 wherein the implementing in at least one device at least one user-health test function that is structurally distinct from at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device comprises: implementing in the at least one device at least one calculation test function that is structurally distinct from the at least one application whose primary function is different from symptom detection, the at least one application being operable on the at least one device. 7. The method of claim 1 wherein the implementing in at least one device at least one user-health test function that is stru... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00 G01N033/48
US20080249374 | A1 | 2008-10-09 | METHOD AND APPARATUS TO ENABLE MULTIPLE METHODS OF CLINICAL ORDER INPUT INTO A HEALTHCARE IT APPLICATION | 2007-04-06 | The present invention relates to a system for facilitating the input of clinical orders into a healthcare application. This system can provide a variety of ways for a user to interface with the system and a variety of ways to facilitate data input into an order form. The system can facilitate selection of an order form by selecting an order form for the user, or by providing the user with a list of order forms that the user can select an order form from. The system facilitates the input of data into an order form in several ways. This system can use stored information to pre-populate some of the information fields in an order form. Data input can also be facilitated by the use of predictive typing, keywords, and drop-down lists. The system can also facilitate data input by taking free form text, parsing the text and entering the text into at least one information fields on an order form. | 1. A system for clinical order entry comprising: a user interface accepting input data from a user;a memory comprising stored data; anda processor using said input data and said stored data to facilitate selection of an order form, wherein said order form has at least one information field, and wherein said processor provides facilitated entry of data into at least one of said information fields. 2. The system for clinical order entry of claim 1 wherein said processor facilitates selection of said order form using said input data and said stored data by at least one of selecting an order form for said user and providing said user with a list of order forms, wherein said user can select an order form from said list of order forms. 3. The system for clinical order entry of claim 1 wherein said facilitated entry of data into at least one of said information fields comprises at least one of predictive typing, the use of keywords and providing a drop-down list for said information fields. 4. The system for clinical order entry of claim 3 wherein said information field has a category, and wherein at least one of said category of said information field, said order form, said stored data, said input data, and data in other of said information fields is used to form said drop-down list, and wherein said user can select order information from said drop-down list in order to enter data into at least one of said information fields. 5. The system for clinical order entry of claim 4 wherein said category of said information field comprises at least one of patient name, age, date of birth, medical reference number, height, weight, allergies, diagnosis, orderable, dose, formulation, route, frequency, PRN, duration, refills, samples, procedure, test result, test order, and insurance information. 6. The system for clinical order entry of claim 1 wherein said stored data used to form said drop-down list comprises at least one of previous orders for a patient, previous patient information, and clinical order information, wherein said clinical order information comprises orderable test information, and orderable drug information, wherein said orderable drug information comprises at least one of the orderable drug, doses, formulations, routes of delivery, frequency, duration, refills, and further instructions for each diagnosis. 7. The system for clinical order entry of claim 4 wherein said drop-down list comprises only said clinical order information that is commonly ordered. 8. The system for clinical order entry of claim 1 wherein said user can customize said order form by at least one of adding additional said information fields, selecting further instructions, and adding special instructions. 9. The system for clinical order entry of claim 1 wherein said user interface comprises at least one of a keyboard, mouse, touchpad, and voice interface. 10. The system for clinical order entry of claim 1 wherein said system can identify a possible adverse drug inte... | General Electric Company | A61B005/00
US20080287810 | A1 | 2008-11-20 | Miniature actuator mechanism for intravascular optical imaging | 2006-05-02 | The present invention relates to a new intravascular imaging device based on a Shape Memory Alloy (SMA) actuator mechanism embedded inside an elongate member such as a guide wire or catheter. The present invention utilizes a novel SMA mechanism to provide side-looking imaging by providing movement for an optical coherence tomography (OCT) element. This novel SMA actuator mechanism can be easily fabricated in micro-scale, providing an advantage over existing imaging devices by offering the ability to miniaturize the overall size of the device. Because the device does not require a rotating shaft or fiber optic along the length of the catheter, it also allows for a more flexible catheter or guide wire, and provides room for other interventional devices. The device simplifies the manufacture and operation of OCT by allowing a straight fiber optic directed by an independent, oscillating reflector or prism controlled by the actuator mechanism located only in the distal tip of the device. A variation uses the actuator mechanism to rotate only the distal end of the optical fiber, eliminating the need to spin the entire fiber via a remote mechanism. Also disclosed are methods of using the same. | 1. A side-looking intravascular optical coherence tomography apparatus comprising: an elongate member having a proximal end and a distal end, wherein at least a portion of said distal end is at least partially transparent to light energy;an actuator mechanism disposed in said distal end, said actuator mechanism comprising a first anchor, a second anchor, at least one movable element, a first SMA actuator connected to said first anchor and at least one movable element, and a deformable component connected to said second anchor and at least on movable element, wherein said anchor elements are secured relative to said elongate member;an optical fiber having a proximal and a distal end, said distal end of said optical fiber disposed in said distal end of said elongate member substantially parallel to a longitudinal axis of said elongate member; anda reflector connected to said movable element, said reflector oriented to reflect light energy from a distal tip of said optical fiber through said transparent portion of said distal end at an angle of between about 15° to about 165° relative to said longitudinal axis of said elongate member;wherein said first SMA actuator has an activated and a deactivated state; andwherein said movable element and mirror move in a first direction relative to the elongate member upon activation of said first SMA actuator. 2. The apparatus of claim 1, wherein said deformable component comprises a second SMA actuator; wherein said second actuator has an activated and a deactivated state; andwherein activation of said second SMA actuator following deactivation of said first SMA actuator moves said movable element and reflector relative to the elongate member in a second direction of movement which is counter to said first direction of movement. 3. The apparatus of claim 1, wherein said deformable component is elastic or superelastic; wherein said deformable component has a relaxed state and a deformed state;wherein said deformable component is in a relaxed state when said first SMA actuator is deactivated;wherein said movement of said movable element and reflector in said first direction upon activation of said first SMA actuator deforms said elastic or superelastic deformable component; andwherein following deactivation of said first SMA, said elastic or superelastic deformable component substantially returns to said relaxed state, said movable element and reflector moving in a second direction of movement which is counter to said first direction of movement. 4. The apparatus of claim 2, wherein said first and second direction of movement is rotational about the longitudinal axis of said elongate member, or substantially parallel to the longitudinal axis of said elongate member. 5. The apparatus of claim 3, wherein said first and second direction of movement is rotational about the longitudinal axis of said elongate member, or substantially parallel to the longitudinal axis of said elongate member. 6. The apparatus of c... | - | A61B005/1459
US07462491 | B2 | 2008-12-09 | Methods and compositions for diagnosis and monitoring of prostate cancer progression by detection of serum caveolin | 2004-09-15 | Serum cav-1 is disclosed as a biomarker for prostate cancer that has the power to differentiate between prostate cancer and BPH patients. | 1. A method of diagnosing prostate cancer comprising detecting caveolin-1 in the serum of a subject suspected of or susceptible to developing prostate cancer, wherein a serum caveolin-1 level that is significantly higher than the serum caveolin-1 of a pool of subjects without prostate cancer is indicative of prostate cancer. 2. The method of claim 1, wherein the caveolin-1 is detected in the HDL3 fraction of the serum. 3. The method of claim 1, wherein the caveolin-1 is detected by an anti-caveolin-1 antibody in an assay. 4. The method of claim 3, wherein the assay is a sandwich immunoassay. 5. The method of claim 3, wherein the antibody is a rabbit antibody. 6. The method of claim 3, wherein the antibody is a human antibody. 7. The method of claim 1, wherein the subjects without prostate cancer have benign prostatic hyperplasia. | Baylor College of Medicine | G01N033/48 G01N033/53 A61B005/055
US07471973 | B2 | 2008-12-30 | Determining a surface geometry of an object | 2005-04-21 | Methods for determining a surface geometry of an object, including determining a first projection matrix based on a first imaging device, determining a second projection matrix based on a second imaging device, obtaining at least one first two-dimensional (2D) image of the object using the first imaging device, obtaining at least one second 2D image of the object using the second imaging device, determining a contour of the object in the first 2D image and the second 2D image, and, based on the at least two contours, the first projection matrix, and the second projection matrix, reconstructing 3D data associated with the surface of the object. In one embodiment, the object can be a heart. | 1. A method for determining a surface geometry of an object, the method comprising: determining a first projection matrix based on a first imaging device,determining a second projection matrix based on a second imaging device,obtaining at least one first two-dimensional (2D) image of the object using the first imaging device,obtaining at least one second 2D image of the object using the second imaging device,determining at least two contours, the at least two contours comprising a contour of the object in the first 2D image and a contour of the object in the second 2D image,based on the at least two contours, the first projection matrix, and the second projection matrix, reconstructing 3D data associated with the surface of the object;registering the surface of the object with a second object surface; andemploying a boundary element method to determine the geometry between the surface of the object and the second object surface. 2. A method according to claim 1, wherein obtaining at least one first two-dimensional (2D) image of the object using the first imaging device include using at least one of fluoroscopy and a digital camera. 3. A method according to claim 1, wherein obtaining at least one first two-dimensional (2D) image of the object using the first imaging device, and obtaining at least one second 2D image of the object using the second imaging device, include using multi-planar fluoroscopy. 4. A method according to claim 1, wherein determining a contour includes determining a contour based on at least one of a user provided designation and an automatic computation. 5. A method according to claim 1, wherein reconstructing 3D data includes computing a fundamental matrix based on the first projection matrix and the second projection matrix. 6. A method according to claim 1, wherein reconstructing 3D data includes forming a meshed surface based on the first projection matrix, the second projection matrix, and at least one matching contour point pair based on the at least two contours. 7. A method according to claim 1, wherein reconstructing 3D data includes creating splines based on 3D points reconstructed from the first projection matrix, the second projection matrix, and at least one matching contour point pair based on the at least two contours. 8. A method according to claim 7, further including convening the spline data to spherical coordinates. 9. A method according to claim 7, further including at least one of extrapolating and fitting the spline data. 10. A method according to claim 9, further including convening the data to Cartesian coordinates. 11. A method according to claim 7, further including fitting the spline data based on a bicubic spline least squares fit. 12. A method according to claim 1, wherein the surface of the object is a heart surface, and the second object surface is an electrode vest surface. 13. A method according to claim 12, wherein the second imaging device is the same device as the first imaging... | Case Western Reserve University | A61B005/044
US07447543 | B2 | 2008-11-04 | Pathology assessment with impedance measurements using convergent bioelectric lead fields | 2005-02-15 | A method and apparatus for measuring impedance for pathology assessment in a living being using convergent bioelectric lead fields is disclosed, including injecting a current between first and second electrodes implanted in a body of a living being, where the first and second electrodes define a first electric lead field oriented between the first and second electrodes. A potential difference is measured between third and fourth electrodes implanted in the body, where the potential difference results from the current injected between the first and second electrodes. The third and fourth electrodes define a second electric lead field oriented between the third and fourth electrodes. The first and second electric lead fields converge near an assessment site within the body, but are substantially separated otherwise. An impedance value is calculated based on the potential difference and the current injection, and is used to assess a pathology near the assessment site. | 1. A method of measuring impedance for pathology assessment in a living being, the method comprising: injecting a current between first and second electrodes implanted in a body of a living being, the first and second electrodes defining a first electric lead field in the body, the first electric lead field oriented between the first and second electrodes;measuring a potential difference between third and fourth electrodes implanted in the body, the potential difference resulting from the current injected between the first and second electrodes, the third and fourth electrodes defining a second electric lead field in the body, the second electric lead field oriented between the third and fourth electrodes, wherein the first and second electric lead fields converge near an assessment site within the body, but are substantially separated otherwise; andcalculating an impedance value based on the potential difference and the current injection, and using the impedance value to assess a pathology near the assessment site. 2. The method of claim 1, wherein the assessment site is not within a heart chamber. 3. The method of claim 2, wherein the assessment site is a left lung and the pathology is pulmonary edema. 4. The method of claim 2, wherein the assessment site is a left ventricular wall of a heart and the pathology is myocardial ischemia. 5. The method of claim 2, wherein the assessment site is a left ventricular wall of a heart and the pathology is left ventricular hypertrophy. 6. The method of claim 1, wherein the second and fourth electrodes are positioned near a left ventricle of a heart, the first electrode is positioned in a pectoral region, and the third electrode is positioned in a right atrium of the heart. 7. The method of claim 1, wherein the second and fourth electrodes are positioned near a left ventricle of a heart, the first electrode is positioned in a pectoral region, and the third electrode is positioned near a brachiocepahlic vein. 8. The method of claim 1, wherein the second and fourth electrodes are positioned near a left ventricle of a heart, the first electrode is positioned in a pectoral region, and the third electrode is positioned in a superior vena cava. 9. The method of claim 1, wherein the second and fourth electrodes are positioned near a left ventricle of a heart, the first electrode is positioned in a right ventricle of the heart, and the third electrode is positioned in a right atrium of the heart. 10. The method of claim 1, wherein the second and fourth electrodes are positioned near a left ventricle of a heart, the first electrode is positioned in a right ventricle of the heart, and the third electrode is positioned in a superior vena cava. 11. The method of claim 1, wherein the second and fourth electrodes are separated by about fifteen millimeters. 12. The method of claim 1, wherein the second and fourth electrodes are separated by about twenty-five millimeters. 13. The method of claim 1, wherein the ... | Regents of the University of Minnesota | A61B005/053
US07541602 | B2 | 2009-06-02 | System and method for noninvasively monitoring conditions of a subject | 2007-06-04 | A method and system are presented for use in determining one or more parameters of a subject. A region of interest of the subject is irradiated with acoustic tagging radiation, which comprises at least one acoustic tagging beam. At least a portion of the region of interest is irradiated with at least one electromagnetic beam of a predetermined frequency range. Electromagnetic radiation response of the at least portion of the region of interest is detected and measured data indicative thereof is generated. The detected response comprises electromagnetic radiation tagged by the acoustic radiation. This enables processing of the measured data indicative of the detected electromagnetic radiation response to determine at least one parameter of the subject in a region corresponding to the locations in the medium at which the electromagnetic radiation has been tagged by the acoustic radiation, and outputting data indicative of the at least one determined parameter. | 1. A method for use in determining one or more parameters of a subject, said method comprising; (a) irradiating a region of interest of the subject with acoustic tagging radiation, the acoustic tagging radiation comprising at least one acoustic tagging beam being a coded continuous acoustic wave in the form of a predetermined function of at least one parameter of the acoustic radiation varying over time during a measurement time interval, said predetermined time function having a narrow autocorrelation;(b) irradiating at least a portion of the region of interest with at least one electromagnetic beam of a predetermined frequency range;(c) detecting an electromagnetic radiation response of said at least portion of the region of interest and generating data indicative thereof, said response comprising electromagnetic radiation tagged by the acoustic radiation, thereby enabling processing said data indicative of the detected electromagnetic radiation response, to determine at least one parameter of the subject in a region corresponding to the locations in the medium at which the electromagnetic radiation has been tagged by the acoustic radiation, and output data indicative of the at least one determined parameter. 2. The method according to claim 1, wherein said continuous acoustic wave is coded by modulating said at least one wave parameter selected from the following: phase, amplitude, frequency, frequency gradient, phase jump, duty cycle, chirp gradient. 3. The method according to claim 1, wherein said continuous acoustic wave is coded by modulating its phase. 4. The method according to claim 3, wherein the phase modulated continuous acoustic wave is generated as a sequence of pulses having equal or not values of at least one of the following parameters: duration, amplitude, frequency, frequency gradient, phase jump, duty cycle and chirp gradient. 5. The method according to claim 1, wherein said acoustic tagging beam is produced by operating an acoustic transducer with at least one coding signal in the form of an arbitrary sequence to thereby modulate output of the acoustic transducer in at least one of frequency, phase and amplitude parametric domains. 6. The method according to claim 5, wherein said arbitrary sequence is a pseudo random sequence with a narrow autocorrelation. 7. The method according to claim 5, wherein said at least one coding signal includes at least one of Barker and Golay codes. 8. The method according to claim 5, wherein said at least one coding signal includes short pulses of high amplitude and separation periods of low amplitude. 9. The method according to claim 8, wherein the separations periods are selected such that the phase of light propagating through the media during each pulse is uncorrelated to the phase of light propagating through the media during the timely adjacent pulses. 10. The method according to claim 1, comprising processing said data indicative of the detected electromagnetic response by cross... | Or-Nim Medical Ltd. | A61B005/05
US08060175 | B2 | 2011-11-15 | System and apparatus for collecting physiological signals from a plurality of electrodes | 2007-06-15 | A system and apparatus for obtaining physiological data from a patient. The system and apparatus comprising at least one electrode disposed to collect physiological data from the patient and an electrode connection device having a conductive array formed by a plurality of conductive regions and a plurality of nonconductive regions. The conductive regions are suitable to be connected to the at least one electrode. The electrode may be connected to a first conductive region of the electrode communication device and wherein upon movement by the patient or the conductive array, the electrode is connected to a second conductive region of the conductive array. | 1. A system for obtaining physiological data from a patient, the system comprising: a plurality of physiological electrodes positionable on the patient and independently movable with respect to each other, each of the physiological electrodes operable to obtain physiological data from the patient, each electrode including a first conductive element positionable in contact with the patient to obtain the physiological data from the patient and a second conductive element disposed to transmit the physiological data obtained by the first conductive element;a conductive array having a plurality of spaced conductive regions formed along a generally planar contact surface, the conductive array positioned over the plurality of physiological electrodes such that the second conductive element of each physiological electrode of the plurality directly and removably contacts at least one of the conductive regions of the conductive array; andat least one communication path connected to the conductive array to conduct at least the physiological data from the patient to a data monitor;wherein the second conductive element of each electrode directly contacts at least one electrical connection region and transmits collected physiological data through the second conductive element to the at least one electrical connection region to which the second conductive element directly contacts and upon movement by the patient or the plurality of electrical connection regions, the second conductive element directly contacts an alternative electrical connection region and transmits collected physiological data from the second conductive element to the alternative electrical connection region. 2. The system of claim 1 wherein the plurality of conductive regions are spaced from each other by non-conductive regions, wherein the non-conductive regions are smaller than the second conductive element of each of the plurality of electrodes such that the second conductive element of each of the plurality of electrodes is in contact with at least one of the conductive regions when the conductive array is placed over the plurality of electrodes. 3. The system of claim 1 wherein the at least one communication path includes a plurality of data wires each having a first end connected to one of the conductive regions. 4. The system of claim 1 wherein the at least one communication path includes a single communication bus connected to each of the conductive regions, the single communication bus being configured to communicate with the data monitor. 5. The system of claim 1 wherein each of the electrodes includes electrode identifying circuitry operable to produce a signal indicative of an anatomical location on the patient to which the electrode is to be attached. 6. An apparatus for obtaining physiological data from a patient, the apparatus comprising: a plurality of electrodes configured to attach to a patient with a first conductive element and collect physiological data from the pat... | General Electric Company | A61B005/04
US08112141 | B2 | 2012-02-07 | Method and device for quick press on EEG electrode | 2008-05-22 | Embodiments relate to methods and systems for monitoring bioelectric potentials. In some instances, an electrode is applied to a patient's skin. The electrode may be at least partly inserted into the patient's skin, such as by inserting at least part of one or more teeth underneath the skin. | 1. An electrode comprising: a main portion defining a plane; a plurality of legs extending from said main portion and away from the plane; and a plurality of protrusions extending from each of said legs wherein said legs are configured to bend, upon application of a force, towards the plane so as to define a substantially flat shape in an attack configuration and to return towards an initial configuration upon release of said force, wherein said electrode comprises a shape memory material. 2. The electrode of claim 1, wherein said legs are bendable such that the distal end of each leg moves away from the main portion upon application of the force. 3. The electrode of claim 1, wherein the length of said protrusions is between about 300 to about 600 microns. 4. The electrode of claim 1, wherein said protrusions comprise sharp points configured to penetrate skin of the patient. 5. The electrode of claim 1, wherein said electrode comprises stainless steel. 6. The electrode of claim 1, wherein said electrode comprises an alloy. 7. The electrode of claim 1, wherein said electrode comprises a nitinol. 8. The electrode of claim 1, wherein said electrode is between about 0.002 and about 0.010 inches thick. 9. An electrode system, the system comprising: a plurality of electrodes according to claim 1; andsupport structure to support the plurality of electrodes,wherein said system is configured to be worn on a patient's head. 10. The electrode system of claim 9, further comprising a power supply. 11. The electrode system of claim 9, further comprising at least one of an amplifier, a filter and an A/D converter. 12. The electrode system of claim 9, further comprising a wireless transmitter. 13. The electrode system of claim 9, wherein said support structure comprises a plurality of straps configured to be wrapped around the patient's head. 14. A kit comprising: a plurality of electrodes according to claim 1;support structure to support the plurality of electrodes;a wireless transmitter; anda receiver configured to receive signals from the wireless transmitter,wherein said system is configured to be worn on a patient's head. 15. An electrode comprising: a main portion defining a plane; a plurality of legs extending from said main portion and away from the plane; and a plurality of protruding portions extending from each of said legs and configured to stabilize the electrode upon attaching the electrode to a patient, wherein said legs are configured to bend, upon application of a force, towards the plane so as to define a substantially flat shape in an attack configuration and to return towards an initial configuration upon release of said force, wherein the electrode comprises a shape-memory material. 16. The electrode of claim 15, wherein at least two of said protruding portions are configured to be separated by a first distance upon application of a force and to be separated by a second distance upon release of said force, said first distance ... | Persyst Development Corporation | A61B005/04
US08226581 | B2 | 2012-07-24 | Bite force gauge | 2008-07-25 | A bite force measuring device includes a handle having first and second ends. A measurement shaft extends outwardly from the handle first end and includes a distal end relative to the handle. A test plate is positioned on the measurement shaft. The measuring device includes a flexible reservoir positioned inside the measurement shaft that is in contact with the test plate that contains a generally incompressible fluid such that the flexible reservoir undergoes pressure changes when the test plate bears against the reservoir. A plunger shaft is operatively connected to the reservoir and is pushed outwardly therefrom when pressure in the reservoir is increased. The plunger shaft may include indicia indicative of a measurement of a relative bite force. | 1. A patient bite force measuring device, comprising: a handle having first and second ends and at least one side wall extending therebetween, said handle defining an interior space;a measurement shaft fixedly attached to said handle first end and extending outwardly therefrom and having a distal end relative to said handle;a test plate mounted adjacent said distal end of said measurement shaft;a flexible reservoir positioned inside said measurement shaft in operative communication with said test plate and containing a generally incompressible fluid, said reservoir configured to experience a pressure change when said test plate bears against said reservoir;wherein said incompressible fluid is hydraulic fluid;an elongate channel in fluid communication with said reservoir and extending linearly in said interior space of said handle;a plunger shaft having one end operatively positioned in said channel and configured for relative movement therein when a pressure of said fluid is increased or decreased;wherein: said handle second end defines an opening;said plunger shaft includes a second end configured to extend through said opening, said plunger shaft moving outwardly through said opening when a relative pressure in said channel is increased and moving inwardly when a relative pressure in said channel is decreased;a compression spring positioned adjacent said opening and being operatively connected to said plunger shaft, said spring being compressed when said plunger shaft is moved outwardly through said opening, said spring being normally biased to urge said plunger shaft inwardly relative to said handle; andwherein: said plunger shaft includes indicia proximate said plunger shaft second end that is indicative of a relative degree of pressure such that the patient bite force upon said test plate is visually measurable as said plunger shaft extends from said handle second end. 2. The measuring device as in claim 1, wherein said handle is constructed of a hard plastic material. 3. A patient bite force measuring device, comprising: a handle having first and second ends and at least one side wall extending therebetween, said handle defining an interior space;a measurement shaft extending outwardly from said handle first end and having a distal end relative to said handle;a test plate mounted adjacent said distal end of said measurement shaft;a flexible reservoir positioned inside said measurement shaft in operative communication with said test plate and containing a generally incompressible fluid, said reservoir configured to experience a pressure change when said test plate bears against said reservoir;an elongate channel in fluid communication with said reservoir and extending linearly in said interior space of said handle;a plunger shaft having a first end operatively positioned in said channel and configured for relative movement therein when a pressure of said fluid is increased or decreased;said handle second end defines an opening;said plunger... |  | A61B005/103 A61B005/117 G01L001/00 G01L005/00
US08123685 | B2 | 2012-02-28 | System for monitoring a number of different parameters of a patient in a bed | 2006-09-29 | In order to provide a simple and reliable monitoring technique for a number of different parameters of a patient in a bed (2), a system (1) is provided, said system (1) comprising: at least one first sensor device (25) adapted to measure a first force component acting in a first direction (12), said first force component corresponding to a first parameter of the patient, at least one second sensor device (36) adapted to measure a second force component acting in a second direction (13, 14), said second force component corresponding to a second parameter of the patient, and a data device (16) adapted to acquire the measured data from the at least one first sensor device (25) and the at least one second sensor device (36) and to process those data in order to provide information about the number of parameters of the patient. | 1. A system for monitoring a plurality of physiological parameters of a patient in a bed, the system comprising: a base frame;a patient support;a plurality of sensor units mounted between the base frame and the patient support to support the patient support, each of the sensor units including: a first sensor device which measures a first force component in a first direction, the first sensor device being attached to one of the patient support and the base frame;a rail mounted to the other of the patient support and the base frame;a slide which is supported on the rail and movable relative to the rail, wherein the first sensor device is in contact with the slide;a second sensor device attached adjacent the rail to interact with the slide to measure a second force component in a second direction, the second direction being different from the first direction; anda data device which receives measured force component data from the sensor units and processes the force component data to provide information about the plurality of physiological parameters of the patient. 2. The system as claimed in claim 1, further including: a resilient element which biases the slide against the second sensor device. 3. The system as claimed in claim 1, in which the first sensor device comprises a strain gauge sensor. 4. The system as claimed in claim 1, in which the second direction of the second force component measured by the second sensor device is in a horizontal direction. 5. The system as claimed in claim 1, in which the second sensor device comprises a piezo sensor. 6. The system as claimed in claim 1, wherein the first direction of the component measured by the first sensor device is vertical. 7. The system as claimed in claim 6, wherein the sensor units are arranged such that the second force component measured by at least one of the sensor units is in a first horizontal direction and the second force component force component measured by at least one other of the sensor devices is in a second horizontal direction transverse to the first horizontal direction. 8. A system for monitoring at least respiration and heart rate parameters of a patient in a bed, the system comprising: a base frame;sensor units mounted to the base frame, each sensor unit including: a first sensor device which measures a first force component acting in a first direction on a slide which is movable in a second direction orthogonal to the first direction, the first sensor generating a first signal indicative of the first force component;a second sensor device which measures a second force component acting in the second direction on the slide and generates a second signal indicative of the second force component;a patient support mounted to and supported by the sensor units and wherein each sensor unit enables relative movement between the base frame and the patient support via the slide; anda data device which receives the first and second signals from the sensor units and processes... | Koninklijke Philips Electronics N.V. | A61B005/00 A61B005/02 A61B005/08 A61B005/103 A61B005/117 A47C031/00 A47B071/00
US08182429 | B2 | 2012-05-22 | Enhanced functionality and accuracy for a wrist-based multi-parameter monitor | 2006-10-30 | A method for adjusting a blood pressure measurement to compensate for a pressure offset due to obtaining the blood pressure measurement at an elevation other than heart level is illustrated. The method includes measuring a blood pressure at a location above or below a heart level and adjusting the measured blood pressure in accordance with a pressure difference between the location at which the blood pressure is obtained and the heart level. | 1. A physiological measurement system comprising: a wrist mounted unit configured to be mounted to a wrist of a patient;a blood pressure sensor that measures a blood pressure relative to atmospheric pressure at the wrist, the blood pressure sensor being carried by the wrist mounted unit to abut the wrist;a pressure compensation system that determines an offset pressure between the wrist and a heart level, the pressure compensation system including: a pressure sensor that senses pressure in a fluid filled cavity that extends from a level of the blood pressure sensor to the heart level relative to atmospheric pressure; anda component which combines the blood pressure measured relative to atmospheric pressure and the cavity fluid pressure relative to atmosphere to generate a measure of blood pressure at the heart level;a plurality of physiological parameter sensors configured to be mounted to the patient, each sensor sending sensed physiological parameters over a wireless body area network;a transceiver mounted in the wrist mounted unit and coupled to the wireless body area network to at least receive the sensed physiological parameters;a communication component mounted in the wrist mounted unit and connected with the transceiver and the pressure sensor to receive the sensed physiological parameters and the blood pressure at heart level therefrom, the communication component wirelessly transmitting the sensed physiological parameters and the blood pressure at heart level from the wrist mounted unit to a remote medical network. 2. The system as set forth in claim 1, the pressure sensor being disposed in contact with the fluid in the cavity, to measure a pressure within the cavity relative to atmospheric pressure. 3. The system as set forth in claim 2, wherein the pressure sensor is disposed at an end of the cavity that is positioned at one of: about the heart level, andabout blood pressure sensor level. 4. The system as set forth in claim 1, wherein the pressure sensor includes a solid state transducer, an amplifier, and an analog-to-digital converter which converts the heart level blood pressure to a digital blood pressure value. 5. The system as set forth in claim 1, wherein the fluid filled cavity includes a flexible tube. 6. The system as set forth in claim 1, further including: a call button attached to the wrist mounted unit and connected with the communication unit for transmitting a call for assistance to the medical network. 7. The system as set forth in claim 1, further including: a tracking unit disposed in the wrist mounted unit and connected with the communications unit which communicates patient location to the medical network. 8. The system as set forth in claim 1, further including: a component connected with the wrist unit which generates an electrical signal identifying the monitored patient. 9. A physiological measurement system comprising: a wrist unit which extends around a patient's wrist;a blood pressure sensor mounted ... | Koninklijke Philips Electronics N.V. | A61B005/02 A61B005/04
US09138171 | B2 | 2015-09-22 | System and method for detecting swallowing activity | 2005-06-17 | A system and method for detecting swallowing activity is provided. In an embodiment, a method includes receiving an electronic signal from an acceleronic signal from an accelerometer that represents swallowing activity, extracting at least two features from the signal, classifying the signal as a type of swallowing activity based on the extracted features, and generating an output of the classification. Exemplary activities include swallows, aspirations, movement and vocal artifacts. By indicating whether an activity is a swallow or an aspiration, the manner in which a patient afflicted with an increased likelihood for aspirations is fed can be adjusted to increase the likelihood of achieving a swallow instead of an aspiration during feeding. In turn this could reduce hospitalizations for aspiration pneumonia in patients with acute or chronic injury. | 1. A method for detecting swallowing activity including aspirations and swallows comprising the steps of: receiving an electronic acceleration signal generated by a sensor positioned on the throat of a patient, said signal representing swallowing activity;extracting features from said signal, wherein said features comprise stationarity, normality and dispersion ratio;wherein the extracting the features comprises determining a mean absolute deviation (“S1”) of the electronic acceleration signal according to the following equation:S1=1n⁢∑i=1n⁢xi-med⁡(x)determining an interquartile range (“S2”) of the electronic acceleration signal; andextracting the dispersion ratio according to the following equation:S1S2;applying a radial basis neural network based on the features for classifying said signal as representing an aspiration when an aspiration is detected and as representing a swallow when a swallow is detected; andgenerating an output representing one of a swallow and an aspiration. 2. The method of claim 1 wherein the features extracted from said signal consist of the stationarity, the normality and the dispersion ratio. 3. A method for detecting swallowing activity including aspirations and swallows comprising the steps of: receiving an electronic acceleration signal generated by a sensor positioned on the throat of a patient, said signal representing swallowing activity;extracting from said signal, wherein said features comprise stationarity, normality and dispersion ratio;applying a radial basis neural network for classifying said signal as representing an aspiration when an aspiration is detected and as representing a swallow when a swallow is detected; andgenerating an output representing one of a swallow and an aspiration;wherein said extracting the stationarity comprises:dividing said signal into a plurality of non-overlapping bins;determining a total number of total number of reverse arrangements, (ATotal) in a mean square sequence is determined;extracting said stationarity feature (z), determined according to the following equation:where:z=ATotal-μAσAwhere:μA is the mean number of reverse arrangements expected for a stationary signal of the same length, andσA is the standard deviation for an equal length stationary signal;wherein said extracting the dispersion ratio comprises determining a mean absolute deviation (“S1”) of the electronic acceleration signal according to the following equation:S1=1n⁢∑i=1n⁢xi-med⁡(x)determining an interquartile range (“S2”) of the electronic acceleration signal; andextracting the dispersion ratio according to the following equation:S1S2. 4. A method for detecting swallowing activity including aspirations and swallows comprising the steps of: receiving an electronic acceleration signal generated by a sensor positioned on the throat of a patient, said signal representing swallowing activity;extracting features from said signal, wherein said features comprise stationarity, normality and dispersion ratio;... | Holland Bloorview Kids Rehabilitation Hospital | A61B005/11 A61B005/00
US09907893 | B2 | 2018-03-06 | System of medical treatment units and peripheral devices | 2006-12-01 | The invention relates to a system of a plurality of medical treatment units that are respectively associated with a patient, and a plurality of peripheral appliances also respectively associated with a patient. The peripheral appliances transmit control signals and patient signals to the treatment units, while the treatment units receive the control signals and patient signals from the peripheral appliances. The present invention ensures that the treatment units do not receive patient signals from other peripheral appliances that are not associated with the particular patient. The operating staff allocate a peripheral appliance to a patient and confirm the successful association of the peripheral appliance with the treatment unit allocated to the patient. | 1. A system for matching multiple peripheral devices and medical treatment units with the correct patients comprising: a plurality of peripheral devices and a plurality of medical treatment units comprising: a first medical treatment unit and a first peripheral device being allocated to a first patient, and a second medical treatment unit and a second peripheral device being allocated to a second patient, wherein the first medical treatment unit is a first blood treatment apparatus comprising a first dialysis machine having a first extracorporeal blood circuit, and the second medical treatment unit is a second blood treatment apparatus comprising a second dialysis machine having a second extracorporeal blood circuit,wherein the first peripheral device comprises a first blood pressure monitor that is configured to monitor the first patient's blood pressure and is wirelessly connected to the first dialysis machine of the first medical treatment unit, and the second peripheral device comprises a second blood pressure monitor that is configured to monitor the second patient's blood pressure and is wirelessly connected to the second dialysis machine of the second medical treatment unit,each peripheral device comprising: a patient monitor having components configured to monitor a patient and configured to measure blood pressure of the patient;a control signal sending device configured to send control signals to the dialysis machine of the respective medical treatment unit,a patient signal sending device configured to send patient signals comprising data relating to the patient's blood pressure to the dialysis machine of the respective medical treatment unit,a first monitoring device configured to monitor an allocation, of the respective first or second blood pressure monitor, to the patient,a memory storage device for non-volatile storage of an identification signal corresponding to the blood pressure monitor of the peripheral device, anda log-on device configured to send a log-on signal and the identification signal from the blood pressure monitor of the peripheral device to the respective dialysis machine of the respective medical treatment unit, said log-on device cooperating with the control signal sending device and the first monitoring device; andeach medical treatment unit comprising: a control signal receiving device configured to receive control signals from the control signal sending device of the respective peripheral device,a patient signal receiving device configured to receive patient signals from the patient signal sending device of the respective peripheral device,a memory storage device for non-volatile storage of a reference signal corresponding to the dialysis machine of the medical treatment unit allocated to the blood pressure monitor of the respective peripheral device,an acknowledgement device configured to allow an operator to input the acknowledgement of an allocation of the respective blood pressure monitor of the respective p... | Fresenius Medical Care Deutschland GmbH | G06Q050/22 G06Q050/24 A61M001/16 A61M001/14 A61M001/34 A61M001/36
US09808574 | B2 | 2017-11-07 | Insertion device | 2008-06-25 | An insertion apparatus and a method for use with a device for delivery of a therapeutic fluid into a body of a patient and/or for sensing of a bodily analyte are disclosed. The apparatus includes a housing adapted for loading therein at least one cannula cartridge unit having a protective member. The protective member accommodates at least one penetrating cartridge having a subcutaneously insertable element and a penetrating member. The apparatus includes a displacement mechanism capable of protracting the penetrating cartridge towards the body of the patient, where protraction of the penetrating cartridge results in insertion of the subcutaneously insertable element into the body of the patient. | 1. A subcutaneous insertion apparatus comprising: a housing configured to receive at least one cartridge unit, the at least one cartridge unit comprising a protective member containing therein at least one subcutaneously insertable element and at least one penetrating member, the housing including: a first housing opening adapted for at least loading therethrough the at least one cartridge unit into the housing; anda second housing opening adapted for at least protracting therethrough the at least one insertable element and the at least one penetrating member; anda displacement mechanism contained within the housing and configured to at least protract the at least one insertable element and the at least one penetrating member through the second housing opening toward the body of the patient for subcutaneous insertion of the at least one insertable element,wherein the displacement mechanism is configured to be primed for activation prior to the housing receiving the at least one cartridge unit; andthe insertion apparatus is configured to couple to a skin-securable cradle that comprises a tubular protrusion emerging upwardly from the cradle that defines therein a well, the well comprising a plurality of penetration tunnels for directing the subcutaneous insertable element toward a plurality of angles with respect to a plane of a bottom housing opening, and the insertion apparatus is configured to align the subcutaneous insertable element with any of the plurality of penetration tunnels of the well, so as to insert the subcutaneous insertable element through the aligned tunnel. 2. The insertion apparatus according to claim 1, wherein the displacement mechanism is further configured to retract the penetrating member into the protective member subsequent to the subcutaneous insertion of the insertable element. 3. The insertion apparatus according to claim 1, wherein the displacement mechanism comprises a rod for manually forcing the at least one insertable element and the at least one penetrating member toward the body of the patient. 4. The insertion apparatus according to claim 1, wherein the displacement mechanism includes a loadable spring, wherein unloading of the spring results in protraction of the at least one insertable element and the at least one penetrating member toward the body of the patient. 5. The insertion apparatus according to claim 4, wherein the displacement mechanism further comprises a second spring, wherein unloading of the second spring results in retraction of the penetrating member into the protective member subsequent to subcutaneous insertion of the insertable element. 6. The insertion apparatus according to claim 1, further comprising a user-actuated button for actuating the displacement mechanism when the cartridge unit is loaded therein. 7. The insertion apparatus according to claim 1, further comprising a safety means for preventing inadvertent actuation of the displacement mechanism. 8. The insertion apparatus... | Roche Diabetes Care, Inc. | A61M005/142 A61B005/00 A61M005/158 A61M005/42 A61B005/145 A61M005/14
US20080243012 | A1 | 2008-10-02 | Method of compressing electrocardiogram data and electrocardiogram telemetry system using the same | 2007-03-29 | In an electrocardiogram telemeter, a data acquirer is operable to acquire electrocardiogram data. A data compressor is operable to compress the electrocardiogram data with either a wavelet transform, a Huffman coding, or an arithmetic coding, thereby generating compressed electrocardiogram data adapted to be transmitted to a remote receiver which is configured to reconstruct the electrocardiogram data. | 1. A method of compressing an electrocardiogram data, comprising: acquiring the electrocardiogram data;subjecting the acquired electrocardiogram data to a frequency filtering; andcompressing the electrocardiogram data, which has been subjected to the frequency filtering, with either a wavelet transform, a Huffman coding, or an arithmetic coding. 2. The method as set forth in claim 1, wherein: the wavelet transform includes dividing the electrocardiogram data into a plurality of data pieces which are respectively associated with different frequency ranges; andthe wavelet transform includes at least one of: deleting at least one of the divided pieces;deleting at least a part of at least one of the divided pieces, in which noises having a wave height less than a prescribed value are superposed; anddeleting at least a part of at least one of the divided pieces, in which noises having a wave height no less than the prescribed value are superposed. 3. An electrocardiogram telemeter, comprising: a data acquirer, operable to acquire electrocardiogram data;a filter, operable to subject the electrocardiogram data acquired by the data acquirer to a frequency filtering; anda data compressor, operable to compress the electrocardiogram data, which has been subjected to the frequency filtering, with either a wavelet transform, a Huffman coding, or an arithmetic coding, thereby generating compressed electrocardiogram data adapted to be transmitted to a remote receiver which is configured to reconstruct the electrocardiogram data. 4. The electrocardiogram telemeter as set forth in claim 3, wherein: the data compressor is operable to divide the electrocardiogram data into a plurality of data pieces which are respectively associated with different frequency ranges; andthe data compressor is operable to perform the wavelet transform by at least one of: deleting at least one of the divided pieces;deleting at least a part of at least one of the divided pieces, in which noises having a wave height less than a prescribed value are superposed; anddeleting at least a part of at least one of the divided pieces, in which noises having a wave height no less than the prescribed value are superposed. | NIHON KOHDEN CORPORATION | A61B005/0402
US20080243023 | A1 | 2008-10-02 | Device For Preventing Bruxism | 2005-06-22 | A device for preventing bruxism includes a carrier (1) which is intended for receiving in a mouth of a user and which includes at least a part of an electronic bio-feedback system. The bio-feedback system includes at least one signal sensor (2) which is sensitive to electrical muscle signals and which is coupled to a control unit. The control unit is designed and adapted to generate a correcting stimulus via an electrode (2) when a muscle signal is detected above a predetermined threshold level. The bio-feedback system is herein provided with elements for individually adjusting and storing the threshold level. | 1. Device for preventing bruxism, comprising a carrier intended for receiving in a mouth of a user, which carrier comprises at least a part of an electronic bio-feedback system, which bio-feedback system comprises at least one signal sensor which is sensitive to electrical muscle signals and which is coupled to a control unit designed and adapted to generate a correcting stimulus via an electrode when a muscle signal is detected above a predetermined threshold level, characterized in that the bio-feedback system is provided with means for adjusting and storing the threshold level. 2. Device as claimed in claim 1, characterized in that the threshold value is adjustable to a level of between about 5 millivolts and 50 millivolts. 3. Device as claimed in claim 1, characterized in that the bio-feedback system is provided with means for placing the device in a learning mode in which an output signal from the signal sensor is stored as threshold signal. 4. Device as claimed in claim 1, characterized in that the at least one signal sensor is able and adapted to detect a first muscle signal in a first frequency domain and a second muscle signal in a second frequency domain, and that the bio-feedback system comprises a threshold value per frequency domain. 5. Device as claimed in claim 4, characterized in that the first and second frequency domain lie respectively in a frequency range of about 1 to 20 Hz and about 90 to 600 Hz. 6. Device as claimed in claim 1, characterized in that the control unit comprises a processing unit which is able and adapted to determine a running average of a muscle signal measured during a rest state and to store this average as offset value. 7. Device as claimed in claim 6, characterized in that the processing unit is able and adapted to decrease a measured muscle signal by the offset value in an operating state and present this as usable signal to the control unit. 8. Device as claimed in claim 1, characterized in that the control unit comprises clock means for generating a control signal only after a determined time period has elapsed during which the input signal lies above said threshold value. 9. Device as claimed in claim 8, characterized in that said time period is adjustable between 1 and 5 seconds. 10. Device as claimed in claim 8, characterized in that the control unit comprises adding means which are able and adapted to find the sum of time periods during which said threshold value has been exceeded while subtracting intervening periods during which said threshold value has not been reached. 11. Device as claimed in claim 1, characterized in that the carrier is provided with an electronic element which is able to comprise an electronically detectable identification value associated with the carrier, and that the control device comprises read means which are able to record the identification value of a carrier coupled thereto and compare it to a value stored in the control device. 12. Device as claimed in ... | - | A61B005/04
US20080306346 | A1 | 2008-12-11 | Diagnostic tool detecting the degradation status of Von Willebrand Factor multimers | 2007-06-07 | A method in which the cleavage profile and size distribution of von Willebrand factor (VWF) multimers is analyzed, includes: providing a sample medium of human body fluids comprising a plurality of VWF multimers of different size; enrichment or purification of the VWF multimers by cryoprecipitation or chromatography to obtain a separated preparation of the VWF multimers from said sample medium; exposing the separated preparation of VWF multimers to a light source to produce signals obtained by vibrational spectroscopy; detecting said signals; transformation by mathematical alogrithms; generation of patterns based on computing of data of original resonance spectra and determining the cleavage profile and the size distribution of said separated VWF multimers by chemometrics; and acquisition of a databank obtained from healthy individuals for identifying subjects at risk of developing at least one of the following diseases: sepsis, coagulopathy, thrombotic disease, infection, and inflammation. | 1. A method of analyzing the cleavage profile and size distribution of von Willebrand factor (VWF) multimers, comprising: providing a sample medium or human body fluids comprising a plurality of VWF multimers of different size; enrichment or purification of VWF by cryoprecipitation or chromatography from said sample medium; exposing said separated preparation of VWF multimers to a light source to produce signals obtained by vibrational spectroscopy; detecting said signals; transformation by mathematical algorithms; generation of patterns based on computing of data of original resonance spectra and determining the cleavage profile and the size distribution of said separated VWF multimers by chemometrics and acquisition of a databank obtained from healthy individuals identifying subjects at risk of developing thrombotic disease and/or identifying subjects at risk of developing infection, inflammation in particular sepsis and/or coagulopathy. 2. The method of claim 1 wherein analysis of the cleavage profile and size distribution of von Willebrand factor (VWF) multimers is used for therapeutically monitoring and decision making for these subjects. 3. The method of claim 1 wherein analysis of the cleavage profile and size distribution of von Willebrand factor (VWF) multimers is used for screening of subjects at risk of developing thrombotic disease, infection, inflammation in particular sepsis and/or coagulopathy. 4. The methods of claim 1 wherein analysis of the cleavage profile and size distribution of von Willebrand Factor (VWF) multimers from said samples is used for comparing spectra corresponding to a set of VWF abnormalities and/or disease for diagnostic approaches. 5. The methods of claim 1 wherein analysis of the cleavage profile and size distribution of von Willebrand Factor (VWF) multimers from said samples is used for comparing spectra detecting the VWF-degrading activity of proteolytic enzymes, in particular of the ADAMTS13 protease. 6. The methods of claims 1-5 wherein analysis of the cleavage profile and size distribution of von Willebrand Factor (VWF) multimers from said samples is used for early diagnosis and/or differential diagnosis and/or monitoring of the disease. 7. Methods of claims 1-6 wherein analysis of the cleavage profile and size distribution of von Willebrand Factor (VWF) multimers from said samples is used for correlational analysis with non-spectroscopic data. 8. Methods of claims 1-7 wherein the plurality of signals of said sample is analysed by the use of non-supervised classification analysis, in particular comprising hierarchical clustering and principal component analysis. 9. Methods of claims 1-7 wherein the plurality of signals of said samples is analysed by the use of supervised classification analysis, in particular comprising K-nearest neighbour analysis, nearest mean classifier, linear discrimination analysis, artificial neural networks, as well as support vector machines. 10. Methods of claims 1-9 w... | - | A61B005/00 G01N033/48
US20080260229 | A1 | 2008-10-23 | System and method for segmenting structures in a series of images using non-iodine based contrast material | 2007-05-25 | A method and system of defining a boundary of a part of a blood vessel in an image captured by an ex vivo imager, where such part of the blood vessel is free from iodine-based contrast material. | 1. A method comprising defining a boundary of a part of a blood vessel in an image of a series of images, said series of images captured by an ex vivo imager, said part of said vessel in said image being free of iodine-based contrast material. 2. The method as in claim 1, comprising designating a seed area in said image. 3. The method as in claim 2, wherein said designating comprises marking an area of said vessel at which to stop a segmentation of said vessel. 4. The method as in claim 2, comprising clustering pixels in a portion of said image containing said seed area, said portion comprising less than all of said image. 5. The method as in claim 1, comprising designating non-uniform ranges of image intensity levels into which pixels are clustered. 6. The method as in claim 1, wherein said defining comprises: identifying said blood vessel; andcomparing a grayscale scoring of an area in said image to a plurality of stored grayscale scores of samples of said blood vessel. 7. The method as in claim 1, comprising defining a boundary of a plurality of vessels in said image. 8. The method as in claim 1, comprising: a first clustering of a first plurality of pixels in said image into a first set of clusters; anda second clustering of a second plurality of pixels into a second set of clusters, said second plurality of pixels being a subset of said first plurality of said pixels. 9. The method as in claim 8, wherein said second clustering comprises selecting a plurality of cluster center means of said second set of clusters, said selection based on an intensity of a pixel in a seed area. 10. The method as in claim 8, wherein said second clustering of said second plurality of pixels into a second set of clusters comprises clustering said second plurality of pixels into more than four clusters. 11. The method as in claim 1, comprising: a first clustering of a first plurality of pixels in said image into a first set of clusters; anda second clustering of less than all of said first plurality of pixels into a second set of clusters. 12. The method as in claim 1, comprising: mapping an isolable contour region of a cluster of pixels, said pixels in said cluster having a range of image intensity ranges, andselecting from among a plurality of said isolable contour regions having pixels in said range of image intensity levels, a region including a pixel overlapping a pixel in said seed area. 13. The method as in claim 2, comprising: recording a coordinate of a pixel within said seed area,an image intensity of said pixel; anda designation of said image within said series of images; anddefining a boundary around said area around said pixel, said boundary around said area selected from the group consisting of a boundary box and a convex hole. 14. The method as in claim 1, wherein defining said boundary of said blood vessel comprises defining a boundary of an outer wall of said blood vessel. 15. The method as in claim 1, wherein said defining comprises... | - | A61B005/00
US20080249385 | A1 | 2008-10-09 | ISOLATED INTRAVENOUS ANALYTE MONITORING SYSTEM | 2007-04-04 | A continuous intravenous analyte monitoring system includes an amperometric biosensor detecting an analyte concentration in the blood, a controller receiving a signal from the biosensor and computing the concentration, and an isolation device isolating the biosensor from EMI. A CPU may be coupled to the controller via the isolation device for continuous output of sensed concentration to a display unit. Isolated circuits may include a temperature sensor transmitting biosensor temperature to the controller for correction of the computed concentration, a multiplexer combining biosensor and temperature sensor signals, and an A/D converter converting multiplexed input to the controller. The biosensor may be a multi-electrode sensor having a working electrode immobilizing an glucose oxidase enzyme to detect blood glucose concentration. The biosensor and temperature sensor may be located in vivo using a catheter for continuous monitoring. | 1. An analyte monitoring system, comprising: a biosensor sensing an analyte concentration in vivo and outputting a signal corresponding to the analyte concentration;a controller receiving the signal and computing the analyte concentration therefrom; andan isolation means coupled to the controller and isolating the biosensor from electrical and magnetic noise. 2. The analyte monitoring system of claim 1 further comprising a CPU coupled to the controller via the isolation means and receiving the computed analyte concentration for output to a display. 3. The analyte monitoring system of claim 2 further comprising an isolated circuit and a non-isolated circuit, the isolation means separating the isolated circuit from the non-isolated circuit, the isolated circuit including the biosensor and the controller, and the non-isolated circuit including the CPU. 4. The analyte monitoring system of claim 3 wherein the isolation means provides an isolation space between the isolated circuit and the non-isolated circuit between about 0.3 in. and about 1.0 in. 5. The analyte monitoring system of claim 4 wherein the isolation means comprises an optical isolator. 6. The analyte monitoring system of claim 4 wherein the isolation means comprises a DC/DC converter. 7. The analyte monitoring system of claim 2 further comprising a display unit coupled to the CPU, the isolation means electrically and magnetically isolating the display unit from the biosensor. 8. The analyte monitoring system of claim 7 wherein the CPU continually updates the computed analyte concentration for output to the display unit. 9. The analyte monitoring system of claim 1 wherein the biosensor comprises a four-electrode sensor. 10. The analyte monitoring system of claim 9 wherein the four-electrode sensor includes at least one enzyme electrode. 11. The analyte monitoring system of claim 10 wherein the enzyme electrode immobilizes glucose oxidase. 12. The analyte monitoring system of claim 3 further comprising a temperature sensor in the isolated circuit coupled to the controller. 13. The analyte monitoring system of claim 12 wherein the controller corrects the computed analyte concentration according to temperature sensed by the temperature sensor. 14. The analyte monitoring system of claim 13 wherein the sensed temperature represents temperature of the biosensor. 15. The analyte monitoring system of claim 12 further comprising a multiplexer in the isolated circuit, the multiplexer receiving the signal from the biosensor and a second signal from the temperature sensor and transmitting them as multiplexed output to the controller. 16. The analyte monitoring system of claim 15 further comprising an analog-to-digital converter in the isolated circuit, converting analog multiplexed output to digital data for the controller. 17. An isolated intravenous analyte monitoring system, comprising: an isolated sensing circuit including an in vivo biosensor sensing an analyte level in blood and... | - | A61B005/05
US20080275367 | A1 | 2008-11-06 | Systems and methods for mapping intra-body tissue compliance | 2008-04-23 | Robotic instrument systems and methods for generating a geometric map of an area of body tissue which is correlated with a tissue characteristic such as tissue compliance or related property. The system comprises a robotically controlled catheter which is controlled by a robotic instrument driver. A force sensor system is provided generates force signals responsive to a force applied to the distal end of the catheter. A position determination system is also provided which generates position signals responsive to the location of the distal end of the catheter. A computer is configured to receive and process the force signals and position signals to generate a geometric map of an area of body tissue correlated to the tissue compliance of different regions of the body tissue or a tissue characteristic determinable from the tissue compliance. | 1. A robotically controlled instrument system, comprising: an elongate instrument having a distal end surface adapted to make contact with an internal body tissue surface;a force sensor associated with the instrument, wherein the force sensor generates force signals responsive to a force applied to the distal end surface of the instrument upon contacting a tissue surface, the force signals corresponding to a respective compliance of tissue being contacted by the instrument distal end surface;a position determining system which generates position data indicative of a position of the distal end of the instrument;a processor operatively coupled to the force sensor and position determining system, and configured to process the respective force signals and position signals to generate a geometric map of an area of body tissue correlated with a tissue characteristic related to tissue compliance. 2. The system of claim 1, wherein the tissue characteristic is tissue compliance. 3. The system of claim 1, wherein the instrument comprises a robotically controlled guide instrument. 4. The system of claim 3, wherein the force sensor is coupled to a distal end portion of the guide instrument. 5. The system of claim 1, wherein the instrument comprises a catheter carried in a robotically controlled guide instrument. 6. They system of claim 5, wherein the force sensor is coupled at a proximal end portion of the catheter. 7. The system of claim 1, further comprising a display coupled to the processor, wherein the processor causes regions of the body tissue area in the map having differences in tissue compliance to be visually highlighted on the display. 8. The system of claim 7, wherein the body tissue regions having differences in tissue compliance to be visually highlighted using different colors or shades of colors on the map. 9. The system of claim 1, wherein the position determining system comprises at least one position sensor disposed on the instrument. 10. A method of mapping a relative tissue compliance of an area of body tissue, comprising: robotically maneuvering an elongate instrument within an internal body space;advancing a distal end surface of the instrument into contact with each of a plurality of locations on the area of body tissue;sensing the compliance of the tissue at each of the plurality of locations on a tissue surface;determining a position of each of the plurality of tissue surface locations;generating a geometric map of the area of tissue representative of a characteristic of the tissue related to a compliance of the tissue. 11. The method of claim 10, wherein the tissue characteristic is tissue compliance. 12. The method of claim 10, wherein tissue compliance is measured using a force sensor coupled to the instrument. 13. The method of claim 12, wherein the force sensor is coupled to a proximal portion of the instrument. 14. The method of claim 12, wherein the force sensor is coupled to a distal portion of the instrument. | Hansen Medical, Inc | A61B005/00 A61B017/00
US20080287748 | A1 | 2008-11-20 | System and method for physiological data readings, transmission and presentation | 2008-02-05 | Systems, methods and computer program products for a non-invasive, automated method of receiving, collecting, storing and transmitting a person's physiological data and activity levels for the purposes of determining the well-being of a person are disclosed. Aspects of the present invention are also directed to making additional health status determinations based on the historical information and trends of the collected data. Such aspects of the present invention readily lend themselves to incremental component and functionality modifications, which would allow for any number of physiological data to be collected through the introduction of any number of non-invasive sensors. This results in higher accuracy, reliability, and utility of the collected information, further solidifying the uniqueness and desirability of the invention, for consumer and clinical Wireless Body Area Networks (WBAN) applications. | 1. A system for facilitating the receipt, collection, analysis and transmission of physiological data comprising: a. at least one physiological sensor capable of being attached to a user utilizing the system;b. at least one computer system, in operative communication with said at least one physiological sensor and wherein said at least one computer system contains data aggregation and transmission software;c. a short range transceiver, in operative communication with said at least one computer system;d. a long range transceiver, in operative communication with said at least one computer system;e. a mechanical device, in operative communication with said at least one computer system, capable of determining the orientation and physical activity of the user; andf. a position locator for locating said user. 2. The system of claim 1, wherein the mechanical device is one of the following devices: an accelerometer and an inclinometer. 3. The system of claim 1, further comprising a clock for synchronizing all the devices within the system. 4. The system of claim 1, further comprising an antenna in operative communication with said short-range transceiver and said long-range transceiver. 5. The system of claim 1, further comprising a microphone, wherein said microphone enables duplex voice communication with a remote location. 6. The system of claim 1, wherein the system is in wireless communication with a pre-defined remote location. 7. The system of claim 6, wherein said communication with said pre-defined remote location is via a wireless communications protocol selected from a group consisting of Cellular, ZigBee, Wireless (802.11a/b/g/n), Wi-Fi, ANT, Bluetooth, Ultra wide Band, and custom application protocol. 8. A method for receiving, collecting, storing and transmitting human physiological data using a system comprising: a. receiving physiological data indicative of the health status of a user utilizing a system comprising at least one physiological sensor, wherein said at least one sensor is capable of being attached to the body of said user;b. analyzing the physiological data received from said at least one physiological sensor by utilizing at least one computer system, wherein said at least one computer system is in operative communication with said at least one physiological sensor;c. storing the received physiological data on said at least one computer system;d. processing the physiological data, by said at least one computer system, in preparation for further action on said physiological data; ande. transmitting, utilizing a long range transceiver, the physiological data to a pre-defined remote location. 9. The method of claim 8, wherein step (d) comprises compressing the physiological data. 10. The method of claim 8, wherein step (d) comprises encrypting the physiological data. 11. The method of claim 8 further comprising: sensing orientation and physical activity of said user;generating data regarding the orientation and physical... | - | A61B005/00 G06Q050/00
US20080287767 | A1 | 2008-11-20 | Biometric Sensor | 2006-11-21 | A bio metric sensor device (1) comprises: —a thin, flexible, layered sensor body (2); —a conductive sense plate (21); —a first non-conductive layer (41) between the sense plate and an outer surface (7); —a conductive shield plate (53) having a passage opening (54), overlaying the sense plate; —a second non-conductive layer (51) between the sense plate and the shield plate; —conductive circuit lines (73) on an inner surface (6); —a non-conductive separation layer (61, 71) between the shield plate and the circuit lines; —a signal processing circuit (100) mounted on the inner surface (6), the circuit (100) comprising a differential amplifier (110) having an input (111); —a conductive interconnector (82) crossing the second non-conductive layer (51) and the separation layer (61, 71), extending through the passage opening (54) of the shield plate (53), coupling the sense plate (21) and said input (111) of said amplifier (110). | 1. Biometric sensor device (1), suitable for capacitively sensing bio-electrical signals, the device comprising: a thin, flexible, layered sensor body (2) having an inner surface (6) and an outer surface (7) opposite the outer surface, the body (2) comprising a first body part (3) which comprises:an electrically conductive sense plate (21) in a first conductive layer (42);a first non-conductive layer (41) between the sense plate (21) and the outer surface (7);an electrically conductive shield plate (53) in a second conductive layer (52), overlaying the sense plate (21) at the side opposite the outer surface, and having a size at least corresponding to the size of the sense plate and preferably projecting beyond the contours of the sense plate;the shield plate (53) having a passage opening (54);a second non-conductive layer (51) between the sense plate (21) and the shield plate (53);electrically conductive circuit lines (73, 74, 75, 125) in a conductive circuit layer (72) on the inner surface (6);a non-conductive separation layer (61, 71) between the shield plate (53) and the electrically conductive circuit lines (73, 74, 75);the device further comprising:an electronic signal processing circuit (100) having circuit components (17) mounted on the inner surface (6), the circuit (100) comprising at least one differential amplifier (110) having a first input (111) and an output (114);a first electrically conductive interconnector (82) crossing the second non-conductive layer (51) and the separation layer (61, 71), extending through the passage opening (54) of the shield plate (53), for coupling the sense plate (21) to the first input (111) of said amplifier (110). 2. Device according to claim 1, wherein each non-conductive layer is made of capton. 3. Device according to claim 1, wherein the sensor body (2) is implemented as a stack of flexfoil layers (40; 50; 60; 70) attached to each other, each flexfoil layer comprising the combination of at least one conductive layer (42; 52; 62; 72) and at least one non-conductive layer (41; 51; 61; 71). 4. Device according to claim 1, further comprising an electrically conductive shield ring (22) around the sense plate (21) in the first conductive layer (42), the shield ring (22) being electrically connected to the shield plate (53). 5. Device according to claim 4, further comprising a series of second electrically conductive interconnectors (81) crossing the second non-conductive layer (51), each second interconnector (81) contacting the shield ring (22) and the shield plate (53). 6. (canceled) 7. Device according to claim 1, wherein the shield plate (53) is electrically connected to the output (114) of the amplifier (110). 8. Device according to claim 7, wherein the amplifier's output (114) is connected to the amplifier's inverting input (112), and wherein the sense plate (21) is connected to the amplifier's non-inverting input (111). 9. Device according to claim 7, further comprising a third electricall... | - | A61B005/0408
US20080287774 | A1 | 2008-11-20 | Magnetic Resonance Imaging and Spectroscopy Means and Methods Thereof | 2006-11-02 | The present invention discloses neurochemical agents and biochemical agents for human or mammalian neuro- and body-metabolic imaging, comprising chemicals involved in neuronal or glial function, neuromodulatory processes in the brain of said human or mammalian, vascular function, or organ specific metabolic processes; said neurochemical and biochemical agents are labeled with stable isotopes selected from a group including carbon-13, nitrogen-15, deuterium, fluorine-19 or a combination thereof in predetermined positions, so as to enhance the detectability of the agents and their metabolic successors. | 1-40. (canceled) 41. High T1 neurochemical and biochemical contrast agents (HTNCs) for imaging metabolic processes and activities in the brain or body of either human or otherwise mammalian (patient); said HTNCs comprising chemicals involved in neuronal or brain function or neuromodulatory processes in the brain of said patient, vascular function, or organ specific metabolic processes; said HTNCs are labeled with stable isotopes selected from a group consisting of deuterium, carbon-13, nitrogen-15, fluorine-19 (2H, 13C, 15N, 19F) or a combination thereof in predetermined positions, so as to enhance the detectability of both the agents and their metabolic successors. 42. The HTNCs of claim 41, selected from a group consisting of: a. molecules of metabolic processes selected from a group consisting of choline, betaine, acetylcholine, acetate, aspartate, N-acetylaspartate, creatine, L-tyrosine, L-DOPA, dopamine, norepinephrine, epinephrine, vanillylmandelic acid (VMA), homovanillic acid (HVA), 3-O-methyldopamine, 3-O-methylnorepinephrine, 3-O-methylepinephrine, dopaquinone, L-tryptophan, 5-hydroxy-tryptophan, serotonin, 5-hydroxyindole acetaldehyde, 5-hydroxyindole acetic acid, melatonin, glutamate, arginine, citrulline, N-acetylcitrulline, argininosuccinate, kynurenic acid (KYNA), 7-chlorokynurenic acid (7-Cl—KYNA), kynurenine, 4-chlorokynurenine, pharmacologically acceptable salts thereof, or any combination thereof;b. molecules used in drugs selected from a group consisting of psychiatric or neuroprotective drugs, blood flow modulating drugs, mood altering drugs; with and drugs selected from a group consisting of rivastigmine, rasagiline, methylphenidate, amphetamine, tacrine, donepezil, metrifonate, fluoxetine, sertraline, paroxetine, fluvoxamine, citalopram, escitalopram, venlafaxine, nefazodone, mirtazapine, bupropion, cianopramine, femoxetine, ifoxetine, milnacipran, oxaprotiline, sibutramine, viqualine, clozapine, fenclonine, dexfenfluramine, chlorpromazine, methamphetamine, prazosin, terazosin, doxazosin, trimazosin, labetalol, medroxalol, tofenacin, trazodone, viloxazine, riluzole, pharmacologically acceptable salts thereof, or any combinations thereof being preferable;c. molecules used as either PET or SPECT contrast agents; with molecules selected from a group consisting of ligands for dopamine receptors and transporters, serotonin receptors and transporters, acetylcholine receptors and transporters, norepinephrine receptors and transporters, beta-amyloid peptide and its imidazopyridinylbenzeneamine and benzothizolylbenzeneamine derivatives ligands, pharmacologically acceptable salts thereof, or any combinations thereof being preferable; and,d. molecules that upon hydrogenation yield said HTNCs; with molecules selected from a group consisting of (2-hydroxyethenyl)trimethylammonium, (2-hydroxyethynyl) trimethylammonium, (S)-2-amino-3-(5-hydroxy-1H-indol-3-yl)propenoic acid, (S)-2-amino-3-(3,4-dihydroxyphenyl)propenoic acid, 2-amino-2-en... | - | A61B005/055
US20080243020 | A1 | 2008-10-02 | Physiological Signal Collecting And Monitoring Device And System | 2006-10-08 | A physiological signal collecting and monitoring system includes a physiological signal sensing device (10) and a computer device (48). The sensing device (10) includes a physiological signal sensing unit for detecting physiological signals from a subject, a circuit arrangement including a processor for executing a preloaded program to receive the physiological signals from the sensing unit, process thereof and store resultant physiological information into a memory, and a carrier (14) for carrying the sensing unit and the circuit arrangement to attach to the subject. The computer device (48) having a standard computer communication port is used for further processing the physiological information. Furthermore, the sensing device (10) further includes a communication interface to communicate with the standard computer communication port, and the interface to communicate with the processor and memory are integrated into a removable module (16), and the removable module (16) is detachable from the sensing device for acting as a computer dongle so that the removable module (16) is capable of electrically connecting with the computer device (48) to download the physiological information stored in the removable module (16) into the computer device (48). | 1. A physiological signal collecting and monitoring system, comprising: a physiological signal sensing device, comprising: an airflow physiological signal sensing unit for detecting airflow physiological signals from a subject;a circuit arrangement, comprising a processor and a memory, wherein the processor executes a preloaded program to receive the physiological signals from the physiological signal sensing unit, process thereof and store resultant physiological information into the memory; anda flexible carrier for carrying the airflow physiological signal sensing unit and the circuit arrangement to detect the airflow physiological signal at the nostrils and the mouth of the subject; anda computer device, having a standard computer communication port, for further processing the physiological information,whereinthe sensing device further comprises a communication interface to communicate with the standard computer communication port, and the communication interface together with the processor and the memory are integrated into a removable module; andthe removable module is detachable from the sensing device for acting as a computer dongle so that the removable module is capable of electrically connecting with the computer device to download the physiological information stored in the removable module into the computer device. 2. A physiological signal collecting and monitoring system as claimed in claim 1, wherein the airflow physiological signal sensing unit is an airflow pressure sensor or a thermal sensor. 3. A physiological signal collecting and monitoring system as claimed in claim 2, wherein the physiological signal sensing device further comprises a sensor/electrode connecting port for connecting with an external sensor/electrode, and said external sensor/electrode is selected from a group consisting of a respiratory sensor, a blood oxygenation sensor, a snoring sensor, an ECG electrode, an EMG electrode, an EOG electrode and an EEG electrode. 4. A physiological signal collecting and monitoring system as claimed in claim 1, wherein the removable module is separated from the physiological signal sensing device by setting a connector therebetween, tearing, cutting or other separating methods. 5. A physiological signal collecting and monitoring system as claimed in claim 1, wherein the circuit arrangement further comprises a signal processing circuit having functions of filtering and amplifying for processing the detected physiological signals, and/or the circuit arrangement further comprises an A/D (analog-to-digital) converter for converting the physiological signals into a digital form. 6. A physiological signal collecting and monitoring system as claimed in claim 1, wherein the circuit arrangement further comprises a RF module for executing a wireless input and output. 7. A physiological signal collecting and monitoring system as claimed in claim 6, wherein the RF module is implemented to wirelessly output an alarm signal when the... | - | A61B005/08
US20080275365 | A1 | 2008-11-06 | Methods of Transferring Data to a Medical Test Device | 2007-05-04 | Disclosed are various preferred embodiments for dynamic transfer of information from a test sensor to an analyte medical test device. Exemplary embodiments include various containers, systems and methods. | 1. A method of transmitting data, the method comprising: providing discrete surface features on a surface of a container indicative of predetermined parameters for at least one test strip disposed in the housing;inserting the container into a test port of the test device;removing the container out of the test port; andreading the discrete surface features as the container is moving relative to the test port during one of the removing and inserting steps to provide parameters specific to the test strip. 2. The method of claim 1, wherein the reading comprises recognizing the surface features during the inserting. 3. The method of claim 2, wherein the reading comprises recognizing the surface features during the removing. 4. The method of claim 1, wherein the reading comprises decoding data encoded by the surface features during the inserting and removing: comparing data during the removing with data during the inserting; andoutputting a signal indicative of any difference in the data during the removing and data during the inserting. 5. The method of claim 1, further comprising retaining the test strip to the sampling port subsequent to separation of the container from the test device. 6. The method of claim 1, wherein the surface features comprise; a first plurality of discrete surface features disposed on the outer surface to define a first data line; anda second plurality of discrete surface features disposed on the outer surface in a repeating sequence to define a clock line. 7. The method of claim 1, wherein the reading comprises recognizing transitions between the second plurality of discrete surface features to coordinate reading of the first data line. 8. The method of claim 7, wherein the reading comprises correlating the transitions of the clock line to transitions between the first plurality of discrete surface features; andproviding binary data from the correlating. 9. The method of claim 7, wherein the reading further comprises recognizing from the second plurality of discrete surfaces a predetermined geographic region for the at least one test strip. 10. The method of claim 7, wherein the first plurality of discrete surface features includes a first plurality of low and high reflectance areas disposed on the outer surface to define a first data line; and the second plurality of discrete surface features includes a second plurality of low and high reflectance areas disposed on the outer surface in a repeating sequence along a second perimeter to define timing intervals for the first data line. 11. The method of claim 10, wherein the data lines further comprise information selected from a group consisting essentially of calibration parameters, calibration code, geographical data, error checking data, lot number, manufacturing date, expiration date, batch number, manufacturer's name, chemical composition, ingredients, and combinations thereof 12. The method of claim 8, wherein each of the first and second plurality of discrete... | - | A61B005/00 C12M001/00 C12M003/00 C12Q001/00 C12Q001/54 G01N001/12
US20080255430 | A1 | 2008-10-16 | PORTABLE DEVICE WITH BIOMETRIC SENSOR ARRANGEMENT | 2007-04-16 | The present invention is directed to a biometric sensor system 100 for measuring one or more physical or behavioral characteristic of a user, which system comprises: an information presenting unit 20 provided with a first sound presenting unit 22 and a second sound presenting unit 24 for operatively presenting sounds to a user, a portable device 10 provided with a sound producing unit 11, 40, 44, 46 for producing sounds to said sound presenting units 22, 24, and a biometric control unit 42 for controlling said measurement. The information presenting unit 20 comprises at least a first biometric sensor unit 32 and a second biometric sensor unit 32′ arranged to operatively enable a differential measurement of one or more physical or behavioral characteristic of a user, and the biometric control unit 42 is arranged to operatively control the biometric sensor units 32, 32′ so as to provide a differential measurement. | 1-32. (canceled) 33. A biometric sensor system, comprising: an information presenting unit provided with a first sound presenting unit and a second sound presenting unit to operatively present sounds to a user, wherein the information presenting unit includes a first biometric sensor unit and a second biometric sensor unit to provide signals to generate a differential measurement of one or more physical or behavioral characteristics of a user of the biometric sensor system;a portable device provided with a sound producing unit to provide sounds to the first and second sound presenting units, the portable device including a third biometric sensor unit to obtain another measurement of the one or more physical or behavioral characteristics of the user; anda biometric control unit to provide an average measurement of the one or more physical or behavioral characteristics of the user based on the differential measurement and the other measurement. 34. The biometric sensor system of claim 33, wherein the biometric control unit is completely or at least partly arranged in the information presenting unit. 35. (canceled) 36. The biometric sensor system of claim 33, further comprising a fastening arrangement to attach at least a part of the information presenting unit to a body of the user. 37. The biometric sensor system of claim 33, wherein each of the sound presenting units operatively attaches to an ear of the user. 38. The biometric sensor system of claim 33, wherein the first biometric sensor unit is attached to or included in the first sound presenting unit, and the second biometric sensor unit is attached to or included in the second sound presenting unit. 39. The biometric sensor system of claim 33, wherein at least one of the biometric sensor units operatively senses a heart pulse, an EEG, an EKG, or a body temperature of the user. 40. The biometric sensor system of claim 39, wherein the biometric control unit operatively presents the sensed heart pulse in the form of audible pulse beats via the information presentation unit. 41. The biometric sensor system of claim 40, wherein the biometric control unit and the sound producing unit operatively present information to the user simultaneously via the information presentation unit. 42. The biometric sensor system of claim 33, wherein the portable device is a cellular phone that operatively connects to the information presenting unit via a wired or a wireless connection. 43. The biometric sensor system of claim 33, wherein the information presenting unit is an earphone that operatively connects to the portable device via a wired or a wireless connection. 44. The biometric sensor system of claim 33, further comprising another information presentation unit in the form of a display to present visible information to the user associated with the characteristic of the user. 45-51. (canceled) 52. A method of measuring a physical or behavioral characteristic of a user using a biometric sensor s... | SONY ERICSSON MOBILE COMMUNICATIONS AB | A61B005/117
US07604629 | B2 | 2009-10-20 | Multi-parameter infection monitoring | 2007-04-19 | Methods for monitoring an infection in proximity to an implantable medical device include monitoring a first indicator of infection in proximity to a medical device implanted in a patient; monitoring a second patient parameter; and determining whether the combination of the first indicator and the second parameter are indicative of infection in proximity to the implanted device. If a determination is made that the combination of the first indicator and the second parameter are indicative of infection, an alert may be issued. | 1. A method comprising: monitoring, via an implantable medical device, a first indicator of infection in proximity to the medical device implanted in a patient;monitoring a second patient parameter that provides information regarding whether the first indicator is indicative of infection, wherein the second patient parameter is not an indicator of infection; anddetermining whether the combination of the first indicator and the second parameter are indicative of infection in proximity to the implanted device,wherein determining whether the combination of the first indicator and second parameter are indicative of infection in proximity to the implanted device comprises: determining whether a first value associated with the first indicator crosses a first threshold; anddetermining whether a second value associated with the second patient parameter crosses a second threshold,wherein the combination is determined to be indicative of infection if the first value crosses the first threshold and the second value does not cross the second threshold. 2. The method of claim 1, further comprising issuing an alert if the combination of the first indicator and the second parameter are indicative of infection in proximity to the implanted device. 3. The method of claim 1, wherein monitoring the second patient parameter comprises monitoring a parameter related to patient activity. 4. The method of claim 3, wherein monitoring a parameter related to patient activity comprises monitoring heart rate. 5. The method of claim 3, wherein monitoring the first indicator of infection comprises monitoring temperature. 6. The method of claim 1, further comprising determining whether a value associated with the first indicator of infection crosses a threshold, wherein monitoring the second patient parameter commences when the value associated with the first indicator of infection crosses the threshold. 7. The method of claim 1, wherein monitoring the second patient parameter comprises monitoring the second parameter in discrete time intervals. 8. The method of claim 7, further comprising determining whether a value associated with the first indicator of infection crosses a threshold, wherein the frequency with which the second parameter is monitored is increased if the value associated with the first indicator crosses the threshold. 9. A computer readable medium comprising instructions that when implemented cause an implantable medical device to: monitor a first indicator of infection in proximity to the medical device implanted in a patient;monitor a second patient parameter that provides information regarding whether the first indicator is indicative of infection, wherein the second patient parameter is not an indicator of infection; anddetermine whether the combination of the first indicator and the second parameter are indicative if infection in proximity to the implanted devices,wherein the determination of whether the combination of the first indicator and secon... | Medtronic Inc. | A61K009/22
US20080294037 | A1 | 2008-11-27 | Apparatus and Method for Guided Chronic Total Occlusion Penetration | 2008-05-23 | An apparatus and method for guided penetration of a chronic total occlusion in a blood vessel are disclosed. The invention is directed to an apparatus that facilitates accurate placement of a drilling tip within a body lumen using ultrasound-based detection to determine the position of the intravascular catheter relative to the vessel occlusion and vessel walls. | 1. An apparatus for image-based guidance and for traversing an obstruction in a body lumen comprising: a catheter having a proximal end and a distal end and a longitudinal bore therethrough forming a proximal opening and a distal opening;a guide wire having a proximal end, a distal end, and a therapeutic tip at said distal end;a vibration transducer, said vibration transducer being capable of causing vibration of said therapeutic tip;a power source for energizing said vibration transducer to generate vibrations or oscillations comprising at least one frequency, said power source being connected to a controller for controlling said vibrational frequency via controlling energy supplied by said power source, wherein said at least one vibrational frequency creates detectable signals;an imaging system comprising one or more receivers for collecting said detectable signals, and an imaging screen, said controller further comprising a processor for processing said signals into an image on said image screen;wherein said body lumen has walls and wherein said imaging system permits an operator to visualize and position said therapeutic tip relative to said walls and said obstruction. 2. The apparatus of claim 1, wherein said vibration transducer is a piezoelectric micromotor. 3. The apparatus of claim 2, wherein said power source energizes said piezoelectric micromotor with ultrasonic energy and said imaging system is an ultrasound imaging system. 4. The apparatus of claim 3, wherein said imaging system produces 3-dimensional images. 5. The apparatus of claim 2, wherein said piezoelectric motor is an oscillating ceramic motor. 6. The apparatus of claim 2, wherein said piezoelectric motor is attached to said distal end of said catheter and comprises a longitudinal bore through which said guide wire is positioned. 7. The apparatus of claim 2, wherein said piezo-electric motor is attached to said guide wire. 8. A method for guiding an endovascular device through an obstruction in a blood vessel, comprising the steps of: a) providing a device, said device comprising: i) a guide wire having a proximal end, a distal end and a therapeutic tip at said distal end; ii) a catheter having a proximal end, a distal end, and a longitudinal bore therethrough; iii) a piezoelectric micromotor, said micromotor capable of generating one or more vibrational frequencies when energized by a power source and capable of causing said therapeutic tip to vibrate at said one or more vibrational frequencies; iv) an imaging system comprising one or more receivers for receiving vibrational frequency signals and an imaging screen;b) introducing said guide wire into a blood vessel having vessel walls and an obstruction, and advancing said guide wire until said therapeutic tip of said guide wire contacts said obstruction, wherein said catheter is slidably mounted on said guide wire, said guide wire passing through said longitudinal bore of said catheter;c) advancing said catheter over... | - | A61B005/05 A61B017/22
US20080319300 | A1 | 2008-12-25 | Medical Imaging System | 2007-01-05 | Medical imaging system, comprising a movable part (10) composed of a work surface (11) and of a bracket (12) supporting an arm (13) that is equipped with a radiation emitter (14) and a radiation receiver (15), said movable part (10) additionally comprising: means (16) allowing said arm (13) to slide vertically along the bracket (12), and means of rotation (41 to 48) for turning said bracket (12) about a vertical axis (12a). The system can comprise a fixed part (50) with at least one support (51) that can receive part of a patient's body and allows an imaging incidence to be reproduced. | 1-19. (canceled) 20. A medical imaging system comprising a mobile part (10) composed of a pillar stand (12) supporting an arm (13) equipped with a radiation emitter (14) and a radiation receiver (15), said mobile part (10) comprising moreover: means (16) for allowing said arm (13) to slide vertically along the pillar stand (12); andmeans of rotation (41 to 48) of said pillar stand (12) about a vertical axis (12a);said mobile part (10) moreover comprising a work surface (11). 21. A system according to claim 20, moreover comprising means of rotation (16) of the arm (13) about a horizontal axis (16a). 22. A system according to claim 20, wherein the pillar stand (12) is of a fixed size and the arm (13) is at a fixed distance from the pillar stand (12). 23. A system according to claim 20, moreover comprising means for stabilizing the pillar stand (12) and/or the arm (13) and/or the mobile part (10) in a any position. 24. A system according to claim 20, wherein the work surface (11) comprises a cavity (65) in which the receiver (15) can be fitted in order to be fixed on said work surface (11). 25. A system according to claim 24, wherein the receiver (15) is used as a support for at least one part of the body of a patient. 26. A system according to claim 20, wherein the arm (13) equipped with an emitter (14) and a receiver (15) is arranged to frame, horizontally and close to the ground, a support (51) for one part of the body of a patient. 27. A system according to claim 20, characterized in that it comprises moreover means of protection (52) of at least one part of the body of a patient from the radiation emitted by the emitter (14) and/or a collision with at least one element of the system. 28. A system according claim 20, moreover comprising an adjustable screen (61) for displaying a medical image. 29. A system according claim 20, moreover comprising computer means for processing a medical image. 30. A system according to claim 20, wherein the radiation is X-rays. 31. A system according to claim 20, moreover comprising means of motorization (41) making it possible to position at least one element of said system. 32. A system according to claim 20, wherein the receiver (15) comprises at least one digital sensor making it possible to visualize an element of a size less than or equal to 200 μm. 33. A system according to claim 20, moreover comprising a fixed part (50) comprising at least one support (51) which can receive a part of the body of a patient and/or means of reproducing an imaging view. 34. A system according to claim 20, wherein in one position the arm (13) is vertical, the radiation emitter (14) is at the top of the arm (13), and the radiation receiver (15) rests on the work surface (11). 35. A system according to claim 20, wherein in one position the arm (13) is vertical, the radiation emitter (14) and the radiation receiver (15) being opposite the work surface in relation to the pillar stand (12). 36. A system according to... | DESIGNERS DEVELOPERS DISTRIBUTORS ASSOCIATES (D3A) MEDICAL SYSTEMS | A61B005/00 A61B006/00
US20080312558 | A1 | 2008-12-18 | FLEXIBLE CONDUIT INSERTION MEDICAL DEVICE | 2008-06-12 | A flexible conduit insertion medical device includes a flexible medical device conduit and an insertion mechanism. The flexible medical device conduit includes an elongated framework formed from a flexible material (e.g., Nitinol) with a body portion, sharp head, distal end and proximal end. The flexible medical device conduit also includes a flexible tube at least partially jacketing the elongated framework between the distal end and the proximal end. Moreover, the sharp head is disposed at the distal end and is configured for subcutaneous skin insertion and the elongated framework and flexible tube define at least one conduit between the elongated framework and the flexible tube, the conduit having an opening at the distal end. The insertion mechanism is operatively connected to the flexible medical device conduit and configured to insert a portion of the flexible medical device conduit, including at least the sharp head and the opening, into a user's skin target site. | 1. A flexible conduit insertion medical device comprising: an elongated framework formed from a flexible material, the elongated framework having: a body portion;a sharp head;a distal end; anda proximal end; anda flexible tube at least partially jacketing the elongated framework between the distal end and the proximal end, wherein the sharp head is disposed at the distal end;wherein the elongated framework and flexible tube define at least one conduit therebetween with at least one opening therealong; andwherein the sharp head is configured for target site insertion; andan insertion mechanism operatively connected to, and integrated with, the flexible medical device conduit, the insertion mechanism configured to insert a portion of the flexible medical device conduit, including at least the sharp head and the opening, into a target site. 2. The flexible conduit insertion medical device of claim 1 wherein the insertion mechanism is configured to subcutaneously insert a portion of the flexible medical device conduit, including at least the sharp head and the opening into a skin target site. 3. The flexible conduit insertion medical device of claim 1 wherein the at least one opening is at the distal end. 4. The flexible conduit insertion medical device of claim 1, wherein the flexible material is Nitinol. 5. The flexible conduit insertion medical device of claim 4 wherein the distal end of the flexible medical device conduit is formed to have a no load state curved shape. 6. The flexible conduit insertion medical device of claim 1 wherein the insertion mechanism includes a guide channel configured for movement of the flexible medical device conduit therethrough during insertion. 7. The flexible conduit insertion medical device of claim 6 wherein the guide channel is configured to provide a flexible medical device conduit deployment angle of approximately 45 degrees with respect to the target site. 8. The flexible conduit insertion medical device of claim 6 wherein the guide channel is configured to provide a flexible medical device conduit deployment angle of approximately 45 reverse degrees with respect to the target site. 9. The flexible conduit insertion medical device of claim 8 wherein the guide channel is further configured to prevent deflection and buckling of the flexible medical device conduit during insertion. 10. The flexible conduit insertion medical device of claim 1 further including a conduit guide configured to prevent deflection and buckling of the flexible medical device conduit during insertion. 11. The flexible conduit insertion medical device of claim 1 further including a connection port. 12. The flexible conduit insertion medical device of claim 1 further including a housing. 13. The flexible conduit insertion medical device of claim 1 further including a connector in liquid communication with the proximal end of the flexible medical device conduit. 14. The flexible conduit insertion medical device of claim 1 furt... | - | A61B005/00
US20080287755 | A1 | 2008-11-20 | METHOD AND APPARATUS FOR TREND ALERT CALCULATION AND DISPLAY | 2008-05-14 | Embodiments of the present invention provide methods, apparatuses, and systems associated with detecting, monitoring, analyzing, and/or displaying analyte level trends in a body. In an embodiment, there is provided a method comprising obtaining a plurality of analyte measurements from a body using an analyte sensing device, and selecting two or more measurements of the plurality of measurements to calculate and display on a device associated with the analyte sensing device a rate of change of the analyte over a period of time spanning the two or more measurements. | 1. A method, comprising: measuring with an analyte sensing device a plurality of analyte values of an individual for a plurality of points of time over a defined time period;determining for the plurality of points of time over the defined time period a rate of change of the plurality of analyte values; anddisplaying a graphical representation of the rate of change, wherein said determined rate of change is represented as a best fit rate of change from among a plurality of predefined options indicative of relative rates of change, and wherein the best fit rate of change is highlighted. 2. The method of claim 1, wherein said defined time period has a duration of less than 10 minutes. 3. The method of claim 1, wherein said defined time period has a duration of 5 minutes to 30 minutes. 4. The method of claim 1, wherein said graphical representation is displayed on an electronic monitoring unit. 5. The method of claim 1, wherein said graphical representation is displayed on a personal computer. 6. The method of claim 1, wherein said graphical representation depicts a plurality of predefined options indicative of relative rates of change. 7. The method of claim 1, further comprising a secondary notification indicating a current or future condition based on the determined rate of change. 8. The method of claim 7, wherein said secondary notification is displayed as a textual description of the current or future condition. 9. The method of claim 7, wherein said secondary notification comprises an audible alert or alarm. 10. The method of claim 1, further comprising determining the best fit rate of change in accordance with a selection rule incorporating one or more factors to select one rate of change as the best fit of change as opposed to another rate of change. 11. The method of claim 1, wherein highlighting the best fit rate of change comprises highlighting more than one rate of change when the determined rate of change is not an exact match to a single predefined best fit rate of change. 12. A method, comprising: obtaining a plurality of analyte measurements from a body using an analyte sensing device; andselecting two or more measurements of the plurality of measurements to calculate and display on a device associated with the analyte sensing device a rate of change of the analyte over a period of time spanning the two or more measurements. 13. The method of claim 12, wherein the displayed rate of change is a representation of the calculated rate of change. 14. The method of claim 12, wherein the displayed rate of change is a best fit representation of the calculated rate of change. 15. The method of claim 12, wherein the rate of change is displayed on the device graphically. 16. The method of claim 12, wherein the rate of change is displayed on the device textually. 17. The method of claim 12, further comprising a secondary notification indicating a current or future condition based on the determined rate of change. 18. The method ... | ISENSE CORPORATION | A61B005/145
US20080319276 | A1 | 2008-12-25 | Computational user-health testing | 2008-06-02 | Methods, apparatuses, computer program products, devices and systems are described that carry out detecting user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection; mapping the user data from the interaction between the user and the at least one device-implemented application whose primary function is different from symptom detection to at least one user-health test function set; and selecting at least one user-health test function in response to the at least one user-health test function set. | 1-73. (canceled) 74. A system comprising: a device configured to detect user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection; a user data mapping unit; anda user-health test function selection module. 75. The system of claim 74 wherein the device configured to detect user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: a data detection module. 76. The system of claim 74 wherein the device configured to detect user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: a data capture module. 77. The system of claim 74 wherein the device configured to detect user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: a user input device. 78. The system of claim 74 wherein the device configured to detect user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: a user monitoring device. 79. The system of claim 74 wherein the device configured to detect user data from an interaction between a user and at least one device-implemented application whose primary function is different from symptom detection comprises: a user interface. 80. The system of claim 74 wherein the user data mapping unit comprises: at least one user-health test function set. 81. The system of claim 74 wherein the at least one user-health test function set comprises: a mental status analysis module. 82. The system of claim 74 wherein the at least one user-health test function set comprises: a cranial nerve function analysis module. 83. The system of claim 74 wherein the at least one user-health test function set comprises: a cerebellum function analysis module. 84. The system of claim 74 wherein the at least one user-health test function set comprises: an alertness or attention analysis module. 85. The system of claim 74 wherein the at least one user-health test function set comprises: a visual field analysis module. 86. The system of claim 74 wherein the at least one user-health test function set comprises: a neglect or construction analysis module. 87. The system of claim 74 wherein the at least one user-health test function set comprises: a memory analysis module. 88. The system of claim 74 wherein the at least one user-health test function set comprises: a speech or voice analysis module. 89. The system of claim 74 wherein the at least one user-health test function set comprises: a body movement, eye movement, or pupil movement analysis module. 90. The system of claim 74 wherein the at least one user-health test function set comprises: a face pattern... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00 G06Q050/00
US20080281189 | A1 | 2008-11-13 | MEDICAL GUIDING SYSTEM | 2008-05-02 | A medical guiding system of the invention includes: a detection portion for detecting at least one of position and orientation of a medical instrument; a storage portion for storing multiple reference image data obtained from a human body in multiple states and including anatomical positional information of at least one of viscus and organ; a state selection portion for selecting, among the stored reference image data, the reference image data obtained in a state coincident with or approximate to a state of the subject in using the medical instrument; and a guide image creation portion for creating a guide image showing at least one of an anatomical position, shape, and orientation of the medical instrument with respect to the subject, based on at least one of the position and the orientation of the medical instrument detected by the detection portion and reference image data stored in the state selection portion. | 1. A medical guiding system comprising: a detection portion for detecting at least one of a position and an orientation of a medical instrument;a storage portion for storing a plurality of reference image data, the plurality of reference image data being obtained from a human body in a plurality of states before the medical instrument is used with respect to a subject, and including anatomical positional information of at least one of a viscus and an organ that corresponds to each of the plurality of states;a state selection portion for selecting, among the plurality of reference image data stored in the storage portion, the reference image data obtained in a state coincident with or approximate to a state of the subject in using the medical instrument; anda guide image creation portion for creating a guide image showing at least one of an anatomical position, shape, and orientation of the medical instrument with respect to the subject, based on at least one of the position and the orientation of the medical instrument detected by the detection portion and the reference image data stored in the state selection portion. 2. The medical guiding system according to claim 1, wherein the plurality of reference image data stored in the storage portion are obtained in states where a human body is set in a plurality of body positions different from one another, andthe state selection portion selects the reference image data obtained from a human body in a body position coincident with or approximate to a body position of the subject in using the medical instrument. 3. The medical guiding system according to claim 2, wherein the medical instrument is an ultrasound probe to be introduced into a body of the subject, for scanning ultrasound to obtain an ultrasound signal for creating an ultrasound tomographic image, andthe reference image data is obtained at least in a state where the human body is set in a left lateral position. 4. The medical guiding system according to claim 1, wherein the plurality of reference image data stored in the storage portion are obtained, concerning a predetermined one or a plurality of physical features, from a plurality of human bodies corresponding to a plurality of different conditions of the one or the plurality of physical features, andthe state selection portion selects the reference image data obtained from the human bodies having the physical feature coincident with or approximate to a physical feature of the subject in using the medical instrument. 5. The medical guiding system according to claim 4, wherein the physical feature includes at least one of body size, body height, sex, age, anamnesis, and surgical history. 6. The medical guiding system according to claim 4, wherein the medical instrument is an ultrasound probe to be introduced into a body of the subject, for scanning ultrasound to obtain an ultrasound signal for creating an ultrasound tomographic image. 7. The medical guiding system according to claim... | OLYMPUS MEDICAL SYSTEMS CORPORATION | A61B005/05
US20080255473 | A1 | 2008-10-16 | SYSTEMS, METHODS, AND DEVICES FOR SAMPLING BODILY FLUID | 2008-02-25 | A fluid sampling system is disclosed comprising a fluid drawing device and a fluid sampling device. In one embodiment, the fluid sampling device comprises a base portion having a canula, a handle, and a ridge extending around at least a portion of an outer surface; and a top portion having a flange adapted to engage said ridge to couple the top portion to the base portion. In another embodiment, the fluid sampling device comprises a base portion having a canula, a handle, a mounting portion, and a test strip coupled to the base portion at least partially within the mounting portion. In yet another embodiment, the fluid sampling device comprises a test strip housing for receiving an end of a test strip therein. Extending from an end of the test strip housing is a blunt canula. | 1. A fluid sampling device comprising: a blunt canula with a lumen extending therethrough;a platform in fluid communication with said lumen of said blunt canula, said platform comprising a mounting portion having opposing retaining walls; anda fluid monitoring device coupled to said mounting portion of said platform, said fluid monitoring device being in fluid communication with said lumen of said blunt canula. 2. The medical device of claim 1, wherein said platform further comprises an aperture and a groove adjacent said aperture, said aperture being in fluid communication with said lumen of said blunt canula. 3. The medical device of claim 1, wherein said fluid monitoring device comprises a test strip having electrical connections to facilitate analysis of a fluid sample received by said fluid sampling device. 4. The medical device of claim 1, wherein said fluid monitoring device is coupled to said mounting portion of said platform with at least one of an adhesive and a mechanical fastener. 5. The medical device of claim 1, wherein said fluid monitoring device is integrally formed with said platform. 6. The medical device of claim 1, further comprising a housing that is adapted to protect from damage and prevent contamination of said blunt canula, said platform, and said fluid monitoring device, said housing comprising: a substrate that lies adjacent said platform and said fluid monitoring device;a railing extending from said substrate; anda cover coupled to said railing, said cover substantially encloses said blunt canula to prevent contamination of said blunt canula prior to use. 7. The medical device of claim 6, wherein said cover of said housing is formed of a resilient material adapted to extend around said blunt canula after use. 8. A method for sampling a bodily fluid, comprising the steps: providing a fluid sampling device, said fluid sampling device comprising:a blunt canula having a lumen extending therethrough, said blunt canula being adapted for insertion into a sample port of a central line;a base portion in fluid communication with said blunt canula, said base portion comprising a mounting portion and an aperture, each of said mounting portion and said aperture being in fluid communication with said lumen of said blunt canula; andmeans for detecting properties of a fluid sample coupled to said base portion at least partly within said mounting portion, said means for detecting being in fluid communication with said lumen of said blunt canula;inserting said blunt canula into said sample port to create fluid communication between said central line and said lumen of said blunt canula to enable fluid within said central line to flow into said fluid sampling device; and analyzing the fluid sample received within said fluid sampling device. 9. The method of claim 8, further comprising the step of drawing a bodily fluid into said central line prior to inserting said blunt canula into said sample port. 10. The method of claim 8, w... | - | A61B005/00
US07444875 | B1 | 2008-11-04 | Real time visualization of shear wave propagation in soft materials with sonoelastography | 2006-02-17 | An ultrasound system visualizes shear wave propagation in real time by slowing down the propagation of the shear wave as seen by the ultrasound probe. The shear wave source propagates shear waves into the medium at a frequency ω. The ultrasound probe is vibrated by a vibrator at the frequency ω−Δω, where Δω is much smaller than ω. The wave propagation as seen by the ultrasound probe is slowed down by a factor Δω/ω. An appropriate value of Δω allows real-time visualization of the wave propagation. Variations include electronically producing a virtual vibration and the use of multiple shear wave sources. | 1. A method for visualizing shear wave propagation in a medium, the method comprising: (a) introducing a shear wave from a shear wave source into the medium, the shear wave having a shear wave frequency and a shear wave phase velocity;(b) detecting a propagation of the shear wave through the medium by using an ultrasound probe to produce an ultrasound signal;(c) introducing a vibration into the ultrasound signal at a second frequency;(d) from the ultrasound signal having the vibration, deriving a modulation wave which represents a propagation of the shear wave relative to the ultrasound probe, the modulation wave having a modulation wave phase velocity which varies as a difference between the shear wave frequency and the second frequency; and(e) producing an image representing the shear wave propagation from the modulation wave derived in step (d). 2. The method of claim 1, wherein step (e) comprises producing an image in real time of the modulation wave. 3. The method of claim 2, wherein step (e) further comprises estimating, from the image of the modulation wave, a local vibration phase at each pixel of the image of the modulation wave. 4. The method of claim 3, wherein a plurality of said images in real time of the modulation wave are produced, and wherein the local vibration phase is estimated from a Fourier transform carried out at a time trace of each pixel from the plurality of images of the modulation wave. 5. The method of claim 3, wherein step (e) further comprises performing a phase unwrapping on the local vibration phase to produce a smooth phase surface. 6. The method of claim 5, wherein step (e) further comprises performing a noise reduction on the smooth phase surface. 7. The method of claim 5, wherein step (e) further comprises converting the smooth phase surface to a wave pattern with an artificially increased spatial density. 8. The method of claim 1, wherein step (c) comprises physically vibrating the ultrasound probe at the second frequency. 9. The method of claim 8, wherein step (b) comprises applying a layer of ultrasound gel between the medium and the ultrasound probe such that the ultrasound probe is acoustically coupled to the medium through the ultrasound gel as the ultrasound probe vibrates. 10. The method of claim 1, wherein step (c) comprises electronically producing an equivalent virtual vibration in the ultrasound signal. 11. The method of claim 1, wherein step (a) comprises introducing multiple shear waves into the medium. 12. A system for visualizing shear wave propagation in a medium, the system comprising: a shear wave source for introducing a shear wave into the medium, the shear wave having a shear wave frequency and a shear wave phase velocity;an ultrasound probe for detecting a propagation of the shear wave through the medium to produce an ultrasound signal, a vibration at a second frequency being introduced into the ultrasound signal;an image processing device, in communication with the ultrasoun... | University of Rochester | A61B008/00 A61B005/103
US07440789 | B2 | 2008-10-21 | Electrode structure for measuring electrical responses from the human body | 2004-11-16 | The invention relates to an electrode structure (10) for attachment to a more extensive measuring structure (11), in order to measure electrical responses from the human body. The electrode structure (10) includes a conductive electrode (1). According to the invention, the electrode (1) is shaped to be thin in the thickness direction of the electrode structure (10), and the electrode structure (10) is equipped with a hole (6) and the electrode (1) is located at the edge of the hole (6), in such a way that its longitudinal axis is essentially parallel to the plane of the measurement subject. | 1. An electrode structure for attachment to a more extensive measuring structure, in order to measure electrical responses from the human body, the electrode structure comprising: a inner surface to contact with a measurement subject;an outer surface opposed a predetermined distance to the inner surface;a peripheral surface connecting the inner surface and the outer surface;a central opening in the inner surface extending through the outer surface;an opening in the peripheral surface extending though the electrode structure to meet the central opening; andan electrode fitted in the opening in the peripheral surface, the electrode having a measuring lead, an electrode pellet and a conductor connecting the measuring lead and the electrode pellet, an end of the electrode pellet being positioned where the opening in peripheral surface meets the central opening and a portion of the measuring lead extending beyond the peripheral surface. 2. The electrode structure according to claim 1, wherein the electrode is formed from silver/silver-chloride (Ag—AgCl), in order to form electrically stable interfaces between the measurement subject and the measuring electronics. 3. The electrode structure according to claim 2, wherein the conductor connecting the measuring lead and the electrode pellet is made of pure silver (Ag). 4. The electrode structure according to claim 2, wherein a thickness of the electrode structure anywhere between the inner surface and the outer surface is less than 5 mm. 5. The electrode structure according to claim 1, wherein a thickness of the electrode structure anywhere between the inner surface and the outer surface is less than 5 mm. 6. The electrode structure according to claim 5, wherein the thickness of the electrode structure anywhere between the inner surface and the outer surface is less than 2 mm. 7. The electrode structure according to claim 1, wherein the outer surface is configured to receive a locking piece and the electrode structure attaches to the more extensive measuring structure via the locking piece. 8. The electrode structure according to claim 7, wherein the outer surface has a curved opening and the locking piece locks into the curved opening. 9. The electrode structure according to claim 1, wherein a diameter of the central opening is in a range of between 2 nm and 4 mm. 10. The electrode structure according to claim 1, wherein the electrode pellet is cylindrical in shape and an axial direction in which the electrode pellet extends is essentially parallel to a plane of a subject being measured. 11. The electrode structure according to claim 10, wherein the electrode pellet has a cross-section along any plane perpendicular to the axial direction in which the electrode pellet extends with a surface area that is less than 15 mm2. 12. The electrode structure according to claim 11, wherein the cross-section of the electrode pellet is less than 4 mm2. 13. The electrode structure according to claim 1, wher... | Nexstim Oy | A61B005/04
US07465274 | B2 | 2008-12-16 | Diagnostic apparatus for analyzing arterial pulse waves | 2004-06-16 | A pulse waveform of a living body is detected and is analyzed using peak point data or timing data to perform a diagnosis of a condition of the body. | 1. A pulse wave analyzing apparatus comprising: p1 pulse wave detecting means for detecting a pulse waveform from a living body; and analyzing means for outputting at least one of peak point data or timing data of the pulse waveform detected by said pulse wave detecting means; andwherein the analyzing means comprises: differential means for differentiating the pulse waveform detected by the pulse wave detecting means; inflection detecting means for detecting a timing when an output from said differential means is “0”;count means for counting a sampling clock; andmeans for outputting peak point data or timing data for waveform parameters on the basis of the count value counted by the count means and the amplitude of pulse waveform detected by the detecting means when the output from the differential means is “0”. 2. A pulse wave analyzing apparatus in accordance with claim 1, wherein said living body is a living human body. 3. A pulse wave analyzing apparatus comprising: pulse wave detecting means for detecting a pulse waveform from a living body; andanalyzing means for outputting at least one of peak point data or timing data of the pulse waveform detected by said pulse wave detecting meanswherein said analyzing means comprises: display means for displaying the pulse waveform detected by said pulse wave detecting means;designating means for designating a peak point of the pulse waveform displayed in display means; andmeans for outputting peak point data or timing data for waveform parameters on the basis of the peak point designated by the designating means. 4. A pulse wave analyzing apparatus in accordance with claim 3, wherein said living body is a living human body. 5. A diagnostic apparatus comprising: analyzing means for creating waveform parameters from information representing a pulse wave of a living body; anddiagnostic means for performing a diagnosis of a condition of the living body on the basis of the waveform parameters,wherein said analyzing, means outputs as waveform parameters, peak point data or timing data of the pulse wave obtained from the living body; andwherein said analyzing means comprises: differential means for differentiating the pulse waveform parameters;inflection detecting means for detecting a timing when an output from said differential means is “0”; andcount means for counting a sampling clock; andmeans for outputting peak point data or timing data for the pulse waveform parameters on the basis of the count value counted by the count means and the amplitude of the pulse waveform parameters when the output from said differential means is “0”. 6. A diagnostic apparatus as in claim 5, wherein the diagnostic means perform a diagnosis of a stress level of a living body. 7. A diagnostic apparatus as in claim 5, wherein the stress level is a physiological stress level. 8. A diagnostic apparatus as in claim 5, wherein the stress level is a psychological stress level. 9. A diagnostic apparatus as in claim 5, wherei... | Seiko Epson Corporation | A61B005/00
US07782358 | B2 | 2010-08-24 | Measuring human movements—method and apparatus | 2007-06-08 | A method and apparatus are disclosed for measuring human movement of a test subject during a standard coordination and balance test. The test subject is placed in a selected test position, e.g. standing on one leg for a selected test period. A sensor device is attached to at least one body member of the test subject. The sensor generates and transmits a signal representative of movement of the body member during the test (hereinafter “the test”). The test subject is viewed through a viewfinder of a visual recorder. The recorder and sensor device are activated and generate image and sensor signals which are transmitted to an analyzer for determining the movement of the body members during the test. A coordination and balance score is calculated from the test based on the body movement, body member location and test time. The score is compared a standard associated with the test, and provides an indicator of the test subject's coordination and balance skill. | 1. A method, comprising: causing, at least in part, recording images of a test subject during a selected static balance test for a selected time period by visual recorder;causing, at least in part, reception from at least one sensor attached to at least one body member of the test subject a signal representative of static balance of the body member during the test; andanalyzing image data of the images and sensor data of the signal to determine static balance of the test subject during the test by: aligning one test specific static balance position template over the test subject, and detecting positions of the body member of the test subject that overlap one outline of the position template during the test. 2. The method claim 1, further comprising: including the template aligned in a viewfinder with the test subject during the test wherein the images of the test subject are recorded within the template and the at least one sensor is activated during the test. 3. The method of claim 1, further comprising: calculating a static balance score for the test subject during the test. 4. The method of claim 3, wherein the static balance score is calculated based upon at least an amount of movement, a time period and a body member location. 5. The method of claim 1, further comprising: comparing results of the test to a standard associated with the test. 6. The method of claim 1, further comprising: generating a signal to activate the at least one sensor. 7. The method of claim 6, causing, at least in part, installing a start time stamp in the signal; and causing, at least in part, transmission of the signal to activate the at least one sensor. 8. The method of claim 6, wherein causing, at least in part, installing a stop time stamp in the signal to deactivate the at least one sensor. 9. The method of claim 1, further comprising: causing, at least in part, displaying of an analysis of static balance derived from the image and sensor data corresponding to the at least one body member of the test subject during the test. 10. The method of claim 1, further comprising: causing, at least in part, transmission of the image and sensor data via at least one short-range communication protocol. 11. The method of claim 1, wherein the visual recorder is selected from the group consisting of a portable camera phone, a stationary time-controlled camera, a digital camcorder, and a digital camera. 12. The method of claim 1, wherein the sensor is selected from the group consisting of an accelerometer, an angular rate sensor, a gyroscope, and a magnetometer. 13. The method of claim 1, wherein a test position for a test subject is selected from the group consisting of standing on one leg, one leg standing with different hand positions, one leg standing with different torso positions, and one leg standing with different hand positions and eyes blinded. 14. The method of claim 1, further comprising: causing, at least in part, transmission of the image and sensor ... | Nokia Corporation | H04N007/18 H04N005/262 A61B005/103
US07467005 | B2 | 2008-12-16 | Method and apparatus for non-invasively measuring blood flow and for acquiring and processing an ECG signal | 2003-11-10 | The flow of blood in a vessel (G) in a living being (P) which is exposed to a magnetic field (B) is determined by measuring the potential difference (ΔV) between at least two points (P1, P2) next to the vessel (G). The potential difference (ΔV) is produced on the basis of the Hall effect by the moving charge carriers (Q) in the blood in the presence of the magnetic field (B). When recording an ECG, the cardiac output can be determined in this manner under the influence of a magnetic field (B). At the same time, the ECG signal can be corrected in the knowledge of the component (T) of the ECG signal (E) which can be attributed to the Hall effect. | 1. A non-invasive method for acquiring and processing an electro-physiological ECG signal from a patient who is exposed to an extraneous magnetic field which produces interference in said ECG signal, said method comprising steps of: recording a signal by measuring potential differences between at least two points on a person's body surface near a blood vessel, the signal being an ECG signal overlaid with a magnetic interference component,determining a portion of said ECG signal curve in which the signal is overlaid with a signal component produced on the basis of the Hall effect on charge carriers in the blood flowing in the vessel, and subsequentlyprocessing the signal curve and eliminating the component of the signal which is brought about by the Hall effect in said portion of the signal curve. 2. A method according to claim 1, comprising a further step of determining the cardiac output on the basis of the determined signal curve profile in said portion of the signal curve. 3. An apparatus for acquiring and processing an electro-physiological signal, said apparatus comprising an ECG arrangement having at least two electrodes for measuring a potential difference between at least two points on a surface of the body near a blood vessel and for producing a potential signal, the signal being an electro-physiological ECG signal overlaid with interference signals due to the magnetic field, andmeans for ascertaining a component of the potential signal curve which is produced on the basis of the Hall effect by the movement of charge carriers in the vessel. 4. An apparatus according to claim 3, further comprising means for automatically determining the flow of blood on the basis of the determined component of the potential signal curve. 5. An apparatus according to claim 3, further comprising means for suppressing the component of the potential signal curve. 6. A non-invasive method for operating an electro-physiological instrument and for evaluating an electro-physiological signal acquired from a body exposed to a magnetic field by measuring potential differences between at least two points on a body surface near a blood vessel, the potential being overlaid with interference signals due to the magnetic field, said method comprising steps of determining a component of the signal curve which is caused by the Hall effect on charge carriers in blood flowing in the blood vessel and subsequentlyderiving from the component of the signal curve which is caused by the Hall effect a value which is characteristic of the flow of blood in a vessel. | Schiller AG | A61B005/05
US07569017 | B2 | 2009-08-04 | Diagnosing peripheral vascular disease by monitoring oxygen saturation changes during an ischemia phase | 2008-07-09 | Peripheral vascular disease is diagnosed through measurements of oxygen saturation. In a specific implementation, peripheral vascular disease is diagnosed based on changes in oxygen saturation in tissue. Ischemia is induced, and then measurements of changes in oxygen saturation in tissue are made. Based on changes in oxygen saturation during the induced ischemia phase, a diagnosis is provided of whether a patient has or does not have peripheral vascular disease. | 1. A method of diagnosing peripheral vascular disease by causing a body part of a patient to experience occlusion, which comprises an ischemia phase, the method comprising: measuring oxygen saturation in tissue of the patient during the ischemia phase;analyzing the oxygen saturation in the tissue during the ischemia phase; anddiagnosing the patient as having peripheral vascular disease according to changes in the oxygen saturation in the tissue during the ischemia phase. 2. The method of claim 1 wherein the analyzing the oxygen saturation comprises calculating a rate of change of oxygen saturation during the ischemia phase. 3. The method of claim 2 wherein the rate of change is calculated at a midpoint of the ischemia phase. 4. The method of claim 2 wherein the diagnosing the patient comprises determining if the rate of change crosses a threshold. 5. The method of claim 1 wherein the analyzing the oxygen saturation comprises calculating a time for oxygen saturation to reach a specified percentage during the ischemia phase. 6. The method of claim 5 wherein the diagnosing the patient comprises determining if the time for oxygen saturation to reach a specified percentage crosses a threshold. 7. The method of claim 1 wherein the analyzing the oxygen saturation comprises: calculating a rate of change of oxygen saturation during the ischemia phase; andcalculating a time for oxygen saturation to reach a specified percentage during the ischemia phase. 8. The method of claim 7 wherein the diagnosing the patient comprises determining if the rate of change and the time for oxygen saturation during the ischemia phase cross thresholds. 9. A method of diagnosing peripheral vascular disease by causing a body part of a patient to experience occlusion, which comprises an induced ischemia phase, the method comprising: measuring oxygen saturation of tissue of the patient during the induced ischemia phase; anddiagnosing the patient as having peripheral vascular disease based on changes in the oxygen saturation of the tissue during the induced ischemia phase. 10. The method of claim 9 comprising: providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured rate of change in the oxygen saturation during the induced ischemia phase crosses a threshold value. 11. The method of claim 9 comprising providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured time for the oxygen saturation to crosses a percentage value during the hyperemia phase exceeds a specific time value. 12. The method of claim 9 comprising: providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured time, from a starting time value to an ending time value, for the oxygen saturation to cross a percentage value exceeds a specific time value, wherein the ending time is within the induced ischemia phase. 13. The method of claim 9 compris... | ViOptix Inc. | A61B005/00 A61B005/02
US07582060 | B2 | 2009-09-01 | Diagnosing peripheral vascular disease by monitoring oxygen saturation changes during a hyperemia phase | 2008-07-09 | Peripheral vascular disease is diagnosed through measurements of oxygen saturation. In a specific implementation, peripheral vascular disease is diagnosed based on changes in oxygen saturation in tissue. Ischemia is induced, and then measurements of changes in oxygen saturation in tissue are made. Based on changes in oxygen saturation during a hyperemia phase, a diagnosis is provided of whether a patient has or does not have peripheral vascular disease. | 1. A method of diagnosing peripheral vascular disease by causing a body part of a patient to experience occlusion, which comprises an ischemia phase, the method comprising: measuring oxygen saturation in tissue of the patient during a hyperemia phase after the ischemia phase;analyzing the oxygen saturation in the tissue during the hyperemia phase; anddiagnosing the patient as having peripheral vascular disease according to changes in the oxygen saturation in the tissue during the hyperemia phase. 2. The method of claim 1 comprising: determining a midpoint of the hyperemia phase. 3. The method of claim 1 wherein the analyzing the oxygen saturation comprises calculating a rate of change of oxygen saturation during the hyperemia phase. 4. The method of claim 3 wherein the rate of change is calculated at approximately a midpoint of the hyperemia phase. 5. The method of claim 3 wherein the diagnosing the patient comprises determining if the rate of change crosses a threshold. 6. The method of claim 5 wherein the rate of change is calculated at approximately a midpoint of the hyperemia phase. 7. The method of claim 1 wherein the analyzing the oxygen saturation comprises calculating a time for oxygen saturation to reach a specified percentage during the hyperemia phase. 8. The method of claim 7 wherein the diagnosing the patient comprises determining if the time for oxygen saturation to reach a specified percentage crosses a threshold. 9. The method of claim 1 wherein the analyzing the oxygen saturation comprises: calculating a rate of change of oxygen saturation during the hyperemia phase; andcalculating a time for oxygen saturation to reach a specified percentage during the hyperemia phase. 10. The method of claim 9 wherein the diagnosing the patient comprises determining if the rate of change and the time for oxygen saturation during the hyperemia phase cross thresholds. 11. The method of claim 1 comprising: providing a visual indication of a positive or negative diagnosis of peripheral vascular disease on a display. 12. A method of diagnosing peripheral vascular disease by causing a body part of a patient to experience occlusion, which comprises an induced ischemia phase, the method comprising: measuring oxygen saturation of the tissue of the patient during a hyperemia phase after inducing ischemia; anddiagnosing the patient as having peripheral vascular disease based on changes in the oxygen saturation of the tissue during the hyperemia phase after inducing ischemia. 13. The method of claim 12 comprising: providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured rate of change in the oxygen saturation during the hyperemia phase crosses a threshold value. 14. The method of claim 12 comprising providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured time for the oxygen saturation to crosses a percentage value during the hypere... | ViOptix Inc. | A61B005/00 A61B005/02
US07468039 | B2 | 2008-12-23 | Adjustable tension cuff assembly | 2007-05-25 | A cuff assembly for placement around a body vessel comprises a cuff member sized to substantially surround the body vessel, and a tether assembly engaged with the cuff member. The tether assembly is sized to at least substantially encircle the cuff member, and includes a locking mechanism for releasably maintaining a selected tension when the tether assembly encircles the cuff member. The tether assembly may comprise a tether and a tab, wherein the tab is engaged with the cuff member and has an aperture therethrough. The distal end of the tether is engaged with the tab. The locking mechanism may comprise a plurality of stop members disposed along a surface of the tether. The stop members are sized relative to the aperture to enable passage of at least a portion of the tether therethrough to encircle the cuff member when the stop members are aligned at a first angle relative to the aperture, and to substantially fix a position of the tether at a tension around the cuff member when the stop members are aligned at a second angle relative to the aperture. | 1. A cuff assembly for placement around a body vessel, comprising: a cuff member sized to substantially surround said body vessel; anda tether assembly comprising a tether and a tab, said tab engaged with said cuff member and having an aperture therethrough, said tether having a proximal end and a distal end, said tether distal end engaged with said tab, said tether assembly sized for at least substantially encircling said cuff member, said tether assembly further including a locking mechanism for releasably maintaining a selected tension when said tether assembly encircles said cuff member, said locking mechanism comprising a plurality of stop members disposed along a surface of said tether, said stop members sized relative to said aperture to enable passage of at least a portion of said tether therethrough to encircle said cuff member when said stop members are aligned at a first angle relative to said aperture, and to substantially fix a position of said tether at a tension around said cuff member when said stop members are aligned at a second angle relative to said aperture. 2. The cuff assembly of claim 1, further comprising a lead-in portion along a length of said tether, said lead-in portion positioned for initial passage through said aperture. 3. The cuff assembly of claim 2, wherein said lead-in portion is configured for passage through said aperture and for hindering withdrawal in a reverse direction through said aperture. 4. The cuff assembly of claim 3, wherein said lead-in portion is substantially coil-shaped. 5. The cuff assembly of claim 1, further comprising an extension member juxtaposed between said cuff member and said tether assembly. 6. The cuff assembly of claim 5, wherein said extension member is collapsible into a contracted condition, further comprising at least one dissolvable suture member for maintaining said extension member in said contracted condition for a selected time period. 7. The cuff assembly of claim 1, wherein said stop members comprise a generally square-shaped profile. 8. The cuff assembly of claim 1, wherein said stop members comprise a generally triangular profile. 9. The cuff assembly of claim 1, further comprising a probe operationally engaged with said cuff member for providing a signal corresponding to a fluid flow in said body vessel. 10. A method for monitoring fluid flow in a body vessel, comprising: providing a cuff assembly, said cuff assembly comprising a cuff member sized to at least substantially encircle said body vessel, a probe member operationally engaged with said cuff member and alignable for providing a signal corresponding to fluid flow in said vessel, and a tether assembly engaged with said cuff member and sized for at least substantially surrounding said cuff member, said tether assembly having an aperture therein and a plurality of stop members disposed along a surface thereof and receivable through said aperture to a selected length for maintaining a selected tension whe... | Cook Vascular Incorporated | A61B005/04
US07662086 | B2 | 2010-02-16 | Detection and/or monitoring of diastolic heart failure | 2004-06-17 | In an implantable medical device for detecting and/or monitoring the progression of diastolic heart failure (DHF), and a corresponding method, a parameter is measured that is indicative of left ventricular ejection fraction (LVEF), and a variable is also measured that is indicative of the workload of the patient, and a relation is determined between LVEF and the workload, and DHF is detected and/or the progression of DHF is monitored, dependent on this relation. | 1. An implantable medical device for detecting and/or monitoring progression of diastolic heart failure DHF of a patient, comprising a first measurement unit that measures in vivo a parameter indicative of left ventricular ejection fraction LVEF, an implantable second measurement unit that measures in vivo a variable indicative of the workload of the patient, an implantable determination unit that determines a relation between LVEF and the workload of the patient, and an implantable DHF unit that detects DHF and/or monitors progression of DHF from said relation. 2. The device according to claim 1, wherein said second measurement unit is an accelerometer sensor. 3. The device according to claim 1 wherein said second measurement unit is a unit that measures the intrinsic heart rate. 4. The device according to claim 1 wherein said second measurement unit is a unit that measures minute ventilation. 5. The device according to claim 1 wherein said first measurement unit is an impedance measuring means for measuring the unit that measures an impedance between electrodes (4,6,8,10,12;26,28,30,32,34,36) implanted in the patient's heart. 6. The device according to claim 5, wherein said impedance measuring unit measures a cardiogenic impedance for detecting opening and closing of the aortic valve. 7. The device according to claim 1 wherein said first measurement unit is a unit that measures a pre-ejection period PEP starting with onset of QRS and ending with opening of the aortic valve, and left ventricular ejection time LVET starting with opening and ending with closing of the aortic valve, and that estimates LVEF from PEP/LVET. 8. The device according to claim 7, wherein said unit that measures a parameter indicative of LVEF calculates LVEF from a formula LVEF=k3−k4×(PEP/LVET)where k3 and k4 denote predetermined numerical constants. 9. The device according to claim 7 wherein said unit that measures a parameter indicative of LVEF comprises a sensor for detecting opening and closing of the aortic valve. 10. The device according to claim 9, wherein said sensor is a pressure sensor. 11. The device according to 9, wherein said sensor is a sound sensor for picking-up heart sounds. 12. The device according to claim 1 wherein said determination unit determines said relation by modeling LVEF as a linear function of the workload, A, LVEF=k1+k2*A where k1 and k2 denote two patient dependent, numerical constants. 13. The device according to claim 12, wherein said DHF unit calculates the constant K2 from the formulak2=n⁢∑i=1n⁢⁢Ai⁢Ei-(∑i=1n⁢⁢Ai)⁢(∑i=1n⁢⁢Ein⁢∑i=1n⁢⁢Ai2-(∑i=1n⁢⁢Ai)2where Ai and Ei denote measured workload and ejection fraction values respectively and n the number of measurement values. 14. The device according to claim 12 comprising a memory is provided for storing k2 values. 15. The device according to claim 14, comprising a transmitter that transmits stored k2 values to an external receiver upon request. 16. The device according to claim... | St. Jude Medical AB | A61N001/00
US07462157 | B2 | 2008-12-09 | Method and apparatus for detecting endometriosis | 2005-03-08 | A vaginal probe includes a metal detection sensor for detecting the presence of ferrous-laden endometriotic tissues. The probe is electrically connected to external electronic circuitry for interpreting the readings taken by the sensor and for providing a qualitative or quantitative measurement of the amount of endometriotic tissues present. The probe is inserted in either the vagina or peritoneal cavity of a patient to detect the location and the relative amounts of iron or ferrous-laden tissues in the pelvic area of the patient. Using the probe and monitoring the readings over time as part of pelvic exams assists in determining the early onset or progression of endometriosis and/or the response to one or more therapies being provided to the patient to treat the disease. | 1. A method for detecting endometriosis, the method comprising the steps of: a) providing a probe having detection means for detecting the presence of ferrous-laden tissue, said probe suitable for in vivo use, the detection means further comprising means for generating a magnetic field and means for detecting disruptions in the magnetic field when the probe is in the presence of ferrous-laden tissue;b) placing said probe in the vagina or peritoneal cavity of a patient;c) moving said probe in the vagina or peritoneal cavity of said patient until said probe detects the presence of ferrous-laden tissue; andd) measuring the amount of ferrous-laden tissue, qualitatively or quantitatively, present in the vagina or peritoneal cavity of said patient. 2. The method as set forth in claim 1 further comprising the steps of: a) taking an initial measurement of the amount of ferrous-laden tissue present in the vagina or peritoneal cavity of said patient; andb) taking at least a second measurement of the amount of ferrous-laden tissue present in the vagina or peritoneal cavity of said patientfollowing a period of time to determine if the amount of ferrous-laden tissue in said patient has changed or remained the same. 3. The method as set forth in claim 2 wherein said at least second measurement is taken after said patient has undergone a medical treatment for endometriosis whereby said at least second measurement is taken to determine the efficacy of said treatment. 4. A method for detecting endometriosis, the method comprising the steps of: a) providing a probe having detection means for detecting the presence of ferrous-laden tissue, said detection means comprising means for generating a magnetic field and for detecting disruptions in the magnetic field when the probe is in the presence of ferrous-laden tissue, said probe suitable for in vivo use;b) placing said probe in the vagina or peritoneal cavity of a patient;c) moving said probe in the vagina or peritoneal cavity of said patient until said probe detects the presence of ferrous-laden tissue;d) taking an initial measurement of the amount of ferrous-laden tissue, qualitatively or quantitatively, present in the vagina or peritoneal cavity of said patient; ande) taking at least a second measurement of the amount of ferrous-laden tissue present in the vagina or peritoneal cavity of said patient following a period of time to determine if the amount of ferrous-laden tissue in said patient has changed or remained the same. 5. The method as set forth in claim 4 wherein said at least second measurement is taken after said patient has undergone a medical treatment for endometriosis whereby said at least second measurement is taken to determine the efficacy of said treatment. | - | A61B005/00
US07435231 | B2 | 2008-10-14 | Biological sample device receiver | 2002-10-24 | A blood collection tube holder assembly is disclosed that includes a needle assembly including a piercing end and a non-piercing end and a generally cylindrical housing including a distal end engageable with the needle assembly, a proximal end adapted to receive a blood collection tube, and a sidewall extending between the proximal and distal ends. A cover is movably associated with the proximal end of the cylindrical body and movable between a closed position covering the proximal end and an open position opening the proximal end. | 1. A blood collection tube holder assembly, comprising: a generally cylindrical housing including a distal end engageable with a needle assembly, a proximal end adapted to receive a blood collection tube, a sidewall extending between the proximal and distal ends, and an axial passageway extending through said distal end and having a peripheral surface;the needle assembly including a piercing end, a non-piercing end and at least first and second engagement surfaces between said piercing and non-piercing ends for engaging said peripheral surface of said passageway, wherein said first and second engagement surfaces are axially disposed relative to one another with at least one of said engagement surfaces comprising a radially extending ring;wherein said radially extending ring is engaged in a fluid tight interference fit against the peripheral surface of said passageway. 2. The holder assembly of claim 1wherein the peripheral surface of the passageway defines a slot for receiving and securing said radially extending ring of the needle assembly. 3. The holder assembly of claim 2, wherein the slot in the passageway is defined between a first inwardly extending radial ledge and a second inwardly extending radial ledge axially spaced from the first inwardly extending ledge. 4. The holder assembly of claim 1 further comprising a cover movably associated with the proximal end of a cylindrical body and movable between a closed position covering the proximal end and an open position opening the proximal end. 5. The holder assembly of claim 4, wherein said cover is attached to said housing by a hinge. 6. The holder assembly of claim 5 further comprising a closure member that has a first end attached to the cover and a second end attached to the housing, the closure member being disposed to hold the cover in either the open or closed position and to move the cover preferentially from an intermediate position to the open or closed position. 7. The holder assembly of claim 1 wherein said first and second engagement surfaces of said needle assembly includes a threaded portion proximally disposed relative to said radially extending ring. 8. The holder assembly of claim 1 wherein said needle assembly further includes a second radial ring spaced from said radially extending ring. 9. The holder assembly of claim 1 wherein said ring is engaged in snap engagement with the peripheral surface of said passageway. 10. The holder assembly of claim 9 wherein said peripheral surface of said passageway includes an inwardly extending annular internal rib. 11. The holder assembly of claim 10 wherein said annular internal rib is non-continuous. 12. The holder assembly of claim 10 wherein said peripheral surface of said passageway further includes an internal shoulder. 13. The holder assembly of claim 12 wherein said radially extending ring is engaged in snap engagement between said rib and said shoulder. 14. The holder assembly of claim 1 wherein said needle assembl... | Fenwal, Inc. | A61B005/00
US08882670 | B2 | 2014-11-11 | Apparatus and method for measuring constituent concentrations within a biological tissue structure | 2006-07-06 | Apparatus for minimally invasively measuring concentrations of constituents contained within a biological tissue structure includes a microwave energy source arranged generate a range of microwave frequencies, a first antenna coupled to the microwave energy source and arranged to transmit at least a portion of the microwave energy into the tissue structure, a second antenna arranged to receive at least a portion of the microwave energy transmitted through the tissue structure, a signal processor arranged to determine the resonant frequency of the received microwave energy, and a data processor arranged to provide an output of the concentration of constituents within the biological tissue structure according to the determined resonant frequency. | 1. An apparatus for minimally invasively measuring concentrations of constituents contained within a biological tissue structure, the apparatus comprising: a microwave energy source arranged to generate a range of microwave frequencies;a first antenna coupled to the microwave energy source and arranged to transmit at least a portion of the microwave energy into the tissue structure;a second antenna arranged to receive at least a portion of the microwave energy transmitted through the tissue structure;a signal processor arranged to determine the resonant frequency of the received microwave energy; anda data processor arranged to provide an output of the concentration of constituents within the biological tissue structure according to the determined resonant frequency. 2. The apparatus according to claim 1, wherein the signal processor is arranged to measure the magnitude response of the ratio of the received microwave energy to the transmitted microwave energy and determine the frequency at which a minima or maxima in the magnitude response occurs, said frequency being the resonant frequency. 3. The apparatus according to claim 2, wherein the signal processor is arranged to determine the 3 dB bandwidth of the magnitude response for the frequency of the minima or maxima and thereby derive the Q factor of the biological tissue structure. 4. The apparatus according to claim 3, wherein the data processor is arranged to correlate the derived value of Q factor to a constituent concentration value. 5. The apparatus according to claim 1, wherein the signal processor is arranged to measure the phase response of the ratio of the received microwave energy to the transmitted microwave energy and determine the frequency at which a minima or maxima in the phase response occurs, said frequency being the resonant frequency. 6. The apparatus according to claim 1, wherein the first and second antennas comprise a single transceiver wherein the received microwave energy comprises reflected microwave energy. 7. The apparatus according to claim 6, wherein the apparatus further comprises a reflector plate arranged to reflect microwave energy transmitted from the single antenna back to said antenna. 8. The apparatus according to claim 1, wherein the microwave energy source is arranged to generate microwave energy over a range of frequencies such that at the resonant frequency the biological tissue structure forms a single wave resonance cavity. 9. The apparatus according to claim 1, wherein the microwave energy source is arranged to generate microwave energy over a range of frequencies such that at the resonant frequency the biological tissue structure forms a half wave resonance cavity. 10. The apparatus according to claim 1, wherein the microwave energy source is arranged to generate microwave energy within the frequency range of 8 GHz to 18 GHz. 11. The apparatus according to claim 10, wherein the microwave source is arranged to generate microwave energy wit... | Credent Medical Limited | A61B005/00 A61B005/05 A61B005/145
US08983575 | B2 | 2015-03-17 | Device and method for the determination of the position of a catheter in a vascular system | 2005-02-18 | The invention relates to a device and a method for the determination of the position of a catheter in a vascular system (8). In this, the measured positions (r1, r2) of two magnetic localizers at the tip of a catheter are displaced by correction vectors (k1, k2) while optimizing a quality dimension. The quality dimension includes a component taking account both of the deviation of the measured positions (r1, r2) from the vascular layout and of the deviation of the associated orientation (r2−r1) from the orientation of the vascular layout according to a vascular map. In addition, the quality dimension may include components which evaluate the measured shape of the catheter compared to the vascular map. An additional correction step can further ensure that the corrected positions (r1′, r2″) correspond to the preset fixed distance (d) of the localizers (4, 5). | 1. A device for determining a position of an instrument in a vascular system, comprising: two localizers fitted to the instrument, each of the localizers incorporating a magnetic field sensor configured to measure a plurality of spatial positions of the localizers used to determine an orientation of the instrument and/or a shape of a section of the instrument between the two localizers;a memory to store a vascular map representing a vascular layout of the vascular system in which the instrument is positioned; anda processor configured to correct the measured plurality of spatial positions of the localizers taking into account a distance of the spatial positions from the vascular layout represented on the vascular map, a difference between the spatial positions of the two localizers, and a quality dimension including a first weighted component indicating a deviation of the measured plurality of spatial positions of the two localizers from the vascular layout as represented by the vascular map, and a second weighted component indicating a deviation of the measured orientation of the instrument and/or shape of the section from the vascular layout as represented by the vascular map. 2. The device as claimed in claim 1, wherein the processor is further configured to calculate a locally continuous transformation from respective corrections of the plurality of spatial positions. 3. The device as claimed in claim 1, wherein the localizers are attached to the instrument in a known relative position used to correct the measured plurality of spatial positions. 4. The device as claimed in claim 3, wherein the quality dimension is used in accordance with the vascular layout, so that the corrected plurality of spatial positions adopt the known relative position. 5. The device as claimed in claim 1, wherein the processor is further configured to output a warning if the corrected plurality of spatial positions includes one of an orientation of the instrument and/or a shape of the section of the instrument that deviates by more than a preset limit value from the indicating orientation of the instrument and/or shape of the section. 6. The device as claimed in claim 1, wherein the plurality of spatial positions relative to the vascular layout is verified. 7. The device according to claim 1, further comprising an imaging device configured to generate the vascular map. 8. A method for determining a position of an instrument in a vascular system, the method comprising acts of: providing two localizers attached to the instrument each of the localizers incorporating a magnetic field sensor and a vascular map representing a vascular layout of the vascular system in which the instrument is positionedmeasuring a plurality of spatial position of the localizers;determining an orientation of the instrument and/or shape of a section of the instrument between at least two localizers; andcorrecting the measured plurality of spatial positions using a distance of the spatia... | Koninklijke Philips N.V. | A61B005/05 A61B005/06 A61B019/00 A61B006/12
US08758020 | B2 | 2014-06-24 | Periodic evaluation and telerehabilitation systems and methods | 2008-05-09 | The present invention relates to a system and method for periodic evaluation and telerehabilitation in which a portable device is adapted to be held or coupled to a body part. For example, the portable device can be a cell phone. The body part can be, for example, the hand, ankle, torso, forearm or upper arm. An animated or virtual reality sequence forming a videogame runs or the portable device. Upon interaction of the user with the videogame, movement of the body part is captured as exercise data. The exercise data can be forwarded and stored at a remote clinical server. The system provides monitoring and scheduling of a therapy session by a clinician/doctor/therapist to control the exercises, length and/or difficulty of the exercises based on recorded data. In one embodiment, vital signs or body temperature are monitored by sensors on the portable device or by vital sign monitors associated with the user. The vital sign data can be synchronized with the exercise data at the remote clinical server, and used to improve patient safety during home therapy sessions. | 1. A system for periodic evaluation and telerehabilitation of a disabled user having a body part, comprising: a portable processing device; said portable device adapted to be held or attached to said body part of said disabled user;an animated or virtual reality sequence forming a videogame, said videogame being processed on said portable device to provide directions to said disabled user so as to elicit movements of said body part by said disabled user;means for capturing exercise data resulting from interaction of said user with said videogame;means for measuring a baseline of said disabled user based on a subset of the directions to said disabled user that direct said disabled user to move parts of said disabled user's body;means for mapping said measured baseline with said videogame for tuning said videogame by adjusting amplification of said means for capturing exercise data to provide mapping between real movement of said disabled user and movement of said avatar during said videogame so that movement of the avatar appears greater than the real movement of said disabled user;remote clinical server means for receiving said exercise data; andstoring means for storing said exercise data at said remote clinical server, the exercise data including errors and game duration;exercise scheduler and difficulty setting means for setting a therapy session using said portable device in accordance with said exercise data including errors and game duration, wherein said therapy session controls a length and a difficulty of said videogame being processed on said portable processing device in accordance with said exercise data including errors and game duration. 2. The system of claim 1 wherein a clinician/doctor/therapist can access said therapy session on a web portal to follow a history of a therapy session of said user. 3. The system of claim 1 wherein a clinician/doctor/therapist can access said therapy session and coach a remote said user in real time. 4. The system of claim 1 further comprising: one or more vital sign monitors, said vital sign monitors capturing vital sign data for a vital sign of said disabled user and said remote clinical server means receiving said vital sign data. 5. The system of claim 4 further comprising means for synchronizing said exercise data with said vital sign data at said remote clinical server. 6. The system of claim 5 wherein said portable device includes a sensor to measure body temperature data and said remote clinical server means receiving said body temperature data. 7. The system of claim 6 further comprising watchdog means, said watchdog means interrupting said therapy session if said vital sign data and/or said body temperature data, and/or duration of play of said videogame in said therapy session exceeds a threshold and further comprising means for notifying an authorizer of said therapy session. 8. The system of claim 4 wherein said portable device includes a removable memory card for storing said ex... | - | A63B069/00
US08647268 | B2 | 2014-02-11 | Patient monitoring apparatus | 2008-08-24 | A monitoring system for a person in a building includes a plurality of wireless bases positioned in the building; a body mounted temperature sensor; a body mounted heart signal sensor; and a wearable device coupled to the temperature sensor and to the heart signal sensor and having a wireless transceiver adapted to communicate with the wireless bases; and an accelerometer. | 1. A monitoring system for a person in a building, comprising: a plurality of 802-protocol wireless bases positioned in the building;a sensor mounted on the person's finger in order to detect blood-oxygen levels and pulse rate; anda wristwatch wearable device coupled to the sensor and having a wireless 802-protocol transceiver adapted to communicate with the wireless bases, and a recognizer in the wearable device, the recognizer including one of: a neural network, fuzzy logic, Hidden Markov Models (HMMs), dynamic time warping engine, wherein the recognizer receives data from one or more sensors and detects a predetermined condition based on the heart signal sensor, and wherein the wearable device sends an assistance request for the person based on the predetermined condition, and wherein the wearable device provides the person with home automation control. 2. The system of claim 1, wherein the wearable device determines position based on triangulation. 3. The system of claim 1, wherein the wearable device determines position based on RF signal strength and RF signal angle. 4. The system of claim 1, further comprising a device to detect a confirmatory signal from the person before sending an assistance request. 5. The system of claim 4, wherein the device comprises a switch. 6. The system of claim 1, comprising a heart rate sensor or a blood pressure sensor or an EKG sensor. 7. The system of claim 1, comprising an EKG sensor having a plurality of EKG nodes attached to a plurality of locations on the body of the person. 8. The system of claim 1, wherein the wireless transceiver communicates over an 802 protocol. 9. The system of claim 1, further comprising computer readable code to transmit a help request to a remote computer. 10. The system of claim 1, further comprising computer readable multimedia communication code to allow a user and a remote person to audibly communicate with each other. 11. The system of claim 10, wherein the remote person includes a doctor, a nurse, a medical assistant, or a caregiver. 12. The system of claim 1, further comprising code on a computer readable medium to store and analyze patient information. 13. The system of claim 12, wherein the patient information includes medicine taking habits, eating and drinking habits, sleeping habits, or exercise habits. 14. The system of claim 1, wherein the wireless transceiver complies with 802.15 hardware requirements and uses a non-802.15 stack. 15. The system of claim 1, comprising code on a computer readable medium to store and analyze information relating to the person's ambulation. 16. The system of claim 1, comprising a global positioning system (GPS) receiver to detect movement and where the person falls. 17. The system of claim 1, comprising code on a computer readable medium to map the person's location onto an area for viewing. 18. The system of claim 1, comprising: one or more cameras positioned to capture a video of the patient; anda server coupled to... | - | A61B005/00 G06F019/00
US08740819 | B2 | 2014-06-03 | Method and apparatus for measuring representational motions in a medical context | 2008-03-20 | Disclosed is an apparatus and corresponding method to capture individual representational motions made by a person while performing a neuropsychological test. A user of the apparatus and method may be allowed to view, edit, classify, and retrieve any captured and recorded motions made by the person. In one embodiment, a capture unit, such as a digitizer, captures the motions made by a person while performing a Clock Drawing Test, and the captured data may then be reported in a myriad of ways. The apparatus and method also provide a means to standardize administration and scoring of writing utensil and paper based neuropsychological testing administered to a person or control and test groups for domestic or international applications. By analyzing individual representational motions made by a person while performing a neuropsychological test, unexpectedly more information may be gleaned from the test than was available to physicians from only a completed drawing. | 1. A computer implemented method for processing graphomotor motion in a diagnostic task, the method comprising: receiving, at a computing device, first data representing graphomotor motion made by a first subject during a first execution of a graphomotor diagnostic task performed in a medical context, wherein the data representing the motion comprises data representing dynamic characteristics of said motion;processing, using the computing device, prior data characterizing motion made during one or more prior executions of the diagnostic task to determine characteristics of one or more motion element classes;processing, using the computing device, the first data representing the motion to identify a plurality of elements in the motion, including classifying at least some of said elements according to the one or more motion element classes; anddetermining, using the computing device, using the processed first data and the determined characteristics of the one or more motion element classes, a quantitative, qualitative and/or categorical assessment of the first execution of the diagnostic task associated with neurocognitive mechanisms underlying said motion. 2. The method of claim 1 further comprising soliciting the first subject to make the graphomotor motion and recording the first data representing said motion. 3. The method of claim 1 wherein the prior data includes data characterizing one or more prior executions by the first subject of the diagnostic task. 4. The method of claim 2 further comprising identifying, using the processed first data and the determined characteristics of the one or more motion element classes, changes in the first subject's executions of the graphomotor diagnostic task over time. 5. The method of claim 1 wherein the prior data includes data characterizing one or more prior executions by subjects other than the first subject of the diagnostic task. 6. The method of claim 5 wherein the processing to determine characteristics of the one or more motion element classes includes establishing a performance baseline, and determining the assessment of the first diagnostic task by the subject is based on the performance baseline. 7. The method of claim 1 wherein the diagnostic task includes a clock drawing test. 8. The method of claim 7 wherein the motion element classes include one or more of hour numbers, an hour hand, a minute hand, a second hand, a clock face outline, and a clock face center point. 9. The method of claim 1 wherein the dynamic characteristics include one or more of a change of position relative to a reference point, an acceleration of motion, a rate of motion, a time of making a motion relative to a time of making another different motion, a starting position and an ending position relative to an expected starting position and an expected ending position, and one or more via point positions relative one or more corresponding via point positions between a starting position and an ending position. 10.... | Massachusetts Institute of Technology Lahey Clinic Foundation, Inc. | A61B005/103 A61B005/11
US20080312560 | A1 | 2008-12-18 | Transducer Unit, Device Arrangement and a Method Utilizing the Device Arrangement for Creating and Presenting an Estimate of Bone Mass Development | 2005-05-25 | The invention relates to a method and device arrangement for determining and presenting the change of bone mass caused by accelerations having an impact on the body and representing the amount and intensity of exercise. The device arrangement according to the invention includes a transducer unit, a data transfer medium, a server and a database connected to it. The invention also relates to a transducer unit used in the device arrangement and a method utilizing the device arrangement for measuring and analyzing the intensity, amount and effectiveness of the exercise. | 1. A method for determining and presenting a forecast of the bone mass development caused by the exercise performance of an exercising person (19), concerning one part of the skeletal structure, which method comprises: a step (42), in which a transducer unit (11) measuring acceleration data, carried along by the exercising person (19), continuously measures the acceleration maximums experienced by the supportive organs of the body of the exercising person,characterized in that for drawing up an estimate concerning one part of the skeletal structure, the method further comprises:a step (43), in which the acceleration maximums measured by the transducer unit (11) from a certain period of time are classified as belonging to an acceleration level class, where the classification comprises: defining before measurement an upper and lower limit for an acceleration measurements range based on one part of the skeleton structure;dividing the defined acceleration measurements range in equal acceleration level classes; andclassifying during measurement each measured acceleration maximum to a certain acceleration level class, in which case a classified maximum adds occurrence number (N) by one in that acceleration level class;a step (44), in which the classified acceleration maximums are compared to reference data measured from normal population;a step (45), in which it is checked if the classified acceleration measurement level classes of the exercising person (19) are higher or lower than the reference data, anda step (46, 48) in which an estimate of the bone mass development concerning one part of the skeletal structure is presented to the exercising person (19) on the basis of the comparison. 2. The method according to claim 1, characterized in that a measuring range for acceleration maximums is between +0.3 g and +12 g and that the width of one acceleration level class belonging to it is of the order of 0.3 g. 3. The method according to claim 1, characterized in that the numbers of occurrence (N) included in selected acceleration level classes are compared to corresponding numbers of occurrence measured from the normal population, and that a forecast for the development of bone mass of one part of the skeletal structure of the exercising person (19) is calculated from the result of the comparison. 4. (canceled) 4. The method according to claim 1, characterized in that a graph is formed of the gathered numbers of occurrence (N), with the acceleration classes as the argument and the relative numbers of occurrence of the acceleration level classes as the function, and that the area between two acceleration levels defined by the graph of the function is determined and that the area determined is compared to the corresponding area obtained from measurements of the normal population, and that the result of the comparison is used to draw up a forecast of the development of the bone mass of one part of the skeletal structure of the exercising person (19). 5.... | NEWTEST OY | A61B005/103
US20080242942 | A1 | 2008-10-02 | Fetal movement calculator | 2007-03-27 | Mother needs to calculate the fetal movement manually several times per day. No any instrument which can let mother save these jobs in the past.;This instrument combined the function of LCD alarm clock, stop watch, and calculator to automatically calculate the daily fetal movements, and converts to the dairy and weekly average figures; and show each figure on the LCD screen. | 1. Claim the patent that covers the design which combines the function of LCD alarm clock, stop watch, and calculator to automatically calculate the daily and weekly fetal movement average figures. 2. Claim the patent which adds more accessory functions such as the remote control keys etc. based on the above (1) basic function design. 3. Claim the patent which is based the above (1) design but to calculate other physical data such as Blood pressure, pulse movement average figures etc. | - | A61B005/00
US20080294024 | A1 | 2008-11-27 | Glucose meter system and monitor | 2007-05-24 | A handheld portable glucose meter includes a glucose sensor having a sensor output related to glucose in a blood sample. A display is configured to display information to a user. A manual input is configured to receive user input data from the user. Communication circuitry is configured to send and/or receive data to and/or from a remote location. A controller controls operation of the handheld portable glucose meter. | 1. A handheld portable glucose meter, comprising: a glucose sensor having a sensor output related to glucose in a blood sample;a display configured to display information to a user;a manual input configured to receive user input data from the user;a remote I/O (input/output) configured to send and receive data to and from a remote location; anda controller configured to send data to the remote location based upon the user input data, is configured to display information on the display based upon data received from the remote location. 2. The apparatus of claim 1 including a memory which contains an address identifying the handheld portable glucose meter. 3. The apparatus of claim 2 wherein data sent to the remote location is associated with the address stored in the memory. 4. The apparatus of claim 2 wherein data received from the remote location is associated with the address stored in the memory. 5. The apparatus of claim 2 wherein the address comprises 32 bits. 6. The apparatus of claim 2 wherein the address comprises 128 bits. 7. The apparatus of claim 1 wherein the remote I/O is configured to communicate over a network. 8. The apparatus of claim 1 wherein the remote I/O is configured for wireless communication. 9. The apparatus of claim 1 wherein the controller is configured to receive information related to dosage information for the user. 10. The apparatus of claim 9 wherein the controller is configured to display information on the display based upon dosage information and the sensor output from the blood glucose sensor. 11. The apparatus of claim 10 wherein the controller is further configured to display dosage information on the display based upon user input from the manual input. 12. The apparatus of claim 1 wherein controller receives dosage information from the user through the manual input and is configured to transmit data related to the dosage information to the remote location over the remote I/O. 13. The apparatus of claim 1 wherein the display is configured to display at least one soft key and wherein the user input received through the manual input is a function of a value of the soft key. 14. The apparatus of claim 1 wherein the manual input is associated with the display to thereby provide a touch screen display. 15. The apparatus of claim 1 including an audio output. 16. The apparatus of claim 15 wherein audio which is output from the audio output is based upon data received from the remote location. 17. The apparatus of claim 1 including an audio input. 18. The apparatus of claim 17 wherein data sent to the remote location is based upon audio received through the audio input. 19. The apparatus of claim 1 wherein the remote I/O is configured to communicate over a cellular data network. 20. The apparatus of claim 1 wherein the remote I/O is configured to send and receive messages in accordance with an SMS (short messaging service). 21. The apparatus of claim 1 including a memory configured to store progr... | - | A61B005/05 A61B017/52 G01R015/00 G01R021/00
US20080262389 | A1 | 2008-10-23 | MONITORING CONDITIONS OF A PATIENT'S URINARY SYSTEM | 2008-05-13 | A monitoring system and method are presented for use in monitoring a condition of a patient's urinary system. The monitoring system comprises an acoustic assembly comprising at least one acoustic receiver adapted for continuously receiving acoustic signals during a patient's urination and generating data indicative thereof. The monitoring system also includes a control unit that is in communication with said acoustic assembly. The control unit is configured and operable for analyzing said generated data indicative of the continuously received acoustic signals during a patient's urination, obtaining a time variation of the acoustic signal during the urination and determining a corresponding spectral data of the acoustic signal. The control unit further analyzes the spectral data and, upon detecting at least one first signal peak corresponding to a condition of turbulence in the urine flow, determining a relation between said first signal peak and a second signal peak corresponding to a condition of laminar urine flow. Based on said relation, the control unit determines the condition of a patient's low urinary system and generating output data indicative thereof. | 1. A monitoring system for use in monitoring a condition of a patient's urinary system, the monitoring system comprising: (a) an acoustic assembly comprising at least one acoustic receiver adapted for continuously receiving acoustic signals during a patient's urination and generating data indicative thereof; and(b) a control unit in communication with said acoustic assembly, the control unit being configured and operable for analyzing said generated data indicative of the continuously received acoustic signals during a patient's urination, obtaining a time variation of the acoustic signal during the urination and determining a corresponding spectral data of the acoustic signal, analyzing the spectral data and, upon detecting at least one first signal peak corresponding to a condition of turbulence in the urine flow, determining a relation between said first signal peak and a second signal peak corresponding to a condition of laminar urine flow, and based on said relation determining the condition of a patient's low urinary system and generating output data indicative thereof. 2. The system of claim 1, wherein said spectral data includes a Strouhal frequency range. 3. The system of claim 1, wherein said spectral data includes a frequency range of about 20-1000 Hz. 4. The system of claim 1, wherein the second signal peak corresponding to condition of laminar urine flow is in a frequency range of 70-150 Hz. 5. The system of claim 1, wherein the first signal peak corresponding to the turbulent urine flow is in a frequency range of 150-1000 Hz. 6. The system of claim 1, wherein the control unit is configured and operable for analyzing the spectral data by determining a time variation of the relation between the first and second signal peaks, a frequency of at least the first signal peak varying with time during the urination. 7. The system of claim 1, wherein the control unit is configured and operable to determine the relation between the first and second signal peaks by calculating at least one of the following: a ratio between amplitudes of the first and second signals, and a ratio between frequencies of the first and second signals, and time variations of these ratios during urination and/or during successive urinations. 8. The system of claim 1, wherein the control unit comprises a memory utility for storing reference data comprising a given value or a range of values for at least one of the following parameters: an urethral diameter, urethral length, and elasticity of an urethral wall. 9. The system of claim 1, wherein the control unit is configured and operable to apply a predetermined model to the spectral data, said model being based on a given value or a range of values for at least one of the following parameters: an urethral diameter, urethral length, and elasticity of an urethral wall. 10. The system of claim 1, wherein the control unit is configured and operable to process and analyze the relation between the first and second si... | P. SQUARE MEDICAL LTD. | A61B007/00 A61B005/00
US20080287756 | A1 | 2008-11-20 | Pulse oximetry relational alarm system for early recognition of instability and catastrophic occurrences | 2008-05-16 | A relational pulse oximetry alarm system and method is presented for earlier identification of the occurrence of an adverse clinical event. The system includes a pulse oximeter based microprocessor alarm system which provides an alarm based on a relational conformation of a plurality of time series and further based on the recognition of specific dynamic patterns of interaction between a plurality of corresponding and related time series including the occurrence of pathophysiologic divergence of two or more time series outputs. The processor is programmed to compare a first time series to a second time series to produce a comparison result, to identify a relationship between the first time series and the second time series, to identify a relational threshold breach, and to output an alarm based on the relational threshold breach. The system can include an oximeter testing system for predicting the timeliness of the response of the alarm of a pulse oximeter to the occurrence of an adverse clinical event. | 1. A pulse oximeter based system for the recognition of patterns in a plurality of time series, the system comprising a processor programmed to: process a first time series to identify a first primary threshold breach based on the first time series;process a second time series;compare at least a portion of the first time series to at least a portion of the second time series to produce a comparison result;identify a relational threshold breach based at least in part on the comparison result; andoutput an alarm based on at least one of the primary and the relational threshold breaches. 2. An oximeter testing system for testing an oximeter, the system comprising: a processor programmed to generate an output indicative of a time series of oxygen saturation values, the output simulating a time series of oxygen saturation values associated with an occurrence of an adverse clinical event, and programmed to calculate a difference between a time of onset of the occurrence and a time of onset of the alarm; andan oxygen saturation simulator for presenting the output to the oximeter to determine a delay between the time of onset of the occurrence and the time of onset of the alarm. 3. The system of claim 2, wherein the output is simulating a time series of oxygen saturation values associated with a respiratory arrest. 4. The system of claim 2, wherein the output is simulating a time series of oxygen saturation values associated with a respiratory arrest occurring with a lung volume at functional residual capacity at the time of onset of the occurrence. 5. The system of claim 2, wherein the output is simulating a time series of oxygen saturation values associated with a respiratory arrest occurring with an alveolar oxygen concentration at about 21% at the time of onset of the occurrence. 6. The system of claim 2, wherein the processor is programmed to identify the time of onset of the alarm. 7. The system of claim 2, wherein the processor is programmed to measure the difference between the time of onset of the occurrence and the time of onset of the alarm. 8. The system of claim 2, wherein the processor is programmed to output an indication of the difference. 9. The system of claim 2, wherein the processor is programmed to output the difference. 10. The system of claim 2, wherein the oximeter includes a probe configured to interface with the oxygen saturation simulator via a member, the member configured to present a secondary dynamic variation simulating a primary dynamic variation in oxygen saturation, the secondary dynamic variation being based on the output, wherein the processor is configured to present the output to the member and the member is configured to present the dynamic variation to the probe. 11. The system of claim 2, comprising an audio sensor configured to automatically detect the onset of the alarm. 12. A method of processing data comprising a plurality of time series, the method comprising: searching a first one of the pluralit... | - | A61B005/00
US20080300619 | A1 | 2008-12-04 | RECTAL BALLOON APPARATUS WITH RADIATION SENSOR AND/OR MARKERS | 2008-06-18 | A rectal balloon apparatus has a shaft with a fluid passageway extending therethrough. A balloon is affixed over an end of the shaft such that the fluid passageway communicates with an interior of the balloon. The balloon has a laterally flat surface thereon when inflated. The laterally flat surface of the balloon has a longitudinal groove formed thereon. The balloon has a radiation sensor positioned within the groove and a plurality of fiducial markers positioned on a surface of the balloon. A stopper is slidably mounted on the shaft. | 1. A rectal balloon apparatus comprising: a shaft having a fluid passageway extending therethrough, said shaft having indicia formed on an exterior surface thereof;a stopping means slidable longitudinally along said shaft so as to reside at a desired indicia, said stopping means for fixing an amount of movement of said shaft into the rectum; anda balloon affixed over an end of said shaft such that said fluid passageway communicates with an interior of said balloon, said balloon movable between a non-inflated condition and an inflated condition. 2. The apparatus of claim 1, said balloon having a sensing means affixed theron, said sensing means for sensing an amount of radiation delivered thereto. 3. The apparatus of claim 1, said balloon having a laterally flat surface when in said inflated condition, said balloon having a longitudinal groove formed in said laterally flat surface. 4. The apparatus of claim 3, said groove having a sensing means affixed within said groove, said sensing means for sensing an amount of radiation delivered thereto. 5. The apparatus of claim 1, said balloon having a plurality of fiducial markers formed on a surface thereof. 6. The apparatus of claim 5, said fiducial markers affixed or formed on different surfaces of said balloon. 7. The apparatus of claim 6, said balloon having a groove formed thereon, said fiducial markers comprising a first plurality of fiducial markers positioned on one side of said groove and a second plurality of fiducial markers positioned on an opposite side of said groove. 8. The apparatus of claim 6, said fiducial markers comprising a first set of fiducial markers positioned on a top surface of said balloon and a second set of fiducial markers positioned on a bottom surface of said balloon. 9. The apparatus of claim 1, said stopping means comprising: a semispherical member slidably mounted on said shaft, said semispherical member having a curved surface facing said balloon. 10. A rectal balloon apparatus comprising: a shaft having a fluid passageway extending therethrough; anda balloon affixed over an end of said shaft such that said fluid passageway communicates with an interior of said balloon, said balloon movable between a non-inflated condition and an inflated condition, said balloon having a groove formed thereon when in said inflated condition, said groove having a sensing means affixed therein, said sensing means for sensing an amount of radiation delivered thereto. 11. The apparatus of claim 10, said groove having a depth such that said sensing means has an outer surface that is flush with a surface of said balloon or inwardly of said surface of said balloon. 12. The apparatus of claim 10, said balloon having a laterally flat surface thereon when in said inflated condition. 13. The apparatus of claim 10, said shaft having indicia formed on an exterior surface thereof. 14. The apparatus of claim 13, said shaft having a stopping means slidable longitudinally along said shaft ... | - | A61M029/00 A61B005/05
US07460900 | B1 | 2008-12-02 | Method and apparatus for detecting ischemia using changes in QRS morphology | 2005-09-14 | A method for operating an implantable medical device includes determining the difference of the absolute value of the voltage of a test QRS complex and the voltage of a baseline QRS template at a plurality of corresponding sample points and detecting ischemia if the sum of the differences at the plurality of sample points is greater than an ischemia detection threshold. | 1. A method for detecting ischemia, comprising: generating a baseline QRS template;determining absolute value of difference of voltage of a test QRS complex and voltage of the baseline QRS template at a plurality of sample points; anddetecting ischemia if a sum of the differences at the plurality of sample points is greater than an ischemia detection threshold. 2. The method of claim 1 wherein generating a baseline QRS template comprises generating an ensemble average of a plurality of digitized baseline QRS complexes. 3. The method of claim 1 further comprising aligning the test QRS complex with the baseline QRS template. 4. The method of claim 1 wherein aligning the test QRS complex with the baseline QRS template comprises aligning maximum amplitude of the test QRS complex with maximum amplitude of the baseline QRS template. 5. The method of claim 1 further comprising generating and storing an ischemia burden metric. 6. The method of claim 1 further comprising adjusting cardiac pacing parameters to increase heart perfusion in response to detection of ischemia. 7. The method of claim 1 wherein generating a baseline QRS template comprises generating a baseline QRS template at each of a plurality of heart rates. 8. The method of claim 7 further comprising determining current heart rate and wherein determining total difference of absolute value of amplitude of a test QRS complex and amplitude of the baseline QRS template at a plurality of sample points comprises determining total difference of absolute value of amplitude of a test QRS complex and amplitude of the baseline QRS template corresponding to the current heart rate. 9. The method of claim 1 further comprising notifying patient of detection of ischemia. 10. The method of claim 1 wherein generating a baseline QRS template comprises generating a baseline QRS template for each of a plurality of sensing vectors and wherein determining total difference of absolute value of amplitude of a test QRS complex and amplitude of the baseline QRS template at a plurality of sample points comprises determining total difference of absolute value of amplitude of a test QRS complex sensed with each of the plurality of sensing vectors with the amplitude of the corresponding baseline QRS template for that sensing vector. 11. The method of claim 10 wherein detecting ischemia if the absolute value of the difference is greater than an ischemia detection threshold comprises detecting ischemia if the absolute value of the difference for each of the plurality of sensing vectors is greater than an ischemia detection threshold. 12. The method of claim 1 wherein generating a baseline QRS template comprises generating a short term moving average of a plurality of digitized baseline QRS complexes. 13. The method of claim 1 wherein generating a baseline QRS template comprises generating a long term moving average of a plurality of digitized baseline QRS complexes. 14. An implantable medical device, comprising:... | Pacesetter, Inc. | A61B005/04
US20080294056 | A1 | 2008-11-27 | BISPECTRAL PEROPERATIVE OPTICAL PROBE | 2008-05-23 | The general field of the invention is that of peroperative optical probes designed to assist surgeons in performing medical procedures. The optical probes of the fluorescence type according to the invention are designed to be used on living tissues where the diseased areas have been marked by a fluorescent marker. They have dual lighting. The first situated in the red or near-infrared spectrum is necessary for achieving the fluorescence of the marked areas and for obtaining an image exploitable by a camera. The second situated in the visible spectrum is necessary for illuminating the marked areas with visible light, thus making the surgeon's work easier. The visible lighting can be punctiform or can be provided by an image projector. In the latter case, the projected image illuminates only the diseased areas. | 1. A peroperative optical probe comprising: a first light source emitting a first light beam intended to illuminate a biological tissue surface, of which at least one area, marked area, has previously been fixed by a fluorescent marker, said beam being emitted at a first wavelength corresponding substantially to the wavelength of excitation of the marking fluorophore, said fluorophore emitting at a second wavelength;a first optical device comprising a first lens and a photosensitive surface, said lens forming, on said photosensitive surface, a first image of the biological tissue surface, the first device being connected to means for processing the image issuing from the photosensitive surface;optical means that are different from the first optical device and that are designed to illuminate at least a part of the biological tissue surface, the surface part illuminated by the optical means preferably corresponding to at least one of the marked zones identified by the means for processing the image. 2. The peroperative optical probe according to claim 1, wherein the optical means comprise at least one image projector which, at the command of a control device analyzing the signal issuing from the processing means, generates a second image at at least a third wavelength situated in the visible spectrum, said second image being representative of the marked area, of which second image a second lens forms a third image on the biological tissue surface, the second device being arranged in such a way that this third visible image is of the same dimensions as and in coincidence with at least a marked area of the biological tissue surface. 3. The peroperative optical probe according to claim 2, wherein the optical means comprise a splitter allowing the second wavelength to be divided spectrally from the third wavelength, said splitter being arranged in such a way that the parts of the optical axes of the first lens and of the second lens that are situated between the biological tissue surface and said splitter are common. 4. The peroperative optical probe according to claim 1, wherein the first optical device comprises a first optical filter transparent to the second wavelength. 5. The peroperative optical probe according to claim 2, wherein the image projector comprises a second optical filter transparent to the third wavelength. 6. The peroperative optical probe according to claim 1, wherein the optical means comprise at least a third light source emitting a third beam at the first wavelength and arranged in such a way as to illuminate the centre of the biological tissue surface in a quasi-punctiform manner. 7. The peroperative optical probe according to claim 6, wherein the optical means comprise a fourth light source emitting a fourth beam at a fourth wavelength situated in the visible spectrum, and arranged in such a way as to illuminate the centre of the biological tissue surface in a quasi-punctiform manner in coincidence with the third beam. ... | COMMISSARIAT A L'ENERGIE ATOMIQUE | A61B005/00
US20080312524 | A1 | 2008-12-18 | Medical Sensor Having Electrodes and a Motion Sensor | 2006-12-02 | A medical sensor having at least one electrode configured to be placed on a patient for medical monitoring and a motion sensor integrated in the medical sensor with the electrode, the motion sensor configured to detect patient motion and provide electrical signals in response thereto. | 1. A medical sensor which attaches to the body of a patient, comprising: at least one electrode located on a substrate and configured to be attached to a patient for medical monitoring;a motion sensor located on the substrate and configured to detect patient motion and provide electrical signals in response thereto; anda processor electrically coupled to the electrode and motion sensor to receive signals therefrom. 2. The medical sensor of claim 1 wherein the motion sensor is integrated onto the common substrate with the electrode. 3. (canceled) 4. The medical sensor of claim 1 wherein the at least one electrode comprises three or more electrodes. 5. The medical sensor of claim 1 wherein the motion sensor comprises a piezoelectric motion sensor configured to translate patient motion into electrical signals. 6. The medical sensor of claim 5 wherein the piezoelectric motion sensor is formed from a polyvinylidene fluoride (PVDF) layer integrated onto the substrate. 7. The medical sensor of claim 6 wherein the electrode comprises a plurality of electrodes attached to the substrate, and wherein the piezoelectric motion sensor comprises a PVDF layer laminated onto the substrate. 8. The medical sensor of claim 7, wherein the piezoelectric motion sensor comprises first and second PVDF layers laminated in opposition to each other on opposite sides of the substrate. 9. (canceled) 10. The medical sensor of claim 1, further comprising: an adhesive layer configured to adhere the medical sensor to the patient. 11. (canceled) 12. A medical sensor, comprising: a plurality of electrodes configured for attachment to a patient and operable to electrically couple electrical signals to and from the patient; andan integral motion sensor configured to sense patient motion and provide signals in response thereto,wherein the integral motion sensor and the plurality of electrodes are located on a common substrate of the medical sensor. 13. (canceled) 14. (canceled) 15. The medical sensor of claim 12 wherein the integral motion sensor comprises a piezoelectric motion sensor configured to translate patient motion into electrical signals. 16. (canceled) 17. The medical sensor of claim 13, further comprising: an adhesive layer configured to adhere the medical sensor to the patient. 18-22. (canceled) | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/04
US20080287772 | A1 | 2008-11-20 | Motion Compensation in PET Reconstruction | 2008-05-15 | Method and apparatus motion compensation in Positron Emission Tomography (PET) imaging. PET and Magnetic Resonance Imaging (MRI) are performed simultaneously and the latter is used to determine motion of the subject during the acquisition process. The PET image data are then corrected according to the subject motion so determined. | 1. A method of correcting a set of PET data for movement of a subject during PET data acquisition, said method comprising the steps of: simultaneously performing a PET scan and an MRI scan;processing the results of the MRI scan to generate an MRI image and to determine motion of the subject during the scans andcorrecting the data acquired from the PET scan for subject motion so detected. 2. A method according to claim 1 comprising the steps of: acquiring MRI image data over a sequence of discrete time periods, T0-Tn, during acquisition of the PET data, to produce a sequence of corresponding MRI datasets, D0-Dn;selecting an MRI dataset D0, acquired over one of the time periods T0 as a reference;for each dataset, D1, calculating a geometric transformation matrix Mi representing the change in subject position between those represented by Di and D0 respectively;applying each matrix Mi to the PET data acquired during Ti to produce a set of corrected PET data andreconstructing the corrected PET data to produce a set of PET images, corrected for movement of the subject during data acquisition. 3. A method according to claim 2, where the Geometric Transformation matrices are used, during reconstruction, to modify the Geometric value assigned to the probability of each event, associated with a particular pair of detector bins, occurring in a particular voxel. 4. A method according to claim 2, wherein the Geometric transformation matrices are used to reassign each event to a particular detector bin pair prior to image reconstruction. 5. Apparatus for correcting a set of PET data for movement of a subject during PET data acquisition comprising means for simultaneously performing a PET scan and an MRI scan on a subject and a processor arranged to: receive data resulting from the PET and MRI scans;process the data resulting from the MRI scan to determine movement of the subject during the scan andcorrect the data resulting from the PET scan according to the movement so determined. | Siemens Medical Solutions USA, Inc. | A61B005/055 A61B006/00
US20080245692 | A1 | 2008-10-09 | Carrying assembly with analyte-determining instrument | 2008-03-26 | A combination of carrying assembly with analyte-determining device. The analyte-determining device includes housing, a display and at least one user-input mechanism. The carrying assembly is adapted to magnetically attach the analyte-determining device. One of the analyte-determining device and the carrying assembly includes at least one magnet. The other of the analyte-determining device and the carrying assembly includes a magnetic-attracting portion. The at least one magnet is attracted to the at least one magnetic-attracting portion and attaches the analyte-determining device and the carrying assembly with each other. | 1. A combination of carrying assembly with analyte-determining device, the combination comprising: an analyte-determining device including housing, a display and at least one user-input mechanism; anda carrying assembly being adapted to magnetically attach the analyte-determining device,wherein one of the analyte-determining device and the carrying assembly includes at least one magnet, the other of the analyte-determining device and the carrying assembly including at least one magnetic-attracting portion, the at least one magnet being attracted to the at least one magnetic-attracting portion and attaching the analyte-determining device and the carrying assembly with each other. 2. The combination of claim 1, wherein the carrying member is a billfold. 3. The combination of claim 1, wherein the carrying assembly includes the at least one magnet and the analyte-determining device includes the at least one magnetic-attracting portion. 4. The combination of claim 3, wherein the at least one magnet is located within the material that forms the carrying assembly. 5. The combination of claim 1, wherein the analyte-determining device includes the at least one magnet and the carrying member includes the at least one magnetic-attracting portion. 6. The combination of claim 1, further including a lancing device and wherein at least one of the lancing device and the carrying assembly includes at least one magnet, the other of the lancing device and the carrying assembly including the at least one magnetic-attracting portion, the at least one magnet being attracted to the at least one magnetic-attracting portion and attaching the lancing device and the carrying assembly with each other. 7. A method of using a carrying assembly and an analyte-determining device, the method comprising the acts of: providing an analyte-determining device including housing, a display and at least one user-input mechanism;providing a carrying assembly being adapted to magnetically secure the analyte-determining device, wherein one of the analyte-determining device and the carrying assembly includes at least one magnet, the other of the analyte-determining device and the carrying assembly including at least one magnetic-attracting portion; andplacing the analyte-determining device into the carrying assembly such that the at least one magnet is attracted to the at least one magnetic-attracting portion and attaches the analyte-determining device and the carrying assembly with each other. 8. A combination of carrying assembly with analyte-determining device, the combination comprising: an analyte-determining device including housing, a display and at least one user-input mechanism;a removable pouch being adapted to magnetically secure the analyte-determining device; anda carrying assembly being adapted to magnetically secure the removable pouch,wherein the analyte-determining device and the removable pouch includes at least one magnet, the other of the analyte-determining devic... | Bayer HealthCare LLC | A45C015/00 A45C013/02 A61B005/05
US20080242975 | A1 | 2008-10-02 | DEVICE AND METHOD FOR LOCATING A CANNULA THAT IS INSERTED INTO A BODY | 2008-03-26 | A light pointer (16) is mounted on the proximal end of a cannula (10) by use of an adapter (14) for positioning the cannula (10) that is inserted into a body by use of an ultrasonic head. The light pointer (16) emits a light beam (18) in a direction axially parallel to the cannula (10). When the cannula (10) is inserted, the light beam (18) produces a light spot on the bodily surface on which the ultrasonic head is placed. | 1. A device for positioning a cannula (10) inserted into a body by use of an ultrasonic head which may be placed on the bodily surface (22), comprising: a light pointer (16) proximally situated on the cannula (10) which emits a light beam (18) in a distal direction in a plane encompassing a longitudinal axis of the cannula (10) and produces a light spot (24) on the bodily surface (22) in a distal direction in front of a puncture site (20) of the cannula (10) for placement of the ultrasonic head. 2. The device according to claim 1, wherein the light pointer (16) is a laser pointer. 3. The device according to claim 1, wherein the light pointer (16) is detachably mounted on a proximal projection (12) of the cannula (10) by use of an adapter (14). 4. The device according to claim 1, wherein the light beam (18) from the light pointer (16) extends axially parallel to the longitudinal axis of the cannula (10). 5. The device according to claim 1, wherein the light spot (24) has a two-dimensional shape. 6. The device according to claim 5, wherein the light spot (24) has the shape of a cross, in particular having an arm which intersects an axis of the cannula (10) and an arm which extends perpendicular thereto. 7. A method for positioning a cannula inserted into a body by use of an ultrasonic head which may be placed on the bodily surface, comprising the steps of: directing a light pointer mounted on the cannula to emit a light beam that extends in a distal direction and in a plane that encompasses a longitudinal axis of the cannula;inserting the cannula into the bodily surface at an angle a such that the light beam produces a light spot on the bodily surface in a distal direction in front of a puncture site of the cannula,placing the ultrasonic head on the bodily surface at the light spot. 8. The method according to claim 7, wherein the light pointer is mounted on the cannula such that the light beam from the light pointer extends axially parallel to the longitudinal axis of the cannula. 9. The method according to claim 7, wherein the light pointer produces a light spot on the bodily surface which is formed in two dimensions, and upon insertion of the cannula, further comprising the step of rotating the cannula about its longitudinal axis in such that a plane defined by the light beam from the light pointer and the cannula is oriented perpendicular to the bodily surface, and the light spot is symmetrical with respect to this plane. 10. The method according to claim 9, wherein light spot is formed in the shape of a cross. | PAJUNK GMBH & CO. KG BESITZVERWALTUNG | A61B005/06 A61N005/067
US20080287829 | A1 | 2008-11-20 | COLLECTION MEANS AND A METHOD FOR COLLECTING CORD BLOOD | 2008-04-03 | The invention provided herein comprises a kit, where the kit includes at least one collecting receptacle for collecting cord blood, and a pre-measured non-liquid anticoagulant which is optionally pre-loaded into the receptacle. The anticoagulant is preferably selected from the group consisting of heparin or heparin type additives, citrates, EDTA (ethylenediaminetetraacetic acid), and oxalates and any other anticoagulant that can be used to decrease the clotting ability of the blood. The invention provided herein also comprises a method for collecting cord blood, where the method comprises collecting a desired amount of cord blood from the umbilical cord using a collection receptacle of the invention, wherein the collection receptacle is pre-loaded with a pre-measured amount of non-liquid anticoagulant. | 1: A kit for collecting a desired amount cord blood comprising: at least one collecting receptacle for collecting cord blood; anda pre-measured non-liquid anticoagulant;wherein the anticoagulant is pre-loaded into the receptacle. 2: The kit of claim 1 wherein the collecting receptacle is selected from the group consisting of a syringe, bag, beaker and pouch. 3: The kit of claim 1 wherein the anticoagulant is selected from the group consisting of heparin, heparin type additives, citrates, EDTA (ethylenediaminetetraacetic acid), and oxalates. 4: The kit of claim 1 wherein the anticoagulant is lyophilized or in powder form. 5: The kit of claim 1 wherein the anticoagulant is lyophilized heparin or sodium citrate. 6: The kit of claim 1 wherein the collecting receptacle is a bag and the non-liquid anticoagulant is lyophilized heparin, wherein the lyophilized heparin is in the amount of 5 or about 5 heparin pellets. 7: The kit of claim 1 wherein the collecting receptacle is a syringe and the non-liquid anticoagulant is lyophilized heparin, wherein the lyophilized heparin is in the amount of 5 or about 5 heparin pellets. 8: The kit of claim 1 wherein the collecting receptacle is a syringe, the non-liquid anticoagulant is lyophilized heparin, wherein the lyophilized heparin is in the amount of 5 or about 5 heparin pellets, and the amount of cord blood collected is less than or equal to 30 ccs. 9: The kit of claim 1 further comprising at least one of the following components: one or more containers for maternal blood collection;one or more sterile wipes;one or more needles;one or more trays;one or more identification labels;one or more instructions;one or more gauze pads;one or more needle holders; andwherein all of the components of the kit are stored in a sterile container prior to use. 10: The kit of claim 1 wherein the anti-coagulant is pre-loaded in a sterile manufacturing environment. 11: A kit for collecting a desired amount of cord blood comprising: at least one collecting receptacle for collecting cord blood; anda pre-measured lyophilized anticoagulant; wherein the anticoagulant is pre-loaded into the receptacle, and wherein the anticoagulant is heparin or sodium citrate; andthe collecting receptacle is a syringe or a bag; andwherein the amount of cord blood collected is less than or equal to 30 ccs. 12: The kit of claim 11 wherein the pre-measured lyophilized anticoagulant is 5 or about 5 heparin pellets. 13: The kit of claim 11 further comprising at least one of the following components: one or more containers for maternal blood collection;one or more sterile wipes;one or more needles;one or more trays;one or more identification labels;one or more instructions;one or more gauze pads;one or more needle holders; andwherein all of the components of the kit are stored in a sterile container prior to use. 14: The kit of claim 11 wherein the anti-coagulant is pre-loaded in a sterile manufacturing environment. 15: A method for collecting ... | - | A61B005/153
US20080319343 | A1 | 2008-12-25 | Fetal blood sampling device | 2008-01-04 | A fetal scalp attachable blood-sampling device, which includes an annular tube the outer wall and the axial opening of which generally taper from an enlarged proximal end to an opposed smaller distal end. An annular ring-like space is defined in the wall of the tube and surrounds the distal end of the tube and is also open at its distal end. The tube has a continuous outer wall which smoothly engages with the outside of the annular ring-like space without presenting any sharp corners. The device is insertable by the attending physician into the vaginal canal of a patient for engaging at the distal end of the device with the scalp of a fetus while the patient is in labor. An air suction tube extends within the tube wall from a point external to the device and connects to the interior of the annular sleeve. Accordingly, a vacuum applied to the suction tube causes the distal open end of the sleeve to securely engage the fetus scalp to enable blood sampling by the physician via the internal tube. A handle is secured to the tube at the physician accessible end of the device to facilitate emplacement of the device by the physician into the vaginal canal of the patient and its proper engagement with the fetal scalp. The handle is mounted at the outer surface of the tube as to be angled with respect to the axis of said device in the direction toward the rear of the device, to further facilitate use of the handle during manipulation of the device by the physician. Vacuum application means are connected to the air suction tube at a point within the convenient reach of the physician while the physician is emplacing the device in contact with the fetus scalp. This enables the physician to control application of suction at the distal end of the device as it is emplaced at the fetus scalp. | 1. A fetal scalp attachable blood sampling device for use by an attending physician at a patient in labor, comprising: (a) an annular tube, the outer wall and the axial opening of which generally taper from an enlarged proximal end to an opposed smaller distal end;(b) an annular ring-like space being defined in the wall of the said tube and surrounding the distal end of said tube, and being open at its distal end;(c) the wall of said tube having a continuous outer surface which smoothly engages with the outside of said annular ring-like space without presenting any sharp corners;(d) the said tube being insertable by the attending physician into the vaginal canal of said patient for passing into the cervical opening and engaging at the distal end of said device with the scalp of a fetus while the patient is in labor;(e) an air suction tube extending within the wall of said tube from a point external to said device to said annular sleeve; said suction tube being in gas communication with the interior volume of said annular ring-like space, whereby a vacuum applied to said suction tube causes the distal open end of said ring-like space to securely engage the fetus scalp to enable blood sampling by said physician via the interior of the internal tube; and(f) handle means secured to the outer surface of said tube toward the physician accessible end of said device to facilitate emplacement of the device by the physician into the vaginal canal of the patient and its proper engagement with the fetal scalp. 2. A device in accordance with claim 1, further including vacuum application means connected to said air suction means at a point within the convenient reach of the physician while emplacing the device in contact with the fetus scalp, for enabling the physician to control application of suction at the distal end of said device as it is emplaced at the fetus scalp. 3. A device in accordance with claim 2, further including a light holding compartment formed in the wall of said tube at the enlarged open end of the tube for enabling insertion of a light source to illuminate the patient and fetal regions where the physician is inserting and positioning the said device. 4. A device in accordance with claim 3, wherein the said compartment is open at its end remote from the enlarged end of said tube, to facilitate insertion of a said light source in order to achieve accurate illumination at the portion of the fetal scalp which is to be pricked by the lance, and to facilitate subsequent rapid and accurate positioning of a capillary tube to draw off the blood sample from the pricked point of the scalp. 5. A device in accordance with claim 1, wherein said handle is mounted at said tube as to be angled with respect to the axis of said device in the direction toward the rear of the device, to further facilitate use of the handle during manipulation of the device by the physician. | - | A61B005/15
US07450002 | B2 | 2008-11-11 | Method and apparatus for monitoring human activity pattern | 2006-01-17 | A method and apparatus for monitoring a human activity pattern irrespective of the wearing position of the sensor unit by a user and a direction of the sensor unit are provided. The method for monitoring an inertia movement signal according to a movement of a user using a sensor unit attached to the user; detecting a direction of the sensor unit from the inertia movement signal; detecting a wearing location of the sensor unit by using acceleration and direction; determining the activity pattern of the user from inertia sensors; and delivering physical activity data corresponding to at least one caloric consumption, number of steps, and movement distance. | 1. A method for monitoring an activity pattern comprising: sensing an inertia movement signal according to a movement of a user using a sensor unit attached to the user;detecting a direction of the sensor unit from the inertia movement signal;detecting a wearing location of the sensor unit, by using the inertia movement signal and the direction; anddetermining the activity pattern of the user from the inertia movement signal by reflecting the wearing location. 2. The method of claim 1, wherein the direction is detected from a DC component of the inertia movement signal by obtaining a roll angle and a pitch angle of the sensor unit with respect to a planar surface of the sensor unit. 3. The method of claim 2, wherein the inertia movement signal is compensated for the direction by multiplying an AC component of the inertia movement signal by a rotational transform matrix defined as the pitch angle and the roll angle. 4. The method of claim 2, wherein the sensor unit comprises a tilt sensor and the roll angle and the pitch angle are detected by the tilt sensor. 5. The method of claim 2, wherein the sensor unit comprises a terrestrial magnetism sensor, and the direction is detected by further obtaining a yaw angle swept by the terrestrial magnetism sensor from a reference line on a planar surface. 6. The method of claim 2, wherein the sensor unit further includes a gyroscope and the direction is detected by further obtaining an integration of angular speed in a horizontal plane. 7. The method of claim 1, wherein detecting the wearing location comprises: by using the direction, transforming the inertia movement signal output from the sensor unit into an inertia movement signal in space fixed coordinates; andobtaining the wearing location of the sensor unit from differences between signal patterns of the transformed inertia movement signal according to locations of the sensor unit. 8. The method of claim 1, wherein determining the activity pattern comprises: extracting dynamic parameters of inertia movement at a predetermined sensor unit location; anddetermining the activity pattern from a distribution of the values of extracted dynamic parameters. 9. The method of claim 1, further comprising: analyzing the activity pattern according to the determined activity pattern. 10. The method of claim 9, further comprising: if the activity pattern is determined to be walking or running, detecting a second wearing location of the sensor unit; andif the second detected wearing location is changed from a previous wearing location, analyzing again the activity pattern at the second detected wearing location. 11. The method of claim 10, wherein in analyzing the activity pattern, at least one of a caloric consumption during the activity, a number of steps, and a moving distance is measured by using the inertia movement signal. 12. The method of claim 11, wherein the caloric consumption is measured by obtaining a vector magnitude of an integrated accelerati... | Samsung Electronics Co., Ltd. | G08B001/08 H04Q007/00 A61B005/103 A61B005/117
US07450982 | B2 | 2008-11-11 | Magnetic resonance imaging system and method | 2004-03-12 | A magnetic resonance imaging system is provided wherein measurement data (k space data) is extracted easily (and instantaneously) after imaging a plurality of times phases, and an image can be displayed in dynamic MRA measurement employing a contrast medium. Dynamic measurement is carried out, including extracting the time phase evaluation values (e.g. the origin data in k space) in respective time phases and then automatically extracting a data set including a time phase where the time phase evaluation value reaches a specified threshold value. | 1. A magnetic resonance imaging apparatus including: measurement control means for dividing a k space into a high repetitive-frequency measurement area containing an origin of the k space and measured at a high repetitive-frequency and a plurality of low repetitive-frequency measurement areas not containing the origin and measured at a low repetitive-frequency, and obtaining k space data by repeating measurement of said high repetitive-frequency measurement area and measurement of at least one of said low repetitive-frequency measurement areas between said measurements of the high repetitive-frequency measurement area in a predetermined measurement order;signal processing means for reconstructing an image by using all or a portion of the k space data as an image reconstruction set; anddisplay means for displaying the resulting image;wherein said signal processing means acquires a time phase evaluation value from said high repetitive-frequency measurement area, determines a time phase at which said time phase evaluation value reaches a predetermined threshold value or greater, andsaid measurement control means rearranges the predetermined measurement order of some of said high repetitive-frequency measurement area and said at least one of said low repetitive-frequency measurement areas in such a manner that a measurement period of said high repetitive-frequency measurement area contains said time phase. 2. A magnetic resonance imaging apparatus according to claim 1, wherein at least one low repetitive-frequency measurement area constituting said image reconstruction set is a measurement area measured immediately before or immediately after said high repetitive-frequency measurement area constituting said image reconstruction set. 3. A magnetic resonance imaging apparatus according to claim 1, wherein selection of each of said measurement areas constituting said image reconstruction set is made in such a manner as to contain the whole area of the k space. 4. A magnetic resonance imaging apparatus according to claim 1, wherein said signal processing means predicts a timing of said time phase from a time change of said time phase evaluation value, and said measurement control means controls the measurement order of each of said high repetitive-frequency measurement areas on the basis of the timing predicted. 5. A magnetic resonance imaging apparatus according to claim 1, wherein said time phase evaluation value is a substantial peak value of the k space data in said high repetitive-frequency measurement area. 6. A magnetic resonance imaging apparatus according to claim 5, wherein said threshold value is at least 1.8 times a base line value of said time phase evaluation value. 7. A magnetic resonance imaging apparatus according to claim 5 wherein said threshold value is at least 80% of a maximum value of said time phase evaluation value. 8. A magnetic resonance imaging apparatus according to claim 1, wherein said time phase evaluation value is ... | Hitachi Medical Corporation | A61B005/055
US07467003 | B2 | 2008-12-16 | Dual electrode system for a continuous analyte sensor | 2006-10-04 | Disclosed herein are systems and methods for a continuous analyte sensor, such as a continuous glucose sensor. One such system utilizes first and second working electrodes to measure additional analyte or non-analyte related signal. Such measurements may provide a background and/or sensitivity measurement(s) for use in processing sensor data and may be used to trigger events such as digital filtering of data or suspending display of data. | 1. An analyte sensor configured for insertion into a host for measuring an analyte in the host, the sensor comprising: a first working electrode disposed beneath an active enzymatic portion of a membrane;a second working electrode disposed beneath an inactive-enzymatic or non-enzymatic portion of a membrane; anda non-conductive material located between the first working electrode and the second working electrode, wherein the first working electrode and the second working electrode are intertwined. 2. The sensor of claim 1, further comprising a reference electrode, wherein the reference electrode is configured to provide electrical conductance and structural support. 3. The sensor of claim 1, further comprising a reference electrode, wherein the reference electrode is configured to provide electrical conductance and a diffusion barrier. 4. The sensor of claim 1, further comprising at least one of a reference electrode and a counter electrode. 5. The sensor of claim 4, wherein the at least one of the reference electrode and the counter electrode, together with the first working electrode and the second working electrode, integrally forms at least a portion of the sensor. 6. The sensor of claim 4, wherein the at least one of the reference electrode and the counter electrode is located at a position remote from the first working electrode and the second working electrode. 7. The sensor of claim 4, wherein a surface area of at least one of the reference electrode and the counter electrode is at least six times a surface area of at least one of the first working electrode and the second working electrode. 8. The sensor of claim 1, wherein each of the first working electrode and the second working electrode comprises a wire with a diameter from about 0.001 inches to about 0.010 inches. 9. The sensor of claim 1, wherein each of the first working electrode and the second working electrode comprises a bulk material or plated wire. 10. The sensor of claim 1, wherein the non-conductive material is conformally coated around a circumference of a substantial portion of at least one of the first working electrode and the second working electrode. 11. The sensor of claim 1, wherein the non-conductive material comprises an insulating material selected from the group consisting of: parylene, fluorinated polymers, polyethyleneterephthalate, PTFE, ETFE, polyurethane, polyethylene, polyimide, and silicone. 12. The sensor of claim 1, wherein the sensor comprises a width of no more than about 0.015 inches. 13. The sensor of claim 1, wherein the first and second working electrodes are at least one of bundled, coiled and twisted. 14. The sensor of claim 13, further comprising a non-conductive material surrounding at least a portion of the first and second working electrodes. 15. The sensor of claim 1, wherein, wherein the non-conductive material is configured to provide an insulative property and structural support. 16. The sensor of claim 1, wherein the me... | DexCom, Inc. | A61B005/00
US07465277 | B2 | 2008-12-16 | System and methods for conducting multiple diagnostic hearing tests | 2003-05-15 | Systems and methods for conducting multiple diagnostic hearing tests are provided. The systems comprise an insertion probe for insertion into an ear of a patient, at least one transducer for providing test sounds to the ear of the patent through the insertion probe, and at least one transducer for transmitting test results from the ear of the patient. The systems are capable of conducting an audiometric test and at least one of an acoustic immittance test and an otoacoustic emission test. More preferably, the systems are capable of conducting an audiometric test, a tympanogram test, an acoustic reflex test, and an otoacoustic emission test using a single ear probe. | 1. A system for conducting an audiometric hearing test, an otoacoustic emission hearing test and an acoustic immittance hearing test in a first ear of a patient, said system comprising: a first insertion probe comprising a sealing surface for engaging the external auditory canal of said first ear and providing an airtight seal within said canal;at least one speaker for providing to said first ear through said first insertion probe an audiometric test sound, an otoacoustic emission test sound, and an acoustic immittance test sound;a reversible compressor for delivering a gas to and removing a gas from said first ear through said first insertion probe;a first microphone for receiving an otoacoustic emission sound from said first ear through said first insertion probe; anda second microphone for receiving an acoustic immittance test result sound from said first ear through said first insertion probe. 2. The system of claim 1 wherein said otoacoustic emission hearing test comprises a distortion product otoacoustic emission test, said otoacoustic emission test sound comprises a sound having a first frequency and a second frequency, and said first microphone is a microphone for receiving a distortion product otoacoustic emission. 3. The system of claim 1 wherein said acoustic immittance hearing test comprises a test selected from the group consisting of a tympanogram and an acoustic reflex test. 4. The system of claim 1 wherein said acoustic immittance test result sound is selected from the group consisting of a tympanogram test result sound and an acoustic reflex test result sound. 5. The system of claim 1 further comprising at least one ambient noise microphone for receiving ambient noise from the patient environment. 6. The system of claim 3 wherein said acoustic immittance hearing test comprises both a tympanogram test and an acoustic reflex test. 7. The system of claim 1 further comprising a second insertion probe element for insertion in a second ear of the patient, and wherein said reversible compressor is coupled to said second insertion probe element and is capable of delivering a gas to and removing a gas from said second ear though said second insertion probe. 8. The system of claim 7, further comprising: at least one speaker for providing to said second ear though said second insertion probe an audiometric test sound, an otoacoustic emission test sound, and an acoustic immittance test sound;a third microphone for receiving an otoacoustic emission sound from said second ear through said second insertion probe; anda fourth microphone for receiving an acoustic immittance test result sound from said second ear though said second insertion probe. 9. The system of claim 1 further comprising a bone vibrator for conducting a bone conduction hearing test. 10. A system for conducting an audiometric hearing test and a second hearing test selected from the group consisting of an otoacoustic emission hearing test and an acoustic immittance hear... | Tympany, LLC | A61B005/00
US07462153 | B2 | 2008-12-09 | Method and system for modeling cardiovascular disease using a probability regession model | 2004-07-23 | A method and system for modeling cardiovascular disease using a probability regression model is provided. A parameter estimate of a probability regression model for cardiovascular disease can be generated using predictors derived from cardiovascular sound signals and disease status information. A probability of cardiovascular disease can be generated using a probability regression model that includes a predictor derived from cardiovascular sound signals. | 1. A processor-readable medium comprising code representing instructions to cause a processor to: generate at least one parameter estimate of a probability regression model for cardiovascular disease at least partially based on (1) a plurality of predictors generated from cardiovascular sound signals of a plurality of subjects and (2) disease status information for cardiovascular disease of the plurality of subjects;determine a probability of cardiovascular disease for a new subject with the probability regression model, the probability of cardiovascular disease being generated at least partially based on the at least one parameter estimate and another plurality of predictors generated from cardiovascular sound signals of the new subject; andupdate the at least one parameter estimate at least partially based on disease status information associated with the new subject. 2. The processor-readable medium of claim 1, further comprising code representing instructions to cause a processor to: determine a probability of cardiovascular disease for a second new subject with the probability regression model, the probability of cardiovascular disease for the second new subject being generated at least partially based on the updated at least one parameter estimate. 3. A processor-readable medium comprising code representing instructions to cause a processor to: generate a probability of cardiovascular disease using a probability regression model:Pr⁡(Y❘X,b)=12⁢π⁢∫s=-∞X′⁢b⁢exp(-s22)⁢ⅆs;wherein X is at least one predictor generated from cardiovascular sound signals;wherein b is at least one parameter estimate previously generated from the probability regression model based on cardiovascular sound signals of a plurality of subjects; andwherein Pr(Y¦X, b) is the probability of the existence of cardiovascular disease Y given the at least one predictor X and the at least one previously generated parameter estimate b. 4. The processor-readable medium of claim 3, wherein the at least one previously generated parameter estimate b is previously generated from the probability regression model also based on clinical data of the plurality of subjects. 5. A processor-readable medium comprising code representing instructions to cause a processor to: generate at least one parameter estimate of a probability regression model for cardiovascular disease at least partially based on (1) a predictor generated from cardiovascular sound signals of at least one subject and (2) disease status information for cardiovascular disease;wherein the probability regression model is one of a logit model, a multinomial probit model, a multinomial logit model, an ordered probit model, an ordered logit model, a Weibull model, a Cox proportional hazards model, an exponential model, a log-logistic model, a lognormal model, and a Kaplan-Meier model. 6. The processor-readable medium of claim 5, further comprising code representing instructions to cause a processor to: update the at least one pa... | Sonomedica, Inc. | A61B005/02
US07450980 | B2 | 2008-11-11 | Intracorporeal substance measuring assembly | 2005-03-25 | An intracorporeal substance measuring assembly to be provided in an embedded-type substance sensor for detecting and measuring an intercorporeal analyte includes: a detection layer containing at least one fluorescent indicator for generating fluorescence according to the concentration of the analyte; and an optical separation layer which is provided on the detection layer, is optically opaque, permits the analyte to penetrate therethrough, and prevents the penetration therethrough of at least one of living body substances possibly deteriorating the detection layer and/or obstructing the fluorescence. | 1. An intracorporeal substance measuring assembly provided in an embedded type substance sensor for detecting and measuring an intracorporeal analyte, comprising: a detection layer containing at least one fluorescent indicator containing phenyl boronic acid that generates fluorescence according to the concentration of said analyte; andan optical separation layer on one side of said detection layer that is optically opaque and permits said analyte to penetrate therethrough;wherein said analyte is a saccharide, and said fluorescent indicator is represented by formula (1):where Q is a fluorescent residual group; at least one of R1, R2, and R3 is an active group bound to a base material of said detection layer; R4, R5, R6, and R7 are each at least one substituent group selected from the group consisting of hydrogen atom, alkyl, alkenyl, allyl, allylalkenyl, substituted alkyl, oxyalkyl, acyl, and halogen atom. 2. The intracorporeal substance measuring assembly as set forth in claim 1, wherein Q is a fluorescent residual group containing anthracene. 3. The intracorporeal substance measuring assembly as set forth in claim 1, wherein said optical separation layer is comprised of an opaque substance and a base material for supporting said opaque substance. 4. The intracorporeal substance measuring assembly as set forth in claim 3, wherein said opaque substance is comprised of at least one substance selected from the group consisting of carbon black, fullerene, and carbon nanotube. 5. The intracorporeal substance measuring assembly as set forth in claim 1, wherein said optical separation layer is a polymeric porous film formed by a phase transition method. 6. The intracorporeal substance measuring assembly as set forth in claim 1, wherein said optical separation layer does not permit intracorporeal components that deteriorate said detection layer to penetrate therethrough. 7. The intracorporeal substance measuring assembly as set forth in claim 6, wherein said optical separation layer inactivates said intracorporeal components. 8. A method of producing an intracorporeal substance measuring assembly provided in an embedded type substance sensor for detecting and measuring an intracorporeal saccharide analyte the method comprising: providing a detection layer containing at least one phenyl boronic acid-containing fluorescent indicator that fluoresces according to the concentration of said analyte; and an optical separation layer disposed on one side of said detection layer that is optically opaque and permits said analyte to penetrate therethrough, wherein said fluorescent indicator is represented by formula (1):where Q is a fluorescent residual group: at least one of R1, R2, and R3 is an active group to be bound to said base material of said detection layer: R4, R5, R6, and R7 are each at least one substituent group selected from the group consisting of hydrogen atom, alkyl, alkenyl, allyl, allylalkenyl, substituted alkyl, oxyalkyl, acyl, and halogen... | Terumo Kabushiki Kaisha | A61B005/1455
US07768387 | B2 | 2010-08-03 | Method and apparatus for providing dynamic multi-stage signal amplification in a medical device | 2008-04-14 | Methods and apparatus for providing multi-stage signal amplification in a medical telemetry system are provided. | 1. An apparatus, comprising: a first amplifier having at least one input terminal and an output terminal, the at least one input terminal coupled to a signal source, the output terminal configured to provide a first output signal;a second amplifier having at least one input terminal and an output terminal, the at least one input terminal coupled to the output terminal of the first amplifier, the output terminal of the second amplifier configured to provide a second output signal;a processor operatively coupled to receive the first output signal and the second output signal;an analog to digital (A/D) converter coupled to the output terminals of the first and second amplifiers; anda processor operatively coupled to the A/D converter for processing the one or more signals received at the one or more first amplifier output terminal and the second amplifier output terminal; wherein the first output signal is a predetermined ratio of the second output signal;wherein the first output signal and the second output signal are associated with a monitored analyte level of a user;wherein the processor is configured to compare the one or more signals received at the one or more first amplifier output terminal and the second amplifier output terminal to a predetermined threshold value; and furtherwherein the processor is configured to process a signal associated with one of the one or more signals received at the one or more first amplifier output terminal and the second amplifier output terminal when another signal associated with the other one of the one or more signals received at the one or more first amplifier output terminal and the second amplifier output terminal exceeds the predetermined threshold value. 2. The apparatus of claim 1 wherein the first amplifier includes a transimpedance amplifier. 3. The apparatus of claim 1 wherein the monitored analyte level includes glucose level. 4. The apparatus of claim 1 wherein the at least one input terminal of the first amplifier includes an inverting input terminal. 5. The apparatus of claim 4 further including a reference signal source coupled to a noninverting input terminal of the first amplifier. 6. The apparatus of claim 1 wherein the second amplifier includes a gain of approximately two. 7. The apparatus of claim 1 wherein the first output signal is associated with a signal level from the signal source. 8. The apparatus of claim 1 wherein the analog to digital (A/D) converter is a 12 bit A/D converter. 9. The apparatus of claim 1 wherein the predetermined threshold value includes approximately 4,000 bits. 10. The apparatus of claim 1 wherein the signal source is configured to receive an analyte sensor signal. 11. The apparatus of claim 10 wherein the analyte sensor signal corresponds to a monitored analyte level. 12. The apparatus of claim 10 wherein the analyte sensor signal includes a glucose sensor signal. 13. The apparatus of claim 1 wherein the signal source is operatively coupled to an ... | Abbott Diabetes Care Inc. | G08B029/00
US07455648 | B2 | 2008-11-25 | Living body inspection apparatus | 2004-03-01 | A living body inspection apparatus including an oscillation coil which passes an AC current, a detection coil which detects an AC magnetic field generated from the detection coil, an amplification circuit which amplifies a voltage generated by the magnetic field induced by the detection coil, a detecting unit for detecting the output signal of the amplification circuit, a low pass filter to which the output signal of the detecting unit is input, a unit for setting the oscillation coil and detection coil in first and second regions of the living body, a recording unit for recording the output of the low pass filter while the first region and the second region of the living body are moving and a displaying unit for displaying the data recorded in the recording unit or results of analysis of the recorded data. | 1. A living body inspection apparatus comprising: an AC generation circuit which generates an AC voltage having a predetermined frequency;converting means for converting said AC voltage output from said AC generation circuit to an AC current;an oscillation coil which passes said AC current output from said converting means;a detection coil which detects an AC magnetic field generated from said oscillation coil;an amplification circuit which amplifies a voltage generated by the magnetic field induced by said detection coil;detecting means for detecting the output signal of said amplification circuit with the same frequency as that of said AC generation circuit or a double harmonic frequency;a low pass filter to which the output signal of said detecting means is input;means for setting said oscillation coil in a first region of the living body;means for setting said detection coil in a second region of said living body;recording means for recording the output of said low pass filter while said first region and said second region of said living body are moving; anddisplaying means for displaying the data recorded in said recording means or results of an analysis of said recorded data, whereinsaid output of said low pass filter contains the continuous information related to the distance between the first region and the second region of the living body;calculating means for converting a voltage value of said data recorded in said recording means to a distance corresponding to the distance between said first region and said second region;primary differentiating means for primary differentiating the waveform of the voltage value of said data recorded in said recording means or the waveform obtained by said calculating means;secondary differentiating means for secondary differentiating the waveform of the voltage value of said data recorded in said recording means or the waveform obtained by said calculating means; andsaid display means further configured for displaying results obtained in said primary and secondary differentiating means as information of velocity and acceleration on said display means. 2. The living body inspection apparatus according to claim 1, wherein said oscillation coil and said detection coil are configured to be attached to different fingers. 3. The living body inspection apparatus according to claim 1, further comprising calculating means for converting a voltage value of said data recorded in said recording means to a distance corresponding to the distance between said first region and said second region, wherein said distance is displayed on said displaying means. 4. The living body inspection apparatus according to claim 1, further comprising timing detecting means for detecting tapping timing from the output waveform of said primary differentiating means, wherein a waveform of said detected timing is displayed. 5. The living body inspection apparatus according to claim 4, further comprising peak detecting means for dete... | Hitachi, Ltd. | A61B005/103 A61B005/05
US07911209 | B2 | 2011-03-22 | Head coil and neurovascular array for parallel imaging capable magnetic resonance systems | 2005-02-22 | A head coil for use with a parallel-imaging compatible MR system is disclosed, as is a method of making, and a neurovascular array (NVA) equipped with, same. The head coil includes conductive rings and rods configured to produce a plurality of electrically-adjacent primary resonant substructures about a birdcage-like structure, with each such primary resonant substructure including two rods neighboring each other and the short segment of each of the first and second rings interconnecting them. The primary resonant substructures are isolated from each other via a preamplifier decoupling scheme and an offset tuning scheme thereby enabling each primary resonant substructure (i) to receive an MR signal from tissue within its field of view and (ii) to be operatively couplable to one processing channel of the MR system for conveyance of the MR signal received thereby (iii) while being simultaneously decoupled from the other primary resonant substructures. | 1. A neurovascular array configured for use with a magnetic resonance (MR) system capable of parallel-imaging via a plurality of processing channels, the neurovascular array comprising: (a) a head coil having: (I) a first electrically conductive ring,(II) a second electrically conductive ring, and(III) a plurality of rods electrically interconnecting said first and said second rings in order to form a birdcage-like structure therewith,wherein said rods and said first and said second rings are configured to produce a plurality of electrically-adjacent primary resonant substructures about the birdcage-like structure, with each of said primary resonant substructures: (A) constituting a coil element including two of said rods neighboring each other and a corresponding short segment of each of said first and said second rings interconnecting them while electrically sharing one of said rods with each of its neighboring primary resonant substructures, so that each of said primary resonant substructures is enabled to receive magnetic resonance signals from tissue within a separate field of view thereof, and(B) providing a source impedance considerably higher than a load impedance to which said primary resonant substructure connects creating a resonant circuit therewith so as to enable said primary resonant substructure (i) to be operatively couplable to one processing channel of the MR system for conveyance of the magnetic resonance signals received thereby (ii) while simultaneously being at least partially decoupled from the other of said primary resonant substructures of said head coil;(b) an anterior coil having in proximity to said head coil at least one other coil element configured for receiving magnetic resonance signals from tissue within a separate field of view thereof;(c) a posterior coil having in proximity to said head coil at least one other coil element configured for receiving magnetic resonance signals from tissue within a separate field of view thereof; and(d) an interface configured for enabling said coil elements of said head coil, said anterior coil and said posterior coil to be selectively interconnected to the processing channels of the MR system so that the neurovascular array is selectively operable in a plurality of modes. 2. The neurovascular array of claim 1 wherein said plurality of modes includes a neurovascular phased array mode in which said interface enables: (a) each pair of said coil elements of said head coil to be interconnected with a separate one of the processing channels of the MR system; and(b) each of said other coil elements of said anterior and said posterior coils to be interconnected with a separate one of the processing channels of the MR system. 3. The neurovascular array of claim 2 wherein said head coil has eight of said coil elements and said anterior and said posterior coils each have two of said other coil elements, configured for use with the MR system equipped with at least eight of the processing... | Medrad, Inc. | G01V003/00 A61B005/05
US07894887 | B2 | 2011-02-22 | Biosignal measuring equipment | 2008-03-26 | A signal-to-noise ratio and measurement precision is increased in electrode units disposed on the left and right sides of a steering wheel. A plurality of electrode units is disposed in the left and right handholds of a steering wheel. The contact impedances of all the electrode units are measured. A pair of left and right electrode units to be used to measure an electrocardiographic signal is designated from among the electrode units whose measured contact impedances are less than or equal to a first threshold. The results of measurement of an electrocardiographic signal by the designated electrode units are added in order to minimize noise. An electrode unit with high contact impedance is used to measure induction noise and remove the induction noise component from the electrocardiographic signal measurement result. | 1. Biosignal measuring equipment comprising: a plurality of electrode units disposed in left and right positions of a steering wheel; anda measuring means for measuring a biosignal associated with a subject using designated ones of the plurality of electrode units, the subject capable of holding the steering wheel using the left and right positions,wherein the measuring means measures a plurality of biosignals using the designated ones of the plurality of electrode units and determines final information on the basis of the plurality of measured biosignals, whereineach of the plurality of electrode units includes a first electrode and a second electrode; andthe measuring means includes: a differential amplification means for producing a biosignal by amplifying a difference between two input signals;an impedance measurement means for measuring an impedance between two electrodes;a switching circuit unit that connects the first and second electrodes of the each of the each of the plurality of electrode units to one of the differential amplification means, the impedance measurement means, and a reference potential; andan arithmetic control means for measuring a plurality of biosignals from the subject by connecting the electrodes of each of the electrode units to the one of the differential amplification means, impedance measurement means, and the reference potential via the switching circuit unit, and for deducing a final complex. 2. The biosignal measuring equipment according to claim 1, wherein the plurality of electrode units are disposed in the left and right positions such that a left palm and a right palm of the subject simultaneously come into contact with multiple ones of the plurality of electrode units when the subject holds the steering wheel. 3. The biosignal measuring equipment according to claim 2, wherein left ones and right ones of the plurality of electrode units are disposed within respective left and right circular areas having a diameter is 15 cm or less, around the left and right positions respectively. 4. The biosignal measuring equipment according to claim 2, wherein: the plurality of electrode units disposed near left and right joint portions where a ring portion and left and right spoke portions of the steering wheel are joined, the plurality of electrode units disposed in the circumferential direction of the ring portion; andat least one of the plurality of electrodes is formed to protrude from the ring portion toward the spoke portion. 5. The biosignal measuring equipment according to claim 1, wherein the measuring means is mounted in the steering wheel. 6. The biosignal measuring equipment according to claim 1, wherein the measuring means includes a plurality of differential amplification means, a number of the differential amplification means corresponding to a number of the plurality of electrode units disposed in each of the handholds. 7. The biosignal measuring equipment according to claim 6, wherein the measuri... | Denso Corporation | A61B005/04
US08010177 | B2 | 2011-08-30 | Intraoperative image registration | 2007-04-24 | A system to allow registration of a patient to image space is disclosed. The registration can be performed without touching a plurality of fiducial points on the patient or in the image data. The registration process can eliminate manual steps of image to patient registration and image to image registration. | 1. A navigation system for registering an image space to a physical space relative to a patient for a surgical procedure, comprising: a tracking system including a localizer and a tracking device;a system operable to access intraoperative image data of the patient that was acquired at a determined position relative to the patient in the physical space;a processor operable to at least one of determine a position of the tracking device in the physical space, determine an intraoperative image data fiducial point in the intraoperative image data, determine a preoperative image data fiducial point in a preoperative image data, correlate the intraoperative image data and the preoperative image data with the intraoperative image data fiducial point and the preoperative image data fiducial point, or combinations thereof;wherein the preoperative image data is operable to be registered to the physical space to correlate points in the preoperative image data with corresponding points in the physical space based upon the intraoperative image data that was acquired at the determined position relative to the patient;wherein the registered preoperative image data allow navigation of the tracking device in the physical space relative to the preoperative image data;wherein the processor is operable to determine the position of an intraoperative imaging system to determine a position relative to the determined intraoperative fiducial image data point in the physical space. 2. The system of claim 1, further comprising: a preoperative imaging system operable to acquire the preoperative image data of the patient; andan intraoperative imaging system operable to acquire the intraoperative image data of the patient. 3. The system of claim 2, wherein the intraoperative imaging system and the preoperative imaging system are a single imaging system. 4. The system of claim 1, further comprising: a display device operable to display at least one of the intraoperative image data, the preoperative image data, or combinations thereof. 5. The system of claim 4, further comprising: an instrument, wherein the tracking device is operable to track an instrument position. 6. The system of claim 5, wherein the instrument is at least one of a deep brain stimulation electrode, a micro-electrode, a catheter, an ablation catheter, or combinations thereof. 7. The system of claim 4, wherein the display device is operable to display an icon representing an instrument superimposed on the registered preoperative image data for the navigation of the tracking device. 8. The system of claim 1, further comprising: a fiducial marker operable to be imaged and form the intraoperative image data fiducial point and the preoperative image data fiducial point. 9. The system of claim 1, further comprising: a second processor operable to at least one of determine a position of the tracking device in the physical space, determine an intraoperative image data fiducial point in the intraoperative imag... | Medtronic, Inc. | A61B005/05
US08280136 | B2 | 2012-10-02 | Method and apparatus for detecting intraventricular dyssynchrony | 2006-09-18 | A method and apparatus for detecting intraventricular dyssynchrony is provided. In one embodiment, the method includes: a) receiving information from an observer viewing a source image to define a plurality of points associated with muscle boundaries of a ventricle, b) extracting muscle boundaries for a series of source images to define a corresponding series of contour shape images using a Fourier descriptor model, and c) deforming each contour shape based at least in part on image forces, external forces, and internal forces to track movement of the muscle boundaries over time. In another embodiment, the method includes: a) characterizing the series of source images as a corresponding series of contour shape images representative of movement of muscle boundaries of the ventricle over time, b) classifying the series of contour shape images in a dyssynchronous class or a non-dyssynchronous class, and c) pre-diagnosing the ventricle as dyssynchronous or non-dyssynchronous. | 1. A method of detecting intraventricular dyssynchrony for a ventricle of a beating heart, the method including: a) receiving initialization information from an observer viewing a first source image of a sequence of source images of a ventricle of a beating heart over at least one complete cardiac cycle to define a single vector constructed by concatenating Fourier shape descriptors extracted from a plurality of points associated with muscle boundaries of the ventricle over time;b) extracting muscle boundaries for each source image of the sequence of source images to define a corresponding spatio-temporal characterization of a complete source image sequence with contour shapes corresponding to the muscle boundaries using a Fourier descriptor model; andc) deforming the single vector of Fourier shape descriptors representing each contour shape in the complete source image sequence based at least in part on image forces, external forces, and internal forces to track movement of the muscle boundaries for the ventricle over time in the complete source image sequence. 2. The method of claim 1 wherein the initialization information includes at least four points. 3. The method of claim 1 wherein the first source image corresponds to the end of diastole for the cardiac cycle. 4. The method of claim 1 wherein the muscle boundaries include an endocardial wall and an epicardial wall. 5. The method of claim 1 wherein the set of source images include short-axis cine magnetic resonance images. 6. The method of claim 1 wherein the Fourier descriptor model includes:s⁡(t)=x⁡(t)+jy⁡(t)=∑k=-NN⁢Sk⁢ⅇj⁢⁢kt,with⁢⁢j=-1,where each complex number Sk=ak+jbk, with ak, bkεR, is treated as a rotating phasor with magnitude ¦Sk¦=√{square root over (ak2+bk2)} and phase angleϕk=tan-1⁡(bkak). where: x(t) and y(t) are projections, with t ε[0,2π]; andak and bk are real numbers. 7. The method of claim 6 wherein each pair of descriptors Sk and S−k describes an ellipse. 8. The method of claim 1 wherein the image forces include edge-based forces and texture-based forces. 9. The method of claim 8 wherein the edge-based forces have a limited zone of attraction. 10. The method of claim 1, the deforming in c) associated with image forces further including; d) defining edges of the contour shapes in the complete contour shape image sequence using an optimal zero crossing operator model to determine edge-strength vectors and gradient vectors for associated pixels of the contour shape images;e) deforming edges of the contour shapes in the complete contour shape sequence using an averaging mask in conjunction with a range of pixels along, the edges defined by the optimal zero crossing operator model;f) determining a length of the contour shapes in the complete contour shape sequence using an approximation model; andg) determining edge-based image forces for the contour shapes the complete contour shape sequence based at least in part on the determined length for the corresponding contou... | The Ohio State University | G06K009/00 A61B006/00
US08357095 | B2 | 2013-01-22 | Non-invasive treatment of fascia | 2006-10-18 | Various methods and devices for minimally invasive treatment and prevention of conditions of the fascia are provided. In one aspect, a method includes providing an acoustic wave source effective to deliver a focused acoustic wave to a target site within a patient's body focusing an acoustic wave through a patient's skin such that at least one location in the patient's fascia is fenestrated in a desired pattern. | 1. A method for non-invasive treatment and prevention of fascial conditions, comprising: focusing an acoustic wave to a target site of one or more muscles within a patient's body to treat compartment syndrome, the one or more muscles having fascia disposed therearound; andapplying the acoustic wave to the target site through a patient's skin to fenestrate one or more cuts in a desired pattern in at least one location in the patient's fascia. 2. The method of claim 1, further applying a medium to the patient's skin adjacent to the target site. 3. The method of claim 1, further comprising non-destructible heating of tissue adjacent to the target site. 4. The method of claim 1, further comprising imaging the fascia using an imaging source that is integrated with the acoustic wave source. 5. The method of claim 4, wherein the imaging source is an ultrasound source. 6. The method of claim 1, wherein the acoustic wave is delivered to the target site at a frequency in the range of about 0.5 MHz to about 5 MHz. 7. The method of claim 1, wherein the acoustic wave is delivered to the target site at a frequency in the range of about 1 MHz. 8. The method of claim 1, wherein the acoustic wave source is focused to form a focal spot in the range of about 0.3 mm to about 3.0 mm. 9. The method of claim 1, further comprising adjusting a peak power of the acoustic source in the range of about 10 W to about 100 W prior to fenestration of the fascia. 10. The method of claim 1, wherein the frequency of the acoustic wave source is adjusted to the range of about 10 W to about 20 W prior to fenestration of the fascia. 11. A method for non-invasive treatment and prevention of compartment syndrome, comprising: inducing cavitation into a target region of one or more muscles within a patient's body, the one or more muscles having fascia disposed therearound, thus forming one or more cuts in the fascia using an acoustic treatment source to treat compartment syndromefocusing the acoustic treatment source; andadjusting a frequency of the acoustic treatment source such that the fascia is penetrated. 12. The method of claim 11, further comprising imaging the target region using an imaging device that is integrated with the acoustic treatment source. 13. The method of claim 12, wherein the imaging device is an ultrasound source. 14. The method of claim 11, further comprising moving the acoustic treatment source such that cuts or a pattern of fenestrations are formed on the target region. 15. The method of claim 11, wherein the step of inducing cavitation further includes generating acoustic waves in the range of about 0.5 MHz to about 5 MHz. 16. The method of claim 11, wherein the step of inducing cavitation further includes generating acoustic waves in the range of about 1 MHz. 17. The method of claim 11, wherein the treatment source is focused to form a focal spot on the target region in the range of about 0.3 mm to about 3.0 mm. 18. The method of claim 11, wher... | The General Hospital Corporation | A61B008/00
US09289149 | B2 | 2016-03-22 | Apparatus and method for wirelessly transferring and storing medical data | 2007-05-10 | A medical diagnostic/monitoring system is disclosed herein. The medical diagnostic/monitoring system includes a data acquisition device having a first near field communication device, and a data storage device wirelessly connected to the data acquisition device. The data storage device includes a second near field communication device. The first near field communication device and the second near field communication device are collectively configured to wirelessly transfer power and data from the data acquisition device to the data storage device. | 1. A method for wirelessly transferring physiological data to a data storage device comprising: obtaining physiological data with a transducer of a data acquisition device, the transducer configured to obtain the physiological data from a body of a patient;generating an electromagnetic field between a first near field communication device of the data acquisition device and a second near field communication device of a data storage device;transferring power from the data acquisition device to the data storage device across the electromagnetic field;transferring the physiological data from the data acquisition device to the data storage device across the electromagnetic field andstoring the physiological data at the data storage device. 2. The method of claim 1, further comprising implementing the electromagnetic field to wirelessly transfer information from a radio frequency identification memory card of the data storage device to the data acquisition device. 3. The method of claim 1, wherein the transducer is an electrode and the physiological data is a cardiac electrical signal. 4. The method of claim 3, further comprising implementing the electrode to generate an analog signal proportional to the sensed cardiac electrical signal, and providing the analog signal from the electrode to the data acquisition device. 5. The method of claim 1, further comprising: storing the physiological data at a radio frequency identification memory card of the storage device. 6. The method of claim 1, wherein the data storage device comprises a radio frequency identification memory card comprising a plurality of alphanumeric characters configured to visually convey a selectable type of information. 7. The method of claim 1, wherein the data storage device comprises a radio frequency identification memory card comprising graphic image configured to visually convey information pertaining to the identity of a patient. | General Electric Company | A61B005/00 A61B005/0432 A61N001/372
US20080269578 | A1 | 2008-10-30 | Method for Detecting an Analyte Within the Body of a Patient or an Animal | 2006-06-30 | The invention relates to a method for detecting an analyte within tissue of a patient or an animal. In the laboratory medicine it is often necessary to determine the amount of analytes within body fluids such as blood. For that purpose a blood withdrawal is carried out by experienced medical personnel, and the blood is analysed in the laboratory. This ex-vivo measurement of analytes is uncomfortable for the patient. Moreover, there is an unwanted delay between the blood withdrawal and the time at which the measurement values are obtained from the laboratory. In order to avoid these disadvantages it is suggested to detect analytes 1 in the body fluids by using laser-induced breakdown spectroscopy (LIBS). In this way it is possible to generate a plasma 6 within tissue, e.g. skin 3, of the body 2 in a certain depth below the surface of the skin 3, to detect the plasma light 16, and to extract the desired information from an analysis of the detected plasma light. This method can be carried out by the patient at home making it more acceptable to him or her. | 1. Method for detecting an analyte (1) within tissue of a patient (2) or of an animal by means of LIBS, the method comprising the following steps: a) exposing tissue (3) of the patient or the animal with electromagnetic radiation (4),b) the electromagnetic radiation having a wavelength such that it is transmitted by the outmost layer (5) of the skin,c) the intensity of the electromagnetic radiation being chosen to generate a plasma (6) below the outmost layer of the skin,d) detecting the electromagnetic radiation (16) emitted by the plasma,e) identifying at least one spectral line characteristic of the analyte. 2. Method according to claim 1, characterized in that the analyte is contained in a body fluid within the tissue, and whereby the body fluid is preferably blood. 3. Method according to claim 1, characterized in that the electromagnetic radiation has a wavelength in the near infrared region. 4. Method according to claim 3, characterized in that the electromagnetic wavelength is between about 600 nm and 1200 nm, preferably between about 800 nm and 1200 nm. 5. Method according to claim 1, characterized in that the intensity is above 5×1011 W/cm2. 6. Method according to claim 1, characterized in that the pulse length is within the femtosecond to picosecond range. 7. Method according to claim 1, characterized in that it is applied for the detection and/or prediction of arrhythmias. 8. Method according to claim 1, characterized in that it is applied for diuretic therapy. 9. Method according to claim 1, characterized in that it is used for the diagnosis of heart attacks. 10. Method according to claim 1, characterized in that it is applied for the diagnosis of micronutrition deficiencies. 11. Use of an apparatus for a laser-induced breakdown spectroscopy (LIBS) for the in-vivo detection of an analyte in the body of a patient or an animal. 12. Use according to claim 11, characterized in that the analyte is contained in a body fluid, in particular in blood. | KONINKLIJKE PHILIPS ELECTRONICS , N.V. | A61B005/1455
US20080287749 | A1 | 2008-11-20 | Method and Apparatus for Remote Patient Monitoring | 2006-10-30 | A method and apparatus for patient monitoring are described. The method and apparatus include gathering data from a patient medical device (104) and providing the data for analysis. Messages may be provided to the patient via a patient terminal (103). | 1. An apparatus, comprising: a patient terminal (103);a medical device (104) adapted to garner measurements from a patient and to transmit data from the measurements;a clinician terminal (102) adapted to receive manual inputs or audio inputs, or both; anda server (101) adapted to receive the data and the inputs, wherein the server is operative to transfer the inputs to the patient terminal. 2. An apparatus as recited in claim 1, wherein the patient terminal (103) comprises a stationary device (107) and a remote device (108) adapted to communicate with the stationary device. 3. An apparatus as recited in claim 1, wherein the clinician terminal comprises a stationary device (109) and a remote device (110) adapted to communicate with the remote device (110). 4. An apparatus as recited in claim 1, wherein the server (101) further comprises a processor (200) and the processor (200) is adapted to convert the manual and the audio inputs into text and to transmit the text to the patient terminal (103). 5. An apparatus as recited in claim 1, wherein the medical device (104) is adapted to transmit the data to the patient terminal (103) and the patient terminal (103) is adapted to transmit the data to the server (101) or the medical device (104) is adapted to transmit the data to the server (101) without transmitting the data to the patient terminal (103). 6. (canceled) 7. An apparatus as recited in claim 1, wherein the server (101) is adapted to analyze the data and, based on the analysis, to transmit a message to the patient terminal. 8. An apparatus as recited in claim 1, wherein the data include a time and a date, or a signal quality, or both of each of the measurements. 9. A method, comprising: measuring a vital sign;transmitting data from the measuring to a server (101);transmitting the data from the server to a clinician terminal (102);based on the data, inputting a message to the clinician terminal (102); andproviding the message at the patient terminal (103). 10. A method as recited in claim 9, wherein the measuring vital sign further comprises measuring a time and a date of the measuring. 11. A method as recited in claim 10, wherein the inputting the message further comprises one of either: reviewing the data and inputting a text message at the clinician terminal (102) or reviewing the data and inputting a voice message at the clinician terminal (102). 12. (canceled) 13. A method as recited in claim 9, wherein the transmitting to the server (101) further comprises one of either transmitting the data to the patient terminal (103) and transmitting the data from the patient terminal (103) to the server or not transmitting the data to the patient terminal (103). 14. (canceled) 15. A method as recited in claim 9, wherein the transmitting to the clinician terminal (102) further comprises transmitting the data to a stationary device (109) of the clinician terminal (102), and transmitting the data from the stationary device (109) to a remote... | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61B005/00
US20080242955 | A1 | 2008-10-02 | RELIABILITY IN DETERMINATION OF CLINICAL STATE OF A SUBJECT | 2007-03-30 | The invention relates to the determination of a diagnostic index indicative of the clinical state, especially nociceptive state, of a subject. In order to improve the quality of the determination process, desired input signals derived from the subject are monitored and the quality of the process is improved based on the monitoring. This is implemented by increasing user awareness of the reliability of the index or by controlling the measurement set-up. Quality information indicative of the reliability of the current index values may be produced based on the input signals. The quality information may then be employed to give an indication of the current reliability of the index and a warning if the reliability of the diagnostic index becomes compromised. | 1. A method for ascertaining the clinical state of a subject, the method comprising: acquiring at least one physiological signal from a subject;deriving at least one parameter sequence from at least one desired physiological signal belonging to the at least one physiological signal;performing an index determination process, thereby to form a diagnostic index based on the at least one parameter sequence, wherein the diagnostic index serves as a measure of the clinical state of the subject and wherein the index determination process includes indicating the diagnostic index to a user;monitoring at least one input signal, wherein the at least one input signal belongs to a group including the at least one physiological signal and the at least one parameter sequence; andenhancing, based on the monitoring, the quality of the index determination process. 2. A method according to claim 1, wherein the enhancing includes: producing, based on the at least one input signal, quality information indicative of the reliability of the diagnostic index; andgiving, based on the quality information, an indication of the current reliability of the diagnostic index. 3. A method according to claim 2, wherein the monitoring comprises: determining changes occurring in a selected physiological signal belonging to the at least one physiological signal; anddefining changes occurring in a selected parameter sequence belonging to the at least one parameter sequence; andthe producing comprisescomparing the changes occurring in the selected physiological signal with the changes occurring in the selected parameter sequence; andgenerating the quality information based on the comparing. 4. A method according to claim 3, wherein the determining includes determining changes occurring in the selected physiological signal, wherein the selected physiological signal is a blood pressure (BP) signal. 5. A method according to claim 3, wherein the defining includes defining changes occurring in the selected parameter sequence, wherein the selected parameter sequence is a plethysmographic amplitude (PGA) signal. 6. A method according to claim 2, wherein: the producing includes calculating the quality information as the correlation between a selected physiological signal and a selected parameter sequence, wherein the selected physiological signal belongs to the at least one physiological signal and the selected parameter sequence belongs to the at least one parameter sequence; andthe giving includes displaying a warning message when the correlation fulfills a predetermined condition. 7. A method according to claim 2, wherein: the monitoring includes determining an auxiliary signal based on at least one of the at least one physiological signal; andthe producing includes producing the quality information, in which the quality information represents the correlation between the auxiliary signal and a selected parameter sequence belonging to the at least one parameter sequence. 8. A method a... | - | A61B005/00
US20080241065 | A1 | 2008-10-02 | Systems and methods for the detection and analysis of in vivo circulating cells, entities, and nanobots | 2007-03-27 | An improved circulating cell counter for generating light, and for delivering this light to a site in vivo for determining the presence, absence, concentration or count of a target cell, in which a light source such as a laser diode (121) and integrated optics (153) produce a beam transmitted to an in vivo target region (165), such as a capillary bed with flowing cells in a living tissue. Based upon the movement of cells in and out of this region, a circulating cell count (192) is generated, allowing determination of the presence, absence, concentration or count of the target cell. Use with optical, magnetic, or nanobot contrast agents, and methods of use are also described. | 1. A noninvasive in vivo circulating cell counting system, comprising: a detector, said detector functionally coupled to a target region and further arranged to detect a signal within a living entity, said contrast signal representative of a contrast agent present in a target cell; anda counter, which determines when a target cell passes through a field of view of said detector, said target cell passage created by a cell movement within said living entity, for determining a target cell estimate, measure, count, presence, absence, degree, or level. 2. The system of claim 1, wherein said contrast agent is a ferrite bead, and said detector is comprised of a magnetic field detector. 3. The system of claim 1, wherein said contrast agent is an optical contrast agent, said detector is comprised of a photodetector, and said system further comprises a light source, said light source optically coupled to said target region. 4. The system of claim 1, wherein said contrast agent is targeted to circulating bacteria. 5. The system of claim 1, wherein said contrast agent is targeted to ovarian cancer using a folate receptor. 6. The system of claim 1, wherein said contrast agent is targeted to prostate cancer using a PSMA extracellular membrane protein. 7. The system of claim 1, wherein said contrast agent is located in an injected and circulating micelle, said micelle operating as said target cell, and further wherein said contrast signal is induced in said micelle by contact with a selected cell type, protein, pH, or other trigger. 8. A method of noninvasively monitoring a parameter related to the in vivo presence, absence, count, or concentration of a target cell type within a living entity, comprising the steps of: emitting electromagnetic radiation into a target region of the entity, the emitted radiation selected to interact with a reporter agent and/or target cells present in the living entity and moving through the region;detecting over time or space a target signal returning from said region; anddetermining a parameter related to the presence, absence, or concentration of the target cell in circulation within the living entity based upon a temporal change or distribution of the target signal within the region over time. 9. A noninvasive in vivo circulating cell counting system, comprising: a light emitter, said emitter optically coupled to a target region located within a living entity;a light detector, said detector optically coupled to the target region and further arranged to detect light from said emitter after having interacted with said target region; anda counter, configured to determine when a target cell has passed into or out of a field of view of said detector, said target cell motion created by a cell movement within said living entity. 10. A method of noninvasively monitoring the in vivo presence, absence, count, or concentration of a target cell type within a living entity, involving: providing an optical contrast reporter agent;a... | - | A61B005/145
US20080287891 | A1 | 2008-11-20 | Device for irrigating and inspecting a wound | 2008-05-06 | The invention relates to a device for more effectively performing wound irrigation and cleansing. The device consists of a coupling member that may be releasably coupled by threaded engagement with the dispensing outlet of a manually compressible container having a reservoir of cleansing fluid. The coupling member has an upstream end portion containing a void region which is in fluid communication with the reservoir after coupling. A magnifying lens is carried by the coupling member and is located intermediate the upstream and downstream ends of the coupling member. The magnifying lens has a multiplicity of orifices adjacent its peripheral edge where the orifices communicate with both the void region and a chamber region. The chamber region is located downstream of the lens and has an opening at its downstream end to permit cleansing fluid to be delivered to the wound upon manually compressing the container. After cleansing the wound, the coupling member is de-coupled from the container and the wound inspected through the magnifying lens for contaminants or debris. | 1. A device for both irrigating a laceration with irrigation fluid contained in the reservoir of a manually compressible container of the type having a neck portion with a dispensing outlet in fluid communication with the container reservoir and for examining the laceration for residual contaminants and debris, comprising: (a) a coupling member having an axis of elongation, an upstream end portion, and a downstream end portion, where said upstream end portion is so dimensioned and proportioned to permit said upstream end portion to releasably couple with said dispensing outlet, said upstream end portion of said coupling member having a axially extending void region therein where said reservoir is in fluid communication with said axially extending void region upon coupling of said coupling member with said dispensing outlet;(b) a magnifying lens having a peripheral edge carried by said coupling member intermediate said upstream and downstream end portion, said magnifying lens forming a downstream boundary of said axially extending void region;(c) a splash shield having a downstream terminal end carried by said coupling member, said splash shield having an axially extending chamber region therein, where said downstream end portion has a radially extending opening communicating with said chamber region and where said magnifying lens forms an upstream boundary of said chamber region, and where said magnifying lens has an orifice adjacent said peripheral edge to permit fluid communication between said axially extending void region and said axially extending chamber region. 2. The device recited in claim 1 where said coupling member is radially symmetrical to said axis of elongation. 3. The device recited in claim 1 where said radially extending opening is located adjacent said downstream terminal end of said splash shield. 4. The device recited in claim 1 where said splash shield has a multiplicity of drain ports located adjacent said downstream terminal end. 5. The device recited in claim 1 where said peripheral edge is so contoured and proportioned such that said orifice forms a part of said peripheral edge. 6. An improved system for both irrigating and examining a wound of the type having a manually compressible container, a dispensing outlet in the neck portion of said container, and a reservoir within said container for holding cleansing fluid, the improvement comprising: (a) a coupling member having an axis of elongation, an upstream end portion, and a downstream end portion, where said upstream end portion is so dimensioned and proportioned to permit said upstream end portion to releasably couple with said dispensing outlet, said upstream end portion of said coupling member having a axially extending void region therein where said reservoir is in fluid communication with said axially extending void region upon coupling of said coupling member with said dispensing outlet;(b) a magnifying lens having a peripheral edge carried by said couplin... | - | A61M035/00 A61B005/00
US20080287823 | A1 | 2008-11-20 | Index Determination | 2006-07-19 | A method of determining an index indicative of the presence, absence or degree of left ventricular hypertrophy in a subject. The method includes using a processing system to determine a measured impedance value for at least one body segment. For each body segment the measured impedance values are used to determine at least one impedance parameter, which are then used to determine a fat-free mass for the subject. The fat free mass can then be used as the index. | 1-19. (canceled) 20. A method of determining an index indicative of the presence, absence or degree of left ventricular hypertrophy in a subject, the method including, in a processing system: (a) determining a measured impedance value for at least one body segment;(b) for each body segment, and using the measured impedance values, determining at least one impedance parameter value;(c) using each determined impedance value to determine a fat-free mass for the subject; and,(d) determining the index at least in part using the fat-free mass. 21. A method according to claim 20, wherein the method includes, in the processing system determining the index using the fat-free mass and an indication of a measured left ventricular mass. 22. A method according to claim 20, wherein the method includes, in the processing system: (a) comparing the index to a reference; and,(b) determining the presence, absence or degree of LVH using the results of the comparison. 23. A method according to claim 22, wherein the reference includes at least one of: (a) a predetermined threshold;(b) a tolerance determined from a normal population;(c) a predetermined range; and,(d) an index previously determined for the subject. 24. A method according to claim 20, wherein the method includes, in the processing system, displaying at least one of: a) a fat free mass;b) a determined index;c) a ventricular mass;d) normal ranges for the index; and,e) normal ranges for fat free mass; and,f) normal ranges left ventricular mass. 25. A method according to claim 24, wherein the method includes determining the ranges in accordance with subject parameters. 26. A method according to claim 20 wherein the method includes, in the processing system: (a) determining a plurality of measured impedance values for each body segment, each measured impedance value being measured at a corresponding measurement frequency; and,(b) determining the impedance parameters based on the plurality of measured impedance values. 27. A method according to claim 20, wherein the parameter values include R0 and R00, wherein: R0 is the resistance at zero frequency; and,R∞ is the resistance at infinite frequency. 28. A method according to claim 25, wherein the method includes, in the processing system, determining the parameter values using the equation:Z=R∞+R0-R∞1+(jωτ)(1-α)where: Z is the measured impedance at angular frequency ω,τ is a time constant, andα has a value between 0 and 1. 29. A method according to claim 26, wherein the method includes, in the processing system: (a) determining the impedance of each body segment at four discrete frequencies; and, (b) determining values for the parameters by solving the equation using four simultaneous equations. 30. A method according to claim 26, wherein the method includes, in the processing system, determining the parameter values by: (a) determining an impedance locus using the measured impedance values; and, (b) using the impedance locus to determine the paramete... | - | A61B005/053
US20080319331 | A1 | 2008-12-25 | Medical Data Signal Processing Systems | 2006-05-10 | A medical signal processing system for processing medical signal data obtained from a human or animal heart, and displaying a graphical representation of the processed data for providing an alert, the system including: a medical signal data input for inputting one or more ECG signals; a graphical display; and a signal processor coupled to said data input and to said display; and wherein said signal processor is configured to: input medical signal data associated with a spatial position in or on said organ; perform at least one measurement relating to said organ using said medical signal data to provide measurement data; determine from said measurement data whether said measurement is normal or abnormal; and display, on a graphical representation of said heart, a graphical indication of an abnormality at said spatial position on a wall of said heart with which said signal data is associated responsive an abnormal to said determination. | 1. A medical signal processing system for processing medical signal data obtained from a human or animal heart, and displaying a graphical representation of the processed data for providing an alert, the system comprising: a medical signal data input for inputting one or more ECG signals;a graphical display; anda signal processor coupled to said data input and to said display; and wherein said signal processor is configured to:input medical signal data associated with a spatial position in or on said organ;perform at least one measurement relating to said organ using said medical signal data to provide measurement data;determine from said measurement data whether said measurement is normal or abnormal; anddisplay, on a graphical representation of said heart, a graphical indication of an abnormality at said spatial position on a wall of said heart with which said signal data is associated responsive an abnormal to said determination. 2. A medical signal processing system as claimed in claim 1 wherein said abnormality determination determines into which of a plurality of ranges of severity of abnormality said measurement falls; and wherein said graphical indication of said abnormality is responsive to said range determination. 3. A medical signal processing system as claimed in claim 1 wherein said abnormality determination determines into which of a plurality of categories of abnormality said measurement falls; and wherein said graphical indication of said abnormality is responsive to said category determination. 4. A medical signal processing system as claimed in claim 1, wherein said graphical representation is responsive to a history of said abnormality determination. 5. A medical signal processing system as claimed in claim 4 wherein said display further indicates a past abnormality when said determination is normal and was previously abnormal. 6. A medical signal processing system as claimed in claim 4 wherein said graphical representation indicates a duration of a said abnormality. 7. A medical signal processing system as claimed in claim 1 wherein said signal data comprises a plurality of signals associated with a plurality of said spatial positions in or on said organ, wherein said measurement and determination are performed for each of said plurality of signals, and wherein said display graphically indicates, for each of said plurality of signals, an abnormality, responsive to said measurement and determination, at a said spatial position associated with the signal. 8. A medical signal processing system as claimed in claim 7 wherein said graphical indication for a said signal is responsive to an abnormality determination for one or more others of said signals. 9. A medical signal processing system as claimed in claim 1 wherein said graphical indication comprises a representation of a said abnormality which grows in size with duration and/or severity of said abnormality. 10. A medical signal processing system as claimed in claim 1... | - | A61B005/0444 A61B005/044 A61B006/03
US07468034 | B2 | 2008-12-23 | Biological information monitoring apparatus | 2005-03-23 | A biological information monitoring apparatus highly precisely observes a clothed subject who is taking any desired posture. The apparatus acquires information on positional displacements of the surface of a living body, by generating a high-frequency electromagnetic wave to radiate it into air, detecting the electromagnetic wave scattered by the surface of the living body, computationally determining the changes with time of the positional displacements of the surface of the living body from the state of propagation of said electromagnetic wave and computationally determining characteristic quantities of vibrations of the living body such as the pulse or the respiration from the changes with time as biological information. | 1. A biological information monitoring apparatus for monitoring biological information by acquiring information on positional displacements of a surface of a living body, said apparatus comprising: an electromagnetic wave generating unit for generating an electromagnetic wave of a frequency band between 300 GHz and 30 THz and radiating it into air;an electromagnetic wave detecting unit for detecting the electromagnetic wave scattered by the surface of the living body;a first computational-determination unit for computationally determining changes with time of the positional displacements of the surface of the living body from a state of propagation of the electromagnetic wave;a second computational-determination unit for computationally determining characteristic quantities of the living body from the changes with time as biological information; anda memory unit for storing characteristic quantities of a living body in advance,wherein an output signal obtained as biological information from said second computational-determination unit is compared with the characteristic quantities stored in said memory unit to judge physical conditions relating to a blood pressure or a degree of hardening of arteries of the living body, and a result of judgment is displayed by characters or output by voice sounds. 2. The apparatus according to claim 1, the electromagnetic wave is a short pulse that is generated repeatedly and the half-width of the pulse is not greater than 33 psec. 3. The apparatus according to claim 1, wherein the changes with time of the positional displacements of the living body at a plurality of sites are computed by said computational-determination unit by the electromagnetic wave and the state of propagation of the characteristic quantities as computed from the changes with time can be detected. 4. The apparatus according to claim 1, wherein said biological information monitoring apparatus is contained in a wash basin, a toilet, or a chair, and biological information is remotely acquired by said apparatus. 5. A biological information monitoring apparatus for monitoring biological information by acquiring information on positional displacements of a surface of a living body, said apparatus comprising: an electromagnetic wave generating unit for generating an electromagnetic wave of a frequency band between 300 GHz and 30 THz and radiating it into air;an electromagnetic wave detecting unit for detecting the electromagnetic wave scattered by the surface of the living body;a first computational-determination unit for computationally determining changes with time of the positional displacements of the surface of the living body from a state of propagation of the electromagnetic wave;a second computational-determination unit for computationally determining characteristic quantities of the living body from the changes with time as biological information; anda memory unit for storing characteristic quantities of a living body in advance,wherei... | Canon Kabushiki Kaisha | A61B005/02 A61B005/00
US07445599 | B2 | 2008-11-04 | Ultrasonic cellular tissue screening tool | 2002-12-23 | A system to aid in ultrasound scanning of breast tissue includes a fabric covering, which is adapted to hold the breast tissue in place during scanning, an ultrasonic coupling agent applied to one of the covering or the breast tissue, and a lenticular-shaped pad to cover the nipple of the breast tissue. The pad is placed beneath and held in place by the fabric covering, and the covering is capable of absorbing the coupling agent and conducting ultrasonic energy with minimal interference. | 1. A system to aid in ultrasound scanning of breast tissue, comprising a fabric covering for the breast tissue, adapted to hold the breast tissue in place during scanning; andan ultrasonic coupling agent applied to one of the covering or the breast tissue;a lenticular-shaped pad to cover the nipple of the breast tissue, wherein the pad is placed beneath and held in place by the fabric covering;wherein the covering is capable of absorbing the coupling agent and conducting ultrasonic energy with minimal interference. 2. The system of claim 1, wherein the covering is constructed with one or more chain stitched seams. 3. The system of claim 1, wherein the pad is ultrasonically conductive and has different ultrasonic characteristics than breast tissue. 4. A system to aid in ultrasound scanning of breast tissue, comprising: a fabric covering for the breast tissue, adapted to hold the breast tissue in place during scanning; andan ultrasonic coupling agent applied to one of the covering or the breast tissue;wherein the covering is capable of absorbing the coupling agent and conducting ultrasonic energy with minimal interference; andwherein the covering is pre-impregnated with the ultrasonic coupling agent. 5. A system to aid in ultrasound scanning of breast tissue, comprising: a fabric covering for the breast tissue, adapted to hold the breast tissue in place during scanning: andan ultrasonic coupling agent applied to one of the covering or the breast tissue:wherein the covering is capable of absorbing the coupling agent and conducting ultrasonic energy with minimal interference; andwherein the covering provides a pattern that can be used by an optical movement sensor. 6. An aid for ultrasound breast scanning comprising: a fabric covering for the breast tissue, adapted to hold the breast tissue in place during scanning, the covering being constructed with one or more chain stitched seams; anda lenticular-shaped pad to cover the nipple of the breast tissue, the pad being ultrasonically conductive and having different ultrasonic characteristics than breast tissue, said pad placed beneath the fabric covering and held in place by the fabric covering. 7. A system to aid in ultrasound scanning breast tissue, comprising a fabric covering for the breast tissue, adapted to hold the breast tissue in place during scanning;a lenticular-shaped pad placed over the nipple, said pad being ultrasonically conductive and having different ultrasonic characteristics than breast tissue, the fabric covering being positioned over the pad to hold the pad in place during scanning; andan ultrasonic coupling agent applied to one of the covering or the breast tissue;wherein the covering is capable of absorbing the coupling agent and conducting ultrasonic energy with minimal interference. 8. A method to aid in the ultrasonic scanning of breast tissue comprising placing a fabric covering over the breast tissue which is adapted to hold the breast tissue in place during scanning;... | SonoCine, Inc. | A61B008/00 A61B008/12 A61B008/14 A41D001/06 A41D013/02 A41F009/00 G09B023/28 A41C003/04 A41C003/00 A41C003/10 A61B005/05
US07467094 | B2 | 2008-12-16 | System and method for accounting and billing patients in a hospital environment | 2005-04-29 | A system and method for accounting and billing patients in a hospital environment. A hospitalized patient is associated with a patient identifier. Billable services provided to the hospitalized patient are associated with the patient identifier. If the billable service is provided by a physician, a physician identifier is also associated with the billable service. The hospitalized patient is scored and a current procedural terminology (CPT) assignment manager assigns CPT codes to the billable service. A bill generator receives the patient data, patient insurance information, physician insurance information, and CPT codes and generates a bill for the billable services provided to the hospitalized patient. | 1. An accounting and billing system for patients in a hospital environment comprising: a telecommunication network;monitoring stations comprising monitoring equipment, wherein the monitoring equipment comprises instructions for monitoring data elements from geographically dispersed hospitalized patients and sending the monitored data elements to a remote command center via the telecommunications network,wherein the remote command center comprises instructions for: receiving the monitored data elements from the geographically dispersed hospitalized patients;accessing patient data elements indicative of a medical condition associated with each of the geographically dispersed hospitalized patients;establishing patient-specific rules associated with each of the geographically dispersed hospitalized patients; andapplying the patient-specific rules continuously and simultaneously using a rules engine, wherein the rules engine comprises instructions for: selecting data elements from the monitored data elements and the patient data elements associated with a hospitalized patient;applying a patient-specific rule associated with the hospitalized patient to the selected data elements;determining in an automated fashion at the remote command center 24 hours per day 7 days per week whether the patient-specific rule for the hospitalized patient has been contravened; andin the event the patient-specific rule for the hospitalized patient has been contravened, determining in an automated fashion whether intervention is warranted;a decision support module comprising instructions for, in the event the patient-specific rule for the hospitalized patient has been contravened: accessing a decision support algorithm;applying the decision support algorithm to the selected patient data elements of the hospitalized patient; andproviding patient care guidance formulated to return the medical condition of the patient to compliance with the patient specific rule, whereby the treatment of hospitalized patients is standardized; anda bill generator, wherein the bill generator comprises instructions for: receiving a patient ID associated with the hospitalized patient;receiving a current procedural terminology (CPT) code from a CPT code assignment manager for a billable service provided to the hospitalized patient associated with the patient ID, wherein the billable service is selected from a monitoring service and a patient care service; andassembling a bill for the billable service provided to the hospitalized patient associated with the patient ID using the CPT codes. 2. The system of claim 1, wherein the patient ID is associated with a treatment location. 3. The system of claim 1, wherein the patient data are selected from the group consisting of patient identifying information, patient demographic information, patient admission data, patient laboratory results data, patient insurance data, patient guarantor information, a diagnosis code, attending physician data, and patien... | VISICU, Inc. | G06Q010/10 G06Q010/50 A61B005/00
US07458941 | B2 | 2008-12-02 | Multi-purpose vaginal moisture screening | 2006-02-16 | The method of quickly screening for vaginal moisture conditions, that includes providing a manually manipulable element including a probe insertible into the vagina; providing one or more pH indicating first test site or sites on one portion of the probe; providing a vaginosis test site on another portion of the probe; manipulating the probe to transfer vaginal moisture to the sites, and, therefore, externally of the vagina, applying an hydroxide to the vaginosis test site, to come in contact with moisture at the vaginosis test site; and detecting presences of an amine or amines produced at the vaginosis site. | 1. The method of quickly screening for vaginal moisture conditions, that includes: a) providing a manually manipulable generally key shaped element including a probe insertible into the vagina,b) providing one or more pH indicating first test site or sites on one end portion of the probe,c) providing a vaginosis test site on a different portion of the probe, near said one end portion,d) manipulating the probe to transfer vaginal moisture to said sites, simultaneously,e) and, thereafter, externally of the vagina, applying an hydroxide in liquid form to the vaginosis test site, to come in contact with moisture at said vaginosis test site,f) and detecting presence of an amine or amines produced at said vaginosis site,g) wherein said pH test site or sites include i) a pH indicator characterized as producing a color which corresponds to pH level of vaginal moisture contacting the indicator,ii) a local comparison zone having a color corresponding to a predetermined pH level, and positioned and shaped for quick color comparison with the color of said indicator after indicator contact with vaginal moisture,iii) said local comparison zone being substantially adjacent said indicator,h) and wherein said element is provided to include an enlarged substantially flat handle from which the probe projects longitudinally to be gripped between the user's thumb and forefinger, the handle having two laterally shaped edges presented toward said test site or sites and spaced generally longitudinally therefrom to limit probe longitudinal insertion into the vagina, said pH indicator and said local comparison zone being aligned on a line which extends between said two edges, and both having curved edges facing oppositely and convexly away from said line, and protruding at opposite edges of said probe. 2. The method of claim 1 including g) withdrawing the probe from the vagina between said steps d) and e), andh) observing the coloration of said pH indication site as an indication of pH level. 3. The method of claim 1 wherein the probe has opposite sides, said one portion of the probe located at one of said sides, and said different portion of the probe located at another of said opposite sides. 4. The method of claim 1 wherein said hydroxide is selected from the group: i) KOH,ii) milk of magnesia,iii) sea water,iv) baking soda, or mixture thereof with water. 5. The method of claim 1 wherein said detecting step includes a nasal whiff test. 6. The method of claim 1 wherein said probe has opposite sides, and said vaginosis test site is at one of said opposite sides, and said first site or sites is at another of said opposite sides. 7. The method of claim 6 wherein said vaginosis site is elongated along the probe. 8. The method of claim 1 including providing a carrier for said hydroxide, and effecting said application of the hydroxide to the vaginosis test site after said moisture transfer thereto. 9. The method of claim 8 including preliminarily providing a support f... | FemTek LLC | A61B010/00 A61B005/00 B65D081/00
US07440792 | B2 | 2008-10-21 | MR visualization of interventional devices | 2004-01-30 | MR method for generating MR images of an object to be examined located inside an examination volume by means of an MR device and for simultaneously locating an interventional instrument inserted into the examination object, which instrument is equipped with a locating device. For the reconstruction of an MR image (20) MR signals together with device-delivered location signals from the examination volume are then recorded, after which the locating of the instrument takes place by the assessment of the image characteristics conditioned by the location signals. For making it possible to determine the position of the interventional instrument in an extremely rapid way, the invention proposes that the recording of the MR signals be effected by recording successively a plurality of MR partial data records (16, 17, 18) with respective incomplete and preferably non-cartesian sampling of the local frequency spectrum assigned to the examination volume. Two locating data records (21, 22) each are extracted from each of the MR partial data records (16, 17, 18). The actual locating of the instrument then takes place in that a difference image (33) is reconstructed from the two locating data records (21, 22). | 1. An MR method for generating MR images of an object to be examined located inside an examination volume by means of an MR device and for simultaneously locating an instrument also located inside the examination volume, wherein the instrument is equipped with a locating device, by which location signals can be generated in the environment of the instrument, where MR signals together with the location signals from the examination volume are recorded for the reconstruction of an MR image, while the locating device is simultaneously switched to and fro between various states of operation, and wherein the locating of the instrument takes place in that the image characteristics, caused by the switching processes are analyzed, wherein the recording of the MR signals takes place by successively recording at least first and second incompletely sampled partial data records with incomplete sampling of the local frequency spectrum assigned to the examination volume, each incompletely sampled partial data record including a first locating data subrecord and a second locating data subrecord, where the first and second locating data subrecords are extracted from each of the MR incompletely sampled partial data records, so that each of the locating data subrecords is assigned to one of the states of operation of the locating device and where the locating of the instrument takes place by reconstructing difference images from the first and second locating data subrecords of each of the at least first and second incompletely sampled partial data records wherein by combining the at least first and second incompletely sampled partial data records, a complete MR image data record is generated from which the MR image is reconstructed, such that at least two difference images are generated for each MR image. 2. The MR method as claimed in claim 1, wherein the sampling of the local frequency spectrum takes place during the recording of the MR signals in a non-cartesian way. 3. The MR method as claimed in claim 2, wherein the local frequency spectrum is sampled radially or spirally at the time of the recording of the MR partial data records. 4. The MR method as claimed in claim 1, wherein the switching frequency, with which the locating device is switched to and fro between the various states of operation, is half a sampling frequency with which the MR signals are sampled during the recording of the MR partial data records. 5. The MR method as claimed in claim 1, wherein locating the instrument is carried out by finding out pixels in the difference image, the brilliance values of which pixels exceeding a threshold value that can be predefined. 6. A computer useable medium having a program code means embodied therein for controlling a processor of an MR device to perform the method as claimed in claim 1. 7. A magnetic resonance imaging system for imaging an object and an instrument located within the object, said system comprising: an MR imaging device that produce... | Koninklijke Philips Electronics N.V. | A61B005/05
US07431702 | B2 | 2008-10-07 | Device for examination of the motor system of the human or animal body | 2002-07-24 | A device for examination of motor system of human or animal body, through examination and determination of the ability of muscular system to function and perform, include a unit to insert, rest and secure body parts to be examined, and to repeat examinations of the same body parts in a same fixed position, a power source to put weight on muscle and a unit to stimulate or contract the muscles or nerves, a rest and securing element with a splint rest or with an arm rest table, a thumb mould with thumb bridge and a thumb lever, in which the thumb lever is firmly connected to a vertically arranged main unit axle, which connects to the power source; means for determining the type of contraction, to automatically limit the power via stopping mechanism, as well as stimulants for stimulation of muscles, nerves or the central nervous system, mechanical and electrical/electronic measuring devices including dynamometers, acceleration meters, muscle-length meters, sensors for the recording of electrical potential fluctuations, the skin temperature and acoustic signals and vibrations. | 1. A device for the examination of the motor system of the human or animal body, through the examination and determination of the ability of the muscular system to function and perform, comprising means for inserting, resting and securing body parts to be examined; means for repeating examinations of the same body parts in a same fixed position; a power source to put weight on muscle; means for stimulating or contracting the muscles or nerves, a rest and securing element with a splint rest or with an arm rest table; a thumb mould with thumb bridge and a thumb lever, in which the thumb lever is connected to a vertically arranged main unit axle, which connects to the power source; means for determining the type of contraction, to automatically limit the power via stopping mechanism, as well as stimulants for stimulation of muscles, nerves or the central nervous system; and means for measuring resulting physiological reactions and including mechanical and electrical/electronic measuring devices selected from the group consisting of dynamometers, acceleration meters, muscle-length meters, sensors for the recording of electrical potential fluctuations, the skin temperature and acoustic signals and vibrations, wherein when using the rest and securing element with the splint rest, it is fitted vertically on a diagonal casing wall of a casing of the device, and on that is an arm rest with connecting hand rest; arranged further to this, at a right angle to the splint rest and the arm rest, are the back of the hand support with the back of the hand cushion and the adjustment bar, which are vertically adjustable through this adjustment bar via an adapted slot of the splint rest; further there is, also at a right angle to the arm rest, an arm rest and on the hand rest is a wrist support, horizontally adjustable via a carrier and adjustment track and attached securely with a fastening screw; a palm of the hand support, in connection with the wrist support, runs on the same level and attached to it is a finger support for the fingers II-V, which is movable; the hand including the wrist is inserted and secured, so that the back of the hand support, including the back of the hand cushion is moved vertically through an electrical motor with an ammeter/cut out via a spindle, into which the adjustment bar with the small adjustment block meshes and that alternatively, the wrist support on the adjustment track is moved horizontally and locked in place with the adjusting screw; the motor and the spindle are arranged on the side of the splint rest away from the arm rest through spindle tracks and that the back of the hand support, with the connection bar, can be swivelled around a vertical axle via a swivel coupling in the area where the spindle meshes; the finger support is moved towards the index finger with the aid of a small adjustment block on an adjustment track and locked in place with the adjusting screws. 2. A device according to claim 1, wherein the vertical... | - | A61B005/103
US07440795 | B2 | 2008-10-21 | Detecting sleep disorders using an active implantable medical device | 2003-12-19 | An active medical device, in particular an implantable device such as a cardiac pacemaker, defibrillator, cardioverter, or multisite device, including a function for determining sleep disorders. This device can collect the heartbeat rhythm in at least one heart cavity, and analyze the collected rhythm so as to evaluate a value of heartbeat rate accordingly. The device can also discriminate within a collected heart rhythm the spontaneous events of sinusal origin, and to evaluate the heart rate according only to these spontaneous sinusal events, and to analyze the time variations of the aforementioned heart rate to determine the presence of a sleep respiratory disorder accordingly. This analysis can be realized through a search for episodes of bradycardia followed by episodes of tachycardia, or by determination of an index of sinusal variability that is a function of the detected variations, and comparison of this index with a predetermined threshold. | 1. An active implantable medical device comprising: means for collecting a heartbeat rhythm in at least one heart cavity and implantable means for analyzing the collected heartbeat rhythm and evaluating a value of heart rate as a function of the collected rhythm, wherein the means for analyzing the collected heartbeat rhythm further comprises means for discriminating in said collected heartbeat rhythm spontaneous events of a sinusal origin;means for evaluating the heartbeat rate as a function of said discriminated sinusal spontaneous events;means for computing a variation in time of said heartbeat rate, wherein each variation comprises the difference in duration of a successive cardiac interval from a preceding cardiac interval; andmeans for determining from said variations a presence of a sleep respiratory disorder. 2. The device of claim 1, wherein the analyzing means further comprises means for determining an index of sinusal variability as a function of said variations, and means for comparing said index with a predetermined threshold. 3. The device of claim 2, further comprising means for memorizing said determined index of sinusal variability. 4. The device of claim 1, further comprising a physiological sensor having an output signal corresponding to the respiratory activity of the patient, and means for correlating said physiological sensor output signal with an output signal of the means for analyzing the collected heartbeat rhythm. 5. The device of claim 4 wherein the correlating means further comprises means for confirming an apnea condition as a function of said correlation. 6. The device of claim 1, wherein said analyzing means further comprises means for processing said aforementioned variations and determining therefrom cyclic profiles of episodes of bradycardia followed by episodes of tachycardia. | Ela Medical S.A.S. | A61B005/0402
US07450986 | B2 | 2008-11-11 | Non-invasive method and apparatus for determining onset of physiological conditions | 2002-02-28 | The invention relates to the modelling and design of early warning systems for detecting medical conditions using physiological responses. The device comprises sensors for monitoring physiological parameters such as skin impedance, heart rate, and QT interval of a patient, means for establishing when those parameters change, the rate of change of the parameters, and a neural network processor for processing the information obtained by the sensors. The neural network processor is programmed with a fast learning algorithm. When the neural network establishes that a physiological condition is present in the patient an alarm signal will be generated. The invention extends to a method of non-invasive monitoring of a person using a neural network programmed with a fast learning algorithm. A non-invasive hypoglycaemia monitor is specifically described. | 1. A non-invasive method of determining the value, presence or onset of a hypoglycemia condition in a patient known to be susceptible to hypoglycemia comprising the steps of: determining base parameters of a patient at which the onset of hypoglycemia is likely to occur, said base parameters including: skin impedance, heart rate, QT interval, and mean or peak frequency of an alpha wave;continuously monitoring the patient's parameters comprising: skin impedance, heart rate, QT interval, and mean or peak frequency of the alpha wave;establishing whether one or more of those monitored parameters changes, and if so, the rate of change of that parameter or parameters;feeding data obtained from said monitoring and said establishing into a neural network processor programmed with a fast learning algorithm;comparing the data to the base parameters for that patient; andbased on said comparing, causing a signal to be generated either when said neural network establishes conditions that suggest a presence or imminent onset of a hypoglycemic condition, or when said neural network has estimated a value of said hypoglycemic condition. 2. The non-invasive method according to claim 1 wherein the monitoring of the heart rate and QT interval is done with an ECG. 3. The non-invasive method according to claim 1 wherein the monitoring of the alpha wave is done with an EEG. 4. The non-invasive method according to claim 1 wherein the fast learning algorithm has either a magnified gradient function or an optimal gradient function. 5. The non-invasive method according to claim 1 wherein the fast learning algorithm is a robust sliding mode algorithm. 6. An apparatus for generating an alarm when hypoglycemia is present or imminent in a patient known to be susceptible to hypoglycemia, said apparatus comprising: sensors for sensing a patient's parameters comprising: skin impedance, heart rate, QT interval, and mean or peak frequency of an alpha wave;means for establishing when one or more of the sensed parameters changes and when so established, a rate of change of that parameter or parameters;a neural network linked to said sensors and said means for establishing so as to receive a substantially continuous data stream from said sensors and said means for establishing, the neural network being programmed with a fast learning algorithm and adapted to establish when the sensed parameters, and any change to those parameters, for a particular person are such as to indicate a presence or imminent onset of a-hypoglycemic condition;and alarm means linked to said neural network and adapted to be triggered when the presence or imminent onset of said hypoglycemic condition is established. 7. The apparatus according to claim 6 wherein the fast learning algorithm has either a magnified gradient function or an optimal gradient function. 8. The apparatus according to claim 6 wherein the fast learning algorithm is a robust sliding mode algorithm. 9. The apparatus according to claim 6... | Aimedics Pty Limited | A61B005/04
US08059876 | B2 | 2011-11-15 | Cardiac function analysis apparatus, method and program | 2008-05-27 | A position obtainment means obtains the positions of points in the heart of a subject that are present in at least two three-dimensional volume datasets in a group of three-dimensional volume datasets that have been obtained by imaging the heart of the subject at predetermined time intervals, the points anatomically corresponding to each other, and the at least two three-dimensional volume datasets having been obtained by imaging at different time from each other. A function information obtainment means obtains, based on each of the positions of the points in the heart, which have been obtained by the position obtainment means, function information representing the function of the heart at each point within a three-dimensional space. A display means displays, based on the function information, the motion function of the heart at each of the points in an image that has the shape of the heart of the subject. | 1. A cardiac function analysis apparatus comprising: a volume data storage means for storing a group of three-dimensional volume datasets that have been obtained by imaging the heart of a subject at predetermined time intervals;a position obtainment means for obtaining the positions of points in the heart that are present in at least two three-dimensional volume datasets in the group of three-dimensional volume datasets, the points anatomically corresponding to each other, and the at least two three-dimensional volume datasets having been obtained by imaging at different time from each other;a function information obtainment means for obtaining, based on each of the positions of the corresponding points in the heart, which have been obtained by the position obtainment means, three dimensional displacement vectors that represent displacements of the heart among the corresponding points as function information representing the function of the heart at each point within a three-dimensional space; anda display means for assigning colors that respectively represent the direction of the three dimensional displacement vectors to each point of the heart based on the function information, which has been obtained by the function information obtainment means, and displaying the colors assigned to each point at the positions of each of the points in an image that has the shape of the heart of the subject. 2. A cardiac function analysis apparatus, as defined in claim 1, wherein the function information represents an amount of change in the thickness of the cardiac wall of the subject or a rate of increase in the thickness of the cardiac wall of the subject based on the thickness of the cardiac muscle of the subject at each of the points. 3. A cardiac function analysis apparatus, as defined in claim 1, wherein the function information represents a difference between the diameter of a cardiac ventricle in diastole and that of the cardiac ventricle in systole at each of the points. 4. A cardiac function analysis apparatus as defined in claim 1, wherein: the display means assigns color components corresponding to the sizes of x, y, and z components of the displacement vectors as the colors that represent the directions of the displacement vectors. 5. A non-transitory computer readable medium having a program for causing a computer to function as: a position obtainment means for obtaining the positions of points in the heart of a subject that are present in at least two three-dimensional volume datasets in a group of three-dimensional volume datasets that have been obtained by imaging the heart of the subject at predetermined time intervals, the points anatomically corresponding to each other, and the at least two three-dimensional volume datasets having been obtained by imaging at different time from each other;a function information obtainment means for obtaining, based on each of the positions of the corresponding points in the heart, which have been obtain... | Fujifilm Corporation | G06K009/00
US08196590 | B2 | 2012-06-12 | Variable magnetic moment MR navigation | 2008-06-24 | A method of navigating a medical device in an operating region in a subject. The method includes applying a magnetic field to the operating region and changing the magnetic moment of the medical device by selectively changing a physical condition of at least one magnet element in the medical device to change the orientation of the device with respect to the applied magnetic field. | 1. A medical device navigable in an operating region in a subject using an externally applied magnetic field, the device including at least one element of a material whose magnetization can be altered by changing a physical condition of the material. 2. The medical device according to claim 1 further comprising means for changing the physical condition of the material. 3. A medical device navigable in an operating region in a subject using an externally applied magnetic field, the device including at least one element of a thermomagnetic material whose magnetization can be altered by a temperature change that is between 0°C. and 200°C. 4. The medical device according to claim 3 wherein the magnetization of the thermomagnetic material can be altered by a temperature change that is within 100° of body temperature. 5. The medical device according to claim 3 wherein the device includes an optomagnetic material whose magnetization can be altered by the application of light. 6. The medical device according to claim 3 wherein the device includes an piezomagnetic material whose magnetization can be altered by the application of pressure. 7. The medical device according to claim 3 wherein device includes an electri-magnetic element whose magnetization direction can be altered by the application of an electric field. 8. The medical device according to claim 3 wherein the device includes an piezomagnetic element whose magnetization direction can be altered by the application of mechanical stress. 9. The medical device according to claim 3 wherein the thermomagnetic material has a Curie temperature above normal body temperature, and further comprising means for selectively heating the at least one elements above the Curie temperature. 10. The medical device according to claim 3 wherein the thermomagnetic material has a Curie temperature below normal body temperature, and further comprising means for selectively cooling the at least one element below its Curie temperature. 11. The medical device according to claim 3 wherein the plurality of elements are arranged in at least skewed, divergent cylindrical array. 12. A medical device for navigation in an operating region in a subject to which a magnetic field is applied, the device comprising means for changing the magnetic moment at the distal end of the device. 13. A medical device for navigating in an operating region in a subject to which a magnetic field is applied, the device comprising at least one magnet element, and means for selectively changing a physical property of the at least one magnet element to change the magnetic moment of the device. 14. The device according to claim 13 wherein in the means for selectively changing the physical property of the at least one magnet element comprises means for heating the at least one magnet element. 15. The device according to claim 13 wherein in the means for selectively changing the physical property of the at least one magnet element comprises ... | Stereotaxis, Inc. | A61B019/00 A61M035/00
US08262579 | B2 | 2012-09-11 | System and method for prediction and detection of circulatory shock | 2008-05-16 | Systems and methods for prediction and detection of circulatory shock using estimates or measurements of arterial blood pressure, heart rate, stroke volume, cardiac output, total peripheral resistance, cardiac ejection fraction, cardiac contractility and ventricular end-diastolic volume are provided. These estimates and measurements are used to determine a type of circulatory shock. In some embodiments, the type of circulatory shock is determined to be one of septic shock, hypovolemic shock, anaphylactic shock, hemorrhagic shock, and cardiogenic shock. | 1. A system for predicting or detecting circulatory shock, comprising: a blood pressure measuring device;a processor;a display;a user interface; anda memory storing computer-executable instructions, which when executed by the processor cause the processor to: receive measurements of arterial blood pressure from the blood pressure measuring device;compute mean arterial blood pressure from the received arterial blood pressure measurements;compute estimates, based on the received arterial blood pressure measurements, of heart rate, total peripheral resistance, cardiac output, ejection fraction, and ventricular end-diastolic volume; andselect a type of circulatory shock, from a plurality of types of circulatory shock, based on the computed mean arterial blood pressure and the computed estimates. 2. The system of claim 1, further comprising an electrocardiogram device, wherein selecting a type of circulatory shock is also based on measurements from the electrocardiogram device. 3. The system of claim 1, wherein the computer-executable instructions, when executed by the processor, cause the processor to display at least one of the received measurements and the computed estimates. 4. The system of claim 1, wherein the computer-executable instructions, when executed by the processor, cause the processor to display the selected type of circulatory shock. 5. The system of claim 1, wherein the type of circulatory shock is determined to be one of septic shock, hypovolemic shock, anaphylactic shock, and cardiogenic shock. 6. The system of claim 1, wherein the computed mean arterial blood pressure and the computed estimates are each classified as one of low, normal and high. 7. The system of claim 6, wherein the type of circulatory shock is determined to be one of septic shock, hypovolemic shock, anaphylactic shock, and cardiogenic shock. 8. The system of claim 7, wherein the type of circulatory shock is determined to be one or more of: (A) septic shock if the mean arterial blood pressure is low, total peripheral resistance is low, heart rate is high, ejection fraction is normal or high, and cardiac output is low or high;(B) hypovolemic shock if the mean arterial blood pressure is low, total peripheral resistance is high, heart rate is high, ejection fraction is normal or high, end-diastolic volume is low, and cardiac output is low or high;(C) anaphylactic shock if the mean arterial blood pressure is low, total peripheral resistance is low, heart rate is high, ejection fraction is normal or high, and cardiac output is high; and(D) cardiogenic shock if the mean arterial blood pressure is low, total peripheral resistance is high, ejection fraction is low, end-diastolic volume is high, and cardiac output is low. 9. The system of claim 1, 5, 7, or 8, wherein the selecting is done using a lookup table. 10. The system of claim 1, 5, 7, or 8, wherein the selecting is done using population-specific data. 11. The system of claim 1, 5, 7, or 8, wherein the sel... | Massachusetts Institute of Technology | A61B005/02
US07930020 | B2 | 2011-04-19 | Morphology based arrhythmia detection | 2007-04-27 | A medical device system and associated method sample an EGM signal over a processing window having a predetermined time duration. A number of morphology metrics are determined from the sampled EGM signal, and a heart rhythm is detected in response to the morphology metrics without determining depolarization intervals. The morphology metrics include metrics determined from a slope signal derived from the EGM signal in one embodiment. | 1. A method, comprising: sensing a cardiac electrogram (EGM) signal, comprising cardiac depolarization signals, along a plurality of sensing vectors;sampling the EGM signal over a processing window having a predetermined time duration;determining a plurality of morphology metrics and associated weighting factors from the sampled EGM signal, a morphology metric of the plurality of morphology metrics having a first weighting factor associating a first sensing vector of the plurality of sensing vectors and a second weighting factor, different from the first weighting factor, associated with a second sensing vector of the plurality of sensing vectors different from the first sensing vector;updating a value of a rhythm counter based on the associated weighting factors; andclassifying a heart rhythm in response to the updated value of the rhythm counter, the heart rhythm being detected without determining intervals between the cardiac depolarization signals. 2. The method according to claim 1 wherein determining the plurality of morphology metrics comprises determining a slope signal by computing a difference between a plurality of pairs of EGM signal samples. 3. The method according to claim 2 wherein determining the plurality of morphology metrics further comprises determining a low slope content. 4. The method according to claim 2 wherein the slope signal comprises multiple signal peaks corresponding to multiple cardiac depolarizations occurring during the processing window and wherein determining the plurality of morphology metrics further comprises determining a number of signal peaks occurring during the processing window. 5. The method according to claim 1 wherein the EGM signal comprises multiple peaks corresponding to multiple cardiac depolarization signals occurring during the processing window and wherein determining the plurality of morphology metrics comprises aligning the multiple peaks and determining a template corresponding to the aligned multiple peaks. 6. The method according to claim 5 wherein determining the plurality of morphology metrics further comprises comparing each of the multiple peaks to the template. 7. The method according to claim 1 wherein detecting the heart rhythm comprises comparing a morphology metric of the plurality of morphology metrics to a threshold requirement corresponding to a first heart rhythm and increasing a first rhythm counter corresponding to the first heart rhythm in response to the morphology metric meeting the threshold requirement. 8. The method according to claim 7 wherein detecting the heart rhythm further comprises: increasing a second rhythm counter corresponding to a second heart rhythm in response to the morphology metric not meeting the threshold requirement corresponding to the first heart rhythm;comparing the first rhythm counter to the second rhythm counter; anddetecting the heart rhythm in response to the comparing of the first rhythm counter and the second rhythm counter. 9. T... | Medtronic, Inc. | A61B005/0472
US08197844 | B2 | 2012-06-12 | Active electrode for transdermal medicament administration | 2007-06-08 | A transdermal medicament patch includes a biocompatible substrate having a therapeutic face on one side configured for disposition against the skin of a patient, a biocompatible adhesive on the therapeutic face, a planar medicament matrix covering a portion of the therapeutic face, and a release liner covering the portion of therapeutic that is not obscured by the medicament matrix. An aperture formed through the release sheet affords direct access by medicament to the entire surface of the medicament matrix opposite from the therapeutic face of the substrate. An active electrode positioned between the medicament matrix and the therapeutic face of the substrate includes an electrically conductive backing layer positioned against the therapeutic face of the substrate and a pH-control layer covering less than all of the side of the backing layer opposite from the therapeutic face of the substrate. One active electrode design criterion relates the relative size of the pH-control layer to the size of the backing layer; another relates the size of portion of the area of the backing layer that is free of the pH-control lawyer to the size of the pH-control layer. The pH-control layer is made of an electrically conductive material capable of moderating changes in the hydrogen-ion concentration in the medicament matrix during iontophoretic current flow. An electrical contact electrically coupled through the substrate to the backing layer includes a hollow, electrically conductive snap fitting having an open end and a cooperating stud that is inserted into the open end of the snap. | 1. An active electrode for a medicament patch of the type including a flexible, planar biocompatible substrate with a therapeutic face for disposition against the skin of a patient and a planar medicament matrix having on respective opposite sides thereof a skin contact surface and a securement surface, the securement surface being retained against the therapeutic face of the substrate obscuring a concealed portion thereof with the balance of the therapeutic face defining an exposed portion thereof, the skin contact surface of the medicament matrix effecting electrically conductive engagement with the skin of the patient, when the exposed portion of the therapeutic face of the substrate is retained there against and the medicament matrix is permeated by a medicament, wherein said active electrode is capable of assembly in the medicament patch between the medicament matrix and the concealed portion of the therapeutic face of the substrate, and said active electrode comprises: (a) an electrically conductive backing layer having on respective opposite sides thereof a driving face and a securement surface, said securement surface becoming positioned against the concealed portion of the therapeutic face of the substrate, when said active electrode is assembled in the medicament patch; and(b) a pH-control layer on said driving face of said backing layer covering less than all of said driving face of said backing layer, said pH-control layer and the portion of said driving face free of said pH-control layer together electrically engaging the medicament matrix, said pH-control layer being made of an electrically conductive material capable of moderating changes in the hydrogen-ion concentration in the medicament matrix, when the therapeutic face of the substrate is disposed against the skin of the patient, and an electrical potential is imposed between said backing layer and the skin of the patient remote from the medicament matrix. 2. An active electrode as recited in claim 1, wherein said active electrode further comprises an electrical contact electrically coupled to the securement surface of said backing layer, and said electrical contact comprises: (a) a hollow, electrically conductive snap fitting having an open end; and(b) a cooperating, electrically insulative stud, in the assembled condition of said electrical contact said stud being inserted into said snap through said open end thereof. 3. An active electrode for a medicament patch of the type including a flexible, planar biocompatible substrate with a therapeutic face for disposition against the skin of a patient and an absorbent medicament matrix non-releasably retained against a portion of the therapeutic face of the substrate obscuring a concealed portion thereof with the balance of the therapeutic face defining an exposed portion thereof, the medicament matrix effecting electrically conductive engagement with the skin of a patient, when the exposed portion of the therapeutic face is reta... | Activatek, Inc. | A61F013/00 A61F013/02 A61B005/04
US08195275 | B2 | 2012-06-05 | Vessel size imaging for antiangiogenic therapy | 2008-05-21 | A system and method of evaluating the effectiveness of antiangionetic therapy is described. A patient or test animal is treated with an antiangionetic substance. Magnetic resonance imaging data is obtained prior to and subsequent to the treatment. The parameters of the imaging process are configured so that, by administering a intravascular contrast agent, the relative size of the microvascular and the total vascular volumes in a region of interest may be obtained, so as to form a vascular size index. The value of the vascular size index and other pharmacokinetic data obtained by the administration of a diffusible contrast agent are used to assess the efficacy of the treatment or the antiangionetic substance being studied. | 1. A method of assessing angiographic therapy, the method comprising: placing a subject in a magnetic resonance imaging (MRI) device;acquiring a dynamic contrast enhanced (DCE) MRI data set;administering a blood pool contrast agent via a venous catheter during the acquisition of the DCE MRI data set;obtaining an MRI image sequence data set after obtaining the DCE MRI data set; anddetermining a vessel size index value before and after the angiographic therapy with a therapeutic agent, wherein determining the vessel size index value comprises using a susceptibility difference between blood vessels and tissue surrounding the blood vessels or using a spread of a Larmor frequency. 2. The method of claim 1, wherein a volume of the blood pool contrast agent administered is proportional to a body weight of the subject. 3. The method of claim 1, wherein the blood pool contrast agent comprises nanoparticles. 4. The method of claim 1, wherein the blood pool contrast agent comprises macromolecules. 5. The method of claim 1, wherein the blood pool contrast agent comprises a diffusible contrast agent. 6. The method of claim 1, further comprising evaluating the DCE MRI data set, wherein evaluation of the DCE MRI data set comprises the application of pharmacokinetic models to the DCE MRI data set. 7. The method of claim 1, wherein the vessel size index value is parametric in a ratio of a change in a T2* weighted image relaxation rate to a change in a T2 spin echo image relaxation rate. 8. The method of claim 1, wherein during the obtaining of the MRI image sequence data set, transversal relaxation rates R2 and R2* are measured before and after administering the blood pool contrast agent. 9. The method of claim 1, further comprising assessing a relative effectiveness of the therapeutic agent using the vessel size index value. 10. The method of claim 1, wherein the DCE MRI data set and the vessel size index value are used to evaluate the effectiveness of the therapy. 11. A method for treating a patient, the method comprising: acquiring a dynamic contrast enhanced (DCE) MRI data set;administering, via a catheter, an intravascular contrast agent during the acquisition of the DCE MRI data set;obtaining an MRI data set after obtaining the DCE MRI data set, the MRI data set suitable for performing a vascular size imaging analysis;determining, by a processor, a vessel size index value before and after therapy with a therapeutic agent, wherein determining the vessel size index value comprises using a susceptibility difference between blood vessels and tissue surrounding the blood vessels or using a spread of a Larmor frequency in the vascular size imaging analysis; andperforming, by the processor, the vascular size imaging analysis, wherein performing the vascular size imaging analysis comprises measuring a first parameter proportional to a volume of a microvascular network in a region of interest and measuring a second parameter proportional to a volume of all... | Siemens Aktiengesellschaft | A61B005/055 G06K009/00
US08515523 | B2 | 2013-08-20 | Electrochemical device and method for measuring the redox state of the skin | 2006-12-22 | The present inventor relates to an electrochemical device for measuring the redox state of the skin, comprising at least one working electrode, such as a microelectrode, a counter electrode, and a reference electrode, said electrodes all being fixed into a single support intended to allow each of said electrodes to be simultaneously brought into contact with the surface of the skin to be tested, the electrodes being connected, on the one hand, to a means for imposing a defined voltage between the working electrode and the reference electrode, such as a potentiostat, and, or the other hand, to a device for measuring the intensity of the current generated at the working electrode by the detection of redox species. The invention also related to a method of measuring the redox state of the skin. | 1. An electrochemical device for measuring the redox state of the skin, comprising: at least one working electrode,a counterelectrode,a reference electrode,said electrodes all being fastened in a single support intended to allow each of said electrodes to be contacted simultaneously with the surface of the skin to be tested,the electrodes being connected on the one hand to a means allowing a defined voltage to be applied between the working electrode and the reference electrode, and on the other hand to a device for measuring the intensity of the current generated at the working electrode by the detection of redox species. 2. The device of claim 1, wherein the working electrode is a microelectrode. 3. The device of claim 1, wherein the active surface area of the working electrode or the sum of the surface areas of the working electrodes is between 100 and 10,000 μm2. 4. The device of claim 1, wherein the ratio between the active surface area of the working electrode or the sum of the active surface areas of the working electrodes and the surface area of skin tested has an order of magnitude between 10−3 and 10−5. 5. The device of claim 1, wherein the working electrode is ring-, strip-, disk- or ellipse-shaped or is conical. 6. The device of claim 1, wherein the active surface of the working electrode is made of carbon or of metal, the constituent metal being selected from the group consisting of gold, platinum, iridium, tungsten, and palladium, an alloy of at least two of these metals, an alloy of at least two of these metals with an oxide of one of these metals and an alloy of at least one of these metals with a metal oxide. 7. The device of claim 1, wherein the active surface of the working electrode is made of platinized platinum. 8. The device of claim 1, wherein the reference electrode is selected from the group consisting in an Ag/AgCl electrode, a calomel electrode and a platinum-wire reference pseudoelectrode. 9. The device of claim 1, wherein the counterelectrode is a metal wire of platinum or stainless steel, or is of nonoxidizing conductive ceramic. 10. The device of claim 1, wherein the electrodes are fastened in the support by means of an electrochemically and chemically inert insulating resin. 11. The device of claim 1, further comprising a means allowing a predetermined quantity to be contained of ionic conductive liquid which, in the position of use, makes contact with the skin and contact with the support. 12. The device of claim 11, wherein the support and said means for allowing a skin contact liquid to be contained are made from a single piece of material. 13. The device of claim 11, wherein the skin contact liquid is an aqueous, alcoholic, aqueous-alcoholic or oily medium, or an emulsion said medium having been made conductive, if necessary, by salification. 14. The device of claim 11, wherein the contact liquid is buffered and has a pH of 4 to 8. 15. The device of claim 11, wherein the volume of contact liquid ... | L V M H Recherche Ecole Normale Superieure Centre National de la Recherche Scientifique | A61B005/05 A61B005/04
US08750971 | B2 | 2014-06-10 | Wireless stroke monitoring | 2007-08-02 | A heart monitoring system for a patient includes an 802 protocol transmitter; an 802 protocol receiver adapted to communicate with the 802 transmitter, the 802 protocol transmitter and receiver forming a Doppler radar to detect heartbeat motion on a chest; and an analyzer connected to one of the transmitter and receiver. The analyzer calibrates a Doppler radar signal with an actual blood pressure during a training phase to develop a model. The analyzer is configured to use the model with the Doppler radar signal during an operational phase to estimate blood pressure. | 1. A heart monitoring system for a patient, comprising: a Doppler radar formed with an IEEE 802 protocol transmitter and an IEEE 802 protocol receiver to detect heartbeat motion on a chest of the patient, wherein frequency waves are transmitted by the IEEE 802 protocol transmitter to the chest of the patient located a fixed distance from the transmitter, reflected frequency waves are received by the IEEE protocol 802 receiver from the chest of the patient, and the IEEE 802 protocol receiver is configured to generate a Doppler radar signal based on the reflected frequency waves; andan analyzer coupled to one of the transmitter and receiver including computer readable code configured to perform Doppler operations using the IEEE 802 protocol transmitter and receiver, wherein the analyzer calibrates a training Doppler radar signal with a training blood pressure during a training phase to develop a model and wherein the analyzer is configured to use the model with the Doppler radar signal during an operational phase to estimate blood pressure. 2. The system of claim 1, wherein the analyzer identifies one patient from another based on the Doppler radar signal. 3. The system of claim 1, further comprising a sound transducer coupled to the transmitter and a telephone connection to communicate audio over a plain old telephone service (POTS) network or a public switched telephone network (PSTN) network. 4. The system of claim 1, further comprising one of: an electromagnetic detector, an ultrasonic detector, an optical detector, a Hidden Markov Model (HMM) recognizer, a dynamic time warp (DTW) recognizer, a neural network, a fuzzy logic engine, or a Bayesian network. 5. The system of claim 1, wherein the analyzer monitors patient movement using triangulation. 6. The system of claim 1, further comprising an in-door positioning system coupled to one or more mesh network appliances to provide location information using triangulation. 7. The system of claim 1, further comprising a call center coupled to the analyzer to provide a human response. 8. The system of claim 1, wherein the IEEE 802 protocol transmitter communicates an IEEE 802 protocol and wherein the IEEE 802 protocol comprises one of: 802.11 protocol, 802.15 protocol, 802.16 protocol, WiFi protocol, and WiMAX protocol. 9. The system of claim 1, further comprising a wireless router coupled to the IEEE 802 protocol transmitter and wherein the wireless router comprises one of: an 802.11 router, an 802.16 router, a WiFi router, a WiMAX router, a BLUETOOTH router, or an X10 router. 10. The system of claim 1, further comprising a mesh network appliance coupled to a power line to communicate X10 data to and from the IEEE 802 protocol transmitter. 11. The system of claim 1, wherein the IEEE 802 protocol transmitter transmits voice. 12. The system of claim 1, wherein the analyzer determines one of: total body water, compartmentalization of body fluids, cardiac monitoring, blood flow, skinfold thick... | - | A61B005/02 A61B005/024 A61B005/00
US10085680 | B2 | 2018-10-02 | Safety blood collection assembly with indicator | 2008-03-07 | A needle assembly is disclosed. The needle assembly includes a housing having a flash chamber, and having a distal end and a proximal end engageable with a specimen collection container. The assembly includes a cannula having a patient end, a non-patient end, and a sidewall extending therebetween defining a cannula interior. The patient end of the cannula projects at least partially from the distal end of the housing, and the cannula interior is in fluid communication with the flash chamber. The assembly further includes a shield restrainably engaged with a portion of the housing and axially transitionable over the patient cannula from a retracted position in which the patient end is exposed, to an extended position in which the patient end is shielded by at least a portion of the shield, wherein at least a portion of the flash chamber is visible in the retracted position. | 1. A blood collection assembly, comprising: a housing, comprising a flash chamber, the housing having a distal end and a proximal end engageable with a specimen collection container;a cannula having a patient end, a non-patient end, and a sidewall extending therebetween defining a cannula interior, the patient end of the cannula projecting at least partially from the distal end of the housing, the cannula interior in fluid communication with the flash chamber; anda shield including an engagement member and at least one depending arm having a locking structure associated therewith, wherein said locking structure comprises at least one inwardly angled restraint, said locking structure disposed at a location distal from the engagement member, said engagement member comprising a fully-surrounded member engaged with a portion of the housing, the shield being transitionable from a first position to a second position with respect to the housing, wherein the flash chamber includes a top surface which is free from obstructions on the flash chamber and is visible when the shield is in the first position, and wherein in the second position, the shield is disposed over the patient end of the cannula and the locking structure engages at least one ledge located solely at a distal end of the flash chamber at a location that is distal to the engagement member engaged with a portion of the housing to lock the shield in the second position, wherein the flash chamber includes the at least one ledge configured for cooperating with the locking structure to lock the shield in the second position, wherein the at least one ledge is located at a bottom side of the flash chamber so as to leave the top side of the flash chamber free from any ledges. 2. The blood collection assembly of claim 1, wherein the at least one inwardly angled restraint extends in an upward direction from a bottom portion of the at least one depending arm and extends in an upward direction with respect to the blood collection assembly when the shield is in the second position. 3. A blood collection assembly, comprising: a housing, comprising a flash chamber, the housing having a distal end and a proximal end engageable with a specimen collection container;a cannula having a patient end, a non-patient end, and a sidewall extending therebetween defining a cannula interior, the patient end of the cannula projecting at least partially from the distal end of the housing, the cannula interior in fluid communication with the flash chamber; and a shield including an engagement member and at least one depending arm having a locking structure associated therewith, wherein said locking structure comprises at least one inwardly angled restraint, said locking structure disposed at a location distal from the engagement member, said engagement member comprising a fully-surrounded member engaged with a portion of the housing, the shield being transitionable from a first position to a second position with respect ... | Becton, Dickinson and Company | A61B005/00 A61B005/15 A61B005/154 A61B005/153
US09597041 | B2 | 2017-03-21 | Sequential image acquisition with updating method and system | 2007-03-30 | In accordance with embodiments of the present, a method for acquiring an image is provided. The method comprises acquiring initial image data, identifying one or more regions of interest in at least one of the initial image data or in a first image generated from the initial image data. The method further comprises automatically deriving one or more scan parameters based upon the regions of interest, and acquiring second image data using the scan parameters. | 1. A method, comprising: acquiring initial image data using a first imaging system;identifying, using a processor, one or more regions of interest in at least one of the initial image data or in a first image generated from the initial image data;automatically deriving, using the processor, one or more scan parameters for a subsequent image acquisition, wherein automatically deriving the one or more scan parameters comprises executing a change detection routine that detects changes in the one or more regions of interest relative to a previous examination and derives the one or more scan parameters based upon the detected changes; andacquiring second image data using the scan parameters such that the second image data is limited to the one or more regions of interest. 2. The method of claim 1, wherein the regions of interest are identified automatically. 3. The method of claim 1, wherein the second image data is acquired automatically using the one or more scan parameters. 4. The method of claim 1, wherein the second image data is acquired using a different imaging modality than the initial image data. 5. The method of claim 1, wherein the regions of interest comprise at least one of anatomical structures or functional information of physiological processes. 6. The method of claim 1, wherein the regions of interest are identified by a computer aided diagnosis (CAD) algorithm, or by an operator reviewing possible regions of interest automatically selected by one or more routines. 7. The method of claim 1, comprising generating a second image based upon the second image data. 8. The method of claim 7, comprising combining the first image and the second image. 9. The method of claim 1, comprising generating a second image based on at least the first image data and the second image data. 10. An imaging system comprising: a first imaging modality configured to acquire first image data from which a first image may be derived;a processor configured to identify one or more regions of interest in the first image data or the first image, to detect changes in the one or more regions of interest relative to a previous examination, and to automatically derive one or more scan parameters based upon the detected changes; anda second imaging modality configured to acquire second image data based upon the one or more scan parameters. 11. The imaging system of claim 10, wherein the first imaging modality is an X-ray tomosynthesis system and the second imaging modality is an ultrasound (US) system. 12. The imaging system of claim 10, comprising a control system coupled to the first and second imaging modalities, wherein the control system is configured to automatically acquire the first image data and the second image data. 13. A method comprising: acquiring, using a X-ray based imaging modality, initial image data comprising X-ray projection data;identifying, using a processor, regions of interest in at least one of the initial image data or in a first ... | GENERAL ELECTRIC COMPANY | A61B005/00 A61B006/03 A61B006/00 G06K009/20
US20080275315 | A1 | 2008-11-06 | Pigmentary Deposition Portion Remote Diagnosis System | 2005-01-11 | A major object of this invention is to provide a remote diagnosis apparatus, a remote diagnosis system, a user terminal, a program, a diagnosis program, and a storage for easily diagnosing a pigmentary deposition portion. A user of the system picks up the skin having benign nevus pigmentosus that might be melanoma through a dermoscope using a portable telephone with a camera to which the dermoscope is attached by means of an adapter, accesses to the Internet using an Internet connection function of the portable telephone, and sends the picked up skin image to the remote diagnosis apparatus. After receiving the skin image, the remote diagnosis apparatus uses the melanoma diagnosis program to examine the skin image for melanoma and a disease stage of melanoma if there is melanoma and then sends a result to the user. | 1. A remote diagnosis apparatus for diagnosing a skin lesion, which is communicably connected to a first user terminal and a second user terminal provided with a camera device with dermoscope, comprising: a database for storing data which are received from the first user terminal and relate to the skin lesion, anda diagnosis program updated based on the data stored in the database and diagnosing skin images for the skin lesion,the remote diagnosis apparatus receiving a skin image of a pigmentary deposition portion picked up by the camera device with dermoscope from the second user terminal; diagnosing the received skin image for the skin lesion with the use of the diagnosis program; and sending a diagnosis result to a predetermined destination. 2. A remote diagnosis apparatus communicably connected to a first user terminal and a second user terminal provided with a camera device with dermoscope, comprising: a database for storing data which are received from the first user terminal and relate to a skin lesion;a diagnosis program which is updated based on the data stored in the database and diagnoses skin images for the skin lesion; andan image storage for storing the skin images to be diagnosed,the remote diagnosis apparatusreceiving a skin image of a pigmentary deposition portion picked up by the camera device with dermoscope and source information for identifying the pigmentary deposition portion from the second user terminal andstoring the received skin image in the image storage after associating the source information with the skin image, andthe diagnosis program taking out the skin image stored in the image storage to diagnose the skin image for the skin lesion and sending a diagnosis result to a destination designated by the source information. 3. A remote diagnosis apparatus communicably connected to a first user terminal and a second user terminal provided with a camera device with dermoscope, comprising: a data receiving unit for receiving data relating to a skin lesion from the first user terminal;a database for storing data relating to the skin lesion, which stores the received data;a receiving unit for receiving a skin image of a pigmentary deposition portion picked up by the camera device with dermoscope and source information for identifying the pigmentary deposition portion from the second user terminal;an image storage for storing the received skin image after associating the source information with the skin image;a diagnosis program which is updated based on the data stored in the database and takes out the skin image stored in the image storage to diagnose the skin image for the skin lesion; anda sending unit for sending a diagnosis result of the diagnosis program to a destination designated by the source information. 4. A remote diagnosis apparatus communicably connected to a first user terminal and a second user terminal provided with a camera device with dermoscope, comprising: a database for storing data which are receiv... | - | A61B005/00 G06Q050/00
US20080300484 | A1 | 2008-12-04 | Delay insensitive SVD algorithm for perfusion analysis | 2008-04-11 | This document discusses, among other things, a computerized system comprising a perfusion analyzer circuit and a memory circuit to store a time sequence of volumetric image data. The image data is obtained in a time relation to arrival of a contrast agent at an anatomical region associated with the image. The perfusion analyzer circuit is configured to access the stored image data and receive a representation of an artery and a representation of tissue near the artery, shift the representation of the artery in time such that a time of arrival of the contrast at the artery precedes arrival of the contrast at the representation of tissue, automatically compute a perfusion parameter using the shifted representation of the artery, and provide the perfusion parameter as an output to a user or to an automated process. | 1. A computerized system comprising: a memory circuit, configured to store a time sequence of volumetric image data, wherein the stored image data is obtained in a time relation to arrival of a contrast agent at an anatomical region associated with the image; anda perfusion analyzer circuit, in communication with the memory circuit, wherein the perfusion analyzer is configured to: access the stored image data and receive a representation of at least a portion of an artery included in the image data and a representation of tissue near the artery in the image data;shift the representation of the artery in time such that a time of arrival of the contrast at the artery precedes arrival of the contrast at the representation of tissue;automatically compute a perfusion parameter using the shifted representation of the artery; andprovide the perfusion parameter as an output to a user or to an automated process. 2. The system of claim 1, wherein the perfusion analyzer circuit comprises a mean transit time (MTT) computation circuit that is configured to determine a MTT of the contrast using the computed perfusion parameter. 3. The system of claim 2, wherein the MTT computation circuit comprises a filter circuit to inhibit oscillations in determined MTT values. 4. The system of claim 3, wherein the filter circuit is configured to extend a length of the representation of the artery by zero-padding to inhibit oscillations in determined MTT values. 5. The system of claim 1, wherein the perfusion analyzer circuit comprises a MTT computation circuit that is configured to: precondition a system of linear equations to model a representation of infinite zero-padding of the shifted artery to inhibit oscillations in determined MTT valuessolve the preconditioned system of linear equations to compute the perfusion parameter; anddetermine values of MTT of the contrast using the perfusion parameter computed by solving the preconditioned system of linear equations. 6. The system of claim 5, wherein the representation of infinite zero-padding includes a Toeplitz matrix of values of the shifted artery. 7. The system of claim 5, wherein the perfusion analyzer circuit is configured to compute the perfusion parameter by solving a preconditioned system of linear equations using standard singular value decomposition. 8. The system of claim 1, wherein the perfusion analyzer circuit is configured to shift the representation of the artery such that the time of arrival of contrast at the artery curve is at a time equal to zero (t=0) for performing deconvolution. 9. The system of claim 1, wherein the perfusion analyzer circuit is configured to compute a contrast residue in a region of a stored image voxel as a function of time. 10. The system of claim 1, wherein the perfusion analyzer circuit is configured to automatically compute relative cerebral regional blood flow in a region of a stored image voxel using the perfusion parameter. 11. A computerized method comprising: ac... | - | A61B005/026 A61B006/00
US20080249432 | A1 | 2008-10-09 | Diagnosis of Disease By Determination of Elctrical Network Properties of a Body Part | 2004-03-26 | A system and method for detecting and diagnosing disease states in a body part is described. The method is based on a set of electrical impedance measurements taken on the surface of a body part and a representation of the body part in the form of a network of impedances that would result in the same surface measurements as the actual body part. The system includes an electrical data unit for measuring electrical data of the body part, the electrical data unit having a plurality of Ne electrodes. The system also includes a network module for representing the body part by a network, the network having external nodes and internal nodes connected by current pathways. The system further includes an electrical properties module for determining electrical properties of the pathways using the measured electrical data, and a diagnosis module for utilizing the electrical properties to diagnose the possibility of disease in the body part. It is well known, for example, that tissue electrical impedance changes with malignancy. | 1. A method for diagnosing the possibility of disease in a body part, the method comprising measuring electrical data of the body part with a plurality of Ne electrodes;representing the body part by a network, said network having external nodes corresponding to the location of the plurality of the Ne electrodes and internal nodes, wherein the internal and external nodes are connected by current pathways;determining electrical properties associated with the pathways using the measured electrical data; andutilizing the electrical properties to diagnose the possibility of disease in the body part. 2. The method of claim 1, wherein the step of measuring includes measuring an impedance zij of the body part. 3. The method of claim 2, wherein the step of measuring an impedance zij of the body part includes injecting a current ii between an ith electrode chosen from among the plurality and a base electrode taken to be the Neth electrode of the plurality;measuring a resultant voltage Vij between a jth electrode chosen from among the plurality and the base electrode; andrepeating the steps of injecting and measuring to calculate an Ne−1×Ne−1 impedance matrix ZNB with matrix elements zmn given by zmn=Vmn/im where 1≦i≦Ne−1 and 1≦j≦Ne−1. 4. The method of claim 3, wherein the step of determining electrical properties associated with the pathways includes calculating an Ne×Ne admittance matrix YNB from the impedance matrix ZNB. 5. The method of claim 4, wherein the step of determining electrical properties associated with the pathways further includes determining a conductance matrix, G, for the body part by using the admittance matrix YNB. 6. The method of claim 5, wherein the step of utilizing the electrical properties includes obtaining a typical conductance matrix from a population group; andcomparing the conductance matrix for the body part and the typical conductance matrix from a population group to diagnose the possibility of disease in the body part. 7. The method of claim 5, wherein the step of utilizing the electrical properties includes determining a conductance matrix for a homologous body part; andcomparing the conductance matrix for the body part and the conductance matrix for the homologous body part to diagnose the possibility of disease in the body part. 8. The method of claim 7, wherein the step of comparing includes displaying a first representation of the body part having current pathways with associated branch conductances and a second represention of the body part having current pathways with associated branch conductances. 9. The method of claim 5, wherein the step of determining a conductance matrix includes solving the equation Gee−GeiGii−1GeiT32 Re(YNB) where, if Ne is the number of external nodes and Ni is the number of internal nodes, then Gee is an Ne×Ne matrix, Gei is an Ne×Ni matrix and Gii is an Ni×Ni matrix defined byG=(GeeGeiGeiTGii). 10. The method of claim 9, wherein solving the equation includes utilizing at least ... | Z-Tech (Canada) Inc. | A61B005/053
US20080317310 | A1 | 2008-12-25 | Method and system for image processing and assessment of blockages of heart blood vessels | 2007-12-10 | One embodiment discloses a computerized method of assessing deposits and/or blockages in blood vessels in a human, specifically in a human heart. The method may include inputting patient data and creating a computerized interactive model of a heart based on the patient data. Patient data may include a plurality of images of a least a portion of a human heart. Images may include three-dimensional images. An image may be divided into regions. A property of a region may be assessed. A property may include intensity of brightness of a region or a portion of a region. A region may include one or more voxels or one or more pixels. A method may include comparing a property of a region of a heart from a first image to a second image. The first image and the second image may include equivalent regions acquired during different time periods. | 1. A method of imaging blood vessels in a human body, comprising: providing a plurality of three-dimensional images of at least a portion of a human body acquired over a period of time to a computer system, wherein the plurality of images comprises at least a first image and a second image acquired at different times;dividing the first image and the second image into a plurality of regions, wherein each of the regions corresponds between the first image and the second image;assessing a property in a plurality of regions of the body from the first image;assessing the property in a corresponding region of the body from the second image; andcomparing the property of the regions of the body from the first image to the property of the regions of the body from the second image to select either a region from the first image or a corresponding region from the second image; andcreating a third image of at least a portion of human blood vessels using the selected regions. 2. The method of claim 1, wherein the first image and the second image comprise at least a portion of a human heart. 3. The method of claim 1, wherein a region comprises one or more voxels. 4. The method of claim 1, wherein a property comprises an intensity of a region. 5. The method of claim 1, wherein comparing the property of the regions comprises using a mathematical operator to compare the regions. 6. The method of claim 1, wherein comparing the property of the regions comprises using a mathematical operator to compare the regions, and wherein the mathematical operator comprises the operator greater than. 7. The method of claim 1, wherein the third image at least appears three-dimensional. 8. The method of claim 1, wherein the third image is two-dimensional. 9. The method of claim 1, wherein at least a portion of the plurality of three-dimensional images are acquired using computed tomography imaging. 10. The method of claim 1, wherein at least a portion of the plurality of three-dimensional images are acquired using magnetic resonance imaging. 11. A system, comprising: a CPU; anda system memory coupled to the CPU, wherein the system memory stores one or more computer programs executable by the CPU;wherein one or more computer programs are executable to perform a method, comprising: providing a plurality of three-dimensional images of at least a portion of a human body acquired over a period of time, wherein the plurality of images comprises at least a first image and a second image acquired at different times;dividing the first image and the second image into a plurality of regions, wherein each of the regions corresponds between the first image and the second image;assessing a property in a plurality of regions of the body from the first image;assessing the property in a corresponding region of the body from the second image; andcomparing the property of the regions of the body from the first image to the property of the regions of the body from the second image to select... | - | A61B005/00
US20080300469 | A1 | 2008-12-04 | MINIATURE, WIRELESS APPARATUS FOR PROCESSING PHYSIOLOGICAL SIGNALS AND USE THEREOF | 2007-05-31 | The present invention provides a miniature, wireless apparatus for processing physiological signals comprising a signal-receiving element and a signal-processing element, wherein said signal-receiving element receives plural of signals input from external sensors and transmits the signals to said signal-processing element. Then said signal-processing element divides the receiving-time into n equal intervals and corresponds each divided time-interval to signals received by one sensor.;The present invention further provides a method for processing physiological signals comprising receiving the signals by signal-receiving element, dividing the receiving-time into n equal intervals by signal-processing element, and corresponding each divided time-interval to signals received by one sensor. | 1. An apparatus for processing physiological signal comprising: (a) a signal-receiving element;(b) a signal-processing element;wherein said signal-receiving element receives plural of signals input from external sensors and transmits the signals to said signal-processing element; and said signal-processing element divides the receiving-time into n equal intervals and corresponds each divided time-interval to a signal detected by one sensor; andwherein said signal-receiving element is capable of sending a feedback signal to said external sensors so that said plural of signals input from external sensors are received by said signal-receiving element in a synchronized pattern, and that in turn said external sensors are finely tuned by said signal-receiving element. 2. The apparatus as claimed in claim 1, wherein said n is ranged from 1 to 50. 3. The apparatus as claimed in claim 1, which further comprises one or more signal sensors. 4. The apparatus as claimed in claim 3, wherein said signal sensor is selected form the group consisting of wireless electrocardiogram sensor, wireless electroencephalogram sensor, wireless temperature sensor, wireless electromyogram sensor, miniature tension sensor, and acceleration sensor. 5. The apparatus as claimed in claim 3, wherein said signal sensor consists of electrode, amplifier, microcontroller, transceiver module, and power supply. 6. The apparatus as claimed in claim 5, wherein said signal receiving element will guide the signal to the corresponding time-interval if said transceiver module doesn't send signal at corresponding time-interval. 7. The apparatus as claimed in claim 5, wherein said transceiver module is radio interface. 8. The apparatus as claimed in claim 3, which further comprises a signal recorder. 9. The apparatus as claimed in claim 8, wherein said signal recorder is hard disc, floppy disc, miniature hard disc, or flash memory card. 10. The apparatus as claimed in claim 1, which is carried out by a micro-computer system including at least one member selected from the group consisting of personal computer, notebook computer, radio station, and personal digital assistant. 11. The apparatus as claimed in claim 1, which can further analyze the collected data or deliver the collected data to other signal-receiving elements. 12. The apparatus as claimed in claim 11, wherein the data analysis is carried out by sleep-analyzing algorithm and autonomic nervous-analyzing algorithm. 13. The apparatus as claimed in claim 3, wherein said signal-receiving element can further send a feedback signal to said transceiver module in all of said signal sensors, making said transceiver module send signals in a synchronized pattern. 14. The apparatus as claimed in claim 3, wherein said physiological signal processing apparatus and said signal sensors can be integrated by system-on-a-chip. 15. The apparatus as claimed in claim 1, wherein said physiological signals includes physiological signals transmi... | NATIONAL YANG-MING UNIVERSITY | A61B005/00
US20080294022 | A1 | 2008-11-27 | Birthing Medical Monitor | 2004-11-29 | A birth monitoring system (100), comprising a monitor (120) including at least one sensor operative to measure a physiological parameter associated with labor; a posture sensor (130) which generates a signal indicative of posture of a mother (110); and circuitry which generates an output signal dependent on both a measurement of said sensor and a posture measurement of said posture sensor (130). | 1. A birth monitoring system, comprising: a monitor including at least one sensor operative to measure a physiological parameter associated with labor;a posture sensor which generates a signal indicative of posture of a mother; andcircuitry which generates an output signal dependent on both a measurement of said sensor and a posture measurement of said posture sensor. 2. A system according to claim 1, wherein said output signal comprises a signal selectively suppressed responsive to said measured posture. 3. A system according to claim 1, wherein said circuitry is operative to modify said output signal responsive to said posture signal. 4. A system according to claim 1, wherein said circuitry is operative to generate a signal indicative of a change in said physiological parameter responsive to a change in said posture. 5. A system according to claim 1, wherein said physiological parameter comprises a geometry of a cervix. 6. A system according to claim 1, wherein said physiological parameter comprises a fetal head position relative to a birth canal of a mother. 7. A system according to claim 1, wherein said physiological parameter comprises a fetal physiological parameter. 8. A system according to claim 7, wherein said fetal physiological parameter comprises a fetal heart rate. 9. A system according to claim 1, wherein said physiological parameter comprises a maternal physiological parameter. 10. A system according to claim 1, wherein said physiological parameter is changed by a change in posture. 11. A system according to claim 1, wherein a measurement by said sensor is changed by a change in posture. 12. A system according to claim 11, wherein said sensor comprises at least one internal probe and at least one external probe serving as a reference thereto. 13. A system according to claim 1, wherein said output signal is generated if a posture change is not maintained. 14. A system according to claim 1, comprising memory which stores a correspondence between posture and a physiological parameter. 15. A system according to claim 1, wherein said circuitry extracts at least one maternal physiological parameter from said posture sensor. 16. A system according to claim 15, wherein said maternal parameter comprises breathing. 17. A system according to claim 1, wherein said posture sensor comprises an acceleration sensor. 18. A system according to claim 1, wherein said posture sensor is calibrated to be aligned to one or more maternal body axes. 19. A system according to claim 1, wherein said output signal comprises a recommendation to a caregiver regarding posture. 20. A system according to claim 1, wherein said posture sensor is housed together with at least a part of said at least one sensor. 21. A method of generating an output signal indicative of a medical monitoring of a patient in labor, comprising: measuring at least one postural parameter of the patient;measuring at least one physiological parameter associated with labor; a... | BARNEV LTD. | A61B005/00
US20080306407 | A1 | 2008-12-11 | Bandage Pressure Sensor | 2006-03-29 | A pressure sensor for providing an indication of the pressure applied by a bandage to a human or animal body. The sensor comprises an elongate, flexible support strip (63) adapted to be placed between a bandage and the body, the support strip carrying a flat pressure sensitive portion (69) whose electrical properties vary with applied pressure perpendicular to the general plane of the support strip, and also carrying flat, flexible conductors (65, 66) for connecting the pressure sensitive portion to a supply of electricity and to means for providing an indication of the applied pressure (71). In some arrangements an integrated battery, circuitry and display may be provided on the support strip. | 1. A pressure sensor for providing an indication of the pressure applied by a bandage to a human or animal body, the sensor comprising an elongate, flexible support strip adapted to be placed between a bandage and the body, the support strip carrying a flat pressure sensitive portion whose electrical properties vary with applied pressure perpendicular to the general plane of the support strip, and also carrying flat, flexible conductors for connecting the pressure sensitive portion to a supply of electricity and to means for providing an indication of the applied pressure. 2. A pressure sensor as claimed in claim 1, wherein the support strip carries terminals connected to the flexible conductors. 3. A pressure sensor as claimed in claim 2, wherein the terminals are enlarged end portions of the conductors. 4. A pressure sensor as claimed in claim 2, wherein the terminals are adjacent one end of the support strip and the pressure sensitive portion is adjacent the other end of the support strip. 5. A pressure sensor as claimed in claim 2, in combination with a display module which includes a supply of electricity and means for providing an indication of the applied pressure, the display module being connected to the conductors. 6. A pressure sensor as claimed in claim 1, wherein the support strip carries a battery and means for providing an indication of the applied pressure, connected to the conductors. 7. A pressure sensor as claimed in claim 5, wherein the means for providing an indication of the applied pressure includes a visual display adjacent one end of the support strip and the pressure sensitive portion is adjacent the other end of the support strip. 8. A pressure sensor as claimed claim 1, wherein the support strip can be flexed about both longitudinal and transverse axes. 9. A pressure sensor as claimed in claim 1, wherein the pressure sensitive portion is flexible. 10. A pressure sensor as claimed in claim 9, wherein the pressure sensitive portion can be flexed about both longitudinal and transverse axes. 11. A pressure sensor as claimed in claim 1, wherein the pressure sensitive portion is of a material whose conductivity changes with applied pressure. 12. A pressure sensor as claimed in claim 11, wherein the pressure sensitive portion is of a quantum tunnelling composite. 13. A pressure sensor as claimed in claim 1, wherein the thickness of the sensor is no more that 1 mm. 14. A pressure sensor as claimed in claim 13, wherein the thickness of the sensor is no more that 0.5 mm. 15. A pressure sensor as claimed in any preceding claim 1, wherein the pressure sensitive portion has a thickness of less than about 200 microns. 16. A pressure sensor as claimed in claim 15, wherein the pressure sensitive portion has a thickness of between about 125 and 175 microns. 17. A pressure sensor as claimed in claim 15, wherein the pressure sensitive portion has a thickness of less than 100 microns. 18. A pressure sensor as claimed in c... | PRESSURE MEDICAL LTD. | A61B005/103 A61L015/00 A61B019/00
US20080306392 | A1 | 2008-12-11 | BLOOD VESSEL IMAGING APPARATUS | 2008-06-11 | An imaging device used for a blood vessel imaging apparatus includes: a housing; an image sensor mounted in the housing; a control circuit board fixedly held in the housing and provided with the image sensor; an illumination circuit board disposed in the housing so as to be opposed to the control circuit board, and having an opening in a position opposed to the image sensor; a tubular light intercepting member disposed so as to cover the periphery of the image sensor and protrude through the opening; and an optical part attached to the tubular light intercepting member. | 1. An imaging device used for a blood vessel imaging apparatus, comprising: a housing serving as a casing of the imaging device;an image sensor mounted in the housing;a control circuit board fixedly held in the housing and provided with the image sensor;an illumination circuit board disposed in the housing so as to be opposed to the control circuit board, and having an opening in a position opposed to the image sensor; andan optical part attached to a tubular light intercepting member disposed so as to cover a periphery of the image sensor and protrude through the opening. 2. The imaging device according to claim 1, comprising: a visible light cut-off filter disposed in the housing so as to cover an end of the optical part; anda plurality of illumination elements mounted so as to surround an end of the tubular light intercepting member on a surface of the illumination circuit board opposed to the visible light cut-off filter. 3. The imaging device according to claim 2, wherein the illumination elements are light emitting diodes. 4. The imaging device according to claim 3, wherein an LED drive circuit for driving the light emitting diodes by a pulse modulation method is mounted on the illumination circuit board. 5. The imaging device according to claim 4, wherein the light emitting diodes are more than one in number, the more than one light emitting diodes are divided into a plurality of groups, and the plurality of groups of the light emitting diodes are driven by LED drive circuits independent of each other, respectively. 6. The imaging device according to claim 5, wherein the groups of the light emitting diodes are disposed so as to be opposed to each other with the tubular light intercepting member in between. 7. The imaging device according to claim 4, wherein the light emitting diodes are more than one in number, the light emitting diodes are divided into two groups, the light emitting diodes of the two groups are alternately arranged in an annular shape, and the groups of the light emitting diodes are driven by LED drive circuits independent of each other by currents of opposite directions of the annular shape, respectively. 8. A blood vessel imaging apparatus comprising: an imaging device;a housing serving as a casing of the imaging device;an image sensor mounted in the housing;a control circuit board fixedly held in the housing and provided with the image sensor;an illumination circuit board disposed in the housing so as to be opposed to the control circuit board, and having an opening in a position opposed to the image sensor; andan optical part attached to a tubular light intercepting member disposed so as to cover a periphery of the image sensor and protrude through the opening. | FUJITSU LIMITED | A61B005/02
US20080262391 | A1 | 2008-10-23 | INSTRUMENT FOR MEASURING THE MECHANICAL PROPERTIES OF VOCAL TISSUES | 2008-04-22 | The present invention provides an instrument and method for measuring the viscoelastic properties of vocal tissues. A contact tip is attached to the tissue of interest and the linearly oscillated. Some of the linear oscillation results in the application of shear force to the tissue. A position sensor and a force sensor send signals from the instrument to a processor that can be used to calculate material properties of the tissue being tested, such as viscoelastic constants. This instrument and related method can be of assistance for in vivo measurements of tissues for diagnosis or corrective surgery. | 1. An instrument for measuring a characteristic of a tissue, the instrument comprising: a contact tip for attaching to tissue;means for mechanically applying an oscillatory shearing motion to the contact tip at a controlled frequency;a position sensor for measuring the oscillatory shearing motion of the contact tip;a force sensor for measuring a force imposed on the tissue in response to the oscillatory shearing motion; anda processor for calculating a viscoelastic characteristic of the tissue using a set of measurements from the position sensor and the force sensor. 2. The instrument of claim 1 wherein the contact tip includes a suction device for attaching to the tissue. 3. The instrument of claim 2 wherein the suction device further includes a vacuum device that generates a vacuum for connecting the contact tip to the tissue. 4. The instrument of claim 1 wherein the contact tip includes a releasable adhesive for attaching to the tissue. 5. The instrument of claim 1 wherein the means for mechanically applying the oscillatory shearing motion comprises a rotary motor and a transmission for converting a rotary motion of the rotary motor into the oscillatory shearing motion. 6. The instrument of claim 5 wherein the transmission includes an eccentric cam driven by the rotary motor and a yoke, the yoke attached to and providing the oscillatory shearing motion to the contact tip via a shaft connected to both the yoke and the contact tip. 7. The instrument of claim 6 wherein the means for mechanically applying the oscillatory shearing motion further includes a counter-yoke for driving a counter-mass to minimize vibration of the instrument. 8. The instrument of claim 1 wherein the controlled frequency is a frequency in a range of 20 Hz to 1500 Hz. 9. The instrument of claim 1 wherein the position sensor comprises a linear variable differential transformer. 10. The instrument of claim 1 wherein the position sensor comprises an incremental optical encoder with an index pulse signal. 11. The instrument of claim 1 wherein the force sensor includes a pair of stacked piezoelectric discs, one of the pair of stacked piezoelectric discs bearing the force imposed on the tissue and the other of the pair of stacked piezoelectric discs being mechanically blocked from the force imposed on the tissue. 12. The instrument of claim 1 wherein the force sensor includes a strain gage, the strain gage measuring the force imposed on the tissue. 13. The instrument of claim 12 wherein the strain gage is a piezo-resistive strain gage. 14. The instrument of claim 1 wherein the force sensor includes a first sensor and a second sensor, the first sensor receiving the force imposed on the tissue and the second sensor being mechanically blocked from the force imposed on the tissue, and wherein the processor for calculating a viscoelastic characteristic of the tissue includes a circuit for comparing a set of signals from the first sensor and the second sensor to produce a ... | - | A61B005/00
US20080287745 | A1 | 2008-11-20 | Alarm Device for Preventing Cot Death | 2005-07-08 | The invention relates to a device comprising means for the remote transmission of an active control signal for triggering alarm means (2, 10), for a person looking after a child. Said alarm control signal is generated by the control of sensor means which are sensitive to the spatial orientation of a plate mounted in a support (1) to be fixed to the clothing of the child to be watched. The support (1) is fixed to the navel of the child in such a way that the signal triggering the alarm means is automatically generated when the sensor means detect a variation in the orientation of the plate which is higher than a pre-determined threshold. The support (1) is arranged in a remote unit which paralyses the alarm means. | 1-10. (canceled) 11. A cot death prevention device, comprising two units that are movable independently of one another, detector means responsive to the orientation in space of a plate mounted in a support provided in a first one of said units that attaches in a predetermined arrangement to the clothing of the infant to be supervised, means of remote transmission for an alarm control signal that is produced under the control of said detector means and that triggers alarm means incorporated in the second of said units for triggering said alarm means in the event of detection of an orientation variation of said plate that is greater than a predetermined threshold, and further comprising means to hinder the operation of said detector means in producing said alarm control signal when said first unit is in a position of storage in the first unit. 12. A device according to claim 11, wherein said detector means comprise an angular position detector functioning under the effect of gravity in relation to a vertical axis and wherein said support is shaped to facilitate its positioning by attachment to the infant's clothing in an arrangement wherein the angular position detected at least comprises that of the plate in its orientation variations in tilt with respect to a longitudinal plane through the infant's vertebral column. 13. A device according to claim 11, wherein said detector means are associated in said first unit of the device with means of attachment to the clothing of an infant to be supervised in said predetermined arrangement, while said second unit incorporating said alarm means is movable separately from said first unit to put the alarm means within the reach of a person in charge of the infant's supervision. 14. A device according to claim 13, wherein said means of attachment to the infant's clothing function by pinching the belt of a nappy it wears. 15. A device according to claim 11, wherein said second unit s made in the form of a case having a storage space for the first unit in which a magnetic field is adapted to control a magnet switch of an electrical circuit supplying the detector means so as to hinder the functioning of said detector means producing said alarm control signal when said first unit is positioned in said storage space. 16. A device according to claim 14, wherein said first unit is enveloped in a cushion protecting the infant against risks of injury. 17. A device according to claim 11, wherein said alarm control signal is transmitted wirelessly, in the form of electromagnetic waves, from a transmitter integral with said first unit that is controlled by said detector means to a receiver that is integral with said second unit and that triggers the transmission of an alarm signal towards a person in charge of the infant's supervision, said signal being of the visual and/or sound alert type. 18. A device according to claim 11, wherein said detector means are made and arranged for controlling triggering the alarm wh... | - | A61B005/00
US07471970 | B2 | 2008-12-30 | Method for measuring blood oxygen content under low perfusion | 2005-12-22 | A method for measuring blood oxygen content under low perfusion, which is used in a device for measuring blood oxygen content, includes the steps of: initializing the device that is applied with power, collecting and processing data with a driving circuit of light emitting device, a bias circuit, a gain circuit and an A/D sampling circuit, which are controlled under a core control module; calculating blood oxygen saturation based on the collected data with a data processing module which integrates the collected data in a period of time with an area integration method; and outputting from a communication functional module results of the blood oxygen saturation or pulse rate calculated with the data processing module. The method further includes a decision step of deciding the two results acquired from the data processing module with the waveform method and the integration method respectively based on the intensity of the measured signal and generating the final measured result, performed by a decision unit included in the device. By adopting the above method, the disturbance to effective signal by noise can be eliminated. As a result, the measuring accuracy of blood oxygen content under low perfusion can be improved without increasing the production cost for the measuring device. | 1. A method for measuring blood oxygen content under low perfusion, which is used in a device for measuring blood oxygen content, including the steps of: a. initializing the device that is applied with power;b. collecting and processing data with a driving circuit of a light emitting device, a bias circuit, a gain circuit and an A/D sampling circuit, which are controlled under a core control module;c. calculating blood oxygen saturation based on the collected data with a data processing module which integrates the collected data in a period of time with an area integration method, wherein the collected data is cyclical and the period of time for integration includes more than one cycle; andd. outputting from a communication functional module a result of the blood oxygen saturation or pulse rate calculated with the data processing module. 2. The method for measuring blood oxygen content under low perfusion according to claim 1, wherein in the step of c, the data processing module further fits envelope waveform of a pulse wave and finds out the maximum value and minimum value of the pulse wave to calculate the blood oxygen saturation with a waveform method based on the collected data. 3. The method for measuring blood oxygen content under low perfusion according to claim 2, wherein the method, between the steps of c and d, further includes a decision step of deciding between the two results acquired from the data processing module with the waveform method and the integration method respectively based on the intensity of measured signals and generating a final measured result, performed by a decision unit included in the device. 4. The method for measuring blood oxygen content under low perfusion according to claim 3, wherein a precondition for the decision and calculation in the decision step is: A=a*A1+(1−a)*A2,where “A1” is the result of the blood oxygen saturation acquired with the waveform method, “A2” is the result of the blood oxygen saturation acquired with the integration method, “A” is the final measured result of the blood oxygen saturation, and “a” ranges from 1 to 0. 5. The method for measuring blood oxygen content under low perfusion according to claim 2, wherein the device performs the sampling and measuring by making two lights pass through end parts of tissues, wherein one is red light, and the other is infrared light; and in the step of c, the data of the two lights is integrated respectively to calculate the ratio of integration result of the red light to that of the infrared light for replacing the ratio of AC peak value of the intensity of the red light RedAC to that of the infrared light IrAC received during the period of time. 6. The method for measuring blood oxygen content under low perfusion according to claim 5, wherein: in the step of c, the calculation of the data processing module also depends on a forgetting factor λ which ranges from 0 to 1, and the ratio of the AC peak value RedAC to IrAC of the current two ligh... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/1455
US07460895 | B2 | 2008-12-02 | Method for generating a net analyte signal calibration model and uses thereof | 2005-01-24 | A method for generating a net analyte signal calibration model for use in detecting and/or quantifying the amount of an analyte in a test subject. The net analyte signal can be generated by providing a set of in vivo infrared spectra for a test subject during a period in which an analyte concentration is essentially constant; calculating an optimal subspace of spectra that at least substantially describes all non-analyte dependent spectral variance in the in vivo spectra; providing a pure component infrared spectrum for the analyte; and calculating a net analyte signal spectrum from a data set comprising the optimal subspace spectra and the pure analyte spectrum. The net analyte signal calibration model can be used, for example, in measuring the concentration of analyte in a test subject, and/or for evaluating the analytical significance of an in vivo multivariate calibration model. | 1. A method for generating a net analyte signal calibration model for use in detecting an analyte in a test subject, comprising: a) providing a set of in vivo infrared spectra for a test subject during a period in which an analyte concentration is essentially constant;b) calculating an optimal subspace of spectra that at least substantially describes all non-analyte dependent spectral variance in the in vivo spectra of step a);c) providing a pure component infrared spectrum for the analyte; andd) calculating a net analyte signal spectrum from a data set comprising the optimal subspace spectra of b) and the pure analyte spectrum of c), wherein the net analyte signal spectrum identifies one or more in vivo spectral features specific to the analyte. 2. The method of claim 1, wherein the spectra are absorption spectra. 3. The method of claim 2, wherein the absorption spectra are near infrared absorption spectra in the range of from approximately 4000 cm−1 to approximately 5000 cm−1. 4. The method of claim 2, wherein the absorption spectra are near infrared absorption spectra in the range of from approximately 5500 cm−1 to approximately 6500 cm−1. 5. The method of claim 1, wherein the spectra are reflectance spectra. 6. The method of claim 1, wherein the spectra are single-beam spectra. 7. The method of claim 1, wherein the infrared spectra are absorption spectra in the mid infrared spectral range. 8. The method of claim 7, wherein the absorption spectra are in the range of from approximately 1200 cm−1 to approximately 900 cm−1. 9. The method of claim 1, wherein the analyte is a physiological chemical. 10. The method of claim 9, wherein the analyte is glucose, urea, lactate, triglyceride, total protein, cholesterol, or ethanol. 11. The method of claim 9, wherein the physiological chemical comprises at least one C—H, N—H, or O—H molecular bond. 12. The method of claim 9, wherein the analyte is glucose. 13. The method of claim 1, wherein the test subject is a living organism. 14. The method of claim 13, wherein the test subject is a plant. 15. The method of claim 13, wherein the test subject is an animal. 16. The method of claim 15, wherein the animal is non-mammalian. 17. The method of claim 15, wherein the animal is mammalian. 18. The method of claim 13, wherein the test subject is a microbial species. 19. The method of claim 13, wherein the test subject is a human. 20. The method of claim 1, wherein step b) comprises a principle component analysis. 21. The method of claim 1, further comprising reporting the net analyte signal calibration spectrum on a display device. 22. The method of claim 1, further comprising storing the net analyte signal spectrum on a recordable medium. 23. A method for non-invasively measuring the concentration of an analyte in a test subject, comprising: (a) identifying a test subject in need of having an analyte concentration measured;(b) generating an in vivo net analyte signal calibration model by: i) ... | University of Iowa Research Foundation | A61B005/1455 G01J003/00 G12B013/00
US07801618 | B2 | 2010-09-21 | Auto adjusting system for brain tissue stimulator | 2007-06-22 | An implantable neurostimulator for treating disorders such as epilepsy, pain, movement disorders and depression includes a detection subsystem capable of detecting a physiological condition and a therapy subsystem capable of providing a course of therapy in response to the condition. The therapy subsystem includes an auto-adjust module for automatically adjusting one or more parameters of the therapy so that the therapy subsystem can provide an adjusted parameter to the patient and solicit the patient's feedback concerning the adjustment without requiring the presence of, or immediate involvement with, a clinician or physician. The patient feedback can be analyzed by computer, clinician or a combination of both to determine an optimal range of parameters for subsequent courses of therapy. In this manner, information useful in tuning the neurostimulator therapy parameters to optimize them for individual patient can be acquired automatically outside of the traditional clinical setting, saving time and minimizing patient fatigue that otherwise would be experience in marathon, in-clinic tuning sessions. The auto-adjust module also can be configured to prompt the patient to provide feedback even when parameters are not being adjusted, so as to acquire information for a baseline or about any placebo effect when the patient is otherwise expecting changes to the therapy to be made. | 1. A device for delivering therapy to a patient and soliciting information about the therapy from the patient comprising: a therapy module configured to provide a therapy characterized by at least one parameter to the patient;a detection module configured to automatically detect a physiological event in the patient;a controller configured to automatically determine the at least one parameter of the therapy based on the detected physiological event;a prompting module configured to prompt the patient to express the patient's opinion about the therapy; andan input module configured to receive an input corresponding to the patient's expressed opinion;wherein the controller is further configured to automatically determine whether to change the at least one parameter of the therapy based on the detected physiological event and the expression of the patient's opinion about the therapy. 2. The device of claim 1, wherein the controller further comprises settings for changing the therapy. 3. The device of claim 2, wherein the therapy comprises electrical stimulus, the settings for changing the therapy comprising an allowable range of values for the at least one parameter of the electrical stimulus. 4. The device of claim 1, wherein the controller further comprises a memory configured to store information about the changes made to the therapy. 5. The device of claim 4, wherein the controller further comprises a communication module configured to download information about the changes made to the therapy. 6. The device of claim 1, further comprising a communication module configured to communicate with an external reporting device configured to receive and record patient opinions about the adjusted therapy. 7. The device of claim 6, wherein the prompting module is configured to trigger the external reporting device to issue a prompt to the patient. 8. The device of claim 1, wherein the prompting module is configured to prompt the patient a plurality of times each time a change to the therapy is made by the controller. 9. The device of claim 1, wherein the controller is configured to change the at least one parameter a plurality of times, the prompting module being configured to prompt the patient at least once between each change performed by the controller. 10. The device of claim 9, wherein the prompting module is configured to prompt the patient more than once between two of the changes performed by the controller. 11. The device of claim 1, wherein the prompting module comprises a bone conduction microphone configured to provide an audible signal to the patient. 12. The device of claim 1, wherein the therapy module, the detection module, the controller, and the prompting module are configured to be implanted in a human body. 13. A method of delivering electrical stimulation therapy provided to a patient with an implantable neurostimulator and soliciting information about the therapy from the patient comprising: delivering a first course of el... | NeuroPace, Inc. | A61N001/36
US07435224 | B2 | 2008-10-14 | Extraction of heart sound of fetus from heart sound information that comprises plurality of mixtures of plurality of heart sounds of plurality of fetuses | 2004-07-16 | A heart sound analyzer component of an apparatus in one example extracts a heart sound of a fetus from heart sound information that comprises a plurality of mixtures of a plurality of heart sounds of a plurality of fetuses. | 1. An apparatus, comprising: a plurality of acoustic sensors having nonspecific locations on an abdomen of a mother; anda processor that employs the plurality of acoustic sensors to obtain substantially simultaneously a respective plurality of mixtures of a plurality of fetal heart sounds of a respective plurality of fetuses;wherein each of the plurality of fetal heart sound mixtures obtained by the processor from the plurality of acoustic sensors comprises a different mixture of the plurality of fetal heart sounds;wherein the processor extracts any one or more of the plurality of fetal heart sounds that corresponds to a respective any one or more of the plurality of fetuses from the plurality of fetal heart sound mixtures. 2. The apparatus of claim 1, wherein the processor employs a blind source separation algorithm that accounts for respective propagation delays of the plurality of fetal heart sounds within the plurality of fetal heart sound mixtures to extract the any one or more of the plurality of fetal heart sounds that corresponds to the any one or more of the plurality of fetuses. 3. The apparatus of claim 1, wherein the processor employs an independent component analysis procedure to extract the any one or more of the plurality of fetal heart sounds that corresponds to the any one or more of the plurality of fetuses from the plurality of fetal heart sound mixtures. 4. The apparatus of claim 1, wherein the processor employs a neural network construct with a plurality of nodes in number at least equal to a number of the plurality of fetuses to extract the any one or more of the plurality of fetal heart sounds that corresponds to the any one or more of the plurality of fetuses from the plurality of fetal heart sound mixtures. 5. The apparatus of claim 1, wherein the plurality of fetuses comprises a first fetus and a second fetus, wherein the plurality of fetal heart sounds comprises a first fetal heart sound from the first fetus and a second fetal heart sound from the second fetus; wherein the plurality of fetal heart sound mixtures comprises a first fetal heart sound mixture, a second fetal heart sound mixture, and a third fetal heart sound mixture;wherein the first fetal heart sound mixture comprises a contribution from the first fetal heart sound from the first fetus, a contribution from the second fetal heart sound from the second fetus, and a contribution from background noise;wherein the second fetal heart sound mixture comprises: a contribution from the first fetal heart sound from the first fetus, different from the contribution to the first fetal heart sound mixture;a contribution from the second fetal heart sound from the second fetus, different from the contribution to the first fetal heart sound mixture; anda contribution from the background noise, different from the contribution to the first fetal heart sound mixture;wherein the third fetal heart sound mixture comprises: a contribution from the first fetal heart sound from ... | The Board of Trustees of the University of Illinois | A61B005/02
US08165896 | B2 | 2012-04-24 | Compliance data for health-related procedures | 2008-03-31 | Exemplary embodiments provide a verification technique that facilitates administration of a health-related procedure to an intended recipient patient or group of patients. An interface template or signal protocol may be configured to establish suitable matching between the patient and various types of objects used to administer the health-related procedure. In some embodiments real-time monitoring data regarding administration of a health-related procedure to a recipient patient is posted to a patient data record that has restricted read/write access. In some instances the monitoring data is processed to determine compliance or non-compliance based on comparison of the health-related procedure with a predetermined benchmark standard. | 1. A system comprising: circuitry configured for receiving monitoring data of at least one health-related parameter and at least one patient identifier determined from at least one patient interface template having identifying surface indicia;circuitry configured for associating the monitoring data of the at least one health-related parameter with the at least one patient identifier; andcircuitry configured for communicating the monitoring data of the at least one health-related parameter for incorporation with at least one patient data record associated with the at least one patient identifier. 2. The system of claim 1, wherein the circuitry configured for receiving monitoring data of at least one health-related parameter and at least one patient identifier determined from at least one patient interface template having identifying surface indicia comprises: circuitry configured for receiving monitoring data substantiating administration of at least one health-related procedure and at least one patient identifier determined from at least one patient interface template having identifying surface indicia. 3. The system of claim 1, wherein the circuitry configured for receiving monitoring data of at least one health-related parameter and at least one patient identifier determined from at least one patient interface template having identifying surface indicia comprises: circuitry configured for receiving monitoring data of at least one health-related parameter comparable with at least one benchmark standard and receiving at least one patient identifier determined from at least one patient interface template having identifying surface indicia. 4. The system of claim 1, wherein the circuitry configured for receiving monitoring data of at least one health-related parameter and at least one patient identifier determined from at least one patient interface template having identifying surface indicia comprises: circuitry configured for receiving monitoring data of at least one health-related parameter associated with one or more substances and at least one patient identifier determined from at least one patient interface template having identifying surface indicia. 5. The system of claim 1, wherein the circuitry configured for receiving monitoring data of at least one health-related parameter and at least one patient identifier determined from at least one patient interface template having identifying surface indicia comprises: circuitry configured for receiving monitoring data of at least one health-related parameter associated with one or more treatments and at least one patient identifier determined from at least one patient interface template having identifying surface indicia. 6. The system of claim 1, wherein the circuitry configured for receiving monitoring data of at least one health-related parameter and at least one patient identifier determined from at least one patient interface template having identifying surface indicia comprises: circui... | The Invention Science Fund I, LLC | G06F019/00 G06Q010/00
US07919546 | B2 | 2011-04-05 | Color integrated and mobile paint systems for producing paint from a plurality of prepaint components | 2008-08-15 | What is disclosed is a plurality of prepaints which are used in any combination to produce base paints and/or colored paints of varying end-use application characteristics. Additionally, a mobile-paint factory for producing a plurality of lines of architectural coatings from a plurality of prepaints en route to and/or at various locations is disclosed. The mobile paint factory may include a vehicle having a paint production system and a paint mixer located thereon. Furthermore, a method of producing a plurality of lines of user-selected architectural coatings from a plurality of prepaints en route to and/or at various locations is disclosed. The method may include: providing a mobile paint factory; prompting a user to select through the paint production system an interior or exterior application environment, a paint sheen, and a paint quality; automatically dispensing predetermined amounts of appropriate prepaints; and agitating the dispensed prepaints. | 1. A method of producing a plurality of lines of user-selected architectural, industrial and/or specialty coatings from a plurality of prepaints, the method comprising: providing a color integrated paint production system, the color integrated paint production system comprising a paint production apparatus and an integrated colorant system;prompting a user to select through the paint production system an interior or exterior application environment;prompting the user to select through the paint production system a paint sheen;prompting the user to select through the paint production system a paint quality; andautomatically gravimetrically and/or volumetrically dispensing through the paint production system predetermined amounts of appropriate prepaints and colorants, wherein none of the prepaints are architectural, industrial and/or specialty coatings prior to being dispensed and wherein the colorants are dispensed prior to the dispensing of the appropriate prepaints and/or simultaneously with the dispensing of the appropriate prepaints and/or after the architectural, industrial and/or specialty coatings have been formed. 2. The method of claim 1, wherein the appropriate prepaints are dispensed gravimetrically and the colorants are dispensed volumetrically. 3. The method of claim 1, wherein both the appropriate prepaints and the colorants are dispensed volumetrically. 4. The method of claim 1, wherein the appropriate prepaints are selected from the group consisting of a pigment prepaint, a dilutant prepaint, an extender prepaint, a low resin binder prepaint, a high resin binder prepaint, an elastomeric prepaint and/or a fire retardant prepaint. 5. The method of claim 1, wherein the method of producing a plurality of lines of user-selected architectural, industrial and/or specialty coatings from a plurality of prepaints is carried out at a point of sale. 6. The method of claim 1, wherein the color integrated paint production system is mobile and further comprises a vehicle. 7. The method of claim 1, wherein the providing a color integrated paint production system further comprises providing a plurality of portable paint containers for receiving dispensed prepaints and colorants. 8. The method of claim 7, wherein the providing a plurality of portable paint containers comprises providing one of quart containers, one gallon containers, two gallon containers, five gallon buckets, 55 gallon drums, and any combination thereof. 9. The method of claim 1, further comprises providing a paint mixer integrated with the colorant system and the paint production apparatus. 10. The method of claim 9, further comprising agitating the dispensed prepaints and colorants in the paint mixer to mix them to produce the architectural, industrial and/or specialty coatings having the selected application environments, the selected sheens, and the selected qualities. 11. The method of claim 10, wherein the providing a color integrated paint production system furthe... | Microblend Technologies, Inc. | G06F017/00 C09C003/04
US07879213 | B2 | 2011-02-01 | Sensor for in vitro determination of glucose | 2007-10-30 | A process for the manufacture of small sensors with reproducible surfaces, including electrochemical sensors. One process includes forming channels in the surface of a substrate and disposing a conductive material in the channels to form an electrode. The conductive material can also be formed on the substrate by other impact and non-impact methods. In a preferred embodiment, the method includes cutting the substrate to form a sensor having a connector portion and a transcutaneous portion, the two portions having edges that define one continuous straight line. | 1. A sensor for the in vitro determination of the presence and/or level of glucose in a sample, the sensor comprising: a substrate comprising a plurality of depressions formed in response to an applied pressure thereon;a working electrode disposed in one of the depressions, said one of the depressions also including a contact pad for electrical coupling with the working electrode, the contact pad configured to couple with a processing unit, wherein the sensor generates a signal associated with a physiological condition for processing by the processing unit based on the presence and/or level of glucose in a volume sample of about 1 microliter or less. 2. The sensor of claim 1, wherein a depth of the depression is in the range of about 12.5 μm to 75 μm. 3. The sensor of claim 2, wherein a depth of the depression is in the range of about 25 μm to 50 μm. 4. The sensor of claim 1, wherein the sensor is coupled to a skin piercer. 5. The sensor of claim 1, including a tapered end. 6. The sensor of claim 1, including an indentation. 7. The sensor of claim 1, comprising a variable width. 8. The sensor of claim 1, wherein the working electrode comprises gold. 9. The sensor of claim 1, further comprising a second electrode. 10. The sensor of claim 9, wherein the second electrode is a second working electrode. 11. The sensor of claim 10, wherein one or more of the first and the second working electrodes comprise gold. 12. The sensor of claim 9, wherein the second electrode is a reference electrode. 13. The sensor of claim 1 wherein the sample includes blood sample withdrawn from under a skin layer. 14. The sensor of claim 1 wherein the depression including the contact pad has a conductive material to electrically couple to the working electrode. 15. The sensor of claim 14 wherein the conductive material includes one or more of carbon, conductive polymer, metal, alloy or metallic compound. 16. The sensor of claim 1 wherein the signal associated with the physiological condition based on the presence and/or level of glucose in the volume sample of about 1 microliter or less is generated for a physiologically relevant range associated with the condition. 17. A glucose sensor, comprising: a substrate including a plurality of depressions formed in response to an applied pressure on a surface of the substrate; anda working electrode on the substrate disposed in one of the depressions, each of the plurality of depressions formed on the substrate including a respective contact pad on the substrate, wherein one of the contact pads on the substrate is electrically coupled to the working electrode, and wherein the contact pads are configured to electrically couple with a processing unit; wherein the sensor generates a glucose related signal for processing by the processing unit when in fluid contact with a blood sample of about 1 microliter or less in volume. 18. The sensor of claim 17 wherein the depressions are formed based at least in part on an impa... | Abbott Diabetes Care Inc. | G01N027/327
US09962098 | B2 | 2018-05-08 | Heart monitor electrode system | 2008-02-06 | A bioelectric interface for monitoring, detection and transmission of detected ECG data is provided comprising a sensory system for detecting bio-physiological measurements utilizing spatially resolved potential profiles obtained from a localized cluster of sub-electrodes to form constituent sets of miniature sensor arrays. Using only a single macro-electrode, two or more sets of sub-electrode arrays are used to measure bipolar spatial gradients obtained from measured cardiac potentials. The sets of sub-electrodes containing the clusters are optimized to attain measurable gradients of diagnostic value. A minimax procedure allows bio-potential sensory acquisition through a bi-directional digital steering process, which essentially comprises monitoring, detection, selection, grouping, recording and transmission of ECG waveform characteristic data. | 1. A non-invasive method of obtaining local gradient values from a bio-potential source for the purpose of biophysiological signal monitoring by affixing a lead-free, autonomous single macro electrode, without regard to anatomical orientation, to a body surface of a subject to be monitored, the macro electrode comprising a group of N number of sub-electrodes, namely, two or more sub-electrodes arranged in a highly localized constellation capable of detecting highly localized, closely spaced biophysiological electrical potential gradients between two or more constituent sets of the sub-electrodes, each set of sub-electrodes comprising at least one distinct sub-electrode of the two or more sub-electrodes; acquiring the biophysiological signal from two or more of the sub-electrodes within the constellation; parsing data obtained from all sub-electrodes disposed upon the single macro electrode with a microprocessor disposed in communication therewith; and determining with said microprocessor a set of sub-electrodes that contributed to the maximum local gradient according to the anatomical orientation of the macro electrode on the body surface of the subject. 2. The method of claim 1 further comprising the step of remote interrogation of the macro electrode carried out using voice data digital signal packaging. 3. The method of claim 2, wherein said remote interrogation of the macro electrode discerns data relating to biophysiological phenomena occurring in muscle tissue. 4. The method of claim 1, wherein the biophysiological signal originates from the subject's heart. 5. The method of claim 1, wherein the macro electrode is affixed to the body surface of the subject without regard to rotational orientation of the macro electrode. 6. The method of claim 1 further comprising the steps of parsing the data from all of the sub-electrodes; determining which of the sub-electrodes, either alone or in combination with another sub-electrode or other sub-electrodes or a combination of other sub-electrodes, provide a maximum gradient value when contrasted against said another sub-electrode or said other sub-electrodes or said combination of sub-electrodes; and using the so determined maximum gradient value for the purpose of the biophysiological signal monitoring. 7. The method of claim 6 wherein the steps of parsing are conducted using a minimax procedure adapted to each individual subject to be monitored in order to align sub-electrodes in a manner so as to attain a measurable divergence between and amongst biophysiological signals. 8. The method of claim 6 wherein the steps of parsing are conducted using a minimax procedure that allows for bio-potential sensory acquisition through a digital steering process, in which monitoring, selecting, grouping, recording and transmission options are derived from permutations of a plurality of the sub-electrodes. 9. The method of claim 6 wherein the sub-electrodes subtend and delimit an area of no more than a few ... | Global Cardiac Monitors, Inc. | A61B005/04 A61B005/0408
US09398877 | B2 | 2016-07-26 | Method and device for optical imaging of a turbid medium | 2006-11-17 | Disclosed is a method and a device for optical imaging of a turbid medium. This method comprises a reference measurement of reference intensities of light emanating from the turbid medium at a reference blood oxygen saturation (SaO2) in the turbid medium and a reference imaging step for reconstructing a reference image of the turbid medium from the measured reference intensities. Furthermore the method comprises a contrast measurement of the contrast intensity of the light emanating from the turbid medium at a contrast blood oxygen saturation (SaO2) level in the turbid medium and a contrast imaging step for reconstructing a contrast image of the turbid medium from the measured contrast intensities. A comparison is made between the contrast image to the reference image of the turbid medium. | 1. A method for optical imaging of a turbid medium present in a human subject, the method comprising: measuring reference intensities of light emanating from the turbid medium while an O2/N2 gas mixture is supplied to the human subject and while the O2/N2 gas mixture has a reference ratio of partial oxygen pressure (pO2) to partial nitrogen pressure (pN2);reconstructing a reference image of the turbid medium from the measured reference intensities;measuring contrast intensities of the light emanating from the turbid medium while the O2/N2 gas mixture is supplied to the human subject and while the O2/N2 gas mixture has a contrast ratio of partial oxygen pressure (pO2) to partial nitrogen pressure (pN2);reconstructing a contrast image of the turbid medium from the measured contrast intensities; andcomparing the contrast image of the turbid medium to the reference image of the turbid medium,wherein the reference ratio is different from the contrast ratio. 2. The method of claim 1, wherein the turbid medium is irradiated by light from source plurality of optical fibers to generate the reference intensities and the contrast intensities of the light emanating from the turbid medium. 3. The method of claim 1, wherein comparing the contrast image of the turbid medium to the reference image of the turbid medium comprises subtracting the reference image from the contrast image. 4. The method of claim 1, wherein one of the reference ratio and the contrast ratio is greater than 0.26 and another of the reference ratio and the contrast ratio is less than 0.26. 5. A method according to claim 1, wherein the O2/N2 gas mixture is administered via a gas mask. 6. The method of claim 1, wherein the reference ratio and the contrast ratio are each in a range from 0.1 to 0.5. 7. The method of claim 1, wherein the reference image is an average reference image generated from a plurality of measurements of the reference intensities of light emanating from the turbid medium at while the O2/N2 gas mixture has the reference ratio, and wherein the contrast image is an average contrast image generated from a plurality of measurements of the contrast intensities of light emanating from the turbid medium at while the O2/N2 gas mixture has the contrast ratio. 8. An apparatus for optical imaging of a turbid medium present in a human subject, the apparatus comprising: an O2/N2-modulator configured to supply an O2/N2 gas mixture to the human subject and to modulate a ratio of partial oxygen pressure (pO2) to partial nitrogen pressure (pN2) in the O2/N2 gas mixture;one or more detector modules configured to detect light emanating from the turbid medium when the O2/N2-modulator causes the O2/N2 gas mixture to have a reference ratio of partial oxygen pressure (pO2) to partial nitrogen pressure (pN2), and to detect light emanating from the turbid medium when the O2/N2-modulator causes the O2/N2 gas mixture to have a contrast ratio of partial oxygen pressure (pO2) to partial nitrog... | KONINKLIJKE PHILIPS N.V. | A61B005/1455 A61B005/00
US20080287765 | A1 | 2008-11-20 | INTEGRATED RECEIVER FOR CONTINUOUS ANALYTE SENSOR | 2008-07-29 | A system is provided for monitoring glucose in a host, including a continuous glucose sensor that produces a data stream indicative of a host's glucose concentration and an integrated receiver that receives the data stream from the continuous glucose sensor and calibrates the data stream using a single point glucose monitor that is integral with the integrated receiver. The integrated receiver obtains a glucose value from the single point glucose monitor, calibrates the sensor data stream received from the continuous glucose sensor, and displays one or both of the single point glucose measurement values and the calibrated continuous glucose sensor values on the user interface. | 1. A system comprising: a continuous sensor configured for at least partial insertion into a host, wherein the continuous sensor generates continuous sensor data associated with an analyte concentration in the host;a single point monitor external to the host, wherein the single point monitor generates at least one single point value associated with an analyte concentration of a biological sample received from the host; anda processing unit configured to calibrate the continuous sensor data using the single point value from the single point monitor, wherein the processing unit is in data communication with the continuous sensor and configured to receive and process the continuous sensor data from the continuous sensor, and wherein the single point monitor is physically connected to the processing unit. 2. The system of claim 1, wherein each of the continuous sensor and the single point monitor comprise a sensing membrane comprising an enzyme domain, and wherein the enzyme domain of the sensing membrane of the continuous sensor and the enzyme domain of the sensing membrane of the single point monitor are substantially the same. 3. The system of claim 1, wherein the single point monitor is configured to receive a single use test strip. 4. The system of claim 1, wherein the processing unit further comprises programming to control administration of a constituent to the host. 5. The system of claim 4, wherein the programming to control administration of a constituent to the host is configured to determine at least one of an amount of insulin and a time of insulin administration. 6. The system of claim 1, wherein the processing unit further comprises programming to calculate a rate of change of the continuous sensor data and/or directional trend of the continuous sensor data. 7. The system of claim 6, wherein the processing unit further comprises programming to control administration of a constituent to the host based at least in part on a rate of change of the continuous sensor data and/or a directional trend of the continuous sensor data. 8. The system of claim 1, wherein the processing unit further comprises programming to control administration of a constituent to the host based at least in part on the at least one single point value. 9. The system of claim 1, wherein the processing unit is configured to monitor the continuous sensor data to adaptively determine a time for requesting a biological sample. 10. The system of claim 1, further comprising a user interface associated with the processing unit and the single point monitor, wherein the processing unit is configured to prompt the user, via the user interface, to provide a biological sample to the single point monitor at predetermined intervals. 11. The system of claim 1, further comprising a user interface associated with the processing unit and the single point monitor, wherein the processing unit is configured to prompt the user to provide a biological sample to the single point mo... | DexCom, Inc. | A61B005/145
US20080319349 | A1 | 2008-12-25 | System and Related Method For Determining a Measurement Between Locations on a Body | 2006-04-17 | A system and related method for characterizing an effect of a rehabilitation therapy on a body. The apparatus includes a first device, which is configured to be coupled to the body at a first location, and a second device, which is configured to be coupled to the body at a second location that is separated from the first location by a distance. The first device is configured to generate a wireless signal. The second device is configured to detect the wireless signal and to generate data based on the detected wireless signal that is configured to be used to calculate the distance. The distance is used to characterize the effect of the rehabilitation therapy on the body. | 1. A system, comprising: a first device configured to be coupled to a body at a first location and configured to generate a wireless signal; anda second device configured to be coupled to the body at a second location, the second location being disposed apart from the first location by a distance; the second device configured to detect the wireless signal and to generate data based on the detected wireless signal, the data generated being configured to be used to calculate the distance,the distance being used to characterize the effect of the rehabilitation therapy on the body. 2. The system according to claim 1, wherein the distance is a first distance, the wireless signal of the first device is a first wireless signal, the first device configured to generate the first wireless signal at a first frequency, the system further comprising: a third device configured to be coupled to the body at a third location, the third location being disposed apart from the second location by a second distance, the third device configured to generate a second wireless signal at a second frequency,the second device configured to detect the second wireless signal and to generate additional data based on the detected second wireless signal, the additional data generated being configured to be used to calculate the second distance,the second distance being used to characterize the effect of the rehabilitation therapy on the body. 3. The system according to claim 2, wherein the third device is configured to generate the second wireless signal at the same time that the first device is configured to generate the first wireless signal. 4. The system according to claim 3, wherein the second device is configured to detect in a selectable manner a signal selected from the group consisting of the first wireless signal and the second wireless signal. 5. The system according to claim 1, further comprising: the an external device configured to communicate with the second device,the second device configured to communicate the data to the external device, andthe external device configured to calculate the distance based on the data. 6. The system according to claim 5, wherein the external device is configured to calculate an angle of orientation between the second device and the first device based on the data. 7. The system according to claim 5, wherein the external device is configured to communicate with the second device via a wireless communication path selected from the group consisting of a radio frequency path, an electrical current path through the body, a path configured for the communication of modulated sonic waves, a path configured for the communication of modulated ultrasonic waves, and an optical communication path. 8. The system according to claim 5, wherein the external device is configured to calculate a value selected from the group consisting of an average of the first distance over a period of time, a standard deviation of the first distance over a per... | - | A61B005/103
US20080294014 | A1 | 2008-11-27 | PROCESS FOR DIAGNOSING AND TREATING A PSYCHOLOGICAL CONDITION OR ASSIGNING A PERSONALITY CLASSIFICATION TO AN INDIVIDUAL | 2008-05-21 | The present invention is directed to methods of diagnosing a psychological condition and/or the events or perceived events that are the basis of the condition. In another embodiment the present invention is directed to a method of assigning a personality classification to an individual. The methods of the invention preferably comprise the steps of: asking the individual a plurality of questions designed to bring out an emotional response in the individual; observing all repetitive physical movements by the individual while the individual is responding to the questions asked; identifying nonverbal leaks (NVL) from the observed repetitive physical movements reflecting an unresolved perceived trauma in the individual; and diagnosing a psychological condition in the individual or assigning a personality classification to the individual based on the individual's identified NVLs. The present invention is further directed to identifying the underlying cause of the perceived trauma, thereby treating the individual. | 1. A method of diagnosing a psychological condition in an individual comprising the steps of: placing the individual in direct view of a display device having a real-time frontal image of the individual and/or a recording device for review at a later time;asking the individual a plurality of questions designed to bring out an emotional response in the individual while the individual watches the real-time image generated on the display and/or while recording the individuals image;observing all repetitive physical movements by the individual while the individual is responding to the questions asked;identifying from the observed repetitive physical movements nonverbal leaks (NVL) reflecting an unresolved perceived trauma in the individual; anddiagnosing a psychological condition by correlating the individual's NVL to a psychological condition or an underlying perceived trauma. 2. The method of claim 1, further comprising the step of hypnotizing the individual and asking the individual further questions while hypnotized after identifying the individual's NVL. 3. The method of claim 1, further comprising decoding the unconscious meaning of the NVL using one or more symbolic levels, including: Symbolic Level One, (SL-1), the impact of a perceived traumatic event on the individual; Symbolic Level Two (SL-2), the primary emotion the individual has to the perceived traumatic event; and Symbolic Level Three (SL-3), the primary coping strategy to the perceived traumatic event. 4. The method of claim 3, wherein the SL-1 impact is manifested by way one of the following: shock (eyes large), fear (teariness), denial (eyes up, trance, white below eyes), disbelief (eyes closing), pain (tearing, turning away from) and trance (eyes that stare in an unfocused way); the SL-2 primary emotion. 5. The method of claim 3, further comprising identifying the perceived traumatic event. 6. The method of claim 1, wherein the display device is a television or computer monitor and asking the questions at least three times in varied context. 7. The method of claim 1, wherein the real-time frontal image of the individual is generated by a video recorder or digital camera recording the image of the individual. 8. The method of claim 1, wherein a recording is made of the individual during questioning and the step of identifying the individuals NVL comprises sitting with the individual and watching the recording using slow motion and/or still frame to correlate the repetitive physical movements with the questions asked. 9. The method of claim 8, wherein the recording of the image of the individual is from the shoulders up. 10. The method of claim 9, wherein the NVL comprises at least one repetitive physical movement that occurs during questioning selected from the group consisting of: eyes held open, eyes not focused; slack jaw, tongue showing, lowering of head, shallow breathing, shoulders downward, eyes teary, corners of mouth downward, eyes closing, eyebrows up, eyes clos... | - | A61B005/00
US20080262381 | A1 | 2008-10-23 | Infant SID Monitor Based On Accelerometer | 2008-04-14 | Techniques, devices and systems that monitor the orientation and breathing of an infant and wirelessly communicate the orientation/breathing data to a caregiver through a wireless interface to request intervention if an unsafe situation is detected. | 1. An infant monitor system, comprising: a infant sensor module; anda control module in wireless communication with the infant sensor module,wherein the infant sensor module comprises:a base attachable to an infant;three accelerometers on the base to measure accelerations along three different directions, respectively;a signal processor on the base to receive outputs from the three accelerometers and to produce a sensor signal based on the outputs; anda sensor RF transceiver on the base in communication with the signal processor to wirelessly transmit the sensor signal to the control module, andwherein the control module comprises:a controller RF transceiver in wireless communication with the sensor RF transceiver;a controller processor operable to process the sensor signal to measure an orientation and a motion of the infant sensor module and to generate an alert when the measured orientation matches a pre-determined orientation for triggering the alert or when the measured motion matches a pre-determined motion profile for triggering the alert; anda communication interface to communicate the alert from the controller processor to a destination outside the control module. 2. The system as in claim 1, wherein the three accelerometers in the infant sensor module are MEMS devices. 3. The system as in claim 1, wherein the communication interface in the controller module includes at least one of a radio transceiver interface with a cell phone network,a modem connected to a land line,an Ethernet card with a computer network such as the Internet, andan RF transceiver module to wirelessly communicate with a local RF network. 4. The system as in claim 1, wherein the controller processor performs signal filtering on digital data of the sensor signal and fast Fourier transform on the filtered digital data in generating the alert. 5. The system as in claim 1, wherein the controller processor computes a root mean square value of the sensor signal and compares the root mean square value of the sensor signal to a pre-determined threshold value in generating the alert. 6. The system as in claim 1, comprising a temperature sensor in contact with the infant to measure a skin temperature, wherein the controller processor is connected to receive the temperature measurement from the temperature sensor and to produce a temperature alert signal when the temperature is below a low temperature threshold or above a high temperature threshold. 7. A method for monitoring an infant, comprising: attaching a sensor module to the infant to monitor an orientation and a motion of the infant sensor module by using three accelerometers in the sensor module to measure accelerations along three different directions, respectively;processing the outputs from the three accelerometers to produce a sensor signal;wirelessly transmitting the sensor signal to a control module;operating the control module to process the sensor signal to measure to generate an alert when the measured or... | MAGNETO INERTIAL SENSING TECHNOLOGY, INC. | A61B005/11 A61B005/01
US20080269616 | A1 | 2008-10-30 | MIR SPECTROSCOPY OF TISSUE | 2007-11-16 | Disclosed are methods of determining long-term deposition pattern of a compound in tissue. The following steps can be followed: placing tissue against a receptor; directing mid-infrared electromagnetic radiation onto the tissue; quantifying the electromagnetic radiation that is reflected from the tissue to obtain a reflected amount; using a calibration equation to calculate the concentration of a compound from the reflected amount; and using the concentration of the compound to evaluate presence of a clinical condition in the tissue. | 1. A method of accurately measuring body cholesterol levels without inaccuracies related to recent diet, the method comprising: irradiating skin tissue with mid-infrared radiation to excite lipid molecules within superficial layers of the tissue;measuring resulting mid-infrared radiation;comparing the measured radiation to stored reference data to obtain correlation information; andusing the correlation information to determine body lipid levels without inaccuracies related to recent diet. 2. The method of claim 1, wherein measuring resulting mid-infrared radiation comprises measuring reflected mid-infrared radiation. 3. The method of claim 1, wherein measuring resulting mid-infrared radiation comprises measuring transmitted mid-infrared radiation. 4. The method of claim 1, wherein irradiating skin tissue with mid-infrared radiation to excite lipid molecules within superficial layers of the tissue comprises irradiating skin tissue with mid-infrared radiation to excite all cholesterol molecules within superficial layers of the tissue. 5. The method of claim 1, wherein irradiating skin tissue with mid-infrared radiation to excite lipid molecules within superficial layers of the tissue comprises irradiating skin tissue with mid-infrared radiation to excite esterified cholesterol molecules within superficial layers of the tissue. 6. The method of claim 1, wherein irradiating skin tissue with mid-infrared radiation to excite lipid molecules within superficial layers of the tissue comprises irradiating skin tissue with mid-infrared radiation to excite free cholesterol molecules within superficial layers of the tissue. 7. The method of claim 1, wherein irradiating skin tissue with mid-infrared radiation to excite lipid molecules within superficial layers of the tissue comprises irradiating skin tissue with mid-infrared radiation to excite free fatty acids within superficial layers of the tissue. 8. The method of claim 1, wherein irradiating skin tissue with mid-infrared radiation to excite lipid molecules within superficial layers of the tissue comprises irradiating skin tissue with mid-infrared radiation to excite ceramides within superficial layers of the tissue. 9. The method of claim 1, further comprising using a lipid level in a body to assess a disease. 10. The method of claim 9, wherein using the lipid level to assess a disease comprises using body cholesterol to diagnose psoriasis. 11. The method of claim 9, wherein using the lipid level to assess a disease comprises using esterified cholesterol to diagnose psoriasis. 12. The method of claim 9, wherein using the lipid level to assess a disease comprises using a ceramide to diagnose psoriasis. 13. The method of claim 9, wherein using the lipid level to assess a disease comprises using a free fatty acid level to diagnose psoriasis. 14. The method of claim 9, wherein using the lipid to assess a disease comprises using body cholesterol to assess risk of cardiovascular disease. 15. A me... | - | A61B005/1455
US20080287818 | A1 | 2008-11-20 | PRESSURE MEASUREMENT-BASED ISCHEMIA DETECTION | 2007-04-19 | In some aspects ischemia is indicated based on cardiac pressure measurements. For example, ischemia may be indicated based on an increase in an intraventricular electromechanical delay where mechanical contractions are detected by measuring pressure in a cardiac chamber. Ischemia detection also may involve obtaining timing information relating to a mechanical contraction of at least one ventricle to identify an intraventricular dyssynchrony. In this case, pressure measurements may be used to identify the timing of the mechanical contraction. Ischemia may be indicated based on a change in a time interval associated with a systolic interval of a ventricle. Ischemia also may be indicated based on an intracardiac electrogram-based indication of ischemia in conjunction with an increase in mean left atrial pressure and/or an increase in the size of a v-wave. | 1. A method of generating an indication of ischemia associated with a heart of a patient, comprising: determining an electrical depolarization to mechanical contraction delay for a ventricle of a heart;determining whether the delay is more than an expected delay; andgenerating an indication of ischemia in accordance with a determination that the delay is more than the expected delay. 2. The method of claim 1, wherein determining whether the delay is more than the expected delay comprises comparing the delay to a baseline delay value. 3. The method of claim 2, further comprising adjusting the baseline delay value in accordance with an activity level and/or a position of the patient. 4. The method of claim 1, wherein determining whether the delay is more than the expected delay comprises determining that electrical depolarization to mechanical contraction delay for the ventricle has increased. 5. The method of claim 1, wherein determining whether the delay is more than the expected delay comprises determining whether there has been an increase in a percentage of a cardiac cycle associated with electrical depolarization to mechanical contraction delay for the ventricle. 6. The method of claim 1, wherein determining the electrical depolarization to mechanical contraction delay comprises measuring a delay between an R-wave and an upslope of a pressure wave for the ventricle. 7. The method of claim 1, wherein determining the electrical depolarization to mechanical contraction delay comprises measuring left atrial pressure. 8. The method of claim 1, wherein the ventricle comprises a right ventricle or a left ventricle. 9. The method of claim 1, wherein generating the indication of ischemia further comprises corroborating an intracardiac electrogram-based indication of ischemia. 10. A method of generating an indication of ischemia associated with a heart of a patient, comprising: determining timing of an onset of mechanical contraction for at least one ventricle of a heart by measuring pressure associated with the at least one ventricle;identifying an intraventricular dyssynchrony in accordance with the timing of the onset of mechanical contraction for the at least one ventricle; andgenerating an indication of ischemia in accordance with the identification of an intraventricular dyssynchrony. 11. The method of claim 10, wherein determining timing of the onset of mechanical contraction for at least one ventricle comprises determining that contraction of a left ventricle follows contraction of a right ventricle by a time interval that is greater than an expected time interval. 12. The method of claim 10, wherein determining timing of the onset of mechanical contraction for at least one ventricle comprises determining whether there is an increase in a percentage of a cardiac cycle associated with a time period between contraction of a left ventricle and contraction of a right ventricle. 13. The method of claim 10, wherein determining timing of t... | PACESETTER, INC. | A61B005/02
US20080269636 | A1 | 2008-10-30 | System for and Method of Conveniently and Automatically Testing the Hearing of a Person | 2005-06-10 | System and method for conducting a hearing test that is accessible to a mass market of individuals with potential hearing loss. The hearing test is stored on a centrally located computer (140) that is accessible via communications device (121). The system provides step-by-step guidance on the next steps to be taken if hearing loss is found, and provides a means to store and organize the user test data to create a means for reuse of the data. | 1. A remotely accessible data storage media for use in testing hearing of an individual comprising: a plurality of hearing test queries and instruction data, wherein the instruction data includes instructions for operating a local hearing test unit to perform a hearing test using the hearing test queries on the media,wherein the hearing test queries include at least one of a set of frequency versus amplitude hearing test and speech intelligibility queries, wherein the speech intelligibility queries are selectably accessible, andwherein the instruction data includes instructions linked to at least one of the frequency versus amplitude queries and identifying selected user inputs associated with results of the frequency versus amplitude queries and corresponding to selected ones of the speech queries. 2. The media of claim 1 further comprising: a memory for storing user information and user hearing test results obtained from performing the hearing testing with the test queries of the media, wherein the user information and the user hearing test results are generated at the hearing test unit. 3. The media of claim 1 further comprising: incentive data linked to selected ones of the instruction data. 4. The media of claim 3, wherein the incentive data includes an electronic coupon. 5. The media of claim 3, wherein the instruction data indicate an end of a hearing test or a code corresponding to a hearing loss profile for the user obtained based on the results of a preliminary hearing test. 6. The media of claim 1, wherein at least one of the hearing test queries is indexed with a hearing test code, wherein the code represents the results of a preliminary hearing test on the individual. 7. The media of claim 6 further comprising: normal and modified word units having the same index as one of the hearing test codes or one of the selected user inputs identified in the instruction data. 8. A system for performing a hearing test including: a central controller and a hearing test unit, wherein the controller is coupled to a remotely accessible data storage media and wherein the controller and media are located remotely from the hearing test unit;wherein the media comprises: a plurality of hearing test queries and instruction data, wherein the instruction data includes instructions for operating the hearing test unit to perform a hearing test using the hearing test queries on the media,wherein the hearing test queries include at least one of a set of frequency versus amplitude hearing test and speech intelligibility queries, wherein the speech intelligibility queries are selectably accessible, andwherein the instruction data includes instructions linked to at least one of the frequency versus amplitude queries and identifying selected user inputs associated with results of the frequency versus amplitude queries and corresponding to selected ones of the speech queries; andwherein each of the controller and the hearing test unit includes means for provi... | Johnson & Johnson Consumer Companies, Inc. | A61B005/00
US20080255463 | A1 | 2008-10-16 | Method of Obtaining an Estimation of a Person's Aortic Blood Pressure | 2006-01-04 | A method of obtaining an estimation of a person's aortic systolic blood pressure, which method comprises non-invasively measuring the person's brachial systolic blood pressure, non-invasively measuring the person's brachial diastolic blood pressure, obtaining the person's radial augmentation index by measuring the person's radial pulse waveform, and obtaining the estimation of the person's aortic systolic blood pressure from the following equation: aSBP=α+β·rA1x+γ·bSBP+δ·bDBP where the coefficients α, β, γ and δ are constants with approximate values of α=−24.2, β=0.28, γ=0.83, δ=0.17. | 1. A method of obtaining an estimation of a person's aortic systolic blood pressure, which method comprises non-invasively measuring the person's brachial systolic blood pressure, non-invasively measuring the person's brachial diastolic blood pressure, obtaining the person's radial augmentation index by measuring the person's radial pulse waveform, and obtaining the estimation of the person's aortic systolic blood pressure from the following equation: aSBP=α+β·rA1x+γ·bSBP+δ·bDBP where the coefficients α, β, Y and δ are constants with approximate values of α=−24.2=β=0.28Y=0.83Δ=0.17 2. A method according to claim 1 in which the person's brachial systolic blood pressure is non-invasively measured using a sphygmomanometer. 3. A method according to claim 1 in which the person's brachial diastolic blood pressure is non-invasively measured using a sphygmomanometer. 4. A method according to claim 1 in which the person's radial pulse waveform is measured non-invasively. 5. A method according to claim 4 in which the person's radial pulse waveform is obtained using a tonometer. 6. A method according to claim 1 in which the person's radial pulse waveform is measured invasively. 7. A method according to claim 6 in which the person's radial pulse waveform is obtained using by placing a canula in the radial artery and attaching a pressure transducer to the canula. 8. (canceled) | MICRO MEDICAL LTD. | A61B005/022 A61B005/021 A61B005/0215
US20080287809 | A1 | 2008-11-20 | Biological information measuring apparatus and biological information measuring method | 2008-04-23 | A biological information measuring apparatus measures information of a living body by irradiating the living body with light. The apparatus includes: a first light source irradiating the living body with light; a second light source irradiating the living body with light; and a light receiving section constituted by a plurality of light receiving elements that receives reflected light reflected by the living body or transmitting light transmitting through the living body, during light irradiation of the living body. In the biological information measuring apparatus, the first light source and the second light source are disposed so that a first distance between the light receiving section and the first light source is different from a second distance between the light receiving section and the second light source. | 1. A biological information measuring apparatus for measuring information of a living body by irradiating the living body with light, the apparatus comprising: a first light source irradiating the living body with light;a second light source irradiating the living body with light; anda light receiving section constituted by a plurality of light receiving elements that receives reflected light reflected by the living body or transmitting light transmitting through the living body, during light irradiation of the living body, whereinthe first light source and the second light source are disposed so that a first distance between the light receiving section and the first light source is different from a second distance between the light receiving section and the second light source. 2. The biological information measuring apparatus according to claim 1, wherein the first distance and the second distance are determined depending on a depth in the living body to a point to be measured. 3. The biological information measuring apparatus according to claim 1, wherein the first light source and the second light source are configured so as to surround the light receiving section. 4. The biological information measuring apparatus according to claim 3, wherein the first light source and the second light source are configured in an annular form. 5. The biological information measuring apparatus according to claim 4, wherein the first light source and the second light source are configured in a circular form. 6. The biological information measuring apparatus according to claim 5, wherein the first light source and the second light source are configured in a concentrical form. 7. The biological information measuring apparatus according to claim 1, further comprising: an uneven filter disposed on an optical path of the transmitting light or an optical path of the reflected light, whose light transmittance is set so that the intensity of the light received by the light receiving elements is uniform in the entire light receiving section. 8. A biological information measuring method for measuring information of a living body by irradiating the living body with light, the method comprising: providing a biological information measuring apparatus including: a first light source irradiating the living body with light; a second light source irradiating the living body with light; and a light receiving section constituted by a plurality of light receiving elements that receives reflected light reflected by the living body or transmitting light transmitting through the living body, during light irradiation of the living body, the first light source and the second light source being disposed so that a first distance between the light receiving section and the first light source is different from a second distance between the light receiving section and the second light source;acquiring first light intensity distribution information by the light receiving section whil... | SEIKO EPSON CORPORATION NATIONAL UNIVERSITY CORPORATION HOKKAIDO UNIVERSITY | A61B005/00
US20080262331 | A1 | 2008-10-23 | INFECTION MONITORING | 2007-04-19 | Implantable medical devices include a sensor capable of detecting an indicator of infection in proximity to the device. Sensed information regarding the indicator of infection is provided to the device for a time period of between two hours and six months after the device is implanted. A determination may be made as to whether the sensed information is indicative of infection. The determination may be made in the implanted device or in an external device in wireless communication with the implanted device. If the information is indicative of infection, an alert may be issued from the implanted device or an external device. | 1. A method comprising: commencing sensing, in proximity to an implanted medical device, information regarding an indicator of infection;providing the information regarding the indicator to the implanted medical device;determining whether the sensed information is indicative of infection;determining whether a predetermined time period of between 2 hours and 6 months has elapsed since the sensing has commenced; andceasing the sensing if the predetermined time period has elapsed. 2. The method of claim 1, wherein determining whether the sensed information is indicative of an infection is determined within the implanted device. 3. The method of claim 1, further comprising sending the information provided to the implanted device to an external device. 4. The method of claim 3, wherein sending the sensed information provided to the implanted device to an external device comprises sending the information to a programmer device. 5. The method of claim 3, wherein determining whether the sensed information is indicative of an infection is performed in the external device. 6. The method of claim 1, wherein providing information regarding the sensed indicator to an implanted medical device comprises providing the information in discrete time intervals. 8-21. (canceled) 22. The method of claim 1, wherein sensing the information comprises sensing temperature. 23. The method of claim 22, wherein the sensing the information further comprises sensing one or more of impedance, pH, and a biological marker of infection. 24. The method of claim 1, wherein the sensing the information comprises sensing one or more of impedance, pH, and a biological marker of infection. 25. The method of claim 1, further comprising storing the provided information in a memory of the device. 26. The method of claim 1, wherein determining whether the sensed information is indicative of infection comprises determining in discrete time intervals. 27. (canceled) 28. The method of claim 1, wherein sensing the information regarding the indicator of infection comprises sensing the information in discrete time intervals. 29. The method of claim 1, wherein the predetermined time period is a time period of between twelve hours and three months. 30. The method of claim 1, wherein the predetermined time period is a time period of between one day and one month. 31. The method of claim 1, wherein providing information regarding the sensed indicator to an implanted medical device comprises transmitting the information to the device from a sensor operably coupled to the device. 32. The method of claim 1, further comprising issuing an alert if the indicator of infection is indicative of an infection. 33. A computer readable medium containing instructions that when implemented cause an implantable medical device to (i) sense information regarding an indicator of infection; (ii) provide the information regarding the sensed indicator to the implanted medical device; (iii) determine whethe... | MEDTRONIC, INC. | A61B005/145 A61B005/01 A61B005/053
US07700932 | B2 | 2010-04-20 | Method, set, and apparatus for obtaining prints of a part of the human body | 2008-08-14 | A part of a human body is caused to contact the surface layer of a recording medium over which a plurality of metallic grains with an outside size of 200 nm or less are distributed. Then, secretions from the skin surface of the body part are caused to adhere to the surface layer of the recording medium to take the print of the body part. If light is irradiated to the recording medium, specific optical characteristics resulting from the surface structure of the recording medium are obtained, and therefore the color of the recording medium varies between a region having secretions and a region having no secretions. This renders it possible to record a visible print on the recording medium. | 1. A set for obtaining prints of a part of the human body, comprising: a sheet member;a recording medium in which a plurality of metallic grains with an outside size of 200 nm or less are distributed over a surface layer, the recording medium being stuck on the sheet member;a recordable blank region disposed on the sheet member; anda transparent plastic film, comprising a material that has a refractive index different from any refractive index of secretions from human skin, wherein:an end of the transparent plastic film is stuck on the sheet member in advance of taking a print, andthe transparent plastic film except the end is stuck on the surface layer of the recording medium after a print is taken, whereby secretions on the recording medium are covered with the transparent plastic film. 2. A set as defined by claim 1, wherein said transparent plastic film is a transparent film of polyethylene terephthalate or a transparent film of polycarbonate. 3. A set as defined by claim 1, wherein the transparent plastic film is movable to provide positioning prior to taking the print and after taking the print. | FUJIFILM Corporation | G06K005/00 A61B005/103
US07457658 | B2 | 2008-11-25 | Algorithm for accurate three-dimensional reconstruction of non-linear implanted medical devices in VIVO | 2003-03-06 | A computer-implemented software system and method for performing dynamic, three-dimensional reconstruction of a non-linear implanted medical device from biplane x-ray images. A biplane radiographic image of a calibration object is used to compute the relationship between a three-dimensional coordinate system and the two-dimensional coordinate system for each image plane. These relationships are used to superimpose a three-dimensional device model onto each biplane image pair. An iterative, active contour model, initiated by a user-specified curved template, solves for the three-dimensional device centerline coordinates. Device reconstruction is repeated at each time point associated with each image pair. Algorithm output is available for test development, structural analysis or other clinically relevant applications. | 1. A system for reconstructing a dynamic, three-dimensional contour model of a non-linear shaped implanted medical device, comprising: an imaging device generating biplane imaging data of an implanted medical device;a data storage unit storing the generated biplane imaging data;a computer work station processing the stored image data, the work station including a display for displaying acquired biplane images and for displaying the implanted device contour model;a user interface coupled to the work station receiving user-entered data regarding implanted medical device geometry including a user-specified, curved non-linear template corresponding to the non-linear implanted medical device prior to implantation; andmeans resident on the work station for directing the execution of an iterative process of implanted medical device contour model that begins with the user-specified, curved non-linear template,wherein the means resident on the work station for directing the execution of an iterative process of implanted medical device contour model includes means for orienting the template along an image of the implanted medical device based upon user selection of a plurality of points along the image and generates three-dimensional centerline point coordinates for a pair of biplane images of the implanted medical device. 2. The system of claim 1, wherein the selected plurality of points correspond to identifiable landmarks of the implanted medical device. 3. The system of claim 1, wherein at least one point of the selected plurality of points is offset from the implanted medical device. 4. The system of claim 3, wherein the at least one point corresponds to a center of radius of the implanted medical device. 5. The system of claim 1, wherein the implanted medical device is one of a catheter, a stent and a heart valve device. 6. The system of claim 1, wherein the biplane imaging data comprises a plurality of time-paired image planes obtained during an imaging sequence. 7. The system of claim 6, wherein the imaging sequence comprises at least one cardiac cycle. 8. The system of claim 6, wherein the execution of an iterative process of implanted medical device contour model is made at each point in time during an imaging sequence. 9. The system of claim 1, wherein active contour model points along the centerline of the template are projected onto each of two biplane images of the implanted medical device. 10. The system of claim 1, wherein the means resident on the work station for directing the execution of an iterative process of implanted medical device contour model generates three-dimensional centerline point coordinates for a second pair of biplane images of the implanted medical device after generating three-dimensional centerline point coordinates for the first pair of biplane images of the implanted medical device. 11. The system of claim 1, wherein the biplane imaging data comprises a plurality of time-paired image planes obtained during... | Medtronic, Inc. | A61B005/05
US07440791 | B2 | 2008-10-21 | Double echo magnetic resonance imaging sequence and apparatus for the implementation thereof | 2002-11-19 | In a magnetic resonance imaging method and apparatus, a double echo sequence having a gradient echo and a spin echo is implemented with a flow compensation for the gradient echo in at least one direction. The flow compensation gives the double echo sequence an insensitivity to flow phenomena. | 1. A method for operating a magnetic resonance imaging apparatus comprising the steps of: executing a double echo sequence that is a combination of a FISP sequence and a PSIF sequence, including generating a triple gradient pulse in a readout direction having an operation sign that changes from pulse-to-pulse, to read out a gradient echo and a spin echo that are non-superimposed and that are separate from each other; andimplementing a flow compensation for said gradient echo in at least to said readout direction to give said double echo sequence an insensitivity to flow phenomena. 2. A method as claimed in claim 1 comprising implementing said flow compensation in a slice selection direction. 3. A method as claimed in claim 2 comprising, for said flow compensation, generating a triple gradient pulse having an operation sign that changes from pulse-to-pulse. 4. A method as claimed in claim 1 comprising implementing said flow compensation in a phase-coding direction. 5. A method as claimed in claim 4 comprising, for said flow compensation, generating a bipolar gradient pulse. 6. A method as claimed in claim 1 comprising executing said double echo sequence with said flow compensation to obtain data for an orthopedic diagnosis. 7. A magnetic resonance apparatus comprising: a magnetic resonance scanner including a gradient coil system and a radio-frequency system; anda sequence controller configured to operate said magnetic resonance scanner to execute a double echo sequence that is a combination of a FISP sequence and a PSIF sequence, including being configured to generate a triple gradient pulse in a readout direction having an operation sign that changes from pulse-to-pulse, to read out a gradient echo and a spin echo that are non-superimposed and that are separate from each other, and being configured to implement a flow compensation for said gradient echo in at least said readout direction to give said double echo sequence an insensitivity to flow phenomena, to obtain data for producing a magnetic resonance image. 8. A magnetic resonance apparatus as claimed in claim 7 wherein said sequence controller operates said scanner to implement said flow compensation in a slice selection direction. 9. A magnetic resonance apparatus as claimed in claim 8 wherein said sequence controller operates said scanner to generate a triple gradient pulse, for said flow compensation, having an operational sign that changes from pulse-to-pulse. 10. A magnetic resonance apparatus as claimed in claim 7 wherein said sequence controller operates said scanner to implement said flow compensation in a phase-coding direction. 11. A magnetic resonance apparatus as claimed in claim 10 wherein said sequence controller operates said scanner to generate a bipolar gradient pulse for said flow compensation. 12. A magnetic resonance apparatus as claimed in claim 7 wherein said sequence controller operates said scanner to obtain said data for an orthopedic diagnosis. | Siemens Aktiengesellschaft | A61B005/05
US07470234 | B1 | 2008-12-30 | Monitoring device, method and system | 2006-12-03 | A monitoring device (20) for monitoring the vital signs of a user is disclosed herein. The monitoring device (20) is preferably an article (25) having an optical sensor (30) and a circuitry assembly (35). The monitoring device (20) preferably provides for the display of the following information about the user: pulse rate; blood oxygenation levels; calories expended by the user of a pre-set time period; target zones of activity; time; distance traveled; and/or dynamic blood pressure. The article (25) is preferably a band worn on a user's wrist, arm or ankle. | 1. A method of monitoring a user's vital signs, the method comprising: deactivating a processor for a deactivation period of 128 to 640 microseconds;activating an optical sensor during the deactivation period to generate a signal corresponding to the flow of blood through at least one artery of the user;reactivating the processor to process the signal to obtain an average pulse value of a first set of time periods;processing the signal to obtain an average pulse value of a second set of time periods, the second set of time periods less than the first set of time periods;subtracting the average pulse value of the second set of time periods from the average pulse value of the first set of time periods to generate a first filtered pulse data value;processing the first filtered pulse data value to obtain a first heart rate value;using the first heart rate value to filter the subsequent average pulse values generated from the digital signal to obtain a real-time heart rate value of the user;processing the real-time heart rate data of the user for analysis of real-time calories expended by the user and for real-time display of a plurality of the user's vital signs; anddisplaying the plurality of user's vital signs on a display member, the display of the plurality of user's vital signs controlled by the user using a control component. 2. The method according to claim 1 further comprising filtering motion noise from the signal utilizing a motion sensor. 3. The method according to claim 1 wherein the optical sensor comprises a pulsed infrared LED and a photodetector. 4. The method according to claim 1 wherein the optical sensor comprises a pulsed infrared LED, a green light LED and a photodetector. 5. The method according to claim 1 wherein the real-time heart rate data is based on 60 to 180 sample readings and each of the sample readings is based on 4 to 15 sub-sample readings. | Impact Sports Technology, Inc. | A61B005/00
US07447535 | B2 | 2008-11-04 | Mapping the coronary arteries on a sphere | 2004-07-26 | A coronary arteries tree is approximated by a base sphere (32) which is best fitted to vessels centerlines (38). The base surface (32) is gridded to define pixels (52). The base sphere (32) is mapped to fit the centerlines (38) such that a true form surface (56) is determined. A wall thickness to the true form surface (56) is defined, preferably, by a user. A normal of each pixel (52) is searched for grayscale values of voxels. Each pixel (52) is assigned a maximum of grayscale values of voxels within the defined wall thickness intersected by the corresponding normal. The resulting true form surface is undistorted mode of visualization revealing the arteries tree in its context running on the true surface drawn through the vessels. Mapping the assigned grayscale values onto the base sphere (32) visualizes arteries tree on a globe surface (84) which might be rotatably inspected as a globe. Mapping the assigned grayscale values into a flat surface visualizes arteries tree on a two-dimensional map. | 1. A diagnostic imaging system for displaying a vessel tree comprising: a means for defining a base surface, wherein the base surface is a sphere or an ellipsoid, the means for defining a base surface including: a means for determining vessels centerlines from image data;a means for mapping the base surface to the centerlines to define a true form surface;a means for gridding the base surface to define pixels;a means for projecting along a normal of each pixel; anda means for assigning each pixel a grayscale value based on grayscale value of voxels in the image data intersected by a corresponding normal. 2. The system as set forth in claim 1, further including: a means for defining a wall thickness to the true form surface. 3. The system as set forth in claim 2, wherein the grayscale assigning means assigns each pixel a maximum of grayscale values of voxels within the defined wall thickness intersected by the corresponding normal. 4. The system as set forth in claim 1, further including: a means for determining a globe surface including a means for mapping the assigned grayscale values into a spherical surface. 5. The system as set forth in claim 4, further including: a means for projecting the globe surface into a two dimensional surface. 6. The system as set forth in claim 5, wherein the projecting means includes: a matching means which matches coordinates of the spherical surface to coordinates of the two dimensional surface; and2D grayscale processor which assigns each pixel on the two dimensional surface a grayscale value assigned to at least one corresponding pixel on the globe surface. 7. The system as set forth in claim 6, further including: a means for selecting at least one of the true form surface, the globe surface and the two-dimensional surface for displaying on a monitor. 8. A diagnostic imaging apparatus comprising: a scanner which examines a region of a subject including coronary arteries and acquires three-dimensional data;a reconstruction processor for reconstructing the three-dimensional image data into a volumetric three-dimensional image representation;the diagnostic imaging system of claim 1 for converting a portion of the three dimensional image representation into a coronary arteries tree display; anda display connected to the diagnostic imaging system of claim 1 for displaying the coronary arteries tree in a context of the region of interest. 9. A system, comprising: a base surface processor that approximates a spherically-shaped base surface;a volume selector that selects a volume of image data from a volume memory;a centerlines determiner that finds centerlines of vessels in the selected volume of image data;a best fitting surface process that draws a spherically shaped best fitted surface to the determined centerlines;a gridder that spreads a grid over the base surface, thereby gridding the sphere into pixels;a projector that projects a normal from each pixel; andan assigner that assigns each pixel a grayscale... | Koninklijke Philips Electronics N.V. | A61B005/05 G06K009/00
US07435248 | B2 | 2008-10-14 | Medical probes for creating and diagnosing circumferential lesions within or around the ostium of a vessel | 2003-09-26 | The present inventions provide assemblies, probes, and methods for creating circumferential lesions in tissue, e.g., the tissue within or around the ostium of a vessel. An ablation probe with an ablative structure can be placed in contact within or around the ostium of the vessel. A diagnostic probe can be introduced through a lumen within the ablation probe and inserted into the vessel. The energy can be provided to the ablative structure to create a circumferential lesion within or around the ostium of the vessel, and the diagnostic structure can be used to diagnose the tissue to determine whether the circumferential lesion can be properly created. | 1. A probe for ablating tissue, comprising: an outer elongate probe body including a distal ablative structure having an open architecture defining an interior space, wherein the ablative structure tapers downward to a distal tip insertable into the ostium of a pulmonary vein;a lumen extending through the outer probe body, the lumen configured to slidably receive an inner elongate probe body, such that the inner elongate probe body can move independently of the distal ablative structure, the lumen having an exit port out which the inner probe body can extend within the interior space; andone or more ablative elements mounted to the distal ablative structure, wherein the ablative structure is configured for engaging a portion of the pulmonary vein just distal to the ostium of the pulmonary vein, so that the one or more ablative elements are arranged to create a circumferential lesion within the pulmonary vein. 2. The probe of claim 1, wherein the exit port is proximal to the distal ablative structure. 3. A probe assembly for ablating tissue, comprising: an outer probe including an elongate probe body having a distal ablative structure, a lumen having an exit port, and one or more ablative elements mounted to the distal ablative structure, wherein the ablative structure is configured for engaging a portion of an anatomical vessel just distal to the ostium of the vessel, so that the one or more ablative elements are arranged to create a circumferential lesion within the anatomical vessel, wherein the ablative structure tapers downward to a distal tip insertable into the ostium of the vessel; andan inner probe configured to be slidably disposed within the lumen of the outer probe, the inner probe including an elongate probe body having a distal diagnostic structure configured to extend out the exit port, and one or more diagnostic elements mounted to the distal diagnostic structure. 4. The probe assembly of claim 3, wherein the distal ablative structure is an open structure that forms an interior space, and the distal diagnostic structure is configured to extend within the interior space. 5. The probe assembly of claim 3, wherein the anatomical vessel is a pulmonary vein. 6. The probe assembly of claim 3, wherein the exit port is proximal to the distal ablative structure. 7. A medical probe assembly, comprising: an outer probe including an elongate probe body having a distal open helical structure, a lumen having an exit port, and one or more operative elements mounted to the helical structure, wherein the helical structure has at least two coils configured for contacting tissue within an anatomical vessel; andan inner probe configured to be slidably disposed within the lumen of the outer probe, the inner probe including an elongate probe body having a distal diagnostic structure configured to extend out the exit port, and one or more diagnostic elements mounted to the distal diagnostic structure, wherein the diagnostic structure comprises at l... | Boston Scientific Scimed, Inc. | A61B018/14 A61B005/042
US07436511 | B2 | 2008-10-14 | Analyte filter method and apparatus | 2005-07-22 | The invention comprises a method and apparatus for enhancing the analysis of noninvasive spectra, resulting in improved analytical performance. More particularly, the invention comprises a method and apparatus for processing noninvasive spectra with an analyte filter that preferably rejects variation likely to be detrimental to the measurement system, while passing signal that probabilistically is unique to the target analyte. Subsequently, the analyte filtered data are used to estimate an analyte property, such as a glucose concentration, in the presence of noise, interferences, state changes, and/or across analyzers. | 1. A method for determining a property of an analyte using a noninvasive spectrum, comprising the steps of: processing said noninvasive spectrum using an analyte filter yielding a processed spectrum, wherein said analyte filter correlates with frequency response shape of said analyte;generating said analyte filter using a set of spectra from a first set of spectrometers; andestimating said property using a multivariate model and said processed spectrum. 2. The method of claim 1, further comprising the step of: extracting a feature representative of said response shape of said analyte in frequency response. 3. The method of claim 2, wherein said analyte filter enhances said shape in frequency response below a break frequency. 4. The method of claim 3, wherein said analyte filter enhances said shape in frequency responses above a break frequency. 5. The method of claim 2, wherein said feature is characterized in terms of signal and interference. 6. The method of claim 1, wherein said multivariate model comprises any of: principal component regression; andpartial least squares. 7. The method of claim 1, wherein said frequency response shape of said analyte comprises any of: a net analyte signal;an accessible portion of a pure component analyte spectrum; andshape of said analyte after orthogonalization with respect to interference. 8. The method of claim 7, further comprising the step of: calculating said shape of said net analyte signal according to: nas=└I−BB+┘R wherein said net analyte signal, nas, is given by the projection of the pure component analyte signal, R, onto the null space of the interference contained in B, wherein B is a basis set that spans at least a portion of the space of expected variation and I is the identity matrix. 9. The method of claim 7, further comprising the step of: estimating said net analyte signal from a regression vector of said multivariate model. 10. The method of claim 1, further comprising the step of: theoretically generating said analyte filter using frequency response shape of said analyte. 11. The method of claim 10, wherein said step of theoretically generating further comprises the steps of: applying a-priori knowledge of signal of a system from which said noninvasive spectrum is collected; andapplying a-priori knowledge of interference of said system from which said noninvasive spectrum is collected. 12. The method of claim 11, wherein said step of theoretically generating further comprises the step of: generating a multi-stage filter. 13. The method of claim 12, wherein said multi-stage filter comprises: a two-stage filter. 14. The method of claim 13, wherein said two-stage filter comprises: a lowpass filter; anda highpass filter. 15. The method of claim 1, further comprising the step of: preprocessing said noninvasive spectrum prior to said step of processing, wherein preprocessing enhances at least a portion of said shape of said analyte. 16. The method of claim 10, further comprising t... | Sensys Medical, Inc. | G01J003/28 G01J003/42 G01J003/427 G01N033/48 A61B005/00
US07462149 | B2 | 2008-12-09 | Method and apparatus for remote blood alcohol monitoring | 2003-05-19 | A tamper proof transdermal alcohol content monitoring device is made up of analog and digital sides which are securely attached to the human subject to be monitored. The analog side has a sampling circuit which draws a measured insensible skin perspiration sample from the skin of the subject and measured with an electrochemical fuel cell. A distance measurement of the device from the skin of the subject and temperature of the sample are monitored along with the transdermal alcohol content, and converted to digital signals which are transmitted to a modem when the monitor is in proximity to the modem. The signals are stored in the modem and uploaded to a central monitoring station. Automatic alerts may be sent from the central monitoring station to a supervising agency. The supervising agency may also access the information through secured dedicated websites via the Internet. | 1. A system for continuous remote blood alcohol monitoring of a human subject comprising: a monitor device adapted to be attached to the human subject, wherein said monitor device takes a plurality of transdermal alcohol concentration readings at predetermined time intervals according to a first schedule stored in a first memory in said monitor device without active participation by the human subject, wherein said plurality of transdermal alcohol concentration readings are stored in said first memory;a modem in communication with said monitor device, wherein said plurality of transdermal alcohol concentration readings stored in said first memory are communicated from said monitor device to said modem at predetermined time intervals according to said first schedule through radio frequency signals without active participation by the human subject, wherein said plurality of transdermal alcohol concentration readings are stored in a second memory in said modem;a monitor network in communication with said modem, wherein said plurality of transdermal alcohol concentration readings stored in said second memory are communicated from said modem to said monitor network at predetermined time intervals according to a second schedule stored in said second memory through a first communication link without active participation by the human subject, and further wherein said monitor network analyzes said plurality of transdermal alcohol concentration readings to determine if an alert condition exists; anda supervising agency in communication with said monitor network, wherein said monitor network communicates an alert to said supervising agency through a second communication link when said alert condition exists. 2. The system according to claim 1 further comprising: a plurality of monitor devices each of which are adapted to be attached to one of a plurality of human subjects, wherein each of said plurality of monitor devices takes a plurality of transdermal alcohol concentration readings at predetermined time intervals according to a plurality of first schedules stored in a plurality of first memories in said plurality of monitor devices without active participation by said plurality of human subjects, wherein said plurality of transdermal alcohol concentration readings are stored in said plurality of first memories;a plurality of modems, wherein each of said plurality of modems is paired with and in communication with one of said plurality of monitor devices adapted to be attached to said one of a plurality of human subjects, wherein said plurality of transdermal alcohol concentration readings stored in said plurality of first memories are communicated from each of said plurality of monitor devices to its paired said plurality of modems at predetermined time intervals according to said plurality of first schedules through a plurality of radio frequency signals without active participation by said plurality of human subjects, wherein said plurality of transder... | Alcohol Monitoring Systems, Inc. | A61B005/00
US07463918 | B2 | 2008-12-09 | Method and apparatus for receiving data in human body communication system | 2003-12-31 | The present invention provides a method and apparatus for receiving data in a human body communication system. The receiving apparatus, which comprises plural receiving electrodes, selects the optimum pair of receiving electrodes when receiving data, so that it can improve quality of received information and obtain position information of a sensor in the human body. | 1. A method for receiving data in a human body communication system, comprising the steps of: selecting a pair of receiving electrodes from among plural receiving electrodes placed on the surface of the human body;processing a voltage value sensed by and induced at the selected pair of receiving electrodes and storing it in a memory, wherein the voltage is induced by a current flowing in the human body, the current being generated by an electric potential difference between spaced transmitting electrodes located inside the human body;repeating the selecting and processing steps for at least an additional pair of receiving electrodes;performing predetermined operations for values stored in the memory to select an optimum pair of receiving electrodes; andperforming image processing for a value corresponding to the optimum pair of receiving electrodes among the values stored in the memory. 2. The method of claim 1, wherein the step of processing the voltage comprises the sub-steps of: amplifying the voltage value;removing noise in the amplified signal; andconverting the noise-removed analog signal into a digital signal. 3. The method of claim 1, wherein the values stored in the memory are averages of voltage values sampled more than two times or averages of voltage waveforms for a given time. 4. The method of claim 1, wherein a pair of receiving electrodes corresponding to a maximum value among the values stored in the memory is selected as the optimum pair of receiving electrodes. 5. The method of claim 1, further comprising the step of: calculating position information of a sensor from the position of the optimum pair of receiving electrodes and storing it in the memory sequentially. 6. The method of claim 5, wherein a moving path, a speed and a direction of the sensor are extracted from the position information. 7. An apparatus for receiving data in a human body communication system using plural receiving electrodes installable on the surface of the human body, the apparatus comprising: a switching device for selecting a pair of receiving electrodes from among the plural receiving electrodes;a processor for processing a voltage value induced at the pair of receiving electrodes selected by the switching device, wherein the voltage is induced by a current flowing in the human body, and the current is generated by an electric potential difference between two transmitting electrodes located inside the human body;a memory for storing voltage values;a comparator for determining a maximum value among the values stored in the memory;an image processor for performing image processing of data corresponding to an optimum pair of receiving electrodes associated with the maximum value; anda control for controlling the switching device and the comparator. 8. The apparatus of claim 7, wherein the switching device includes a first switching circuit and a second switching circuit to which the plural receiving electrodes are connected respectively. 9. The ... | Korea Institute of Science and Technology | A61B005/05
US07455647 | B2 | 2008-11-25 | Dynamic spacer for total knee arthroplasty | 2003-07-24 | A dynamic spacer is provided for measuring flexion-extension gap during total knee arthroplasty. The dynamic spacer includes a first planar member having a lower tissue engaging surface, and a second planar member having an upper tissue engaging surface. A tensioner is disposed between the first planar member and the second planar member for applying a tensile force acting upon the first and second planar members. The tensioner is attached to the first and second planar members, such that the first planar member is held substantially parallel to the second planar member in the absence of compressive load. The dynamic spacer allows for accurately measuring flexion-extension gaps and angular deviation in flexion indicating the appropriateness of femoral rotation. | 1. A dynamic spacer for measuring flexion and extension gaps during knee arthroplasty, the dynamic spacer comprising: a first planar member having a lower tissue engaging surface and an upper tensioning surface, a second planar member having an upper tissue engaging surface and a lower tensioning surface, and a tensioning means residing entirely between said upper tissue engaging surface and said lower tissue engaging surface for applying a tensile force acting upon said first and said second planar members during the measuring of the flexion and extension gaps, said tensioning means fixedly attached to said upper tensioning surface of said first planar member and said lower tensioning surface of said second planar member, wherein said dynamic spacer is configured to reside completely internally to the knee during said knee arthroplasty. 2. The dynamic spacer of claim 1, further comprising a means for measuring a distance between said lower tissue engaging surface of said first planar member and said upper tissue engaging surface of said second planar member. 3. The dynamic spacer of claim 2, further comprising a means for measuring an angulation of said second planar member as it deviates from parallel with respect to said first planar member. 4. The dynamic spacer of claim 1, wherein said tensioning means comprises a plurality of compressive coil-type springs. 5. The dynamic spacer of claim 1, wherein said tensioning means comprises a plurality of segmental cylindrical spring housings each having an upper section and a lower section, said lower section fixedly attached to said upper tensioning surface of said first planar member and said upper section fixedly attached to said lower tensioning surface of said second planar member, each of said plurality of segmental cylindrical spring housings encapsulating a compressive coil-type spring, said upper section of said segmental cylindrical spring housing having a first diameter and said lower section of said segmental cylindrical spring housing having a second diameter smaller than said first diameter of said upper section, and said lower section of said segmental cylindrical spring housing capable of being slidingly engaged within said upper section of said segmental cylindrical spring housing. 6. The dynamic spacer of claim 5, wherein said lower section of said segmental cylindrical housing comprises graduated indicia for measuring a distance between said lower tissue engaging surface of said first planar member and said upper tissue engaging surface of said second planar member. 7. The dynamic spacer of claim 6, further comprising a means for measuring an angulation of said second planar member as it deviates from parallel with respect to said first planar member. 8. The dynamic spacer of claim 7, wherein said means for measuring the angulation of said second planar member as it deviates from parallel with respect to said first planar member comprises a positioner, said positioner fixedly... | - | A61B005/103 A61B005/117
US07463916 | B2 | 2008-12-09 | Optical measurement apparatus for living body | 2001-10-16 | An optical measurement apparatus for a living body having light transmitting elements for transmitting light to a plurality of positions in areas to be examined within a living body detects light which has been transmitted by the light transmitting elements after passing through the living body at a plurality of positions in the examination area; applies to the living body a load to which the living body responds; calculates signals representing an intensity change of the transillumination detected by the detecting the light at the measuring points based on the positional relation between each light transmitting position and light detecting position when a load is applied to the living body and when a load is not applied; calculates time variations of one signal among two signals calculated at the measuring points based on said positional relation; and displays graphs of the above-calculated signals in correspondence with said positional relation. | 1. An optical measurement apparatus for examination of a living body comprising: light transmitting means for transmitting light to a plurality of positions in an area being examined within a living body;light detecting means for detecting light which has been transmitted by said light transmitting means and has passed through the living body at a plurality of positions in the area being examined;means for calculating signals representing the intensity change of transillumination detected by said light detecting means at measuring points that are determined based on a positional relation between said light transmitting means and said light detecting means, during a period in which a load is applied to the living body and periods in which the load is not applied;means for calculating time variations of at least one signal among two signals calculated at the measuring points that are determined based on the positional relation between each of said light transmitting means and said light detecting means;means for converting the calculated time variations of the light signal at each measuring point into time variations of hemoglobin concentration; anddisplay control means for displaying graphs of the calculated time variations of hemoglobin concentration in correspondence with the positional relation between each of said light transmitting means and said light detecting means; andmeans for selecting and classifying graphs displayed by said display control means on display means so as to display plural groups of graphs. 2. An optical measurement apparatus for examination of a living body according to claim 1, further comprising means for displaying each group of graphs that are classified by said selecting means so as to be distinguishable from other groups of graphs. 3. An optical measurement apparatus for examination of a living body according to claim 1, comprising means for displaying only one particular group of graphs among the plural of groups of graphs selected and classified by said selecting means. 4. An optical measurement apparatus for examination of a living body according to claim 1, comprising means for averaging measured values of signals of the respective groups of graphs classified by said selecting means and means for displaying to said display means the average data of the groups of graphs averaged by said averaging means. 5. An optical measurement apparatus for examination of a living body according to claim 4, wherein said averaging means includes means for converting the calculated average values into data of graphs to be displayed to said display means. 6. An optical measurement apparatus for examination of a living body according to claim 5, comprising means for calculating a correlation between standard data representing characteristics of one selected group of graphs and data calculated by said averaging means, and means for displaying values representing the calculated correlation in the vicinity of the displayed graph... | Hitachi Medical Corporation | A61B005/1455
US07440793 | B2 | 2008-10-21 | Apparatus and method for removing abnormal tissue | 2004-07-22 | A computer assisted, minimally invasive method and apparatus for surgically removing abnormal tissue from a patient, for example, from a breast, are disclosed. The method involves imaging of the breast to locate the abnormal tissue, and determining a volume encapsulating the abnormal tissue and including a margin of healthy tissue. Based on the volume, a sequence of movements of a surgical instrument for tissue removal device is planned, so as to predictably excise the desired volume of tissue. | 1. A computer-assisted method of removing abnormal tissue from a region of interest by way of a reduced number of incisions, said method comprising: generating a three-dimensional (3-D) map of said region of interest comprising abnormal tissue;determining a tissue volume, based on said 3-D map, said tissue volume including the abnormal tissue and an adjacent margin of healthy tissue;determining, based on said tissue volume, at a computing device a surgical plan including at least one incision point and a series of associated excision points associated with said at least one incision point; andcontrolling movement of a surgical instrument used to excise tissue, to sequentially move to each of said at least one incision point and to each of its associated excision points and excise tissue, resulting in removal of said tissue volume;wherein at least one of said at least one incision point is associated with two or more excision points. 2. The method of claim 1, wherein said region of interest forms part of a breast. 3. The method of claim 2, further comprising restraining said breast. 4. The method of claim 1, wherein said determining a tissue volume comprises determining a volume having a regular geometric shape. 5. The method of claim 4, wherein said regular geometric shape comprises a cylinder. 6. The method of claim 5, wherein said incision point is along a central axis of said cylinder. 7. The method of claim 6, wherein said series of excision points comprises a plurality of axial positions along said central axis. 8. The method of claim 7, wherein said series of excision points comprise a plurality of angular positions associated with each of said axial positions. 9. The method of claim 8, wherein a number of angular positions associated with each of said axial positions depends on a diameter of said cylinder. 10. The method of claim 7, wherein at least some of said series of excision points are at the same axial position along said central axis, and the tissue is assumed to collapse into a void created by said excisions. 11. The method of claim 6, wherein said surgical instrument comprises a vacuum assisted biopsy device. 12. The method of claim 5, wherein said incision point is along an axis parallel to a central axis of said cylinder, proximate an edge of said cylinder. 13. The method of claim 5, further comprising, repeating said generating, and said determining a tissue volume. 14. The method of claim 4, wherein said regular geometric shape comprises a rectangular prism. 15. The method of claim 14, wherein said incision point is along an axis parallel to a central axis of said rectangular prism. 16. The method of claim 1, wherein said generating comprises generating said 3-D map from a plurality of 2-D images of said region of interest. 17. The method of claim 1, wherein said determining a tissue volume comprising using edge detection of said 3-D map to locate boundaries of the abnormal tissue. 18. The method of claim 17,... | - | A61B005/05 A61B008/00 A61N001/30
US07785314 | B2 | 2010-08-31 | Systems and methods for needle access to an intervertebral disc | 2007-06-15 | A catheter with a balloon anchor may be sized to fit through an inner needle and introduced into the patient through the needle. The inner needle may comprise an atraumatic needle with a side port and/or a blunt tip. The catheter may include a radiopaque coil disposed over the distal portion of the catheter body to assist with catheter placement when viewed with fluoroscopy and/or x-ray. The distal portion of the catheter can be very flexible. A balloon inflation tube may terminate proximal to a proximal end of the expandable balloon anchor such that the distal portion of catheter can be flexible. This termination of the inflation tube may also position the inflation tube away from the annulus and nerve roots, so as to avoid irritation of the annulus and rubbing of the nerve roots that may potentially obscure FAD test results. | 1. A catheter for accessing an intervertebral disc, said catheter comprising: an elongate flexible catheter body having a proximal portion, a distal portion, and at least one lumen for introducing substances;an inflatable anchoring balloon on the distal portion of the catheter body;a radiopaque coil comprising a plurality of wound spaced apart loops disposed along the distal portion of the elongated flexible catheter body through the inflatable anchoring balloon from a proximal portion of the inflatable anchoring balloon to a distal portion of the inflatable anchoring balloon and to a distal most point of a distal tip of the elongated flexible catheter body, the spaced apart loops allowing the distal portion to maintain flexibility, the radiopaque coil being configured to permit the catheter to be removed in situ as one-piece if a portion of the catheter breaks where the radiopaque coil is disposed along the elongated flexible catheter body such that the radiopaque coil retains the broken portion of the catheter. 2. A catheter as in claim 1, further comprising: an injection tube extending from the proximal portion to or near the distal tip of the elongate flexible catheter body, andan inflation tube extending from the proximal portion to or near the distal portion of the elongate flexible catheter body;wherein the radiopaque coil is disposed over a distal portion of the injection tube that passes through the inflatable anchoring balloon. 3. A catheter as in claim 2, wherein the radiopaque coil extends from a distal most point of a distal tip of the injection tube to the proximal portion of the elongate flexible catheter body. 4. A catheter as in claim 1, wherein the elongate flexible catheter body comprises a polymeric tube having a central passage which receives the inflation tube and the injection tube. 5. A catheter as in claim 4, wherein the inflation tube comprises a metal tube. 6. A catheter as in claim 4, wherein the injection tube comprises a metal tube. 7. A catheter as in claim 4, wherein the elongate flexible catheter body has an outside diameter capable of passing through a lumen of a 0.9 mm needle. 8. A catheter as in claim 4, wherein at least one of the coil, the injection tube or the inflation tube comprises comprise tungsten, rhenium, molybdenum, tantalum, palladium, cobalt-chromium, tungsten-rhenium, tungsten-carbide or molybdenum-rhenium. 9. A catheter for accessing an intervertebral disc, said catheter comprising: an elongate flexible catheter body having a proximal portion, a distal portion, and at least one lumen for introducing substances;an anchoring balloon comprising an inflatable portion disposed on the distal portion of the catheter body;an inflation tube in the elongate body, said inflation tube having a distal end which terminates proximal to the inflatable portion of the balloon;a radiopaque coil comprising a plurality of wound spaced apart loops disposed along the distal portion of the elongated flexible cat... | Kyphon SÀRL | A61M031/00
US08306599 | B2 | 2012-11-06 | Process and device for monitoring the status of the body fluids of a person | 2008-02-25 | A process and a device are provided for monitoring the status of the body fluids of a person. The device includes a piece of clothing with ventilating ducts, a first temperature sensor (6) and a first moisture sensor (11) for measuring the inlet flow (7) into the piece of clothing and a second temperature sensor (17) and a second moisture sensor (12) for measuring the outlet flow (10) from the piece of clothing. A pressure sensor (13) is provided sensing pressure in the inlet flow. A fan (5) or a compressed-air source is provided for transporting the cooling air into the piece of clothing. A measuring and analyzing unit (14) is connected to the sensors and the fan (5) or the compressed-air source. | 1. A process for monitoring the status of body fluids of a person, the process comprising the steps of: providing a person with a piece of clothing with ventilating ducts;providing a first temperature sensor and a first moisture sensor positioned for measuring inlet flow into the piece of clothing with the first temperature sensor and with the first moisture sensor;providing a second temperature sensor and a second moisture sensor positioned for measuring outlet flow from the piece of clothing with the second temperature sensor and with the second moisture sensor;providing a pressure sensor in the inlet flow;providing a fan or a compressed-air source providing cooling air into the piece of clothing and to provide the inlet flow and the outlet flow;providing a measuring and analyzing unit connected to said first temperature sensor, said first moisture sensor, said second temperature sensor, said second moisture sensor, said pressure sensor and the fan or the compressed-air source;activating the fan or the compressed-air source;measuring pressure established at the air inlet in the inlet flow with the pressure sensor in the inlet flow;determining gas volume flow based on said measured pressure;measuring an inlet flow moisture content of the gas volume flow in the inlet flow with said first moisture sensor;measuring an outlet flow moisture content of the gas volume flow in the outlet flow with said second moisture sensor;determining moisture currently being fed and moisture currently being removed from the gas volume flow based on at least said gas volume flow, said inlet flow measured moisture content and said outlet flow moisture content;determining a loss of moisture by the person from a determined time integral of a difference between the moisture fed in and the moisture removed or from the difference between time integrals of the moisture fed in and the moisture removed, wherein said loss of moisture is compared to a limit value and an alarm signal is sent by the measuring and analyzing unit when the limit value is exceeded. 2. A process in accordance with claim 1, wherein the moisture content is determined either by direct absolute measurements or calculated via a relative moisture measurement with complementary temperature measurement in the inlet flow with said first temperature sensor and in the outlet flow with said second temperature sensor. 3. A process in accordance with claim 1, wherein the gas volume flow is determined on the basis of the current speed of the fan from a characteristic speed-dependent volume flow curve 4. A process in accordance with claim 1, wherein determining a loss of moisture further includes determining moisture gained or released by breathing including measuring a respiratory volume of the person to determine moisture fed in and an absolute moisture content in the breathing air breathed in. 5. A process in accordance with claim 1, wherein said first temperature sensor, said first moisture sensor, said second... | Dräger Safety AG & Co. KGaA | A61B005/00
US08652051 | B2 | 2014-02-18 | Contact pressure sensing apparatus for use with exercise equipment sensors | 2007-04-04 | Contact pressure sensing apparatus for use with exercise equipment sensors are described. An example apparatus includes a sensor to detect a physiological condition of a user of an exercise device through physical contact with the user and a pressure sensor to detect a contact pressure applied by the user to the sensor. | 1. An apparatus for use with an exercise device, comprising: a physiological condition sensor to detect a physiological condition of a user of the exercise device through physical contact with the user;a pressure sensor to detect a contact pressure applied by the user to the physiological condition sensor; anda processor configured to compare the detected contact pressure with a threshold amount of pressure associated with an ability of the physiological condition sensor to detect the physiological condition, wherein the processor is configured to calculate a percentage representative of an accuracy of a reading of the physiological condition by the physiological condition sensor based on the comparison between the detected contact pressure and the threshold amount of pressure. 2. An apparatus as defined in claim 1, wherein the physiological condition is a heart rate. 3. An apparatus as defined in claim 1, wherein the physiological condition sensor to detect the physiological condition comprises an electrode. 4. An apparatus as defined in claim 1, wherein the pressure sensor is located between the physiological condition sensor and a frame of the exercise device. 5. An apparatus as defined in claim 1, wherein the processor is configured to cause the display to present the physiological condition. 6. An apparatus as defined in claim 1, wherein the processor is configured to cause a change in an operation of the exercise machine based on the physiological condition. 7. An apparatus as defined in claim 1, wherein the processor is configured to cause the display to present information associated with the contact pressure. 8. An apparatus as defined in claim 7, wherein the displayed information associated with the contact pressure comprises a difference between the detected contact pressure and the threshold amount of pressure. 9. An apparatus as defined in claim 1, wherein the processor is configured to cause presentation of an indication of how close the user is to attaining the threshold amount of pressure. 10. An apparatus as defined in claim 1, wherein the threshold amount of pressure is a non-zero value. 11. An apparatus as defined in claim 1, wherein the processor is configured to use the comparison to determine whether a sufficient amount of pressure is being applied by the user for the physiological condition sensor to accurately detect the physiological condition. 12. An apparatus as defined in claim 1, wherein the processor is configured to withhold a result of a reading of the physiological condition taken by the physiological condition sensor from the display based on the comparison of the detected contact pressure to the threshold amount of pressure. 13. An apparatus as defined in claim 1, wherein the processor is to cause a display to present an instruction to increase or decrease the contact pressure based on the comparison of the detected contact pressure and the threshold amount of pressure. 14. An apparatus as defined i... | Brunswick Corporation | A61B005/02
US08051720 | B2 | 2011-11-08 | Method and device for measuring the local mechanical resistance of a porous body | 2006-10-31 | Method and device for determining the local mechanical resistance inside of a porous body having a variable density and/or porosity, in particular inside of a porous bone structure comprising the following steps: selecting a site of said porous body destined to receive a fixation element, in particular a screw or a bone implant;drilling a hole into said porous body; andmeasuring a mechanical parameter of the porous structure surrounding said drilled hole by inserting a suitable tool (1) into said drill hole. | 1. A method for determining the local mechanical resistance inside of a porous body having a variable density and/or porosity, comprising the following steps: A) selecting a site of said porous body destined to receive a fixation element, in particular a screw or a bone implant;B) drilling a hole into said porous body;C) measuring a mechanical parameter of the porous structure surrounding said drilled hole by inserting a tool having a blade into said drill hole, wherein the mechanical parameter is measured by pushing the blade into the bottom of the drilled hole to a certain distance and measuring a twisting moment needed to turn the blade once around. 2. The method according to claim 1, wherein said measuring is performed dynamically. 3. The method according to claim 1, wherein said mechanical parameter is the mechanical resistance of the porous structure due to its deformation. 4. The method according to claim 1, wherein said mechanical parameter is the mechanical resistance of the porous structure caused by the torsional force applied through rotating the tool. 5. The method according to claim 4, wherein said torsional force is measured by obtaining free axial rotation of the tool in said porous structure. 6. The method according to claim 1, wherein the twisting angle of the tool is measured which is obtained by applying a pre-selectable maximum torsional force. 7. The method according to claim 1, comprising the further steps of: D) comparing said measured properties with standardized values of said properties obtained with porous structures of different density of porosity;E) displaying the result of the comparison of said measured versus said standardized properties. 8. The method according to claim 1, wherein said measuring is effected by inserting said tool into said hole and impacting it into said surrounding porous structure. 9. The method according to claim 1, wherein several distinctive measurements are taken to get averaged results with standard deviation for higher accuracy. 10. The method according to claim 1, wherein several sites are selected for obtaining several values of the local mechanical resistance at this various sites. 11. The method according to claim 1, wherein the mechanical resistance is measured in the wall of said drilled hole at one or more distinctive locations. 12. The method according to claim 1, wherein a certain number of stepwise measurements are taken in one or in a number n>1 drill holes. 13. The method according to claim 12, wherein said stepwise measurements are taken by alternately predrilling to the depth planned for the next measurement and taking said measurement. 14. The according to claim 12, wherein said stepwise measurements are guided by a wire set coaxial to the central axis of the planned drill hole before the first predrilling, deep enough to stay in place during the whole measurements or even until the fixation element is inserted. 15. The method according to claim 12, wherein s... | AO Technology AG | G01N003/00
US07863896 | B2 | 2011-01-04 | Systems and methods for calibrating functional magnetic resonance imaging of living tissue | 2008-07-09 | Embodiments of the present invention relate to systems and methods for calibrating functional magnetic resonance imaging of living tissue. For example, a calibration method includes providing a contrast signal simulation device configured for generating one or more pre-determined contrast simulation signals selected from a group consisting of BOLD and BOSS contrast simulation signals that simulate at least one contrast signal generated by living tissue of a subject and positioning the contrast signal simulation device adjacent to the subject. The method includes using an imaging instrument to take first and second data sets from the subject's tissue while the contrast signal simulation device is disposed adjacent to the subject and while generating and not generating the one or more pre-determined contrast simulation signals, respectively. The method also includes using the first and second data sets to determine at least one figure of merit of the imaging instrument, the contrast simulation device, and/or the living tissue. | 1. A calibration method comprising the steps of: providing a contrast signal simulation device that is configured for generating one or more pre-determined contrast simulation signals selected from a group consisting of BOLD and BOSS contrast simulation signals, said one or more pre-determined contrast simulation signals simulating at least one contrast signal generated by living tissue of a subject;positioning said contrast signal simulation device adjacent to said subject;using an imaging instrument to take a first data set from said subject's tissue while: (A) said contrast signal simulation device is disposed adjacent to said subject, and (B) said contrast signal simulation device is generating said one or more pre-determined contrast simulation signals;using said imaging instrument to take a second data set from said subject's tissue while said contrast signal simulation device is not generating said one or more pre-determined contrast simulation signals; andusing said first and second data sets to determine at least one figure of merit of at least one of said imaging instrument, said contrast simulation device, or said living tissue. 2. The method of claim 1, wherein said figure of merit of said imaging instrument takes into account in vivo noise associated with said subject. 3. The method of claim 1, wherein said step of using said first and second data sets to determine at least one figure of merit comprises determining statistically reliable differences between said first and second data sets. 4. The method of claim 1, wherein said subject is a human being or a living animal. 5. The method of claim 1, wherein said imaging instrument is a magnetic resonance imaging device. 6. The method of claim 1, further comprising: using said imaging instrument to acquire a third data set of said subject's tissue while said subject is being subjected to a particular task or event;using said imaging instrument to take a fourth data set of said subject's tissue while said subject is not being subjected to said particular task or event;using said third and fourth data sets to determine statistically significant differences between said third and fourth data sets; andusing said at least one figure of merit to assess an accuracy of said statistically significant differences between said third and fourth data sets. 7. The method of claim 1, wherein said figure of merit is selected from a group consisting of a sensitivity, a specificity, and a receiver operating characteristic. 8. The method of claim 1, wherein each of said first and second data sets is comprised spatially of at least one or more voxels. 9. The method of claim 1, wherein said first and second data sets comprise at least one of instrument noise, physiological noise, or data regarding said one or more pre-determined contrast simulation signals. 10. The method of claim 1, wherein using said first and second data sets to determine said at least one figure of merit of said living tissue c... | University of Florida Research Foundation, Inc. | G01V003/00
US08167813 | B2 | 2012-05-01 | Systems and methods for locating a blood vessel | 2007-05-17 | Systems and methods for locating a blood vessel are described. One described method includes receiving a first temperature measure associated with a first point on a surface of a skin; receiving a second temperature measure associated with a second point on the surface of the skin, and determining a differential between the first temperature measure and the second temperature measure. The method further includes generating an output signal indicating the presence of a blood vessel if the differential exceeds a threshold. | 1. A handheld system for locating a blood vessel, the system comprising: a housing configured to be moved across a patient's skin;a temperature sensor disposed within the housing and operable to output a temperature measure of a location on a surface of the patient's skin; anda processor in communication with the temperature sensor and configured to: receive a first temperature measure at a first location from the temperature sensor,receive a second temperature measure at a second location from the temperature sensor, the second location different from the first location,determine a difference between the first temperature measure and the second temperature measure, andgenerate an output signal indicating the presence of a blood vessel if the difference exceeds a threshold,wherein the temperature sensor is configured to output the first and second temperature measures at the first and second locations based on a movement of the housing across the patient's skin. 2. The system of claim 1, further comprising an indicator configured to receive the output signal and indicate the presence of the blood vessel. 3. The system of claim 1, wherein the processor is further configured to determine a direction of the blood vessel. 4. The system of claim 3, wherein the indicator is further configured to indicate the direction of the blood vessel. 5. The system of claim 2, wherein the indicator comprises a visual indicator. 6. The system of claim 2, wherein the indicator comprises an audible indicator. 7. The system of claim 2, wherein the indicator comprises a marker configured to generate a mark on the surface of the skin over the blood vessel. 8. The system of claim 1, further comprising an actuator configured to provide a haptic effect in response to the output signal. 9. The system of claim 8, wherein the haptic effect comprises a vibrotactile haptic effect. 10. A handheld system for locating a blood vessel, the system comprising: a housing configured to be moved across a patient's skin;a first temperature sensor disposed within the housing and operable to output a temperature measure of a first location on a surface of a skin;a second temperature sensor disposed within the housing and operable to output a second temperature measure of a second location on the surface of the skin; anda processor in communication with the first temperature sensor and the second temperature sensor, the processor configured to: receive the first temperature measure from the first temperature sensor,receive the second temperature measure from the second temperature sensor,determine a difference between the first temperature measure and the second temperature measure, andgenerate a first output signal indicating the presence of a blood vessel if the difference exceeds a threshold. 11. The system of claim 10, further comprising an indicator configured to receive the first output signal and indicate the presence of the blood vessel. 12. The system of claim 11, further ... | Immersion Medical, Inc. | A61B005/00
US08208701 | B2 | 2012-06-26 | Method for high-resolution presentation of filigree vessel implants in angiographic images | 2008-04-24 | The invention relates to a method for high-resolution display of vessel implants in angiographic images, featuring the steps: recording at least two images of an object-catheter combination including catheter markers with a medical imaging method; detection of an region of interest created in the form of an area between the balloon markers which completely contains the object to be registered for each recorded image; coarse registration of the region of interest images by registration of the respective pairs of balloon markers of all recorded images; fine registration of the region of interest image content/of the region of interest image by registration of the region of interest content; and arithmetic averaging across the fine-registered images. | 1. A method for creating a high-resolution image of an object in an angiographic image of a patient, comprising: recording at least two images of the object and a catheter combination including catheter markers;detecting a region of interest in an area between the markers completely containing the object in the at least two images;marker-based coarsely registering the region of interest by registering the markers in the at least two images;non-marker-based finely registering the region of interest by registering a content of the region of interest in the at least two images after the marker-based coarse registration to increase a surface of the object on a total surface of the region of interest for generating a fine-registered image; andcreating the high-resolution image of the object by arithmetically averaging across the fine-registered image,wherein the fine registration is undertaken by a method selected from the group consisting of: sum of difference squares, variance of difference, leading-sign-change based measuring, coincident bit counting, and difference-histogram-based measuring. 2. The method as claimed in claim 1, wherein the catheter is a balloon catheter. 3. The method as claimed in claim 1, wherein the area between the markers is a rectangle. 4. The method as claimed in claim 1, wherein the coarse registration is undertaken by moving, turning, or scaling connecting lines of the markers. 5. The method as claimed in claim 1, wherein the at least two images are ECG-triggered recorded. 6. The method as claimed in claim 1, wherein the at least two images are recorded by a magnetic resonance tomography device, an x-ray imaging device, or a projection radiography device. 7. The method as claimed in claim 1, wherein the object comprises a vessel implant of the patient. 8. The method as claimed in claim 7, wherein the vessel implant comprises a braided stent. 9. A device for creating a high-resolution image of an object in an angiographic image of a patient, comprising: an image device that records at least two images of the object and a catheter combination including catheter markers; anda computing device connected to the image device that: detects a region of interest in an area between the markers completely containing the object in the at least two images;marker-based coarsely registers the region of interest by registering the markers in the at least two images;non-marker-based finely registers the region of interest by registering a content of the region of interest in the at least two images after the marker-based coarse registration to increase a surface of the object on a total surface of the region of interest for generating a fine-registered image; andcreates the high-resolution image of the object by arithmetically averaging across the fine-registered image,wherein the fine registration is undertaken by a method selected from the group consisting of: sum of difference squares, variance of difference, leading-sign-chang... | Siemens Aktiengesellschaft | G06K009/00
US08165652 | B2 | 2012-04-24 | Physiological signal detection module, multi-channel connector module and physiological signal detection apparatus using the same | 2008-04-04 | Provided is a physiological signal detection module which has a heating member and a multi-channel connector. The heating member maintains an electrode's contact with a user skin at a constant temperature, and the multi-channel connector module for electrical connection transmits the physiological signal to an external physiological signal detection apparatus. It is possible to prevent the skin from being directly contacted with the electrode in the cold state and have a less impact given by a change of the impedance due to a skin temperature based on the change of the season and external temperature, so that the stable physiological signal can be detected without distortion. | 1. A physiological signal detection apparatus comprising: a plurality of physiological signal detection modules for detecting various physiological signals, each physiological signal detection module using an electrode adapted for contact with a user's the skin; anda multi-channel module for electrical connection to transmit the various physiological signals detected from the electrode to an external physiological signal detection apparatus,wherein each of the physiological signal detection modules comprises:the electrode;an electrode housing detachably connected to the electrode and having a printed circuit board (PCB) mounted with various circuits for matching impedance when the electrode is in use;a heating member interposed between the electrode and the electrode housing, and being supplied with an external power to increase the temperature of the electrode to a certain degree; andan insulator disposed between the heating member and the electrode,wherein the multi-channel connector module comprises:a connector housing that forms a whole body;a plurality of connectors for electrical connection to transmit various physiological signals detected by the plurality of physiological signal detection modules, wherein a portion of each of the connectors is exposed on one side of the connector housing; anda power supply unit for supplying power to the plurality of physiological signal detection modules, wherein the power supply unit is arranged inside the connector housing and electrically connected to the plurality of connectors to detect the various physiological signals. 2. The physiological signal detection apparatus according to claim 1, wherein the power supply unit comprises at least one rechargeable battery. 3. The physiological signal detection apparatus according to claim 2, further comprising a state display unit for monitoring a charging state of the battery. 4. The physiological signal detection apparatus according to claim 2, further comprising a power supply connector for connecting an external power to the battery to supplement the battery with the external power when the battery is low on power. | Electronics and Telecommunications Research Institute | A61B005/04
US07998110 | B2 | 2011-08-16 | Medical device for delivering drug and/or performing physical therapy | 2007-04-25 | Current methods of delivering drug usually require a certain degree of human intervention. For example, a doctor, a paramedic or a patient is required to inject insulin manually when the blood sugar level becomes low. This invention provides a medical device for use with a patient for performing at least one of the functions of delivering drug and performing physical therapy to the patient. The device of this invention includes a sensor for measuring at least one parameter from the patient, an analyzer for analyzing the parameter to determine appropriate function to be carried out on the patient by said medical device, and at least one actuator to perform at least one of the functions of delivering drug and performing physical therapy. | 1. A medical device for use with a patient for performing at least one of the functions of delivering a drug and applying a physical therapy to the patient, the device including: at least one sensor for measuring at least one parameter of the patient, wherein the at least one parameter is indicative of a pain and an inflammation being experienced by the patient;an analyzer for analyzing the parameter to determine appropriate function to be carried out on the patient by said medical device; andat least one actuator to perform at least one of the functions of delivering the drug and applying the physical therapy in response to the analyzed parameter. 2. The medical device of claim 1 further including a drug storage unit for storing at least one drug, said drug storage unit being operatively connected to the actuator for delivering said drug to the patient. 3. The medical device of claim 1, wherein the parameter is selected from a group including electrical current or resistance, temperature, pressure, tension, humidity, flow circulation in blood vessels, pH, intensity of at least one predetermined wavelength of electromagnetic wave, sonic signal, concentration of at least one chemical compound, and their combinations thereof. 4. The medical device of claim 3, wherein said chemical compound is selected from a group including glucose, sodium chloride, dimethyl sulfoxide, azone, isopropyl myristate, propylene glycol, Carbomer resins, gelling agents and their combinations thereof. 5. The medical device of claim 1, wherein the physical therapy is selected from a group including mechanical stimulus, temperature stimulus, electrical stimulus, ultrasonic stimulus, magnetic stimulus, and irradiation of at least one predetermined wavelength. 6. The medical device of claim 1 including an attachment means for attaching the medical device to the patient. 7. A method for use with a patient for performing at least one of the functions of delivering a drug and applying a physical therapy to the patient using a medical device, said method including the steps of: measuring at least one parameter of the patient, wherein the at least one parameter is indicative of a pain and an inflammation being experienced by the patient;analyzing the parameter to determine an appropriate function to be carried out on the patient; andperforming at least one of the functions of delivering the drug and applying the physical therapy in response to the analyzed parameter. 8. The method of claim 7 wherein the medical device includes a drug storage unit for storing at least one drug, said drug storage unit being operatively connected to the actuator for delivering said drug to the patient. 9. The method of claim 7, wherein the parameter is selected from a group including electrical current or resistance, temperature, pressure, tension, humidity, flow circulation in blood vessels, pH, intensity of at least one predetermined wavelength of electromagnetic wave, sonic signal, concentr... | Hong Kong Polytechnic University | A61M031/00
US07993279 | B2 | 2011-08-09 | Methods and systems for implantably monitoring external breathing therapy | 2008-07-15 | An implantable device is used to monitor one or more conditions associated with an external breathing therapy delivered to the patient. The device may monitor therapy parameters including therapy effectiveness, impact of the therapy on the patient, therapy usage, compliance with a prescribed usage, therapy interactions, and/or other parameters. | 1. A method of monitoring respiration therapy delivered to a patient, comprising: sensing one or more conditions associated with patient-external breathing therapy;analyzing a sensed respiratory signal; andimplantably monitoring the patient-external breathing therapy based on the one or more sensed conditions, wherein implantably monitoring comprises monitoring the patient's use of the patient-external breathing therapy based on the respiratory signal analysis, wherein monitoring the patient's use comprises comparing a respiratory signal sensed during patient use of the patient-external breathing therapy with a sensed respiratory signal during non-use of the patient-external breathing therapy to detect signal features that indicate usage of the patient-external breathing therapy. 2. The method of claim 1, wherein sensing the one or more conditions comprises sensing one or more conditions associated with positive airway pressure breathing therapy. 3. The method of claim 1, wherein sensing the one or more conditions comprises sensing transthoracic impedance of the patient. 4. The method of claim 1, wherein analyzing the sensed respiratory signal comprises analyzing a sensed waveform morphology of the sensed respiratory signal. 5. The method of claim 1, wherein analyzing the sensed respiratory signal comprises detecting a pressure notch indicative of flow controlled breathing therapy usage in a waveform of the respiratory signal. 6. The method of claim 1, wherein implantably monitoring the patient-external breathing therapy comprises implantably monitoring an effectiveness of the patient-external breathing therapy. 7. The method of claim 6, wherein implantably monitoring the effectiveness of the patient-external breathing therapy comprises: detecting disordered breathing episodes; andmonitoring the effectiveness of the patient-external breathing therapy based on the detected disordered breathing episodes and the patient's use of the patient-external breathing therapy. 8. The method of claim 1, wherein implantably monitoring the patient's use of the patient-external breathing therapy comprises comparing the sensed respiratory signal with a respiratory signal indicating usage of the patient-external breathing therapy. 9. The method of claim 1, wherein implantably monitoring the patient use of the patient-external breathing therapy comprises comparing the patient use of patient-external breathing therapy with a prescribed use of the patient-external breathing therapy. 10. The method of claim 1, further comprising providing one or more alert signals associated with the patient use of the patient-external breathing therapy. 11. The method of claim 1, wherein implantably monitoring the patient-external breathing therapy comprises implantably monitoring an impact of the patient-external breathing therapy on the patient. 12. The method of claim 1, wherein implantably monitoring the patient-external breathing therapy comprises implantably monitori... | Cardiac Pacemakers, Inc. | A61B005/08
US09754078 | B2 | 2017-09-05 | Haptic health feedback monitoring | 2007-06-21 | A haptic health feedback monitor includes a health parameter monitor that detects a health parameter. A haptic feedback generator receives the health parameter and compares it to a predetermined level. If the health parameter reaches or exceeds the level, a type of haptic feedback to generate is determined. The type of feedback may depend on which predetermined level is reached or exceeded. The haptic feedback generator then generates the determined type of haptic feedback. | 1. A method of monitoring health of a person comprising: detecting a first health parameter and a second health parameter of the person with a health parameter monitor;comparing the first health parameter to a first predetermined level;determining a first type of haptic feedback to generate based on the comparing the first health parameter to the first predetermined level, the first type of haptic feedback being vibrotactile haptic feedback,the vibrotactile haptic feedback having a plurality of parameters,the plurality of parameters including a magnitude parameter, a frequency parameter, and a duration parameter, andat least one of the plurality of parameters varying based on the comparing the first health parameter to the first predetermined level;generating a first haptic effect, based on the first type of haptic feedback, for a user with a first haptic feedback generator, the user being different from the person with the health parameter monitor;comparing the second health parameter to a second predetermined level;determining a second type of haptic feedback to generate based on the comparing the second health parameter to the second predetermined level;generating a second haptic effect, based on the second type of haptic feedback, with the first haptic feedback generator; andguiding the user to the person after the generating the first haptic effect by providing a series of haptic effects to assist in pinpointing a location of the person. 2. The method of claim 1, wherein the first health parameter is at least one of heart rate, blood pressure, blood glucose level, respiratory level, temperature, blood alcohol level, blood iron level, blood salinity, blood electrolyte level or hormone level. 3. The method of claim 1, further comprising: generating the first haptic effect for the person with a second haptic feedback generator. 4. The method of claim 1, further comprising: transmitting a signal that indicates the location of the person with the health parameter monitor. 5. A haptic health feedback monitor comprising: a health parameter monitor that detects a first health parameter and a second health parameter of a person; anda haptic feedback generator coupled to the health parameter monitor,the haptic feedback generator including a processor that performs the operations of receiving the first health parameter,performing a comparison of the first health parameter to a first predetermined level,determining a first type of haptic feedback based on the comparison of the first health parameter to the first predetermined level,generating a first haptic effect, based on the first type of haptic feedback, to a user different from the person,receiving the second health parameter,performing a comparison of the second health parameter to a second predetermined level,determining a second type of haptic feedback based on the comparison of the second health parameter to the second predetermined level,generating a second haptic effect, based on the seco... | Immersion Corporation | A61B005/00 A61H001/00 G06F019/00 G06F003/01 A61B005/024
US09468412 | B2 | 2016-10-18 | System and method for accuracy verification for image based surgical navigation | 2007-06-22 | Certain embodiments of the present invention provide for a system and method for assessing the accuracy of a surgical navigation system. The method may include acquiring an X-ray image that captures a surgical instrument. The method may also include segmenting the surgical instrument in the X-ray image. In an embodiment, the segmenting may be performed using edge detection or pattern recognition. The distance between the predicted location of the surgical instrument tip and the actual location of the surgical instrument tip may be computed. The distance between the predicted location of the surgical instrument tip and the actual location of the surgical instrument tip may be compared with a threshold value. If the distance between the predicted location of the surgical instrument tip and the actual location of the surgical instrument tip is greater than the threshold value, the user may be alerted. | 1. A method for assessing an accuracy of a surgical navigation system, the method comprising: acquiring an X-ray image capturing a surgical instrument tip;tracking a tracked location of the surgical instrument tip;segmenting the surgical instrument tip in the X-ray image to determine an imaged location of the surgical instrument tip; andcomputing a distance between the tracked location of the surgical instrument tip and the imaged location of the surgical instrument tip to form a navigation error of the surgical navigation system. 2. The method of claim 1, comprising comparing the distance between the tracked location of the surgical instrument tip and the imaged location of the surgical instrument tip with a threshold value. 3. The method of claim 2, comprising alerting a user if the navigation error of the surgical navigation system is greater than the threshold value. 4. The method of claim 1, wherein the computing the distance between the tracked location of the surgical instrument tip and the imaged location of the surgical instrument tip comprises: computing a back-projection line stemming from the imaged location of the surgical instrument tip; andcomputing a distance between the back-projection line and the tracked location of the surgical instrument tip. 5. The method of claim 1, wherein the computing the distance between the tracked location of the surgical instrument and the imaged location of the surgical instrument tip comprises: projecting 3D coordinates of the tracked location of the surgical instrument tip onto an image to form a tracked projection; andcomputing a distance between the tracked projection of the surgical instrument tip and an imaged projection of the surgical instrument tip on the image. 6. The method of claim 1, wherein the tracking the tracked location of the surgical instrument tip comprises tracking the tracked location of the surgical instrument tip with an electromagnetic tracking system. 7. The method of claim 1, wherein the tracking the tracked location of the surgical instrument tip comprises tracking the tracked location of the surgical instrument tip according to a planned trajectory of the surgical instrument tip. 8. A system for assessing an accuracy of a surgical navigation system, the system comprising: an X-ray unit configured to acquire an X-ray image capturing a surgical instrument tip;a tracking system configured to track a tracked location of the surgical instrument tip; anda computer unit configured to segment the surgical instrument tip in the X-ray image to form an imaged location of the surgical instrument tip, the computer unit configured to compute the distance between the tracked location of the surgical instrument tip and the imaged location of the surgical instrument tip to form a navigation error of the surgical navigation system. 9. The system of claim 8, wherein the computer unit is configured to compare the navigation error of the surgical navigation system with a threshold v... | GENERAL ELECTRIC COMPANY | A61B005/05 A61B006/12 G06T007/60 A61B005/06
US20080249400 | A1 | 2008-10-09 | Intraoperative Imaging Of Hepatobiliary Structures | 2007-10-05 | The invention provides methods for visualizing hepatobiliary structures, or lesions in a liver, intraoperatively by use of fluorescent dyes. | 1. A method of visualizing a hepatobiliary structure in a subject during a surgical operation, comprising: (a) within eight hours prior to, or during, said surgical operation, administering by subcutaneous injection, intramuscular injection, or slow, continuous intravenous infusion, a dye which fluoresces at an emission wavelength when said dye is contacted with an excitation wavelength;(b) exposing said hepatobiliary structure during said operation to a source of illumination comprising said excitation wavelength such that the fluorescent dye fluoresces; and(c) detecting the fluorescence of said dye,thereby visualizing said hepatobiliary structure during said surgical operation. 2. A method of claim 1, wherein said hepatobiliary structure is the bile duct, the common hepatic duct or the right hepatic duct. 3. A method of claim 1, wherein said hepatobiliary structure is the gall bladder or bile duct and the surgical operation is a cholecystectomy. 4. A method of claim 3, wherein said gall bladder or bile duct, or both, are inflamed. 5. A method of claim 1, wherein said hepatobiliary structure is the liver and the surgical operation is a resection of the liver for a living-related liver transplantation. 6. A method of claim 1, wherein said hepatobiliary structure is the liver and the surgical operation is transplantation of a donor liver. 7. A method of claim 1, wherein said hepatobiliary structure is the common bile duct and the surgical operation is removal of a stone from said bile duct. 8. A method of claim 1, wherein said surgical operation is by a laparoscopic instrument. 9. A method of claim 1, wherein said dye is administered one hour or less prior to said surgical operation. 10. A method of claim 1, wherein said dye is administered at the time of anesthesia induction. 11. A method of claim 1 wherein said dye is a near infrared dye. 12. A method of claim 11, wherein the near infrared dye is indocyanine green. 13. A method of determining, during a surgical operation to remove a patient's gall bladder, whether the patient's common bile duct is open following removal of the gall bladder, comprising: (a) after the gall bladder has been removed, but during the surgical operation, administering systemically a bolus of a dye which fluoresces at an emission wavelength when said dye is contacted with an excitation wavelength;(b) exposing said common bile duct to a source of illumination comprising said excitation wavelength such that the fluorescent dye fluoresces; and(c) detecting the presence or absence of fluorescence of said dye throughout said common bile duct,wherein presence of fluorescence of said dye throughout said common bile duct indicates that the common bile duct is open and wherein absence of fluorescence of said dye throughout said common bile duct indicates that the common bile duct is not open. 14. A method of claim 13, wherein said surgical operation is by a laparoscopic instrument. 15. A method of claim 13, wherein... | University of Rochester Medical Center | A61B005/00
US20080262376 | A1 | 2008-10-23 | WIRELESS SENSOR SYSTEM FOR MONITORING SKIN CONDITION USING THE BODY AS COMMUNICATION CONDUIT | 2008-04-17 | Devices and methods for measuring a local skin parameter or the presence or concentration of an analyte present in a biological medium are disclosed. A monitoring system comprising disposable sensor components and a network component for the collection of sensor information and for relaying this information for remote access and analysis is disclosed, where the sensor components and the network component communicate using the wearer as a signal propagation medium. | 1. A wireless sensor system using the body as a conduit for signal propagation, comprising: a wearable network component configured to use the body as a conduit for signal reception by using capacitive coupling between the network component and the body, the network component configured to receive a signal from a wearable sensor component, wherein the signal is indicative of a local skin parameter or an analyte in contact with skin. 2. The sensor system of claim 1, wherein the signal is indicative of the local skin parameter, and wherein the local skin parameter is pH, hydration, conductivity, temperature, or salinity. 3. The sensor system of claim 1, wherein the signal is indicative of an analyte in contact with skin, and wherein the analyte is water vapor or is present in feces or urine where the sensor component is located. 4. The sensor system of claim 1, wherein the network component is further configured to transmit the signal to an external device through a wireless network. 5. The sensor system of claim 1, further comprising: at least one wearable sensor component configured to be in contact with skin, wherein the sensor component generates the signal and transmits the signal to the network component by using capacitive coupling between the sensor component and the body. 6. The sensor system of claim 5, wherein the sensor component is configured to simultaneously sense two or more different local skin parameters or analytes in contact with skin. 7. The sensor system of claim 5, wherein the sensor component comprises a power source. 8. The sensor system of claim 5, wherein the sensor component is configured to extract power from its environment. 9. The sensor system of claim 5, the sensor component comprising: a sensing element; andcircuitry for generating the signal and transmitting the signal to the network component. 10. The sensor system of claim 5, further comprising: a device configured to receive the signal from the network component and to display, monitor, analyze, store, or report the signal. 11. A sensor device configured to be in contact with skin and uses the body as a conduit for signal propagation, comprising: a sensing element configured to sense a local skin parameter or an analyte in contact with skin; andcircuitry for generating a signal indicative of the local skin parameter or the analyte that is in contact with skin, and for transmitting the signal to a network component by using capacitive coupling between the sensor device and the body. 12. The sensor device of claim 11, wherein the signal is indicative of the local skin parameter, and wherein the local skin parameter is pH, hydration, conductivity, temperature, or salinity. 13. The sensor device of claim 11, wherein the signal is indicative of an analyte in contact with skin, and wherein the analyte is water vapor or is present in feces or urine where the sensor device is located. 14. The sensor device of claim 11, wherein the sensor device is configur... | PROACTIVE HEALTH DEVICES, INC. | A61B005/053
US20080294058 | A1 | 2008-11-27 | Wearable Device, System and Method for Measuring a Pulse While a User is in Motion | 2005-08-10 | A system, device and method are provided for monitoring a user's pulse, the system MAIN including a medical center server to provide commands to a wireless mobile pulse monitoring de-vice, the monitoring device including a pulse sensor sub-system including at least two pulse sensors adapted to accurately measure a pulse while a user is mobile. | 1. A system for monitoring a user's pulse, comprising a medical center server to provide commands to a wireless mobile pulse monitoring device, said device including a pulse sensor sub-system including at least two pulse sensors adapted to accurately measure a pulse while a user is mobile. 2. The system of claim 1, wherein said pulse sensor sub-system includes a sensor, an amplifier, and an analog-to-digital converter. 3. The system of claim 1, wherein said pulse sensor sub-system includes a Digital Signal Processing unit. 4. The system of claim 1, wherein said pulse sensor sub-system includes one or more analog filters. 5. The system of claim 1, wherein the functioning of said pulse sensor sub-system may be remotely configured by said server. 6. The system of claim 1, comprising an array of sensors for measuring physiological parameters. 7. The system of claim 1, comprising an array of sensors for measuring environmental parameters. 8. The system of claim 1, comprising an array of sensors for measuring physiological parameters and environmental parameters. 9. The system of claim 1, comprising one or more sensors selected from the group consisting of piezoelectric sensors and optical sensors. 10. The system of claim 1, wherein said monitoring device is to function in one or more of keeper mode, extended mode, and emergency mode. 11. The system of claim 1, wherein said monitoring device is to measure one or more selected parameters continuously and/or intermittently. 12. The system of claim 1, wherein said monitoring device is to automatically send a warning message to said medical center server if parameters measured exceed a selected threshold. 13. The system of claim 1, wherein said remote configuration of said mobile monitoring device includes remotely implementing selected software update(s). 14. The system of claim 1, wherein communications between said monitoring device and said medical center are encrypted. 15. The system of claim 1, wherein communications between said monitoring device and said medical center are authenticated. 16. The system of claim 1, wherein communications between said monitoring device and said medical center are encrypted and authenticated. 17. The system of claim 1, comprising a synchronization module. 18. The system of claim 1, comprising an update module. 19. The system of claim 1, comprising a data encryption module and a data authentication module. 20. A device for measuring pulse signals while a user is mobile, the device comprising: a pulse sensor sub-system including two or more pulse sensors adapted to accurately measure a pulse while the user is mobile; anda main controller to enable controlling of said pulse sensor sub-system. 21. The device of claim 20, wherein said pulse sensor sub-system includes at least one sensor to substantially measure pulse signals and at least one sensor to substantially measure pulse artifact signals. 22. The device of claim 20, comprising one or more senso... | - | A61B005/02
US20080262328 | A1 | 2008-10-23 | Pulse Oximetry Grip Sensor and Method of Making Same | 2006-01-23 | A probe for use with a pulse oximeter apparatus is disclosed, as are a fixture for and a method of making the probe. The probe includes a housing and two fiber optic bundles. The first bundle has an emitter portion at one end, and is used for conducting light from a source thereof to the emitter portion from which the light is transmitted for transillumination through a body part. The second bundle is for conducting the transilluminated light incident upon a detector portion thereof to the pulse oximeter apparatus. The housing is overmolded onto the bundles so that the emitter and detector portions are securely positioned diametrically opposite each other across an opening defined by the housing. The overmolding makes the alignment of the emitter and detector portions largely impervious to movement of the body part when placed within the opening and thus between the emitter and detector portions. | 1. A probe for use with a pulse oximeter apparatus, the probe comprising: (a) a first fiber optic bundle having an emitter portion at one end thereof, said first fiber optic bundle for conducting light from a source thereof to said emitter portion from which the light is transmitted for transillumination through a body part;(b) a second fiber optic bundle having a detector portion at one end thereof, said second fiber optic bundle for conducting the transilluminated light incident upon said detector portion to said pulse oximeter apparatus; and(c) a housing wherein said emitter and said detector portions of said first and said second fiber optic bundles, respectively, are overmolded so that said emitter and said detector portions are securely positioned diametrically opposite each other across an opening defined by said housing thereby making an alignment of said emitter and said detector portions largely impervious to movement of the body part when placed within said opening of said housing and thus between said emitter and said detector portions therein. 2. The probe claimed in claim 1 wherein said emitter and said detector portions are securely positioned diametrically opposite each other a predetermined distance apart, with said predetermined distance being dependent upon at least a size of said insert member. 3. The probe claimed in claim 1 wherein said emitter and said detector portions are securely positioned diametrically opposite each other a predetermined distance apart, with said predetermined distance being dependent upon at least a size of the body part. 4. The probe claimed in claim 1 wherein the body part said opening is shaped to accommodate includes at least one of a finger, an earlobe, and a nose. 5. The probe claimed in claim 1 wherein said housing is made of a flexible material. 6. A fixture for making a probe for a pulse oximeter apparatus, the fixture comprising: (a) a base member;(b) an opposing member; and(c) an insert member for placement between said base and said opposing members and therewith defining a cavity of predetermined shape wherein a housing of the probe is formable via introduction of a flowable material therein via at least one conduit defined by at least one of said members, said members further defining two channels for holding first and second fiber optic bundles, respectively, as said housing is overmolded thereabout so that an emitter portion of said first fiber optic bundle and a detector portion of said second fiber optic bundle are securely positioned diametrically opposite each other across an opening defined by said housing thereby making an alignment of said emitter and said detector portions largely impervious to movement of a body part when the body part is positioned within said opening of said housing and thus between said emitter and said detector portions therein during use of the probe. 7. The fixture claimed in claim 6 wherein said emitter and said detector portions are securely posi... | MEDRAD, INC. | A61B005/1455
US07434459 | B2 | 2008-10-14 | Context acquisition based on load sensing | 2003-09-26 | Methods and systems are described for sensing force information at a plurality of points on a continuous surface, and detecting contextual information about an object on the surface based on the distribution of force on the surface. For example, a plurality of sensors may be included below different points on such a substantially continuous surface, where the sensors are operable to sense force information. Then, a processor connected to the sensors may determine a location of the object relative to the surface, based on the force information from the sensors. | 1. A method comprising: sensing force information at a plurality of points on a substantially continuous surface, the force information related to an object on the surface;calculating a distribution of force exerted by the object, relative to the surface, based on the force information; anddetermining a location of the object on the surface, based on the distribution of force,wherein computing a center of force comprises: computing a total force applied to the surface by the object;computing the center of force of the object on the surface;determining a first component of the center of force by computing a first quotient of a first sum and a total force applied to the surface by the object, the first sum obtained by summing a first force applied at a first point and a second force applied at a second point; anddetermining a second component of the center of pressure, perpendicular to the first component, by computing a second quotient of a second sum and the total force, the second sum being obtained by summing the first force applied at the first point and a third force applied at a third point. 2. The method of claim 1 comprising: sensing a variation in the force information; anddetecting a change in the location of the object, relative to the surface, based on the variation in the force information. 3. The method of claim 1 comprising: computing the first and second components of the center of force of the object at a first period of time;computing the first and second components of the center of force of the object at a second period of time; anddetecting a new location of the object on the surface, based on a change in the first and second components from the first period to the second period. 4. The method of claim 1 comprising: sensing the force information at a later period of time; anddetecting a second location of a second object on the surface based on the force information. 5. The method of claim 4 comprising: identifying the object as a person; andtracking the position of the person. 6. The method of claim 4 comprising: identifying the first and second objects as people; andtracking a center of activity of the people. 7. The method of claim 1 in which sensing force information includes measuring a voltage level at a plurality of load cells, each of the load cells corresponding to each of the plurality of points. 8. The method of claim 7 comprising compensating for the force applied by the surface to the load cells. 9. A method comprising: sampling force information at points on a continuous surface during a plurality of time intervals;identifying an interaction between an object and the surface based on the sampled information;computing an average force on the surface during each of the time intervals;comparing a first average weight during a starting interval to a second average weight during an ending interval;determining that a first variability during the starting interval meets or exceeds a variability threshold value;co... | SAP Aktiengesellschaft | A61B005/00 G01L001/00 G01L013/00 G01L015/00 G01G007/00 G01G009/00 G08B021/00
US07458989 | B2 | 2008-12-02 | Intraoperative joint force measuring device, system and method | 2005-06-30 | A surgical device for joint replacement surgery includes an intraoperative joint head having a stem attachment structure where the stem attachment structure may be removably attachable to a stem and a force sensor housed by the joint head for measuring in vivo forces during surgery. The in vivo forces may be generated by one or more of tension provided by soft tissue, load application during surgery, limb movement during surgery, and a combination thereof. Also, the joint head may provide a cavity where the stem attachment structure houses the force sensor and may be removably insertable in the joint head cavity. A method of performing joint replacement surgery includes installing an intraoperative joint head having a force sensor housed by the joint head and a stem attachment structure, the stem attachment structure being removably attachable to a stem, measuring joint forces, and adjusting an implant parameter based on the measured joint forces. | 1. A surgical device for joint replacement surgery, comprising: an intraoperative joint head having a cavity for accepting the insertion of a stem attachment structure;a stem attachment structure being removably attachable to a stem and removably insertable in the joint head cavity;a plurality of force sensors disposed within the joint head for measuring in vivo forces during surgery;wherein the force sensors are placed at the periphery of the joint head cavity and the distance between adjacent force sensors is uniform, and at least one force sensors is located coaxial with a longitudinal axis of the stem such that the plurality of force sensors measure joint component forces to determine the resultant joint vector; andwherein the in vivo forces are generated from the group consisting of tension provided by soft tissue, load application during surgery, limb movement during surgery, and a combination thereof. 2. The device according to claim 1, wherein the plurality of force sensors that are placed at the periphery of the joint head cavity comprise three force sensors, each force sensor being located approximately 90° from an adjacent force sensor. 3. The device according to claim 1, wherein the joint head is at least partially spherically shaped. 4. The device according to claim 3, wherein the joint head is shaped as an anatomical structure from the group consisting of a femoral head, a humeral head, and a combination thereof. 5. The device according to claim 1, wherein the stem is the implant and the joint head is a trial joint head dimensioned as an implant for measuring in vivo forces during surgery. 6. The device according to claim 2, wherein the stem attachment structure includes a recess for receiving insertion of the stem. 7. The device according to claim 1, wherein the force sensor is a strain sensor. 8. A kit for assisting joint replacement procedures, the kit comprising: an intraoperative joint head having a cavity for accepting the insertion of a stem attachment structure;a stem attachment structure being removably attachable to a stem and removably insertable in the joint head cavity;a plurality of force sensors disposed within the joint head for measuring in vivo forces during surgery;wherein the force sensors are placed at the periphery of the joint head cavity and the distance between adjacent force sensors is uniform, and at least one force sensors is located coaxial with a longitudinal axis of the stem such that the plurality of force sensors measure joint component forces to determine the resultant joint vector;wherein the in vivo forces are generated from the group consisting of tension provided by soft tissue, load application during surgery, limb movement during surgery, and a combination thereof; andan intraoperative joint head tool. 9. The kit according to claim 8, wherein the joint head includes at least one series of joint heads, each joint head providing at least one unique dimension relative to the joint heads i... | University of Florida Rearch Foundation, Inc. | A61F002/32 A61F002/36 A61B005/00 A61B005/103
US07785265 | B2 | 2010-08-31 | Monitoring device and method | 2008-04-04 | The present invention relates to a method and an apparatus for providing information relating to disturbances and/or disease pattern based on a patient's breathing parameters. Preferably, the apparatus has multiple functionalities and provides information relating to different disturbances and/or disease patterns. In particular, the present invention relates to an apparatus for identifying different indications of a patient, the apparatus comprising a sensor, for sensing signals indicative of a patient's breathing, a processing unit, for processing and/or analyzing the signals, and display means for displaying the outcome of the processing and/or analysis. The corresponding method for identifying different indications of a patient comprises the steps of sensing signals indicative of a patient's breathing, processing and/or analyzing the signals, and displaying the outcome of the processing and/or analysis. | 1. An apparatus for identifying different indications of a patient, the apparatus comprising a sensor configured to sense signals indicative of a patient's breathing;a processing unit configured to process and/or analyze the signals by generating flow and snore channel signals;a snore detection module configured to calculate a snoring index based on the flow and snore channel signals;an Obstructive Sleep Apnea (OSA) detection module configured to calculate an apnea/hypopnea index based on the flow channel signal;a periodic breathing and/or Cheyne-Stokes (CS) detection module configured to calculate a periodic breathing index and/or CS index based on the flow channel signal and output from the OSA detection module; anda display configured to display data related to output from the snore, OSA, and periodic breathing and/or CS detection modules. 2. The apparatus according to claim 1, wherein the display is further configured to display data related to output from the snore, OSA, and periodic breathing and/or CS detection modules as logical YES or NO indicators, the logical YES or NO indicators corresponding the presence or absence of patient snore, OSA, and periodic breathing and/or CS, respectively. 3. The apparatus according to claim 2, wherein the indicators comprise LEDs. 4. The apparatus according to claim 1, wherein the apparatus is a closed unit. 5. The apparatus according to claim 1, wherein the apparatus is a handheld and/or pocket size device. 6. The apparatus according to claim 1, wherein the apparatus comprises an energy source. 7. The apparatus according to claim 6, wherein the energy source is an accumulator or a rechargeable battery. 8. The apparatus according to claim 1, wherein the apparatus further comprises at least one A/D converter. 9. The apparatus according to claim 1, wherein the snore detection module includes a filter and an envelope generator. 10. The apparatus according to claim 9, wherein the filter has a lower frequency cutoff of about 10 Hz. 11. The apparatus according to claim 10, wherein the filter has an upper frequency cutoff of about 50-100 Hz. 12. The apparatus according to claim 1, wherein the apparatus is a single-use apparatus. 13. The apparatus according to claim 1, further comprising first and second A/D converters, the first A/D converter being usable in the generation of the flow channel signal and the second A/D converter being usable in the generation of the snore channel signal. 14. A portable handheld diagnostic apparatus, comprising: a sensor configured to sense signals indicative of a patient's breathing;a processing unit configured to generate flow and snore channel signals based on the signals sensed by the sensor;a snore detection module configured to calculate a snoring index based on the flow and snore channel signals;an Obstructive Sleep Apnea (OSA) detection module configured to calculate an apnea/hypopnea index based on the flow channel signal;a periodic breathing and/or Cheyne-S... | Map Medizin-Technologie GmbH | A61B005/08
US20080255469 | A1 | 2008-10-16 | Method for Monitoring the Depth of Anesthesia | 2007-04-12 | A method for monitoring the depth of anesthesia is provided for detecting the conscious state of one being anesthetized in order to facilitate an anesthesiologist to predict exactly the dosage of an anesthetic required. At first, an original electroencephalogram (EEG) is taken from one being tested. Then, the original electroencephalogram is analyzed by approximate entropy to obtain its approximate entropy value. Next, the approximate entropy value is multiplied by 1000/17, and the corrected value is assumed as the predicted value of depth of anesthesia. The predicted value of depth of anesthesia represents degree of the conscious state or the depth of anesthesia for the one being tested. The higher the predicted depth of anesthesia value, the more conscious the one being tested is, i.e., in a shallower depth of anesthesia. On the other hand, the lower the predicted depth of anesthesia value, the less conscious the one being tested is, i.e., in a deeper depth of anesthesia. | 1. A method for predicting the depth of anesthesia, comprising steps of: step 1: obtaining original electroencephalogram (EEG) from one being tested;step 2: calculating the approximate entropy value from the original electroencephalogram (EEG) using following formula: Approximat Entropy=Φm(r)−Φm+1(r);whereinΦm(r)=(N-m+1)-1·∑i=1N-m+1lnCim(r);Cim(r)=(number of x(j) such that d[x(i),x(j)]≦r)/(N−m+1);x(i)=[u(i), . . . , u(i=m−1)];x(j)=[u(j), . . . , u(j=m−1)];u(i),u(2) . . . u(N) are time sequence data;wherein N is the length of data cycle; m is the number of data comparison;r is a noise filtering coefficient; andstep 3: multiplying the approximate entropy value obtained in step 2 by 1000/17, and using thus corrected value as the predicting value of depth of anesthesia. 2. A method for predicting the depth of anesthesia as recited in claim 1, wherein the sampling time of the original electroencephalogram (EEG) is 1/256 to 1/128 second/time. 3. A method for predicting the depth of anesthesia as recited in claim 1, wherein the N value is 1024. 4. A method for predicting the depth of anesthesia as recited in claim 1, wherein the m value is 2. 5. A method for predicting the depth of anesthesia as recited in claim 1, wherein the r value is 0.2. 6. A method for predicting the depth of anesthesia as recited in claim 1, wherein the predicting value of depth of anesthesia represents the degree of consciousness state or the depth of anesthesia of one being tested. | YUAN ZE UNIVERSITY | A61B005/04
US07458937 | B2 | 2008-12-02 | Method and system for assessing breathing effectiveness via assessment of the dynamic arterial pressure wave using the oscillometric measurement technique | 2005-01-11 | The present invention specifies a method and system for assessing breathing effectiveness via dynamic assessment of respiratory sinus arrhythmia and the consequent arterial pressure wave using the oscillometric measurement technique. Arterial pressure wave peak pressure, valley pressure, amplitude, and rate are characterized and displayed for diagnostic and remedial purposes. The dynamic change in respiratory sinus arrhythmia and resultant arterial pressure wave is visually presented in real time as user biofeedback. A respiratory sinus arrhythmia generator generates varying frequencies to which the user synchronizes their breathing cycle for purposes of increasing arterial pressure wave amplitude with a consequent increase in breathing depth and decrease in breathing frequency. An instructive method in the application of the present invention is also specified. | 1. A method for assessing breathing optimality of a patient comprising: measuring a varying breathing induced arterial pressure wave of a patient;assessing optimality of a breathing process of the patient based on the measured arterial pressure wave; andproviding a characterization to the patient of the optimality of the patient's breathing process based on the measured arterial pressure wave. 2. The method of claim 1 wherein measuring the varying breathing induced arterial pressure wave of the patient comprises measuring a varying pulse pressure of the patient via oscillometry. 3. The method of claim 1 wherein measuring the varying breathing induced arterial pressure wave of the patient comprises measuring one or more from the group consisting of a pulse pressure amplitude of the patient, pulse pressure periodicity of the patient, and coherence of the patient. 4. The method of claim 1 wherein measuring the varying breathing induced arterial pressure wave of the patient comprises measuring oscillometric characteristics of the arterial pressure wave of the patient over a measurement interval including multiple breathing cycles; and wherein assessing optimality of a breathing process of the patient based on the measured arterial pressure wave comprises assessing changes in the arterial pressure wave pressure of the patient over the measurement interval using the oscillometric characteristics. 5. The method of claim 1 wherein assessing the optimality of a breathing process of the patient is performed for an approximate 60 second period defining an arterial pressure wave assessment cycle and further comprises: sustaining a pulse oscillation for the approximate 60 second period by gradually reducing a cuff pressure; andmeasuring multiple arterial pressure wave cycles of the patient during the arterial pressure wave assessment cycle. 6. The method of claim 1 wherein measuring the varying breathing induced arterial pressure wave of the patient further comprises detecting cardiac systole of the patient; anddetecting pulse oscillations of the patient corresponding to the cardiac systole occurring at the peak of the breathing induced arterial pressure wave. 7. The method of claim 1 wherein measuring the varying breathing induced arterial pressure wave of the patient comprises: measuring an oscillometric pulse oscillation amplitude of the patient; andmeasuring an arterial pressure wave amplitude of the patient as a function of the measured oscillometric pulse oscillation amplitude of the patient, andwherein providing a characterization to the patient of the optimality of the patient's breathing process based on the measured arterial pressure wave comprises displaying the arterial pressure wave amplitude of the patient on a display for diagnostic and remedial purposes. 8. The method of claim 1 wherein measuring the varying breathing induced arterial pressure wave of the patient comprises: measuring an oscillometric pulse oscillation periodicity of the... | Coherence LLC | A61B005/00 A61B005/02
US20080275369 | A1 | 2008-11-06 | Arrangement and Method for Determining Muscular Contractions in an Anatomical Organ | 2005-03-24 | The invention relates to an arrangement for determining the degree of muscular contraction of a generally tubular anatomical organ (4), said arrangement comprising a longitudinally extending catheter (9) being adapted for introducing into said organ (4), and a measuring unit (11; 12) for cooperating with said catheter (9) and being adapted for assessing said degree of contraction. According to the intention, said catheter (9) is provided with a sensor arrangement (10; 22; 25) for indicating the geometrical configuration of said catheter (9), and said measuring unit (11; 12) is provided with means for determining the degree of longitudinal muscular contraction of said organ (4) based on values indicating said geometrical configuration. The invention also relates to a catheter for of generally longitudinal extension and being adapted for introducing into a generally tubular anatomical organ. The invention also relates to a method for assessing the degree of said muscular contraction. | 1. Arrangement for determining the degree of muscular contraction of a generally tubular anatomical organ (4), said arrangement comprising: a longitudinally extending catheter (9) being adapted for introducing into said organ (4), and a measuring unit (11; 12) for cooperating with said catheter (9) and being adapted for assessing said degree of contraction, characterized in that said catheter (9) is provided with a sensor arrangement (10; 22; 25) for indicating the geometrical configuration of said catheter (9), and that said measuring unit (11; 12) is provided with means for determining the degree of longitudinal muscular contraction of said organ (4) based on values indicating said geometrical configuration. 2. Arrangement according to claim 1,characterized in that said sensor arrangement (10) comprises a plurality of radiologically opaque markers (10) arranged on said catheter (9), and in that said measuring unit (11; 12) comprises a fluoroscopy scanning device (11) adapted for detecting the positions of said markers (10). 3. Arrangement according to claim 1,characterized in that said sensor arrangement (10) comprises a plurality of strain gauges (22) arranged in the catheter (9) for detecting bending strain acting upon said catheter (9). 4. Arrangement according to claim 1,characterized in that said sensor arrangement (10) comprises at least one optical fiber (24) being adapted for detecting bending thereof. 5. Arrangement according to claim 1,characterized in that said sensor arrangement (10) comprises a plurality of transponders (25) which cooperate with a sender/receiver arrangement (26, 27) for providing information regarding the position of each transponder. 6. Arrangement according to claim 1, characterized in that said catheter (9) also comprises means (19) for side-hole manometry in said organ (4), for assessing circumferential muscular contractions of said organ (4). 7. Arrangement according to claim 1, characterized in that it is adapted for determining the degree of contraction of muscular contraction of a generally tubular organ (4) in a gastrointestinal system of a human or animal. 8. Catheter (9) of generally longitudinal extension and being adapted for introducing into a generally tubular anatomical organ (4), characterized in that said catheter (9) is provided with a sensor arrangement (10; 22; 25) for indicating the geometrical configuration of said catheter (9). 9. Catheter (9) according to claim 8,characterized in that said sensor arrangement comprises a plurality of radiologically opaque markers (10). 10. Catheter (9) according to claim 8,characterized in that said sensor arrangement comprises a plurality of strain gauges (22) for detecting bending strain acting upon said catheter (9). 11. Catheter (9) according to claim 8,characterized in that said sensor arrangement comprises at least one optical fiber (24) being adapted for detecting bending thereof. 12. Catheter (9) according to claim 8,characterized in that... | - | A61B005/103
US20080281219 | A1 | 2008-11-13 | Method and System for Assessing Lung Condition and Managing Mechanical Respiratory Ventilation | 2008-02-21 | The present invention discloses a novel non-invasive, bedside system and method to monitor parameters associated lung changes. A novel approach for monitoring the operation of the respiratory system of a subject is provided. There is also provided a method for objectively evaluating the benefit of one mode of ventilation over another, and for assessing the differences in regional lung vibration during different modes of mechanical ventilation. The method comprises recording one or more signals from the subject, the signal varying in time according to operation of the respiratory system; and; processing the recorded signals to obtain a predetermined functional thereof presenting one or more time-varying energy functions of the subject, an abnormality in the one or more energy functions being indicative of a suspected abnormality in the operation of the respiratory system. The signals may be acoustic signals recorded by a plurality of acoustic sensors placed over the subject's thorax or back, and the at least one time-varying energy function is obtained from one or more specific regions of lung or by summing/averaging the time-dependent acoustic signals of the plurality of sensors indicative of the whole lungs. | 1. A method for use in monitoring the respiratory system of a subject, the method comprising: (a) recording one or more signals from the subject, the signal varying in time according to operation of the respiratory system; and;(b) processing the recorded signals to obtain a predetermined functional thereof presenting one or more time-varying energy functions of the subject, an abnormality in said one or more energy functions being indicative of a suspected abnormality in the operation of the respiratory system. 2. A method according to claim 1, wherein the signals are acoustic signals. 3. A method according to claim 2, wherein the signals are recorded by a plurality of acoustic sensors placed over the subject's thorax or back, and said at least one time-varying energy function is obtained from one or more specific regions of lung. 4. A method according to claim 3, comprising summing or averaging the time-dependent acoustic signals of the plurality of sensors indicative of the whole lungs. 5. A method according to claim 1, wherein the one or more energy functions of the subject is/are displayed in form of one or more graphs. 6. A method according to claim 1, wherein the one or more energy functions of the subject is/are displayed in form of still or dynamic digital image, or in form of succession of still images or frames. 7. A method according to claim 6, wherein the still or dynamic digital image is indicative of the regional distribution of vibrations in the lungs. 8. A method according to claim 6, comprising analyzing said dynamic image thereby providing data indicative of the intensity and distribution of vibration within lungs in real-time. 9. A method according to claim 1, wherein the one or more energy functions of the subject is/are displayed in form of graph and in form of image. 10. A method according to claim 9, comprising analyzing said at least one energy function graph to select an appropriate still image in a dynamic image recording. 11. A method according to claim 1, wherein said processing of the one or more energy functions comprises determining a degree of correlation between one or more parameters of the energy function and one or more corresponding parameters of certain reference energy function. 12. A method according to claim 11, wherein said one or more parameters of the energy function include at least one of the following: geographical distribution and/or intensity of energy, synchronization and/or balance between lungs, signal periodicity and/or signal symmetry of the function. 13. A method according to claim 1, comprising utilizing results of said processing of the energy function for optimizing the operational mode of mechanical ventilation. 14. A method according to claim 1, comprising utilizing results of said processing of the energy function for selecting optimized parameters set for a specific mode. 15. A method according to claim 13, comprising utilizing results of said processing of the energy fun... | DeepBreeze Ltd. | A61B005/08
US20080312518 | A1 | 2008-12-18 | ON-DEMAND ANALYTE MONITOR AND METHOD OF USE | 2007-06-14 | An analyte monitor is provided with a sensor unit body configured for mounting on tissue, a sensor configured to detect an analyte in a fluid in the sensing area, an output device configured to communicate a result from the sensor to a user; and a user input device coupled with the sensor and the output device, wherein the monitor is configured to communicate a result to the user through the output device only after the user input device is activated. Systems, sensors and methods associated with the monitor are also disclosed. | 1. An analyte monitor comprising: a sensor unit body configured for mounting on tissue;at least one tissue piercing element extending from the sensor body, the tissue piercing element comprising a distal opening, a proximal opening, and an interior lumen extending between the distal and proximal openings;a sensing area located on the sensor unit body in fluid communication with the proximal opening of the tissue piercing element;a sensor configured to detect an analyte in a fluid in the sensing area;an output device configured to communicate a result from the sensor to a user; anda user input device coupled with the sensor and the output device, wherein the monitor is configured to communicate a result to the user through the output device only after the user input device is activated. 2. An analyte monitor according to claim 1, wherein the monitor is configured to communicate only a predetermined number of results to the user. 3. An analyte monitor according to claim 2, wherein the predetermined number is at least 2. 4. An analyte monitor according to claim 2, wherein the predetermined number is no more than 200. 5. An analyte monitor according to claim 1, wherein the monitor is configured to communicate only a predetermined number of results per a predetermined time period to the user. 6. An analyte monitor according to claim 5, wherein the predetermined number and time period are at least one result per hour. 7. An analyte monitor according to claim 5, wherein the predetermined number and time period are no more than 6 per hour and at least 5 minutes apart from each other. 8. An analyte monitor according to claim 1 further comprising a handheld unit that wirelessly communicates with the sensor unit body, wherein the handheld unit houses the output device. 9. An analyte monitor according to claim 8, wherein the handheld unit houses the user input device. 10. An analyte monitor according to claim 1, wherein the sensor unit body houses the output device. 11. An analyte monitor according to claim 1, wherein the sensor unit body houses the input device. 12. An analyte monitor according to claim 1, wherein the output device comprises a numeric display. 13. An analyte monitor according to claim 1, wherein the input device comprises a push button. 14. An analyte monitor according to claim 1, wherein the output device and the input device comprise a single touch screen display. 15. An analyte monitor according to claim 1, wherein the sensor is configured to detect a presence of the analyte. 16. An analyte monitor according to claim 1, wherein the sensor is configured to detect a concentration or amount of the analyte. 17. An analyte monitor according to claim 1, wherein the sensor is configured to detect a glucose concentration. 18. An analyte monitor according to claim 1, wherein the at least one tissue piercing element comprises a plurality of micro-needles. 19. An analyte monitor according to claim 1, wherein the lumen is configure... | ArKal Medical, Inc | A61B005/00
US20080239920 | A1 | 2008-10-02 | Wireless transmission of temperature data | 2007-03-27 | An identification input device and a temperature detector that detects body temperature data. The identification input device may be an optical touch pen. The infrared touch pen may be affixed to the temperature detector using a hook and connector or a snap fit connector. In use, the temperature detector computes body temperature data and the identification input device determines an identifier. The identifier is unique to a user. After obtaining the body temperature data and identifier, a transmitter may transmit body temperature data and the identifier over a wireless communications path to a processing unit. By transmitting data over a wireless communications path, an individual's data or a large group of data may be analyzed and viewed via a display unit. | 1. A method of collecting temperature data, comprising: determining an identifier of an individual for each of plural individuals;detecting body temperature data of the individual associated with the identifier; andtransmitting the body temperature data and the identifier over a wireless communications path to a processor. 2. The method of claim 1 wherein determining an identifier of the individual is through an identification reader. 3. The method of claim 2 wherein the identification reader is a touch pen. 4. The method of claim 3 wherein the touch pen is a bar code reader. 5. The method of claim 3 wherein the touch pen is affixed to a temperature detector using a connector. 6. The method of claim 2 wherein the identification reader includes a bar code reader. 7. The method of claim 1 wherein detecting body temperature data of the individual further comprises scanning across the temporal artery to obtain body temperature data. 8. The method of claim 1 further comprising storing the body temperature data and the identifier in memory of the processor. 9. The method of claim 1 further comprising processing the body temperature data using the processor. 10. The method of claim 9 further comprising displaying a result for the body temperature data to a user of the processor. 11. An apparatus for collecting a temperature data, comprising: an identification input device;a temperature detector; anda transmitter that transmits body temperature data from the temperature detector and an identifier from the identification input device over a wireless communications path to a processing unit. 12. The apparatus of claim 11 wherein the identification input device is an identification reader. 13. The apparatus of claim 12 wherein the identification input device is a touch pen. 14. The apparatus of claim 13 wherein the touch pen is a bar code reader. 15. The apparatus of claim 13 wherein the touch pen is affixed to the temperature detector by using a connector. 16. The apparatus of claim 11 wherein the identification input device includes a bar code reader. 17. The method of claim 11 wherein the temperature detector scans across the temporal artery to obtain body temperature data. 18. The apparatus of claim 11 further comprising memory in the processor to store the body temperature data and the identifier. 19. The apparatus of claim 11 further comprising the processing unit processes the body temperature data. 20. The apparatus of claim 19 further comprising a display unit of the processor that displays a result for the body temperature data to a user. 21. An apparatus for collecting a temperature data, comprising: means for determining an identifier of an individual for each of plural individuals;means for detecting body temperature data of the individual associated with the identifier; andmeans for transmitting the body temperature data and the identifier over a wireless communications path to a processor. | - | G11B007/00 A61B005/00
US20080281212 | A1 | 2008-11-13 | Transseptal monitoring device | 2008-03-14 | An implantable medical device includes a first sensor array movable between a collapsed configuration and a deployed configuration. The first sensor array includes a first main portion and at least one rigid sensor movably coupled to the first main portion. A coupler is operatively coupled to the first sensor array and configured to couple the implantable medical device with respect to a heart chamber. | 1. An implantable sensing unit comprising: an anchoring mechanism; anda first sensor coupled to said anchoring mechanism, said first sensor configured to sense at least one of a physical, chemical, and physiological parameter of at least one heart chamber. 2. An implantable sensing unit in accordance with claim 1 wherein said anchoring mechanism comprises at least one of a substrate, a post, a tethering member, a hook, a helical wire, a barbed member and a screw. 3. An implantable sensing unit in accordance with claim 1 wherein said anchoring mechanism comprises a post defining said first end of said anchoring mechanism, and a coupling member configured to couple said implantable sensing unit with respect to the heart chamber. 4. An implantable sensing unit in accordance with claim 1 wherein said anchoring mechanism comprises a tethering member, said implantable sensing unit further comprising a second sensor coupled to a second end of said anchoring mechanism opposing said first end, and said tethering mechanism biasing said second sensor towards said first sensor. 5. An implantable sensing unit in accordance with claim 4 wherein a surface of said first sensor is positioned within a plane normal to a central axis of said tethering member. 6. An implantable sensing unit in accordance with claim 1 wherein said first sensor comprises one of a pressure sensor, an optical sensor, a biochemical sensor, a protein sensor, a motion sensor, an accelerometer, a gyroscope, a temperature sensor, a chemical sensor, a pH sensor, and a genetic sensor. 7. An implantable sensing unit in accordance with claim 1 wherein said first sensor comprises one of a capacitive pressure sensing device, an inductive pressure sensing device, a piezoelectric pressure sensing device and a piezoresistive pressure sensing device. 8. An implantable medical device comprising: a first sensor array movable between a collapsed configuration and a deployed configuration, said first sensor array comprising a first main portion and at least one rigid sensor movably coupled to said first main portion; anda coupler operatively coupled to said first sensor array and configured to couple said implantable medical device with respect to a heart chamber. 9. An implantable medical device in accordance with claim 8 wherein said first sensor array comprises a plurality of rigid sensors, each sensor of said plurality of rigid sensors movably coupled to said first main portion. 10. An implantable medical device in accordance with claim 8 further comprising a second sensor array operatively coupled to said coupler, said second sensor array movable between a collapsed configuration and a deployed configuration, and comprising a second main portion and at least one rigid sensor movably coupled to said second main portion. 11. An implantable medical device in accordance with claim 10 wherein said coupler has a first end and an opposing second end, said first sensor array coupled to said first en... | - | A61B005/02
US20080249386 | A1 | 2008-10-09 | Systems, Methods, and Computer Program Product for Improved Management of Medical Procedures for Patients on Medical Protocols | 2008-04-03 | Techniques are described for improved monitoring and control of a patient's physiologic status, such as blood glucose control according to a protocol for use in a patient care unit, such as an intensive care unit. Instructions to medical staff are adapted to factor in variations in responses to protocol recommendations. For example, a patient's physiological data as a result of a blood analysis is submitted as input to a program. An instruction for a medical procedure for the patient in response to the physiological data is automatically provided by the program. A confirmation response that the medical procedure has been accomplished and results of following the medical procedure are requested by the program. The time when the medical procedure was accomplished and the results of following the medical procedure are evaluated by the program to determine whether to adapt a subsequent instruction. | 1. A method for adapting management of medical procedures for a patient assigned to a medical protocol, the method comprising: receiving periodically at a management process a report of a patient's monitored physiological data and confirmed medical procedures from a patient's monitoring station;adapting an instruction for management of the medical procedures for the patient in response to an evaluation of the patient's periodic reports; andsending the adapted instruction from the management process to the patient's monitoring station or to selected other monitoring stations based on the evaluation. 2. The method of claim 1 further comprises: receiving an updated patient report of the patient's monitored physiological data in response to a first medical procedure according to the adapted instruction, wherein the first medical procedure is a member of a first set of medical procedures which can be accomplished by a member of the medical staff having authorization for the first set of medical procedures; andadapting a second instruction for patient management in response to an evaluation of the updated patient's report, wherein the second adapted instruction includes a request for a second medical procedure based on the medical protocol and a request for confirmation that the second medical procedure has been accomplished, wherein the second medical procedure is a member of a second set of medical procedures. 3. The method of claim 2, further comprising: sending the second adapted instruction from the management process to the patient's monitoring station; andconfirming the second medical procedure according to the second adapted instruction has been accomplished by the member of the medical staff having authorization for the second set of medical procedures. 4. The method of claim 2 wherein the first set of medical procedures includes taking a blood glucose reading and wherein the second medical procedure includes changing an insulin infusion drip rate. 5. The method of claim 3 further comprises: sending an alert to the patient's monitoring station indicating the second medical procedure has not been confirmed within a specified period; andsending repeatedly an escalating the alert to the patient's monitoring station until the confirmation is received. 6. The method of claim 1 further comprises: sending an alert to the selected other monitoring stations in response to not receiving the patient's periodic reports at the management process, the alert requesting that operation of the patient's monitoring station be checked. 7. The method of claim 1 further comprises: determining whether multiple alerts are to be sent to one or more monitoring stations within a specified time interval;reducing the number of alerts sent to the same monitoring station without loss of information; andstaging selected groups of alerts to be sent to the one or more monitoring stations to allow the available medical staff time to respond to each individual alert. 8. T... | Pronia Medical Systems, LLC | A61B005/145 G06Q050/00 A61M005/14
US07462155 | B2 | 2008-12-09 | Objective determination of chronic pain in patients | 2004-10-27 | A method for objectively determining if a patient suffers chronic pain by initially observing a multiplicity of persons who are free of chronic pain and subjecting them to a given pain stimulus. Radiofrequency (RF) emissions generated by the brains of the persons who are free of chronic pain while they are subjected to the pain stimulus are measured and a chart in the form of an average f-MRI image for all such persons is prepared which statistically reflects an average of RF emissions generated by the brains of the persons who are free of chronic pain while subjected to the pain stimulus. The same pain stimulus is then applied to the patient, RF emissions generated by a brain of the patient while the pain stimulus is applied are sensed, and the RF emissions from the patient's brain are compared to the chart to determine if the RF emissions from the patient's brain are statistically different from the approximately average RF emissions from the persons who are free of chronic pain on the chart. If so, it can be concluded without any conscious input from the patient that the patient suffers chronic pain. The method can be used to verify insurance or other compensation claims by persons who assert that they suffer chronic pain that might prevent them from performing certain functions such as work. | 1. A method for determining the presence of chronic pain in a patient claiming to suffer chronic pain, comprising applying a pain stimulus to a part of the body of the patient claiming to suffer chronic pain, observing neuron activity in the brain caused by the pain stimulus, comparing observed neuron activity to neuron activity caused by the stimulus when applied to a plurality of persons, wherein each control person is free of pain, and deciding that the patient claiming to suffer chronic pain suffers chronic pain if the neuron activity observed in the patient while the stimulus is applied is judged to be statistically significantly increased from the neuron activity in the plurality of patients who are free of pain, wherein the deciding step is performed without the subjective input from the patient. 2. A method according to claim 1 wherein observing neuron activity in the patient's brain comprises observing physical emissions generated by the brain of the patient. 3. A method according to claim 2 wherein observing physical emissions comprises observing radiofrequency (RF) emissions generated by the brain. 4. A method according to claim 2 wherein observing neuron activity in the RF emissions comprises observing an intensity of the RF emissions and comparing them to an intensity of RF emissions from the plurality of control persons who are free of pain. 5. A method according to claim 3 wherein observing RF emissions comprises observing a pattern of RF emissions generated by the patient's brain, and wherein comparing comprises comparing the observed RF emission pattern in the patient's brain with an RF emission pattern for the control persons who are free of pain. 6. A method according to claim 5 wherein observing patterns of RF emissions includes observing patterns of RF emissions in the patient and in the plurality of control persons who are free of pain in prefrontal areas of the brains. 7. A method according to claim 3 wherein comparing comprises comparing f-MRIs of the patient with a standard f-MRI based on RF emissions by the brains of the control persons who are free of pain. 8. A method according to claim 7 wherein observing neuron activity in the brain of the patient includes generating an f-MRI of RF emissions by the patient's brain. 9. A method according to claim 8 including preparing a standard f-MRI for RF emissions by the brains of control persons who are free of pain which is a statistical summary of the f-MRIs of the control persons who are free of pain, and wherein comparing comprises comparing the f-MRJ of the patient with the standard f-MRI. 10. A method according to claim 9 wherein comparing comprises processing the patient's f-MRI and the standard f-MRI in a computer and generating a computer output which quantifies a difference between the patient's f-MRI and the standard f-MRI. 11. A method according to claim 10 including judging if the difference between the patient's f-MRI and the standard f-MRI is statisticall... | - | A61B019/00 A61B005/05 A61B005/04 A61B005/00
US07448996 | B2 | 2008-11-11 | Method and apparatus for remotely monitoring the condition of a patient | 2003-04-15 | A patient monitoring system provides enhanced functional capability relative to known systems and provides a wireless communication link between a patient monitoring device, worn by a patient, and a local hub. The patient monitoring system is adapted to monitor various patient physiological characteristics, such as blood pressure, pulse rate, blood glucose, weight, pulse oximetry and others. The data from the patient monitoring device is wirelessly transmitted to a local hub, which, in turn, is configured to automatically transfer the data to a remote server, for example, over a public or private communications network. In one embodiment of the invention, the server is configured as a web portal to selectively allow access to such patient physiological data by designated third parties, such as physicians, clinicians, relatives and the patient themselves. | 1. A patient monitoring system comprising: one or more patient monitoring devices for one or more patients, each patient monitoring device including a physiological transducer monitoring physiological characteristic data of said one or more patients and storing said physiological characteristic data in a device memory, and a first wireless communication system automatically transmitting and synchronizing said physiological characteristic data to a common hub over a first communication link,wherein said physiological data is associated with a time stamp representing a time reading from an internal microprocessor taken when said physiological characteristic data is recorded,wherein said synchronizing includes comparing said time stamp with a time stamp received from a common hub and transmitting said physiological characteristic data and said associated time stamp to said hub only when said time stamp indicates a time after said time stamp received from said hub,wherein said physiological characteristic data and associated time stamp are deleted from said device memory when said associated time stamp represents a time earlier than said time stamp received from said hub;wherein said internal microprocessor receives an indication of the current time from said hub and updates its time to the current time received from said hub;a common hub remote from said patient monitoring devices and including a second communication system automatically receiving and synchronizing said physiological characteristic data from said patient monitoring devices and exchanging messages with said patient monitoring devices over said first communication link from said one or more patient monitoring devices and automatically transmitting and synchronizing said physiological characteristic data, received from said patient monitoring devices, to a remote server over a second communication link,wherein synchronizing said hub with said patient monitoring device includes transmitting a time stamp stored at said hub to said patient monitoring device,wherein said time stamp stored at said hub is previously received from said remote server;wherein synchronizing said hub with said remote server includes receiving a time stamp from said remote server and replacing said time stamp stored at said hub with said time stamp from said remote server, andwherein said hub receives a current time indication from said remote server and updates an internal hub microprocessor to said current time indication, wherein a time indication from said internal hub microprocessor is later sent to said patient monitoring device to update the current time at said patient monitoring device; anda remote server including a third communication system receiving said physiological characteristic data from said hub over said second communication link, which automatically synchronizes said physiological characteristic data from said common hub and exchanges messages with said common hub over said second communicatio... | Carematix, Inc. | A61B005/00
US08317718 | B2 | 2012-11-27 | Methods of testing digestive functions using both a breath test and a scintigraphy test, and methods of using a breath test as an overall digestive health assessment | 2008-05-07 | Methods for assessing an overall digestive health of a patient are provided, which include administering a breath test that uses a label incorporated into proteins, carbohydrates, and lipids and further incorporating the labeled material into a meal which, when cooked, facilitates the binding of the label material to the solid phase matrix of the meal. Methods for assessing digestive functions of a patient are also provided, which includes concurrently administering a breath test and a scintigraphy test that measure gastric emptying in a stomach. | 1. A method of assessing digestive functions of a patient, wherein the digestive functions include digestive functions that occur post-stomach, the method comprising: concurrently administering a breath test and a scintigraphy test that measure gastric emptying in a stomach, wherein the breath test and the scintigraphy test have a diagnostic concordance to one another, wherein the breath test uses a label incorporated into proteins, carbohydrates, and lipids;obtaining a breath test result, the breath test result being either normal or abnormal;obtaining a scintigraphy test result, the scintigraphy test result being either normal or abnormal; andclassifying the patient's digestive functions as normal, stomach abnormal, or post-stomach abnormal, wherein said normal classification is used when both the breath test result and the scintigraphy test result are normal, wherein said stomach abnormal classification is used when both the breath test result and the scintigraphy test result are abnormal, and wherein said post-stomach abnormal classification is used when the breath test result is abnormal and the scintigraphy test result is normal. 2. The method of claim 1, wherein the classifying the patient's digestive function further comprises classifying the patient's digestive function as both stomach abnormal and post-stomach abnormal when the breath test result and scintigraphy test result are each abnormal and the breath test data is inconsistent with the scintigraphy test data. 3. The method of claim 1, wherein the classifying the patient's digestive function as stomach abnormal comprises classifying the digestive function as stomach abnormal only when the breath test data is consistent with the scintigraphy test data. 4. The method of claim 1, wherein the obtaining a breath test result comprises selecting a normal time point result or an abnormal time point result at a series of time points, the series of time points including a 45 minute time point, a 90 minute time point, a 120 minute time point, a 150 minute time point, and a 180 minute time point, the selecting including: selecting a normal time point result if the breath test yields a measurement that is between a 12.9 and 32.9 kPCD range or selecting an abnormal time point result if the breath test is outside of the 12.9 and 32.9 kPCD range at the 45 minute time point;selecting a normal time point result if the breath test yields a measurement that is between a 33.1 and 63.2 kPCD range or selecting an abnormal time point result if the breath test is outside of the 33.1 and 63.2 kPCD range at the 90 minute time point;selecting a normal time point result if the breath test yields a measurement that is between a 44.2 and 72.5 kPCD range or selecting an abnormal time point result if the breath test is outside of the 44.2 and 72.5 kPCD range at the 120 minute time point;selecting a normal time point result if the breath test yields a measurement that is between a 49.3 and 74.5 kPCD range or sel... | Advanced Breath Diagnostics, LLC | A61B005/00 A61B005/08 A61B005/103 A61B005/117 A61K051/00 A61K049/00 A61M036/14
US08231530 | B2 | 2012-07-31 | Methods and systems for assessing gastric emptying | 2008-05-07 | Gastric emptying classification systems, displays, methods for classifying gastric emptying data, methods for treating patients, and methods for developing gastric emptying classification systems are disclosed. | 1. A method for developing a gastric emptying zone classification system, the method comprising: performing gastric emptying tests on a series of healthy patients;obtaining gastric emptying measurements for the series of healthy patients at a desired time point;using said gastric emptying measurements to form a reference interval for said desired time point;using an upper fractile or a lower fractile of the reference interval as a cutoff point for demarcating normal gastric emptying measurements from abnormal gastric emptying measurements;measuring biologic variation for the series of healthy patients at the desired time point to obtain biologic variation measurements;calculating a normal zone, a borderline zone, and an abnormal zone, the borderline zone being calculated around the cutoff point using the biologic variation measurements. 2. The method of claim 1, wherein said reference interval is a reference interval selected from the group consisting of 95% reference interval, a 90% reference interval, and an 85% reference interval. 3. The method of claim 1, wherein said upper fractile is an upper 2.5% fractile and said lower fractile is a lower 2.5% fractile. 4. The method of claim 1, wherein the using an upper fractile or a lower fractile of the reference interval as a cutoff point for demarcating normal gastric emptying measurements from abnormal gastric emptying measurements comprises using an upper fractile of the reference interval as an upper cutoff point for demarcating normal gastric emptying measurements from rapid gastric emptying measurements. 5. The method of claim 1, wherein the using an upper fractile or a lower fractile of the reference interval as a cutoff point for demarcating normal gastric emptying measurements from abnormal gastric emptying measurements comprises using a lower fractile of the reference interval as a lower cutoff point for demarcating normal gastric emptying measurements from delayed gastric emptying measurements. 6. The method of claim 1, wherein the measuring biologic variation for the series of healthy patients comprises: administering to healthy patients the same gastric emptying test on a first day and a second day to obtain first day gastric emptying measurements and second day gastric emptying measurements;subtracting the first day gastric emptying measurements from the second day gastric emptying measurements to obtain biologic variation measurements. 7. The method of claim 1, wherein the calculating a normal zone, a borderline zone, and an abnormal zone comprises: using the biologic variation measurements to calculate a standard deviation;using the standard deviation to calculate a borderline zone; anddesignating a zone above or below the borderline zone as the abnormal zone and an opposite zone as the normal zone. 8. The method of claim 1, wherein the abnormal zone is a rapid emptying zone, and the calculating a normal zone, a borderline zone, and a rapid emptying zone comprises: using the bi... | Advanced Breath Diagnostics, LLC | A61B005/00 A61B008/00 A61B005/08 A61K051/00
US07938846 | B2 | 2011-05-10 | Femoral compression device | 2007-06-22 | A femoral compression device (1) for compressive bearing against the femoral artery of a patient, comprising a base plate (2), an inflatable air cushion (4), and an electronic manometer (8) connected to the inflatable air cushion for measurement of the current pressure difference between the pressure prevailing inside the inflatable air cushion and the ambient air pressure. The femoral compression device can further comprise a mechanical device (11; 46) which prevents an air opening leading to the inflatable air cushion from being closed before the manometer is zeroed. In another embodiment, an electronic manometer is provided such that a current characteristic of an electric signal representing the current pressure difference can be compared with a corresponding characteristic which was obtained at zero pressure difference and which was stored in the electronic manometer, wherein the manometer cannot be zeroed as long as the current characteristic deviates more than a predetermined amount from the stored characteristic. | 1. A femoral compression device for compressive bearing against the femoral artery of a patient, comprising: a base plate;an inflatable air cushion attached to the base plate;a first extension extending from a first side of the base plate;a second extension extending from a second side of the base plate, said second side being opposite to the first side;a belt adapted to be attached to an end of the first extension and being adapted to extend around the patient's body to be attached to an end of the second extension;a pump connected to the inflatable air cushion for inflation of the inflatable air cushion;a manometer connected to the inflatable air cushion for measurement of a current pressure difference between pressure prevailing inside the inflatable air cushion and ambient air pressure, wherein the manometer is an electronic manometer; anda mechanical device comprising a closing portion, and a removable portion configured to prevent the manometer or circuits thereof from being electrically powered while preventing the closing portion from closing an air opening leading to the inflatable air cushion. 2. The femoral compression device according to claim 1, wherein the closing portion comprises a vent knob, by which the air opening leading to the inflatable air cushion can be closed or opened, wherein said vent knob at a time of manufacture of the femoral compression device is positioned such that the air opening is open, and wherein the removable portion is configured to prevent the vent knob from being maneuvered to a closed position before the manometer is zeroed. 3. The femoral compression device according to claim 2, wherein the removable portion comprises a removable cap, which is positioned over the vent knob and which is provided with a member that mechanically prevents the manometer from being zeroed by preventing the manometer or circuits thereof from being electrically powered as long as the removable cap has not been removed. 4. The femoral compression device according to claim 2, wherein the removable portion comprises a removable safety pin, which is positioned in engagement with the vent knob and which has a securing end portion that is configured to mechanically prevent the manometer from being zeroed by preventing the manometer, or circuits thereof, from being electrically powered as long as the removable safety pin has not been removed. 5. The femoral compression device according to claim 4, wherein the removable safety pin, as long as the removable safety pin is in place, is configured to ensure that the vent knob cannot be moved such that the air opening is closed. 6. The femoral compression device according to claim 1, wherein the electronic manometer is provided with electrical circuits such that a current characteristic of an electric signal representing the current pressure difference can be compared with a corresponding characteristic which, at a time of manufacture of the femoral compression device, was obtained at... | Radi Medical Systems AB | A61B017/00
US08166834 | B2 | 2012-05-01 | In-vitro measurement of catamenial tampon systems | 2007-06-21 | An apparatus for testing of medical products such as a feminine hygiene product, is presented. The apparatus includes a body, a pump and a vaginal canal assembly. The body includes an internal chamber and a bottom surface having a bore open to the internal chamber. The pump provides a fluid to the body. In one embodiment, the fluid is a menses simulant. The vaginal canal assembly includes an interior canal accepting the product. The vaginal canal assembly includes a passage providing the fluid to the interior canal. The apparatus includes a pressure regulator controlling pressure exerted on the vaginal canal assembly from a volume of air within the internal chamber. The apparatus includes a stand. The stand includes a retaining device and a locking device. The locking device cooperates with the retaining device to selectively secure the body in at least one of a rotational position and an angular position. | 1. An apparatus for in-vitro testing of a feminine hygiene product, comprising: a body having an internal chamber, the body including a bottom surface having a central bore open to the internal chamber;a pump coupled to the body, the pump providing a fluid of interest;a vaginal canal assembly disposed within the internal chamber, the vaginal canal assembly including: a wall defining an interior canal, the interior canal having an open end and a closed end, the open end of the interior canal being coaxial writhe central bore, the interior canal accepting the feminine hygiene product; and a passage in the wall for providing the fluid of interest to the interior canal;a bracket assembly disposed within the internal chamber, the bracket assembly having upright portions, the upright portions accepting the vaginal canal assembly and retaining the vaginal canal assembly at a predetermined orientation within the internal chamber; andwherein the bracket assembly further includes: a pin; and the upright portions include a plurality of opposing bores aligned along a vertical portion of the upright portions, the opposing bores accepting the pin wherein the pin is positioned within corresponding opposing bores for selectively adjusting the orientation of the vaginal canal assembly. 2. The apparatus of claim 1, wherein the vaginal canal assembly further includes a retaining plate portion, and wherein when the vaginal canal assembly is disposed in the body the retaining plate portion seals the central bore. 3. The apparatus of claim 2, further including: a pressure regulator coupling the body to an air supply; wherein the pressure regulator controls pressure exerted on the vaginal canal assembly from a volume of air within the internal chamber. 4. The apparatus of claim 2, further including: a retaining ring disposed about an external surface of the vaginal canal assembly, the retaining ring and retaining plate portion of the vaginal canal assembly cooperating to seal the central bore of the body. 5. The apparatus of claim 1, further including a stand supporting the body, the stand comprising: a retaining device for rotational and angular coupling of the body to the stand; and a locking device cooperating with the retaining device for selectively securing the body in at least one of a rotational position and an angular position. 6. The apparatus of claim 1, wherein the pump is comprised of a variable rate metering pump. 7. The apparatus of claim 6, further including: a control coupled to the pump, the control controlling delivery of the fluid for simulating at least one of intermittent and variable flow rates in the vaginal canal assembly. 8. The apparatus of claim 1, wherein the fluid is one of water, syngyna fluid and a menses simulant. 9. The apparatus of claim 1, wherein the vaginal canal assembly further includes a support portion accepting the pin of the bracket assembly, and wherein a location of the pin within the plurality of opposing bores adj... | Playtex Products, LLC | G01N015/00 G01N003/08 A61F013/20
US08831709 | B2 | 2014-09-09 | Method for 3-dimensional fluorescence tomographic imaging | 2005-09-26 | There is provided a method for determining the concentration of a fluorophore in a medium using moments of order k of the fluorescence signal. The method allows higher fidelity 3-dimensional reconstructions of the fluorophore in the medium. The method can be applied in imaging of fluorophores in biological tissues. | 1. A method for determining a concentration of a fluorophore in a medium as a function of three spatial coordinates of a three dimensional frame of reference, said method comprising: acquiring fluorescence-based measurements from three or more source-detector pairs;obtaining, for each source-detector pair, a moment of order k from fluorescence-based measurements acquired at that source-detector pair; anddetermining said concentration as a function of said three spatial coordinates based on said moments of order k and on a weighting coefficient for each source-detector pair. 2. The method as claimed in claim 1 wherein said fluorescence-based measurements are described by a photon diffusion equation and wherein said weighting coefficient is based on said photon diffusion equation. 3. The method as claimed in claim 2 wherein said photon diffusion equation is a normalized first order fluorescent Born approximation equation. 4. The method as claimed in claim 3 wherein said fluorescence-based photon diffusion equation is a time-domain photon diffusion equation and wherein said measurements define a temporal point spread function (TPSF). 5. The method as claimed in claim 1 wherein said moments are selected from 0, 1st, 2nd moments and combinations thereof. 6. The method as claimed in claim 5 wherein said moments are normalized moments. 7. The method as claimed in claim 1 wherein one of said three spatial coordinates represents depth relative to a surface of said medium. 8. The method as claimed in claim 1 wherein said medium is a biological tissue. 9. The method as claimed in claim 8 wherein said biological tissue is selected from brain tissue and breast tissue. 10. The method as claimed in claim 8 wherein said fluorophore is an endogenous fluorophore. 11. The method as claimed in claim 8 wherein said fluorophore is an exogenous fluorophore. | Softscan Healthcare Group Ltd. | A61B006/00 G01N021/47 G06T011/00 G01N021/64 A61B005/00
US09008743 | B2 | 2015-04-14 | Method and apparatus for providing data processing and control in medical communication system | 2008-04-14 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A method, comprising: receiving a signal associated with a monitored analyte level;retrieving a prior signal associated with an analyte monitoring device;analyzing the prior signal and the received signal based on a predetermined signal characteristic including a determination of the received signal below a predetermined level after a determination of the prior signal above a predetermined threshold;generating an output data based at least in part on the analysis of the prior signal and the received signal; andtransitioning from a sensor insertion state to a sensor removal state associated with the analyte monitoring device based at least in part upon a response to the output data. 2. The method of claim 1, wherein the predetermined signal characteristic includes a determination of a plurality of prior signals above the predetermined level. 3. The method of claim 2, wherein the predetermined level includes one of approximately 9 ADC counts or approximately 18 ADC counts. 4. The method of claim 2, wherein the predetermined signal characteristic includes a determination of the predetermined number of prior signals above the predetermined level for a second predetermined time period. 5. The method of claim 1, wherein the predetermined signal characteristic includes a signal transition from above the predetermined level to below the predetermined level. 6. The method of claim 1, wherein the output data includes a notification alert. 7. The method of claim 1, wherein the output data includes an indicator to start one or more processing timers associated with a respective one or more data processing routines. 8. The method of claim 7, wherein the one or more processing timers includes a respective one of a calibration timer, or a sensor expiration timer. 9. The method of claim 1, including presenting the output data. 10. The method of claim 9, wherein presenting the output data includes one or more of visually presenting the output data, audibly presenting the output data, vibratorily presenting the output data, or one or more combinations thereof. 11. The method of claim 1, wherein the analyte level includes a monitored glucose level. 12. The method of claim 1, wherein the predetermined signal characteristic includes a determination of a predetermined number of received signals below the predetermined level. 13. The method of claim 12, wherein the predetermined signal characteristic includes the determination of the predetermined number of received signals below the predetermined level for a first predetermined time period. 14. The method of claim 13, wherein the first predetermined period of time includes one of approximately 10 seconds, 30 seconds, less than 30 seconds, or greater than 30 seconds. 15. The method of claim 13 wherein the predetermined signal characteristic further includes a determination of a predetermined number of prior signals above the predetermined level for a second predetermined time period. 16. An apparatus,... | Abbott Diabetes Care Inc. | A61B005/05 A61B005/145 A61B005/1486 A61B005/00 G06F019/00 A61B005/1495
US10111608 | B2 | 2018-10-30 | Method and apparatus for providing data processing and control in medical communication system | 2008-04-14 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A method for notifying a user about an alarm condition without causing interruption to a predetermined routine, the method comprising: executing a concurrent passive notification routine on an analyte monitoring device, the concurrent passive notification routine comprising monitoring analyte levels received from a sensor for the alarm condition;executing the predetermined routine comprising one or more of performing a finger stick blood glucose test, calibrating a sensor unit, configuring a device setting, reviewing historical glucose data, reviewing historical alarms, events, or entries in a data log, managing a data communication setting, transferring data to a data processing terminal, or reviewing a higher priority alarm condition;in response to detecting the alarm condition by the concurrent passive notification routine during the execution of the predetermined routine, outputting to a user interface a first indication associated with the alarm condition, wherein the first indication is outputted during the execution of and without interruption to the predetermined routine; andin response to a termination or a completion of the predetermined routine, outputting to the user interface by the concurrent passive notification routine a second indication for the alarm condition associated with the first indication. 2. The method of claim 1, wherein the first indication includes one or more of a visual, audible, or vibratory indicator. 3. The method of claim 1, wherein the second indication includes one or more of a visual, audible, or vibratory indicator. 4. The method of claim 1, wherein the first indication includes a temporary indicator, and further, wherein the second indication includes a predetermined alarm associated with the detected alarm condition. 5. The method of claim 1, wherein the first indication is active during the execution of the predetermined routine, and is inactive at the end of the predetermined routine. 6. The method of claim 1, wherein the second indication is active at the end of the predetermined routine. 7. The method of claim 1, wherein each of the first indication and the second indication includes one or more of a visual text notification alert, a backlight indicator, a graphical notification, an audible notification, or a vibratory notification. 8. The method of claim 1, wherein the predetermined routine is executed to completion without interruption. 9. The method of claim 1, wherein when the alarm condition is detected during the execution of the predetermined routine, the first indication is output as a temporary alarm notification that does not disrupt the predetermined routine. 10. An apparatus for notifying a user about an alarm condition without causing interruption to a predetermined routine, the apparatus comprising: a user interface; anda data processing unit operatively coupled to the user interface, the data processing unit configured to: execute a concurrent passive notification routine on... | ABBOTT DIABETES CARE INC. | A61B005/00 G06F019/24 A61B005/145 A61B005/1486 A61B005/1495
US09278203 | B2 | 2016-03-08 | CT-enhanced fluoroscopy | 2008-03-26 | Real-time, enhanced imaging of remote areas, too minute for CT imaging, is made possible through a probe having a radiopaque tip as well as radiopaque volume markers. When deployed, the markers outline the space containing the tip such that both the tip and the volume containing the tip are viewable on a fluoroscope. This device may be used in conjunction with or independently of 3-D volumes created from CT scans and 3-D tip sensors. | 1. A method of navigating a device to an area of interest and imaging the area of interest in a branched structure comprising: using a first imaging modality to obtain at least one image of the branched structure;generating a recommended path inside the branched structure to the area of interest;imaging at least one tracing element projecting from a device in the branched structure using a second imaging modality, the at least one tracing element able to move with biological movement of the branched structure;deploying additional tracing elements to fluoroscopically illuminate locations of anatomical features by releasing a plurality of markers tethered to and radially extending from the device;tracing an actual three-dimensional path traveled by the device and the additional tracing elements by acquiring a plurality of images using the second imaging modality;registering at least one image from the first imaging modality to at least one image from the second imaging modality by comparing the actual three-dimensional path of the device with the recommended path;superimposing the at least one image from the first imaging modality and the at least one image from the second imaging modality; andnavigating the device through the branched structure to the area of interest by viewing a location of the device in a real-time image obtained with the second imaging modality of the tracing element relative to the at least one image from the first imaging modality. 2. The method of claim 1 wherein using a first imaging modality to obtain an image of the branched structure comprises using CT to obtain at least one image of the branched structure. 3. The method of claim 2 wherein using CT to obtain at least one image of the branched structure comprises using CT to obtain at least one image of the branched structure prior to navigating the device through the branched structure. 4. The method of claim 2 wherein using CT to obtain at least one image of the branched structure comprises using CT to obtain a plurality of images of the branched structure. 5. The method of claim 4 wherein using a first imaging modality to obtain at least one image of the branched structure further comprises assembling the plurality of images into a three dimensional model of the branched structure. 6. The method of claim 1 wherein imaging at least one tracing element attached to a device in the branched structure using a second imaging modality comprises imaging at least one radiopaque portion of the device. 7. The method of claim 6 wherein imaging at least one radiopaque portion of the device comprises imaging at least one radiopaque portion of the device using a fluoroscope. 8. The method of claim 1 wherein registering at least one image from the first imaging modality to at least one image from the second imaging modality comprises aligning anatomical features visible in the at least one image from the first imaging modality with corresponding anatomical features visible in ... | Covidien LP | A61B005/05 A61M031/00 A61B006/12 A61B006/00 A61M025/00 A61M025/01 A61M025/06
US10624572 | B2 | 2020-04-21 | Devices that monitor penetration of an instrument in an anatomical structure | 2005-02-11 | A device to monitor penetration of an instrument in an anatomical structure including at least two electrodes, a source of current supplying the at least two electrodes, and means for measuring impedance between the electrodes, wherein the electrodes are located on the penetration instrument, wherein the first electrode has a contact surface coinciding with a distal surface of the penetration instrument and the second electrode has a contact surface coinciding with a lateral surface of the penetration instrument, and wherein the contact surfaces are dimensioned to have a coinciding and constant contact surface as a function of a degree of penetration of the penetration instrument in the anatomical structure. | 1. A device that monitors penetration of a penetration instrument in an anatomical structure having at least two zones of different impedances comprising: a penetration instrument adapted to form a hole into said anatomical structure having said at least two zones of different impedances, the penetration instrument extending along a longitudinal axis between opposed proximal and distal ends. the penetration instrument having a surface including a cylindrical lateral surface extending from the proximal end, and a frustoconical distal surface at the distal end, the cylindrical lateral surface having a diameter measured transversally with respect to the longitudinal axis;at least two electrodes located on the penetration instrument, each of the at least two electrodes having a contact surface coinciding with the surface of the penetration instrument, wherein the at least two electrodes comprise at least one first electrode having the contact surface coinciding with the distal surface of the penetration instrument, and at least two second electrodes having the respective contact surfaces coinciding with the lateral surface of the penetration instrument, wherein the contact surface of each second electrode is annular, the second electrodes forming longitudinally spaced annular contacts;a source of current adapted to supply the first and second electrodes;an impedance measuring device adapted to measure the impedance between the first and second electrodes during penetration of the penetration instrument into said anatomical structure;a signalling device adapted to produce a signal at a time a variation in impedance is detected by the impedance measuring device when the penetration instrument passes from a first of said zones having a first impedance to a second of said zones having a second impedance to inform in real time a practitioner of the position of the penetration instrument;wherein the contact surfaces of the first and second electrodes are dimensioned to be constant as a function of a degree of penetration of the penetration instrument in the anatomical structure, said contact surfaces having a dimension measured from the proximal end to the distal end of the penetration instrument along the longitudinal axis that is smaller than the diameter of the cylindrical lateral surface. 2. The device according to claim 1, wherein the at least two electrodes comprise at least two first electrodes that are coaxially arranged and separated from each other by insulation. 3. The device according to claim 1, wherein the at least two electrodes comprise at least two first electrodes that are symmetrical with respect to the longitudinal axis of the penetration instrument. 4. The device according to claim 1, wherein the at least two electrodes further comprise a third electrode partially covering the lateral surface of the penetration instrument. 5. The device according to claim 1, wherein the signal is a sound signal, the signalling device being adapted ... | SpineGuard | A61B005/053 A61B005/00 A61B017/16 A61B090/00
US20080312519 | A1 | 2008-12-18 | Examination unit with an integrated mini-laboratory analysis unit | 2008-06-10 | An examination unit, which can be used especially in the area of acute and intensive diagnostics in hospitals (e.g. in emergency rooms, intensive care units, operating theaters etc.) or in ambulances is provided. The examination unit which at least has an ultrasonic device, patient monitor, ECG, ventilator and/or resuscitation device for monitoring the vital functions and/or for the immediate medical treatment of emergency patients with symptoms of acute cardiovascular disease, myocardial infarction, angina, apoplectic stroke, etc. The device furthermore according to the invention includes an integrated mini-laboratory analysis unit for the biochemical and/or cell-biology investigation of blood samples or tissue biopsies. | 1.-15. (canceled) 16. An examination unit, comprising: an integrated mini-laboratory analysis unit for a biochemical and/or cell-biology investigation of blood samples or tissue biopsies of a patient; andat least one device selected from the group consisting of patient monitor, ultrasonic device, ECG unit, ventilator device and resuscitation device. 17. The examination unit as claimed in claim 16, wherein the examination unit is configured as an emergency combi-set for at least one action selected from the group of producing a first diagnosis, monitoring a vital function, initiation of life-supporting, and immediate therapeutic measures for emergency patients with acute cardiovascular diseases, myocardial infarction, stroke or diabetes mellitus symptoms. 18. The examination unit as claimed in claim 17, wherein the emergency combi-set comprises a stationary component selected from the group consisting of “all-in-one” ultrasonic device, patient monitor, ECG, ventilator and resuscitation device. 19. The examination unit as claimed in claim 17, wherein the emergency combi-set comprises a portable component selected from the group consisting of “all-in-one” ultrasonic device, patient monitor, ECG, ventilator and resuscitation device. 20. The examination unit as claimed in claim 19, wherein the portable emergency combi-set is fitted with an integrated mobile power supply unit in the form of a battery, an accumulator, a fuel cell or a solar panel as a power-supply. 21. The examination unit as claimed in claim 16, wherein the integrated mini-laboratory analysis unit provides a capillary electrophoresis. 22. The examination unit as claimed in claim 16, wherein the integrated mini-laboratory analysis unit is suitably provides gene tests for the identification of high-risk patients with respect to a genetic predisposition to the occurrence of a disease. 23. The examination unit as claimed in claim 16, wherein the disease pertains to a cardiovascular disease, a myocardial infarction, angina, an apoplectic stroke, arteriosclerosis or diabetes mellitus. 24. The examination unit as claimed in one of claims 18, further comprises a permanently-integrated ultrasonic unit used only by the emergency combi-set itself. 25. The examination unit as claimed in one of claims 18, further comprises an ultrasonic unit which is integrated, but constructed as an independent module, which can be used both by the emergency combi-set itself or by external technical medical equipment. 26. The examination unit as claimed in one of claims 18, further comprises a connecting socket for plugging in a connecting cable to a connectable external ultrasonic transducer. 27. The examination unit as claimed in one of claims 18, further comprises a reception part for the wireless reception of ultrasonic signals from an external ultrasonic transducer via a wireless interface. 28. The examination unit as claimed in claim 27, wherein the wireless interface is an infrared interface. ... | SIEMENS AKTIENGESELLSCHAFT | A61B005/05 C12Q001/68
US20080306410 | A1 | 2008-12-11 | Methods and apparatuses for measuring pressure points | 2008-06-05 | Pressure sensing methods, systems, and computer program products for detecting and monitoring pressure in selectable areas of interest include a sensing system to determine in-sole foot pressure of a user in sports training and monitoring applications. | 1. A sensing system, comprising: a transducer to continually measure pressure of each of a plurality of points in an area of interest, the transducer including:a compressible layer, andfirst and second flexible conductive layers, between which the compressible layer is disposed;a transmitting/receiving device disposed proximate to the transducer to wirelessly transmit the measured data. 2. The system according to claim 1, wherein each first and second layer includes an electrode grid. 3. The system according to claim 2, further including: a selector to turn on and off selected points of the electrode grid to variably measure the pressure from the selected points of the area of interest. 4. The system according to claim 1, wherein said plurality of points of interest comprise a plurality of parts of a foot selected from the group consisting of a forefoot area, a midfoot area, and a hindfoot area. 5. The system according to claim 4, wherein said group further comprises one or more of a plurality of phalanges, one or more of a plurality of metatarsals, one or more of a plurality of phalangeal joints, a ball of said foot, one or more of a plurality of tarsal bones forming an arch of said foot, a plantar fascia, a talus, calcaneus, and a subtalar joint. 6. The system according to claim 3, wherein the selector turns on and off the selected points of the electrode grid dynamically in real-time. 7. The system according to claim 1, further including: a data compressor to compress the measured data before transmitting. 8. The system according to claim 1, wherein the transducer is embedded in a shoe sole. 9. The system according to claim 1, wherein the compressible material comprises a compressible conductive foam. 10. The system according to claim 9, wherein the compressible conductive foam comprises a material suitable for electrostatic discharge (ESD). 11. The system according to claim 1, further including: a host computer to wirelessly receive the transmitted data and output the received data in a user readable format. 12. The system according to claim 1, further comprising an electronic game coupled to receive the measured data and adapt said game accordingly. 13. The system according to claim 1, further comprising diagnostic means for interpreting the measured data and recommending changes to said pressure points. 14. The system according to claim 13, further comprising an orthotic to make said recommended changes. 15. The system according to claim 1, further comprising tracking means for interpreting the measured data and recommending changes to a training program. | 24/8 LLC | A61B005/103
US20080287747 | A1 | 2008-11-20 | Flexible Electronic Device | 2006-02-28 | According to the present invention there is provided a device that may be worn by a person, or fitted to an apparatus for monitoring, sensing or controlling one or more parameters such as, but by no means limited to, moisture, temperature, heart rate or respiration rate, or controlling various parameters of electronic functions. The device includes: a) a flexible substrate that can be worn by a person or fitted to an apparatus and, wherein the flexible substrate requires disposal or replacement with a fresh flexible substrate or requires maintenance that prevents continuous use of the device; andb) an electrical circuit having separate first and second parts, wherein the first part is carried by the flexible substrate and the first and second parts can be releasably connected together to form an electrical connection and disconnected to allow the parts to be reused or disposed as desired, and wherein one or both of the first and second parts includes any one or combination of a controller for controlling a parameter, a monitor for monitoring a parameter or a sensor for sensing a parameter. | 1. A device that may be worn by a person or fitted to an apparatus for monitoring, sensing or controlling one or more parameters such as, but by no means limited to, moisture, temperature, heart rate, respiration rate, strain, movement, light intensity, sound, pressure or impact forces, or an electrical function, wherein the device includes: a) a flexible substrate that can be worn by a person or attached to an apparatus and, wherein the flexible substrate requires disposal or replacement with a fresh flexible substrate or requires maintenance that prevents continuous use of the flexible substrate; andb) an electrical circuit having separate first and second parts and one or both of the first and second parts includes any one or a combination of a controller for controlling a parameter, a monitor for monitoring a parameter or a sensor for sensing a parameter, wherein the first part is carried by the flexible substrate and includes a first pair of contact surfaces and the second part includes a second pair of contact surfaces and the first and second pairs of contact surfaces can be connected together to form an electrical connection therebetween and thereafter disconnected to allow the parts to be reused or disposed as desired, and wherein an electrical conductive adhesive is present on the contact surfaces of either one or both of the first and second pairs and electrical current can be conducted therethrough when the first and second pairs are connected. 2. The device according to claim 1, wherein the first part of the circuit is directly attached to the flexible substrate. 3. The device according to claim 1, wherein the first part of the circuit is able to flex such that both the flexible substrate and the first part of the circuit is able to be flexed. 4. The device according to claim 1, wherein the first part of the circuit includes printed circuitry that is formed on or embedded in the flexible substrate. 5. The device according to claim 1, wherein the first part of the circuit includes an electrical conductive material that is on or embedded in the flexible substrate. 6. The device according to claim 5, wherein the conductive material is a filament or yarn that has been woven, knitted or incorporated in the flexible substrate by way of a layered structure. 7. The device according to claim 1, wherein the second part of the circuit is directly or indirectly connected to a second flexible substrate. 8. The device according to claim 1, wherein the second part of the circuit is housed by a rigid enclosure. 9. The device according to claim 1, wherein the second part of the circuit is carried or supported on another flexible substrate. 10. The device according to claim 1, wherein the first and second parts are able to be manually connected and disconnected. 11. The device according to claim 10, wherein the first and second parts are, in addition to the conductive adhesive, held together in electrical connection by any one or a combinati... | - | A61B005/00
US20080292046 | A1 | 2008-11-27 | Bronchopulmonary medical services system and imaging method | 2008-04-29 | A bronchopulmonary medical services system is provided in order to offer medical services to a patient in a single location. In one aspect, the system is provided with a patient alignment device and a C-arm imaging device. A medical services suite may be equipped with various medical service devices such that the patient receives medical services in a single location. | 1. A method for providing bronchopulmonary medical services to a patient at a single location of a medical services suite, comprising: positioning the patient via an alignment device;providing a medical service by producing an image of the positioned patient, the image produced via an X-ray imaging system having multi-axis movement in order to vary position of the patient image;determining from the previously provided medical service whether a subsequent medical service is to be provided to the patient;providing the subsequent medical service to the patient when the determination is to provide the subsequent medical service, the subsequent medical service becoming the previously provided medical service; andrepeating the determining and the providing of the subsequent medical service until a determination to not provide the subsequent medical service. 2. The method as claimed in claim 1, wherein the image is a reconstructed 3D image of a region of interest of the patient. 3. The method as claimed in claim 1, wherein the image is a reconstructed combination of 2d AND 3D image of a region of interest of the patient. 4. The method as claimed in claim 1, wherein the subsequent medical service is a subsequent image of the patient. 5. The method as claimed in claim 4, further comprising applying a contrast prior to producing the image or the subsequent image is produced 6. The method as claimed in claim 4, wherein the subsequent image selected from the group consisting of angiographic, intra-arterial angiography, bronchoscopic, morphology, and functional. 7. The method as claimed in claim 1, further comprising repositioning the patient in accordance with the subsequent medical service. 8. The method as claimed in claim 1, further comprising transferring the patient via the alignment device to the positioning the patient and transferring the patient from the alignment device after a determination to not provide the subsequent medical service. 9. The method as claimed in claim 8, wherein the transfer is via a robot. 10. The method as claimed in claim 1, wherein the subsequent medical service is a bronchoscopic procedure or a percutaneous procedure. 11. The method as claimed in claim 1, wherein the X-ray imaging device comprises a C-arm angiography device or a C-arm CT angiography device. 12. The method as claimed in claim 1, wherein an endoscopic bronchoscopy device is spatially combined to the X-ray imaging device. 13. The method as claimed in claim 12, wherein the subsequent medical service is selected from the group consisting of subsequent image, biopsy, cauterization, local-thermo, chemotherapy, bronchoalveolar lavage, stent insertion, removal of an object lodged in an airway of the patient. 14. The method as claimed in claim 12, further comprising: recording the X-ray image;spatially combining the endoscopic bronchoscopy device to the X-ray imaging arrangement; andrecording an bronchoscopic image of the patient via a navigation. 15. T... | - | A61B006/00 A61B005/05
US20080255438 | A1 | 2008-10-16 | SYSTEM FOR MONITORING PHYSIOLOGICAL CHARACTERISTICS | 2008-06-25 | Apparatuses and methods for medical monitoring physiological characteristic values such as blood glucose levels for the treatment of diabetes, are presented. The apparatuses and methods provide for preventing any negative consequence in the operation of a monitor and/or infusion device as a result of disorientation that may occur from waking from slumber with a low blood glucose level. In addition, a graphical display is disclosed incorporating a variety of enhancements which readily conveys to the user historical as well as real time information regarding the measured characteristic value. | 1. A device, comprising: a sensor input capable of receiving a signal from a sensor, the signal being based on a sensed physiological characteristic value of a user;a memory for storing a plurality of measurements of the sensed physiological characteristic value of the user from the received signal from the sensor; anda display for presenting a graphical representation of the plurality of measurements of the sensed physiological characteristic value, wherein the graphical representation comprises one or more trend indicators, indicating an approximate rate trend in the sensed physiological characteristic value over a recent series of the plurality of measurements. 2. The device of claim 1, wherein the graphical representation displays real time measurements of the sensed physiological characteristic value. 3. The device of claim 1, wherein the received signal from the sensor is transmitted via RF telemetry. 4. The device of claim 1, wherein the device is an infusion pump. 5. The device of claim 1, wherein the physiological characteristic value is a blood glucose measurement. 6. The device of claim 1, wherein the physiological characteristic value measurement is displayed in a graphical format or a text format or a combination thereof. 7. The device of claim 1, wherein the graphical representation comprises a color coded set of indicators. 8. The device of claim 1, wherein the graphical representation comprises a vertical cursor which allows a user to alternatively select measurements of the sensed physiological characteristic value from the graphical format or from the text format. 9. The device of claim 1, wherein the recent series of the plurality of measurements comprises four most recent measurements. 10. The device of claim 1, wherein the one or more trend indicators includes an indicator of a current measurement. 11. The device of claim 1, wherein the one or more trend indicators comprise a color coded set of indicators. 12. The device of claim 11, wherein the physiological characteristic value is a blood glucose measurement and the color coded set of indicators comprises a grid of nine indicators including three rate trend categories including going down, stable and going up and three current measurement categories including low, in zone and high where a single rate trend category and a single current measurement category are indicated simultaneously. 13. The device of claim 12, wherein going down and low and going up and high are indicated in red; going down and high and going up and low are indicated in yellow;going down and in zone, going up and in zone, stable and high and stable and low are indicated in orange; andstable and in zone is indicated in green. 14. The device of claim 1, wherein the one or more trend indicators comprise five symbols indicating a fast increasing rate, a moderate increasing rate, a negligible rate change, a moderate decreasing rate and a fast decreasing rate, respectively. 15. The device of clai... | MEDTRONIC MINIMED, INC. | A61B005/145
US20080287771 | A1 | 2008-11-20 | SURGICAL NAVIGATION SYSTEM WITH ELECTROSTATIC SHIELD | 2007-10-26 | A surgical navigation system is disclosed herein. The surgical navigation system includes a tracking system, a field generator operatively connected to the tracking system, a field sensor operatively connected to the tracking system, and an electrostatic shield circumscribing the field sensor. The electrostatic shield is adapted minimize capacitive coupling between the field generator and the field sensor. | 1. A surgical navigation system comprising: a tracking system;a field generator operatively connected to the tracking system;a field sensor operatively connected to the tracking system; andan electrostatic shield circumscribing the field sensor, wherein the electrostatic shield is adapted minimize capacitive coupling between the field generator and the field sensor. 2. The surgical navigation system of claim 1, further comprising a structural portion of a housing assembly adapted to retain the field sensor. 3. The surgical navigation system of claim 2, wherein the electrostatic shield comprises a coating applied to the structural portion of the housing assembly. 4. The surgical navigation system of claim 2, wherein the electrostatic shield comprises a paint applied to the structural portion of the housing assembly. 5. The surgical navigation system of claim 1, wherein the electrostatic shield is sufficiently resistive to prevent the formation of eddy currents caused by the interaction of the electrostatic shield with a magnetic field produced by the field generator. 6. The surgical navigation system of claim 1, wherein the electrostatic shield comprises a material having a resistance in the range of one to ten ohms as measured at a distance of one centimeter. 7. The surgical navigation system of claim 1, wherein the field sensor is disposed within the electrostatic shield such that the electrostatic shield completely surrounds the field sensor. 8. A surgical navigation system comprising: a computer;an imaging device connected to the computer;a display connected to the computer;a tracking system connected to the computer, said tracking system adapted to estimate a position and/or orientation of a medical instrument;a field generator connected to the tracking system;a field sensor connected to the tracking system; andan electrostatic shield completely surrounding the field sensor, wherein the electrostatic shield is adapted minimize capacitive coupling between the field generator and the field sensor. 9. The surgical navigation system of claim 8, wherein the imaging device comprises one of a CT imaging device; a MR imaging device; a PET imaging device; an ultrasound imaging device; and an X-ray imaging device. 10. The surgical navigation system of claim 8, wherein the display is configured to graphically convey the estimated position and/or orientation of the medical instrument. 11. The surgical navigation system of claim 8, further comprising a structural portion of a housing assembly adapted to retain the field sensor. 12. The surgical navigation system of claim 11, wherein the electrostatic shield comprises a coating applied to the structural portion of the housing assembly. 13. The surgical navigation system of claim 8, wherein the electrostatic shield is sufficiently resistive to prevent the formation of eddy currents caused by the interaction of the electrostatic shield with a magnetic field produced by the field generator. 14. A... | GENERAL ELECTRIC COMPANY | A61B006/00 A61B005/055
US20080319271 | A1 | 2008-12-25 | PERSONALIZED HEALTH MANAGEMENT TOOL | 2007-06-19 | A personalized health management tool provides personalized health management information to one or more patients or members of a health plan for use both by the patient and the patient's healthcare provider(s) to improve the quality and efficiency of interactions between the patient and the provider(s). The personalized health management tool includes information relating to a patient's medical conditions and may include integrated guidelines that provide guidance on management and treatment of co-morbidities of the patient. The tool further includes information relating to the self-management of the patient's condition(s), information relating to the medical history of a patient, and/or information relating to the recommended care and treatment required for the patient's condition(s). For example, the health management tool may include EBM guidelines regarding the medical condition(s), integrated guidelines for multiple condition management, and/or information relating to resources related to the patient's medical condition(s) that are available to the patient for assistance with healthcare, financial issues, and/or social issues. A method for providing personalized health management information to one or more patients or members of a health plan includes identifying patients or members having one or more specific medical conditions; gathering personalized health data for each identified patient or member; and providing each identified patient or member with a personalized health management tool including information on the specific medical conditions, self-management of the specific medical conditions, treatment of the medical conditions for use by the patient's healthcare providers, and the personalized health data gathered for the respective identified patient or member. | 1. A personalized health management tool, comprising: information relating to at least one medical condition of a patient;information relating to self-management of the at least one medical condition; andmedical history information for the patient. 2. The personalized health management tool of claim 1, wherein the information relating to the at least one medical condition of a patient comprises evidence-based medicine guidelines. 3. The personalized health management tool of claim 1, further including information relating to the care for the at least one medical condition for the patient. 4. The personalized health management tool of claim 1, wherein the information relating to at least one medical condition of a patient comprises integrated guidelines relating to more than one condition of the patient. 5. The personalized health management tool of claim 1, further including information relating to available resources on healthcare assistance, financial assistance, or social issues assistance. 6. The personalized health management tool of claim 1, further including health insurance information for the patient. 7. The personalized health management tool of claim 1, further including health care provider information for the patient. 8. The personalized health management tool of claim 1, wherein the information relating to the medical history for the patient comprises information gathered from at least one of the patient, the physician or other care provider, and insurance claim data. 9. The personalized health management tool of claim 1, wherein the at least one medical condition comprises diabetes. 10. The personalized health management tool of claim 9, wherein the at least one medical condition farther comprises congestive heart failure. 11. The personalized health management tool of claim 1, wherein the information relating to the self-management of the at least one medical condition comprises tools for tracking patient progress. 12. The personalized health management tool of claim 1, further including a self-monitoring tool. 13. The personalized health management tool of claim 1, wherein the information relating to the at least one medical condition of a patient comprises at least one of background information relating to the at least one medical condition, treatments relating to the at least one medical condition, and information relating to daily living with the at least one medical condition. 14. A method for providing personalized health management information to one or more members of a health plan comprising: identifying one or more members of a health plan having a specific medical condition;gathering personalized health data for each identified member; andproviding each identified member with a personalized health management tool including information on the specific medical condition, information on self-management of the specific medical condition, and the personalized health data gathered for the respective identified me... | - | A61B005/00
US20080242976 | A1 | 2008-10-02 | Electric field tomography | 2007-03-30 | Methods for evaluating motion of a tissue, such as of a cardiac location, e.g., heart wall, via continuous field tomography are provided. In the subject methods, a continuous field (e.g., an electrical, mechanical, electromechanical, or other field) sensing element is stably associated with the tissue location. A property of the applied continuous field is detected with the sensing element to evaluate movement of the tissue location. Also provided are systems, devices and related compositions for practicing the subject methods. The subject methods and devices find use in a variety of different applications, including cardiac resynchronization therapy. | 1. A method for evaluating movement of a tissue location in a subject, said method comprising: (a) generating a continuous electric field so that said tissue location is present in said continuous field; and(b) detecting a change in property of the continuous electric field at said tissue location to evaluate movement of said tissue location. 2. The method according to claim 1, wherein said evaluating comprises converting the detected change in property to a measurement of distance, location, or motion of the tissue location relative to a second location. 3. The method according to claim 2, wherein said movement is evaluated by calculating a motion between said tissue location and a second location. 4. The method according to claim 3, wherein said continuous field is generated from said second location. 5. The method according to claim 1, wherein said detecting comprises obtaining a signal from a sensing element stably associated with said tissue location, wherein said signal is induced in said sensing element by movement of said tissue location in said continuous field. 6. The method according to claim 1, wherein said detecting comprises determining a value for said property at least twice over a duration of time to evaluate movement of said tissue location. 7-9. (canceled) 10. The method according to claim 1, wherein said electric field is an oscillating electrical conduction current field. 11-15. (canceled) 16. The method according to claim 1, wherein said continuous electric field is generated between a source and at least one sensing element. 17. The method according to claim 1, wherein said continuous electric field is generated between a source and a ground, and said change in property is detecting by at least one sensing element that is not said ground. 18. The method according to claim 1, wherein said property is chosen from amplitude, phase and frequency. 19. The method according to claim 18, wherein said property is amplitude. 20. The method according to claim 19, wherein said detecting comprises detecting amplitude signals having the same phase and frequency. 21. The method according to claim 18, wherein said property is frequency. 22. The method according to claim 21, wherein said evaluating comprises determining velocity based on frequency. 23. The method according to claim 5, wherein said sensing element comprises at least one electrode. 24. The method according to claim 23, wherein said sensing element comprises two or more closely spaced electrodes. 25. The method according to claim 24, wherein said detecting comprises (a) measuring a local gradient of the electric field between the closely spaced electrodes; and(b) measuring a change in the value of the field. 26. The method according to claim 25, wherein said evaluating comprises calculating a location or motion of said tissue location based on both the measured gradient and the measured change of the value. 27. The method according to claim 1, wherein said ... | Proteus Biomedical, Inc. | A61B005/04
US20080319284 | A1 | 2008-12-25 | METHOD FOR TRAINING A USER IN RECOGNITION OF THE USER'S BODILY FLUID ANALYTE CONCENTRATION AND CONCENTRATION TRENDS VIA USER-PERCEIVED SENSATIONS | 2007-06-25 | A method for training a user in recognition of the user's bodily fluid analyte concentration and concentration trends includes monitoring a user's bodily fluid analyte concentration (for example, a user's whole blood glucose concentration) with an analyte monitoring module of a device for training a user in recognition of the user's bodily fluid analyte concentration. The method also includes generating, at a predetermined time interval, using a generation module of the device, a user-perceived sensation (e.g., a user-perceived vibratory sensation) that is proportional the user's bodily fluid analyte concentration in a predetermined manner. | 1. A method of training a user in recognition of the user's bodily fluid analyte concentration and concentration trends comprising: monitoring a user's bodily fluid analyte concentration with an analyte monitoring module of a device for training a user in recognition of the user's bodily fluid analyte concentration; andgenerating a user-perceived sensation, at predetermined time interval, that is proportional the user's bodily fluid analyte concentration in a predetermined manner using a sensation generation module of the device. 2. The method of claim 1 wherein the generating step generates a user-perceived vibratory sensation. 3. The method of claim 2 wherein a duration of the user-perceived vibratory sensation is proportional to the user's bodily fluid analyte concentration in predetermined proportional manner. 4. The method of claim 2 wherein an amplitude of the user-perceived vibratory sensation is proportional to the user's bodily fluid analyte concentration in predetermined proportional manner. 5. The method of claim 2 wherein the user-perceived vibratory sensation is proportional to the user's bodily fluid analyte concentration in predetermined proportional manner based on a count of user-perceived vibratory sensations. 6. The method of claim 1 wherein the user-perceived sensation is proportional to the user's bodily fluid analyte concentration in a predetermined step-wise manner. 7. The method of claim 1 wherein the monitoring step monitors a user's bodily fluid glucose concentration. 8. The method of claim 1 the generating step also generates a pre-alert user-perceived sensation prior to generating the user-perceived sensation. | LifeScan Scotland, Ltd. | A61B005/00
US20080279776 | A1 | 2008-11-13 | Photosensitizers and MRI Enhancers | 2005-04-07 | The present invention relates to the use of a compound of formula 3 or a salt thereof for the manufacture of a medicament or phototherapeutic agent for the treatment of acne, Aids, viral hepatitis, diabetic retinopathy, infection with sars virus, coronary artery stenosis, carotid artery stenosis, intermittent claudication, Asian (chicken) flu virus, cervical dysplasia or cancer of the blood, cervix, naso-pharynx, trachea, larynx, bronchi, bronchioles, bladder, esophagus, stomach, rectum, colon, prostate, hollow organs, bile duct, ureter, kidney, uterus, vaginal or other female adnexa. The invention also relates to methods of treating these diseases.;The present invention further relates to the use of a compound of formula 3 or a salt thereof for the manufacture of a photodiagnostic agent for the detection of the above diseases, as well as atherosclerosis, multiple sclerosis, diabetes, arthritis, rheumatoid arthritis, a fungal, viral, chlamydial, bacterial, nanobacterial or parasitic infectious disease, HIV, hepatitis, herpes simplex, herpes zoster, psoriasis, a cardiovascular disease, or a dermatological condition. The invention also relates to methods of detecting these diseases by photodiagnosis.;The present invention further relates to a method of cold sterilising a surgical or other device, comprising the steps of: providing a compound of formula 3 or a salt thereof on the device and subjecting the device to irradiation or sound.;The present invention further relates to a compound of formula 3 or a salt thereof, linked or attached to a magnetic element. Such a compound may be used as an MRI enhancer. The present invention also relates to a method of carrying out an MRI scan using such an MRI enhancer. | 1. Use of a compound for the manufacture of a medicament for the treatment of acne, Aids, viral hepatitis, diabetic retinopathy, infection with sars virus, coronary artery stenosis, carotid artery stenosis, intermittent claudication, Asian (chicken) flu virus, cervical dysplasia or cancer of the blood, cervix, naso-pharynx, trachea, larynx, bronchi, bronchioles, bladder, esophagus, stomach, rectum, colon, prostate, hollow organs, bile duct, ureter, kidney, uterus, vaginal or other female adnexa, wherein the compound is of formula 3wherein M is a metal atom in the M(II) oxidation state, a metal halide, a metal oxide or a silicon with two substituents,each R1, R2, R3, R4, R5, R6, R7, R8, R9, R10, R11, R12, R13 and R14 is independently hydrogen, (CH2)n—CHO, (CH2)n—CO2R15 or a C1-C6 saturated or unsaturated alkyl group optionally substituted with one or more of —OH and —NH2,n is 0, 1, 2 or 3, andeach R15 is independently hydrogen, lithium, sodium, potassium, magnesium, calcium, a C1-C6 saturated or unsaturated alkyl group optionally substituted with one or more of —OH and —NH2, or a naturally occurring amino acid. 2. Use of a compound for the manufacture of a phototherapeutic agent for the use in photodynamic therapy or cytoluminescent therapy for the treatment of acne, Aids, viral hepatitis, diabetic retinopathy, infection with sars virus, coronary artery stenosis, carotid artery stenosis, intermittent claudication, Asian (chicken) flu virus, cervical dysplasia or cancer of the blood, cervix, naso-pharynx, trachea, larynx, bronchi, bronchioles, bladder, esophagus, stomach, rectum, colon, prostate, hollow organs, bile duct, ureter, kidney, uterus, vaginal or other female adnexa, wherein the compound is of formula 3wherein M is a metal atom in the M(II) oxidation state, a metal halide, a metal oxide or a silicon with two substituents,each R1, R2, R3, R4, R5, R6, R7, R8, R9, R10, R11, R12, R13 and R14 is independently hydrogen, (CH2)n—CHO, (CH2)—CO2R15 or a C1-C6 saturated or unsaturated alkyl group optionally substituted with one or more of —OH and —NH2,n is 0, 1, 2 or 3, andeach R15 is independently hydrogen, lithium, sodium, potassium, magnesium, calcium, a C1-C6 saturated or unsaturated alkyl group optionally substituted with one or more of —OH and —NH2, or a naturally occurring amino acid. 3. Use of a compound for the manufacture of a photodiagnostic agent for the detection of acne, Aids, viral hepatitis, diabetic retinopathy, infection with sars virus, coronary artery stenosis, carotid artery stenosis, intermittent claudication, Asian (chicken) flu virus, cervical dysplasia, or cancer of the blood, cervix, naso-pharynx, trachea, larynx, bronchi, bronchioles, bladder, esophagus, stomach, rectum, colon, prostate, hollow organs, bile duct, ureter, kidney, uterus, vaginal or other female adnexa, or atherosclerosis, multiple sclerosis, diabetes, arthritis, rheumatoid arthritis, a fungal, viral, chlamydial, bacterial, nanobacterial or parasitic infec... | Photo Diagnostic Devices (PDD) Limited | A61K049/06 A61K031/555 A61K049/00 A61B005/055 C07D209/56 A61L002/18
US20080249537 | A1 | 2008-10-09 | Skin Resurfacing Device | 2004-08-13 | A skin resurfacing device for peeling the outermost layer of the skin for renewing the skin surface and repairing skin damage including a housing, a skin sensor, and a skin treater. The skin treater consists of an abrasive tip, a first end, a transparent portion, a first filter, and a second end, all of which are detachable and replaceable. The skin treater is connected to a vacuum source by a tubular hose. The vacuum provides both closeness of contact between the abrasive tip and the user's skin and the suctioning of skin debris peeled off. The skin treater, especially the abrasive tip is made with common material and mass production process so that they are disposable and economic. | 1. A skin resurfacing device comprising a) a skin treater, andb) a vacuum source connected to the skin treater,wherein the skin treater comprises a hollow tube having a first end, a second end, an abrasive tip detachably fixed on the first end, a first filter that is provided inside the tube between the first end and the second end, wherein the abrasive tip comprises abrasive particles, and wherein the vacuum source is connected to the second end. 2. The skin resurfacing device of claim 1 wherein the abrasive particles of the abrasive tip comprises aluminum oxide crystals, silicon carbide crystals or silicon oxide crystals having a predetermined range of size. 3. The skin resurfacing device of claim 2 wherein the predetermined range of the abrasive particles size is from about sixty (60) mm to about one hundred fifty (150) mm. 4. The skin resurfacing device of claim 2 wherein the abrasive tip is made by pressure molding and heat treating abrasive particles. 5. The skin resurfacing device of claim 2 wherein the abrasive tip has an flat annular portion that contacts the skin of a user, and wherein a suction hole is provided in the annular portion through which air is sucked. 6. The skin resurfacing device of claim 2 wherein the abrasive tip is coated with liquid ceramic material. 7. The skin resurfacing device of claim 6 wherein the liquid ceramic material is colored with a different color for a different size of abrasive crystals. 8. The skin resurfacing device of claim 1 wherein the abrasive tip comprises a roller that protrudes from the flat annular portion so that the roller contacts and rolls on the skin of a user. 9. The skin resurfacing device of claim 1 further comprising a skin sensor that measures the oiliness of the skin of a user. 10. The skin resurfacing device of claim 9 wherein the intensity of the vacuum provided by the vacuum source is controlled according to the measured oiliness by the skin sensor. 11. The skin resurfacing device of claim 10 wherein a user can adjust the intensity of the vacuum. 12. The skin resurfacing device of claim 1 further comprising a timer that controls operation time of the device. 13. The skin resurfacing device of claim 1 wherein the tube has a transparent portion so that the filter is visible outside. 14. The skin resurfacing device of claim 1 further comprising a second filter between the skin treater and the vacuum source. 15. The skin resurfacing device of claim 14 wherein the second filter comprises a container that comprises an open end, a lid that plugs the open end, an inlet pipe passing through the lid, an outlet pipe passing through the lid, a filter element that is fixed to the outlet pipe, wherein the container is detachable from the lid. 16. The skin resurfacing device of claim 1 wherein the abrasive tip comprises a generally cylindrical housing having an open end and a closed end, an abrasive member having a first surface and a second surface, and a rod member extending fro... | - | A61B017/50 A61B005/00
US07460897 | B1 | 2008-12-02 | Patient interface for spectroscopy applications | 2007-04-05 | A patient interface and a method of installing a patient interface for a spectroscopy kit on a patient. The patient interface includes a butterfly-shaped patient interface with a cap on top of the patient interface and in proximity to an optical head. | 1. A patient interface for a tissue measurement instrument, the patient interface comprising: a flexible base member having a central region with a longitudinal axis, the base member having an outer perimeter defining a distal concave locating feature, the outer perimeter further defining a pair of opposing wings extending from the central region of the base member, each wing having a convex outer perimeter, the wings generally symmetric about the longitudinal axis, the longitudinal axis substantially bisecting the locating feature;wherein the base member has two openings for light transmission therethrough, the openings lying generally along the longitudinal axis and defining a sensor region of the central region;wherein the wings extend longitudinally along the outer perimeter of the base member adjacent the sensor region; andan opaque cap positioned on the central region of the base member and in proximity to an optical head. 2. The patient interface of claim 1 further comprising an adhesive coating on an underside of the patient interface. 3. The patient interface of claim 2 further comprising a bifurcated release liner covering the adhesive coating. 4. The patient interface of claim 3 wherein the release liner covering is made of a spectrally flat material that minimizes light reflection. 5. The patient interface of claim 1 further comprising a connector and cable. 6. The patient interface of claim 1 wherein the base member and wings comprise materials that breathe, transmit moisture vapor, block ambient light or a combination thereof. 7. The patient interface of claim 1 wherein the cap is a foam cap. 8. The patient interface of claim 7 wherein the cap is removable. 9. The patient interface of claim 1 wherein the base and wings are removable. 10. The patient interface of claim 1 wherein base member is made of a spectrally flat material that minimizes light reflection. 11. A method of installing a patient interface for a tissue measurement instrument comprising the steps of: providing a flexible base member having a central region with a longitudinal axis, the base member having an outer perimeter defining a distal concave locating feature, the outer perimeter further defining a pair of opposing wings extending from the central region of the base member, the wings generally symmetric about the longitudinal axis, the base member having one or more openings for light transmission therethrough, the one or more openings lying generally along the longitudinal axis and defining a sensor portion of the central region, wherein the wings extend longitudinally along the outer perimeter of the base member adjacent the sensor portion; and a cap positioned on the central region of the base member and surrounding an optical head;connecting one end of a cable to the optical head and the other end of the cable to a monitor; andaffixing the patient interface onto a patient. 12. The method of claim 11 further comprising applying on skin of the patie... | Hutchinson Technology Incorporated | A61B005/1455
US07462347 | B2 | 2008-12-09 | Fluorescent membrane intercalating probes | 2006-02-03 | The invention relates to a family of cyanine dyes which fluoresce in the far red and near infra red wavelengths of the spectrum and preferably possess lipophilic side chains. The dyes of the invention are soluble in commercially available membrane staining vehicles, are useful as probes for rapidly staining lipophilic structures such as membranes in cells or isolated from cells, and are well retained therein. Methods of using the dyes to detect stained cells both in vivo and in vitro are also disclosed. | 1. An in vivo method for assessing the shedding rates of mature surface epithelial cells of a warm blooded animal comprising the steps of: labeling mature surface epithelial cells at a target site of the animal with a cyanine dye, and monitoring the site for the presence or absence of the label following said labeling step, wherein the cyanine dye is a compound having the structural formula:wherein Z is a biologically compatible counterion. 2. A compound having the structural formulawherein Z is a biologically compatible counterion. 3. An in vivo method for assessing the shedding rates of mature surface epithelial cells of a warm-blooded animal, comprising the steps of: labeling mature surface epithelial cells at a target site of the animal with a cyanine dye;exciting the target site with a light source; andmonitoring the site for the presence of absence of fluorescence resulting therefrom;wherein the cyanine dye is a compound having the structural formula:wherein Z is a biologically active counterion. 4. The method of claim 3 wherein the site is a mucosal surface. 5. The method of claim 4 wherein the mucosal surface lines the surface of the gastrointestinal tract, the respiratory tract, or the genitourinary tract. 6. The method of claim 4 wherein the site is a mucosal surface of the stomach. 7. The method of claim 4 wherein the site is a mucosal surface of the colon. 8. The method of claim 3 wherein cell shedding rates are detected by observing changes in the level of the label at the site at a pre-selected time following said labeling step. 9. The method of claim 3 wherein the excitation light has a wavelength of from about 600 nm to about 900 nm. 10. The method of claim 3 wherein the epithelial cells are labeled by direct application of a labeling composition to the site. 11. A method for labeling a cell, comprising the steps of: contacting the cell with a compound having the structural formula:wherein Z is a biologically compatible counterion. 12. The method according to claim 11 wherein said labeling composition comprises a pharmaceutically acceptable vehicle. | Phanos Technologies, Inc. | A61B010/00 A61B005/00 A61B008/00
US07460903 | B2 | 2008-12-02 | Method and system for a real time adaptive system for effecting changes in cognitive-emotive profiles | 2003-09-12 | A means and method for inducing a temporary physiological state-of-mind to effect persistent changes to the cognitive-emotive profile of an individual, which is adaptable for neurofeedback and “mental-state” therapeutic and non-therapeutic interventions. The system comprises an EEG Recording Module (ERM), a Neurodynamics Assessment Module (“NAM”), and a Transcranial Magnetic Stimulation module (“TMS”) for acquiring and manipulating bioelectrical and/or EEG data, defining a cognitive-emotive profile, and mapping the cognitive-emotive profile to selectively control transcranial magnetic stimulation to drive therapeutic and non-therapeutic stimulus interventions. A bi-directional feedback feature is provided to further enhance the performance of the system to effect prolonged changes. | 1. A method of adaptive intervention for effecting persistent changes in the cognitive-emotive profile of an individual, comprising the steps of: selectively acquiring a plurality of bioelectric signals of the individual to determine current psychological state;comparing current psychological state to a set of templates or indices to extract a multi-dimensional cognitive-emotive profile based on the bioelectric signals;mapping the cognitive-emotive profile onto a set of commands;controllably delivering brain stimulation commands to the individual to drive therapeutic and non-therapeutic stimulus intervention; andeffecting a prolonged change in the individual's cognitive-emotive profile which persist over long timeframes and outlast stimulation. 2. The method of claim 1, wherein the effected by transcranial magnetic stimulation (TMS) which results in behavioral changes. 3. The method of claim 2, wherein the TMS signal can be delivered at one or more sites of the individual's body simultaneously and the effects of the TMS on the behavioral changes outlast the short time of actual stimulation and persist over longer time frames. 4. The method of claim 1, wherein the bioelectric signal is an electroencephalogram (EEG) signal. 5. The method of claim 4, wherein the EEG signal is recorded from multiple recording sites from the scalp of the individual using a portable headset. 6. The method of claim 5, wherein the portable headset includes a matrix of EEG sensors and magnetic field coils oriented over specific areas of the brain of the individual. 7. The method of claim 4, further comprising the decomposition of the EEG signal into a plurality of signal subcomponents including: frequency domain subcomponents;time domain subcomponents; and/or spatial domain subcomponents. 8. The method of claim 7, wherein the frequency domain subcomponents are selected from a group consisting of a mu rhythm, a theta rhythm, an alpha rhythm, and a beta rhythm. 9. The method of claim 7, wherein the time domain subcomponents are selected from a group consisting of event-related potentials (ERPs) including N1, P3, and steady state visual evoked response (SSVER). 10. The method of claim 7, wherein the spatial domain subcomponents are selected from a group derived from special algorithmic transformation of the EEG signal. 11. The method of claim 7, wherein the frequency and time domain subcomponents are analyzed using one of a group of signal processing algorithms consisting of a variable epoch frequency decomposition (VEFD), a fast Fourier transform (FFT), and independent component analysis (ICA). 12. The method of claim 7, further comprising identifying and classifying feature clusters from the plurality of signal subcomponents. 13. The method of claim 12, further comprising creating a brain/computer interface feature map (BFM) from a feature cluster identified through one of a group of transformation algorithms consisting of: a discriminant optimization analysis;... | - | A61B005/04
US07460901 | B2 | 2008-12-02 | Procedure for deriving reliable information on respiratory activity from heart period measurement | 2003-05-30 | The present invention describes a method of tracking minute ventilation from R-R intervals that are based on the acquisition of ECG signal. Sub-methods (RDF1), which are mainly based on heart rate variability, are used, wherein the respiratory frequency is determined from a pattern of rhythmic changes in heart beat data. The sub-methods (RDF1-n) are combined using an expert function, wherein each sub-method (RDF1-n) determines an estimate of the respiratory frequency. | 1. A method for deriving reliable information on respiratory activity from a pattern of rhythmic changes in heart beat data, the method comprising the steps of: measuring an electrocardiographic (ECG)-signal;obtaining RR-interval data from the ECG-signal;calculating respiratory frequency estimates from the RR-interval data using two or more sub-methods, at least one sub-method being based on heart rate variability;combining the calculated respiratory frequency estimates in a health apparatus using an expert function including a weight vector across frequencies for each frequency estimate and obtaining a resulting respiratory frequency estimate directly in the combination made by said expert function. 2. Method according to claim 1, characterized in that one sub-method is optimal for a steady condition. 3. Method according to claim 1, characterized in that at least one sub-method comprises determination of temporal changes in the respiratory frequencies by defining temporal deviations in the distribution of frequencies. 4. Method according to claim 1, characterized in that post-correction of the respiratory period is made with a chosen method. 5. Method according to claim 1, characterized in that an expert function comprises a reliability weightings for each sub-method as a function of the calculated respiratory frequency by the respective sub-method. 6. Method according to claim 1, characterized in that at least one sub-method is controlled dynamically according to the heart rate level. 7. Method according to claim 1, characterized in that the expert function makes a pre-selection from several different sub-methods with pre-selected criteria. 8. Method according to claim 7, characterized in that the expert function makes a fixed pre-selection for an implementation according to pre-measured empirical data. 9. Method according to claim 1, characterized in that information on tidal volume is calculated using information on RR-interval data. 10. Method according to claim 9, characterized in that information on the ventilation from the heart beat data is derived wherein the following methods (a-c) are used: a. Multiplying heart beat derived values of tidal volume with heart beat derived values on the frequency of breathing,b. Applying information on at least one chosen parameter derived from the heart beat,c. Using a mathematical function to transfer the chosen parameters measures into information on ventilation. 11. Method according to claim 1, characterized in that the method is used in a wearable computer. 12. Method according to claim 1, characterized in that the method is used in a fitness exercise equipment. 13. Method according to claim 1, characterized in that the reliability of the estimate is determined and expressed with the estimate. 14. Method according to claim 1, characterized in that the resulting respiratory frequency is used to determine oxygen consumption. 15. Method according to claim 1, characterized in that the resu... | Firstbeat Technologies Oy | A61B005/02 A61N001/00
US07460902 | B2 | 2008-12-02 | Monitoring of atrial activation | 2005-02-02 | A method and system for monitoring atrial activation in a patient is provided. In the method, a moment of occurrence and a wave form are predicted for a subsequent P wave based on a sequence of P waves already received in ECG signal data obtained from the patient. The process then checks whether the subsequent P wave really occurs in the ECG signal data within a first time window around the predicted moment of occurrence. If this is the case, the process further monitors whether the P wave is followed by a QRS complex in the ECG signal data. | 1. A method for monitoring atrial activation in an individual patient, the method comprising the steps of: receiving ECG signal data obtained from a patient;based on a first sequence of P waves received in the EGG signal data, predicting a moment of occurrence for a subsequent P wave in the ECG signal data;estimating a reference wave form for the subsequent P wave based on a second sequence of P waves received in the ECG signal data;checking whether said subsequent P wave occurs in the ECG signal data within a first time window around the predicted moment of occurrence;monitoring whether a subsequent QRS complex occurs in the ECG signal data within a second time window ending after the first time window; andindicating the detection of an AV block if the subsequent QRS complex fails to occur in the ECG data within the second time window,wherein the monitoring step is performed upon detecting in the checking step that said subsequent P wave occurred in the ECG signal data within the first time window. 2. A method according to claim 1, wherein the predicting step includes the steps of: calculating an estimate for a subsequent PP interval;defining the moment of occurrence for the subsequent P wave based on the estimate. 3. A method according to claim 2, wherein the calculating step includes a sub-step of calculating a first predetermined number of moving averages, each moving average representing a moving average of a second predetermined number of previous PP intervals calculated at a moment of time specific to the moving average concerned. 4. A method according to claim 3, wherein the calculating step further includes a sub-step of determining a weighted sum of the moving averages. 5. A method according to claim 4, further comprising a step of determining weight values for the weighted sum of the moving averages. 6. A method according to claim 5, wherein the calculating sub-step includes calculating three moving averages. 7. A method according to claim 6, wherein the determining step includes a sub-step of adapting the weight values to a second order polynomial. 8. A method according to claim 1, wherein the estimating step includes calculating an average wave form based on the second sequence of P waves, wherein the average wave form represents the reference wave form for the subsequent P wave. 9. A method according to claim 1, wherein the checking step includes the steps of: extracting a signal data sequence from the ECG signal data, the signal data sequence representing a time interval having a length substantially equal to that of the reference wave form and including the moment of occurrence predicted in the predicting step and;correlating the signal data sequence with the reference wave form, whereby a correlation coefficient indicative of the relationship of the signal data sequence and the reference wave form is obtained; anddeciding, based on the correlation coefficient, whether the monitoring step is to be performed. 10. A method ... | The General Electric Company | A61B005/04 A61B005/0452 A61B005/0472
US07460904 | B2 | 2008-12-02 | Wireless systems and methods for the detection of neural events using onboard processing | 2003-10-09 | A neural event electrode interface is configured to communicate with electrodes that capture electrical signals that are associated with at least one neuron. A neural event waveform processor is configured to extract neural events from the electrical signals. A wireless transmitter is configured to wirelessly transmit the neural events that are extracted from the electrical signals. | 1. A system for detecting neural events comprising: a neural event electrode interface that is configured to communicate with electrodes that capture electrical signals that are associated with at least one neuron;a neural event waveform processor that is configured to extract single neuron action potentials with waveforms that exceed 500 Hz from the electrical signals; anda wireless transmitter that is configured to wirelessly transmit the single neuron action potentials with waveforms that exceed 500 Hz that are extracted from the electrical signals. 2. A system according to claim 1 wherein the neural event waveform processor is configured to extract the single neuron action potentials with waveforms that exceed 500 Hz from the electrical signals by performing time/amplitude window discrimination of the electrical signals. 3. A system according to claim 1 wherein the neural event waveform processor is configured to extract the single neuron action potentials with waveforms that exceed 500 Hz from the electrical signals by applying template matching to the electrical signals. 4. A system according to claim 1 wherein the neural event waveform processor is further configured to time-stamp the single neuron action potentials with waveforms that exceed 500 Hz and wherein the wireless transmitter is configured to wirelessly transmit the single neuron action potentials with waveforms that exceed 500 Hz that are extracted from the electrical signals and associated time-stamps. 5. A system according to claim 1 wherein the neural event waveform processor is further configured to identify a neuron that is responsible for the single neuron action potentials with waveforms that exceed 500 Hz and wherein the wireless transmitter is configured to wirelessly transmit the single neuron action potentials with waveforms that exceed 500 Hz that are extracted from the electrical signals and the associated neuron identification. 6. A system according to claim 1 wherein the neural event waveform processor is configured to extract the single neuron action potentials with waveforms that exceed 500 Hz from the electrical signals by detecting occurrence of single neuron action potentials with waveforms that exceed 500 Hz in the electrical signals and by assigning the single neuron action potentials with waveforms that exceed 500 Hz to a category, and wherein the wireless transmitter is configured to wirelessly transmit the single neuron action potentials with waveforms that exceed 500 Hz that are extracted from the electrical signals by transmitting the category. 7. A system according to claim 1 further comprising a wireless receiver that is configured to wirelessly receive instructions for configuring the neural event waveform processor to use a selected category of algorithms to extract the single neuron action potentials with waveforms that exceed 500 Hz, and to provide the instructions to use the selected category of algorithms to extract the single neuron acti... | Wake Forest University Health Sciences | A61B005/04
US07470288 | B2 | 2008-12-30 | Telemetric tibial tray | 2004-03-31 | A telemetric tibial tray includes a lower plate that defines a plurality of cylindrical transducer cavities having circular load diaphragms. An upper plate is attached to the lower plate through support posts projecting from the load diaphragms. The support posts have a circular cross-section and a diameter of about 5.0 mm. The lower plate further defines wiring channels communicating between the transducer cavities and a central cavity housing the circuit board for the telemetric tray. The wiring channels are arranged at a 45 degree angle relative to the sagittal plane of the knee joint in order to reduce the effects of the wiring channel intersection on the strain sensitivity of the tray. Each transducer cavity includes a radial strain gage array with four pairs of radially aligned strain gages, each pair aligned at a 45 degree angle relative to the sagittal plane of the knee joint. Each pair of strain gages includes an inner gage positioned at the point of maximum positive micro-strain across the diaphragm when loaded, and an outer gage positioned at the point of maximum negative micro-strain, to thereby increase the differential strain measured by the gages and increase the strain sensitivity of the tibial tray. | 1. In a telemetric knee prosthesis adapted to measure forces transmitted across the knee joint, the knee prosthesis having a femoral component, a tibial bearing member in articulating contact with the femoral component, a tibia engaging member and a tibial tray engaged to the tibial bearing member and the tibia engaging member, the tibial tray comprising: an upper plate having a portion configured for engaging the tibial bearing member;a lower plate having a portion configured for engaging the tibia engaging member, said lower plate spaced apart from said upper plate and defining a plurality of cavities opening away from said upper plate, each of said cavities including a diaphragm adapted to flex when subjected to a load normal to the diaphragm;a plurality of support posts, each connected between said upper plate and said diaphragm of a corresponding one of said plurality of cavities, wherein said support posts are circular in cross-section and have a diameter that is about ⅓ the diameter of the corresponding one of said plurality of cavities; anda force sensing element disposed within each of said plurality of cavities and operable to produce an output signal in response to flexing of said diaphragm. 2. The tibial tray of claim 1, wherein said support posts each have a diameter of about 5.0 mm. 3. In a telemetric knee prosthesis adapted to measure forces transmitted across the knee joint, the knee prosthesis having a femoral component, a tibial bearing member in articulating contact with the femoral component, a tibia engaging member and a tibial tray engaged to the tibial bearing member and the tibia engaging member, the tibial tray comprising: an upper plate having a portion configured for engaging the tibial bearing member;a lower plate having a portion configured for engaging the tibia engaging member, said lower plate spaced apart from said upper plate and defining a plurality of cylindrical cavities opening away from said upper plate, each of said plurality of cylindrical cavities including a circular diaphragm adapted to flex when subjected to a load normal to the diaphragm and an outer wall;a plurality of support posts, each connected between said upper plate and said diaphragm of a corresponding one of said plurality of cylindrical cavities; anda force sensing element disposed within each of said plurality of cavities and operable to produce an output signal in response to flexing of said diaphragm, said force sensing element including four pairs of radially aligned strain gages, the strain gages of each pair arranged to measure differential strain in a radial direction and includes an inner gage mounted on said diaphragm adjacent the center of said circular diaphragm and an outer gage mounted on said diaphragm immediately adjacent said outer wall of said cylindrical cavity, wherein said circular diaphragm exhibits a micro-strain behavior under load that produces a maximum magnitude at a radial location from the center of said circul... | DePuy Products, Inc. | A61F002/38 A61B005/103
US07433728 | B2 | 2008-10-07 | Dynamic metal immunity by hysteresis | 2003-05-29 | Apparatus is provided for tracking an object in a body of a patient in the presence of an interfering article. A set of one or more radiators are adapted to generate an energy field at a fundamental frequency in a vicinity of the object. A position sensor, fixed to the object, is adapted to generate a signal responsive to the energy field. A control unit has access to a database of one or more harmonic frequency patterns associated with one or more respective specific types of interfering articles. The control unit is adapted to (a) receive the signal, (b) detect a pattern of harmonic frequencies of the fundamental frequency present in the signal responsive to an interaction of the interfering article with the energy field, (c) compare the pattern to harmonic frequency patterns stored in the database, (d) identify the interfering article responsive to the comparison, (e) correct the signal responsive to the identification of the interfering article and a magnitude of one or more of the harmonic frequencies in the detected pattern, and (f) determine position coordinates of the object responsive to the corrected signal. | 1. Apparatus for tracking an object in a body of a patient in the presence of an interfering article, comprising: a set of one or more radiators, which are adapted to generate an energy field at a fundamental frequency in a vicinity of the object;a position sensor, fixed to the object, which is adapted to generate a signal responsive to the energy field; anda control unit, with access to a database of one or more harmonic frequency patterns associated with one or more respective specific types of interfering articles, the control unit adapted to:receive the signal,detect a pattern of harmonic frequencies of the fundamental frequency present in the signal responsive to an interaction of the interfering article with the energy field,compare the pattern to harmonic frequency patterns stored in the database,identify the interfering article responsive to the comparison,correct the signal responsive to the identification of the interfering article and a magnitude of one or more of the harmonic frequencies in the detected pattern, anddetermine position coordinates of the object responsive to the corrected signal. 2. Apparatus according to claim 1, wherein the interfering article is ferromagnetic, and wherein the control unit is adapted to identify the ferromagnetic interfering article. 3. Apparatus according to claim 1, wherein the position sensor comprises at least one of: a Hall effect device, a coil, and an antenna. 4. Apparatus according to claim 1, wherein the fundamental frequency is between about 200 Hz and about 12 kHz, and wherein the radiators are adapted to generate the energy field at the fundamental frequency. 5. Apparatus according to claim 1, wherein the control unit is adapted to detect the pattern of harmonic frequencies by calculating a ratio of an amplitude of a first harmonic of the fundamental frequency present in the signal to an amplitude of a second harmonic of the fundamental frequency present in the signal, and to compare the pattern to harmonic frequency patterns stored in the database by comparing the calculated ratio to ratios stored in the database. 6. Apparatus according to claim 1, wherein the control unit is adapted to correct the signal by calculating an interfering effect of the interfering article on the signal, and by removing the interfering effect from the signal. 7. Apparatus according to claim 6, wherein the control unit is adapted to remove the interfering effect by subtracting, in a frequency domain, the interfering effect from the signal. 8. Apparatus according to claim 6, wherein the control unit is adapted to calculate the interfering effect by: calculating a ratio of an amplitude of at least one of the harmonic frequencies of the fundamental frequency present in the signal to an amplitude of the fundamental frequency present in the signal, andcomparing the calculated ratio to ratios stored in the database. 9. Tracking apparatus, comprising: an object, adapted to be inserted in a body of a patient;a ... | Biosense, Inc. | A61B005/05
US07822465 | B2 | 2010-10-26 | Device and method for image-based device performance measurement | 2007-04-25 | Devices and methods for providing in vivo performance measurements of an orthopedic implant involve an orthopedic implant having at least two radiopaque markers associated with a movable portion of the implant. The markers are positioned at known locations and in a first known and predetermined relationship to each other. The markers provide visible references upon application of a medical imaging technique to show the relative position of the markers according to a second, changed relationship to each other after implantation of the orthopedic implant. At least one measurement of performance of the implant can be determined from the first and second relationships of the markers. | 1. A method, comprising: providing an orthopedic implant device having at least one movable portion, wherein said device includes at least two radiopaque markers integrated with said device in a first known, predetermined relationship;implanting said device adjacent at least one vertebra of a spinal segment of a patient;applying a medical imaging technique to said device at a first moment in time, wherein said applying provides at least one image of said markers in a second, changed relationship resulting from movement of said implant;evaluating the relative positions of said markers in said second, changed relationship;measuring at least one indication of performance of said device based at least on said second, changed relationship; anddetermining the expected life span of said implant device based on said indication of performance. 2. The method of claim 1, wherein said measuring includes measuring the bending and rotation of said implant device. 3. The method of claim 1, comprising changing at least one property of said implant device based on said measured indication of performance. 4. The method of claim 1, comprising mandating at least one recovery activity for the patient based on said measured indication of performance. 5. The method of claim 1, comprising: applying said medical imaging technique to said implant device at a second moment in time, wherein said applying provides at least one image of said markers in a third, changed relationship resulting from movement of said implant device;evaluating the relative positions of said markers in said third, changed relationship; andmeasuring at least one indication of performance of said implant device based at least on said third, changed relationship. 6. The method of claim 5, wherein the first moment in time corresponds to a first position of the spinal segment and the second moment in time corresponds to a second position of the spinal segment. 7. The method of claim 5, comprising correlating each of said relationships with at least one physical parameter of the patient. 8. The method of claim 7, wherein said at least one physical parameter is selected from the group consisting of pain, muscle activation, and vertebral spacing. 9. A method, comprising: providing an orthopedic implant device having at least one movable portion, wherein said device includes at least two radiopaque markers integrated with said device in a first known, predetermined relationship;implanting said device adjacent at least one vertebra of a spinal segment of a patient;applying a medical imaging technique to said device at a first moment in time, wherein said applying provides at least one image of said markers in a second, changed relationship resulting from movement of said implant;evaluating the relative positions of said markers in said second, changed relationship;measuring at least one indication of performance of said device based at least on said second, changed relationship; andpredicting the exp... | Warsaw Orthopedic, Inc. | A61B005/05
US07942827 | B2 | 2011-05-17 | Minimally invasive allergy testing system | 2006-07-11 | An allergy testing system includes encapsulated allergens, a microneedle array, and an activation system coupled to the microneedle array and the encapsulated allergens such that the encapsulated allergens are moved into contact with a subject as the microneedle array is moved from a resting position to a penetrating position. A method for determining a degree of reaction to one or more allergens by a patient in a minimally invasive manner is also disclosed. Penetration of one or more microneedles into a skin of the patient is caused. Each of the penetrations into the skin is exposed with an allergen from each of the one or more microneedles. One or more images are captured of each of the penetrations into the skin. Each of the captured images are analyzed to assess the degree of reaction to the specific allergen. Allergic reactivity data is output for at least one of the allergens. | 1. An allergy testing system, comprising: encapsulated allergens;a microneedle array;an activation system coupled to the microneedle array and the encapsulated allergens such that the encapsulated allergens are moved into contact with a subject as the microneedle array is moved from a resting position to a penetrating position;at least one imaging system configured to capture a plurality of images of penetration sites by the microneedle array;at least one image sensor;an imaging optics system which focuses the images of penetration sites on the at least one image sensor;a second image sensor;a second imaging optics system which focuses the images of the penetration sites on the second image sensor, wherein the image sensors are used to determine a topographic profile of an allergic reaction; andan analyzer coupled to the at least one imaging system and configured to analyze the plurality of images to determine a time rate of change of an allergy response to one or more of the allergens. 2. The allergy testing system of claim 1, wherein the microneedle array pierces the encapsulated allergens while moving to the penetrating position. 3. The allergy testing system of claim 1, wherein the encapsulated allergens are coupled between the activation system and the microneedle array, and the encapsulated allergens are released by a pressure exerted on the encapsulated allergens by the activation system as the microneedle array is moved to the penetrating position. 4. The allergy testing system of claim 1, wherein at least one microneedle in the microneedle array comprises a hollow needle. 5. The allergy testing system of claim 1, wherein at least one microneedle in the microneedle array comprises a grooved needle. 6. The allergy testing system of claim 1, wherein at least one microneedle in the microneedle array comprises a corrugated needle. 7. The allergy testing system of claim 1, wherein the microneedle array comprises at least one needle of a first penetration depth and at least one needle of a second penetration depth which is different from the first penetration depth. 8. The allergy testing system of claim 1, wherein the microneedle array comprises at least one needle with a cross-section that is selected from the group consisting of: square, rectangular, triangular, and circular. 9. The allergy testing system of claim 1, wherein the microneedle array comprises at least one needle with a varying cross-section. 10. The allergy testing system of claim 1, wherein the microneedle array comprises silicon. 11. The allergy testing system of claim 1, wherein the microneedle array comprises glass. 12. The allergy testing system of claim 1, wherein the microneedle array comprises quartz. 13. The allergy testing system of claim 1, wherein the microneedle array comprises a substantially transparent material. 14. The allergy testing system of claim 1, further comprising a package that at least partially encloses the microneedle array and the enca... | Infotonics Technology Center, Inc. | A61B005/00
US07945461 | B2 | 2011-05-17 | Prescription compliance monitoring system | 2008-02-15 | A method for monitoring patient compliance to medical therapy that includes repeatedly measuring therapy compliance of a patient, uploading the therapy compliance measurements to a personal computer, connecting a central server using the personal computer, sending the uploaded therapy compliance measurements from the personal computer to the central server, providing a number of predetermined compliance threshold amounts, comparing the sent therapy compliance measurements with the number of predetermined compliance thresholds corresponding to the patient, determining a need and level of intervention based on the result of the comparing and reporting the determined need for intervention to a health professional corresponding to the patient and determined level. | 1. A method for monitoring patient compliance to medical therapy using a device having a central processing unit and a server having a central processing unit, comprising: repeatedly measuring therapy compliance of a patient by obtaining a voluntary compliance response from the patient using the device and storing information of the compliance on the device;uploading the therapy compliance measurements stored on the device to a central server via a wireless connection;selecting a number of individualized compliance threshold amounts for the patient, the individualized compliance threshold amounts being uniquely selected for the patient based on characteristics of the medical therapy;comparing, using the server, the uploaded therapy compliance measurements with the number of individualized compliance thresholds corresponding to the patient;identifying, using the server, instances where patient compliance fails to meet the individualized compliance thresholds; andreporting a failure to meet the individualized compliance thresholds to a care-giving party enabling action to be taken by that party. 2. The method according to claim 1, wherein the action is one of calling the patient, emailing the patient, speaking to the patient at a future checkup, sending an ambulance, visiting the patient, or calling a contact of the patient. 3. The method according to claim 1, further comprising: determining incentives and penalties based on the result of the comparing. 4. The method according to claim 1, further comprising: determining any continuation of therapy based on the result of the comparing. 5. The method according to claim 1, wherein the repeatedly measuring therapy compliance is accomplished using a NexDose device. 6. The method according to claim 1, wherein the central server is a SQL server. 7. The method according to claim 1, wherein a personal computer serves as a buffer between a collecting device used to repeatedly measuring therapy and a server. 8. The method according to claim 1, wherein the repeatedly measuring therapy compliance uses a PDA or cell phone equipped with software that records performance data and has the capacity to upload information to the server. 9. The method according to claim 1, wherein the number of predetermined compliance threshold amounts include, one of, a length of exercise, blood pressure, heart rate, weight, blood glucose, user entered thresholds or drug regimens. 10. The method according to claim 1, wherein the number of predetermined compliance threshold amounts include, one of, taking a dose of medication within a specified interval where many such intervals are scheduled, taking a physiologic measurement within a specified interval where many such intervals are scheduled, meeting with a healthcare provider within a specified interval where many such intervals are scheduled and performing exercise or physical therapy within a specified interval where many such intervals are scheduled. 11. The method acco... | VivoNex, L.L.C. | G06Q010/00 G06Q050/00
US08078273 | B2 | 2011-12-13 | Detection and classification of neuromuscular late wave activity for the assessment of neuromuscular function | 2007-03-30 | A method for the assessment of neuromuscular function by the detection and classification of late wave activity, comprising: (i) pre-processing an ensemble of traces so as to attenuate noise, PFI and baseline disturbances; (ii) identifying time segments of late wave activity in the ensemble of traces; (iii) classifying the time segments of late wave activity into F-wave, A-wave and Repeater segments by exploiting the variable morphology of F-waves across the traces and the fixed morphology of A-waves and Repeater waves across the traces; and (iv) searching the traces in the vicinity of the F-wave trace segments on a trace-by-trace basis so as to identify the F-wave onset latency. | 1. A method for detecting and classifying neuromuscular late wave activity so as to assess neuromuscular function in a patient, the method comprising: (i) electrically stimulating a nerve;(ii) measuring the electrical responses of the muscle innervated by the stimulated nerve, and recording the responses as an ensemble of traces;(iii) pre-processing, using a computer, an ensemble of traces so as to attenuate at least one of noise, power-line frequency interference (PFI) and baseline disturbances;(iv) identifying, using the computer, time segments of late wave activity in the ensemble of traces, wherein identifying time segments of late wave activity in the ensemble of traces comprises calculating a differential feature component and a mean feature component at every time point along the ensemble of traces and comparing the differential feature component and the mean feature component against thresholds so as to identify time segments of late wave activity in the ensemble of traces;(v) classifying, using the computer, the time segments of late wave activity into F-wave, A-wave and Repeater wave segments by exploiting the variable morphology of F-waves across the traces in the ensemble of traces and the fixed morphology of A-waves and Repeater waves across the traces in the ensemble of traces;(vi) searching, using the computer, the traces in the vicinity of the F-wave trace segments on a trace-by-trace basis so as to identify the F-wave onset latency; and(vii) using the identified F-wave onset latency to assess neuromuscular function in the patient. 2. A method according to claim 1 wherein the F-wave latency is determined using the ensemble of traces when the A-wave segments mask the onset of the F-waves. 3. A method according to claim 1 wherein the step of pre-processing an ensemble of traces so as to attenuate at least one of noise, power-line interference (PFI) and baseline disturbances comprises: a. reducing electronic noise by eliminating high frequencies on a trace-by-trace basis;b. estimating the amplitude and the phase of the power-line interference (PFI) components, at power-line frequency and harmonics of the power-line frequency, by a least squares estimation, and then using the estimated parameters to model the power-line interference (PFI) and thereafter remove the power-line interference (PFI) from the trace by subtracting the modeled power-line interference (PFI) from the trace; andc. calculating a sample-by-sample median trace from all the traces in the ensemble, and then modeling the baseline disturbance using a polynomial fit of the median trace. 4. A method according to claim 3 wherein a fixed fraction of samples are rejected, and then the polynomial fit is repeated with a reduced number of samples, in order to reduce the effect of late waves on the polynomial fit. 5. A method according to claim 4 wherein the rejected samples are those where the error between the median trace and the model is high. 6. A method according to c... | Neurometrix, Inc. | A61B005/04
US07962348 | B2 | 2011-06-14 | Apparatus, method and software for developing electronic documentation of imaging modalities, other radiological findings and physical examinations | 2008-02-15 | A system for guiding a user's physical examination of a patient and for generating a transcript. The system comprises: a screen requesting patient identification information, wherein the patient identification information excludes a patient name and further requesting patient medical history information; a screen requesting vital signs information; a screen requesting selection of an anatomical region of the patient for examination, wherein the user selects an anatomical region responsive to the request; a screen providing a plurality of possible observed symptoms related to the selected anatomical region, wherein the user selects one or more of the possible observed symptoms; a screen presenting possible findings that may be observed during the physical examination, wherein the user selects one or more of the possible findings based on the physical examination and a screen displaying a patient transcript responsive to information entered into the previous screens. | 1. A system for guiding a user's examination of a patient and for generating a transcript, the system comprising: a processor in communication with an output display device and a user interface, the processor programmed to output to the display device:(a) a screen requesting patient identification information, wherein the patient identification information excludes a patient name and further requesting patient medical history information;(b) a screen requesting vital signs information;(c) a screen presenting a plurality of symptoms in text form, wherein the user can identify one or more observed symptoms as selected from among the plurality of symptoms or the user can indicate that no symptoms are observed;(d) in response to the user identifying one or more observed symptoms, a screen presents a plurality of anatomical regions for examination, the plurality of anatomical regions presented in text form, wherein the user selects an anatomical region based on the one or more observed symptoms;(e) a screen presents anatomical subregions within a selected anatomical region, the anatomical subregions presented in text form, wherein the user selects an anatomical subregion based on the one or more observed symptoms;(f) a plurality of screens displays a rendering of the selected anatomical subregion and possible findings associated with a condition of the patient as related to the selected anatomical subregion, the renderings depicting normal and abnormal conditions of the selected anatomical subregion to assist the user with selecting one or more findings, and wherein the computer presents more-detailed renderings responsive to the user rolling a cursor over regions of the renderings;(g) wherein as the user selects findings, additional renderings are presented related to the selected findings;(h) wherein one or more of the screens include an icon for allowing the user to switch from the physical examination to an examination of a radiographic image, the processor for receiving user selections from the user interface and generating a patient transcript responsive to information entered into the screens (a) through (h). 2. The system of claim 1 wherein the patient transcript includes information regarding a selected rendering that best describes a patient's condition, the selected rendering using standard nomenclature. 3. The system of claim 2 wherein the renderings comprise skeletal line drawings, anatomical images and photographs. 4. The system of claim 2 wherein the user consults patient diagnostic test results and compares the diagnostic test results with the renderings to select one or more findings, wherein the diagnostic test results comprise a radiographic image, a magnetic resonance imaging image and a positron emission tomography image. 5. The system of claim 1 further comprising a screen permitting the user to select medical references and differential diagnoses related to the selected anatomical region. 6. The system of claim 1 wherein th... | - | G06Q010/00 G06Q050/00
US08452369 | B2 | 2013-05-28 | CMOS compatible microneedle structures | 2008-03-06 | The present invention provides an electronic device for sensing and/or actuating, the electronic device comprising at least one microneedle (10) on a substrate (1), each of the microneedles (10) comprising at least one channel (7, 8) surrounded by an insulating layer (6). The present invention also provides a method for making such an electronic device for sensing and/or actuating. | 1. An electronic device for at least one of sensing and actuating, the electronic device comprising: at least one microneedle on a substrate lying in a plane, wherein each of the at least one microneedle is perpendicular to the plane in which the substrate is lying, wherein each of the at least one microneedle comprises at least one channel surrounded by an insulating layer, the channel having a longitudinal axis which is directed in a direction substantially perpendicular to the plane in which the substrate is lying, wherein at least one microneedle comprises at least one channel that is at least partly filled with a filling material. 2. The electronic device according to claim 1, wherein the at least one channel that is at least partly filled with a filling material is completely filled with the filling material. 3. The electronic device according to claim 1, wherein at least one microneedle comprises at least one channel that is hollow. 4. The electronic device according to claim 1, wherein the filling material is a conductive material selected from the group consisting of Cu, Al, and W. 5. The electronic device according to claim 1, further comprising at least one active element or at least one passive element on top of at least one microneedle. 6. The electronic device according to claim 5, wherein the at least one active or the at least one passive element comprises a chip or electronic circuitry. 7. The electronic device according to claim 1, wherein the insulating layer has a thickness of between about 50 nm and about 1000 nm. 8. The electronic device according to claim 1, wherein the at least one microneedle has a diameter between about 100 nm and about 6 μm, a height between about 150 nm and about 50 μm, and an aspect ratio between about 0.5 and about 10. 9. The electronic device according to claim 1, further comprising at least two microneedles on the substrate, wherein a further insulating layer covers at least a part of the substrate and lies between the microneedles. 10. The electronic device according to claim 9, wherein either one or both of the insulating layer and the further insulating layer comprise at least one material that can serve as an electrical insulator or as a diffusion barrier. 11. The electronic device according to claim 9, wherein either one or both of the insulating layer and the further insulating layer comprise at least one material selected from the group consisting of silica, silicon carbide, and silicon nitride. 12. The electronic device according to claim 1, further comprising electronic circuitry in the substrate. 13. The electronic device according to claim 1, further comprising at least one microfluidic channel in the substrate. 14. The electronic device according to claim 1, wherein at least one microneedle comprises two channels, wherein each of the two channels is at least partly filled with the filling material, wherein the two channels are insulated from each other, and wherein each of ... | IMEC Katholieke Universiteit Leuven, K.U. Leuven R&D | A61B005/04
US10114086 | B2 | 2018-10-30 | Hybrid PET/MR imaging systems | 2008-08-21 | A hybrid imaging system includes a magnetic resonance scanner and a second modality imaging system disposed in the same radio frequency isolation space. The second modality imaging system includes radiation detectors configured to detect at least one of high energy particles and high energy photons. In some embodiments a retractable radio frequency screen is selectively extendible into a gap between the magnetic resonance scanner and the second modality imaging system. In some embodiments shim coils are disposed with the magnetic resonance scanner and are configured to compensate for distortion of the static magnetic field of the magnetic resonance scanner produced by proximity of the second modality imaging system. | 1. A hybrid imaging system comprising: a radio frequency isolation space defined by a Faraday cage;a magnetic resonance scanner disposed in the radio frequency isolation space;a second modality imaging system disposed in the same radio frequency isolation space as the magnetic resonance scanner, the second modality imaging system being one of the group consisting of a positron emission tomography (PET) scanner and a gamma camera, the magnetic resonance scanner and the second modality imaging system being physically separate units that are spaced apart by a gap; andsecond modality electronics associated with radiation detectors of the second modality imaging system, the second modality electronics configured to reduce radio frequency interference generated by the second modality electronics at one or more magnetic resonance frequencies of the magnetic resonance scanner by at least one of: the second modality electronics being clocked at a frequency selected to suppress generation of radio frequency interference at the one or more magnetic resonance frequencies of the magnetic resonance scanner,the second modality electronics being notch filtered to block generation of radio frequency interference at the one or more magnetic resonance frequencies of the magnetic resonance scanner,wherein no radio frequency shield is disposed between the magnetic resonance scanner and the second modality imaging system. 2. The hybrid imaging system as set forth in claim 1, further including: magnetic shielding arranged to reduce magnetic fields from the magnetic resonance scanner at the radiation detectors of the second modality imaging system to less than a magnetic field magnitude that results in degradation of performance of radiation detectors of the second modality imaging system. 3. The hybrid imaging system as set forth in claim 2, wherein the magnetic shielding includes: active shielding coils disposed on the second modality imaging system and configured to generate a compensatory magnetic field that at least partially cancels the static magnetic field B0 from the magnetic resonance scanner at radiation detectors of the second modality imaging system. 4. The hybrid imaging system as set forth in claim 1, further including: magnetic shielding arranged to redirect magnetic fields from the magnetic resonance scanner at the radiation detectors to a direction less interfering with radiation detectors of the second modality imaging system. 5. The hybrid imaging system as set forth in claim 1, the hybrid imaging system further comprising: rails for moving the second modality imaging system respective to the magnetic resonance scanner between a first position and a second position, the first position being relatively closer to the magnetic resonance scanner and the second position being relatively further from the magnetic resonance scanner. 6. The hybrid imaging system as set forth in claim 5, wherein the magnetic resonance scanner includes: switchable magneti... | KONINKLIJKE PHILIPS N.V. | A61B005/05 A61B006/00 G01R033/28 G01R033/3875 G01R033/421 G01R033/422 G01R033/48
US20080300475 | A1 | 2008-12-04 | Laryngoscope with Time Indicating Means and Method for Use Thereof | 2004-12-09 | The present invention discloses a laryngoscope (1) comprising time indicating means (20). This medical device comprising a handle (2) having a distal portion and a proximal portion; a blade (3) adapted to be inserted into the respiratory tract of the patient; said blade comprising a distal portion and a proximal portion and said distal portion being in communication with said proximal portion of the handle (2); time indicating means (20) adapted to indicate that medical procedure of endotracheal intubation has been terminated or is to be terminated soon. The laryngoscope may also include sensing means (42). The present invention also discloses a method for endotracheal intubation of a patient by means of the aforesaid laryngoscope (1). | 1. A self-contained laryngoscope comprising time indicating means for indicating a time associated with an endotracheal intubation procedure , additionally comprising; a. a handle having a distal portion and a proximal portion; and,b. a blade adapted to be inserted into a respiratory track of a patient;said blade comprising a distal portion and a proximal portion; said distal portion Is in communication with the proximal portion of the handle. 2. The laryngoscope according to claim 1, wherein the time indicating means is integrated in the handle of the laryngoscope. 3. The laryngoscope according to claim 1, wherein the time indicating means is attached to the laryngoscope in a non-integrated manner. 4. The laryngoscope according to claim 1, wherein the time indicating means is adapted to display a remaining time for the procedure. 5. The laryngoscope according to claim 4, wherein the remaining time is displayed continuously or discontinuously. 6. The laryngoscope according to claim 1, wherein the time indicating means is adapted to display remaining seconds before a termination of the procedure. 7. The laryngoscope according to claim 1, wherein the time indicating means is adapted to display a stage of the procedure. 8. The laryngoscope according to claim 1, wherein the time indicating means is attached to the handle of the laryngoscope. 9. The laryngoscope according to claim 8, wherein the time indicating means is characterized by a bracelet-like rounded structure adapted to grasp the handle in at least one location. 10. The laryngoscope according to claim 8, wherein the time indicating means is characterized by a plug-like or a cup-like structure adapted to connect to the handle. 11. A time indicating means adapted to indicate a termination time of an endotracheal intubation medical procedure; said means is being adapted to be attached in an integrated or non-integrated manner to a laryngoscope. 12. The time indicating means according to claim 11, adapted to display a remaining time until the procedure is to be terminated. 13. The time indicating means according to claim 11, wherein the time is displayed continuously or discontinuously. 14. The time indicating means according to claim 11, wherein the time indicating means is adapted to display remaining seconds before termination of the procedure. 15. The time indicating means according to claim 11, adapted to display a stage of the procedure. 16. The time indicating means according to claim 11, adapted to be attached to a handle portion of the laryngoscope. 17. The time indicating means according to claim 16, characterized by a bracelet-like rounded structure adapted to grasp the handle in at least one location. 18. The time indicating means according to claim 16, characterized by a plug-like or a cup-like structure adapted to connect to the handle. 19. A laryngoscope comprising non-pressure sensing means, comprising; a. a handle having a distal portion and a proximal portion... | - | A61B001/267 G04B047/00 A61B005/145 A61B005/01 A61B005/08
US20080300500 | A1 | 2008-12-04 | APNEA DETECTION USING A CAPNOGRAPH | 2007-05-30 | A method for diagnosis includes receiving a signal indicative of a partial pressure of CO2 in air expired by a patient during sleep. The signal is processed so as to detect a breathing-related event from a group of events consisting of apneas and hypopneas, and to classify the event as a central event or an obstructive event. | 1. A method for diagnosis, comprising: receiving a signal indicative of a partial pressure of CO2 in air expired by a patient during sleep;processing the signal so as to detect a breathing-related event from a group of events consisting of apneas and hypopneas, and to classify the event as a central event or an obstructive event; andgenerating a record of an occurrence and classification of the event. 2. The method according to claim 1, wherein processing the signal comprises detecting a repetitive waveform in the signal prior to the event, and comparing the signal during the event to the detected waveform. 3. The method according to claim 2, wherein comparing the signal comprises detecting an apnea responsively to an interruption of the repetitive waveform. 4. The method according to claim 3, wherein processing the signal comprises classifying the apnea as central or obstructive responsively to a level of the signal during the interruption. 5. The method according to claim 4, wherein detecting the repetitive waveform comprises finding peak and baseline values of the repetitive waveform, and wherein classifying the apnea comprises identifying a central apnea when the level is closer to the peak value than to the baseline value, and identifying an obstructive apneas when the level is closer to the baseline value than to the peak value. 6. The method according to claim 2, wherein detecting the repetitive waveform comprises finding shape parameters of the waveform, and wherein comparing the signal comprises detecting a hypopnea responsively to a change in one or more of the shape parameters. 7. The method according to claim 6, wherein the shape parameters comprise a respective slope and a respective duration of one or more phases of the waveform, and wherein processing the signal comprises classifying the hypopnea as obstructive responsively to a change in the respective slope and the respective duration of at least one of the phases. 8. The method according to claim 2, wherein detecting the repetitive waveform comprises finding a peak level of the waveform/ and wherein comparing the signal comprises detecting a hypopnea responsively to an increase in the peak level over multiple cycles of the repetitive waveform. 9. The method according to claim 8, wherein processing the signal comprises identifying a pattern of Cheyne-Stokes breathing responsively to a succession of alternating increases and decreases of the peak level. 10. The method according to claim 1, wherein generating the record comprises detecting and recording an occurrence of Cheyne-Stokes breathing. 11. The method according to claim 10, and comprising determining a prognosis of heart failure (HF) in the patient based on the occurrence of the Cheyne-Stokes breathing. 12. Apparatus for diagnosis, comprising: a sensor, which is configured to be coupled to a body of a patient during sleep and to output a signal indicative of a partial pressure of CO2 in air expired by a patient; ... | WIDEMED LTD. | A61B005/08
US20080249426 | A1 | 2008-10-09 | Palm type electrocardiograph | 2007-04-05 | The present invention provides a palm type electrocardiograph having a case, in which a push button is provided at the front side thereof, a USB port is installed along a lateral side thereof, and two metal holes are assembled at the back thereof. A conductive sticker is also designed with two metal electrodes being penetrately assembled in the center thereon. The conductive sticker has three layers, the first layer is foam, the second layer is conductive graphite, and the third layer is conductive adhesive. An internal circuit is provided in the palm type electrocardiograph, including a microprocessor, an A/D converter, a program memory and a flash memory. | 1. A palm type electrocardiograph, comprising: a case, wherein a push button is provided at the front side thereof, a USB port is installed along a lateral side thereof, and two metal holes are assembled at the back thereof;a conductive sticker, wherein two metal electrodes being penetrately assembled in the center thereon;an internal circuit, including a microprocessor, an A/D converter, a program memory and a flash memory. 2. The palm type electrocardiograph according to claim 1, wherein the conductive sticker has three layers, a first layer is foam, a second layer is conductive graphite, and a third layer is conductive adhesive; the conductive graphite and the conductive adhesive are each formed into two pieces, and the two metal electrodes are each penetrately assembled respectively through the two pieces so that the two metal electrodes do not short; the conductive adhesive contains AgCl. 3. The palm type electrocardiograph according to claim 1, wherein after the two metal electrodes are plugged into the two metal holes respectively and the conductive sticker is attached around a human heart, push the push button so that the internal circuit accepts electrocardio signals via the two metal electrodes for processing and storing in the flash memory; thereafter a USB wire can be used for connecting the USB port and a computer to display the electrocardio signals stored in the flash memory. 4. The palm type electrocardiograph according to claim 1, wherein a microprocessor PS 2101 of Phison Electronics Corp. is used to act as said microprocessor of the palm type electrocardiograph; AD8607ARMZ (A/D converter) and AD7468BRTZ (amplifier) of ADI Company are used to function as the A/D converter; EN29LV512 (EEPROM) of Eon Silicon Solution Inc. is used to act as the program memory; TH58NVG2S3BTG00 of TOSHIBA is used to function as the flash memory; a rechargeable battery UP042035-PK1 of Usance Technology Co., Ltd is used as a direct current power supply. 5. The palm type electrocardiograph according to claim 3, wherein a operation steps of the palm type electrocardiograph is as follows: first to push the push button for enabling power; a program begins accepting electrocardio signals via the two metal electrodes; the electrocardio signals is processed in the miroprocessor and stored in the flash memory; if a push button signal is received again, it means that a user stops the operation, so the program terminates; if no push button signal is received, then calculate if electrocardio signals are recorded for five minutes, if five minutes are reached, stop the operation, otherwise continue recording. | GTMRI, Inc. | A61B005/0404
US20080243011 | A1 | 2008-10-02 | BLOOD PRESSURE MONITOR | 2008-03-31 | The blood pressure monitor 100 includes plural case members 111 to 116 that are rotatably connected with one another to form a housing unit housing a cuff. The plural case members 111 to 116 are connected with plural torque hinges 140 or plural rotation hinges 150, and rotate to be transformed between a closed state in which the case members 111 to 116 are collapsed with the cuff housed herein, and an open state in which the case members 111 to 116 are opened. | 1. A blood pressure monitor housing a cuff therein, wherein at least four case members are rotatably connected with one another in a ring-shape to form a housing unit that houses the cuff, andthe case members rotate to be transformed between a closed state in which the case members are collapsed with the cuff enclosed therein, and an open state. 2. The blood pressure monitor according to claim 1, wherein at least one of the case members includes a display. 3. The blood pressure monitor according to claim 1, wherein an elastic body that exerts a force in a direction that opens the cover is provided on the main body. 4. The blood pressure monitor according to claim 2, wherein an adjusting mechanism that adjusts the force exerted by the elastic body is provided. 5. The blood pressure monitor according to claim 1, wherein the cuff includes a cuff ring, a blood-flow restricting bag outside the cuff ring, and a cuff band outside the blood-flow restricting bag. 6. The blood pressure monitor according to claim 5, wherein the cuff band has a circumferential extension rate lower than an axial extension rate. 7. The blood pressure monitor according to claim 5, wherein the cuff ring includes two openings through which support members are respectively provided, andthe support members are fitted to the main body and/or the cover. 8. The blood pressure monitor according to claim 1, further comprising a first locking mechanism that includes a first locking member that is provided on one of the case members that are adjacent in a state in which the cover is open, anda first hook that is provided on the other of the case members that are adjacent in the state in which the cover is open, and at which the first locking member locks, whereinthe first locking mechanism maintains the state in which the cover is open. 9. The blood pressure monitor according to claim 8, wherein the first locking mechanism includes a first button that unlocks the first locking mechanism. 10. The blood pressure monitor according to claim 1, further comprising a second locking mechanism that includes a second locking member that is provided on one of the case members that oppose each other in a state in which the cover is closed,a second hook that is provided on the other of the case members that oppose each other in the state in which the cover is closed, whereinthe second locking mechanism maintains the state in which the cover is closed. 11. The blood pressure monitor according to claim 10, wherein the second locking mechanism includes a second button that connects to and unlocks the second locking member. 12. A blood pressure monitor, comprising: a cuff;a plurality of case members that surround an outer circumference of the cuff;a plurality of hinges that rotatably connects the case members with one another to transform the case members between a open state in which the cuff is open and a stowed state in which the cuff is enclosed; anda lock member that locks a part of the ca... | CITIZEN HOLDINGS CO., LTD. | A61B005/02
US20080312886 | A1 | 2008-12-18 | System and Method for Estimating Substance Concentrations in Bodily Fluids | 2006-11-23 | A first aspect of the invention relates to a method for estimating the optical response of a subject comprising blood, the method comprising forming a model whereby the spectral response of a subject to incident light can be estimated, the model including a plurality of parameters, each parameter corresponding to the influence on the spectral response of a respective biochemical substance, directing light of a known spectral content at the subject, detecting the spectral response of the subject to the light and repeatedly performing the steps of: i. modelling the spectral response of the subject to the light by means of the model; ii. comparing the modelled spectral response with the detected response; and iii. altering the model in dependence on that comparison by varying one or more of the said parameters. | 1. A method for estimating the optical response of a subject comprising blood, the method comprising: forming a model whereby the spectral response of a subject to incident light can be estimated, the model including a plurality of parameters, each parameter corresponding to the influence on the spectral response of a respective biochemical substance;directing light of a known spectral content at the subject;detecting the spectral response of the subject to the light; andrepeatedly performing the steps of:i. modelling the spectral response of the subject to the light by means of the model;ii. comparing the modelled spectral response with the detected response; andiii. altering the model in dependence on that comparison by varying one or more of the said parameters. 2. A method as claimed in claim 1, wherein the step of altering the model is only performed if the modelled spectral response is closer to the detected response than the modelled spectral response of the current model. 3. A method as claimed in claim 1, wherein light comprising multiple wavelengths is directed at the subject and the spectral response of the subject is detected for each respective wavelength. 4. A method as claimed in claim 3, wherein the light directed at the subject is generated by multiple devices, each being capable of generating light consisting of a single respective wavelength. 5. A method as claimed in claim 3, wherein the number of multiple wavelengths is greater than or equal to the number of parameters included in the model. 6. A method as claimed in claim 1, wherein the steps of modelling, comparing and altering are performed repeatedly until a difference between the modelled spectral response and the detected response is below a predetermined threshold. 7. A method as claimed in claim 1, wherein the step of altering the model is performed using a neural network. 8. A method as claimed in claim 1, wherein the spectral response of the subject is detected in light passing through at least two different paths through the subject. 9. A method as claimed in claim 1, wherein the model includes parameters representing the light scattering and light absorption properties of the biochemical components. 10. A method as claimed in claim 1, wherein the parameters include the concentrations of one or more of fat, water and glucose in the subject. 11. A method as claimed in claim 3, wherein the multiple wavelengths include 2.10 μm, 2.27 μm and 2.32 μm. 12. A method for manufacturing a non-invasive sensor for estimating the concentration of a biochemical component in a subject comprising blood, the method comprising forming an altered model by the method of any preceding claim;forming a processing device configured to apply the model so as to estimate, in dependence on the spectral response of a subject to light of a known spectral content, the influence on the spectral response of the biochemical component; andproviding output means whereby the processing devic... | CITY UNIVERSITY | A61B005/00 G06F017/10
US20080292164 | A1 | 2008-11-27 | SYSTEM AND METHOD FOR COREGISTRATION AND ANALYSIS OF NON-CONCURRENT DIFFUSE OPTICAL AND MAGNETIC RESONANCE BREAST IMAGES | 2007-08-27 | A method for joint analysis of non-concurrent magnetic resonance (MR) and diffuse optical tomography (DOT) images of the breast includes providing a digitized MR breast image volume comprising a plurality of intensities corresponding to a 3-dimensional (3D) grid of voxels, providing a digitized DOT breast dataset comprising a plurality of physiological values corresponding to a finite set of points, segmenting the breast MR image volume to separate tumorous tissue from non-tumorous tissue, registering a DOT breast dataset and the MR image volume and fusing said registered DOT and MR datasets, wherein said fused dataset is adapted for analysis. | 1. A method for joint analysis of non-concurrent magnetic resonance (MR) and diffuse optical tomography (DOT) images of the breast, comprising the steps of: providing a digitized MR breast image volume comprising a plurality of intensities corresponding to a 3-dimensional (3D) grid of voxels;providing a digitized DOT breast dataset comprising a plurality of physiological values corresponding to a finite set of points;segmenting said breast MR image volume to separate tumorous tissue from non-tumorous tissue;registering said DOT breast dataset and said MR image volume; andfusing said registered DOT and MR datasets, wherein said fused dataset is adapted for analysis. 2. The method of claim 1, wherein said physiological values include one or more of total hemoglobin concentration, blood oxygenation saturation, and light scattering data. 3. The method of claim 1, wherein segmenting said breast MR image volume comprises: selecting at least one axial, at least one coronal, and at least one sagittal slice of said MR image volume;selecting 3 different seed points in each selected slice, said seed points representative of fatty breast tissue, non-fatty breast tissue, and non-breast tissue;determining a probability that a random walker starting at an unselected point reaches one of said selected seed points; andlabeling each unselected point according to the seed point with a highest probability to create a mask file, wherein each point in each slice is labeled as fatty breast tissue, non-fatty breast tissue, or non-breast tissue. 4. The method of claim 3, further comprising resampling said DOT dataset into a 3D volume of voxels corresponding to said MR grid of voxels. 5. The method of claim 4, comprising incorporating said mask file into said DOT dataset. 6. The method of claim 1, wherein registering said DOT breast dataset to said MR image volume comprises: generating a 2D sagittal projection signature from said MR image and from said DOT dataset;registering said DOT sagittal signature and said MR sagittal signature;generating a 2D coronal projection signature from said MR image and from said DOT dataset;registering said DOT coronal signature and said MR coronal signature;generating a 2D axial projection signature from said MR image and from said DOT dataset;registering said DOT axial signature and said MR axial signature, wherein said 3D registration mapping is defined in terms of said 2D sagittal, coronal, and axial registrations. 7. The method of claim 6, wherein said steps of generating a 2D projection signature and registering said signatures for each of said sagittal, coronal, and axial projections is repeated for a predetermined number of iterations. 8. The method of claim 6, wherein said 2D projection signatures are generated from a maximum intensity projection. 9. The method of claim 6, wherein one of said DOT and MR signatures is a moving signature and the other is a fixed signature, and wherein registering a DOT signature and an MR sig... | Siemens Corporate Research, Inc. | A61B005/055
US20080281163 | A1 | 2008-11-13 | APPARATUS AND METHOD FOR ACQUIRING MEDICAL DATA | 2007-05-10 | A medical diagnostic/monitoring system is disclosed herein. The medical diagnostic/monitoring system includes a data collection system having a sensor, and a collection device connected to the sensor. The collection device includes a first near field communication device. The medical diagnostic/monitoring system also includes a signal acquisition device wirelessly connected to the data collection system. The signal acquisition device includes a second near field communication device. The first near field communication device and the second near field communication device are collectively configured to wirelessly transfer data from the data collection system to the signal acquisition device. | 1. A medical diagnostic/monitoring system comprising: a data collection system comprising: a sensor; anda collection device connected to the sensor, said collection device including a first near field communication device; anda signal acquisition device wirelessly connected to the data collection system, said signal acquisition device including a second near field communication device;wherein said first near field communication device and said second near field communication device are collectively configured to wirelessly transfer data from the data collection system to the signal acquisition device. 2. The medical diagnostic/monitoring system of claim 1, wherein the medical diagnostic/monitoring system comprises an electrocardiogram system. 3. The medical diagnostic/monitoring system of claim 1, wherein the medical diagnostic/monitoring system comprises an electroencephalogram system. 4. The medical diagnostic/monitoring system of claim 1, wherein the data collection system is disposable. 5. The medical diagnostic/monitoring system of claim 1, wherein the sensor comprises a plurality of electrodes. 6. The medical diagnostic/monitoring system of claim 1, wherein the collection device includes an amplifier configured to amplify signals from the sensor. 7. The medical diagnostic/monitoring system of claim 1, wherein the collection device includes a converter configured to convert analog data from the sensor into digital data. 8. The medical diagnostic/monitoring system of claim 1, wherein the collection device is configured to be inexpensive and lightweight. 9. The medical diagnostic/monitoring system of claim 1, wherein the signal acquisition device includes a monitor. 10. The medical diagnostic/monitoring system of claim 1, wherein the data collection system includes a conductor configured to couple the sensor with the collection device. 11. An electrocardiogram system comprising: a data collection system comprising: a sensor;a conductor connected to the sensor; anda collection device connected to the conductor, said collection device including a first near field communication device, wherein said collection device is configured to be inexpensive and lightweight; anda signal acquisition device wirelessly connected to the data collection system, said signal acquisition device including a second near field communication device;wherein said first near field communication device and said second near field communication device are collectively configured to wirelessly transfer power from the signal acquisition device to the data collection system and to wirelessly transfer data from the data collection system to the signal acquisition device. 12. The electrocardiogram system of claim 11, wherein the data collection system is disposable. 13. The electrocardiogram system of claim 11, wherein the sensor comprises a plurality of electrodes. 14. The electrocardiogram system of claim 11, wherein the conductor comprises a plurality of lead wire... | GENERAL ELECTRIC COMPANY | A61B005/00
US20080269664 | A1 | 2008-10-30 | System and Method For Controlling Traversal of an Igested Capsule | 2006-01-16 | A treatment system (1600, 1900) is provided for traversing the alimentary tract. The system (1600, 1900) includes an ingestible capsule, which includes a gas pressurizing module (1602)) providing a gas and at least one balloon (1604, 1901) in fluid communication with the gas pressurization module (1602). The capsule further includes an exhaust channel (1610) in fluid communication with a respective balloon of the at least one balloon (1604, 1901), and a depressurizing closure member (1608) for selectively controlling flow of gas between the balloon (1604, 1901) and the ambient surroundings of the capsule. The system further includes control circuitry (906) for controlling the depressurizing closure member (1608). | 1. A treatment system (1600, 1900) for traversing the alimentary tract comprising: an ingestible capsule comprising:a gas pressurizing module (1602) providing a gas;at least one balloon (1604, 1901) in fluid communication with the gas pressurization module (1602);an exhaust channel (1610) in fluid communication with a respective balloon of the at least one balloon (1604);a depressurizing closure member (1608) for selectively controlling flow of gas between the respective balloon and the ambient surroundings of the capsule (1600, 1900) through the exhaust channel (1610); andcontrol circuitry (906) for controlling the depressurizing closure member (1608). 2. The treatment system (1600, 1900) according to claim 1, wherein the capsule further comprises a pressurizing closure member (1606) for controlling flow of fluid between the gas pressurizing module (1602) and the at least one balloon (1604, 1901), wherein the control circuitry (906) controls the pressurizing closure member (1606) for controlling inflation of the at least one balloon (16041901). 3. The treatment system (1600, 1900) according to claim 2, wherein at least one of the pressurizing closure member (1606) and the depressurizing closure member (1608) includes a MEMS microvalve and an associated actuator. 4. The treatment system (1600, 1900) according to claim 2, wherein at least one of the pressurizing closure member (1606) and the depressurizing closure member (1608) includes a hatch and a MEMS motor for opening and closing the hatch. 5. The treatment system (1600, 1900) according to claim 2, wherein at least one of the pressurizing closure member (1606) and the depressurizing closure member (1608) includes an artificial muscle including a polymer which selectably contracts and expands in response to an electrical stimulus for facilitating achievement of the open and closed state of the respective closure member (1606m 1608). 6. The treatment system (1600, 1900) according to claim 1, the capsule further comprising a catheter (1904) for combining with one of the balloons of the at least one balloon to form a balloon catheter. 7. The treatment system (1600, 1900) according to claim 1, a balloon of the at least one balloon (1604, 1901) including an outer microporous membrane with holes infused with a medicament, and wherein upon inflation of the balloon the medicament weeps from the membrane to the ambient surroundings of the capsule. 8. The treatment system (1600, 1900) according to claim 1, wherein the capsule further comprises: a medicament delivery system (901) for outputting a medicament; anda perimetrical lumen 1630) having an open end and having a plurality lf apertures (1631), wherein the perimetrical lumen (1630) is wound around a respective balloon of the at least one balloon (1604, 1901);wherein the open end of the perimetrical lumen (1630) receives medicament output by the medicament delivery system (901) and the medicament is output through the apertures (1631) to the a... | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61M037/00 A61B018/20 A61B005/103
US20080242943 | A1 | 2008-10-02 | SYSTEM AND METHOD OF PATIENT MONITORING AND DETECTION OF MEDICAL EVENTS | 2007-03-28 | A system of patient health condition monitoring includes a device configured to measure a health parameter of a patient and a computer. The computer is programmed to receive an input based on the measured health parameter, determine a first moving average value for a first period of time based on the measured health parameter and determine a second moving average value for a second period of time based on the measured health parameter, the second period of time different than the first period of time. The computer is further programmed to calculate a difference between the first and second moving average values and store the difference in computer memory. | 1. A system of patient health condition monitoring comprising: a device configured to measure a health parameter of a patient;a computer programmed to: receive an input based on the measured health parameter;determine a first moving average value for a first period of time based on the measured health parameter;determine a second moving average value for a second period of time based on the measured health parameter, the second period of time different than the first period of time;calculate a difference between the first and second moving average values; andstore the difference in computer memory. 2. The system of claim 1 wherein the computer is further programmed to: compare the difference to a range of non-alarm values; andtrigger an alarm if the difference falls outside of the range of non-alarm values. 3. The system of claim 2 wherein the computer is further programmed to display the alarm to a user. 4. The system of claim 1 wherein the device is located remotely from the computer. 5. The system of claim 4 wherein the device is connected to the computer via a communications link. 6. The system of claim 5 wherein the computer is programmed to receive the input directly from the device over the communications link. 7. The system of claim 6 were in the device is further configured to automatically communicate the measured health parameter to the computer over the communications link. 8. The system of claim 4 wherein the computer is programmed to receive the input via a telephone system. 9. The system of claim 1 wherein the first period of time is 3 days and wherein the second period of time is 60 days. 10. The system of claim 1 wherein the health parameter is one of a weight, a blood pressure, a pulse, and a blood sugar of the patient. 11. The system of claim 1 wherein the first and second moving average values are determined based on one of a simple moving average, a rolling average, a weighted moving average, and an exponential moving average. 12. A method of patient monitoring comprising: calculating a short-term moving average value based on a measured patient health parameter;calculating a long-term moving average value based on the measured patient health parameter;comparing the short-term moving average value to the long-term moving average value; andstoring a result of the comparison to database on a computer readable storage medium. 13. The method of claim 12 further comprising comparing the result of the comparison to a threshold. 14. The method of claim 13 further comprising alerting a medical staff member if the result of the comparison crosses a threshold. 15. The method of claim 13 wherein the measured patient health parameter is weight and wherein the threshold is approximately 5 pounds. 16. The method of claim 12 further comprising removing stored results from the database immediately before and after a decompensation event while the results are greater than or equal to zero. 17. The method of claim 12 further co... | - | A61B005/00
US20080306355 | A1 | 2008-12-11 | Method and System for Monitoring Gastrointestinal Function and Physiological Characteristics | 2008-05-08 | A system and method for evaluating gastrointestinal motility and, optionally, other physiological characteristics (e.g., pulse rate) that can be effectively employed to acquire one or more signals associated with acoustic energy (i.e. sound) emanating from an abdominal region of a body and determine at least one gastrointestinal parameter or event based on the acoustic energy signal(s) is described. The gastrointestinal parameter can include a gastrointestinal event, including gastrointestinal mixing, emptying, contraction and propulsion, and gastrointestinal transit time, or a gastrointestinal system disorder, including reflux disease, irritable bowel disease, ulcerative colitis, constipation, diarrhea, and a migrating motor complex disorder. | 1. A system for monitoring gastrointestinal function of a subject, comprising: at least one sensor mountable proximate a body region of the subject, said sensor being adapted to sense acoustic energy and generate at least one acoustic energy signal representing said acoustic energy; anda processing unit adapted to receive said acoustic energy signal, said processing unit being further adapted to process said acoustic energy signal and determine the occurrence of at least one gastrointestinal parameter therefrom. 2. The system of claim 1, wherein said gastrointestinal parameter comprises an event selected from the group consisting of gastrointestinal mixing, emptying, contraction and propulsion, and gastrointestinal transit time. 3. The system of claim 1, wherein said gastrointestinal parameter comprises an event associated with a gastrointestinal system disorder, said gastrointestinal system disorder being selected from the group consisting of reflux disease, irritable bowel disease, ulcerative colitis, constipation, diarrhea, and a mitigating motor complex disorder. 4. The system of claim 1, wherein said sensor generates a plurality of acoustic energy signals representing said acoustic energy. 5. The system of claim 4, wherein said processing unit is adapted to receive and process said plurality of acoustic energy signals and determine the occurrence of at least a first gastrointestinal parameter therefrom. 6. The system of claim 5, wherein said first gastrointestinal parameter comprises gastrointestinal transit time. 7. A system for monitoring gastrointestinal function of a subject, comprising: at least one acoustic energy sensor mountable proximate a body region of the subject, said acoustic energy sensor being adapted to sense acoustic energy generated by the subject and generate at least one acoustic energy signal representing said acoustic energy;at least one spatial parameter sensor mountable proximate a body region of the subject, said spatial parameter sensor being adapted to monitor at least one spatial parameter associated with the subject's body and generate at least one spatial parameter signal representing said spatial parameter; anda processing unit adapted to receive said acoustic energy and spatial parameter signals, said processing unit being further adapted to determine the occurrence of at least one gastrointestinal parameter as a function of said acoustic energy and spatial parameter signals. 8. The system of claim 7, wherein said spatial parameter sensor comprises a motion sensor that is adapted to monitor motion of the subject's body. 9. The system of claim 8, wherein said spatial parameter comprises said motion of the subject's body. 10. The system of claim 7, wherein said spatial parameter sensor comprises an orientation sensor that is adapted to monitor orientation of the subject's body. 11. The system of claim 10, wherein said spatial parameter comprises said orientation of the subject's body. 12. The system ... | SMITHKLINE BEECHAM CORPORATION | A61B005/02 A61B007/00 A61B005/11
US20080287808 | A1 | 2008-11-20 | APPARATUS, PROBE AND METHOD FOR PROVIDING DEPTH ASSESSMENT IN AN ANATOMICAL STRUCTURE | 2007-09-12 | Apparatus and method can be provided for obtaining information regarding at least one portion of an anatomical structure. For example, using at least one first arrangement, it is possible to forward at least one first electromagnetic radiation to the at least one portion. In addition, using at least one second arrangement, it is possible to detect at least one second electromagnetic radiation from the sample, the second electromagnetic radiation being related to the first electromagnetic radiation. The second arrangement can be used to obtain data associated with the second electromagnetic radiation at a plurality of integration times. Further, with at least one third arrangement, it is possible to determine at least one motion characteristic as a function of depth within the portion using the second electromagnetic radiation. The motion characteristic can be determined by obtaining data associated with the second electromagnetic radiation. The data can be speckle data detected by the second arrangement. In addition, using the third arrangement, it is possible to determine optical properties of the portion based on the data. Furthermore, with a fourth arrangement, it is possible to receive the first electromagnetic radiation and generate a plurality of third electromagnetic radiations to irradiate multiple points on the portion. The second arrangement can be used to detect the second electromagnetic radiation from the sample based on the third electromagnetic radiations, and the third arrangement can also be used to determine a plurality of motion characteristics as a function of depth for a plurality of location of the at least one portion as a function of the second electromagnetic radiation. | 1. An apparatus for obtaining information regarding at least one portion of an anatomical structure, comprising: at least one first arrangement which is configured to forward at least one first electromagnetic radiation to the at least one portion;at least one second arrangement which is configured to detect at least one second electromagnetic radiation from the at least one sample, the at least one second electromagnetic radiation being related to the at least one first electromagnetic radiation, wherein the at least one second arrangement obtains data associated with the at least one second electromagnetic radiation at a plurality of integration times; andat least one third arrangement which is configured to determine at least one motion characteristic as a function of depth within the at least one portion using the at least one second electromagnetic radiation, wherein the at least one third arrangement determines the at least one motion characteristic by obtaining data associated with the at least one second electromagnetic radiation, the data being speckle data detected by the at least one second arrangement. 2. The apparatus according to claim 1, wherein the data is an image of at least one section of the at least one portion. 3. The apparatus according to claim 1, wherein the at least one third arrangement is further configured to determine a contrast of the data. 4. The apparatus according to claim 3, wherein the contrast is determined as function of a distance of illumination from a point of contact of the at least one first electromagnetic radiation on the at least one portion. 5. The apparatus according to claim 1, wherein the at least one third arrangement is further configured to determine at least one decorrelation function of the data collected at the integration times. 6. The apparatus according to claim 5, wherein the at least one decorrelation function is determined as function of a distance of illumination from a point of contact of the at least one first electromagnetic radiation on the at least one portion. 7. The apparatus according to claim 5, wherein the at least one third arrangement is further configured to determine at least one time constant of the at least one decorrelation function of the data collected at the integration times. 8. The apparatus according to claim 1, wherein the at least one third arrangement is further configured to determine optical properties of the at least one portion based on the data. 9. The apparatus according to claim 8, wherein the at least one third arrangement is further configured to determine the at least one motion characteristic of at least one section of the at least one portion as a function the optical properties and a time constant of the at least one decorrelation function of the data collected at the integration times. 10. The apparatus according to claim 8, wherein the optical properties are determined based on the data collected at the integration times using a diffus... | The General Hospital Corporation | A61B005/00
US20080294060 | A1 | 2008-11-27 | DEVICES AND METHODS FOR DISEASE DETECTION, MONITORING AND/OR MANAGEMENT | 2007-05-21 | Embodiments of the invention are related to methods and devices for respiratory or cardiac disease detection, monitoring, and/or management. In an embodiment, the invention includes a method of calculating a pulmonary function parameter of a subject. The method can include obtaining a first signal indicative of lung volume change during breathing from a first sensor, obtaining a second signal indicative of distending pressure from a second sensor, and calculating the pulmonary function parameter based on the first signal and the second signal. In an embodiment, the invention includes a method of monitoring pulmonary or cardiac disease status. In an embodiment, the invention includes an implantable medical device. The implantable medical device can include a first sensor configured to produce a first signal indicative of lung volume change during breathing, a second sensor configured to produce a second signal indicative of intrapleural pressure, and a processor configured to calculate lung compliance or pulmonary resistance based on the first signal and the second signal. Other aspects and embodiments are provided herein. | 1. A method of determining a pulmonary function parameter of a subject, the method comprising obtaining a first signal indicative of lung volume change during breathing from a first sensor;obtaining a second signal indicative of lung distending pressure from a second sensor, wherein at least one of the first and second sensors are chronically implanted; andcalculating the pulmonary function parameter based on the first signal and the second signal. 2. The method of claim 1, the pulmonary function parameter selected from the group consisting of lung compliance, pulmonary resistance, pressure-volume loop area, and pressure-volume loop centroid. 3. The method of claim 1, further comprising prompting the subject to perform a respiratory maneuver. 4. The method of claim 1, further comprising electrically stimulating contraction of the diaphragm. 5. The method of claim 1, further comprising electrically stimulating the phrenic nerve. 6. The method of claim 1, wherein both the first sensor and the second sensor are chronically implanted. 7. The method of claim 1, wherein obtaining a first signal indicative of lung volume change during breathing from a first sensor comprises obtaining an impedance signal from an impedance sensor. 8. The method of claim 1, wherein obtaining a second signal indicative of distending pressure from a second sensor comprises obtaining a pulmonary artery pressure signal from a pressure sensor. 9. The method of claim 1, wherein obtaining a second signal indicative of distending pressure from a second sensor comprises obtaining a pleural pressure signal. 10. The method of claim 1, further comprising transmitting the first signal and second signal to a processing unit. 11. The method of claim 1, wherein obtaining the first signal and obtaining the second signal are performed during normal tidal breathing of the patient. 12. The method of claim 1, wherein obtaining the first signal and obtaining the second signal are only performed in response to a triggering signal. 13. The method of claim 12, the triggering signal related to posture of the subject, wherein the triggering signal indicates that the subject is in a supine position. 14. The method of claim 12, the triggering signal related to the time of day. 15. The method of claim 12, the triggering signal related to the activity level of the subject, wherein the triggering signal indicates that the subject is at a particular level of activity. 16. A method of monitoring pulmonary or cardiac disease status, the method comprising: obtaining a first signal indicative of lung volume change during breathing with a first sensor;obtaining a second signal indicative of distending pressure with a second sensor, wherein at least one of the first and second sensors is chronically implanted;calculating lung compliance and/or pulmonary resistance based on the first signal and the second signal; andmonitoring lung compliance and/or pulmonary resistance values over a period of tim... | CARDIAC PACEMAKERS, INC. | A61B005/087
US20080262319 | A1 | 2008-10-23 | Device for Measuring Brain Parameters | 2006-01-21 | A device (1) is used for measuring brain parameters. This device has a catheter sensor unit (2) which can be implanted distally in a minimally-invasive manner in the parenchyma and/or the ventricles, guided through the cranial bone (4). The sensor unit (2) can be fixed proximally to the cranial bone. For this purpose a proximal circumferential portion of the sensor unit (2) has a self-cutting outer screw thread (5) for fastening the sensor unit (2) to the cranial bone (4). The result is a device which can be fastened in a secure, permanent and at the same time inexpensive manner preventing movement relative to the cranial bone. | 1. Device (1) for measuring brain parameters comprising a catheter sensor unit (2), which can be implanted distally in a minimally-invasive manner in at least one of the parenchyma and the ventricles, guided through the cranial bone (4),can be fixed proximally to the cranial bone (4), wherein a proximalcircumferential portion of the sensor unit (2) has a self-cutting outer screw thread (5) for fixing the sensor unit (2) to the cranial bone (4). 2. Device according to claim 1, wherein the outer screw thread (5) is an integral component of the sensor unit (2). 3. Device according to claim 1, wherein the outer screw thread (5) is configured such that the sensor unit (2) can be sealed from the cranial bone (4) by means of said outer screw thread. 4. Device according to claim 3, wherein the outer screw thread (5) is made of a sealing material. 5. Device according to claim 3, wherein the outer screw thread (5) has a sealing coating. 6. Device according to claim 1, wherein, at a proximal end portion (9), the sensor unit (2) has a stop (10) which can be positioned against the cranial bone (4). | Raumedic AG | A61B017/58 A61B005/00
US20080249425 | A1 | 2008-10-09 | Method and Apparatus for Defining Cardiac Time Intervals | 2005-08-30 | A method and apparatus for defining cardiac time intervals using the pre-systolic diastolic filling period and the proceeding systolic period. The diastolic period can begin at the end of the T wave and terminate at the onset of the proceeding Q wave as defined in electrocardiograph signals. The systolic period can begin at onset of the Q peak and terminate at the end of the T wave as defined in electrocardiograph signals. | 1. A method of performing heart monitoring, the method including the step of defining cardiac time intervals using the pre-systolic diastolic filling period and the proceeding systolic period. 2. A method as claimed in claim 1 wherein said diastolic period begins at the T wave and terminates at onset of the proceeding QRS, as defined in electrocardiograph signals. 3. A method as claimed in claim 1 wherein said systolic period begins at the onset of the Q wave and terminates at the T wave peak, as defined in electrocardiograph signals. 4. A method of monitoring the heart, the method including the step of utilising the pre-systolic diastolic filling period and the proceeding systolic period to derive an indicator of heart performance. 5. A method as claimed in claim 4 further comprising the step of utilising said indicator in treatment of the heart. 6. A method for measuring the cardiac period of a heart, the method comprising the steps of: (a) extracting a waveform indicative of cardiac activity within the heart;(b) determining a pre-systolic diastolic filling period and the proceeding systolic period; and(c) combining the pre-systolic diastolic filling period and the proceeding systolic period and outputting the combined result as a measure of cardiac period. 7. An apparatus for measuring the cardiac period of a heart, the apparatus comprising: a transducer for measuring physical or electrical activity associated with the heart;processing unit for extracting a measure of a pre-systolic diastolic filling period and the proceeding systolic period;display for outputting the combination of the pre-systolic diastolic filling period and the proceeding systolic period as a measure of cardiac period. 8. An apparatus as claimed in claim 7 wherein said transducer is a continuous wave Doppler transducer. | USCOM LIMITED | A61B005/0468
US20080269619 | A1 | 2008-10-30 | Method and Means for Measuring Systolic Blood Pressure in the Ankle | 2005-11-04 | A flexible pad for measuring systolic ankle blood pressure comprises two or more pairs of light emitting diodes and photo detectors. The pairs are disposed in parallel. The detectors are adapted for detecting light emitted by the respective diode into tissue and reflected from there. The pad further comprises conductor means for providing power to the light emitting diodes from a power source and means selected from conducting means and wireless means for putting the detectors in communication with electronic equipment for detector signal analysis. Also disclosed is the combination of the pad and an inflatable cuff, and a method for measuring systolic ankle blood pressure. | 1. A method for measuring systolic blood pressure in the ankle of a subject, comprising providing an assembly including an ankle cuff, a flexible measuring pad comprising two pairs or more of light emitting diodes and photo detectors, and an electronic control unit in communication with the pad, the pad being optionally fixed at the cuff;positioning the pad or, if the pad is fixed at the cuff, the combination of pad and cuff in contact with the skin of the ankle region of the subject so as to dispose one pair in proximity of the anterior tibial artery and substantially parallel with it and the other pair in proximity of the posterior tibial artery and substantially parallel with it and, if the pad and cuff are separate, positioning the cuff around the pad that is in contact with the skin of the ankle region;inflating the cuff to a pressure sufficient for stopping blood flow through the anterior and posterior tibial arteries;deflating the cuff while making the diodes emit light;recording the light reflected from the tissue by the photo detectors during deflation;recording the cuff pressure during deflation;analysing the recorded light signal to identify the cuff pressure at which of blood flow in the anterior and/or posterior tibial arteries is resumed. 2. The method of claim 1, comprising positioning of the pad or the combination of pad and cuff on the ankle in a manner so as to dispose the photo detectors downstream of the light emitting diodes in respect of the blood flow in said arteries. 3. The method of claim 2, comprising positioning the pad in respect of the cuff so as to avoid compression of ankle tissue distally of the pad. 4. A flexible pad for measuring systolic ankle blood pressure, comprising two or more pairs of light emitting diodes and photo detectors, said pairs being disposed in parallel, the detectors being adapted for detecting light emitted by the respective diode into tissue and reflected from there, the pad further comprising conductor means for providing power to the light emitting diodes from a power source and for putting the detectors in communication with electronic equipment for detector signal analysis. 5. The pad of claim 4 comprising a means for detector signal amplification. 6. The pad of claim 5, wherein the detector signal amplification means comprises an AMP board. 7. The pad of claim 4 of about rectangular shape, its short sides being designed to extend in a proximal/distal direction when mounted on the leg of a patient. 8. The pad of claim 4, of a length allowing it to extend around at least half of the circumference of the ankle of an adult person. 9. The pad of claim 7 comprising two sections of about rectangular shape, the line of division being parallel to the short sides and intended for disposition along the margo anterior tibiae in a mounted position. 10. The pad of claim 9, wherein each section comprises one or more pairs of light emitting diodes and one or more pairs of photodetectors. 11. ... | - | A61B005/022
US20080287751 | A1 | 2008-11-20 | APPARATUS FOR DETECTING HUMAN PHYSIOLOGICAL AND CONTEXTUAL INFORMATION | 2007-10-29 | The invention comprises an apparatus for determining the contextual or physiological status of the individual wearing the apparatus. The apparatus is designed to be consumable or disposable. In most embodiments the invention comprises an adhesive housing. In some embodiments, two different sensors are secured to the housing. The apparatus is in electronic communication with a processing unit that can derive analytical status data by using the data received from the two sensors. In some embodiments, the processing unit is programmed to control other devices, or is programmed to trigger an event. In still other embodiments, the apparatus is in electronic communication with a separate computing device, which may contain the processing unit. | 1. An apparatus for monitoring an individual's physiological or contextual status, said apparatus able to be disposed of after use, said apparatus comprising: a. a housing having a surface, said surface comprising an adhesive material, said adhesive material enabling said housing to be removably attached to selected portions of said individual's body;b. a first sensor secured to said housing, said first sensor capable of generating heart-related data of said individual;c. a second sensor secured to said housing, said second sensor capable of generating temperature-related data of said individual; andd. a processing unit in electronic communication with said first and second sensors, said processing unit programmed to use both of said heart-related data from said first sensor and said temperature-related data from said second sensor to derive analytical status data. | - | A61B005/02
US20080262557 | A1 | 2008-10-23 | OBESITY MANAGEMENT SYSTEM | 2007-04-19 | A system, apparatus and method for obesity management through remote monitoring and feedback are described. In one embodiment of the present invention, periodic measurements of parameters affecting the obesity of a user are transmitted to a server through a user interactive communication device. Based on the information of the user, a customized feedback action is transmitted to the particular user. In another embodiment of the present invention an implantable apparatus having a neuro-stimulator device provides a therapeutic effect by stimulating a feeling of satiety to eating. | 1. An obesity management system, the system comprising: a server;a detecting device configured to intermittently measure various physical and eating activities of a user;a user interactive communication device configured to intermittently receive information on the physical and eating activity from the detecting device;wherein the user interactive communication device is configured to communicate with the server for sending information on physical and eating activity of the user, and receiving a feedback action. 2. The system of claim 1, wherein the information comprises physical movement of the user. 3. The system of claim 1, wherein the information comprises food intake detail of the user. 4. The system of claim 1, wherein the information comprises muscle contraction and respiration pattern detail of the user. 5. The system of claim 1, wherein the detecting device is implanted in the user. 6. The system of claim 1, wherein the detecting device is wearable by the user. 7. The system of claim 1, wherein the detecting device is self powered. 8. The system of claim 1, wherein the user interactive communication device is a personal digital assistance device. 9. The system of claim 1, wherein the user interactive communication device is a personal computing device. 10. The system of claim 1, wherein the user interactive communication device is a remotely programmable device configured to receive the feedback action inform of a script. 11. The system of claim 1, wherein the period of detecting information by the detecting device is variable. 12. The system of claim 1, wherein the feedback action is generated using an expert application linked with the server, in response to the information maintained at the server. 13. The system of claim 1, wherein an automated feedback action is generated using a virtual coach application available at the server. 14. The system of claim 12, wherein the expert application is a medical care provider application. 15. The system of claim 12, wherein the expert application is a dietician and physical trainer application. 16. The system of claim 12, wherein the expert application is a coaching program content provider application. 17. The system of claim 1, further comprising a neuro-stimulator configured to stimulate a feeling of satiety to eating. 18. An apparatus for physical and eating activity detection comprising: a microprocessor;a rechargeable power storage device;a power source;an accelerometer configured to detect physical movement of a user;a diagnostic impedance meter configured to detect food intake details of the user;a radio frequency transmitter configured to transmit information on physical movement of the user and food intake details of the user on instructions of the microprocessor. 19. The apparatus of claim 18, wherein the power source is a kinetic energy power generator to generate current for providing power to the apparatus. 20. The apparatus of claim 18, wherein the power source... | - | A61N001/00 A61B005/00 G08B023/00
US07435225 | B2 | 2008-10-14 | Method and arrangement for measuring breath gases of a patient | 2006-11-22 | Method and arrangement for measuring breath gases of a patient in which the patient is connected to a breath gas unit by using a sampling line to enable flowing a sample of the breathing gases with a pre-determined sample flow rate to the breath gas measuring unit. The sample flow rate is adjusted in case it is not within limits set by the patient's breathing rate. | 1. Method for measuring breath gases of a patient, the method comprising following step: connecting the patient to a breath gas measuring unit by using a sampling line to enable flowing a sample of the breath gases of the patient with a pre-determined sample flow rate to the breath gas measuring unit for measurement of the sample, andadjusting the sample flow rate to within limits set by the patient's breathing rate. 2. The method of claim 1 wherein the relationship between the sample flow rate and the patient's breathing rate is linear. 3. The method of claim 1 wherein the relationship between the sample flow rate and the patient's breathing rate is logarithmic. 4. The method of claim 1 wherein the relationship between the sample flow rate and the patient's breathing rate is exponential. 5. The method of claim 1 wherein the relationship between the sample flow rate and the patient's breathing rate is higher degree polynomial. 6. The method of claim 1 wherein the sample flow rate is increased while the patient's respiration rate increases and decreased while the patient's respiration rate decreases. 7. The method of claim 1 wherein the method further comprises the step of connecting the patient to a ventilator by using ventilator tubes. 8. Arrangement for measuring breath gases of a patient which arrangement comprises a sampling line and a breath gas measuring unit, the sampling line being arranged to enable flowing a sample of the patient's breath gases with a pre-determined sample flow rate to the breath gas measuring unit for measurement of the sample, the breath gas measuring unit being provided with a flow control arrangement, wherein the flow control arrangement adjusts the sample flow rate to within limits set by the patient's breathing rate. 9. The arrangement of claim 8 wherein the breath gas measuring unit comprises a water trap for separating water and mucus from the sample of the patient's breath gases, a pump for enabling the flow of the sample of the breath gases from the patient through the sampling line to the breath gas measuring unit, a gas sensor for measuring the breath gas sample, a flow measurement unit for measuring sample flow and a CPU for analyzing the results measured and for controlling the pump for adjusting the sample flow. 10. The arrangement of claim 8 wherein the arrangement further comprises a medical monitor display showing essential measuring results. 11. The arrangement of claim 10 wherein the essential measuring results comprise the gas curves, Et- and Fi-values measured and respiration rate. 12. The arrangement of claim 11 wherein the breath gas measuring unit is arranged to measure the rise and fall time from the gas curve measured to take care that the curve has a right shape. 13. The arrangement of claim 8 wherein the arrangement further comprises a ventilator and ventilator tubes, the ventilator tubes being arranged to connect the patient to the ventilator. 14. The arrangement of claim 8, wh... | General Electric Company | A61B005/08
US07559903 | B2 | 2009-07-14 | Breathing sound analysis for detection of sleep apnea/popnea events | 2007-03-28 | Apparatus for use detection of apnea includes a microphone mounted in the ear of the patient for detecting breathing sounds and a second external microphone together with an oximetric sensor. A transmitter at the patient compresses and transmits the signals to a remote location where there is provided a detector module for receiving and analyzing the signals to extract data relating to the breathing. The detector uses the entropy or range of the signal to generate an estimate of air flow while extracting extraneous snoring and heart sounds and to analyze the estimate of air flow using Otsu's threshold to detect periods of apnea and/or hypopnea. A display provides data of the detected apnea/hypopnea episodes and related information for a clinician. | 1. Apparatus for use in analysis of breathing of a patient during sleep for detection of apnea comprising: a microphone arranged to be located on the patient for detecting breathing sounds;a transmitter at the patient for transmitting signals from the breathing sounds to a remote location;a detector module for receiving and analyzing the signals to extract data relating to the breathing;the detector module being arranged to analyze the signals to generate an estimate of air flow while extracting extraneous sounds related to snoring or heart;the detector module being arranged to analyze the estimate of air flow to detect periods of apnea or hypopnea;and a display of the detected apnea or hypopnea episodes and related information for a clinician;wherein the detector module is arranged to generate an estimate of air flow during breathing by calculating a function representing the range of the signals which is defined as the log of the difference between minimum and maximum of the signal within each window of data. 2. The apparatus according to claim 1 wherein the microphone is arranged so as to be attached on the neck of the patient so as to receive sounds in the neck emanating from the breathing of the patient. 3. The apparatus according to claim 1 wherein there is provided a second microphone to collect extraneous sounds from the patient and a sensor for oximetric signals for transmission to the remote receiver. 4. The apparatus according to claim 1 wherein there is provided a second microphone arranged to receive sounds from the patient in the vicinity of the patient so as to be sensitive to snoring and wherein the detector module is arranged to use adaptive filtering to extract the signals relating to the snoring from the signals including both the breathing sounds and the snoring sounds. 5. The apparatus according to claim 1 wherein sleep apnea or hypopnea is detected by comparing the function to a threshold. 6. The apparatus according to claim 5 wherein the threshold is defined as the minimum of the Otsu's threshold and the average value of the function within each data window. 7. The apparatus according to claim 6 wherein the Otsu's threshold is defined as the threshold which maximizes the between class variance. 8. The apparatus according to claim 5 wherein the display includes a display of airflow versus time is plotted with apnea and hypopnea episodes marked in. 9. The apparatus according to claim 8wherein the display includes oximetry data plotted in association with the estimated airflow. 10. The apparatus according to claim 9 wherein the display is capable of zoom-in and zoom-out functions in the same window for both airflow and oximetry data simultaneously. 11. The apparatus according to claim 9 wherein the display is capable of playing the breathing and snoring sounds in any zoomed-in or zoomed-out data window. 12. The apparatus according to claim 9 wherein the display is capable of displaying the extracted information abou... | TR Technologies Inc. | A61B005/08 A61B007/00
US07785248 | B2 | 2010-08-31 | Steerable lesion excluding heart implants for congestive heart failure | 2008-02-19 | Devices, systems, and methods for treating a heart of a patient may make use of one or more implant structures which limit a size of a chamber of the heart, such as by deploying a tensile member to bring a wall of the heart toward (optionally into contact with) a septum of the heart. | 1. A system for treating a heart, the heart having a first chamber bordered by a septum and a wall, the heart having a second chamber separated from the first chamber by the septum, the system comprising: an implant having a septum anchor, a wall anchor, and a tension member to apply tension between the septum and wall when the implant is deployed so as to bring the wall and septum into engagement; anda deployment catheter releasably supporting at least a portion of the implant for deploying the implant in the heart, the deployment catheter comprising or receiving an adhesion inducing surface for directing energy or a material toward a region of endocardium bordering the first chamber so as to induce adhesions along the region, the tension being applicable so as to approximate the wall and the septum along the endocardium region. 2. The system of claim 1, wherein at least a portion of the implant is biodegradable or is removable while the adhesions restrain a size of the first chamber and effectively exclude scar tissue from the first chamber. 3. The system of claim 1, further comprising: at least one additional implant, the at least one additional implant being deployable separated from the implant, the implants, when deployed, providing a region of engagement extending across a lower portion of the ventricle so as to exclude lower portions of the septum and wall from the ventricle. 4. The system of claim 3, wherein at least one of the anchors comprises a plurality of arms defined by axial cuts in a tube, wherein radial expansion of the anchors is induced by bending the arms radially outwardly with the axially-oriented surface comprising a first portion of each arm and being supported by a longer angled portion of the arm. 5. The system of claim 4, further comprising a bladder material disposable within the anchor, the bladder material defining a chamber, a port in communication with the chamber, and a fill fluid introducible into the chamber through the port so as to support the axially-oriented surface. 6. The system of claim 5, wherein the fill fluid comprises a fluid that hardens after being introduced into the anchor. 7. The system of claim 5, wherein the fill fluid or the implant comprises a bioactive material releasably supported by the anchors or tension member. 8. The system of claim 3 wherein the deployment catheter comprises or receives a tissue identifier for characterization of a tissue of the first chamber. 9. The system of claim 8, wherein the tissue identifier comprises a pacing electrode to facilitate characterization of tissue. | BioVentrix (a CHF Technologies Company) | A61F001/00
US07460905 | B2 | 2008-12-02 | System for displaying vital sign data | 2006-01-12 | Each of a plurality of electroencephalographs is adapted to measure electroencephalograms of a patient and to be connected to a communication network. Each of the electroencephalographs is operable to analyze the electroencephalograms to generate electroencephalogram information. Each of a plurality of cameras is adapted to monitor a state that the patient is subjected to the electroencephalogram measurement to transmit the monitored state as image information to associated one of the electroencephalographs. A central control monitor is connected to the communication network and operable to simultaneously displaying the electroencephalogram information and a part of the image information which are transmitted from each of the electroencephalographs via the communication network. | 1. A system for displaying vital sign data, comprising: a plurality of electroencephalographs, each of which is adapted to measure electroencephalograms of a patient and to be connected to a communication network, each of the electroencephalographs being operable to generate electroencephalogram information which is information obtained by successively analyzing a plurality of electroencephalograms of uniform prescribed time intervals;a plurality of cameras, each of which is adapted to monitor a state of the patient subjected to the electroencephalogram measurement, and to transmit the monitored state as image information to an associated one of the electroencephalographs; anda central control monitor, connected to the communication network and operable to simultaneously display for a plurality of patients, in a first display region, the electroencephalogram information and a part of the image information which are transmitted from each of the electroencephalogram graphs via the communication network. 2. The system as set forth in claim 1, wherein: each of the electroencephalographs comprises a storage adapted to store the electroencephalograms and the image information; andthe central control monitor is operable to receive a request including information indicative of a time point, and to display, in a second display region, the electroencephalograms and the image information of at least the time point which are read out from the storage in one of the electroencephalographs. 3. The system as set forth in claim 1, wherein the central control monitor is operable to issue an alarm when a prescribed event is detected in at least one of the electroencephalographs. 4. The system as set forth in claim 1, wherein the image information is obtained at a first frame rate and the part of the image information is extracted at a second frame rate which is lower than the first frame rate. 5. The system as set forth in claim 1, wherein the electroencephalogram information is obtained by subjecting the electroencephalograms to frequency analysis. 6. The system as set forth in claim 5, herein the frequency analysis is fast Fourier transform. 7. The system as set forth in claim 6, wherein the electroencephalogram information is a density spectral array of the electroencephalograms. 8. The system as set forth in claim 1, wherein the electroencephalogram information has a smaller data amount than that of the electroencephalograms. 9. The system as set forth in claim 1, wherein: the electroencephalogram information has at least one channel, the number of said at least one channel being less than the number of channels of the electroencephalograms; andthe electroencephalogram information has an appearance different from that of the electroencephalograms. | Nihon Kohden Corporation | A61B005/04
US08187197 | B2 | 2012-05-29 | Cerebral perfusion monitor | 2005-06-15 | A method of estimating cerebral blood flow includes obtaining a measure of time-varying blood volume in the head, using an impedance plethysmography (102 and 104), obtaining a measure of time-varying blood volume in the scalp, and using the time-varying blood volume in the head and scalp to estimate cerebral blood flow. | 1. A method of estimating cerebral blood flow, comprising: a) obtaining a measure of time-varying blood volume in the head, using impedance plethysmography electrode units that pass current through the head;b) obtaining a measure of time-varying blood volume in the scalp using a plethysmography sensor; andc) weighting the measures of time-varying blood volume in the head and time-varying blood volume in the scalpd) estimating the cerebral blood flow from the weighted measures of time-varying blood volume in the head and time-varying blood volume in the scalp. 2. A method according to claim 1, wherein obtaining a measure of time-varying blood volume in the scalp comprises using photoplethysmography. 3. A method according to claim 1, wherein estimating the cerebral blood flow comprises estimating the relative cerebral blood flow as it changes over time. 4. A method according to claim 1, wherein the measures of time-varying blood volume are weighted to have at least approximately the same value at a time in the cardiac cycle when the blood pressure is falling. 5. A method according to claim 1, wherein the measures of time-varying blood volume are weighted to have approximately equal power spectra at frequencies for which the cross-power spectrum between the measures of time-varying blood volume is relatively high. 6. A method according to claim 1, wherein obtaining a measure of blood volume in the head using impedance plethysmography comprises: a) passing a current through the head using two current-carrying electrodes; andb) measuring a voltage across the head, associated with the current, using two voltage-measuring electrodes. 7. A method according to claim 6, also including applying to the head an annular electrode surrounding at least one of the current-carrying electrodes, and maintaining the annular electrode at a same voltage as the current-carrying electrode it surrounds, thereby suppressing radial current from said current-carrying electrode. 8. A method according to claim 6, wherein the voltage-measuring electrodes are distinct from, and substantially electrically decoupled from, the current-carrying electrodes. 9. A method according to claim 6, wherein obtaining a measure of blood volume in the head using impedance plethysmography comprises placing the two current-carrying electrodes on the left and right temples respectively. 10. A method according to claim 6, wherein obtaining a measure of blood volume in the head using impedance plethysmography comprises placing each of the two voltage-measuring electrodes on the head in a position adjacent to a different one of the current-carrying electrodes. 11. A method according to claim 10, wherein obtaining a measure of blood volume in the scalp using photoplethysmography comprises placing a photoplethysmography sensor on the head adjacent to one of the current-carrying electrodes and to the voltage-measuring electrode which is adjacent to said current-carrying electrode. 12. A unit ... | Orsan Medical Technologies Ltd. | A61B005/02 A61B005/04
US08092385 | B2 | 2012-01-10 | Fluid access interface | 2006-05-23 | The present invention relates generally to systems, apparatuses, and methods for obtaining a fluid sample from a patient. In particular, the present invention relates to a various types of fluid access interfaces for enabling contact between a patient blood sample and blood parameter sensors for the measurement of physiological parameters and blood constituents. | 1. A blood constituent measuring device comprising: at least one fluid access interface configured to automatically interface with a patient blood access line at a sampling port to access a blood sample therefrom and to separate the blood sample from the patient blood access line, said at least one fluid access interface comprising a container configured to contain the blood sample separated from the patient blood access line;a plurality of test substrates, wherein the at least one fluid access interface is configured to transfer each of a plurality of blood samples to each of the plurality of test substrates, wherein the at least one fluid access interface comprises a plurality of fluid access interfaces wherein each fluid access interface is configured to transfer a blood sample to a corresponding test substrate of the plurality of test substrates; andat least one pump element configured to direct the separated blood sample to the corresponding test substrate. 2. The device of claim 1, wherein the at least one fluid access interface is movable with respect to the sample port. 3. The device of claim 1, wherein the device is configured to transfer the blood samples to the corresponding test substrates using a predetermined schedule. 4. The device of claim 1, further comprising a housing containing said plurality of test substrates. 5. The device of claim 1, further comprising an infusion source coupled to the patient blood access line. 6. The device of claim 1, wherein the at least one pump element comprises a flexible portion configured to direct the separated blood sample to the test substrate. 7. The device of claim 6, wherein the flexible portion is configured to contract. 8. The device of claim 7, wherein the flexible portion is configured to expand. 9. The device of claim 1, wherein the container comprises a tube. 10. The device of claim 1, wherein the container comprises a valve. 11. The device of claim 1, wherein the container comprises a cylinder. 12. The device of claim 1, wherein the test substrate comprises a glucose test substrate. 13. The device of claim 1, wherein the at least one pump element is further configured to pull the blood sample from the sampling port. | IntelliDx, Inc. | A61B005/00
US08551008 | B2 | 2013-10-08 | Method and apparatus for determining vascular health conditions | 2007-10-12 | The present invention relates to assessing and measuring vascular and endothelial function. A vasostimulant is provided to a patient to stimulate hemodynamic activity in a selected extremity and vascular function is assessed by monitoring a change in a blood flow, skin temperature and/or blood oxygen content at the selected extremity and assessing the patient's vascular function based upon the monitoring. | 1. A method for assessing arterial reactivity in a patient, comprising: a) providing a vasodilating stimulant to a patient by creating an arterial occlusion at a body part of said patient using one or more pressure cuffs thereby stimulating arterial reactivity in a first region of said body part of said patient;b) simultaneously monitoring a patient's skin surface temperature at the first region and at a second control region before the arterial occlusion, during the arterial occlusion, and after the arterial occlusion is removed and until an equilibrium temperature is achieved using a temperature measurement device that does not substantially apply any one of thermal energy, positive pressure and negative pressure to the skin surface; wherein said second control region and said first region are located at contralateral body parts; andc) identifying a baseline temperature at the first region before the arterial occlusion;d) identifying a peak temperature at the first region after the arterial occlusion; ande) assessing arterial reactivity in said patient based on temperature changes at the second control region and a rebound temperature change related to said baseline temperature and said peak temperature at the first region. 2. The method of claim 1 for assessing arterial reactivity further comprising: a) creating the arterial occlusion for a predetermined period of time; andb) ceasing the occlusion after the predetermined period of time. 3. The method of claim 1, wherein providing a vasodilating stimulant comprises occluding blood flow to at least one of: (i) a leg, (ii) ankle, (iii) toe, (iv) arm, (v) wrist, and (vi) a finger. 4. The method of claim 1, wherein said monitoring occurs from a time prior to the beginning of said compression until a time after the hemodynamic activity has reached equilibrium. 5. The method of claim 1, wherein the monitoring comprises measuring changes in temperature at one of the patient's fingertips. 6. The method of claim 5, further comprising plotting the measured changes in temperature as a function of time. 7. A method for assessing arterial reactivity in a patient comprising: a) measuring a first equilibrium skin temperature at a first region of a body part of said patient using a temperature measurement device that does not substantially apply any one of thermal energy, positive pressure, and negative pressure to said first region of the patient;b) compressing arteries of said body part of the patient using a pressure cuff such that arterial blood flow in said first region of said body part is occluded;c) simultaneously measuring skin temperature at the first region and at a second control region of said body part during cuff compression, after release of the cuff and until a second temperature equilibrium is achieved; wherein the second control region does not undergo arterial blood flow occlusion said cuff compression;d) identifying a peak temperature at the first region after release of the cuff; an... | - | A61B005/02
US08540632 | B2 | 2013-09-24 | Low profile antenna for in body device | 2008-05-23 | Low profile antennas for ingestible devices, such as ingestible event markers, are provided. Aspects of the ingestible devices of the invention include low profile signal transmission antennas configured such that they are flexible or degrade after use, allowing the antenna and attached components to pass easily through the body. Embodiments of the low profile antennas are configured to emit a detectable signal upon contact with a target physiological site. Also provided are methods of making and using the devices of the invention. | 1. An ingestible event marker comprising: an identifier that emits a signal when activated at a target site, wherein the identifier comprises a signal generation element comprising a low profile antenna,wherein the low profile antenna is enclosed within an encapsulating material such that the encapsulating material holds the low profile antenna in a first configuration prior to and during signal transmission, andwherein the encapsulating material is dissolvable after ingestion and after signal transmission to release the low profile antenna to a second configuration. 2. The ingestible event marker according to claim 1, wherein the low profile antenna is more flexible when in the second configuration than when in the first configuration. 3. The ingestible event marker according to claim 2, wherein the flexible second configuration allows the antenna to pass easily through the body after signal transmission. 4. The ingestible event marker of claim 1, wherein the low profile antenna comprises a first and a second transmission element separated by flexible conductor. 5. The ingestible event marker according to claim 4, wherein the first and second transmission elements are electrodes. 6. The ingestible event marker according to claim 5, wherein the low profile antenna comprises a coil in the first configuration. 7. The ingestible event marker according to claim 6, wherein the low profile antenna is linear in the second configuration. 8. The ingestible event marker according to claim 7, wherein each antenna ranges in length from 100 μm to 10 cm. 9. The ingestible event marker according to claim 5, wherein the low profile antenna forms a loop in the second configuration. 10. The ingestible event marker according to claim 1, wherein the identifier is activated upon contact with fluid at the target site. 11. The ingestible event marker according to claim 1, wherein the target fluid is stomach fluid. 12. The ingestible event marker according to claim 1, wherein the identifier comprises an integrated circuit. 13. The ingestible event marker according to claim 1, wherein the identifier comprises a power source. 14. The ingestible event marker according to claim 13, wherein the power source comprises a battery. 15. The ingestible event marker according to claim 1, wherein the low profile antenna is configured to conform to ingestible pill dimensions. 16. The ingestible event marker according to claim 1, further comprising an active agent stably associated with the signal generation element. 17. The ingestible event marker according to claim 1, wherein the low profile antenna has a greater effective area when in the first configuration than when in the second configuration. 18. The ingestible event marker according to claim 1, wherein the encapsulating material is at least a part of a pill wall. 19. The ingestible event marker according to claim 1, wherein the encapsulating material is coupled to a surface of a pill. 20. A system comprising... | Proteus Digital Health, Inc. | A61B005/07
US08433111 | B2 | 2013-04-30 | Method for image data recording with a medical modality designed for image data recording and associated medical modality | 2008-03-20 | A method for image data recording with a medical modality is provided. Protocols for image data recording using the at least two different image recording techniques are provisioned in a common database. The protocols are organized on the basis of body regions and/or clinical pictures. A body region and/or a clinical picture are selected and at least one protocol is displayed as a function of the selection. At least one display protocol for image data recording from the database is selected by a user and/or automatically, and image data is recorded as a function of the at least one selected protocol. | 1. A method for image data recording with a medical modality, which is designed for image data recording using at least two different image recording techniques, the method comprising: provisioning image data recording protocols for at least two different image recording techniques in a common database, the protocols being organized on the basis of at least one of body regions and clinical pictures;selecting at least one of a body region and a clinical picture;displaying at least one protocol as a function of the selected at least one of the body region and clinical picture;selecting at least one display protocol for image data recording from the database; andrecording image data as a function of the at least one selected protocol. 2. The method as claimed in claim 1, wherein the database is stored in a console, the console being designed for at least one of patient registration and control of the image data recording. 3. The method as claimed in claim 1, wherein at least one of the protocol, data, and selection fields associated with at least one of the protocol and data are displayed by a software interface on a display. 4. The method as claimed in claim 1, wherein at least one image data recording protocol for at least one of positron-emission tomography, computed tomography, single-photon-emission computed tomography, magnetic resonance imaging, and X-ray based techniques is provided in the database. 5. The method as claimed in claim 1, wherein the at least one protocol is provided in a tree-like structure. 6. The method as claimed in claim 1, further comprising: providing the at least one protocol taking into account at least one relationship resulting from the at least two image recording techniques. 7. The method as claimed in claim 6, wherein the at least one relationship is at least one relationship for recording of positron-emission image data, the at least one relationship including at least one of, a requirement for a previous magnetic resonance image data record to allow calculation of an attenuation correction,a requirement to record a magnetic resonance scout image data for planning image data recording by positron-emission tomography,a relationship between a reconstruction volume for the positron-emission tomography and measurement field changes during a magnetic resonance image data recording, anda definition of automatic multiplanar reconstructions of positron-emission image data using two-dimensional magnetic resonance image data. 8. The method as claimed in claim 1, wherein the at least one protocol is provided as a function of a selection of at least one protocol in at least one of combinations and sequences, which are governed by the at least two different image recording techniques. 9. The method as claimed in claim 8, wherein a protocol with a long measurement time is provided for selection of a protocol for image data recording by positron-emission tomography, the positron-emission tomography having at least a lon... | Siemens Aktiengesellschaft | G06F017/30 A61B005/00
US08386027 | B2 | 2013-02-26 | Skin permeation device for analyte sensing or transdermal drug delivery | 2008-04-25 | Devices, systems, kits and methods for increasing the skin's permeability controlled by measured skin electrical parameter are described herein. They may be used for transdermal drug delivery and/or analyte extraction or measurement. The controlled abrasion device contains (i) a hand piece, (ii) an abrasive tip, (iii) a feedback control mechanism, (iv) two or more electrodes, and (v) an electrical motor. The feedback control mechanism may be an internal feedback control mechanism or an external feedback control. The kit contains the controlled abrasion-device, one or more abrasive tips, optionally with a wetting fluid. The method for increasing the skin's permeability requires applying the controlled abrasion device to a portion of the skin's surface for a short period of time, until the desired level of permeability is reached. Then the abrasion device is removed, and a drug delivery composition or device or an analyte sensor is applied to the treated site. | 1. A controlled abrasion device comprising a hand piece, an abrasive tip, a feedback control mechanism, and an electrical motor, wherein the abrasive tip does not contain a micro-needle,wherein the feedback control mechanism comprises (a) a source electrode, (b) a return electrode and (c) a controller,wherein the source electrode is located in the abrasive tip,wherein the return electrode is located at the proximal end of the hand piece,wherein the abrasive tip connects directly or indirectly to the motor such that the motor is able to cause the abrasive tip to move, and wherein the abrasive tip is attachable to and removable from the proximal end of the hand piece, andwherein the source electrode and the return electrode are in electrical communication with the controller, andwherein the feedback control mechanism measures the conductance through the skin in real time and calculates in real time the rate of change in the conductance over time when the device is applied to the skin to control the level of skin abrasion. 2. The device of claim 1, wherein the feedback control mechanism is an internal feedback control mechanism. 3. The device of claim 1, wherein the abrasive tip comprises a wetting fluid. 4. The device of claim 1, wherein the abrasive tip comprises a material selected from the group consisting of conductive and non-conductive materials. 5. The device of claim 4, wherein the abrasive tip comprises a conductive material, and wherein the abrasive tip is the source electrode. 6. The device of claim 5, wherein the conductive material comprises perforations. 7. The device of claim 5, wherein an outer wall of the proximal end of the hand piece comprises the return electrode. 8. The device of claim 1, wherein the abrasive tip is a disposable abrasive tip, and wherein the device further comprises a cup that surrounds the abrasive tip. 9. The controlled abrasion device of claim 1, wherein the return electrode surrounds the abrasive tip. 10. The controlled abrasion device of claim 9, wherein the abrasive tip comprises a wetting fluid. 11. The device of claim 1, wherein the electrical motor is a rotary, direct current (DC) motor. 12. The device of claim 1, further comprising a spring loaded motor shaft that provides a downward force on the abrasive tip, when the abrasive tip is in contact with the skin surface. 13. A method for reducing the impedance of a tissue site comprising applying an abrasive tip of a controlled abrasion device to the tissue site, wherein the device comprises a hand piece, the abrasive tip, a feedback control mechanism, and an electrical motor, wherein the feedback control mechanism comprises (a) a source electrode, (b) a return electrode and (c) a controller,wherein the source electrode is located in the abrasive tip, and wherein the return electrode is located at the proximal end of the hand piece,wherein the abrasive tip connects directly or indirectly to the motor such that the motor is able to cause the a... | Echo Therapeutics, Inc. | A61B005/05 A61B005/00 A61B017/20 A61N001/30 A61N001/00
US08798711 | B2 | 2014-08-05 | Shielding of catheter handle | 2008-07-09 | Electrical apparatus includes a probe, having a proximal end and a distal end. The probe includes a sensor, which outputs a sensor signal, and a first connector at the proximal end of the probe, electrically coupled at least to the sensor. A probe adapter includes a second connector, which is arranged to mate with the first connector, and a third connector, for coupling to a console. A shield includes a material of high magnetic permeability, which is configured to enclose an internal volume containing the first and second connectors when the probe is connected to the probe adapter. | 1. Electrical apparatus, comprising: a probe, having a proximal end and a distal end, the probe comprising: a sensor, which outputs a sensor signal; anda first connector at the proximal end of the probe, electrically coupled at least to the sensor;a probe adapter, comprising: a second connector at a distal end of the probe adapter, which is arranged to mate with the first connector of the probe; anda third connector located proximal from the first connector and the second connector, and coupled to a console; anda shield, comprising a mu-metal material of high magnetic permeability, the shield being made of unitary tubular form which is configured to enclose an internal volume containing the first connector and second connectors when the probe is connected to the probe adapter, the unitary tubular form of the shield preventing leakage of electromagnetic fields from the first connector and the second connector when the probe is connected to the probe adapter. 2. The apparatus according to claim 1, wherein the sensor further comprises at least one amplifier, which is contained in the internal volume and is coupled to process the sensor signal so as to output an amplified signal via the third connector to the console. 3. The apparatus according to claim 1, wherein the sensor comprises a position sensor, which is operative to generate the sensor signal responsively to a magnetic field applied externally to the body. 4. The apparatus according to claim 1, wherein the probe comprises a catheter, and wherein the distal end is configured for insertion into a heart of the subject. 5. The apparatus according to claim 1, wherein the probe comprises a twisted pair of wires, which are coupled to convey the sensor signal from the sensor to the first connector, and wherein the wires in the twisted pair are separated only within the internal volume of the shield for connection to respective pins of the first connector. 6. The apparatus according to claim 1, and further comprising a handle, which contains the second connector and the shield, and wherein the proximal end of the probe is configured to be plugged into and removed from the handle. 7. A method for producing an electrical device, comprising: providing a probe, having a proximal end and a distal end, which comprises a sensor, which outputs a sensor signal, and a first connector at the proximal end of the probe, electrically coupled at least to the sensor;providing a second connector, for use in a probe adapter to mate with the first connector of the probe;coupling a third connector located proximal of the second connector to connect the probe adapter to a console; andconfiguring a shield, comprising a mu-metal material of high magnetic permeability and being made of unitary tubular form, so as to enclose an internal volume containing the first connector and the second connector&amp; when the probe is connected to the probe adapter, the unitary tubular form of the shield preventing leakage of electr... | Biosense Webster, Inc. | A61B005/05
US08571632 | B2 | 2013-10-29 | Open architecture imaging apparatus and coil system for magnetic resonance imaging | 2008-06-13 | Apparatus and method for using radio frequency coil systems for magnetic resonance imaging within an open architecture apparatus is provided. The MRI coil system includes a support structure with an open architecture in which secondary support structures, compression systems and plates containing RF coil systems may be introduced. These structures and RF coils can be moved relative to the patient, or removed entirely from the system. In one embodiment the system consists of a tabletop coil system, while another embodiment consists of a dedicated stretcher design. | 1. An RF coil for use with an MRI scanner having a main magnetic field to perform MRI in a volume of interest comprising: a first RF coil array configured to be adjacent to the volume of interest comprising a first plurality of RF coil elements, the first RF coil array having a reception sensitivity to a B1 field substantially oriented in a first direction throughout the volume of interest wherein the first direction is configured to be orthogonal to the main magnetic field of the MRI scanner; anda second RF coil array, configured to be positionable on the opposite side of the volume of interest, the second RF coil array moveably coupled to the first RF coil array to enable the first RF coil array and second RF coil array to be moved together for immobilization of the volume of interest there between, the second RF coil array comprising a second plurality of RF coil elements, the second RF coil array having a reception sensitivity to a B1 field oriented in a direction substantially orthogonal to the first direction and to the main magnetic field of the MRI scanner throughout the volume of interest. 2. The RF coil of claim 1, wherein: the first plurality of RF coil elements is disposed in a substantially coplanar geometry;the second plurality of RF coil elements is disposed in a substantially coplanar geometry; andthe second RF coil array is oriented substantially parallel to the first RF coil array. 3. The RF coil of claim 2, wherein the first RF coil array comprises at least one of a loop coil elements, a saddle coil element, a butterfly coil element, a dipole, a solenoid coil, and a transmission line coil element. 4. The RF coil of claim 2, wherein the first RF coil array comprises a first loop coil element and a second loop coil element, the first and second loop coil elements being arranged in an overlapping geometry. 5. The RF coil of claim 4, wherein the first and second loop coil elements are overlapped along a first axis, each of the first and second loop coil elements has a long axis and a short axis, and the long axis of each coil element is aligned along a second axis substantially perpendicular to the first axis, wherein when positioned for use with an MRI scanner, the first axis is substantially parallel to the main magnetic field of the MRI scanner main magnetic field and the second axis is substantially perpendicular to the main magnetic field. 6. The RF coil of claim 4, wherein the first and second loop coil elements are overlapped along a first axis, each of the first and second loop coil elements has a long axis and a short axis, and the long axis of each coil element is aligned along a second axis substantially perpendicular to the first axis, wherein when positioned for use with an MRI scanner, the first axis is substantially perpendicular to the magnetic field of the MRI scanner and the second axis is substantially parallel to the magnetic field of the MRI scanner. 7. The RF coil of claim 2, wherein the first RF coil ar... | Hologic, Inc. | A61B005/055 A47B013/00
US20080294061 | A1 | 2008-11-27 | Health care gaming device and method using the same | 2007-08-20 | A health care gaming device includes: a case, which can be placed on a platform or be handheld by an user; at least one communicating tube provided in the case; an inflow/outflow unit connected to the communicating tube, for the user to exhale and inhale gas; at least one sensor connected to the communicating tube, for sensing gas inflow and outflow and detecting gas flow; a processing unit connected to the sensor, for measuring pulmonary function and executing games; a display unit connected to the processing unit; and a keying unit connected to the processing unit, wherein the sensor generates a sensed deformation due to gas flowing in the communicating tube; the processing unit analyzes the sensed deformation and converts the analytical result to measurement data for pulmonary function or uses the analytical result as game data for proceeding games; and the display unit displays the measurement data and game pictures. The invention also discloses a method using such health care gaming device. | 1. A health care gaming device, which measures pulmonary function and uses games for lung training, comprising: a case, which can be placed on a platform or handheld by an user;at least one communicating tube provided in the case;an inflow/outflow unit connected to the communicating tube, for the user to exhale and inhale gas;at least one sensor connected to the communicating tube, for sensing gas inflow and outflow and detecting gas flow;a processing unit connected to the sensor, for measuring pulmonary function and executing games;a display unit connected to the processing unit; anda keying unit connected to the processing unit, for the user to select functions of the gaming device, set game parameters, and play games;wherein the sensor generates a sensed deformation due to gas flowing in the communicating tube; the processing unit analyzes the sensed deformation and converts the analytical result to measurement data for pulmonary function or uses the analytical result as game data for proceeding games; and the display unit displays the measurement data and game pictures. 2. The health care gaming device as described in claim 1, wherein the measurement data of pulmonary function includes forced vital capacity, forced expiratory volume in the first second, and total forced expiratory volume. 3. The health care gaming device as described in claim 1, wherein the inflow/outflow unit comprises a blowing device having a converging outlet. 4. The health care gaming device as described in claim 1, wherein the inflow/outflow unit comprises: an inhalation device having a diverging outlet and a flexible conduit. 5. The health care gaming device as described in claim 1, wherein the inflow/outflow unit is a detachable assembly. 6. The health care gaming device as described in claim 5, wherein the inflow/outflow unit is able to accept inflow of other gases. 7. The health care gaming device as described in claim 5, wherein the inflow/outflow unit is worn on the ear or the head of the user. 8. The health care gaming device as described in claim 1, wherein the sensor is disposed inside the communicating tube. 9. The health care gaming device as described in claim 1, wherein the sensor is an elastic or piezoelectric sensing element. 10. The health care gaming device as described in claim 1, wherein the connection between the display unit and the processing unit is wireless. 11. The health care gaming device as described in claim 1, wherein the display unit is a computer, a television or a display screen. 12. The health care gaming device as described in claim 1, wherein the game parameter setting includes: setting gas inflow/outflow data to be game data, for controlling the moves of on-screen targets in the game. 13. The health care gaming device as described in claim 1, further comprising: a hand-pressing device, which is a hollow compressible container and connected to the communicating tube, for game operation and hand exercise. 14. The health ca... | - | A61B005/091 A63F009/24
US20080300493 | A1 | 2008-12-04 | Optical microprobe for blood clot detection | 2008-06-06 | The invention is devices and related methods for detecting blood clots in a blood vessel. An optical microprobe is configured to illuminate a blood vessel with electromagnetic radiation corresponding to the near-infrared portion of the electromagnetic spectrum. The optical microprobe has a pair of fiber optic strands configured for transmission spectroscopy to obtain the absorption spectrum generated by the components within the blood vessel. Because blood clots generate a detectable and unique spectrum, the presence or absence of the blood clot is determined by examining the blood vessel absorption spectrum. A specially-designed holder is configured to stably position the optical microprobe relative to the blood vessel and is used to facilitate precise blood clot detection along a length of blood vessel. | 1. An optical microprobe for non-invasively detecting blood clots in a blood vessel by transmission spectroscopy, said optical microprobe comprising: a. an optical source for generating electromagnetic radiation having a wavelength range corresponding to the wavelength of near infrared light and capable of being absorbed by a blood clot;b. a first fiber optic strand capable of transmitting the electromagnetic radiation generated by the optical source, the strand having a proximal end and a distal end, wherein the proximal end is optically connected to the optical source, and the distal end is capable of illuminating the blood vessel with the electromagnetic radiation;c. a second fiber optic strand capable of transmitting the electromagnetic radiation generated by the optical source, the strand having a proximal end and a distal end, wherein the distal end is capable of collecting the electromagnetic radiation that has illuminated and passed through the blood vessel; andd. a holder having a first holding arm connected to the first fiber optic strand distal end and a second holding arm connected to the second fiber optic strand distal end, wherein the holder stably positions the distal portions of the first and second fiber optic strands in a diametrically opposed configuration and separated by a separation distance. 2. The optical microprobe of claim 1 wherein each of the holding arms has a bottom end, the optical microprobe further comprising: a. a first holding tip connected to the first holding arm bottom end; andb. a second holding tip connected to the second holding arm bottom,wherein, the first holding tip is connected to the distal end of the first fiber optic strand, and the second holding tip is connected to the distal end of the second fiber optic strand. 3. The optical microprobe of claim 2, wherein the holding tip further comprises an orifice, and the distal end of the fiber optic strand is disposed within the orifice. 4. The optical microprobe of claim 1, further comprising means for selecting the separation distance. 5. The optical microprobe of claim 4, wherein the separation distance is selected from a range of about 0.5 mm to about 2 cm. 6. The optical microprobe of claim 1, further comprising a micromanipulator connected to the holder for controllable positioning of the distal ends of the fiber optic strands. 7. The optical microprobe of claim 1, further comprising an optical detector optically connected to the second fiber optic strand proximal end for detecting the electromagnetic radiation collected by the second fiber optic strand distal end. 8. The optical microprobe of claim 7, further comprising an analyzer for determining the intensity of electromagnetic radiation at a wavelength or wavelength range corresponding to the wavelength absorbed by a blood clot. 9. The optical microprobe of claim 8, wherein the wavelength range is selected from between about 600 nm and 1000 nm. 10. The optical microprobe of claim 8, wh... | - | A61B005/02
US20080249375 | A1 | 2008-10-09 | Arrangement and Method for Evaluating Operational Effectiveness of Implantable Medical Electrode Leads for Different Lead Placements | 2004-11-02 | In a method and an arrangement for evaluating operational effectiveness of an implantable medical device for different lead placements associated with the medical device, a measuring unit records signals that are characteristic of cardiac activity at respectively different lead positions, and these signals are stored. A processor accesses the stored signals and, from the stored signals, determines a measure of cardiac activity at each of the lead positions. The recorded signals may be intracardiac ECG signals, surface ECG signals, heart sound signals obtained from a microphone, or impedance signals. The lead position at which the best hemodynamic behavior of the heart is identified from the analysis of the stored signals, and is determined as being the optimum site for placement of the electrode leads. | 1-41. (canceled) 42. In a device comprising a plurality of electrode leads configured to interact in vivo with tissue of the patient, during said electrode leads being respectively, selectively positionable at a plurality of different positions in the patient during an implantation procedure, the improvement of an arrangement for evaluating effectiveness of operation of said electrode leads with said electrode leads placed at respectively different positions in the patient, said arrangement comprising: a measuring unit that obtains samples, with said electrode leads being respectively placed at different ones of said positions, selected from the group consisting of impedance samples, microphonic heart sound samples, surface ECG samples, and intracardiac ECG samples;a storage unit connected to said measuring unit that stores said samples; anda processor having access to said storage unit that analyzes said samples and, from said samples, identifies one of said positions of said electrode leads that produces most favorable hemodynamic behavior of the heart of the patient. 43. A device as claimed in claim 42 wherein said measuring unit obtains said impedance samples, and wherein said processor analyzes said impedance samples and determines a rate of change in impedance represented by said samples and, from said rate of change of impedance, identifies, as said one of said positions, a position of said electrode leads at which a highest ventricular contractibility occurs. 44. A device as claimed in claim 42 wherein said measuring unit obtains said impedance samples, and wherein said processor identifies a maximum and a minimum of impedance represented by said impedance samples, and identifies said one of said positions as a position of said electrode leads at which a highest ventricular ejection fraction, determined from said maximum and said minimum of said impedance, occurs. 45. A device as claimed in claim 42 wherein said measuring unit obtains said impedance samples, and wherein said processor identifies, as said one of said positions, a position of said electrode leads at which a most synchronized contraction of right and left chambers of the heart occurs, determined from said impedance. 46. A device as claimed in claim 42 comprising a stimulation pulse generator connected to said electrode leads that emits stimulation pulses that are delivered to said tissue via said electrode leads to pace the heart of the patient, and wherein said measuring unit obtains said impedance samples, and wherein said processor determines an impedance of the heart from said impedance samples and determines a pre-ejection period from said impedance as a time interval starting with a paced QRS complex or a ventricular stimulation and ending when impedance in the ventricle reaches a predetermined value, and identifies said one of said lead positions as a position of said electrode leads at which a minimum of said pre-ejection period occurs. 47. A device as claimed ... | - | A61B005/02
US20080306381 | A1 | 2008-12-11 | Method for optimizing the time sequence of examinations with the aid of a tomography system having a control unit by using contrast agents, a tomography system and contrast agent injector | 2008-05-23 | A method, a tomography system and a contrast agent injector for optimizing the time sequence of contrast agent examinations are disclosed, an automatically controlled text announcement system being used for instructing the behavior of a patient before or during the examination, and it being possible for a multiplicity of different announcement texts of different announcement duration to be temporally triggered and announced by the text announcement system. In at least one embodiment, on the basis of the announcement duration of one announcement text to be announced during the examination, the temporal behavior of this text relative to the work sequence of the contrast agent injector and/or of the tomography system is calculated, and inserted in the timing of the contrast agent injector via an electronic connection between the tomography system and the contrast agent injector is one time pause that leads to a synchronization of the work sequence of the contrast agent injector with the announcement duration of the current text announcement in such a way that the text announcement ends at a predetermined triggering instant of the combined work sequence of the tomography system and contrast agent injector. | 1. A method for optimizing the time sequence of examinations with the aid of a tomography system having a control unit by using contrast agents, the method comprising: setting a prescribed time controlled work sequence for operation of an electronically controlled contrast agent injector, the contrast agent injector being connected in control terms to the tomography system at least in such a way that a scanning procedure is started upon actuation of a starting apparatus at least one of at the contrast agent injector and at the tomography system;setting a prescribed time controlled work sequence for the operation of the tomography system, an automatically controlled text announcement system being used for instructing the behavior of a patient before or during the examination, and a multiplicity of different announcement texts of different announcement duration being temporally triggered and announced by the text announcement system, the temporal work sequences of a contrast agent injector and tomography system being interlinked;calculating, on the basis of the known announcement duration of at least one announcement text currently to be announced during the planned examination, the temporal behavior of this text relative to the work sequence of at least one of the contrast agent injector and the tomography system; andinserting, at least in the timing of the contrast agent injector via an electronic connection between the tomography system and the contrast agent injector, at least one time pause that leads to a synchronization of the work sequence of the contrast agent injector with the announcement duration of the current text announcement in such a way that the text announcement ends at a predetermined triggering instant of the combined work sequence of the tomography system and contrast agent injector. 2. The method as claimed in claim 1, wherein, when calculating the temporal behavior of the announcement text relative to the work sequence of at least one of the contrast agent injector and the tomography system, the announcement duration of a starting text announced at the beginning of the examination is compared with the prescribed time interval between the beginning of the injection and the start of a monitoring phase, and in the event of a longer lasting announcement duration, there is inserted, between the receipt of the start command and the beginning of the injection, a pause that corresponds to the difference between the announcement duration and the time interval between the beginning of the injection and the start of the monitoring phase. 3. The method as claimed in claim 1, wherein, for an announcement of a text between a first and a second triggering event in the combined work sequence of the tomography system and contrast agent injector, the time between the two triggering events is compared with the announcement duration of the announcement text, and for a longer lasting announcement text an additional pause, corresponding to the ... | - | A61B005/05
US20080242951 | A1 | 2008-10-02 | Effective low-profile health monitoring or the like | 2007-05-29 | Systems, methods, computer program products, and media are described for obtaining data from occupational or leisure intercommunication at least between a device and a user and signaling a data distillation indicating a health status change of the user at least partly based on the data from the occupational or leisure intercommunication; or for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user and applying one or more data extraction criteria to the health-status-indicative data surreptitiously captured from the interaction between the device and the user. | 1. A method comprising: receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user; andapplying one or more data extraction criteria to the health-status-indicative data surreptitiously captured from the interaction between the device and the user. 2-10. (canceled) 11. The method of claim 1 in which receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user comprises: receiving the health-status-indicative data substantially contemporaneously with the interaction. 12-17. (canceled) 18. A system comprising: means for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user; andmeans for applying one or more data extraction criteria to the health-status-indicative data surreptitiously captured from the interaction between the device and the user. 19-34. (canceled) 35. A system comprising: circuitry for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user; andcircuitry for applying one or more data extraction criteria to the health-status-indicative data surreptitiously captured from the interaction between the device and the user. 36. The system of claim 35 in which the circuitry for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user comprises: circuitry for obtaining an indication of an activity status change of an application program. 37. The system of claim 35 in which the circuitry for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user comprises: circuitry for detecting one or more anomalous performances from the interaction between the device and the user. 38. The system of claim 35 in which the circuitry for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user comprises: circuitry for obtaining at least a portion of the health-status-indicative data from occupational intercommunication between the device and the user. 39. The system of claim 35 in which the circuitry for receiving health-status-indicative data surreptitiously captured from an interaction between a device and a user comprises: circuitry for receiving other data captured from an interaction between the device and another user. 40. The system of claim 35 in which the circuitry for applying one or more data extraction criteria to the health-status-indicative data surreptitiously captured from the interaction between the device and the user comprises: circuitry for presenting a comparison between other data and the health-status-indicative data surreptitiously captured from the interaction between the device and the user. 41. The system of claim 35 in which the circuitry for applying one or more data extraction criteria to the health-status-indicative data surreptitiou... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00
US20080287754 | A1 | 2008-11-20 | System for Processing Decaying Periodic Physiological Signals | 2008-08-01 | A system for processing decaying periodic physiological signals comprises at least a sensor, a physiological signal-digitizing device, and a digital signal processing unit to establish a parameterized database according to Pji, Sji, the relationship between Pji and Sji and the change with respect to time or locations of the sensors on the human body for evaluating the physiological or mental conditions. The system is applicable to processing and utilizing the decaying periodic physiological signals, such as heartbeats, respiration and ECG. | 1. A system for processing decaying periodic physiological signals including: at least a sensor, for sensing the decaying periodic physiological signals of a human body;a physiological signal-digitizing device, for receiving and digitizing the decaying periodic physiological signals; anda digital signals processing unit, for digitalizing and decomposing decaying periodic physiological signals D of the human body into poles P and residuals S, wherein the poles P represent the decay rate of signals, the residuals S represent the contribution value of the P, and by taking j as the serial number of signals, Dj represent periodic signals with different characters, or the same periodic signals in different measuring time or locations, and Dj, P and S satisfy the following formula:Dj=∑i=1mSij-Pijwherein m represents the maximum number of poles from the decaying periodic physiological signals to be decomposed and is collected from a human body, Pji represents the ith pole of the jth periodic human signals or jth location of the sensor, and Sji represents the contribution value or residual value of the pole; thus, obtaining the values of the poles Pji and residuals Sji by decomposing the decaying periodic physiological signals by using a numerical method. 2. The system for processing decaying periodic physiological signals according to claim 1 further comprising a parameterized database that is established according to Pji, Sji, the relationship between Pji and Sji and the change with respect to time or locations by the numerical method. 3. The system for processing decaying periodic physiological signals according to claim 1 further comprising an evaluating unit for receiving the data from the parameterized database and for outputting an evaluating result for estimating the physiological or mental conditions of said human body. 4. The system for processing decaying periodic physiological signals according to claim 1 further comprising an external stimulator for stimulating the body to generate decaying periodic physiological signals or change the natural decaying periodic physiological signals, wherein the decaying periodic physiological signals are obtained from said human body exerted by the external stimulator and then the decaying periodic physiological signals are collected and digitized. 5. The system for processing decaying periodic physiological signals according to claim 4, wherein the external stimulation generated by said external stimulator includes voltage, electromagnetic waves, ultrasonic waves, heat, or pressure. 6. The system for processing decaying periodic physiological signals according to claim 1, wherein the value of Pji or Sji are used as the signatures of the physiological or mental conditions of said human body. 7. The system for processing decaying periodic physiological signals according to claim 1, wherein the digital signals processing unit further comprises a first calculating unit for calculating ratios or statisti... | - | A61B005/00
US07459144 | B2 | 2008-12-02 | Process for preparation of MR contrast agents | 2004-12-02 | The present invention provides a process for the preparation of an MR contrast agent, said process comprising: i) obtaining a solution in a solvent of a hydrogenatable, unsaturated substrate compound and a catalyst for the hydrogenation of said substrate compound;ii) introducing said solution in droplet form into a chamber containing hydrogen gas (H2) enriched in para-hydrogen (p-1H2) and/or ortho-deuterium (o-2H2) whereby to hydrogenate said substrate to form a hydrogenated imaging agent;iii) optionally subjecting said hydrogenated imaging agent to a magnetic field having a field strength below earth's ambient field strength;iv) optionally dissolving said imaging agent in an aqueous medium;v) optionally separating said catalyst from the solution of said imaging agent in said aqueous medium;vi) optionally separating said solvent from the solution of said imaging agent in said aqueous medium; andvii) optionally freezing the solution of said imaging agent in said aqueous medium. | 1. A hydrogenation apparatus comprising a hydrogenation chamber having a liquid outlet into a conduit leading to a liquid droplet generator inlet to a solvent removal chamber, said hydrogenation chamber having a hydrogen inlet and a solution inlet provided with a further liquid droplet generator,said conduit including a catalyst removal chamber between said hydrogenation chamber and said solvent removal chamber and being provided with a liquid inlet, said solvent removal chamber being provided with a gas outlet and with a liquid outlet. 2. An apparatus of claim 1, wherein said hydrogenation apparatus is further provided with magnetic shielding such that the magnetic field within at least part of said hydrogenation chamber and/or within at least part of said conduit is &lt;50 μT. 3. An apparatus of claim 2, wherein said magnetic field is &lt;1 μT. 4. An apparatus of claim 2 An, wherein said magnetic field is &lt;0.1 μT. 5. An apparatus of claim 1, wherein said conduit is provided with a liquid inlet between said hydrogenation chamber and said catalyst removal chamber. | GE Healthcare AS | A61B005/055 B01J010/00
US07448999 | B1 | 2008-11-11 | Trans-septal intra-cardiac lead system | 2005-02-07 | An apparatus for and method of measuring pressure through a septum in a patient's heart is disclosed. A lead inserted into the right side of a heart is routed through the septum to gain access to the left side of the heart. The lead includes an attachment structure that secures the lead to one or both of the septal walls. The attachment structure may include at least one protruding tine, membrane, inflatable balloon, involuted spiral or J-lead that engage one or more sides of the septum. The lead also includes one or more sensors for measuring cardiac pressure on the left side of the heart and, as necessary, the right side of the heart. | 1. A method of measuring pressure in a patient comprising: providing a lead with at least one sensor;positioning the lead at a location in the patient at which a thoracic pressure or approximate thoracic pressure may be measured;measuring thoracic pressure or approximate thoracic pressure using the at least one sensor;measuring pressure at least one other location in the patient to provide a relative thoracic pressure measurement or a relative approximate thoracic pressure measurement;wherein positioning the lead comprises:routing the lead into the coronary sinus of the patient;piercing a coronary sinus wall of the patient; androuting the lead through the coronary sinus wall and into a pericardial space of the patient. 2. The method of claim 1 comprising using the at least one sensor to measure pressure in the pericardial sac of the patient to obtain an approximate thoracic pressure measurement. 3. The method of claim 1 comprising routing the lead through the pericardial sac into a thoracic space of the patient and using the at least one sensor to measure thoracic pressure. | Pacesetter, Inc. | A61B005/02
US07449002 | B1 | 2008-11-11 | Steerable guide wire for delivering an implantable medical device | 2006-04-06 | A guide wire assembly for facilitating placement of a medical device comprises (a) a longitudinally extending guide wire comprising a plurality of wire portions disposed adjacent to each other, each of the plurality of wire portions having a distal extremity and a proximal extremity, the distal extremities of the plurality of wire portions being joined, and (b) an actuator operatively associated with the proximal extremities of the plurality of wire portions for applying longitudinally directed, differential push-pull forces to the proximal extremities for deflecting a distal end portion of the guide wire. | 1. A guide wire assembly for facilitating placement of a medical device, the assembly comprising: a longitudinally extending guide wire comprising a plurality of wire portions disposed adjacent to each other, each of the plurality of wire portions having a distal extremity and a proximal extremity, the distal extremities of the plurality of wire portions being joined, each of the plurality of wire portions having an outer surface extending from the distal extremity to the proximal extremity, a substantial longitudinal length of the outer surfaces of the plurality of wire portions are configured to abut each other; andan actuator operatively associated with the proximal extremities of the plurality of wire portions for applying longitudinally directed, differential push-pull forces to said proximal extremities for deflecting a distal end portion of the guide wire. 2. The assembly of claim 1 wherein: the distal extremities of the plurality of wire portions are joined by a weld. 3. The assembly of claim 1 wherein: the distal extremities of the plurality of wire portions are joined by a cap encapsulating said distal extremities. 4. The assembly of claim 1 wherein: the distal extremities of the plurality of wire portions are joined by means of a filament wound about the distal extremities. 5. The assembly of claim 1 wherein: the actuator comprises a pivotable member operatively associated with the proximal extremities of the plurality of wire portions for applying said push-pull forces. 6. The assembly of claim 1 further comprising: a sheath disposed about said guide wire, the deflectable distal end portion of the guide wire projecting from a distal extremity of the sheath. 7. The assembly of claim 1 wherein: the plurality of wire portions comprise at least two wire portions. 8. The assembly of claim 1 wherein: each wire portion of the plurality of wire portions has a non-circular cross section, wherein the wire portions provide differential deflections of the distal end portion of the guide wire along different deflection axes. 9. The assembly of claim 1 wherein: the guide wire is segmented in cross section to define the plurality of wire portions. 10. The assembly of claim 1 wherein: the plurality of wire portions have a combined circular cross sectional shape. 11. A guide wire assembly for facilitating placement of a medical device, the assembly comprising: a guide wire comprising (a) a first longitudinally extending wire portion having an outer surface extending from a distal extremity to a proximal extremity and (b) a second longitudinally extending wire portion having an outer surface extending from a distal extremity to a proximal extremity, said second wire portion being disposed adjacent to said first wire portion such that the outer surface of the first longitudinally extending wire portion abuts the outer surface of the second longitudinally extending wire portion throughout a substantial longitudinal length of the first and secon... | Pacesetter, Inc. | A61B005/00
US20080319348 | A1 | 2008-12-25 | DIAGNOSIS CATHETER FOR INTERSTITIAL CYSTITIS | 2008-06-05 | Provided is a diagnosis catheter 101 for an interstitial cystitis wherein a catheter 101 is connected to a current perception threshold inspection apparatus 4 to diagnose the interstitial cystitis, comprising: a cylindrical catheter body 111 made of a soft flexible material and including a distal end section A to be left in the bladder 102 and a body section B disposed adjacent to the distal end section A and defined in a proximal end side; at least one electrode 114 provided on a distal end of the catheter body 111; and a lead wire 115 inserted into the catheter body 111 so that first end of the lead wire is connected to the electrode 114 and second end of the lead wire is drawn out of a proximal end of the catheter body 111 to be connected to the current perception threshold inspection apparatus 4. | 1. A diagnosis catheter for an interstitial cystitis wherein the catheter is connected to a current: perception threshold inspection apparatus to diagnose the interstitial cystitis, comprising: a cylindrical catheter body made of a soft flexible material and including a distal end section to be left in the bladder and a body section disposed adjacent to said distal end section and defined in a proximal end side;at least one electrode provided on a distal end of said catheter body; anda lead wire inserted into said catheter body so that first end of said lead wire is connected to said electrode and second end of said lead wire is drawn out of a proximal end of said catheter body to be connected to said current perception threshold inspection apparatus. 2. A diagnosis catheter for an interstitial cystitis according to claim 1, characterized in further comprising: a balloon mounted on an outer periphery of said catheter body near a boundary between said distal end section and said body section; a fluid supply passage formed inside said catheter body in communication with said balloon; and an injection section capable of injecting a predetermined fluid through said fluid supply passage. 3. A diagnosis catheter for an interstitial cystitis according to claim 2, characterized in that both end sections of said balloon adhere to an exterior surface of said catheter body, and a width of adhesion varies circumferentially along said exterior surface of said catheter body. 4. A diagnosis catheter for an interstitial cystitis according to claim 2, characterized in that said balloon has a varied thickness circumferentially along said exterior surface of said catheter body. 5. A diagnosis catheter for an interstitial cystitis according to claim 2, characterized in that said catheter body has a varied exterior wall thickness circumferentially along said exterior surface thereof. 6. A diagnosis catheter for an interstitial cystitis according to claim 3, characterized in that said catheter body has a varied exterior wall thickness circumferentially along said exterior surface thereof. 7. A diagnosis catheter for an interstitial cystitis according to claim 4, characterized in that said catheter body has a varied exterior wall thickness circumferentially along said exterior surface thereof. | Tsukada Medical Research Co., Ltd. | A61B005/20 A61B018/18 A61M025/00
US20080319283 | A1 | 2008-12-25 | METHOD AND APPARATUS FOR MEASURING SKIN TEXTURE | 2008-06-19 | Among various methods, apparatuses, and media, a number of methods are provided for measuring skin surface texture. One such method includes illuminating an area of skin with polarized light, and obtaining a measurement of light returned by the illuminated area of skin in a first and a second waveband. The method includes processing the measurement of light in the first waveband to determine an estimated expected level of light in the second waveband returned by the illuminated area of skin utilising a model of the interaction of light with at least one chromophore in the skin. A measurement of the surface texture of the imaged illuminated area of skin can be determined on the basis of a difference between the estimated and actual levels of light in said second waveband returned by the illuminated area of skin. | 1. A method of measuring skin surface texture comprising: illuminating an area of skin with polarized light;obtaining a measurement of light returned by the illuminated area of skin in a first and a second waveband, wherein the measured light in the first waveband is light having a different polarity to the light with which said area of skin is illuminated and the measured light in the second waveband comprises light having the same and different polarities of light as the light with which said area of skin is illuminated;processing the measurement of light in the first waveband to determine an estimated expected level of light in said second waveband returned by the illuminated area of skin utilising a model of the interaction of light with at least one chromophore in the skin; anddetermining a measurement of the surface texture of the imaged illuminated area of skin on the basis of a difference between the estimated and actual levels of light in said second waveband returned by the illuminated area of skin. 2. The method of claim 1 wherein said first waveband comprises a waveband corresponding to visible light. 3. The method of claim 2 wherein said first waveband comprises a waveband corresponding to red light. 4. The method of claim 1 wherein said second waveband comprises a waveband corresponding to infra-red light. 5. The method of claim 1 wherein said at least one chromophore in the skin comprises melanin. 6. The method of claim 1, further comprising: obtaining a measurement of light returned by the illuminated area of skin in a third waveband, wherein the measured light in the third waveband is light having a different polarity to the light with which said area of skin is illuminated;wherein processing the first measurement of light to determine an estimated level of light in said second waveband returned by the illuminated area of skin comprises:processing the measurements of light in said first and third wavebands to determine an estimated expected level of light in said second waveband returned by the illuminated area of skin utilising a model of the interaction of light with a first and a second chromophore in the skin. 7. The method of claim 6, wherein said a first and a second chromophore comprise melanin and haemoglobin. 8. An apparatus for measuring skin surface texture, the apparatus comprising: a light source operable to illuminate an area of skin with polarized light;a detector operable to: obtain a measurement of light returned by an illuminated area of skin in a first waveband and a second waveband, wherein the measured light in the first waveband is light having a different polarity to the light with which said area of skin is illuminated by said light source, and the measured light in the second waveband comprises light having the same and different polarities of light as the light with which said area of skin is illuminated by said light source; anda processor operable to: process an obtained measurement of light in ... | Astron Clinica Limited | A61B005/00
US20080300508 | A1 | 2008-12-04 | Diagnostic patch | 2007-05-30 | The Present invention generally relates to medical diagnostics devices and more specifically to a diagnostic patch attachable to a subject's skin for performing a rapid blood test.;A well-known format for performing assays where a sample is applied to a test strip impregnated with assay specific reagents. | 1. A diagnostic device for performing a diagnostic blood test of a subject comprising a flexible patch attachable to the subject's skin, a puncturing unit for producing a blood sample of the subject, and an at least one test element embedded in said patch in a flow communication with said blood sample. 2. The diagnostic device of claim 1 wherein the device comprises a display window through which test results can be read. 3. The diagnostic device of claim 2 wherein said test results are visibly read. 4. The diagnostic device of claim 1 wherein said patch comprises a first flexible sheet comprising an aperture located opposite the puncturing unit and a second flexible sheet overlaying said first flexible sheet and wherein said test element is sandwiched between said first and second flexible sheets in a fluid communication with said aperture. 5. The diagnostic device of claim 1 wherein the puncturing unit comprises a lancing element mounted within a flexible housing. 6. The diagnostic device of claim 1 wherein said at least one test element is a diagnostic strip adapted for a lateral flow assay of a whole blood sample. 7. The diagnostic device of claim 5 wherein said assay is an immunoassay, or an enzymatic assay, or a biochemical assay or a chemical assay. 8. The diagnostic device of claim 5 wherein said assay is a positive/negative assay. 9. The diagnostic device of claim 6 wherein said diagnostic strip is configured for detecting the presence and/or concentration of an analyte in a sample applied thereto. 10. The diagnostic strip of claim 9 wherein said analyte is a blood borne pathogen. 11. The device of claim 9 wherein said diagnostic strip is impregnated with one or more reagents disposed along the strip, said reagents are selected to provide a visible signal when said analyte is present in said sample. 12. The diagnostic device of claim 6 wherein said diagnostic strip comprises a whole blood separation zone for entrapping and retaining red blood cells. 13. The diagnostic device of claim 6 wherein said diagnostic strip comprises a sample receiving zone, a reaction zone and a detection zone. 14. The device of claim 1 further comprising a reservoir of a releasable reagent solution adapted to release the reagent solution to facilitate running the diagnostic test. 15. The device according to claim 14 wherein said reservoir of a releasable reagent solution comprises a blister made of liquid impermeable film encapsulating said reagent solution. 16. The device according to claim 15 wherein said blister is accommodated within the puncturing unit. 17. The device according to claim 15 wherein said blister is accommodated within a second puncturing unit located upstream the puncturing unit for producing the blood sample of the subject and wherein the patch comprises a channel connecting between said two puncturing units to form a flow communication between the reagent solution released from said blister and the blood sample. 18. The de... | Inverness Medical Switzerland GmbH | A61B005/157
US07726206 | B2 | 2010-06-01 | Foot pressure alert and sensing system | 2007-11-01 | This invention relates to a system that continuously monitors pressure and force on the foot, analyzes and visualizes the pressure and force exerted on said foot in real-time. The invention measures pressure and force applied to a plurality of sensors placed in various points of an orthotic, shoe, shoe lining, insert, sock or sock type device. If the sensors detect pressure or force, the sensors send a signal to a microcomputer processor located in the shoe that subsequently analyzes and sends the data wireless to either a handheld electronic device, a personal computer, an electronic data capture system or a software program. The handheld electronic device or the personal computer then displays the data to an operator of the device or computer instantaneously; while the data capture system or program forwards the information to a handheld device and/or personal computer. | 1. Footwear capable of sensing and alerting a user to pressure upon a foot comprising: a plurality of primary pressure sensors located in the footwear;a plurality of secondary pressure sensors located in the footwear, and adapted to be activated only when one or more of the primary pressure sensors are defective; anda medical node located in the footwear connected to the primary pressure sensors and the secondary pressure sensors, the medical node receiving data from the primary pressure sensors and the secondary pressure sensors, processing the data to generate pressure data, and transmitting the pressure data. 2. The footwear of claim 1 wherein the medical node transmits the pressure data to a base station. 3. The footwear of claim 2 wherein the base station analyzes the pressure data and generates an event alert if a predetermined condition occurs. 4. The footwear of claim 3 wherein a predetermined condition occurs when the pressure data indicates that there is pressure above a predetermined level and for longer than a predetermined time period at one or more of the plurality of primary pressure sensors. 5. The footwear of claim 3 wherein the event alert includes a media alert to an electronic device. 6. The footwear of claim 3 wherein the event alert includes a text message to a mobile phone. 7. The footwear of claim 2 wherein the base station calibrates the medical node and the pressures sensors. 8. The footwear of claim 1 wherein a rechargeable battery is connected to the medical node. 9. A system for sensing and alerting a user to pressure upon a foot comprising: footwear containing a plurality of primary pressure sensors, a plurality of secondary pressure sensors and a medical node connected to the primary pressure sensors, and the secondary pressure sensors, one or more of the secondary pressure sensors being activated only when one or more of the primary pressure sensors become defective, the medical node receiving data from the primary pressure sensors and the secondary pressure sensors and processing the received data to generate pressure data; anda first base station connected to the medical node to receive the pressure data, the first base station adapted to analyze the pressure data and generate an event alert when the pressure data indicates that there is pressure above a predetermined level or for longer than a predetermined time period at one or more of the pressure sensors. 10. The system of claim 9 wherein the first base station includes a data storage module to store the pressure data received from the medical node. 11. The system of claim 9 wherein the event alert includes a media alert to an electronic device. 12. The system of claim 9 wherein the event alert includes a text message to a mobile phone. 13. The system of claim 9 wherein the base station calibrates the medical node, the primary pressures sensors, and the secondary pressure sensors. 14. The system of claim 9 wherein a rechargeable battery is connec... | The Regents of The University of California Medisens Wireless, Inc. | G01D007/00
US07455643 | B1 | 2008-11-25 | Continuous non-invasive blood pressure measurement apparatus and methods providing automatic recalibration | 2004-12-20 | A blood pressure measurement system is configured to perform a calibration automatically when a calibration condition is satisfied. The calibration condition is based upon one or more parameters of pulse waves of a subject. The parameters may include pulse wave area; a time difference between systolic peak and reflected wave peak or dichrotic notch in the pulse wave and a shape of at least a portion of the pulse wave. | 1. A method for maintaining calibration of a continuous non-invasive blood pressure (CNIBP) measurement system, the method comprising: performing a calibration by: obtaining for a subject a reference blood pressure value, reference CNIBP data and reference pulse wave data, the reference pulse wave data corresponding to one or more reference pulse waves of the subject and having a waveform;calculating and storing coefficients of a CNIBP calculation formula based on the reference blood pressure value and the reference CNIBP data; and,storing a template comprising information characterizing a plurality of characteristics of the waveform of the reference pulse wave data;subsequently, obtaining subsequent pulse wave data characterizing one or more subsequent pulse waves of the subject and evaluating a recalibration condition comparing characteristics of a waveform of the subsequent pulse waves to the template; and,if the subsequent pulse wave data and template satisfy the recalibration condition, repeating performing the calibration of the CNIBP measurement system,wherein the recalibration condition comprises a sensitivity parameter and the sensitivity parameter affects whether or not the subsequent pulse wave data and template satisfy the recalibration condition. 2. A method according to claim 1 wherein the reference CNIBP data comprises differential pulse transit time data or pulse wave velocity data. 3. A method according to claim 2 wherein the differential pulse transit time data comprises first and second signals from first and second sensors, and the reference pulse wave data and subsequent pulse wave data are derived at least in part from one of the first and second signals. 4. A method according to claim 3 wherein the first and second sensors are pulse-oximetry-type sensors. 5. A method according to claim 1 wherein the pulse wave recalibration condition is based, at least in part, on at least one of a first feature, a second feature and a third feature of the reference and subsequent pulse waves: the first feature comprising a pulse wave area of the reference and subsequent pulse waves;the second feature comprising a time difference, ΔT, between two internal features within the reference and subsequent pulse waves; and,the third feature comprising a pulse wave shape of the reference and subsequent pulse waves. 6. A method according to claim 5 wherein the recalibration condition is based on a plurality of the first, second and third features. 7. A method according to claim 5 wherein the recalibration condition is based in part on a pulse amplitude of one or more of the reference pulse waves. 8. A method according to claim 5 wherein the recalibration condition is based in part on a pulse width of one or more of the reference pulse waves. 9. A method according to claim 1 wherein the pulse wave recalibration condition is based, at least in part, on a comparison of a pulse wave area of one or more of the one or more reference pulse waves to... | Nellcor Puritan Bennett Ireland | A61B005/00
US07530956 | B2 | 2009-05-12 | Daytime/nighttime respiration rate monitoring | 2007-06-15 | A device and method can monitor or trend a patient's respiration rate measurements according to the time of day. The device, which may be implantable or external, collects and classifies respiration rate measurements over time. The trended information about particular classes of respiration rate measurements is then communicated to a remote external device, which in turn provides an indication of heart failure decompensation. Examples of classes of respiration rate measurements include a daily maximum respiration rate value, a daily minimum respiration rate value, a daily maximum respiration rate variability value, a daily minimum respiration rate variability value, and a daily central respiration rate value. These respiration rate measurements can be further classified into daytime or nighttime respiration rate measurements. | 1. An apparatus comprising: means for detecting a respiration signal from a subject;means for determining respiration rate or interval measurements from the respiration signal;means for classifying the respiration rate or interval measurements into at least one class of respiration rate or interval measurements and for associating a specified time of day with the at least one class;means for trending at least one of: (1) time of occurrence information about the classified respiration rate or interval measurements of a particular class of respiration rate or interval measurements, or (2) respiration rate or interval measurement information repeatedly, over a period of multiple days, during a specified one of a daytime subperiod or a nighttime subperiod, wherein the specified one of the daytime subperiod or the nighttime subperiod corresponds to the same specified time of day during each of the multiple days; andmeans for providing a heart failure decompensation indication using the trended information. 2. The apparatus of claim 1, wherein the means for classifying the respiration rate or interval measurements into at least one class of respiration rate or interval measurements includes the following classes: a daily maximum respiration rate value, a daily minimum respiration rate value, a daily maximum respiration rate variability value, a daily minimum respiration rate variability value, and a daily central respiration rate value. 3. The apparatus of claim 1, wherein the means for classifying the respiration rate or interval measurements into at least one class of respiration rate or interval measurements includes means for classifying the respiration rate or interval measurements as a specified one of daytime or nighttime respiration rate or interval measurements. 4. The apparatus of claim 3, comprising: means for collecting the specified one of the daytime or nighttime respiration rate or interval measurements;means for performing a statistical analysis on the collected respiration rate or interval measurements; andmeans for providing the heart failure decompensation indication using a result of the statistical analysis. 5. The apparatus of claim 4, comprising means for providing the heart failure decompensation indication by detecting a change over time in a measure of centrality of the collected respiration rate or interval measurements. 6. The apparatus of claim 4, comprising means for providing the heart failure decompensation indication by detecting a change over time in a measure of spread or variance of the collected respiration rate or interval measurements. 7. The method of claim 1, wherein the means for trending information about the classified respiration rate or interval measurements comprises trending information about a time associated with the particular class of respiration rate or interval measurements. 8. An apparatus comprising: a respiration sensor, configured to detect a respiration signal from a subject;a respiratio... | Cardiac Pacemakers, Inc. | A61B005/00
US07462346 | B2 | 2008-12-09 | Light-stabilized(in vivo) stain composition and method of manufacture | 2001-08-28 | Photochemical demethylation reactions in solutions of thiazine dyes, in which the dye molecules act as both sensitizer and substrate, are reduced by quenching triple-state dye molecules, returning them to the unreactive ground state. | 1. A method for at least partially reducing the oxidative demethylation of a thiazine dye, the method comprising the steps of: (a) providing a thiazine dye in solution; and(b) adding a metal ion to the thiazine dye solution; wherein the thiazine dye comprises an inner thiazine ring and two outer rings, wherein at least one of the two outer rings has an -N-methylamino or an -N,N-dimethylamino substituent;wherein the metal ion quenches the thiazine dye to at least partially reduce the oxidative demethylation of the thiazine dye; andwherein the method further comprises the step of purifying the thiazine dye solution by high performance liquid chromatography before the step of adding the metal ion to the thiazine dye solution. | Zila Biotechnology, Inc. | A61B010/00 A61B005/00 A61B008/00
US07591791 | B2 | 2009-09-22 | Diagnostic thimble | 2007-06-21 | A diagnostic device for performing a diagnostic blood test of a subject comprising a thimble-like element adapted to be engaged with the subject's finger, a puncturing unit for producing a blood sample from the subject's finger, and an at least one test element in flow communication with said blood sample. | 1. A diagnostic device for performing a diagnostic blood test of a subject, the device comprising a tubular body mountable on the subject's fingertip, wherein said tubular body comprises: an open end adapted to receive the subject's fingertip;a transparent window in said tubular body;a puncturing unit attached to said tubular body, the puncturing unit comprising a retractable lancing element for producing a blood sample from the subject's fingertip; andat least one diagnostic strip impregnated with assay specific reagents positioned inside said tubular body, the diagnostic strip comprising a sample receiving zone and a detection zone where a visible signal indicative of test results is formed, the diagnostic strip is positioned inside said tubular body such that said sample receiving zone is in fluid communication with the blood sample produced by said puncturing unit and said detection zone is visible through said transparent window. 2. The diagnostic device of claim 1 wherein said diagnostic strip is adapted for a lateral flow assay of a whole blood sample. 3. The diagnostic device of claim 2 wherein said assay is an immunoassay, an enzymatic assay, a biochemical assay or a chemical assay. 4. The diagnostic device of claim 2 wherein said assay is a positive/negative assay. 5. The diagnostic device of claim 1 wherein said diagnostic strip is configured for detecting the presence and/or concentration of an analyte in said blood sample. 6. The diagnostic device of claim 5 wherein the assay specific reagents include one or more reagents disposed along the diagnostic strip, said reagents are selected to provide a visible signal when said analyte is present in said sample. 7. The diagnostic device of claim 6 wherein said analyte is a blood borne pathogen. 8. The diagnostic device of claim 1 wherein said diagnostic strip comprises a whole blood separation zone for entrapping and retaining red blood cells. 9. The diagnostic device of claim 1 wherein said diagnostic strip comprises a reaction zone interposed between said sample receiving zone and said detection zone. 10. The diagnostic device of claim 1 further comprising a reservoir of a releasable reagent solution adapted to release the reagent solution to facilitate running the diagnostic test. 11. The diagnostic device according to claim 10 wherein said reservoir of the releasable reagent solution comprises a blister made of liquid impermeable film encapsulating said reagent solution. 12. The diagnostic device of claim 1 further comprising a safety means to prevent premature activation of the puncturing unit. 13. The diagnostic device of claim 1 wherein the retractable lancing element is a solid pointed element and wherein the blood sample is drawn away from the retractable lancing element into the diagnostic strip. 14. The diagnostic device of claim 1 wherein the retractable lancing element is configured to penetrate through said tubular body. 15. The diagnostic device of claim 1 wher... | Inverness Medical Switzerland GmbH | A61B005/00 B65D081/00 C12M001/34 C12M003/00
US07447533 | B1 | 2008-11-04 | Implantable electronic medical device having an encapsulated optical transducer | 2005-03-24 | An implantable electronic device has a main housing with a leak-proof-sealed housing chamber containing electronic circuitry. An optical module is connected to the main housing, and has a module housing with a leak-proof-sealed transducer chamber isolated from the main housing chamber. The module housing has an optically-transmissive optical element, and the transducer chamber contains an optical transducer connected to the circuitry. The module housing includes a number of signal conductors extending from the transducer chamber to a location external to the transducer chamber, and connected to circuitry in the main housing chamber. | 1. An implantable electronic device comprising: a main housing defining a leak-proof-sealed housing chamber containing electronic circuitry;a header connected to the main housing;an optical module connected to remain in contact with the main housing;the optical module having a module housing defining a leak-proof-sealed transducer chamber isolated from the main housing chamber and the header;the module housing having an optically-transmissive optical element;the transducer chamber containing an optical transducer connected to the circuitry; andthe module housing including a plurality of signal conductors extending from the transducer chamber to a location external of the transducer chamber, and connected to circuitry in the main housing chamber. 2. The device of claim 1 wherein the main housing defines an aperture, and wherein the module housing occupies the aperture and is sealed to the main housing aperture to seal the main housing chamber. 3. The device of claim 2 wherein the main housing has a planar shape with two opposed major faces, and wherein the aperture is defined in one of the major faces. 4. The device of claim 3 wherein the module has a major face flush with the major face of the main housing in which the aperture is defined. 5. The device of claim 1 wherein the main housing includes a connector element operably connected to the signal conductors and having conductors extending into the main housing chamber. 6. The device of claim 5 wherein the connector element sealably occupies an aperture in the main housing. 7. The device of claim 1 wherein the main housing comprises two major portions, and wherein the aperture is defined entirely within one of the major portions. 8. The device of claim 1 wherein the main housing has a generally planar shape, and the header defines lead apertures at a peripheral portion of the main housing, the module being adjacent to the header. 9. The device of claim 1 wherein the module is external to the header. 10. The device of claim 1 wherein the optically transmissive element is apart from the header. 11. The device of claim 8 wherein the optically transmissive element is positioned in a plane defined by a major surface of the main housing. 12. The device of claim 1 further comprising a flange connected to the main housing. 13. The device of claim 12 wherein the flange is a flanged ring and wherein the ring defines a bore, the device further comprising an insert connected to the bore. 14. The device of claim 13 wherein the insert comprises ceramic. 15. The device of claim 13 wherein the insert fully occupies the bore and wherein the insert is gold-blazed to the bore to provide a hermetic seal. 16. The device of claim 13 wherein the plurality of signal conductors are connected to the insert to provide a hermetic seal. 17. The device of claim 13, wherein the insert is plated with a conductive material to provide a shield against electromagnetic interference. 18. The device of claim 1 wher... | Pacesetter, Inc. | A61B005/00
US07468037 | B2 | 2008-12-23 | Apparatus and method for measuring hemodynamic parameters | 2005-02-16 | An apparatus for measuring hydrodynamic parameters, especially for Augmentation Index (Aix) and/or Ejection Duration (ED), by non-invasive, cuff based occlusive, blood pressure measurement, which apparatus comprises occlusive, oscillometric automatic blood pressure meter and units. There can also be a method for non-invasive measurement of hemodynamic characteristics, especially AN and/or ED with an occlusive, pressure sensor cuff, placed on the brachial artery, wherein the apparatus according to the invention samples and analyzes and evaluates the signal flow of the oscillations of pulse waves using an unusual stepwise blood pressure measurement. | 1. Apparatus An apparatus for measuring hemodynamic parameters, by non-invasive, cuff based occlusive, blood pressure measurement, which apparatus comprises occlusive, oscillometric automatic blood pressure meter and units, determining the values of hemodynamic parameters, comprising an oscillation wave separating and storing signal detector for detecting an oscillation signal, the sampling rate thereof is at least 200/heart cycle; and said signal detector has a storage unit resolution that is organized at least 9 bit,a digital anti-filter connected to the signal detector to compensate the distortions rising at sampling, separating and digitizing the oscillation wave, detected by the signal detector,an amplitude arithmetic unit connected to the anti-filter establishing an Augmentation Index (Aix); anda synthesizing unit connected to the anti-filter establishing an Ejection Duration (ED). 2. The apparatus according to claim 1, wherein the sampling rate of the signal detector is 180-220/second. 3. The apparatus according to claim 1, wherein the storage unit storing the signals, generated by the oscillation wave, is organized 10-12 bit. 4. The apparatus according to claim 1 further comprising a time-arithmetic unit establishing a Pulse Wave Velocity (PWV), or integrator unit establishing a Systole Area Index (SAI) and Diastole Area Index (DAI). 5. The apparatus according to claim 1, wherein the amplitude arithmetic unit, the synthesizing unit, a time-arithmetic unit, or an integrator unit are joined to a common program controller, and compiled to an analyzer. 6. The apparatus according to claim 1, further comprising a portable, 24 h ambulatory blood pressure monitor. 7. The apparatus according to claim 1, further comprising a telemedical home care system. 8. The apparatus according to claim 1, further comprising a 24 h blood pressure monitor, which is controlled by a built-in ECG. 9. A method for non-invasive measurement of hemodynamic characteristics comprising the steps of: (a) performing a standard stepwise blood pressure measurement using an occlusive, pressure-sensor cuff placed on the brachial artery;(b) storing systolic blood pressure (SBP), diastolic blood pressure (DBP), and heart rate (HR) values;(c) subsequently setting the cuff to supra-systolic pressure range over the systolic pressure;(d) performing a pressure oscillometric pulse wave detection at supra-systolic pressure range, receiving oscillation curve and simultaneously by an “anti-filter” process compensating for signal distortions appearing at sampling;(e) calculating an Augmentation Index (Aix) on the basis of the wave amplitudes from the oscillation curves so received; and(f) calculating the Ejection Duration (ED) value on the oscillating curve determining the minimum point after the first reflex wave. 10. The method according to claim 9, wherein the sampling rate of the pressure oscillometric rules wave detection is taken at least 180 samples per second, and the digit... | - | A61B005/02
US07465472 | B1 | 2008-12-16 | Method for developing latent fingerprints | 2007-02-26 | A method for developing latent fingerprints includes placing an apparatus in an area having one or more objects. The apparatus has a breakable capsule containing a cyanoacrylate monomer disposed within an internal chamber of a housing containing a cyanoacrylate catalyst and being adapted for emitting vapors. The housing is impregnated with the cyanoacrylate catalyst or has a layer of cyanoacrylate catalyst on a portion of its interior surface. The apparatus optionally has a breakable capsule breaking mechanism for assisting a user in breaking the breakable capsule within the housing. The user breaks the breakable capsule, causing the cyanoacrylate monomer spilling from the breakable capsule to react with the cyanoacrylate catalyst contained in the housing to create vapor fumes. The vapor fumes contact an object being tested to reveal any latent fingerprints on the object. | 1. A method for developing latent fingerprints on one or more objects, comprising the steps of: (a) placing an apparatus in a target area having one or more objects, said apparatus having a breakable capsule containing a cyanoacrylate monomer disposed within an internal chamber of a housing containing a cyanoacrylate catalyst, said housing being adapted for receiving said breakable capsule and having a means for emitting vapors;(b) breaking said breakable capsule contained within said housing;(c) generating vapors from said cyanoacrylate monomer reacting with said cyanoacrylate catalyst; and(d) developing latent fingerprints on the one or more objects when said vapors contact the objects. 2. The method of claim 1 wherein the target area is selected from a group consisting of a container, box, room, interior of a vehicle, and an open space. 3. The method of claim 1, wherein said step (b) comprises the steps of: (b)(i) engaging a capsule breaking mechanism with said breakable capsule; and(h)(ii) applying pressure to said capsule breaking mechanism to break said breakable capsule. 4. The method of claim 1, wherein said apparatus is contained within a sealed wrapper and further comprising the step of: (e) opening said sealed wrapper prior to said step (a). 5. The method of claim 1, wherein said housing is made of a material impregnated with said cyanoacrylate catalyst. 6. The method of claim 1, wherein said housing further comprising a layer containing said cyanoacrylate catalyst, said layer covering a portion of an interior surface of said housing. 7. The method of claim 6, wherein said layer is selected from the group consisting of a direct application of said cyanoacrylate catalyst to said portion of said interior surface of said housing, and a liner within said internal chamber of said housing. 8. The method of claim 1, wherein said breakable capsule is made of a material selected from the group consisting of glass, plastic, ceramic, composite, and resin coated paper. 9. The method of claim 1, wherein said housing is made of a material selected from the group consisting of paper, fibrous textile, cellulosic fiber, plastic, resin, metal, and composite. 10. The method of claim 1, wherein said means for emitting vapors is selected from the group consisting of said housing having perforations, pores, slots, and one or more open ends. 11. The method of claim 1, wherein said step (b) is performed by a capsule breaking mechanism selected from the group consisting of: a mechanism having a top half and a bottom half pivotally connected, a C-shaped mechanism, a head with a shaft wherein a distal end of said shaft is adapted to engage an exterior surface of said breakable capsule, and said housing incorporating a flange extending into said internal cavity. 12. The method of claim 1, wherein said housing further comprises a resealable side and a means for opening and closing said resealable side. 13. The method of claim 1, wherein said apparatus ... | - | A61B005/117
US07853307 | B2 | 2010-12-14 | Methods, apparatuses, and systems useful in conducting image guided interventions | 2008-06-26 | Methods, apparatuses, and systems relating to image guided interventions on dynamic tissue. One embodiment is a method that includes creating a dataset that includes images, one of the images depicting a non-tissue internal reference marker, being linked to non-tissue internal reference marker positional information, and being at least 2-dimensional. Another embodiment is a method that includes receiving a position of an instrument reference marker coupled to an instrument; transforming the position into image space using a position of a non-tissue internal reference marker implanted in a patient; and superimposing a representation of the instrument on an image in which the non-tissue internal reference marker appears. Computer readable media that include machine readable instructions for carrying out the steps of the disclosed methods. Apparatuses, such as integrated circuits, configured to carry out the steps of the disclosed methods. Systems that include devices configured to carry out steps of the disclosed methods. | 1. A non-transitory processor-readable medium storing code representing instructions to cause a processor to perform a process, the code comprising code to: create a dataset that includes a plurality of images associated with a body part, at least one image from the plurality of images depicting a non-tissue internal reference marker, the at least one image from the plurality of images being associated with a non-tissue internal reference marker position information, the at least one image from the plurality of images being at least 2-dimensional;calculate for each image in the dataset a dataset vector based on a position of the non-tissue internal reference marker and the position of an external reference marker;receive data associated with a current position of the external reference marker and a current position of the non-tissue internal reference marker;calculate a current vector using the data associated with the current position of the external reference marker and the data associated with a current position of the non-tissue internal reference marker;identify a match dataset vector (MDV), the MDV being the dataset vector most similar to the current vector, the MDV being associated with an image from the dataset;receive data associated with a current position of an instrument reference marker coupled to an instrument;apply a transformation associated with the image associated with the MDV to the current position of the instrument reference marker, thus transforming the current position of the instrument reference marker into image space; andsuperimpose a representation of the instrument on the image associated with the MDV. 2. The non-transitory processor-readable medium of claim 1, wherein the code to create the dataset includes code to create a dataset that includes at least one image taken via at least one of fluoroscopy, CINE fluoroscopy, computed tomography (CT), magnetic resonance imaging (MRI), ultrasound, positron emission tomography (PET), or functional magnetic resonance imaging (fMRI). 3. The processor-readable medium of claim 1, wherein the code to create the dataset includes code to create a dataset that includes at least one 3-dimensional image. 4. The processor-readable medium of claim 1, wherein code to create the dataset includes code to create a dataset that includes a plurality of images including gated information. 5. The processor-readable medium of claim 1, further comprising code to: cause threshold segmentation of each image of the plurality of images, the threshold segmentation allowing the code to create the dataset. 6. The processor-readable medium of claim 1, further comprising code to: receive a periodic human characteristic signal. 7. The processor-readable medium of claim 6, wherein the code to receive the periodic human characteristic signal includes code to receive a human characteristic signal calculated using the position of the non-tissue internal reference marker. 8. The processor-readable medium... | Veran Medical Technologies, Inc. | A61B005/05
US07447542 | B2 | 2008-11-04 | Three-dimensional monitoring of myographic activity | 2003-07-11 | A system comprising an electromyogram (EMG) system operative to sense electromyographic activity generated in a muscle, at least one position sensor, and a processor in communication with the EMG system and the at least one position sensor, the processor operative to process data of the EMG system and three-dimensional position and orientation information from the at least one position sensor to provide an output that comprises electromyographic activity data as a function of the three-dimensional position and orientation of the at least one position sensor. | 1. A system comprising: an electromyogram (EMG) system comprising at least one EMG sensor operative to sense electromyographic activity generated in a uterine muscle and output electrical muscular activity signals;at least one position sensor placed near said at least one EMG sensor; anda processor in communication with said EMG system and said at least one position sensor, said processor operative to process electrical muscular activity signals of said EMG system and three-dimensional positions of said at least one EMG sensor from said at least one position sensor to provide an output and display of said electrical muscular activity signals as sensed by said at least one EMG sensor and the three-dimensional positions of said at least one EMG sensor at the same time, and further comprising a cardiotocogram (CTG) monitor in communication with said processor, said CTG monitor comprising a fetal beat-to-beat heart rate (FHR) sensor and a uterine labor activity (TOCO) sensor, and wherein said system indicates onset of premature birth based on the three-dimensional positions of said at least one EMG sensor together with TOCO information from said TOCO sensor. 2. The system according to claim 1 wherein said EMG system comprises at least one reference EMG sensor adapted to sense electromyographic activity generated in a reference muscle. 3. The system according to claim 1, further comprising a monitor coupled to said processor and adapted to display processed information from said processor. 4. The system according to claim 1, further comprising a position sensing system adapted to measure the three-dimensional position and an orientation of said at least one position sensor with respect to a reference position fixed in space. 5. The system according to claim 1, wherein said processor is operative to process data from said CTG monitor in addition to the data of said EMG system and the three-dimensional position information from said at least one position sensor to provide an output and display of electromyographic activity data and CTG data and the three-dimensional position of said at least one EMG sensor at the same time. 6. The system according to claim 1, further comprising a warning device in communication with said processor, operative to issue a warning if processed data processed by said processor is above a predefined limit. | OB Tools Ltd. | A61B005/04 A61B005/103 A61B005/117
US07991462 | B2 | 2011-08-02 | Method and apparatus for monitoring a sedated patient | 2006-02-03 | A method and an apparatus for monitoring a sedated patient, the method comprising the steps of providing a skin conductance signal measured at an area of the patient's skin through a time interval, establishing the existence of at least two fluctuation peaks in the skin conductance signal through said time interval, considering if the amplitudes of fluctuation peaks in the skin conductance signal through said interval, the basal level of the skin conductance signal through said interval and the width of the fluctuation peaks in the skin conductance signal fulfils a predetermined criterion, activating a first output signal which indicates the state of awakening in the patient if said criterion is fulfilled, and activating a second output signal which indicates the state of pain in the patient if said criterion is not fulfilled. | 1. Method for monitoring a sedated patient, comprising the following steps, performed by a control unit that includes a processing unit: providing a skin conductance signal measured at an area of the sedated patient's skin through a time interval,processing the skin conductance signal to establish an existence of at least two fluctuation peaks in the skin conductance signal through said time interval,considering if amplitudes of fluctuation peaks in the skin conductance signal through said interval, basal level of the skin conductance signal through said interval and a width of the fluctuation peaks in the skin conductance signal fulfils the following three predetermined criteria: if an average of said amplitudes exceeds a first limit value, said first limit value having a range of 0.05 μS to 0.20 μS,if said basal level has shown an increase of more than a second limit value, said second limit value having a range of 0.05 μS to 0.3 μS during a recently elapsed time interval in a range of 10 seconds to 30 seconds, andif said width of the fluctuation peaks exceeds a third limit value, said third limit value having a range of 1 second to 5 seconds,the method further comprising the steps of activating a first output signal which indicates a state of awakening in the sedated patient if one of said criteria is fulfilled, andactivating a second output signal which indicates a state of pain in the sedated patient if none of said criteria is fulfilled. 2. Method according to claim 1, wherein said first limit value has a range of 0.07 μS to 0.13 μS. 3. Method according to claim 1, wherein said second limit value has a range of 0.08 μS to 0.12 μS and said recently elapsed time interval is in a range of 12 seconds to 18 seconds. 4. Method according to claim 1, wherein said third limit value has a range of 1.5 seconds to 3 seconds. 5. Method according to claim 1, wherein said step of establishing the existence of at least two fluctuation peaks in the skin conductance signal through said time interval comprises the substep of establishing an existence of a valid peak if a derivative of the skin conductance signal changes sign through a period in the time interval. 6. Method according to claim 5, wherein said derivative is calculated as a difference between two subsequent sample values. 7. Method according to claim 5, wherein an additional criterion is established for when a peak is considered to be a valid peak, including ensuring that an amplitude of the skin conductance signal exceeds an absolute limit value selected from a range of 0.01 uS to 0.02 uS. 8. Method according to claim 5, wherein an additional criterion is established for when a peak is considered to be a valid peak, including ensuring that an increase in the skin conductance signal as a function of time remains below a certain limit. 9. Method according to claim 5, wherein an additional criterion is established for when a peak is considered to be a valid peak, including ensuring that an ... | Med Storm Innovation AS | A61B005/00 A61B005/04 A61B005/05
US08062234 | B2 | 2011-11-22 | Ph sensor | 2005-06-23 | The invention relates to a pH sensor for use in a catheter drainage system. The invention also relates to a catheter, drainage bag or connection unit comprising the pH sensor of the present invention, methods of using the pH sensor and methods of making the pH sensor. | 1. A pH sensor configured for incorporation into a catheter drainage system wherein the pH sensor comprises: a polymer matrix chemically bound to a pH indicator, wherein the pH indicator only responds to sustained contact with a liquid or biofilm at an increased pH. 2. The pH sensor according to claim 1, that is configured to be incorporated into a catheter drainage system at any position provided that during use said pH sensor is visible. 3. The pH sensor according to claim 1 or claim 2 that is formed as a disc for insertion into a drainage bag of a catheter drainage system. 4. The pH sensor according to claim 1 or claim 2, that is formed as an annular ring having a bore allowing fluid to pass through the pH sensor. 5. The pH sensor according to any one of claims 1 and 2, wherein the polymer matrix is polyurethane or cellulose acetate. 6. The pH sensor according to any one of claims 1 and 2, wherein the pH indicator changes colour in response to an increase in pH of a bacterial biofilm that develops on the pH sensor. 7. The pH sensor according to any one of claims 1 and 2, wherein the pH indicator changes colour in response to an increase in pH from pH 6 to pH 7.6 or above. 8. The pH sensor according to any one of claims 1 and 2, wherein the pH indicator is selected from the group consisting of bromocresol purple, chlorophenol red, neutral red, bromothymol blue, cresol red, m-cresol purple, phenol red, 3-(m-)nitrophenol, thymol blue, cresolphthalein, fluorescein, Rosolic acid, and α-naphtholbenzene. 9. The pH sensor according to any one of claims 1 and 2, wherein the pH indicator is selected from the group consisting of neutral red, bromothymol blue, cresol red, phenol red, 3-(m-)fluorescein and rosolic acid. 10. The pH sensor according to any one of claims 1 and 2, wherein the sustained contact is more than 6 hours. 11. The pH sensor according to any one of claims 1 and 2, wherein the sustained contact is more than 9 hours. 12. The pH sensor according to any one of claims 1 and 2, wherein the sustained contact is more than 12 hours. 13. The pH sensor according to claim 1, which additionally comprises a matrix stabiliser. 14. The pH sensor according to claim 13, wherein the matrix stabiliser is a quaternary ammonium compound. 15. The pH sensor according to claim 1, wherein the polymer matrix is cellulose acetate and the pH indicator is bromothymol blue. 16. The pH sensor according to any one of claims 1 and 2, which is provided with a reference material to assist in determining when the pH indicator has responded to a sustained increase in pH. 17. A catheter comprising a pH sensor according to any one of claims 1 and 2 wherein the pH sensor is incorporated at any position in the catheter provided that during use said catheter is visible to a patient or caregiver. 18. A drainage bag comprising a pH sensor according to any one of claim 1, 2, 13, 14, or 15. 19. The catheter according to claim 17, wherein the pH sensor is positione... | University College Cardiff Consultants Limited | A61B005/00 B65D081/00 A61M001/00 A61F005/44 H01H029/02 H01H029/20
US08285369 | B2 | 2012-10-09 | Method for delineation of characteristic points in an electrocardiogram signal, and storage medium, electronic device and electrocardiogram signal analyzing system for implementing the method | 2007-05-30 | A method for delineation of characteristic points in an electrocardiogram signal includes the steps of: (a) processing the electrocardiogram signal to obtain at least one multiscale differential operator wavelet transform thereof; and (b) locating at least one characteristic point in at least one wave segment of the electrocardiogram signal based on the at least one multiscale differential operator wavelet transform. | 1. A method for delineation of characteristic points in an electrocardiogram signal, comprising the step of: (a) processing the electrocardiogram signal to obtain a first order multiscale differential operator wavelet transform, a second order multiscale differential operator wavelet transform, and a third order multiscale differential operator wavelet transform thereof; and(b) locating at least one characteristic point in at least one wave segment of the electrocardiogram signal based on the first order multiscale differential operator wavelet transform, the second order multiscale differential operator wavelet transform, and the third order multiscale differential operator wavelet transform, said at least one characteristic point being at least one of a P-wave peak point, a P-wave on point, a P-wave off point, a QRS-wave peak point, a QRS-wave on point, a QRS-wave off point, a T-wave peak point, a T-wave on point, and a T-wave off point. 2. The method as claimed in claim 1, wherein, in step (b), two characteristic points of the wave segment are located, and include an on point and an off point of the wave segment, the method further comprising the step of: (c) calculating an interval based on the on point and the off point thus located. 3. The method as claimed in claim 2, wherein the interval is one of a PR interval, a ST interval, a QT interval, a P-wave interval, a QRS-complex interval, and a T-wave interval. 4. The method as claimed in claim 1, wherein, in step (b), three characteristic points of the wave segment are located, and include an on point, a peak point, and an off point of the wave segment, the method further comprising the steps of: (c) calculating an interval based on the on point and the off point thus located; and(d) categorizing the wave segment according to morphology of the wave segment falling within the interval. 5. The method as claimed in claim 4, wherein, in step (d), the wave segment is categorized into one of a positive/negative T-wave, an ascending/descending T-wave, and a biphasic T-wave. 6. A non-transitory computer-readable storage medium comprising program instructions for causing a computing apparatus to execute consecutive steps of a method for delineation of characteristic points in an electrocardiogram signal, said method comprising: (a) receiving an electrocardiogram signal;(b) processing the electrocardiogram signal to obtain a first order multiscale differential operator wavelet transform, a second order multiscale differential operator wavelet transform, and a third order multiscale differential operator wavelet transform; and(c) locating at least one characteristic point in at least one wave segment of the electrocardiogram signal based on the first order multiscale differential operator wavelet transform, the second order multiscale differential operator wavelet transform, and the third order multiscale differential operator wavelet transform, said at least one characteristic point being at least... | Taipei Veterans General Hospital | A61B005/04
US07996076 | B2 | 2011-08-09 | Automated polysomnographic assessment for rapid eye movement sleep behavior disorder | 2008-04-02 | Methods and systems for diagnosing or assessing rapid eye movement sleep behavior disorder (RBD). Muscle tone or activity variance during rapid eye movement (REM) and non-rapid eye movement (NREM) sleep intervals of a polysomnogram are compared. A threshold based on the NREM data is used to identify a subject-specific threshold for abnormality in the REM variance. A metric that includes the percentage of REM variance exceeding the threshold relates to RBD. | 1. A method for diagnosing and assessing rapid eye movement sleep behavior disorder (RBD) in a patient, comprising: measuring at least one physiological signal of the patient during at least one rapid eye movement (REM) sleep time interval and at least one non-rapid eye movement (NREM) sleep time interval;determining a REM variance for the physiological signal within the at least one REM sleep time interval and an NREM variance for the physiological signal within the at least one NREM sleep time interval;establishing a threshold based on the NREM variance; andcalculating a percentage of REM sleep time intervals with variance above the threshold, where the steps of determining, establishing and calculating are executed by a computer processor. 2. The method of claim 1, wherein the physiological signal is an electromyogram (EMG). 3. The method of claim 1, wherein the physiological signal is an electrical potential generated by at least one muscle of the patient. 4. The method of claim 3, wherein the at least one muscle is a chin muscle. 5. The method of claim 1, wherein the at least one rapid eye movement (REM) sleep time interval is a plurality of successive time epochs. 6. The method of claim 5, wherein the time epochs are about 3 seconds. 7. The method of claim 1, wherein the measuring includes an automated method for distinguishing REM and NREM sleep time intervals. 8. The method of claim 1, wherein determining at least one of the REM variance or the NREM variance includes using the formula:σ=1/(N-1)⁢∑i=1N⁢(xi-x_)2wherein: σ is the variance,N is a number of samples in the time interval,xi is an ith sample of the signal in the time interval, andx is mean signal in the time interval. 9. The method of claim 1, wherein determining at least one of the REM variance or the NREM variance includes determining the sum of the square amplitudes of the physiological signal within the at least one REM sleep time interval. 10. The method of claim 1, wherein establishing a threshold based on the NREM variance includes: setting an upper limit for a physiological signal background during REM sleep, wherein the upper limit is about two to about six times the 5th percentile of the NREM variance. 11. The method of claim 10, wherein the upper limit is about four times the 5th percentile of the NREM variance. 12. The method of claim 1, wherein the measuring at least one physiological signal of the patient includes signals from two or more nights. 13. The method of claim 1, wherein the measuring at least one physiological signal of the patient occurs following a habituation night of sleep by the patient. 14. The method of claim 1, further comprising classifying the patient as an RBD candidate when the percentage of REM sleep time intervals with variance above the threshold exceeds the threshold during the REM sleep time interval. 15. The method of claim 1, wherein measuring at least one physiological signal of the patient during at least one rapid eye mo... | The Regents of the University of Michigan Michigan Technological University | A61B005/00 A61B005/04
US08029447 | B2 | 2011-10-04 | Multipurpose host system for invasive cardiovascular diagnostic measurement acquisition including an enhanced dynamically configured graphical display | 2007-10-10 | The present invention provides a multipurpose host system for processing and displaying invasive cardiovascular diagnostic measurement data. The system includes a an external input signal bus interface. The bus interface receives data arising from cardiovascular diagnostic measurement sensors. Measurement processing components receive data from particular sensor types. Based on the received data, the processing components render diagnostic measurement parameter values. A multi-mode graphical user interface includes display components corresponding to data received from particular sensor types. The user interface provides recommended action prompts that guide a user through a series of actions. | 1. A multipurpose host system for invasive cardiovascular diagnostic measurement acquisition and display, the system comprising: an external input signal bus interface for receiving data arising from one or more cardiovascular diagnostic measurement sensors attached to a flexible elongate member positioned within a vessel, the data being representative of at least one of pressure within the vessel and flow within the vessel;a plurality of measurement processing components for receiving data from the one or more cardiovascular diagnostic measurement sensors attached to the flexible elongate member and rendering diagnostic measurement parameter values of at least one of the pressure within the vessel and the flow within the vessel according to the received data; anda multi-mode graphical user interface host comprising a set of diagnostic measurement user interfaces including display components corresponding to data output rendered by specified ones of the plurality of measurement processing components, the user interface simultaneously displaying a plurality of buttons associated with actions of the one or more cardiovascular diagnostic measurement sensors attached to the flexible elongate member;wherein the user interface provides recommended action prompts automatically based on the data received from the one or more cardiovascular diagnostic measurement sensors attached to the flexible elongate member, the recommended action prompts providing visual indicators that guide a user through a series of actions by visually accentuating, in series, the plurality of buttons of the user interface associated with the actions of the one or more cardiovascular diagnostic measurement sensors attached to the flexible elongate member such that actuation of each of the plurality of buttons of the user interface results in the associated action of the one or more cardiovascular measurement sensors being executed;wherein the multipurpose host system determines a sensor type for each of the one or more cardiovascular diagnostic measurement sensors connected to the multipurpose host system, wherein the simultaneously displayed plurality of buttons associated with the actions of the one or more cardiovascular diagnostic measurement sensors are based on the sensor type for each of the one or more cardiovascular diagnostic measurement sensors. 2. The multipurpose host system of claim 1 wherein the user interface visually accentuates a particular button of the plurality of simultaneously displayed buttons associated with a particular action of the series of actions by displaying at least one other button of the plurality of simultaneously displayed buttons in a way indicating that execution of the action associated with the at least one other button is prevented. 3. The multipurpose host system of claim 2, wherein the user interface visually accentuates the particular button of the plurality of simultaneously displayed buttons associated with the particular action of... | Volcano Corporation | A61B005/02
US09924886 | B2 | 2018-03-27 | Medical measuring device | 2006-08-09 | The invention relates to a measuring device (1) for the non-invasive measurement of physiological parameters. The measuring device (1) is suited to detect and localize by way of self diagnosis diseases such as, for example, inflammations, tumors or arteriosclerosis. The invention proposes a measuring device (1) with at least one optical measuring unit (100) for the generation of oximetric and/or plethysmographic measuring signals, an evaluation unit (140) processing the measuring signals, and a unit (120, 130) for the acquisition of local tissue parameters such as fat content, water content and/or blood perfusion, with the evaluation unit (140) being designed such that at least one local metabolic parameter is determined, in particular the local oxygen consumption, from the signals furnished by the optical measuring unit and obtained from tissue parameters. Moreover, the measuring device (1) enables the non-invasive determination of the glucose concentration. | 1. A device for determining a local glucose concentration, the device comprising: at least one optical measuring unit, wherein said at least one optical measuring unit produces oximetric measurement signals and plethysmographic measurement signals,an ECG unit, wherein said ECG unit detects an ECG signal by way of two or more ECG electrodes,a bioelectric impedance measuring unit comprising electrodes wherein said bioelectric impedance measuring unit detects bioelectric impedance signals,an evaluation unit, anda heat measuring unit comprising at least one heat sensor, wherein the heat measuring unit determines local heat production, wherein local heat production is a difference arising between an arterial temperature measured by the at least one heat sensor and a temperature the skin surface would reach in case of a perfect thermal insulation,wherein the ECG unit and the bioelectric impedance measuring unit have at least one common electrode,wherein the ECG unit, the bioelectric impedance measuring unit, the at least one optical measuring unit, and the at least one heat sensor are arranged in a common measuring head,wherein said evaluation unit is adapted to determine a local glucose concentration from a combination of:the local heat production determined by the heat measuring unit,a local oxygen consumption determined from the measurement signals produced by the at least one optical measuring unit and the bioelectric impedance signals detected by the bioelectric impedance measuring unit, anda correction factor determined from the ECG signals detected by the ECG unit and the measurement signals produced by the at least one optical measuring unit. 2. The device according to claim 1, wherein the bioelectric impedance measuring unit is adapted to detect comprehensive body tissue parameters. 3. The device according to claim 1, wherein the at least one optical measuring unit includes an optical radiation source and a radiation sensor. 4. The device according to claim 1, further comprising an optical sensor for spatially resolved determination of the skin coloring. 5. The device according to claim 1, wherein the evaluation unit is further adapted to evaluate the variation in respect of time of a plethysmograpic signal detected by the at least one optical measuring unit, and wherein the evaluation unit is further adapted to determine a pulse wave speed from the variation in respect of time of the ECG signal and from the variation in respect of time of the plethysmographic signal. 6. The device according to claim 5, wherein the evaluation unit is adapted to evaluate the variation in respect of time of the pulse wave speed and to ascertain the composition of food ingested by a user of the device based on the basis of the variation in respect of time of the pulse wave speed from the time of food ingestion. 7. The device according to claim 1, wherein said evaluation unit is adapted to determine a blood glucose level of the user of the device using param... | Ingo Flore | A61B005/1455 A61B005/053 A61B005/00 A61B005/021 A61B005/024 A61B005/0456 A61B005/145
US09233255 | B2 | 2016-01-12 | System and method for providing event summary information using an encoded ECG waveform | 2006-09-15 | Systems and methods for transferring medical information from a first medical monitoring device to a second device are provided. Medical information from the first medical monitoring device is encoded as an ECG waveform and the ECG waveform having the encoded medical information is provided to the second device as an input ECG waveform. | 1. A defibrillator, comprising: electrodes configured to be attached to a patient through an electrode terminal and through which medical information of the patient is obtained;a controller coupled to the electrodes and operable to encode the medical information of the patient as an ECG waveform; anda connector coupled to the controller and comprising an adapter configured to be coupled to the defibrillator through the electrode terminal, the adapter further comprising snap connectors configured to be coupled to ECG lead cables and provide the ECG waveform having the encoded medical information therethrough,wherein the medical information includes both of an ECG of the patient and an event summary information for the patient, the controller further operable to switch the electrode terminal between use with the electrodes for obtaining medical information and use with the adapter for providing the ECG waveform having the encoded medical information therethrough. | KONINKLIJKE PHILIPS N.V. | A61N001/39 A61B005/00 A61B005/0402 A61B005/0432
US20080249388 | A1 | 2008-10-09 | SYSTEMS AND METHODS FOR COOLING OF INTRAVENOUS FLUID AND MONITORING OF IN VIVO CHARACTERISTICS | 2008-04-07 | Intravenous sensor systems and methods for determining in vivo characteristics are disclosed, as well as systems for cooling bags of intravenous fluid. The intravenous sensor systems include a probe configured for placement in a peripheral vein. The probe includes one or more sensors for obtaining sensed information, such as body temperature, blood oxygen saturation level and blood pressure. An interface unit determines the in vivo characteristics based upon sensed information received from the one or more sensors. The determined in vivo characteristics may be displayed on the interface unit and/or transmitted to a remote computer. | 1. An intravenous sensor system for monitoring in vivo characteristics, comprising: a probe comprising a sensor housing with one or more sensors for sensing in vivo characteristics; andan interface unit comprising a CPU and a display, the interface unit determining the in vivo characteristics based upon sensed information received from the probe and displaying the in vivo characteristics on the display. 2. The system of claim 1, the probe further comprising a needle for guiding at least a portion of the probe into a patient's peripheral vein. 3. The system of claim 2, the needle being retractable once the at least a portion of the probe is in the vein, the probe comprising a one-way valve at a distal end to prevent blood from entering the probe. 4. The system of claim 2, the needle being removable once the at least a portion of the probe is in the vein, the probe comprising a one-way valve at a distal end to prevent blood from entering the probe. 5. The system of claim 2, the probe being configured to penetrate an intravenous catheter inserted into the vein, the distal tip of the probe extending beyond the intravenous catheter and into the vein. 6. The system of claim 1, the one or more sensors being selected from the group including a temperature sensor, an oxygen sensor, and a blood pressure sensor. 7. The system of claim 7, wherein the in vivo characteristic comprises one or more of body temperature, oxygen level, and blood pressure. 8. The system of claim 1, the interface unit comprising one or both of an audio indicator and a visual indicator, the CPU activating the one or both indicators if any one of the determined in vivo characteristics exceeds a set threshold. 9. The system of claim 8, the interface unit comprising an input device for receiving input from a user to set the threshold. 10. The system of claim 1, wherein the interface unit comprises a transmitter and an antenna for transmitting the in vivo characteristics to a remote receiver. 11. The system of claim 1, wherein the interface unit comprises a memory for storing the in vivo characteristics. 12. A method for intravenous monitoring of in vivo characteristics, comprising: inserting a probe of an intravenous sensor system into a patient's peripheral vein;coupling an interface unit of the intravenous sensor system to the probe; anddetermining the in vivo characteristic based upon sensed information received from one or more sensors within the probe. 13. The method of claim 12, further comprising displaying the in vivo characteristic on a display of the interface unit. 14. The method of claim 12, further comprising activating an indicator to alert a user when one or more of the in vivo characteristics fall outside of a threshold range. 15. The method of claim 14, further comprising setting the threshold range via an input device. 16. The method of claim 12, wherein the in vivo characteristic comprises one or more of body temperature, oxygen level, and blood pressure... | THINKVILLAGE, LLC | A61B005/00 A61F007/00
US20080269581 | A1 | 2008-10-30 | METHOD AND APPARATUS FOR MEASURING BLOOD VOLUME | 2008-04-23 | In one aspect, a conductance catheter is provided for measuring the volume of a fluid. The conductance catheter comprises a series of electrodes and a circuit to compensate for variations in sensitivity of the electrodes in the catheter. In another aspect, a resistivity sensor is provided for determining the resistivity of a fluid. The sensor comprises a series of electrodes spaced such that the total distance between endmost electrodes does not exceed the diameter of the catheter deploying the sensor. | 1. A sensing tip for measuring the volume of a fluid comprising one or more electrodes for measuring said volume and a resistivity sensor disposed on said tip in communication with said fluid to incorporate a current measurement of resistivity during measurement of said volume. 2. The sensing tip according to claim 1 wherein said resistivity sensor comprises a series of resistivity electrodes spaced from each other such that a total distance between endmost ones of said resistivity electrodes does not exceed the diameter of said sensing tip. 3. The sensing tip according to claim 2 wherein said resistivity electrodes are diametrically spaced. 4. The sensing tip according to claim 2 wherein said resistivity electrodes are axially spaced. 5. The sensing tip according to claim 1 configured for continuously obtaining said measurement of resistivity. 6. A system for measuring volume of a fluid comprising a sensing tip having a plurality of pairs of electrodes, each pair of electrodes being connected to a circuit to compensate for variations in sensitivity of respective pairs of electrodes according to the positioning of said respective pairs along said sensing tip. 7. The system according to claim 6 wherein said circuit comprises individual gain and offset circuits to adjust the strength of a respective signal. 8. The system according to claim 7 comprising an output for displaying said respective signal to facilitate adjustment of said gain and offset circuits. 9. The system according to claim 6 comprising a control box for calibrating said sensing tip, said control box including a gain and offset circuit for each pair of electrodes to adjust the strength of a respective signal; and a monitor for displaying said respective signal to facilitate adjustment of said gain and offset circuits. 10. The system according to claim 9 wherein said control box obtains a resistivity measurement from said sensing tip and incorporates a value of resistivity into a calibration measurement. 11. The system according to claim 9 comprising a plurality of cuvettes for containing fluids of differing properties. 12. A method for calibrating a sensing tip used for measuring volume of a fluid comprising inserting said sensing tip into a plurality of cuvettes containing fluids having differing properties, and when in each well, said method comprises: obtaining a plurality of conductance signals using a plurality of electrodes on said sensing tip; adjusting said conductance signals to compensate for variations is sensitivity of said pairs of electrodes due to the positioning of said electrodes along said sensing tip; obtaining a measurement of resistivity; and using said conductance signals and said measurement of resistivity to calibrate said sensing tip. 13. The method according to claim 12 wherein said differing properties comprises differing volumes and said measurement of resistivity is obtained using a resistivity measurement on said sensing tip. 14. The method... | - | A61B005/055 G01F017/00 G01F023/00
US20080281247 | A1 | 2008-11-13 | Autonomic nervous activity monitor, blood processing apparatus, blood collecting apparatus and autonomic nervous activity monitoring method | 2008-07-11 | An autonomic nervous activity monitor capable of predicting occurrence of abnormal reaction depending on autonomic nervous activity and used for medical practice and the like is provided. A blood component collecting apparatus 100 carries out frequency analysis to a pattern of heartbeat which is acquired by an ECG 300 from a donor, and calculates a high frequency component value [msec2] defined as an indicator for parasympathetic activity and a ratio value of a low frequency component to the high frequency component defined as an indicator for sympathetic activity to detect a change in these values on a time-series base. When a judgment circuit 204 predicts that abnormal reaction may occur to a donor after blood collection operation is started, a controller 200 controls a pump 9 to decrease a blood drawing or blood returning speed, or, to stop blood drawing or blood returning operation. | 1. A blood processing apparatus that carries out blood drawing operation for drawing blood from a body and blood returning operation for returning blood to the body, comprising: a biological information acquisition unit which measures heartbeat or pulse of a body to acquire a pattern of the heartbeat or pulse as biological information;a calculation unit which calculates a value [msec2] of a high frequency component (a HF value) defined as an indicator for parasympathetic activity and a value of a ratio of a low frequency component (LF) to the high frequency component (a LF/HF ratio value) defined as an indicator for sympathetic activity on a time-series base by carrying out frequency analysis to the pattern of the heartbeat or pulse acquired by the biological information acquisition unit;a blood drawing unit which carries out the blood drawing operation to the body;a blood processing unit which processes blood drawn from the body by the blood drawing unit;a blood returning unit which returns to the body blood processed by the blood processing unit; anda controlling unit which controls the blood drawing unit, the blood processing unit and the blood returning unit, and which predicts whether abnormal reaction may occur to the body from a predetermined relationship between the HF value and the LF/HF ratio value, based upon the HF values and the LF/HF ratio values calculated by the calculation unit,wherein the controlling unit, when predicted that abnormal reaction may occur to the body after starting the blood drawing operation of the blood drawing unit, reduces a rate for drawing blood from the body according to the blood drawing unit or a rate for returning blood to the body according to the blood returning unit, or, stops the blood drawing operation of the blood drawing unit or the blood returning operation of the blood returning unit. 2. A blood processing apparatus according to claim 1, wherein, when alpha is defined as a predetermined coefficient, the controlling unit predicts that abnormal reaction may occur to the body when a sum of a period of {(the HF value calculated by the calculation unit)>(the coefficient alpha)×(the LF/HF ratio value calculated by the calculation unit)} becomes not less than 75% of a calculation period carried out by the calculation unit. 3. A blood processing apparatus according to claim 2, wherein the coefficient alpha ranges from 10 to 100 and the calculation period ranges from 1.25 to 10 [min]. 4. A blood collecting apparatus that carries out blood collection operation to a body, comprising: a biological information acquisition unit which measures heartbeat or pulse of a body to acquire a pattern of the heartbeat or pulse as biological information;a calculation unit which calculates a value [msec2] of a high frequency component (a HF value) defined as an indicator for parasympathetic activity and a value of a ratio of a low frequency component (LF) to the high frequency component (a LF/HF ratio value) defined a... | Japanese Red Cross Society TERUMO KABUSHIKI KAISHA | A61B005/00 A61M001/00
US20080275364 | A1 | 2008-11-06 | FLASHBACK BLOOD COLLECTION NEEDLE | 2007-12-21 | The invention is a self-venting blood collection needle assembly for the extraction of at least one fluid sample into an evacuated container for laboratory testing, this blood collection needle assembly providing a clear or translucent flashback chamber for blood to flow into, for visualization by the user to confirm successful vein entry. The self-venting mechanism permits escape of air during use, and which, typically, also prevents an outflow of fluid, such as blood. Thus, air under venous pressure will be allowed to escape from the blood collection needle assembly until blood reaches the venting mechanism. | 1. A needle assembly comprising: a hub having an inlet end, an outlet end and a chamber formed between said ends;an inlet cannula having a distal end and a lumen extending therethrough, said inlet cannula being mounted to said hub such that said distal end of said inlet cannula is external of said hub and such that said lumen through said inlet cannula communicates with said chamber;an outlet cannula having a proximal end and a lumen extending therethrough, said outlet cannula being mounted to said hub such that said proximal end of said outlet cannula is external of said hub and such that said lumen of said outlet cannula communicates with said chamber, wherein said inlet and outlet cannula are distinct or comprise a single cannula with an aperture providing communication with said chamber;a multiple sample sleeve mounted over portions of said outlet cannula disposed externally of said hub, wherein said inlet cannula lumen, said chamber, said outlet cannula lumen and the internal surface of said multiple sample sleeve form a fluid passage; anda venting mechanism providing communication between said fluid passage and ambient surroundings, wherein said venting mechanism permits an outflow of air from the needle assembly to said ambient surroundings through said venting mechanism, and wherein said venting mechanism substantially prevents an outflow of fluid from said needle assembly to said ambient surroundings through said venting mechanism,wherein said venting mechanism is located in said passage beyond said proximal end of said outlet cannula. 2. The needle assembly of claim 1, wherein said venting mechanism comprises a vent media. 3. The needle assembly of claim 1, wherein said hub further comprises a non-patient barb, and wherein said venting mechanism extends through said non-patient barb. 4. The needle assembly of claim 3, wherein said venting mechanism comprises a vent media. 5. The needle assembly of claim 3, wherein said venting mechanism comprises a passage through said non patient barb and a venting plug located in said passage. 6. The needle assembly of claim 3, wherein said venting mechanism comprises a slit through said non patient barb and a venting disc. 7. The needle assembly of claim 3, wherein said venting mechanism comprises at least one aperture through said non patient barb, and a venting plug located in said aperture. 8. The needle assembly of claim 1, wherein said hub further comprises a non-patient barb, and wherein at least a portion of said non-patient barb comprises said venting mechanism. 9. The needle assembly of claim 8, wherein said venting mechanism comprises a vent media. 10. The needle assembly of claim 8, wherein at least a portion of said non-patient barb is formed from a porous hydrophobic material or said portion of said non-patient barb comprises a porous material with a hydrophobic surface. 11. The needle assembly of claim 8, wherein said venting mechanism comprises a porous material impregnated w... | BECTON, DICKINSON AND COMPANY | A61B005/15
US20080243032 | A1 | 2008-10-02 | ANALYTICAL DEVICE INCLUDING STERILE PROTECTION | 2008-04-18 | An analytical device is provided comprising a lancet and a test element in an integrated arrangement. The lancet generally comprises a lancet needle and a protective cap wherein the lancet needle can be displaced relative to the protective cap and wherein the protective cap generally completely surrounds the tip of the lancet needle. The protective cap is typically comprised of an elastic or plastic material and is configured to generally seal an opening provided in a chamber of the test element. In certain embodiments, the lancet comprises a conductive material and can be configured for electrical connection to a detection and evaluation device in order to be used to measure the filling level of the sample liquid and/or to serve as a counter or reference electrode. The present invention also relates to a process for producing such an analytical device and the use thereof. | 1. An analytical device comprising a lancet and a test element in an integrated arrangement, the lancet comprising a lancet needle with a tip and a protective cap which generally completely surrounds the lancet needle at least in the area of the tip, wherein the lancet needle can be displaced in at least one direction relative to the protective cap, and wherein the lancet needle is configured to engage the protective cap after displacement of the lancet needle in the at least one direction and thereafter to pull the protective cap during displacement of the lancet needle in a retracting direction, the test element comprising a chamber containing a reagent system, the chamber having a first opening configured to be generally sealed by the protective cap prior to and during displacement of the lancet needle in the at least one direction and configured thereafter to be generally opened upon displacement of the lancet needle in the retracting direction. 2. The analytical device according to claim 1 wherein the protective cap comprises an elastomer. 3. The analytical device according to claim 1 wherein the protective cap encloses the needle tip in a sterile manner. 4. The analytical device according to claim 1 wherein the analytical device comprises structural means for transporting sample liquid. 5. The analytical device according to claim 4 wherein the structural means for transporting sample liquid comprises a structure selected from the group consisting of a capillary gap, a capillary channel, a wick comprising an absorbent material, and a cross-piece made of an absorbent material. 6. The analytical device according to claim 1 wherein the chamber further comprises a second opening and wherein the lancet further comprises a seal configured to generally seal the second opening prior to displacement of the lancet needle in the at least one direction. 7. The analytical device according to claim 6 wherein the protective cap comprises the seal. 8. The analytical device according to claim 6 wherein the lancet further comprises a catching device provided on the lancet needle, the catching device being configured to engage and pull the protective cap when the lancet needle is displaced in the retracting direction, and wherein when the lancet needle is displaced in the retracting direction, the protective cap in turn pulls the seal along with it and opens the second opening, wherein the second opening comprises an air vent. 9. The analytical device according to claim 1 wherein the lancet further comprises a catching device provided on the lancet needle, the catching device being configured to engage and pull the protective cap when the lancet needle is displaced in the retracting direction. 10. The analytical device according to claim 1 wherein upon complete displacement of the lancet needle in the retracting direction, the lancet is configured for electrical connection to a measuring device wherein the tip then comprises an electrode to be used fo... | - | A61B005/157
US20080249395 | A1 | 2008-10-09 | METHOD AND APPARATUS FOR CONTROLLING CATHETER POSITIONING AND ORIENTATION | 2007-04-06 | A method and apparatus for detecting position and orientation of catheter distal magnetic element end while moving in a patient's heart is described. The apparatus includes magnetic sensors for to detect the magnetic field of a generated by the catheter tip. Each sensor transmits the field magnitude and direction to a detection unit, which filters the signals and removes other field sources, such as generated by CGCI coils and external medical hardware. The method allows the measurements of magnitude corresponding to the catheter tip distance from the sensor and the orientation of the field showing the magnetic tip orientation. Since the tip's magnetic field is not necessarily symmetric, the position and orientation computation technique are not independent of each other. Hence, an iterative calculation is used to converge to a solution. The method of determining tip position is calculated by triangulation from each sensor. In one embodiment, the tip orientation is calculated by an intersecting-planes algorithm. The orientation is used to adjust the distances from each sensor, and the process is repeated until convergence for both position and orientation is achieved. The resultant value provides the actual catheter tip position and orientation (AP). The actual position is further filtered by synchronizing the AP measurements with the QRS signal of the heart, allowing the operator and CGCI controller to view the organ as a static object. | 1. A method for determining the position and orientation of an invasive medical device having a magnetic tip with a defined magnetic axis in a patient comprising: obtaining sensor measured data from a plurality of magnetic field sensors, said magnetic field sensors configured to sense a magnetic field produced by said tip;computing first sensor range estimates using, at least in part said sensor measured data, and storing said first sensor range estimate as previous sensor range estimates;computing a first position estimate of the tip using, at least in part said first sensor range estimates, and saving said first position estimate as a previous position estimate;computing a first orientation estimate a magnetic axis of the tip using, at least in part said first sensor range estimates, and saving said first orientation estimate as a previous orientation estimate;computing revised sensor range estimates using, at least in part, said previous orientation estimate, and storing said revised sensor range estimate as said previous sensor range estimates;computing a revised position estimate of the tip using, at least in part, said previous sensor range estimates, and saving said revised position estimate as said previous position estimate;computing a revised orientation estimate of a magnetic axis of the tip using, at least in part, said revised sensor range estimates, and saving said revised orientation estimate as said previous orientation estimate; andrepeating the process of computing revised sensor range estimates, computing a revised position estimate, and computing a revised orientation estimate until a desired accuracy is obtained for said revised position estimate and said revised orientation estimate. 2. The method of claim 1, where computing the revised orientation of the magnetic axis of the tip comprises: computing a first sensor-to-tip vector using the previous position estimate and a sensor position of a first sensor;obtaining a first magnetic field vector measured by said first sensor;computing a vector describing a first magnetic plane, at least in part by, using a cross product of the first sensor-to-tip vector with the first magnetic field vector;computing a second sensor-to-tip vector using the previous position estimate and a sensor position of a second sensor;obtaining second magnetic field vector measured by said second sensor;computing a vector describing a second magnetic plane, at least in part, by using a cross product of the second sensor-to-tip vector with the second magnetic field vector; andcomputing said revised orientation, at least in part by, computing an angular orientation of an intersection of the first magnetic plane and the second magnetic plane. 3. The method of claim 2, further comprising computing an average obtained by first computing said revised orientation when said first sensor and said second sensor comprise a first pair of sensors chosen from a plurality of sensors and then computing said revised orie... | - | A61B005/05
US07471974 | B2 | 2008-12-30 | Method for planning stimulation of hyper/hypometabolic cortical areas | 2003-09-12 | A method for planning the stimulation of hyper/hypometabolic cortical areas includes determining anatomical patient data using an imaging method and detecting positions of (i) the hyper/hypometabolic cortical areas in a patient's anatomy and (ii) a position of a stimulator. The positions of the hyper/hypometabolic cortical areas are registered and/or referenced with respect to the position of the stimulation means. An optimal positioning for the stimulator is determined on the basis of the relative positional information. | 1. A method for planning stimulation of hyper/hypometabolic cortical areas, the method comprising: simulating a field distribution for a stimulator relative to a position of the stimulator;determining a stimulation area for the stimulator relative to a position of the stimulator;determining functional anatomical patient data;determining structural anatomical patient data;navigationally registering the functional anatomical data with the structural anatomical data such that the functional anatomical data are available for navigation;based on the functional anatomical data, detecting positions of the hyper/hypometabolic cortical areas in a patient's anatomy;determining a position of a stimulator;navigationally registering the stimulation area of the stimulator;at least one of (i) registering or (ii) referencing the position of the hyper/hypometabolic cortical areas with respect to the position of the stimulator;determining an optimal positioning for the stimulator on the basis of relative positional information of the hyper/hypometabolic cortical areas and the stimulation area of the stimulator; anddisplaying the optimal positioning for the stimulator. 2. The method as set forth in claim 1, wherein the detecting step is performed using a medical navigation system. 3. The method as set forth in claim 1, wherein the stimulation is planned of hypermetabolic areas related to the manifestation of systemic tinnitus. 4. The method as set forth in claim 1, wherein the functional image detection method includes at least one of (i) functional magnetic resonance detection and (ii) positron emission tomography (PET). 5. The method as set forth in claim 1, wherein the detecting step includes using a navigation system to optically detect arrangements of actively or passively emitting markers arranged on the patient's head and on the stimulation means. 6. The method as set forth in claim 1, wherein the detecting step includes using a navigation system to magnetically or inductively detect (i) at least one of (a) positional coils and (b) oscillating circuits, arranged on the patient's head and on the stimulator. 7. The method as set forth in claim 1, wherein the stimulator includes a cortical stimulation coil. 8. The method as set forth in claim 2, further comprising: outputting detected navigational data together with the determined optimal positioning on an image output. 9. The method as set forth in claim 1, further comprising: calibrating the stimulator. 10. A computer program storage medium comprising a program which, when it is run on a computer or is loaded onto a computer, causes the computer to perform a method in accordance with claim 1. 11. A method of stimulating hyper/hypometabolic cortical areas of a patient, the method comprising: simulating a field distribution for a stimulation coil relative to a position of the stimulation coil;determining a stimulation area for the stimulation coil relative to a position of the coil;determining functio... | BrainLAB AG | A61B005/05
US07728591 | B2 | 2010-06-01 | Imaging region-specific radio frequency coils for MRI | 2006-10-03 | A radio frequency coil includes a non-cylindrical conformal surface (62, 76) that substantially conforms with a magnetic resonance subject. A plurality of conductor loops (60, 71, 72, 73, 74) are disposed in or on the non-cylindrical conformal surface. The plurality of conductor loops are configured to produce a substantially uniform Bi field in the magnetic resonance subject responsive to energizing at a Bi frequency. Optionally, a plurality of load-compensating conductor loops (90) are disposed in or on a compensatory non-cylindrical conformal surface (62) that substantially conforms with a magnetic resonance subject. The plurality of load-compensating conductor loops are configured to produce a non-uniform Bi field that compensates for a loading Bi non-uniformity caused by the magnetic resonance subject. Moreover, the coil may comprise switching means for switching the coil between a first mode of operation (e.g. a volume transmit mode) and a second mode of operation (e.g. a phased array reception mode). | 1. A transmit/receive radio frequency coil for parallel imaging comprising: a non-cylindrical conformal surface that substantially conforms with a magnetic resonance subject;a plurality of conductor loops disposed in or on the non-cylindrical conformal surface;one or more switches connected with the conductor loops, the one or more switches being selectively switched to (i) interconnect the plurality of conductor loops configured as a volume transmit resonator to produce a substantially uniform B1 field across an imaging volume of the magnetic resonance subject responsive to energizing at a B1 frequency in a transmit mode and (ii) de-couple the conductor loops in a receive mode to define a receive coil array of decoupled conductor loops which each receive resonance signals from across the imaging volume, the conductor loops each being resonant at the B1 frequency in both the transmit mode and the receive mode;wherein the plurality of conductor loops when energized at the B1 frequency define a discretized current density across the non-cylindrical conformal surface that corresponds with the substantially uniform B1 field in the magnetic resonance subject, and at least one of: (i) wherein the plurality of conductor loops define the discretized current density when energized in the quasi-static domain, capacitance along the conductors maintaining the defined discretized current density at the B1 frequency, and(ii) wherein at least some of the conductor loops are electromagnetically coupled by mutual inductance therebetween. 2. The radio frequency coil as set forth in claim 1, wherein the plurality of conductor loops define the discretized current density when energized in the quasi-static domain, capacitance along the conductors maintaining the defined current density at the B1 frequency. 3. The radio frequency coil as set forth in claim 1, wherein at least some of the conductor loops are electromagnetically coupled by mutual inductance therebetween. 4. The radio frequency coil as set forth in claim 1, wherein the one or more switches interconnect the plurality of conductor loops in a series interconnected configuration in the transmit mode. 5. The radio frequency coil as set forth in claim 1, wherein the one or more switches galvanically interconnect the plurality of conductor loops in the volume transmit resonator configuration (i) and disconnect the plurality of conductor loops in the receive coil array configuration (ii). 6. The radio frequency coil as set forth in claim 1, wherein the non-cylindrical conformal surface is rigid and non-planar. 7. The radio frequency coil as set forth in claim 1, further including: a plurality of load-compensating conductor loops disposed in or on a compensatory non-cylindrical conformal surface that is configured to substantially conform with a magnetic resonance subject, the plurality of load-compensating conductor loops configured to produce a non-uniform B1 field that compensates for a loading B1 non-un... | Koninklijke Philips Electronics N.V. | G01V003/00
US07464043 | B1 | 2008-12-09 | Computerized method and system for obtaining, storing and accessing medical records | 2000-07-14 | Computerized method and system for managing respective health records of a plurality of patients are provided. The method allows for uploading a progress note of a respective patient. The progress note includes data relative to an encounter between a respective physician and the respective patient. The method further allows for identifying on the progress note respective parameters selectable by the respective physician. A storing step allows for storing the progress note with the identified parameters in a database accessible to a plurality of authorized users. A populating step allows for populating the database with respective progress notes resulting from further encounters between the respective patient and any respective physician so as to create a historical set of progress notes for that respective patient. | 1. A computerized method for managing respective health records of a plurality of patients, said method comprising: uploading a progress note of a respective patient, said progress note comprising data relative to an encounter between a respective physician and the respective patient;identifying on said progress note respective parameters selectable by the respective physician;storing said progress note with said identified parameters in a computer database accessible to a plurality of authorized users; andpopulating said computer database with respective progress notes and respective identified parameters resulting from further encounters between the respective patient and any respective physician so as to create a historical set of progress notes with identified parameters for that respective patient, the set of historical progress notes being interconnectable based on one or more logic operators. 2. The computerized method of claim 1, wherein the identified parameters are selected to convey a snapshot of said encounter. 3. The computerized method of claim 1, wherein the identified parameters are selected from the group consisting of diagnosis and prescription parameters. 4. The computerized method of claim 1, wherein one of the logical operators comprises a chronology-indicative operator. 5. The computerized method of claim 1, wherein one of the logical operators comprises a pathology-indicative operator. 6. The computerized method of claim 1, wherein one of the logical operators comprises a pharmacology-indicative operator. 7. The computerized method of claim 1 further comprising: tracking users accessing information in the database to process respective billing of the accessing users for each access of the database, andallocating fees among entities associated with the respective information accessed by respective users. 8. The computerized method of claim 1, wherein the database is accessible to a plurality of users through a communications network. 9. The computerized method of claim 8, wherein the communications network comprises the Internet. 10. The method of claim 1, further comprising controlling access of the database according to ownership interests assigned to information in the database. 11. A computer readable medium encoded with computer code for managing respective health records of a plurality of patients, the program code causing a computer to execute a method comprising: uploading a progress note of a respective patient, said progress note comprising data relative to an encounter between a respective physician and the respective patient;identifying on said progress note respective parameters selectable by the respective physician;storing said progress note with said identified parameters in a database accessible to a plurality of authorized users; andpopulating said database with respective progress notes and respective identified parameters resulting from further encounters between the respective patient and any r... | - | A61B005/00
US07438690 | B2 | 2008-10-21 | Peak expiratory flow meter capable of measuring continuous expiratory flow signal | 2004-02-20 | Disclosed herein is a peak expiratory flow meter capable of measuring continuous expiratory flow. The peak expiratory flow meter includes a lower casing unit (110), an air expansion unit (130) and an upper casing unit (150). The lower casing unit (110) includes a first body (112) having an open top and open opposite sides and having a first end in which a first expiratory flow entrance (114) is formed, a plurality of rails (118), and a movable plate (122). The air expansion unit (130) includes a pressure transfer tube (132), an elastic plate (136), and a pressure spring (140). The upper casing unit (150) includes a second body (152) adapted to define the air expansion unit and provided with a first end in which a second expiratory flow entrance (154) is formed, and a measurement slot (156), and a scale indicator (160) having a “T” shape. | 1. A peak expiratory flow meter capable of measuring continuous expiratory flow, comprising: a lower casing unit including a first body formed to have an open top and open opposite sides and to have a first end in which a first expiratory flow entrance is formed to be integrated with the first body so as to allow an asthma patient to hold the first expiratory flow entrance in his or her mouth, a plurality of rails arranged on a top of the first body adjacent to a second end of the first body, and a movable plate fitted on the rails;an air expansion unit including a pressure transfer tube arranged on a side of the first expiratory flow entrance, an elastic plate fitted on the pressure transfer tube, and a pressure spring comprised of a first end that is mounted to the elastic plate and a second end that is mounted on a first side surface of the movable plate; andan upper casing unit including a second body coupled with the first body, adapted to define the air expansion unit and provided with a first end in which a second expiratory flow entrance corresponding to the first expiratory flow entrance is formed to be integrated with the second body, and a measurement slot formed on a side of the second expiratory flow entrance and extended to a portion of the second body adjacent to a second end of the second body, and a scale indicator formed to have a “T” shape and to have an upper portion exposed to outside of the measurement slot and hung on a top of the second body and a lower portion supported on an upper portion of a second side surface of the movable plate, the scale indicator being moved by the movable plate. 2. The peak expiratory flow meter according to claim 1, wherein the rails have ends that are fixedly arranged by supporting projections vertically extended from a bottom surface of the first body. 3. The peak expiratory flow meter according to claim 1, wherein the lower casing unit has a pressure transfer hole that is formed on a side thereof to penetrate through a bottom surface of the first body. 4. The peak expiratory flow meter according to claim 3, wherein the pressure transfer tube includes a first transfer tube that is formed to have a closed top and an open bottom to penetrate from the top of the first body through the pressure transfer hole, and a second transfer tube that is extended from an outer circumference of the first transfer tube, arranged in the first body, to be radially broadened towards the movable plate while communicating with the first transfer tube, and then mounted to the elastic plate, the pressure transfer hole being tightly fitted around the first transfer tube, which passes therethrough, without circulation of air therebetween. 5. The peak expiratory flow meter according to claim 4, wherein the first transfer tube extended to an outside of the first body is selectively connected to any one of a typical pressure sensor and a typical electronic circuit for signal analysis, thus performing spirometry. 6.... | - | A61B005/08
US07435232 | B2 | 2008-10-14 | Noninvasive tissue assessment | 2004-09-07 | Methods and apparatus for non-invasively assessing physiological hard of soft tissue of human and other species are described. In a preferred embodiment, tissue is vibrationally stimulated in vivo through a frequency spectrum. The tissue reacts against the stimulus and the reaction is preferably measured and recorded. Based on analytical algorithms or comparisons with previously taken measurements, changes within the tissue can be detected and used for diagnostic purposes. Further embodiments describe the usage of the device and methods for in vivo intra-operative and post-operative implant evaluations and as a therapeutic tool. | 1. A method for tissue quality assessment comprising: (a) engaging a tissue with a handheld device at a contact point and with a contact load, wherein the handheld device comprises a vibration generator for inducing mechanical vibrations over a range of frequencies at the contact point and one or more sensors for measuring the applied force of the mechanical vibrations and for measuring the tissue response at the same contact point;(b) initiating a frequency sweep at the same contact point in (a) with the handheld device;(c) simultaneously performing the frequency sweep over the range of frequencies based on a given set of operational parameters while maintaining the contact load to induce mechanical vibrations in the tissue at the same contact point in (a) andmeasuring the applied force of the mechanical vibrations and the tissue response at the same contact point in (a) with the handheld device to assess tissue quality. 2. The method according to claim 1 wherein the frequency range is between 20 Hz and 1.5 MHz. 3. The method according to claim 2 wherein the frequency range is between 20 Hz and 600 kHz. 4. The method according to claim 2 wherein the frequency range is between 800 Hz and 1.5 MHz. 5. The method according to claim 1 further comprising selecting the operational parameters prior to engaging the tissue with the handheld device. 6. The method according to claim 5 wherein the operational parameters comprise one or more inputs selected from the group consisting of anatomic site, age, sex, body weight and shoe size. 7. The method according to claim 1 wherein providing a user response comprises calculating a relative value based on comparative data from a database. 8. The method according to claim 7 wherein the user response comprises a diagnostic value for osteoporosis. 9. The method according to claim 1 wherein providing a user response comprises calculating a relative value based on an analytical method. 10. The method according to claim 1 wherein the user response comprises a diagnostic value for bone fractures. 11. The method according to claim 1 wherein the vibration generator comprises a shaker. 12. The method according to claim 11 wherein the shaker comprises a device selected from the group consisting of piezoelectric devices, electromechanical devices, hydraulic devices and pneumatic devices. 13. The method according to claim 11 wherein the shaker comprises an electromagnetic device. 14. The method according to claim 1 wherein one or more sensors comprises an impedance transducer. 15. The method according to claim 14 wherein the impedance transducer measures both input and output signals. 16. The method according to claim 14 wherein the impedance transducer comprises one or more selected from the group consisting of accelerometers, displacement transducers, load cells, and combinations thereof. 17. The method according to claim 1 wherein the device further comprises a means for determining the contact load. 18. T... | William Marsh Rice University | A61B005/103 A61B005/117 A61B005/00 A61B007/00 A61B005/04 A61H031/00 A61H001/00 A61B017/00
US20080306357 | A1 | 2008-12-11 | Physiological monitoring system for a computational device of a human subject | 2008-07-17 | A system for monitoring at least one physiological parameter of a human subject. The system of the present invention features a device with which the human subject regularly interacts, and which is connected to the computational device of the human subject for automatic collection of at least one physiological parameter which is also of medical interest. The device features at least one physiological sensor for collecting the measurement of the physiological parameter. The computational device of the human subject then preferably operates a software program to analyze the data which is collected, in order for the human subject to receive an alert when necessary. Alternatively or additionally, the collected data is sent to a remote computational device which is in communication with the computational device of the human subject for analysis. Optionally, the present invention enables the human subject to receive an alert if a deterioration in the physiological condition of the human subject is detected, thereby enabling the human subject to start preventive medical treatment with trained medical personnel as soon as possible. Thus, the awareness of the human subject about any incipient medical problem is immediately improved, which may result in an increased probability of being able to successfully treat and/or otherwise ameliorate those problems. | 1-34. (canceled) 35. A system for non-invasive monitoring of a human subject, comprising: a peripheral device of a host computational device for location in physical contact with the human subject having a regular user interaction in relation to said host computational device;a sensor contained within said peripheral device, configured for sensing the human subject, said sensor automatically collecting data about a physiological parameter of the human subject during said regular user interaction;agent software processed by said host computational device, said agent software configured to receive, from said peripheral device, during said regular user interaction, data relating to said monitoring from said sensor for monitoring the human subject and for analyzing said data relating to said monitoring from said sensor, thereby to determine whether a physiological parameter of the human subject indicates a deterioration in the physiological condition of the subject according to an analysis of said data. 36. The system of claim 35, wherein said peripheral device is contacted by the human subject during normal interactions with said host computational device, such that said data is collected without requiring any additional action by the human subject. 37. The system of claim 35, wherein said agent software determines whether a physiological parameter of the human subject indicates a deterioration in the physiological condition of the subject according to an analysis of said data. 38. The system of claim 36, wherein said agent software determines whether a physiological parameter of the human subject indicates a deterioration in the physiological condition of the subject according to an analysis of said data. 39. The system of claim 37, wherein said agent software is configured to provide an alert to the human subject if said deterioration occurs. 40. The system of claim 35, further comprising: at least one of a remote server in communication with the host computational device for receiving, processing and storing said data for monitoring the human subject, or a manually operated call center for receiving said data and a medical history of said human subject and for diagnosing said human subject according to said data and said medical history. 41. The system of claim 39, wherein said alert is automatically transmitted to said human subject. 42. The system of claim 35, wherein said peripheral device includes a device for transmitting at least one command of the human subject to said host computational device for operating said host computational device. 43. The system of claim 35, wherein said peripheral device is at least one of a computer keyboard and a computer pointing device. 44. The system of claim 43, wherein said computer pointing device is a computer mouse. 45. The system of claim 35, wherein said peripheral device comprises a touch sensitive screen. 46. The system of claim 35, wherein said peripheral device is in physical contact w... | Medic4All INc. | A61B005/00 G06Q050/00
US07444177 | B2 | 2008-10-28 | EKG recording accessory system (EKG RAS) | 2003-11-21 | The invention is a precordial pad for positioning EKG electrodes on a patient for anatomically correct and repeatable placement. Data can be transmitted from the EKG pad of the invention by wire or wireless means. The pad includes a sizing aid, and a positioning device. The invention is also a system for obtaining and sending EKG data. | 1. A disposable EKG precordial pad comprising: a pad body for conductive attachment to a patient's torso, with a first and a second surface, with a plurality of electrodes which are embedded in said pad body but which extend from said pad body before application for contact with a patient, and with an adhesive surface on said second surface for contact with said patient's skin;a temperature sensor;an adhesive cover, which is removable for exposing said adhesive surface before use;at least one electrode extension device, for causing at least one of said plurality of electrodes to extend beyond said adhesive surface when said adhesive cover is removed, wherein said electrode extension device is a biased member mounted between each of said plurality of electrodes and said pad,said electrode extension device is held in biased position by said adhesive cover, andsaid electrode extension device moves at least one of said plurality of electrodes away from said pad when said adhesive cover is removed; anda positioning device for use with a patient for orienting and positioning said pad body for correct positioning on said patient. 2. The disposable EKG precordial pad of claim 1 which further includes a data transmitting module, for sending a plurality of signals from said embedded electrodes to an EKG machine. 3. The precordial pad of claim 2 in which the data transmitting module is a wireless transmitter. 4. The precordial pad of claim 2 in which the data-transmitting module is a wire cable with individual wires. 5. The disposable EKG precordial pad of claim 2 which further includes a circuit layer which is located between said first and said second surface, and which includes electronic connections between said electrodes and said data transmitting module. 6. The disposable EKG precordial pad of claim 5 in which said circuit layer is comprised of an insulating sheet on which is placed electrical connections in the form of conductive pathways. 7. The disposable EKG precordial pad of claim 6 in which said conductive pathways are metallic ink circuitry. 8. The precordial pad of claim 2 in which said electrodes each have a micro-transmitter for wireless transmission of a signal to said data transmitting module. 9. The precordial pad of claim 2 in which said data transmitting module further includes a temperature display for indicating the patient's body temperature. 10. The precordial pad of claim 2 in which the data-transmitting module is a wire harness for connection to an EKG machine. 11. The disposable EKG precordial pad of claim 1, in which said plurality of electrodes includes one or more micro-transmitters for sending a signal from at least one of said plurality of electrodes. 12. The precordial pad of claim 1 in which said electrode extension device is a spring. 13. The precordial pad of claim 1 in which said electrode extension device member is a foam structure. 14. The precordial pad of claim 1 in which said electrodes are packaged u... | - | A61B005/04
US07442169 | B2 | 2008-10-28 | Methods of distinguishing between vasoconstriction and vasodilation as a cause of hypotension | 2005-03-04 | The present invention relates to the use of the arterial pressure waveform recorded invasively or non-invasively, to distinguish between compensatory vasoconstriction with low cardiac output on the one hand from vasodilation from more severe organ damage on the other, as a cause of hypotension in acute emergencies. The waveforms may be subjected to harmonic analysis and the moduli of their harmonic components compared whereby a hypotensive individual can be confirmed to have the higher (second and above) greater than the first harmonic is considered as having vasoconstriction as a cause of hypotension. | 1. A method for providing a diagnosis as to whether hypotension in a patient is due to vasodilation or is associated with compensatory vasoconstriction, said method comprising the steps of: (i) measuring the arterial waveforms from a periphery artery,(ii) performing an harmonic analysis on said measured waveforms and comparing the moduli of the harmonic components,(iii) diagnosing the cause of hypotension as being vasodilation if the modulus of the second harmonic is less than the modulus of the first harmonic, and(iv) diagnosing hypotension as being associated with vasoconstriction if the modulus of the first harmonic is less than the modulus of the second harmonic. 2. A method according to claim 1, part (iii), comprising the further steps of assessing the patient's heart rate and determining the presence of absence of significant arrhythmia, and wherein, in the absence of significant arrhythmia and where the heart rate is less than 120 beats per minute, diagnosing the cause of hypotension as being vasodilation. 3. A method according to claim 1, part (iv) and wherein said harmonic analysis includes the step of determining the ratio of the first harmonic to the sum of the subsequent harmonics, the further step of diagnosing hypotension as being associated with vasoconstriction, if the ratio is much less than 0.5, and diagnosing the cause of hypotension as being vasodilation if the ratio is much greater than 1.0. | AtCor Medical Pty Limited | A61B005/02
US07452332 | B2 | 2008-11-18 | Report on changes over time in blood vessel elasticity indexes and bioinformation output apparatus | 2006-06-16 | A report on changes over time is generated and outputted that presents, in addition to the changes of blood vessel elasticity indexes, the changes over time of other measured values of bioinformation related to vascular disease such as a blood pressure value in time series. This report is useful for a diagnosis with respect to vascular disease from various aspects. | 1. A method of reporting changes over time in blood vessel elasticity indexes to present the changes over time of measured results of blood vessel elasticity indexes with respect to a common subject, comprising: providing a graph area with a horizontal axis representing time and a vertical axis representing values of blood vessel elasticity indexes;plotting in time series each of the blood vessel elasticity indexes by positioning a first indicator at a coordinate defined by a measured time and a value in the graph area;plotting in time series a measured value of bioinformation relative to vascular disease other than the blood vessel elasticity index and measured together with the blood vessel elasticity index in the graph area in association with the corresponding blood vessel elasticity index;representing the measured value of bioinformation by another vertical axis of the graph area; andplotting the measured value of bioinformation with a second indicator located close to the first indicator of the corresponding blood vessel elasticity index and on a coordinate corresponding to the measured value of bioinformation. 2. The method according to claim 1, wherein the measured value of bioinformation is a blood pressure value. 3. The method according to claim 2, wherein the second indicator comprises three marks which are arranged in a vertical direction to present a systolic blood pressure value, a mean blood pressure value, and a diastolic blood pressure value. 4. The method according to claim 3, wherein the second indicator further comprises a straight line connecting the three marks. 5. The method according to claim 3, wherein at least one of the three marks has a different shape or size from the others. 6. The method according to claim 3, wherein the shape or size of the marks changes depending on the relationship between the normal ranges predetermined for each of the systolic blood pressure value, the mean blood pressure value, and the diastolic blood pressure value and the measured values illustrated by the marks. 7. The method according to claim 1, wherein the second indicator has an asymmetry shape, the directionality of which indicates whether the measured value of bioinformation is for right or left side body. 8. The method according to claim 1, wherein the second indicator which is offset to the right or left of the corresponding blood vessel elasticity index indicates whether the measured value of bioinformation is for right or left side body. 9. The method according to claim 1, wherein the graph area is divided into a plurality of areas depending on the values of the blood vessel elasticity index. 10. A bioinformation output apparatus, comprising: obtaining means for obtaining from a storage device, measured values of at least a blood vessel elasticity index and other bioinformation related to one or more vascular diseases which are measured in a common subject;report generating means for generating a report on changes over ti... | Fukuda Denshi Co., Ltd. | A61B005/02 A61B005/04
US07449001 | B2 | 2008-11-11 | Sampling device and method for obtaining samples of internal body substances and method for producing a sampling device | 2002-05-17 | A sampling device for obtaining samples of internal body substances in the digestive system is a swallowable capsule which allows a sample of the body substance to enter the capsule through an inlet which opens in the digestive tract following contact with the substance to be collected. The capsule inlet initially is sealed and, after the capsule has been swallowed, the inlet opens following contact with the body substance. An inner chamber preserves a vacuum or substantial under pressure when the inlet is sealed by a blocking member adjacent to the inlet. The blocking member is elastic and has a flow permitting configuration which admits the body substance into the inner chamber and a flow preventing configuration which blocks the inlet from the inside of the chamber when the pressure difference has been equalized. | 1. A sampling device for obtaining samples of internal body substances in the digestive tract of a patient, said device having a shape of a swallowable capsule (2) which allows a sample of the body substance to enter the capsule through at least one inlet opening (18) which is opened in a predetermined position of the digestive tract following contact with the body substance to be collected, characterised in that the capsule (2) comprises two members, a cap member (4) and a body member (6), which are permanently joined to each other,a capsule wall (3) comprising said inlet opening (18), which initially is sealed and, when the patient has swallowed the capsule (2), is opened in said predetermined position following contact with the body substance to be collected,an inner chamber (5), defined by said capsule wall (3) and disposed to preserve a vacuum or underpressure relative to the environment when the inlet opening (18) is sealed, anda separate blocking member (12) disposed within the inner chamber (5) adjacent to and spaced from the inlet opening (18) in the capsule wall (3), said blocking member (12) being flexible and having at least one laterally located aperture (16) for a through-flow of body substance into the inner chamber (5),said blocking member having a configuration, such that,when the inlet opening (18) has been opened following contact with said body substance, the blocking member (12) has a flow permitting configuration which admits a flow of body substance into the inner chamber (5) as long as a there is a pressure difference between the inner chamber (5) and the external environment of the capsule (2) and a flow preventing configuration which blocks the inlet opening (18) from the inside of the chamber (5) when said pressure difference has been equalised by the flow of body substance into the capsule (2). 2. A sampling device according to claim 1, characterised in that the inlet opening (18) is sealed by a plug member (20) of a material that is dissolved following contact with said body substance. 3. A sampling device according to claim 2, characterised in that the plug member (20) consists of two or more layers of different materials and that each layer is dissolved in different positions in the digestive tract following contact with a specific body substance. 4. A sampling device according to claim 1, characterised in that the blocking member (12) consists of an elastic, self-sealing membrane, which in said flow preventing configuration sealingly bears on the inside of the capsule wall (3), such that an outflow of the body substance in the inner chamber (5) is prevented. 5. A sampling device according to claim 4, characterised in that said blocking member (12) is retained in the capsule wall (3) between said cap member (4) and said body member (6). 6. A sampling device according to claim 5, characterised in that the blocking member (12) is bell-shaped. 7. A sampling device according to claim 6, characterised in that a per... | Fargklamman AB | A61B005/00 A61B010/00 A61F002/02 B65D081/00
US07431699 | B2 | 2008-10-07 | Method and apparatus for third heart sound detection | 2003-12-24 | A cardiac rhythm management system includes a heart sound detector providing for detection of the third heart sounds (S3). An implantable sensor such as an accelerometer or a microphone senses an acoustic signal indicative heart sounds including the second heart sounds (S2) and S3. The heart sound detector detects occurrences of S2 and starts S3 detection windows each after a predetermined delay after a detected occurrence of S2. The occurrences of S3 are then detected from the acoustic signal within the S3 detection windows. | 1. A heart sound detection system, comprising: a cardiac signal input to receive a cardiac signal indicative of ventricular events;an acoustic signal input to receive an acoustic signal indicative of at least second heart sounds (S2) and third heart sounds (S3); anda heart sound detector coupled to the cardiac signal input and the acoustic signal input, the heart sound detector including: an S2 window generator to generate an S2 window after a first predetermined delay starting with each of the ventricular events;an S2 detector to detect S2 using an S2 threshold during the S2 windows;an S3 window generator, coupled to the S2 detector, to generate an S3 window after a second predetermined delay starting with each of the detected S2;an S3 detector to detect S3 using an S3 threshold during the S3 windows;a measurement module to detect an acoustic energy using the acoustic signal; andan S3 threshold generator adapted to determine the S3 threshold based on the acoustic energy. 2. The system of claim 1, further comprising an envelope detector coupled between the acoustic signal input and the heart sound detector, the envelope detector including a rectifier and a low-pass filter coupled to the rectifier. 3. The system of claim 2, further comprising an averaging circuit to improve a signal-to-noise ratio of the acoustic signal by ensemble averaging. 4. The system of claim 3, wherein the averaging circuit is adapted to align segments of the acoustic signal using the ventricular events. 5. The system of claim 3, wherein the averaging circuit is adapted to align segments of the acoustic signal using occurrences of the S2 detection. 6. The system of claim 1, further comprising an S1 detector to detect first heart sounds (S1). 7. The system of claim 6, further comprising a measurement module, coupled to the heart sound detector, to produce parameters related to one or more of S1, S2, and S3. 8. The system of claim 7, wherein the measurement module comprises an amplitude detector to detect one or more of an S1 amplitude, an S2 amplitude, and an S3 amplitude. 9. The system of claim 8, wherein the S2 detector comprises an S2 comparator to compare the acoustic signal to a predetermined S2 threshold. 10. The system of claim 9, wherein the S3 detector comprises an S3 comparator to compare the acoustic signal to the S3 threshold. 11. The system of claim 8, wherein the S2 detector further comprises an S2 threshold generator adapted to determine the S2 threshold based on the detected S1 and S2 amplitudes. 12. The system of claim 8, wherein the S3 threshold generator is adapted to dynamically adjust the S3 threshold. 13. The system of claim 8, wherein the measurement module is adapted to detect one or more of an energy associated with S1, an energy associated with S2, an energy associated with S3, a total acoustic energy during a cardiac cycle, and a total acoustic energy during systole of the cardiac cycle. 14. The system of claim 13, wherein the S3 detec... | Cardiac Pacemakers, Inc. | A61B005/02
US07445600 | B1 | 2008-11-04 | Multi-function, self-service health kiosk | 2005-03-01 | In one aspect, the present invention provides a multi-function, self-service health kiosk. In another aspect, the present invention provides a wheelchair accessible health kiosk. In another aspect, the present invention provides a blood pressure measurement apparatus for use in a health kiosk. | 1. A blood pressure and body fat measurement method, comprising: (1) using a blood pressure measurement apparatus, said blood pressure measurement apparatus comprising: a tubular or semi-tubular support member having an inner surface and an outer surface; andan inflatable cuff for receiving a subject's arm; and(2) using a body fat analyzer, said body fat analyzer comprising: a first infrared emitting diode for emitting infrared radiation having a single wavelength;a second infrared emitting diode for emitting infrared radiation having a single wavelength that is the same as the wavelength of the radiation emitted by the first infrared emitting diode; anda detector positioned midway or about midway between the first and second diodes, whereinsaid body fat analyzer is disposed such that when a subject places an arm fully into the inflatable cuff, the diodes are disposed next to a triceps region of the arm, said step of using the body fat analyzer comprises inserting an arm into the cuff in such a way that the diodes are disposed next to a triceps region of the arm, andsaid body fat analyzer measures body fat using only the single wavelength. 2. The method of claim 1, wherein said cuff has a first end and a tail portion, the first end being fixed to the inner surface of the support member and the tail portion passing through a slit in the support member and being at least partially wound around a roller, said cuff comprising a casing housing an inflatable bladder and a hard yet flexible frame. 3. The method of claim 2, wherein said frame is generally planar when not flexed. 4. The method of claim 2, further comprising a ratchet wheel attached to one end of said roller. 5. The method of claim 4, further comprising a ratchet and pawl configured to lock said gear when the apparatus is in a measurement mode so that said roller can not rotate while the apparatus is in the measurement mode. 6. The method of claim 5, wherein said bladder is configured to inflate in response to the apparatus being placed in the measurement mode. 7. The method of claim 5, wherein the apparatus is placed in the measurement mode in response to activation of a start measurement button. 8. A blood pressure and body fat measurement method, comprising: obtaining a blood pressure measurement apparatus, said blood pressure measurement apparatus comprising an inflatable cuff for receiving a subject's arm; andobtaining a body fat analyzer, said body fat analyzer comprising: a first infrared emitting diode for emitting infrared radiation having a single wavelength;a second infrared emitting diode for emitting infrared radiation having a single wavelength that is the same as the wavelength of the radiation emitted by the first infrared emitting diode; anda detector positioned midway or about midway between the first and second diodes, whereinsaid body fat analyzer is disposed such that when a subject places an arm fully into the inflatable cuff, the diodes are disposed next to a ... | Futrex, Inc. | A61B005/00 A61B006/00
US07447534 | B1 | 2008-11-04 | Method and apparatus for measuring electro-physiological activity | 2005-03-21 | An electro-optic modulator modulates light in response to a life-form bio-potential to provide a bio-potential modulated light output. Modulation can be carried out as a result of capacitive coupling with the bio-potential. Phase, polarization, and intensity light modulation can be used. The light output can then be converted to electronic form for further processing. An optical power splitter is used to split light into multiple portions that are converted to electrical form. A pilot tone is used to achieve high sensitivity and synchronize multiple electro-optic modulators while phase modulation can be used to reduce noise in the bio-potential signal. Low-coherent light may be used as the light source. A housing is provided that can be hermetically-sealed and provided with electro-magnetic shielding. | 1. A method for measuring electro-physiological activity comprising: a) providing a life form having a bio-potential;b) providing an electro-optic modulator;c) providing light to said electro-optic modulator;d) modulating said light in said electro-optic modulator with said bio-potential to produce a bio-potential modulated light wherein said modulating of said light in said electro-optic modulator with said bio-potential is carried out by capacitive coupling with said bio-potential; ande) providing said bio-potential modulated light as a bio-potential modulated light output. 2. The method for measuring electro-physiological activity according to claim 1 further comprising the step of splitting said light into a second light portion with the first light portion being provided to said electro-optic modulator according to step c) of claim 1. 3. The method for measuring electro-physiological activity according to claim 2 further comprising the step of converting said second light portion to electrical form. 4. The method for measuring electro-physiological activity according to claim 1 further comprising the step of splitting said light into a multiple of light portions using an optical power splitter comprising: a) multiple light splitters; andb) a plurality of photodetectors, each photodetector receiving light from separate light splitters of said multiple light splitters and converting it to electrical form. 5. The method for measuring electro-physiological activity according to claim 4 further comprising the step of providing said light from said multiple light splitters to said plurality of photodetectors with a plurality of light fibers. 6. The method for measuring electro-physiological activity according to claim 1 further comprising the step of transforming said bio-potential modulated light output into electronic form. 7. The method for measuring electro-physiological activity according to claim 6 further comprising the step of processing said electronic form to provide a desired output form. 8. The method for measuring electro-physiological activity according to claim 1 further comprising the step of providing a pilot tone that is superimposed on said bio-potential. 9. The method for measuring electro-physiological activity according to claim 1 further comprising the step of phase modulating said light in either unmodulated or bio-modulated form. 10. The method for measuring electro-physiological activity according to claim 1 further comprising at least one of the following steps: a) said light is provided to said electro-optic modulator at an angle to vertical and b) said bio-potential modulated light emerging from said electro-optic modulator is provided at an angle to vertical. 11. The method for measuring electro-physiological activity according to claim 1 further comprising the step of providing said bio-potential to a bio-potential plate that is used to modulate said light in said electro-optic modulator. 12. The method fo... | SRICO, Inc. | A61B005/04
US07761129 | B2 | 2010-07-20 | Measurement head for non-invasive blood analysis | 2005-05-03 | A measurement head has an objective for imaging of a target area such as including a capillary vessel in the skin. The measurement head does not require a lateral shifting of the optical axis of the objective. Transverse relative movements between the objective and a capillary vessel in the skin are performed by mechanically shifting the skin with respect to the objective of the measurement head. Moreover, the measurement head is adapted to host one or more pressure sensors for measuring the contact pressure between the measurement head and the skin. Pressure information may be exploited in order to calibrate a spectroscopic analyzer, and/or to regulate the contact pressure within predefined margins specifying an optimum range of contact pressure for spectroscopic examination of capillary vessels. | 1. A measurement head for a spectroscopic system comprising: fixing means for fixing the measurement head to an area of the surface of the skin of a patient, wherein the fixing means are configured to remain fixed to the area of the surface of the skin;an objective for imaging of a target area located below the surface of the skin, for directing an excitation beam to the target area, and for collecting return radiation from the target area; andlateral moving means for moving the fixing means in a plane substantially perpendicular to an optical axis of the objective. 2. The measurement head according to claim 1, further comprising means for moving the objective relative to the fixing means in a direction substantially parallel to the optical axis of the objective for moving the focal plane of the objective to the target area. 3. The measurement head according to claim 1, wherein the fixing means comprises an adhesive element being adapted to be in contact with the area of the surface of the skin, the adhesive element being further adapted to prevent a relative movement between the area of the surface of the skin and the fixing means. 4. The measurement head according to claim 1, wherein the fixing means comprises a first clamping element and a second clamping element, the first clamping element being an integral part of a housing of the measurement head, the first and the second clamping elements being adapted to exert mechanical stress to the surface of the skin, the first clamping element being configured to move the area of the surface of the skin in a first direction and the second clamping element being configured to move the area of the surface of the skin in a second direction which is opposite the first direction, the first direction and the second direction being perpendicular to the optical axis. 5. The measurement head according to claim 1, wherein the fixing means comprises at least a first and a second magnetic element, the first and the second magnetic element mutually exerting an attractive magnetic force, the first magnetic element being an integral part of a housing of the measurement head and the second magnetic element being adapted to be spaced apart from the first magnetic element by a volume of the tissue of the patient containing the target area. 6. The measurement head according to claim 1, wherein the fixing means comprises means for generating a first barometric pressure between the surface of the skin and the fixing means, the first barometric pressure being substantially smaller than the surrounding barometric pressure. 7. The measurement head according to claim 1, wherein the fixing means comprises at least one disposable element, the at least one disposable element being adapted to be in contact with the area of the surface of the skin. 8. The measurement head according to claim 1, wherein the fixing means is rigidly connected to the objective, the measurement head further comprising at least a first and a seco... | Koninklijke Philips Electronics N.V. | A61B005/145
US07462703 | B2 | 2008-12-09 | Agent for gene transfer | 2003-01-31 | An agent for the transfer of genetic material comprising a genetic material, PEG-containing liposomes, one or more additional polymer particles and a contrasting agent. The invention also is directed toward methods of preparing the agent, and methods of using the agent to affect gene transfer. | 1. A composition suitable for effecting gene transfer, consisting of: (a) one or more genetic materials chosen from the group consisting of DNA and RNA,(b) liposomes consisting of (1) PEG, (2) a steroid, (3) a charged lipid component, (4) genetic material and (5) a carrier liquid,(c) one or more polymer particles selected from the group consisting of nanoparticles, starch particles, gelatin, and Poloxamer,(d) contrasting agent selected from the group consisting of iodine-, gadolinium-, magnetite- and fluorine-containing contrasting agent, and(e) at least one component that assists in packing DNA more tightly. 2. The composition of claim 1, wherein the component that assists in packing DNA more tightly is selected from the group consisting of nuclear capsid protein NCP 7, HMG, polyethylene imine, poly-L-lysine and protamine sulfate. 3. The composition of claim 1, wherein the genetic material encodes one or more proteins selected from the group consisting of herpes simplex-virus thymidine kinase gene (HSVtk), deaminase, NR/CB 1954, purine nucleoside phosphorylase and cytokine genes. 4. The composition of claim 1, wherein the genetic material encodes at least one interleukin (IL) selected from the group consisting of IL-2, IL-4, IL-6, IL-10, L-12 and IL-15. 5. The composition of claim 1, wherein the molar ratio of amphiphile to steroid to charged lipid component is 1:0.3:0.1. 6. The composition of claim 1, wherein the molar ratio of amphiphile to steroid to charged lipid component is 1:1:0.1. 7. The composition of claim 1, wherein the molar ratio of amphiphile to steroid to charged lipid component is or 1:1:0.5. 8. The composition of claim 1, wherein the molar ratio of charged lipid component to contrasting agent is from 2:1 to 10:1. 9. The composition of claim 1, wherein the steroid is chosen from the group consisting of cholesterol, diethoxycholesterol and sitosterol. 10. The composition of claim 1, wherein the charged lipid component is an anion of a lipid selected from the group consisting of diacetyl phosphate, palmitic acid, stearic acid, a phospholipid, and a sphingolipid. 11. The composition of claim 10, wherein the charged lipid component is fluorinated. 12. The composition of claim 1, wherein the polymer particles consist of a 25% aqueous solution of Poloxamer. 13. The composition of claim 1, wherein the liposomes are present in at least one form selected from the group consisting of SUV (small unilamellar vesicles), LUV (large unilamellar vesicles), REV (reverse phase evaporation vesicles), MLV (multilamellar vesicles), and 3β N-(N,N′-dimethylaminoethane)-carbamoyl cholesterol (DAC-chol) liposomes. 14. The composition of claim 1, wherein the polymer particles consist of starch particles having an average diameter of 40 μm to 90 μm. 15. The composition of claim 14, wherein the starch particles have a particle size of 60 μm to 90 μm. 16. The composition of claim 14, wherein the starch particles have a concentration of from 5 ... | Max-Delbruck-Centrum fur Molekulare Medizin | C07H021/04 A61K031/70 A61B005/055
US07634312 | B2 | 2009-12-15 | Visceral fat determining device | 2008-08-22 | A visceral fat determining device is disclosed which can be used easily at home for example and can obtain information on visceral fat. The visceral fat determining device comprises input means for inputting personal data that represents vital characteristics of the patient; a data processing unit capable of performing various calculations; and a display device capable of displaying the personal data and results of the calculations. Upon inputting of a waist size which is a girth of the trunk of the patient, and a hip size which is a girth of the buttocks of the patient, quantitative information on abdominal visceral fat of the patient is given through a calculation based on WHR which is a ratio of the waist size to the hip size (waist size/hip size). | 1. A visceral fat determining device comprising: input means for inputting personal data including an abdominal girth WL which is a circumferential length of a trunk of a patient, a height value TL of the patient, and a gluteal girth HL which is a circumferential length of buttocks of the patient;a data processing unit for storing the personal data and for calculating an estimated value of an abdominal visceral fat cross sectional area VA of the patient;a display device for displaying the personal data and the result of the abdominal visceral fat cross sectional area VA calculation performed by the data processing unit; andimpedance measuring means for measuring a bioelectrical impedance value Z of the patient via electrodes contacted to end portions of the patient;wherein the data processing unit is programmed for storing a first regression coefficient of WHR, and a second regression coefficient of TL2/Z and a first regression constant, WHR being a ratio (WL/HL) of an abdominal girth WL to a gluteal girth HL, TL2/Z being obtained by dividing a squared height of a human TL2 by the human's bioclectrical impedance value Z, the regression coefficients and the regression constant being obtained from statistical processing of correlation among actual measurement values of the abdominal visceral fat cross sectional area VA measured via abdominal tomography of random samples of human bodies, and respective WHR values and TL2/Z values of the human samples;wherein the data processing unit is further programmed for calculating the estimated value of abdominal visceral fat cross sectional area VA of the patient based on the sum of the first regression coefficient multiplied by the WHR value of the patient, the second regression coefficient multiplied by the TL2/Z value of the patient, and the first regression constant. 2. A visceral fat determining method comprising: at a data processing unit, receiving from an input means input of personal data including an abdominal girth WL which is a circumferential length of a trunk of a patient, a height value TL of the patient, and a gluteal girth HL which is a circumferential length of buttocks of the patient;storing the personal data at the data processing unit;measuring, via an impedance measuring means, a bioclectrical impedance value Z of the patient via electrodes contacted to end portions of the patient;at the data processing unit, storing a first regression coefficient of WHR, and a second regression coefficient of TL2/Z and a first regression constant, WHR being a ratio (WL/HL) of an abdominal girth WL to a gluteal girth HL, TL2/Z being obtained by dividing a squared height of a human TL2 by the human's bioclectrical impedance value Z, the regression coefficients and the regression constant being obtained from statistical processing of correlation among actual measurement values of an abdominal visceral fat cross sectional area VA measured via abdominal tomography of random samples of human bodies, and resp... | Yamato Scale Co., Ltd. | A61B005/05
US08024026 | B2 | 2011-09-20 | Dynamic reference method and system for use with surgical procedures | 2007-05-31 | A substrate configured for placement on an internal organ or tissue is provided. In certain embodiments, the substrate conforms to and moves with the internal organ or tissue. Three or more sensor elements are integrated on the substrate. In one implementation, the substrate and associated sensor elements provide dynamic referencing of the internal organ or tissue after registration of the sensor data with images and/or volumetric representations of the internal organ or tissue. | 1. A method for tracking dynamic motion of an organ comprising the acts of: powering three or more sensor components integrated on a substrate placed on an internal organ using a power supply electrically connected to the three or more sensor components;generating a first set of position data for the three or more sensor components, wherein the first set of position data is based on signals or fields generated or detected by the sensor components;generating a second set of position data for the three or more sensor components based on an identification of the three or more sensor components in one or more radiological images or volumetric representations of the internal organ;identifying corresponding sensor components in the first set of position data and the second set of position data; andregistering the first set of position data with the one or more radiological images or volumetric representations based on the identification of corresponding sensor components in the first set of position data and the second set of position data. 2. The method of claim 1 further comprising generating the one or more images or volumetric representations based on image data acquired using at least one of computed tomography, magnetic resonance imaging, ultrasound, or three-dimensional fluoroscopy. 3. The method of claim 1 wherein at least one of generating the first set of position data, generating the second set of position data, identifying corresponding sensor components, or registering the first set of position data with the one or more images or volumetric representations is performed automatically. 4. The method of claim 1 further comprising tracking a surgical instrument using the registered first set of position data and the one or more images or volumetric representations. 5. A method for tracking dynamic motion of an organ comprising the acts of: powering one or more sensor components provided on a substrate placed on an internal organ using a power supply electrically connected to the one or more sensor components;generating position and orientation data for the one or more sensor components, wherein the position and orientation data is based on signals or fields generated by the sensor components;generating a shape model of the substrate based on the position and orientation data;segmenting a region of interest in one or more radiological images or volumetric representations, wherein the region of interest comprises at least one of the internal organ, a portion of the internal organ, or a region proximate or connected to the internal organ;generating a shape model of the region of interest based on the segmentation; andregistering the shape model of the substrate and the shape model of the region of interest. 6. The method of claim 5 further comprising generating the one or more images or volumetric representations based on image data acquired using at least one of computed tomography, magnetic resonance imaging, ultrasound, or three-dimension... | General Electric Company | A61B005/05
US08311609 | B2 | 2012-11-13 | MRI thermometry involving phase mapping and reference medium used as phase reference | 2005-07-29 | A computing unit is arranged to access phase images computed from the image data and representative of the target medium and the reference medium and carry out computing steps to yield temperature difference map. The operation of the computing is controlled by a computer program carrying computing steps programmed as instructions to a processor. The apparatus (1) includes a storage unit (8) arranged to store suitable image data which are accessed by the computer program during the computations. The storage unit (8) stores system parameters of the magnetic resonance apparatus, like the strength of the main field (B0), gyromagnetic ratio (γ), and suitable parameters of a target medium, like screen constant (α), in a file (3). The apparatus (1) further includes a working memory (6), typically based on RAM. An output (9) of the apparatus includes the computed temperature difference map, which can be used to control a suitable medical diagnostic or therapeutic device, or it may be visualized to a user or be stored for archiving purposes. For feed-back purposes, the apparatus includes a display (7) to visualize the computed temperature difference map. Preferably, the temperature difference map is visualized within a suitable interactive graphic user interface (7a). | 1. An apparatus for generating a time dependent phase error corrected temperature difference map comprising: a storage unit configured to store magnetic resonance image data indicative of magnetic resonance signals emitted from an animate object subjected to a magnetic resonance imaging procedure to measure a temporal local temperature variation within the animate object between a first time moment and a second time moment, which magnetic resonance imaging procedure is subject to time dependent phase errors, said animate object comprising a target water-based tissue medium and a reference fat tissue medium;a computing unit configured to:i. access a first phase image generated from the image data and representative of the target water-based tissue medium for the first time moment;ii. access a second phase image generated from the image data and representative of the reference fat tissue medium for the first time moment;iii. access a third phase image generated from the image data and representative of the target water-based tissue medium for the second time moment;iv. access a fourth phase image generated from the image data and representative of the reference fat tissue medium for the second time moment;v. compute a first phase map by subtracting the first phase image from the third phase image for the target water-based tissue medium;vi. compute a second phase map indicative of the time dependent phase error by subtracting the second phase image from the fourth phase image for the reference fat tissue medium;vii. correct the first phase map for the time dependent phase errors by subtracting the second phase map from the first phase map,viii. generate a time dependent phase error corrected temperature difference map based on the corrected first phase map. 2. The apparatus according to claim 1, wherein the apparatus further comprises: an input arranged to forward the image data to the computing unit, the computing unit being further arranged to compute the first phase image, the second phase image, the third phase image, the fourth phase image. 3. The apparatus according to claim 1, wherein the computing unit is further arranged to calculate respective binary masks for the first phase image and for the second phase image. 4. The apparatus according to claim 1, wherein the computing unit is further arranged to fit the second phase image with a smoothing function. 5. The apparatus according to claim 1, further comprising a display for visualization of the temperature difference map. 6. The apparatus according to claim 1, wherein the reference fat tissue medium includes a medium which in the magnetic resonance procedure generates image data with a phase which changes with main magnetic field drift and changes in bulk susceptibility and which remains constant during temperature changes. 7. A magnetic resonance imaging system for computing a time dependent phase error corrected temperature difference map, the system comprising: a magnetic resonan... | Koninklijke Philips Electronics N.V. | A61B005/05
US08147431 | B2 | 2012-04-03 | Measuring member and device for determining the blood flow of the gastrointestinal tract, as well as for registrating the intestinal peristalsis | 2006-02-08 | A measuring member and a device to be used for determining the blood flow of the gastrointestinal tract, includes an intestinal probe, having its distal end provided with a flexible inflatable body and at least one pressure sensor mounted on this supporting body. After bringing the inflatable flexible body to predetermined pressure, the pressure sensor will abut the intestinal wall and the measured pressure development, which represents a measure for the blood flow of the gastrointestinal tract, is transformed into a graphical image with the help of the device and/or is transformed into actual values for the blood flow of the gastrointestinal tract. Further, the measuring member measures the development of the pressure within the inflatable flexible body, which represents a measure for the intestinal peristalsis. | 1. A system for measuring the blood flow of a gastrointestinal tract, comprising (a) an elongated probe having a distal region, the probe configured for placement within the intestine;(b) an inflatable body disposed on the distal region;(c) a sensor operatively associated with the inflatable body, the sensor configured to generate a plethysmographic signal corresponding to time-wise changes in blood flow volume within the intestine; and(d) a device configured to control inflation of the inflatable body to a condition wherein the inflatable body and the sensor abut against a wall surface of the intestine without affecting blood flow within the intestine, the device configured to process the plethysmographic signal to generate a signal corresponding to at least one of the blood pressure and intestinal peristalsis at the location of the sensor within the intestine. 2. The system of claim 1, wherein the sensor is mounted on an outer circumference of the inflatable body. 3. The system of claim 1, wherein the device is configured to filter the plethysmographic signal to derive a signal corresponding to intestinal peristalsis. 4. The system of claim 1, wherein the device is configured to control inflation of the inflatable body to a predetermined pressure. 5. The system of claim 1, wherein the device includes a graphical display, and the plethysmographic signal is displayed on graphical display. 6. The system of claim 3, wherein the device is configured to compare the signal corresponding to intestinal peristalsis threshold value. 7. The system of claim 1, wherein said inflatable body has an elliptical cross-sectional configuration. 8. The system of claim 1, wherein the elongated probe is configured for positioning within the intestine through an oral cavity. 9. The system of claim 1 wherein the elongated probe is configured for positioning within the intestine through the abdominal wall. 10. The system of claim 1 wherein the elongated probe is configured to function as a nutritional probe. 11. A system for monitoring blood flow of a gastrointestinal tract, comprising an elongated probe having a proximal end and a distal region, the probe configured for placement within the intestine;an inflatable body disposed on the distal region, the inflatable body communicating with the proximal end of the probe;an electronic sensor operatively associated with the inflatable body, the electronic sensor configured to generate a plethysmographic signal corresponding to time-wise changes in blood flow volume within the intestine; anda monitor operatively coupled to the inflatable body and the electronic sensor, the monitor configured to control inflation of the inflatable body to a condition so that the electronic sensor abuts against a wall surface of the intestine without diminishing blood flow therein, the monitor configured to process the plethysmographic signal to generate a signal corresponding to at least one of the blood pressure and intestinal peris... | Q-Pidt BV | A61B005/103 A61B005/117 A61B005/02 A61B005/00
US08308662 | B2 | 2012-11-13 | Measuring device | 2007-12-07 | A measuring device for measuring a depth of a tear of a tissue is disclosed. The measuring device comprises a member having a shaft and a distal tip, a plurality of indicia on the shaft of the member for measuring the depth of the tear, and a cannulated member into which the member is disposed. In an example, the measuring device is used to measure the depth of a tear of the rotator cuff tendon. A method of measuring a tear of a tissue using the measuring device is also disclosed. | 1. A measuring device for measuring a depth of a tear of a tissue and a total thickness of the tissue, the measuring device comprising: a member having a shaft and a distal tip, there being a plurality of indicia on the shaft of the member; anda cannulated member into which the member is slidably disposed;wherein the plurality of indicia comprises:a first plurality of indicia disposed on a distal portion of the shaft that is readable to measure the depth of the tear when the distal tip of the member is positioned at the distal margin of the tear and the measuring device is positioned substantially tangential to the tear; anda second plurality of indicia disposed on a proximal portion of the shaft that is readable to measure a total thickness of the tissue when the cannulated member is moved so that a distal end thereof is substantially at a proximal margin of the tissue while maintaining the position of the distal tip at the distal margin. 2. The measuring device as in claim 1 wherein the device is capable of being used to measure the depth arthroscopically. 3. The measuring device as in claim 1 wherein the distal tip is configured to penetrate through the tear. 4. The measuring device as in claim 1 wherein the cannulated member further comprises a window positioned on the cannulated member to view at least a portion of the second plurality of indicia. 5. The measuring device as in claim 1 further comprising a head disposed on the shaft at an end of the shaft opposite the distal tip. 6. The measuring device as in claim 1 wherein the cannulated member further comprises a grip. 7. The measuring device as in claim 1 further comprising a seal positioned on the shaft that stabilizes the shaft within the cannulated member. 8. The measuring device as in claim 1 wherein the tissue is a rotator cuff tendon. 9. The measuring device as in claim 1 wherein the device is made from at least one of stainless steel, titanium, PEEK, or a hard plastic. 10. A method of measuring a tear of a tissue and a total thickness of the tissue using a measuring device, the measuring device including a cannulated member having distal and proximal ends, a member slidably disposed within the cannulated member and having a distal tip and a shaft with distal and proximal portions, there being (i) a first plurality of indicia disposed on a distal portion of the shaft that is readable to measure the depth of the tear when the distal tip of the member is positioned at the distal margin of the tear and the measuring device is positioned substantially tangential to the tear, and (ii) a second plurality of indicia disposed on a proximal portion of the shaft that is readable to measure a total thickness of the tissue when the cannulated member is moved so that a distal end thereof is substantially at a proximal margin of the tissue while maintaining the position of the distal tip at the distal margin, the method of measuring comprising the steps of: a. positioning the device subs... | Arthrex, Inc. | A61B005/103 A61B005/117 A61B005/04
US08086303 | B2 | 2011-12-27 | Cardiac arrhythmias analysis of electrophysiological signals based on symbolic dynamics | 2007-12-14 | The disclosed method analyzes cardiac electrophysiological signals, including ECG and internal cardiac electrograms, based on multi-level symbolic complexity calculation and multi-dimensional mapping. The results may be used to objectively identify cardiac disorders, differentiate cardiac arrhythmias, characterize pathological severities, and predict life-threatening events. Multi-level symbolization and calculation of the electrophysiological signal is used provide better reliability and analysis resolution for identifying and characterizing cardiac disorders. Adaptive analysis of the cardiac signal complexity enables calculation efficiency and reliability with high SNR, and with low calculation volume and power consumption. One dimension (time or frequency domain) and multi-dimension symbolic analysis is used to provide more information of cardiac pathology and high risk rhythm transition to doctors. | 1. A method for predicting cardiac arrhythmia, comprising acquiring patient data elements comprising a patient cardiac signal, the data elements comprising a data stream;establishing a window within the data stream;performing symbolization of data in the data stream window into multi-level symbols individually represented by a plurality of bits;performing a symbolic complexity calculation on the symbolized data;comparing information obtained from the symbolic complexity calculation to a plurality of predetermined thresholds associated with different levels of complexity and corresponding different levels of risk or severity of a medical condition; andproviding information to a user if the information obtained from the symbolic complexity calculation exceeds the predetermined thresholds. 2. The method of claim 1, including adaptively dynamically varying said thresholds in response to signal noise. 3. The method of claim 2, wherein the step of performing symbolic complexity calculation on the symbolized data comprises a multi-level complexity calculation. 4. The method of claim 1, including adaptively dynamically varying said window in response to signal noise. 5. The method of claim 4, wherein the step of providing information to a user comprises providing an audible or visual alarm to the user if the information obtained from the symbolic complexity calculation exceeds a predetermined threshold. 6. The method for predicting cardiac arrhythmia, comprising monitoring patient data elements comprising patient cardiac signals, the data elements comprising a data stream;performing symbolization of the data stream comprising a multi-level symbolization;performing a symbolic complexity calculation comprising a multi-level complexity calculation on the symbolized data;comparing information obtained from the symbolic complexity calculation to a predetermined threshold; andproviding information to a user if the information obtained from the symbolic complexity calculation exceeds the predetermined threshold wherein the step of providing information to a user further comprises:providing a graphical display of frequency domain complexity vs. time domain complexity, the graphical display comprising a predetermined visible normal region associated with normal cardiac signals and a predetermined visible severe region associated with cardiac signals indicative of arrhythmia; andmapping the complexity calculation on the graphical display to provide a visual indication to a user that the cardiac signals are: (a) within the normal region, (b) are approaching the severe region, or (c) are within the severe region. 7. The method of claim 6, wherein the complexity calculation is mapped on the graphical display over time to provide a visible trending of the cardiac signals, the method further comprising providing a risk threshold that identifies to a user when the cardiac signals are trending to the severe region. 8. A machine readable storage device tangibly emb... | Siemens Medical Solutions USA, Inc. | A61B005/0428 A61B005/0452
US08301226 | B2 | 2012-10-30 | Method and apparatus for performing a navigated procedure | 2008-04-28 | A system can be used to navigate or guide an instrument or device into an anatomy of a patient. The navigation can occur with the use of image data acquired of the patient. The image data can be registered to the patient space for navigation. Also, one or more coils can be used for tracking or localization. | 1. A method of tracking an instrument with a navigation system, comprising: generating a first navigation field relative to a patient;navigating an alignment device with the first navigation field before fixing the alignment device to the patient;fixing the alignment device to the patient;generating a second navigation field independent of the first navigation field relative to the patient; andnavigating the instrument relative to the patient only with the second navigation field after ceasing generating the first navigation field. 2. The method of claim 1, further comprising: providing the alignment device including a base; andproviding a first coil integrated into the base;wherein fixing the alignment device to the patient includes fixing the base to the patient;wherein navigating the alignment device includes sensing the first navigation field with the first coil integrated into the base. 3. The method of claim 2, wherein generating the second navigation field includes forming the second navigation field with the first coil. 4. The method of claim 3, further comprising: ceasing generation of the first navigation field prior to generating the second navigation field. 5. The method of claim 4, further comprising: providing the instrument including a guide tube;providing a second coil on the guide tube;navigating the guide tube at least by sensing the second navigation field; andproviding the instrument further including a second moveable portion with a third coil;wherein navigating the instrument includes tracking the location of the guide tube, the second moveable portion, the third coil or combinations thereof in the second navigation field. 6. The method of claim 2, further comprising: forming a burr hole in the patient; andfixing the base at the burr hole. 7. The method of claim 1, further comprising: executing instructions with a processor to cease generating the first navigation field, generating the second navigation field, and tracking the location of the instrument; anddisplaying a graphical representation of the instrument being tracked on a display device. 8. The method of claim 7, further comprising: acquiring image data of the patient; andsuperimposing the graphical representation of the instrument on the image data of the patient. 9. The method of claim 1, further comprising: acquiring image data of the patient;determining a location for positioning the alignment device relative to the patient;tracking the alignment device with the first navigational field to navigate the alignment device to the location relative to the patient; anddisplaying on a display a graphical representation of the alignment device to illustrate a tracked position of the alignment device relative to the location. 10. The method of claim 9, further comprising: inputting an instruction to a processor to determine timing for forming the first navigation field, the second navigation field, or combinations thereof. 11. The method of claim 1, further com... | Medtronic, Inc. | A61B005/05
US08454521 | B2 | 2013-06-04 | Sphygmomanometer | 2005-09-12 | The objective of the present invention is to provide a sphygmomanometer that is easy to use. The sphygmomanometer according to the present invention measures blood pressure in accordance with an oscillation in an artery wall, resulting from an arterial pulse correspondent with a change in cuff pressure. It comprises a cuff that is connected to the sphygmomanometer main body by a tube, a display unit for displaying the results of blood pressure measurements, and an air supply unit for supplying air to, and thus pressurizing, the cuff, which is detachable from the sphygmomanometer main body. The air supply unit is screwed into the sphygmomanometer main body with a screw assembly, and the screwed-in state is preserved by a caulking ring. The air supply unit also comprises a filter for keeping dust from entering the sphygmomanometer main body. | 1. A sphygmomanometer, that measures blood pressure in accordance with an oscillation in an artery wall, resulting from an arterial pulse correspondent with a change in cuff pressure, comprising: a cuff that is connected to a sphygmomanometer main body by a tube;an air supply unit for supplying air to and pressurizing the cuff, the air supply unit being detachable from the sphygmomanometer main body, and forming a single unit when attached to the sphygmomanometer main body;a mode switch for selecting a measurement mode from among a Normal mode that measures blood pressure automatically, a Slow mode that measures blood pressure automatically, and a Stethoscope mode for measuring blood pressure manually via a stethoscope;a control unit for controlling opening and closing of an electromagnetic valve of the sphygmomanometer to adjust a speed at which air is exhausted from the cuff according to the measurement mode selected via the mode switch so that the control unit controls the electromagnetic valve to: i) keep the exhaust speed at a first preset exhaust speed during blood pressure measurement when the Normal mode is selected by the mode switch; ii) keep the exhaust speed at a second preset exhaust speed during blood pressure measurement when the Slow mode is selected by the mode switch, the second preset exhaust speed being approximately half of the first preset exhaust speed; and iii) keep the exhaust speed at a third preset exhaust speed during blood pressure measurement when the Stethoscope mode is selected by the mode switch, the third preset exhaust speed being approximately the same as the second preset exhaust speed;a display unit which displays a blood pressure measurement result and the selected measurement mode,a pressure sensor measuring a pressure of the cuff, anda memory storing pressure values measured by the pressure sensor during depressurization of the cuff,whereinthe control unit determines a systolic blood pressure based on the stored pressure values, identifies a first pressure value from amongst the stored pressure values at which the amplitude dramatically increases for a first time and also identifies a second pressure value from amongst the stored pressure values at which the amplitude dramatically decreases for a first time as seen from a point of maximum amplitude of the pressure,the control unit obtains the first pressure value at which the amplitude dramatically increases for a first time as a first systolic blood pressure candidate value,the control unit obtains the second pressure value at which the amplitude dramatically decreases for a first time as seen from a point of maximum amplitude of the pressure as a second systolic blood pressure candidate value,the control unit determines whether a difference between the first systolic blood pressure candidate value and the second systolic blood pressure candidate value is within a prescribed value, andthe control unit determines a blood pressure corresponding to the first... | Terumo Kabushiki Kaisha | A61B005/02
US08406840 | B2 | 2013-03-26 | Spatially resolved oxymetry | 2006-09-07 | In optical tomography, a calibration of the data may be necessary for image reconstruction. According to an exemplary embodiment of the present invention, the object of interest is used for calibration, wherein the image data is acquired during a highly oxygenated phase of the object of interest and wherein the calibration data is acquired during a low oxygenated phase of the object of interest. This may provide for an improved calibration, resulting in improved image quality. | 1. An optical examination apparatus for optical examination of an object of interest, the optical examination apparatus comprising: a detector unit adapted for detecting first radiation and second radiation from the object of interest in response to light emitted from an optical radiation source propagating through the object of interest, the first radiation and the second radiation providing first detection data and second detection data, respectively; anda determination unit adapted for determining spatially resolved information relating to the object of interest based on the first detection data and the second detection data,wherein the first detection data corresponds to a first oxygen concentration inside the object of interest,wherein the second detection data corresponds to a second oxygen concentration inside the object of interest, which is different from the first oxygen concentration, andwherein the spatially resolved information relating to the object of interest is determined using a ratio of the first detection data and the second detection data for calibration. 2. The optical examination apparatus of claim 1, wherein the optical radiation source comprises a coupler adapted for coupling light emitted by a plurality of lasers or photodiodes, and emitting the coupled light to the object of interest. 3. The optical examination apparatus of claim 1, wherein the determination unit is adapted for reconstructing a two-dimensional or three-dimensional image of the object of interest based on the first detection data and the second detection data. 4. The optical examination apparatus of claim 1, wherein the first detection data corresponds to a systolic phase of the object of interest; and wherein the second detection data corresponds to a diastolic phase of the object of interest. 5. The optical examination apparatus of claim 4, further comprising: an electrocardiogram unit adapted for acquiring electrocardiogram data of the object of interest during the detection of the first radiation and the second radiation by the detector unit;wherein the determination unit is adapted for identifying the systolic phase and therefore the first detection data and for identifying the diastolic phase and therefore the second detection data based on electrocardiogram data. 6. The optical examination apparatus of claim 1, wherein the determination unit is adapted for reconstructing blood dynamics of the object of interest on the basis of a temporal behaviour of the first radiation and the second radiation. 7. The optical examination apparatus of claim 1, wherein the detector unit is adapted for detecting, as the first detection data, first and second transmission data and, as the second detection data, first and second reflection data. 8. The optical examination apparatus of claim 1, adapted as an optical mammography apparatus. 9. The optical examination apparatus of claim 1, wherein the detector unit comprises a spatially resolving detector. 10. Th... | Koninklijke Philips Electronics N.V. | A61B005/00
US09597033 | B2 | 2017-03-21 | Biological-information obtaining apparatus and method thereof | 2008-04-28 | A biological-information obtaining apparatus includes a light-emitting unit configured to emit light, an image-capturing unit configured to capture images, in time sequence, obtained by irradiating a living body with the light emitted and by causing the light to be transmitted through or reflected by the living body, the image-capturing unit being sensitive to at least two color components, an extreme generation unit configure to generate, for each of the captured images in the time sequence, a maximum value and a minimum value of each of the color components in a certain area of the captured image, and an oxygen-saturation calculation unit configured to calculate oxygen saturation on the basis of the maximum value and the minimum value of each of the color components in the certain area of the captured image. | 1. A biological-information obtaining apparatus comprising: light-emitting means for emitting light;image-capturing means for capturing images, in time sequence, obtained by irradiating a living body with the light emitted and by causing the light to be transmitted through or reflected by the living body, the image-capturing means being sensitive to at least two color components;extreme generation means for generating a maximum value and a minimum value, in time sequence, of each of the color components for certain regions of the captured images; andoxygen-saturation calculation means for calculating oxygen saturation on the basis of the maximum value and the minimum value of each of the color components. 2. The biological-information obtaining apparatus according to claim 1, wherein the light-emitting means emits white light. 3. The biological-information obtaining apparatus according to claim 1, wherein the color components include red and blue. 4. The biological-information obtaining apparatus according to claim 1, wherein the extreme generation means generates, for each of the color components, the maximum value and the minimum value in the time sequence from averages of the color component for the entirety of the captured images. 5. The biological-information obtaining apparatus according to claim 1, wherein the extreme generation means generates, for each of the color components, the maximum value and the minimum value in the time sequence from averages of the color component for predetermined regions of the captured images. 6. A biological-information obtaining apparatus comprising: light-emitting means for emitting white light;image-capturing means for capturing images, in time sequence, obtained by irradiating a living body with the light emitted and by causing the light to be transmitted through or reflected by the living body, the image-capturing means being sensitive to at least red and blue color components;extreme generation means for generating a maximum value and a minimum value, in time sequence, of each of the color components for certain regions of the captured images; andoxygen-saturation calculation means for calculating, on the basis of the maximum value and the minimum value of each of the color components, oxygen saturation by solving an equation of {(Rc−Bc)−(Re−Be)}/{Bc−Be}=[(Re−Be)×{S×Eo(λ1)+(1−S)×Er(λ1)}]/[Be×{S×Eo(λ2)+(1−S)×Er(λ2)}], where Rc represents the maximum value of the red component, Bc represents the maximum value of the blue component, Re represents the minimum value of the red component, Be represents the minimum value of the blue component, S represents oxygen saturation, Eo(λ) represents a known absorbance coefficient of oxyhemoglobin at a wavelength λ, Er(λ) represents a known absorbance coefficient of deoxyhemoglobin at a wavelength λ, and λ1 and λ2 represent specific values of wavelength λ. 7. A method of obtaining biological information, the method being performed by a biological-information obt... | Sony Corporation | A61B005/00 A61B005/1455 G06F021/32 G06T007/00
US10478115 | B2 | 2019-11-19 | Handheld home monitoring sensors network device | 2005-10-04 | Light weight personal handheld home monitoring and managing device, which includes a Sound Sensor network/array of Sound Sensor networks combined with an Artificial Neural Network (ANN) and a build in system and methods, making this device an intelligent and portable apparatus to address specific health issues. The combined apparatus is used for managing and/or guidance and/or diagnosing and/or controlling and managing purposes. This present version of the apparatus will address pulmonary disorders and diseases or similar ailments. | 1. A device for obtaining a parameter for nitric oxide (NO) in exhaled air, said device comprising: a housing defining a conduit having an inlet and an outlet, said conduit comprising a Gaussian surface for collecting condensate from the exhaled air, the Gaussian surface formed such that condensate is collected on the Gaussian surface, the conduit further comprising a thermo-electric device arranged for cooling the conduit to generate the condensate from the exhaled air, the conduit further comprising a sensor, said sensor being disposed on the Gaussian surface, said sensor for determining said parameter in the condensate on the Gaussian surface and for generating an electrical signal representing said parameter based on airflow generated by a person blowing in said conduit at said inlet such that the airflow contacts the Gaussian surface and forms the condensate, said sensor having an electrical output for outputting said electrical signal,an analog-to-digital converter having a converter input and a converter output, said converter input being electrically connected to said electrical output of said sensor such that the electrical signal is converted to a digital signal by the analog-to-digital converter, anda processor having a first digital input and a second digital input and a first digital output, said first digital input of said processor being electrically connected to said converter output for receiving said digital signal,wherein said sensor is a NO sensor capable of measuring NO in condensate of exhaled air. 2. The device according to claim 1, further comprising a memory unit having a digital input and a digital output, said digital output of said memory unit connected to said second digital input of said processor for transmitting previously stored reference values identified in a hierarchy. 3. The device according to claim 2, wherein said processor establishing on a basis of said reference values and on a basis of values derived from said digital signal received from said analog-to-digital converter, a health state of said person and presenting information relating to said health state to said person and storing said health state in said hierarchy. 4. The device according to claim 1, wherein said device further comprises a transceiver for transmitting and receiving data representing physiological information determined by external units. 5. The device according to claim 1, further comprising at least one sound sensor. 6. The device according to claim 1, wherein said at least one sound sensor measures ultra sound. 7. The device according to claim 1, further comprising at least one of a temperature sensor, an air pressure sensor, and a moisture sensor. 8. The device according to claim 1, further comprising a sensor for measuring a second parameter selected from the group consisting of: pH, H2O2, CO, O2, and CO2 in condensate of exhaled air. 9. A method for obtaining a physiological parameter regarding an individual using a dev... | SpiroFriend Technology ApS | A61B005/097 A61B005/00 A61B005/087 A61B007/00 A61B005/0205
US20080266118 | A1 | 2008-10-30 | Personal emergency condition detection and safety systems and methods | 2008-03-07 | A drowning or asphyxiation prevention system and methods are set forth, for facilitating drowning prevention of swimmers in bodies of water, such as pools, lakes or the like. The drowning prevention safety system comprises a wearable article worn by a swimmer, an alarm indicator for transmitting an alarm condition. The system may further include an alarm receiving system for receiving the alarm signal from the alarm transmitting device. The invention also relates to a heart rate monitoring system comprising a patch type portion to be adhesively applied to the skin of a user to monitor the electrical activity of the heart and generate heart rate information that is communicated to a separate wearable device, such as a wrist worn device. | 1. A drowning prevention system comprising, a wearable article having at least one sensing system for at least one physiological function of a persons body and/or environmental information or condition around or related to the persons body, and a processing system to identify predetermined variances in the monitored at least one physiological function and/or environmental information or condition related to an emergency condition and possible drowning, and an alarm indicator for issuing an alarm if an emergency condition is detected. 2. A drowning prevention system according to claim 1, further comprising, an alarm receiving system associated with the body of water in which the person is swimming. 3. A drowning prevention system according to claim 2, further comprising, the alarm receiving system having an alarm indicating system for issuing a further alarm upon receipt of an alarm from a wearable article. 4. A drowning prevention system according to claim 1, wherein, the at least one sensing system for at least one physiological function of a persons body and/or environmental information or condition are selected from one or more of the group consisting of pulse or heart rate, peripheral vasoconstriction, blood pressure, blood oxygen, CO2 blood level, body or environment temperature, body sounds or combinations thereof. 5. A drowning prevention system according to claim 1, wherein, the at least one alarm indicator is selected from one or more of the group consisting of audible alarm, visual alarm, ultrasonic alarm, wireless signal alarm or combinations thereof. 6. A drowning prevention system according to claim 2, wherein, the alarm receiving system being selected from one or more of the group consisting of an array of hydrophones, an array of microphones, an array of ultrasonic receivers, a visual alarm receiving system, a wireless signal receiving system or combinations thereof. 7. A drowning prevention system according to claim 2, wherein, the alarm receiving system comprises a plurality of alarm receivers disposed in and/or around a body of water. 8. A drowning prevention system according to claim 1, further comprising, a position detection system associated with the wearable article. 9. A drowning prevention system according to claim 1, further comprising, a sound monitoring system associated with the wearable article. 10. A drowning prevention system according to claim 1, further comprising, a user actuable panic switch associated with the wearable article. 11. A drowning prevention system according to claim 1, further comprising, a user actuable reset switch associated with the wearable article. 12. A drowning prevention system according to claim 1, further comprising, a user actuable panic switch associated with the wearable article. 13. An asphyxiation prevention system comprising, a wearable article having at least one sensing system for at least one physiological function of a persons body and/or environmental information ... | - | G08B023/00 A61B005/08
US20080319294 | A1 | 2008-12-25 | HEALTH MANAGEMENT DEVICES AND METHODS | 2008-06-20 | Methods, devices and systems to detect analyte level in a patient with gestational diabetes and/or provide related therapy management are provided. | 1. A method, comprising: monitoring an analyte level of a subject for gestational diabetic related condition over a predetermined time period;storing a plurality of data associated with the monitored analyte level over the predetermined time period, the plurality of data having one or more parameters associated with the monitored analyte level; andprocessing the stored plurality of data to determine, at least in part, one or more therapy regimen associated with the treatment of gestational diabetes. 2. The method of claim 1 generating one or more output based on the processed stored plurality of data or the therapy regimen. 3. The method of claim 2 wherein the one or more output includes one or more visual output or audible output. 4. The method of claim 2 including providing the generated output to the subject. 5. The method of claim 4 wherein providing the generated output includes displaying the generated output to the subject. 6. The method of claim 1 wherein the one or more parameters includes one or more of the monitored analyte levels, fetal heart rate data, uterine contraction information, diet information, physical activity information, prenatal care information, and medication information. 7. The method of claim 1 wherein the determined one or more therapy regimen includes a recommendation for modification to the diet of the subject, modification to the physical activity of the subject, or modification to the medication dosage information of the subject. 8. The method of claim 1 wherein the determined therapy regimen includes modification to one or more of a modification to a basal rate profile for insulin delivery to the subject. 9. The method of claim 1 wherein the analyte is glucose. 10. A method, comprising: collecting analyte level information over a predetermined time period when one or more condition associated with gestational diabetes is detected; andexecuting one or more computer program to process the collected analyte level information, wherein executing the one or more computer program includes: selecting a predetermined function associated with the detected gestational diabetes;retrieving one or more parameters associated with the collected analyte level information or the monitored gestational diabetes condition;performing data analysis based on the retrieved one or more parameters and the collected analyte level information to generate one or more therapy management information associated with the monitored one or more condition associated with gestational diabetes. 11. The method of claim 10 wherein the one or more computer program is executed on a healthcare provider computer terminal, or a patient computer terminal or a remote terminal. 12. The method of claim 10 including transmitting the collected analyte level information to a remote location. 13. The method of claim 12 wherein the analyte level information is received over the internet. 14. The method of claim 13 wherein the analyte level information ... | Abbott Diabetes Care, Inc. | A61B005/145
US20080287758 | A1 | 2008-11-20 | Illuminator probe with memory device and method | 2008-07-25 | An improved medical illuminator (103) with memory chip (241). The memory element retains information useful in the operation of the illuminator, such as probe identification, probe serial number, use history, calibration details, or other information. Further, the illuminator with information chip can be incorporated into a medical probe or device, such as a pulse oximeter system or probe, as well as to be used in surgical devices or other equipment that require illumination. A method of use is also provided. | 1-23. (canceled) 24. An illuminator medical device, comprising: a probe having a probe body;an illuminator for emitting light; and,an information chip integrated into the probe body and configured to retain information useful in the operation of the device. 25. The device of claim 24, wherein the useful information comprises any one or more of calibration parameters, a reference database, a library of characteristic discriminant features from previously identified tissues, probe identification, probe serial number, use history, calibration details, or other information useful in the operation of the illuminator device. 26. The device of claim 24, wherein the medical device is selected from the group of medical devices consisting of probes for medical endoscopic use, targetable injection needles, catheters, catheters with extendable needles, nibblers, devices with jaws, scissors, probes that measure colon oxygenation, probes that measure by pulse oximetry, probes that measure arterial oxygenation, probes that measure oxygen delivery to the body's core organs, probes that measure in the gastrointestinal system including the oropharynx, nasopharynx, esophagus, stomach, duodenum, ileum, colon, or other gastrointestinal tissues. 27. The device of claim 24, wherein the medical device is an oximeter probe configured for use in performing oxygenation measurements. 28. The device of claim 24, wherein the medical device is an oximeter probe configured for use in performing pulse oximetry. 29. The device of claim 24, wherein the medical device configured for use such inside, or in contact with, living tissue. 30. The device of claim 24, further incorporated into a spectrophotometer system. 31. The device of claim 24, wherein light from the illuminator is pulsed. 32. The device of claim 24, further incorporated into a system in which light from the illuminator is analyzed as time-resolved, frequency-resolved, or spatially-resolved. 33. The device of claim 24, wherein the illuminator is a broadband LED 34. The device of claim 33, wherein the broadband LED is a white LED 35. The device of claim 33, wherein the broadband LED emits at least a portion of an emitted light in the ultraviolet or infrared wavelength range. 36. The device of claim 33, wherein the broadband LED is composed of multiple light emitting elements, such as a combination of different light emitting diodes to produce a spectrum of light. 37. The device of claim 33, wherein the broadband LED is comprised of a combination of different light emitted diodes to produce a broadband and continuous spectrum of light. 38. The device of claim 36, wherein the multiple light emitting elements operate in at least one wavelength band. 39. The device of claim 33, wherein the broadband LED contains a blue LED and a phosphor. 40. The device of claim 33, wherein the illuminator operates to produce at least a portion of its illumination in the infrared spectrum. 41. The device of claim 33, wher... | - | A61B005/1459 A61B005/1455 A61B001/06
US20080306356 | A1 | 2008-12-11 | Vascular status monitoring system | 2008-06-05 | A system including an adjustable strap having at least one sensor for monitoring a physical condition of a patient. The strap bears measurement indicia so that a clinician can establish baseline perimetral measurements of a patient's appendage and later account for changes in swelling that may be attributable to inflammation or edema. The system further includes a user-interactive unit for receiving, storing and processing patient information from the at least one sensor. In addition, the system desirably includes standardized forms for documenting a patient's physical health as a function of his/her sensed physical condition. | 1. A vascular status monitoring system comprising: an adjustable strap adapted for wrapping about an appendage of a patient, said strap carrying at least one physiological sensor for monitoring at least one physiological characteristic of the patient;a computer in communication with said at least one physiological sensor; anda form for recording physiological data monitored by said at least one physiological sensor. 2. The system of claim 1 wherein said computer comprises a portable computer. 3. The system of claim 1 wherein said form is a paper form. 4. The system of claim 1 wherein said form is an electronic form stored on said computer. | - | A61B005/02
US20080275348 | A1 | 2008-11-06 | Monitor device and use thereof | 2008-04-30 | In an apparatus and method for obtaining an indication of energy expenditure by a mammal during exercise, one or more movement transducers (1) each output a respective movement signal related to physical movement. A frequency analysis (3) is performed on at least one of the movement signals to obtain a frequency analysis result. Classification means (5) determines from the frequency analysis result, what class of physical movement is involved in the exercise. Selection means (7) selects a form of calculation according to a class determined by the classification means. A form of calculation selected by the selection means is applied (9) to at least one of the movement signals so as to obtain the energy expenditure indication. | 1. An apparatus for obtaining an indication of energy expenditure by a mammal during exercise, the apparatus comprising: (i) one or more moement transducers, each for outputting a respective movement signal related to physical movement;(ii) analysis means for performing an analysis on at least one of the movement signals to obtain an analysis result;(iii) classification means for determining from the analysis result, what class of physical movement is involved in the exercise;(iv) selection means for selecting a form of calculation according to a class determined by the classification means; and(v) means for applying a form of calculation selected by the selection means to at least one of the movement signals so as to obtain said energy expenditure indication. 2. An apparatus according to claim 1, wherein the analysis means is adapted to perform an analysis of the movement signals which comprises an analysis of frequency components. 3. An apparatus according to claim 2 wherein the analysis means is: adapted to perform an analysis of frequency components which comprises Fourier analysis. 4. An apparatus according to claim 2 wherein the analysis means is: adapted to perform an analysis of frequency components which comprises wavelet analysis. 5. An apparatus according to claim 1, wherein the analysis means is adapted to perform an analysis of the movement signals which comprises mapping of movement vectors to create a vector surface. 6. An apparatus according to claim 1, wherein the classification means is adapted to determine the class of physical movement from the analysis result by comparison of the analysis result with members of a library of stored data sets each indication of a respective different class of analysis result to determine which stored data set best matches the analysis result. 7. An apparatus according to claim 6, further comprising a memory for storing the stored data sets and calibration means for creating the stored data sets from calibration results obtained from the analysis means. 8. An apparatus according to claim 7 wherein the calibration means is adapted to update the stored data sets by means of an adaptive empirical method. 9. An apparatus according to claim 8, wherein the adaptive empirical method uses a Kalman filter or a neural network. 10. An apparatus according to claim 1, wherein said one or more transducers is or are, selected from any of accelerometers, velocity transducers and pedometers. 11. An apparatus according to claim 1, comprising at least two of said transducers arranged to produce a respective movement signal related to physical movement in respective different directions. 12. An apparatus according to claim 1, further comprising one or more secondary transducers, each for outputting a respective secondary indication signal, related to respective one or more physiological parameters. 13. An apparatus according to claim 12, wherein said classification means is arranged also to utilise said... | Conopco, Inc.d/b/a Unilever | A61B005/11 A61B005/01 A61B005/0205
US20080255441 | A1 | 2008-10-16 | Add-On For Invasive Probe | 2005-12-01 | An invasive probe assembly including an invasive probe having a surface area and a tool carrier removably mounted on the invasive probe, including a substrate having an area smaller than 25% of the surface area of the probe and at least one tool mounted on the substrate. The layout of the substrate or of one or more of the at least one tool is optionally substantially different in different sectors of the circumference of the probe. | 1. An invasive probe assembly, comprising: an invasive probe having a surface area; anda tool carrier removably mounted on the invasive probe, including a substrate having an area smaller than 25% of the surface area of the probe and at least one tool mounted on the substrate,wherein the layout of the substrate, or of one or more of the at least one tool is substantially different in different sectors of the circumference of the probe. 2. An assembly according to claim 1, wherein the substrate has a width shorter than 3 cm, measured in an axial dimension of the invasive probe. 3. An assembly according to claim 2, wherein the substrate has a width shorter than 1 cm. 4. An assembly according to claim 1, wherein the substrate substantially entirely surrounds the circumference of the probe. 5. An assembly according to claim 1, wherein the substrate is attached to the probe by an adhesive. 6. An assembly according to claim 1, wherein the substrate includes a protective electrical or heat isolating layer, which is located between the tool and the probe. 7. An assembly according to claim 1, wherein the at least one tool comprises an electrode. 8. An assembly according to claim 7, wherein the substrate is electrically conductive, serving as a part of the electrode. 9. An assembly according to claim 7, wherein the at least one tool comprises only a single electrode. 10. An assembly according to claim 1, wherein the substrate is electrically isolative. 11. An assembly according to claim 1, wherein the at least one tool comprises at least one sensor and at least one electrode. 12. An assembly according to claim 1, comprising a protective sheath that covers the probe and wherein the substrate is mounted on the protective sheath. 13. An assembly according to claim 1, wherein the substrate does not cover a distal 5 mm portion of the probe. 14. An assembly according to claim 1, wherein over most of the length of the substrate, the substrate covers less than 20% of the circumference of the probe. 15. An assembly according to claim 1, wherein the tool comprises a flat electrode and wherein attaching the electrode to the probe increases the cross-section diameter of the probe by less than 2 millimeters. 16. An assembly according to claim 1 wherein the tool comprises an electrode which conforms to the shape of the probe. 17. An assembly according to claim 1, wherein the tool carrier is flexible. 18. An assembly according to claim 17, wherein the tool carrier is more flexible than the invasive probe. 19. An assembly according to claim 1, wherein the tool carrier includes a wire electrically coupled to the tool and extending to a proximal end of the probe. 20. An assembly according to claim 19, comprising the wire is fastened to the probe using a ring at a proximal side of the probe. 21-76. (canceled) | - | A61B005/042
US20080243005 | A1 | 2008-10-02 | Computational user-health testing | 2007-06-11 | Methods, apparatuses, computer program products, devices and systems are described that carry out obtaining user-health data; selecting at least one user-health test function at least partly based on the user-health data; and applying the at least one user-health test function to at least one interaction between at least one user and at least one device-implemented application whose primary function is different from symptom detection. | 1. A method comprising: obtaining user-health data;selecting at least one user-health test function at least partly based on the user-health data; andapplying the at least one user-health test function to at least one interaction between at least one user and at least one device-implemented application whose primary function is different from symptom detection. 2. The method of claim 1 wherein the obtaining user-health data comprises: obtaining user health attribute data. 3. The method of claim 2 wherein the obtaining user health attribute data comprises: obtaining user health condition data or user symptom data. 4. The method of claim 2 wherein the obtaining user health attribute data comprises: obtaining user medication data or user nutraceutical data. 5. The method of claim 1 wherein the obtaining user-health data comprises: obtaining user health measurement data. 6-7. (canceled) 8. The method of claim 1 wherein the obtaining user-health data comprises: obtaining user health testing data. 9. The method of claim 8 wherein the obtaining user health testing data comprises: obtaining user mental health testing data. 10. The method of claim 8 wherein the obtaining user health testing data comprises: obtaining user physical health testing data. 11. The method of claim 1 wherein the obtaining user-health data comprises: obtaining user-health test function output data. 12-17. (canceled) 18. The method of claim 1 wherein the selecting at least one user-health test function at least partly based on the user-health data comprises: selecting at least one mental status test function. 19. The method of claim 1 wherein the selecting at least one user-health test function at least partly based on the user-health data comprises: selecting at least one cranial nerve test function. 20. The method of claim 1 wherein the selecting at least one user-health test function at least partly based on the user-health data comprises: selecting at least one cerebellum test function. 21. The method of claim 1 wherein the selecting at least one user-health test function at least partly based on the user-health data comprises: selecting at least one of an alertness test function, an attention test function, a memory test function, a speech test function, a calculation test function, a neglect test function, a construction test function, or a task sequencing test function. 22-23. (canceled) 24. The method of claim 1 wherein the applying the at least one user-health test function to at least one interaction between at least one user and at least one device-implemented application whose primary function is different from symptom detection comprises: applying the at least one user-health test function to the at least one interaction between the at least one user and the at least one device-implemented application whose primary function is different from symptom detection, the at least one interaction including user input data. 25. The method of claim 1 wherein the... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00
US20080287752 | A1 | 2008-11-20 | EAR CANAL PHYSIOLOGICAL PARAMETER MONITORING SYSTEM | 2008-05-09 | Physiological parameters are monitored from within an ear canal. A sensor module, which is mountable within the ear canal, includes at least one sensor operable to sense the physiological parameters and to generate signals based on the sensed physiological parameters. A processor, which is in communication with the sensor module, processes the signals from the at least one sensor and generates data related to the physiological parameters. | 1. A system for monitoring one or more physiological parameters from within an ear canal, the system comprising: a sensor module securable within the ear canal and including at least one sensor operable to sense the one or more physiological parameters and to generate signals based on the sensed physiological parameters; anda processor in communication with the sensor module and operable to process the signals from the at least one sensor and to generate data related to the one or more physiological parameters. 2. The system of claim 1, and further comprising: a communication module operable to transmit the data from the processor to a remote location. 3. The system of claim 2, wherein the communication module is further operable to receive electronic communications from the remote location. 4. The system of claim 2, wherein the processor compares the data related to each physiological parameter to a corresponding parameter threshold. 5. The system of claim 4, wherein the communication module communicates with the remote location when at least one of the one or more physiological parameters reaches the corresponding parameter threshold. 6. The system of claim 1, and further comprising: an output device in communication with the processor that is operable to generate a sensory output based on the data. 7. The system of claim 6, wherein the output device generates the sensory output when at least one of the one or more physiological parameters reaches a parameter threshold. 8. A system for monitoring physiological parameters, the system comprising: a plurality of physiological sensors each operable to sense at least one physiological parameter;an ear module securable within an ear canal that houses the plurality of physiological sensors and positions the plurality of physiological sensors relative to the ear canal to sense the physiological parameters; anda processor in communication with the plurality of physiological sensors and operable to process the signals from the physiological sensors and to generate data related to the one or more physiological parameters. 9. The system of claim 8, and further comprising: a communication module operable to transmit the data from the processor to a remote location. 10. The system of claim 9, wherein the communication module is further operable to receive electronic communications from the remote location. 11. The system of claim 9, wherein the processor compares the data related to each physiological parameter to a corresponding parameter threshold. 12. The system of claim 11, wherein the communication module communicates with the remote location when at least one of the one or more physiological parameters reaches the corresponding parameter threshold. 13. The system of claim 12, wherein the ear module further includes an audio assembly, and wherein audio signals are communicated between the ear module and the remote location via the audio assembly when at least one of the one or more physiolog... | Mayo Foundation for Medical Education and Research | A61B005/02 A61B005/12
US20080269577 | A1 | 2008-10-30 | INSTRUMENT FOR MEASURING CARBON MONOXIDE POISONING OF HUMANS USING IN VIVO NIRS TECHNOLOGY | 2007-04-25 | Spectral, frequency based non-invasive procedure for determination of blood constituents utilizing in vivo NIRS (Near-Infrared Spectrum) technology, which is the measurement of the near-infrared absorption spectrum within a region of the living human body for the purpose of identifying tissue and blood components and their concentrations and more particularly to novel applications and methodology for determining the optical response, measurements and calculations relating to the concentrations of individual chromophores in the bloodstream and particularly to the level of CO chromophores in the tissues of an animal or human being. | 1. A method for non-invasively determining the relative amounts of a plurality of chromophores in the blood of a human or animal subject, the method comprising the steps of: imparting a wavelength range of light comprising visible and near-infrared components to a light delivery fiberoptic;passing the wavelength range of light through a selected tissue sample of the human or animal subject and receiving an unabsorbed portion of the wavelength range of light in a light collecting fiberoptic;providing a spectrometer for receiving the unabsorbed portion of the wavelength range of light from the light collecting fiberoptic;determining a total absorption spectrum of the selected tissue sample for a plurality of wavelengths selected from the wavelength range of light; andcomparing an absorption of the plurality of chromophores at a first wavelength to an absorption of the plurality of chromophores at one or more additional wavelengths to determine a relative amount of each of the plurality of chromophores in the blood of a human or animal subject. 2. The method for non-invasively determining the relative amount of at least one chromophore in the blood of a human or animal subject as set forth in claim 1 further comprising the step of impinging the selected tissue surface portion with broadband infrared light. 3. The method for non-invasively determining the relative amount of at least one chromophore in the blood of a human or animal subject as set forth in claim 1 further comprising the step of impinging the selected tissue surface portion with broadband infrared light in a range equal to and greater than 450 nm. 4. The method for non-invasively determining the relative amount of at least one chromophore in the blood of a human or animal subject as set forth in claim 1 further comprising the step of impinging different wavelengths of light on a nasal septum of the human or animal subject and detecting the spectral distribution of the light emerging from the selected tissue sample; and determining multiple absorption values of the light in this selected tissue sample. 5. The method of for non-invasively determining the relative amount of at least one chromophore in the blood of a human or animal subject as set forth in claim 1 further comprising the step of determining a first absorption value of the light in the hemoglobin of the arterial blood flow through the selected tissue sample and comparing to at least a second absorption value determined at a different time from the first absorption value, and evaluating the rate of change between the first absorption value and the second absorption value. 6. The method for non-invasively determining the relative amount of at least one chromophore in the blood of a human or animal subject as set forth in claim 1 further comprising the step of determining the relative percentages of oxyhemoglobin, carboxyhemoglobin and deoxyhemoglobin in the blood in accordance with multiple wavelength absorption values. ... | NEUROPHYSICS CORPORATION | A61B005/00
US20080262332 | A1 | 2008-10-23 | INFECTION MONITORING | 2007-04-19 | A method includes monitoring an indicator of infection in proximity to an implanted active therapy delivering medical device and providing information regarding the monitored indicator to the implanted medical device. The method further includes determining whether the information regarding the monitored indicator is indicative of infection in proximity to the implanted medical device. | 1. A method comprising: monitoring an indicator of infection in proximity to an implanted active therapy delivering medical device;providing information regarding the monitored indicator to the implanted medical device; anddetermining whether the information regarding the monitored indicator is indicative of infection in proximity to the implanted medical device. 2. The method of claim 1, wherein determining whether the information regarding the monitored indicator is indicative of infection is performed in the implanted device. 3. The method of claim 1, further comprising transmitting the information regarding the monitored indicator to an external device, wherein determining whether the information regarding the monitored indicator is indicative of infection is performed in the external device. 4. The method of claim 1, wherein the implanted medical device is an infusion device. 5. The method of claim 1, wherein the implanted device is an electrical signal generator. 6. The method of claim 1, wherein monitoring the indicator of infection comprises monitoring temperature. 7. The method of claim 1, wherein monitoring the indicator of infection comprises monitoring pH. 8. The method of claim 1, wherein monitoring the indicator of infection comprises monitoring impedance. 9. The method of claim 1, wherein monitoring the indicator of infection comprises monitoring a biological indicator of infection. 10. An active implantable therapy delivering medical device comprising: a hermetically sealed housing;a power source disposed in the housing;electronics disposed in the housing and operably coupled to the power source, the electronics capable of controlling therapeutic output from the device;a sensor located in proximity to the housing and being operably coupled to the electronics, the sensor capable of detecting information regarding an indicator of infection and transmitting the information to the electronics; anda computer readable medium containing instructions that when implemented by the electronics cause the device to determine whether the information is indicative of infection. 11. The device of claim 10, further comprising an alarm operably coupled to the electronics, wherein the computer readable medium further contains instructions that when implemented by the electronics cause the device to activate the alarm if a determination is made that the information is indicative of infection. 12. The device of claim 10, wherein the active implantable therapy delivering medical device is an infusion device. 13. The device of claim 10, wherein the active implantable therapy delivering medical device is an electrical signal generator. 14. A system comprising: an active implantable therapy delivering medical device comprising: a hermetically sealed housing;a power source disposed in the housing;first electronics disposed in the housing and operably coupled to the power source, the first electronics capable of controlling therapeutic output f... | MEDTRONIC, INC. | A61B005/145 A61B005/01 A61B005/053
US20080269586 | A1 | 2008-10-30 | ELECTROPORATION TO INTERRUPT BLOOD FLOW | 2008-06-27 | A method for disrupting blood flow to undesirable tissue such as cells of a cancerous or non-cancerous tumor is disclosed. It involves the placement of electrodes into or near the vicinity of vessels supplying blood to the undesirable tissue and through the application of electrical pulses causing blood flow disruption. The electric pulses irreversibly permeate the cell membranes, thereby invoking cell death. The irreversibly permeabilized cells are left in situ and are removed by the body immune system. The process may further comprise monitoring blood flow and/or infusion of a material such as a chemotherapeutic agent or marker into the blood. | 1.-36. (canceled) 37. A device for disrupting blood flow, comprising: first and second electrodes that position a blood vessel to be disrupted therebetween;a voltage generator means that applies a voltage between the first and second electrodes in a manner which provides a predetermined electric field around the blood vessel for a time sufficient to disrupt blood flow through the vessel. 38. The device of claim 37, wherein the generator means generates pulses of 100 microseconds ±about 10 microseconds at a voltage gradient in a range of from about 50 volt/cm to about 8000 volt/cm. 39. The device of claim 37, further comprising: a means for adjusting the voltage and pulse duration of the generator means to obtain blood flow disruption. 40. The device of claim 37, further comprising: a means for monitoring blood flow. 41.-53. (canceled) 54. The device of claim 37, wherein the generator is designed to apply electrical pulses for a duration in a range of from about 5 microseconds to about 62 seconds. 55. The device of claim 37, wherein the generator is designed to apply electrical pulses for a period of about 100 microseconds, ±about 10 microseconds. 56. The device of claim 37, wherein the generator is designed so that on activation it applies from about 1 to about 15 pulses. 57. The device of claim 37, wherein the generator is designed so that on activation it applies about eight pulses of about 100 microseconds each in duration. 58. The device of claim 37, wherein the generator is designed so that on activation it applies pulses that produce a voltage gradient in a range of from about 50 volt/cm to about 8000 volt/cm. 59. The device of claim 37, wherein the first electrode is positioned at about 5 mm to 10 cm from the second electrode. 60. The device of claim 37, further comprising: a monitor for monitoring blood flow in blood vessels. 61. The device of claim 60, further comprising: a means for adjusting the electrical pulses based on real time information of the monitored blood flow. 62. The device of claim 37, further comprising: a monitor for monitoring temperature of the identified tissue; anda means for adjusting the electrical pulses to maintain the temperature at 100° C. or less. 63. The device of claim 62, wherein the monitor maintains the temperature at 50° C. or less. 64. A device of claim 37, further comprising: a means for adjusting applied voltage, length of pulses, and number of pulses to obtain blood flow disruption to the target area thereby minimizes damage to non-target tissue. 65. The device of claim 37, further comprising: a means for adjusting duration of applied voltage is in accordance with current-to-voltage ratio to achieve blood flow disruption to a targeted area. 66. The device of claim 65, wherein the means for adjusting, adjusts the current-to-voltage ratio based on temperature to maintain target tissue temperature at 100° C. or less. 67. The device of claim 65, wherein the means for adjusting, adjust... | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA | A61B018/18 A61B005/026 A61B005/01
US20080275307 | A1 | 2008-11-06 | Telemetry System With out of Range Notification Features | 2005-01-28 | An apparatus and method for notifying users of a telemetry device as they move out of the coverage area are provided. As the user of the telemetry device has moved out of the coverage area, the device notifies the user after attempting to establish a connection with the network for a predetermined time period, by transmitting an audio signal and/or displaying a visual signal. | 1. A telemetry apparatus for notifying a user of moving out of a coverage area, comprising: means for establishing a connection with a network in communication with a central monitoring station; andmeans for transmitting an audio signal as the user moves out of the coverage area of the network. 2. The telemetry apparatus of claim 1, further comprising means for collecting patient data of the user. 3. The telemetry apparatus of claim 2, further comprising means for processing the patient data for a subsequent transmission to the central monitoring station. 4. The telemetry apparatus of claim 1, further comprising means for displaying a visual signal as the user moves out of the coverage area of the network. 5. The telemetry apparatus of claim 4, wherein the visual signal is transmitted after a predetermined time period has expired. 6. The telemetry apparatus of claim 4, wherein the visual signal is transmitted after a predetermined number of attempts to reconnect to the network has failed. 7. The telemetry apparatus of claim 1, wherein the audio signal is transmitted after a predetermined time period has expired. 8. The telemetry apparatus of claim 1, wherein the audio signal is transmitted after a predetermined number of attempts to reconnect to the network has failed. 9. A system for notifying a user of moving out of a coverage area, comprising: a central monitoring station wirelessly communicating with a plurality of telemetry devices for processing patient data;wherein, said telemetry device includes:means for establishing a connection with the central monitoring station; andmeans for transmitting an audio or visual signal as the user moves out of the coverage area of the network. 10. The system of claim 9, further comprising means for collecting the patient data of the user. 11. The system of claim 9, wherein the audio or visual signal is transmitted after a predetermined time period has expired. 12. The system of claim 9, wherein the audio or visual signal is transmitted after a predetermined number of attempts to reconnect to the network has failed. 13. A method for notifying a user of moving out of a coverage area, the method comprising the steps of: providing a central monitoring station in communication with a plurality of telemetry devices in a network;continuously collecting patient data from the respective telemetry devices to the central monitoring station;detecting when at least one of the telemetry devices moves out of the coverage area of the network; and,transmitting an audio or visual signal when the user moves out of the coverage area of the network. 14. The method of claim 13, wherein the detecting step further comprises the step of detecting whether at least one of the telemetry devices loses the connection with the central monitoring station for more than a predetermined time period. 15. The method of claim 14, wherein the audio or visual signal is transmitted after a predetermined number of attempts to reconne... | Koninklijke Philips Electronics N.V. | A61B005/00
US20080269580 | A1 | 2008-10-30 | System for Non-Invasive Measurement of Bloold Glucose Concentration | 2006-12-12 | A system and method for non-invasive measurement of glucose concentration in a live subject including a thermal emission spectroscopy (TES) device 10, an optical coherence tomography (OCT) device 20 or near infrared diffuse reflectance (NIDR) device. The TES 10 generates a signal indicative of the absorbtion of glucose, from which the blood glucose concentration is determined and the OCT device 20 generates a signal indicative of the scattering coefficient of a portion of the live subject, from which the blood glucose concentration is determined. The signals generated by the TES and OCT devices along with signals generated by sensors for measuring the body heat and surface temperature of the subject are used in the metabolic heat conformation (MHC) method of determining blood glucose concentration. The system may include a photoacoustic sensor for generating a signal indicative of thermo-elastic skin properties from which the blood glucose concentration is also determined. | 1. A system for the non-invasive measurement of blood glucose concentration in a live subject comprising: a. means for determining the body heat of the subject,b. means for determining the concentration of haemoglobin and oxygenated haemoglobin in the blood of said live subject, andc. means for determining blood flow velocity in respect of said live subject and means for determining blood glucose concentration in said live subject as a function of said body heat, said haemoglobin and oxygenated haemoglobin concentrations and said blood flow velocity; anda plurality of spectroscopic devices each generating a signal indicative of blood glucose concentration, means for determining the blood glucose concentration from the signal indicative of blood glucose concentration and wherein at least one of the spectroscopic devices generates a signal additionally indicative of one or more of:d. concentration of haemoglobin and oxygenated haemoglobin in the blood of the live subject;e. the body heat of the live subject;f. ambient temperature;g. blood flow velocity in respect of said live subject,wherein the signal indicative of one or more of d. to g. are transmitted to at least one of means a. to c. and used to determine the blood glucose concentration. 2. A system according to claim 1 wherein one of the spectroscopic devices comprises x) a detector for detecting the thermal emission spectrum emitted by said live subject and generating a signal indicative of the absorption of glucose. 3. A system according to claim 1 wherein one of the spectroscopic devices comprises y) an irradiator for irradiating a portion of the live subject with a measuring beam and a detector for collecting measuring beam radiation scattered by said live subject and generating a signal indicative of the scattering coefficient of the portion of the subject. 4. A system according to claim 3, wherein the measuring beam is in the near infrared spectrum and/or has multiple wavelengths. 5. A system according to claim 1, wherein one of the spectroscopic devices comprises z) a source for pulsed irradiation of a portion of the live subject and a detector for detecting an acoustic pressure wave caused by the pulsed irradiation and generating a signal indicative of the thermo-elastic skin properties. 6. A system according to claim 3, wherein the signal indicative of the scattering coefficient is used to isolate the thermo-elastic skin properties in the signal indicative of the thermo-elastic skin properties from scattering effects. 7. A system according to claim 2, wherein the spectroscopy device comprises interference filtering means for spatially separating said thermal emission spectrum to create a plurality of spectral patterns and measuring in respect of each of a plurality of said spectral patterns a spectral intensity at a first, reference set of wavelengths, and a second set of wavelengths dependent on glucose or other analyte, and the concentration of glucose or other analyte is det... | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61B005/01 A61B005/1468 A61B008/06
US07465276 | B2 | 2008-12-16 | Identification of disease characteristics using isotope ratios in breath | 2005-11-30 | Methods are disclosed for distinguishing whether an animal is experiencing a bacterial infection or a viral infection. One monitors breath taken from the animal over time to measure the relative amount of a first breath stable isotope to a second breath stable isotope therein over time. A quick change in the isotope ratios within several hours from the likely infection is indicative of a bacterial infection. A delayed change in the isotope ratios, followed by periodic repeated alterations in the ratios, is indicative of viral infection. The methods are particularly efficient when using cavity ringdown spectroscopy for the monitoring. They may be used for monitoring a patient already admitted to a hospital, or for monitoring a patient initially complaining of adverse symptoms, or for triage, or for collectively monitoring a population of animals. | 1. A method of distinguishing whether an isotopically unenriched animal is experiencing a bacterial infection, the method comprising: monitoring breath taken from the animal over time to measure a relative amount of a first isotope to a second isotope therein over time, wherein the first and second isotopes are selected from the group consisting of a pair of 13C and 12C, a pair of 15N and 14N, a pair of 17O and 16O, and a pair of sulphur isotopes; andbased on a comparison of (i) changes to the relative amount of the first isotope to the second isotope over time, with (ii) a known pattern of changes of the relative amount of the first isotope to the second isotope when there is bacterial infection determining whether bacterial infection is present. 2. The method of claim 1, wherein the first and second isotopes are the pair of 13C and 12C. 3. The method of claim 2, wherein the monitoring step uses cavity ringdown spectroscopy to determine the changes in the relative amount. 4. A method of distinguishing whether an animal is experiencing a bacterial infection, the method comprising: monitoring breath taken from the animal over time to measure a relative amount of a first isotope to a second isotope therein over time, wherein the first and second isotopes are selected from the group consisting of a pair of 13C and 12C, a pair of 15N and 14N, a pair of 17O and 16O, and a pair of sulphur isotopes; andbased on a comparison of (i) changes to the relative amount of the first isotope to the second isotope over time, with (ii) a known pattern of changes of the relative amount of the first isotope to the second isotope when there is bacterial infection determining whether bacterial infection is present;wherein the method comprises monitoring breath taken from the animal at at least three different times to measure the relative amount of the first isotope to the second isotope therein over such time; anda first such monitoring time is more than two hours earlier than a third such monitoring time. 5. The method of claim 4, wherein: the method comprises monitoring breath taken from the animal at at least four different times to measure the relative amount of the first isotope to the second isotope therein over such time; anda first such monitoring time is more than three days earlier than another such monitoring time. 6. The method of claim 5, wherein based on a comparison of (i) changes to the relative amount of the first isotope to the second isotope over time, with (ii) a known pattern of changes of the relative amount of the first isotope to the second isotope when there is viral infection determining whether viral infection is present. 7. The method of claim 1, wherein the animal is mammalian. 8. The method of claim 7, wherein the animal is human. 9. A method of distinguishing whether an animal is experiencing a bacterial infection, the method comprising: monitoring breath taken from the animal over time to measure a relative amount of a first iso... | Wisconsin Alumni Research Foundation | A61B005/08 G01N033/497
US07452335 | B2 | 2008-11-18 | Method and apparatus for the detection of neuromuscular disease using disease specific evoked neuromuscular response analysis | 2002-11-06 | A method and apparatus for assessing neuromuscular pathology in an individual, by (a) applying a plurality of stimuli to a peripheral nerve; (b) recording from the peripheral nerve, or from a muscle innervated by the peripheral nerve, at least one response to each stimulus; (c) processing each response into at least one response characteristic; (d) compiling the response characteristics from all of the responses into a plurality of descriptive parameters of the response characteristics; and (e) analyzing the plurality of descriptive parameters according to a function specific to the neuromuscular pathology so as to yield an indicator of the pathology. | 1. A method for the determination of a composite pathology function whose output indicates a neuromuscular pathology in an individual, the method comprising the steps of: (a) selecting a first group of individuals that have the neuromuscular pathology;(b) applying a plurality of stimuli to a peripheral nerve of each individual in said first group of individuals so as to evoke a plurality of responses to said stimuli from said peripheral nerve or from a muscle innervated by said peripheral nerve;(c) recording said responses;(d) processing each of said responses into at least one selected response characteristic;(e) compiling said at least one response characteristic from all of said responses into a first set of selected descriptive parameters characterizing said first group of individuals;(f) selecting a second group of individuals that do not have the neuromuscular pathology;(g) repeating steps (b), (c), (d) and (e) on each individual of said second group of individuals that do not have the neuromuscular pathology to generate a second set of said descriptive parameters characterizing said second group of individuals; and(h) processing said first and second sets of descriptive parameters so as to determine a composite pathology function of said descriptive parameters that is indicative of the neuromuscular pathology, said composite pathology function being characterized in that the composite pathology function integrates the descriptive parameters into a single index represented by the following equation:NCcomposite=ⅇ∑i=1N⁢αi⁢pi1+∑i=1N⁢αi⁢piwhere N represents a total number of predictive nerve conduction parameters, pi is the i-th nerve conduction parameter value and αi is a corresponding weighting coefficient. 2. The method of claim 1 wherein said processing of said first and second sets of descriptive parameters is performed using logistic regression analysis. 3. The method of claim 1 wherein said processing of said first and second sets of descriptive parameters is performed using an artificial neural network. 4. The method of claim 1 wherein said processing of said first and second sets of descriptive parameters is performed using linear discriminate analysis. 5. The method of claim 1 wherein said processing of said first and second sets of descriptive parameters is performed using linear multivariate regression. 6. The method of claim 1 wherein said processing of said first and second sets of descriptive parameters is performed using clustering. | NeuroMetrix, Inc. | A61B005/05 A61B005/04
US07466848 | B2 | 2008-12-16 | Method and apparatus for automatically detecting breast lesions and tumors in images | 2003-12-15 | A method and apparatus for automatically detecting breast tumors and lesions in images, including ultrasound, digital and analog mammograms, and MRI images, is provided. An image of a breast is acquired. The image is filtered and contrast of the image is enhanced. Intensity and texture classifiers are applied to each pixel in the image, the classifiers indicative of the probability of the pixel corresponding to a tumor. A seed point is identified within the image, and a region of interest is grown around the seed point. Directional gradients are calculated for each pixel of the image. Boundary points of the region of interest are identified. The boundary points are passed as inputs to a deformable model. The deformable model processes the boundary points to indicate the presence or absence of a tumor. | 1. A method for automatically detecting breast tumors and lesions in an image comprising: acquiring an image of a breast;filtering the image;applying texture and intensity classifiers to each pixel of the image, the classifiers corresponding to probabilities of the pixel belonging to a lesion or tumor;determining a seed point in the image by retrieving a set of points of interest in the image, selecting a first point from the set of points, calculating a joint probability that the first point corresponds to a tumor, calculating mean joint probabilities that points in a circular region around the first point correspond to a tumor, and designating a point within the circular region having a maximum mean joint probability as the seed point;growing a region of interest around the seed point;calculating directional gradients for each pixel in the image;determining boundary points of the region of interest using the directional gradients; andprocessing the boundary points with a deformable model to determine the presence or absence of a tumor or lesion in the image. 2. The method of claim 1, wherein the step of acquiring the image comprises digitizing the image from an analog mammogram. 3. The method of claim 1, wherein the step of acquiring the image comprises acquiring a digital mammogram, ultrasound, or MRI image of a breast. 4. The method of claim 1, wherein the step of filtering the image comprises removing speckle from the image using a Butterworth filter. 5. The method of claim 4, further comprising enhancing contrast of the image. 6. The method of claim 1, wherein the step of applying texture and intensity classifiers comprises determining intensity and local variance of each pixel of the image. 7. The method of claim 6, further comprising applying a texture probability distribution function to the local variance of the pixel to produce the texture classifier. 8. The method of claim 6, further comprising applying an intensity probability distribution function to the intensity of the pixel to produce the intensity classifier. 9. The method of claim 1, wherein the step of growing the region of interest comprises: adding the seed point to the region of interest; andadding pixels to the region of interest based upon connectivity and values of surrounding pixels. 10. The method of claim 1, wherein the step of determining boundary points comprises scanning the region of interest horizontally and vertically to determine edge points, and combining the edge points. 11. The method of claim 10, further comprising drawing radial lines from the seed point and plotting boundary points corresponding to positions of maximum intensity on the radial lines. 12. The method of claim 11, further comprising removing outliers and local maxima from the boundary points. 13. An apparatus for automatically detecting breast tumors and lesions in an image comprising: a scanner for generating an image of a breast;a filter for filtering the image;texture and inten... | Rutgers, The State University of New Jersey | G06K009/00 G06K009/34 A61B005/05
US07471976 | B2 | 2008-12-30 | Device for measuring electrocardiogram with tapeless format and its method | 2004-01-09 | A device for measuring an electrocardiogram with a tapeless format and its method is provided with electrodes having a flat shape that are contacted by root portions between two fingers of a user's hands or by two fingers of each of the user's hands. The electrocardiogram can be measured directly by that contact, without adding any conductive liquid or gel material. | 1. A device for measuring electrocardiogram with tapeless format comprising: a shell being shaped as a thin and long cube and having a top surface having a left side;a right side;a right upper finger touching area located on the right side of the top surface;a left upper finger touching area located on the left side of the top surface and being parallel with the right upper finger touching area;a bottom surface being opposite to the top surface having; a left side;a right side;a right lower finger touching area located on the right side of the bottom surface; anda left lower finger touching area located on the left side of the bottom surface and being parallel with the right lower finger touching area;a front edge vertically formed between the top and bottom surfaces;a right edge vertically formed between the top and bottom surfaces;a left edge vertically formed between the top and bottom surfaces; anda rear edge vertically formed between the top and bottom surfaces;a right finger gelless electrode with a thin foil shape having a right upper gelless electrode portion embedded in the right upper finger touching area of the top surface and being distant from the right, left and rear edges;a right lower gelless electrode portion formed on the right lower finger touching area of the lower surface and being distant from the right, left and rear edges; anda right middle gelless electrode portion formed on the front edge and connected between the right upper and lower gelless electrodes and being distant from the right and left edges;a left finger gelless electrode with a thin foil shape having a left upper electrode portion embedded in the left upper finger touching area of the top surface and being distant from the right, left and rear edges;a left lower electrode portion formed on the left lower finger touching area of the lower surface and being distant from the right, left and rear edges; anda left middle electrode portion formed on the front edge and connected between the left upper and lower electrodes and being distant from the right and left edges;at least one information display located on the top surface to display a plurality of measured values; anda calculation system mounted in the shell and connected to the two gelless electrodes and the information display for calculating relative electrical information measured from the gelless electrodes and display results on the information display. 2. The device for measuring an electrocardiogram with tapeless format as recited in claim 1, wherein the operating panel has at least one button to set and transfer functions. 3. The device for measuring an electrocardiogram with tapeless format as recited in claim 1, wherein each of the gelless electrodes is made of any conductive metal or rubber. 4. The device for measuring an electrocardiogram with tapeless format as recited in claim 1, wherein information values shown on the information display include at least values of ST segment, QRS interval an... | Kang-Ping Lin | A61B005/0404
US07454244 | B2 | 2008-11-18 | System and method for measuring cross-sectional areas and pressure gradients in luminal organs | 2004-02-19 | The invention comprises a system, catheter and method for measuring the cross-sectional areas and pressure gradients in any hollow organ, such as, for example, blood vessels. One embodiment of such a system includes: an impedance catheter capable of being introduced into a targeted site; a solution delivery source; a constant current source; a balloon inflation control device; and a data acquisition and processing system that receives conductance and/or pressure gradient data from the catheter and calculates the cross-sectional area of the targeted site. In one embodiment, the catheter has an inflatable balloon along its longitudinal axis, thereby enabling the breakup of any materials causing stenosis at the targeted site and/or distention and delivery of an optional stent into the targeted site. | 1. A method for measuring the instantaneous cross-sectional area of a targeted treatment site, comprising: introducing an impedance catheter into a treatment site;providing constant electrical current flow to the treatment site through the catheter;injecting a known volume of a first solution of a first compound having a first conductivity into the treatment site;measuring a first conductance value at the treatment site;injecting a second solution of a second compound having a second conductivity into the treatment site, wherein the second solution has a second volume and wherein the second conductivity does not equal the first conductivity;measuring a second conductance value at the treatment site;calculating the instantaneous cross-sectional area of the treatment site based on the first and second conductance values and the conductivities of the first and second solutions. 2. The method of claim 1, wherein the treatment site comprises a body lumen. 3. The method of claim 2, wherein the body lumen comprises a blood vessel. 4. The method of claim 2, wherein the body lumen comprises a biliary tract. 5. The method of claim 2, wherein the body lumen comprises the esophagus. 6. The method of claim 5, wherein the step of injecting a first solution of a first compound comprises the step of administering said first solution to a patient orally. 7. The method of claim 5, wherein the step of injecting a second solution of a second compound comprises the step of administering said second solution to a patient orally. 8. The method of claim 1, wherein the first compound is NaCl. 9. The method of claim 1, wherein the second compound is NaCl. 10. The method of claim 1, wherein the catheter comprises an inflatable balloon along the longitudinal axis of the catheter. 11. The method of claim 10, further comprising the step of inflating the balloon to breakup any materials causing stenosis at the treatment site. 12. The method of claim 10, wherein the catheter further comprises a stent located over the balloon, said stent capable of being distended to the desired lumen size and implanted into the treatment site. 13. The method of claim 12, further comprising the steps of: distended the stent by inflating the underlying balloon; andreleasing and implanting the stent into the treatment site. 14. The method of claim 1, further comprising the steps of: selecting an appropriately-sized stent based on the cross-section area value of the treatment site; andimplanting the stent into the treatment site. 15. The method of claim 1, wherein the catheter comprising a pressure transducer. 16. The method of claim 15, further comprising the steps of: measuring a first pressure gradient value from the pressure transducer near the treatment site; andcalculating the cross-sectional area of the treatment site based in part on the first gradient pressure value. 17. The method of claim 1, wherein the step of injecting the first solution further includes injecting the f... | Electro-Cat, LLC | A61B005/05 A61B005/117 A61B005/103
US07467012 | B1 | 2008-12-16 | Respiration parameters controlled by heart rate | 2005-06-24 | Heart rate information is used at least in part to obtain one or more parameters for inducing respiration. In various implementations, respiratory parameters, such as a target breathing rate or a target tidal volume may be delivered by an implantable device to a patient during periods of altered respiration, such as sleep apnea or exercise. A respiratory parameter may also be obtained from a physiological variable, a patient's physical activity level, or metabolic demands. | 1. A method, comprising: delivering cardiac stimulation pulses to a patient's heart at a heart rate;obtaining heart rate information for the patient;obtaining a sensed respiratory parameter for the patient;determining a target respiratory parameter based at least in part on the heart rate information;comparing the sensed respiratory parameter to the target respiratory parameter; anddelivering respiratory stimulation if the target respiratory parameter is greater than the sensed respiratory parameter, wherein the delivery of respiratory stimulation is synchronized with the delivery of cardiac stimulation pulses so as to affect respiratory function of the patient without affecting the heart rate, except for every Nth cardiac stimulation pulse, wherein the amplitude of the cardiac pacing stimulation pulse is increased to an energy level sufficient to both stimulate the heart and induce respiration. 2. The method as recited in claim 1, wherein the sensed and target respiratory parameters comprise a breathing rate. 3. The method as recited in claim 2, wherein the target breathing rate is obtained by multiplying a heart rate by a cardiac stroke volume to obtain a cardiac output value, relating the cardiac output value to a respiratory ventilation rate, and dividing the respiratory ventilation rate by a volume of at least part of the patient's lungs. 4. The method as recited in claim 1, wherein the target respiratory parameter comprises a target tidal volume based at least in part on the heart rate information. 5. The method as recited in claim 4, further comprising obtaining the target tidal volume based at least in part on a breathing rate parameter. 6. A device comprising: means for delivering cardiac stimulation pulses to a patient's heart at a heart rate;means for obtaining heart rate information associated with the patient;means for obtaining a native respiratory parameter for the patient;means for determining a target respiratory parameter based at least in part on the heart rate information;means for comparing the native respiratory parameter to the target respiratory parameter; andmeans for delivering respiratory stimulation if the target respiratory parameter is greater than the native respiratory parameter, wherein the delivery of respiratory stimulation is synchronized with the delivery of cardiac stimulation pulses so as to affect respiratory function of the patient without affecting the heart, except for every Nth cardiac stimulation pulse, wherein the amplitude of the cardiac pacing stimulation pulse is increased to an energy level sufficient to both stimulate the heart and induce respiration. 7. The device as recited in claim 6, wherein the native and target respiratory parameters comprise a breathing rate, and means for determining a target respiratory parameter comprise means for squaring a heart rate and dividing the squared heart rate by 360. 8. The device as recited in claim 6, wherein the native and target respiratory parame... | Pacesetter, Inc. | A61N001/05 A61B005/08
US07448141 | B2 | 2008-11-11 | Saddle fit system and method | 2008-02-19 | In an embodiment, a bicycle saddle fit system comprises first and second supports, a substrate coupled to the first and second supports, and a measurement gauge coupled to the substrate. The first and second supports are each configured to receive sit bones of an adult human. At least one of the first and second supports is selectively moveable to locate the sit bones and to define a sit bone width. The substrate is configured to facilitate the selective moveability of the at least one of the first and second supports. The measurement gauge is configured to visually indicate the sit bone width. In another embodiment, a method of fitting a bicycle saddle to a human rider comprises locating the first and second sit bones of the rider, reading a sit bone width from the measurement gauge, and assessing a saddle characteristic based on the rider's sit-bone width. | 1. A bicycle saddle fit system comprising: an impression pad of sufficient size to span the sit bones of an adult human;a substrate layer extending laterally beyond a perimeter of the impression pad; anda measurement gauge comprising a scale, said gauge being positioned relative the impression pad such that the measurement gauge indicates a distance between the impressions between sit-bone impressions made in the impression pad by a person previously seated on the impression pad to enable the distance between impressions to be measured. 2. The fit system of claim 1, wherein the measurement gauge comprises a plurality of marker lines. 3. The fit system of claim 1, wherein the measurement gauge comprises indicia indicating a center point of the measurement gauge. 4. The fit system of claim 1, further comprising a substrate layer fixed to and underlying the impression pad. 5. The fit system of claim 1, wherein the substrate layer includes an exposed portion which extends laterally beyond the perimeter of the impression pad, and the measurement gauge is located on the substrate layer. 6. The fit system of claim 1, wherein the impression pad comprises a layer of foam which is about 5-15 mm thick. | Specialized Bicycle Components, Inc. | A61B005/103
US20080294064 | A1 | 2008-11-27 | Lancing Element, Lancing System and a Method for Skin Detection | 2008-05-09 | In the case of a disposable lancing element for collecting body fluid through the skin comprising a support and a lancing member which distally projects therefrom and can pierce the skin in a lancing movement, it is proposed that a pretensioning means which tightens the skin in the area of the puncture site is arranged on the support and has a contact member which impacts the skin laterally from the lancing member during the lancing movement. | 1. A lancing element for collecting a body fluid through skin, comprising: a support;a lancing member which rigidly projects from the support configured to pierce a skin in a lancing movement; and,a pretensioning means coupled to the support which tightens the skin near a puncture site, the pretensioning means having a contact member that impacts the skin during the lancing movement and moves back during the lancing movement in an opposite direction to the lancing member. 2. The lancing element as in claim 1 where the contact member in an initial state is positioned in a direction of lancing to a side and in front of the lancing member or the contact member and the contact member is located at a same level or downstream of the lancing member and is reset against the support during the lancing movement while in contact with the skin. 3. The lancing element as in claim 2 wherein the contact member is supported on the support via a spring element. 4. The lancing element as in claim 2 wherein the contact member is locked in a reset position that is proximally set back relative to the lancing member by means of a locking mechanism. 5. The lancing element as in claim 2 wherein the contact member is moved together with the support when the support is retracted while reducing or removing a contact pressure on the skin. 6. The lancing element as in claim 1 wherein the pretensioning means has a coupling member that is elastically compressible and held by frictional lock that connects the contact member with the support. 7. The lancing element as in claim 1 wherein the coupling member is formed by a folding arm that is bent having at least one bending point. 8. The lancing element as in claim 1 wherein the pretensioning means is attached as an integral component of the support. 9. The lancing element as in claim 1 wherein the support, the lancing member, and the pretensioning means are formed uniformly from one material as a disposable part. 10. The lancing element as in claim 1 wherein the support, the lancing member, and pretensioning means are formed from a flat substrate that is etched and the pretensioning means is arranged as a part of the substrate in a substrate plane or is bent out of the substrate plane. 11. The lancing element as in claim 1 wherein the contact member has a point shaped or linear edge contour configured to be pressed against the skin laterally next to the lancing member. 12. The lancing element as in claim 1 wherein the lancing member is configured to be inserted into the skin at a lateral distance of less than 3 mm from the contact member. 13. The lancing element as in claim 1, further comprising, a collecting structure that is capillary-active for collecting a body fluid. 14. The lancing element as in claim 1, further comprising, a stop coupled to the support that limits a lancing depth of the lancing member, the stop having a skin contact surface that impacts at a defined proximal distance from the lancing member d... | - | A61B005/151 A61B005/053
US07447541 | B2 | 2008-11-04 | Monitoring subcortical responsiveness of a patient | 2004-06-30 | To monitor subcortical responsiveness of a patient at least one first measurement signal is acquired from a patient and at least one second measurement signal is derived from the at least one first measurement signal. The second measurement signal is indicative of pain-related activity in the patient. A predetermined stimulus signal is then supplied to the patient and a response caused by the stimulus signal in the at least one second measurement signal is monitored in synchronization with the supplying of the stimulus signal. Based on the monitoring step, an indicator of subcortical responsiveness is determined, the indicator being indicative of the patient's level of antinociception. | 1. A method for monitoring subcortical responsiveness of a patient, the method comprising the steps of: acquiring at least one first physiological measurement signal from a patient using at least one sensor, the physiological measurement signal being one of the signals in a group including an RRI signal, a plethysmographic signal, an EMG signal, a blood pressure signal, and a signal representing skin conductivity;deriving at least one second measurement signal from the at least one first measurement signal, said at least one second measurement signal being indicative of pain-related activity in the patient;supplying a predetermined stimulus signal to the patient to cause pain to the patient, wherein the predetermined stimulus signal is supplied in synchronization with a spontaneous rhythm of the patient that is regulated by the autonomous nervous system of the patient;monitoring a response caused by the stimulus signal in the at least one second measurement signal, wherein the monitoring step is synchronized with the supplying step;determining, based on the monitoring step, an indicator of subcortical responsiveness, said indicator being indicative of the patient's level of antinociception; andoutputting the indicator of subcortical responsiveness to a device. 2. A method according to claim 1, further comprising a step of synchronizing the supplying step with at least one of the at least one second measurement signal. 3. A method according to claim 1, wherein the supplying step is synchronized with an external synchronization signal. 4. A method according to claim 1, wherein the supplying step includes supplying a stimulus signal comprising a sequence of stimuli to the patient. 5. A method according to claim 4, wherein the monitoring step includes detecting changes caused by each of said stimuli in the at least one second measurement signal; andthe determining step includes defining the indicator of subcortical responsiveness based on the changes detected. 6. A method according to claim 1, wherein the spontaneous rhythm of the patient is the respiration rhythm of the patient, wherein the synchronizing step includes synchronizing the supplying step with the respiration rhythm of the patient. 7. A method according to claim 1, wherein the acquiring step includes acquiring a single first physiological measurement signal from the patient, whereby a single second measurement signal is obtained in the deriving step. 8. A method according to claim 1, wherein the acquiring step includes a step of measuring a set of physiological signals from the patient. 9. A method according to claim 8, wherein the acquiring step further includes calculating the first physiological measurement signal based on the set of physiological signals. 10. A method according to claim 1, wherein the acquiring steps includes a step of measuring a plurality of first physiological measurement signals from the patient; andthe deriving step includes the steps of normalizing eac... | Instrumentarium Corporation | A61B005/04
US07797042 | B2 | 2010-09-14 | Device for applying a transcutaneous stimulus or for transcutaneous measuring of a parameter | 2007-05-18 | The apparatus for transcutaneous stimulation or transcutaneous measurement of a parameter is a device which fits wholly within the pinna (external ear) of the vertebrate. In order to fit completely within the pinna, the housing is C-shaped and from each end of the housing is a curved section which has a spring action that holds the apparatus inside the ear. This apparatus has a compact design which allows for nerve stimulation in a particularly efficient, ergonomically favorable and easy manipulation for the patient. | 1. An apparatus for transcutaneous stimulation of the skin of a human comprising: a housing configured to fit completely in a pinna of the human;a stimulation unit, housed within the housing, configured to produce an electronic stimulation signal;electronic controls, housed in the housing, configured to control the transcutaneous stimulation;a first curved section having a proximal end and a distal end wherein the proximal end of the first curved section is attached to one end of the housing;a second curved section having a proximal end and a distal end, wherein the proximal end of the second curved section is attached to the other end of the housing, each of the curved sections having a spring action for securing the housing and the curved sections in the pinna;a first electrode mounted at the distal end of the first curved section; anda second electrode mounted at the distal end of the second curved section, wherein the first and second electrodes are electrically connected to the electronic controls and are configured to apply stimulation to the skin. 2. The apparatus of claim 1, wherein the housing has a C-shape. 3. The apparatus of claim 1, wherein the housing and curved sections have an overall C-shape. 4. The apparatus of claim 1, wherein the curved sections are made of wire. 5. The apparatus of claim 1, wherein the electrodes are spherically shaped. 6. The apparatus of claim 1, wherein the housing and the curved sections are designed so that the electrodes rest in a region of or near a tragus. 7. The apparatus of claim 6, wherein the curved sections are designed to rest on the skin above the vagus nerve. 8. The apparatus of claim 1, wherein the housing and the curved sections consist of deformable material. 9. The apparatus of claim 1, wherein the housing consists of a combination of two materials. 10. The apparatus of claim 1, wherein the electronic controls include a rechargeable battery. 11. The apparatus of claim 1, further comprising: an ear phone mounted on the housing. 12. An apparatus for transcutaneous measurement of a parameter from the skin of a human comprising: a housing configured to fit completely in a pinna of the human;a sensing unit, housed within the housing, configured to measure a physiological, a pathophysiological or a physical parameter;electronic controls, housed in the housing, configured to control the measuring of the parameters;a first curved section having a proximal end and a distal end wherein the proximal end of the first curved section is attached to one end of the housing;a second curved section having a proximal end and a distal end, wherein the proximal end of the second curved section is attached to the other end of the housing, each of the curved sections having a spring action for securing the housing and the curved sections in the pinna;a first electrode mounted at the distal end of the first curved section; anda second electrode mounted at the distal end of the second curved section, wh... | Cerbomed GmbH | A61N001/18
US07758519 | B2 | 2010-07-20 | Systems for and methods of assessing lower urinary tract function via sound analysis | 2008-05-19 | Systems for and methods of assessing lower urinary tract function from urinary flow data via sound analysis and user-provided information regarding the lower urinary tract symptoms (LUTS) of frequency, urgency and urge incontinence. Embodiments of the LUTS assessment systems include a computer and a telephone or a digital recording mechanism to capture the sound of one or more urination events, which are stored as audio files in a database. The LUTS assessment systems may include sound analysis software for analyzing the strength and duration of each urination event and may include a web-based software application for viewing the results via the Internet or other network. The database stores information from multiple urination events, and combined with information regarding the lower urinary tract symptoms, serves as an objective tool to assess bladder function and monitor progression of disease or therapy effectiveness. | 1. A method of assessing lower urinary tract function of a person, comprising: capturing the sound of a person's micturition into an air-liquid interface with a sound-capture device;providing information with respect to one or more lower urinary tract symptoms associated with the micturition using the sound-capture device;analyzing the sound of micturition to determine at least one of the following attributes of the micturition: strength and duration; andanalyzing the information with respect to one or more lower urinary tract symptoms together with or without the one or more of said attributes of the micturition for the assessment of lower urinary tract function. 2. A method according to claim 1, wherein the capturing step includes capturing the sound of a person's micturition at a first location and at about the same time providing the sound to a second location that is remote from the first location, the providing step includes providing the information from the first location to the second location, and the analyzing step includes analyzing the sound of micturition at the second location. 3. A method according to claim 1, further including the step of displaying (i) the information and (ii) at least one of the strength and duration attributes, in a user interface. 4. A method according to claim 1, wherein the sound-capture device is a telephone. 5. A method according to claim 1, where in the sound-capture device is microphone not part of a telephone. 6. A method according to claim 2, wherein the sound of micturition and the information is transferred to the second location using a telecommunication network. 7. A method according to claim 1, wherein the providing step includes using an interactive voice response system to assist a user in providing the information. 8. A method according to claim 2, further including analyzing at the second location the information provided over a first period of time to determine frequency of micturition during the first period of time. 9. A method according to claim 1, further including determining frequency of micturition over a first time period based on the time between successive transfers of the sound of micturition during the first time period. 10. A method according to claim 1, wherein the capturing step further includes converting the sound of micturition to a digital file and storing the file in an electronic database. 11. A method according to claim 1, further including determining whether or not the person has bladder overactivity or other lower urinary tract dysfunction based on the information and the sound of micturition. 12. A method according to claim 1, wherein the providing step is performed 1 second to 5 minutes before or after the capturing step. 13. A method according to claim 1, wherein the information includes an identification of the urgency of the micturition. 14. A method according to claim 1, wherein the information includes an indication whether or not the micturition ... | University of Vermont and State Agriculture College | A61B005/00 B65D081/00 A61B005/103 A61B005/117
US07452333 | B2 | 2008-11-18 | Arterial pressure-based, automatic determination of a cardiovascular parameter | 2004-07-14 | One or more cardiovascular parameters is estimated as a function of the arterial pressure waveform, in particular, using at least one statistical moment of pressure waveform having an order greater than one. Arterial compliance, the exponential pressure decay constant, vascular resistance, cardiac output, and stroke volume are examples of cardiovascular parameters that can be estimated using various aspects of the invention. In one embodiment of the invention, not only are the first four moments (mean, standard deviation, skewness, and kurtosis) of the pressure waveform used to estimate the cardiovascular parameter(s) of interest, but also heart rate, statistical moments of a set of pressure-weighted time values, and certain anthropometric patient measurements such as age, sex, body surface area, etc. | 1. A method for determining a cardiovascular parameter of a subject comprising: sensing an input signal that indicates arterial blood pressure;calculating at least one statistical moment of a first function of the input signal having an order of two or higher; andestimating the cardiovascular parameter as a second function of at least one statistical moment of a first function of the input signal having an order greater than one. 2. A method as in claim 1, in which the cardiovascular parameter is arterial compliance. 3. A method as in claim 1, in which the cardiovascular parameter is vascular resistance. 4. A method as in claim 1, in which the cardiovascular parameter is cardiac output. 5. A method as in claim 1, in which the cardiovascular parameter is stroke volume. 6. A method as in claim 1, in which the cardiovascular parameter is a pressure decay constant. 7. A method as in claim 1, in which the step of estimating the cardiovascular parameter comprises computing the statistical moment(s) of a sequence of measured arterial pressure values. 8. A method as in claim 7, further comprising estimating the cardiovascular parameter as a function of the kurtosis of the input signal. 9. A method as in claim 7, further comprising estimating the cardiovascular parameter as a function of the skewness of the input signal. 10. A method as in claim 7, further comprising: measuring a predetermined set of anthropometric parameters of the subject; andestimating the cardiovascular parameter as a function of the measured anthropometric parameters. 11. A method as in claim 7, in which the cardiovascular parameter is arterial compliance. 12. A method as in claim 7, in which the cardiovascular parameter is vascular resistance. 13. A method as in claim 7, further comprising: computing both the standard deviation and at least one statistical moment, having an order greater than two, of the input signal;estimating an arterial compliance value as a function of at least the statistical moment of the input signal having an order greater than two; andcomputing an estimate of stroke volume as a function of the product of the standard deviation and the arterial compliance value. 14. A method as in claim 13, further comprising: computing the kurtosis of the input signal;computing an estimate of stroke volume as a function not only of the product of the standard deviation and the arterial compliance value, but also of an offset factor that is proportional to the kurtosis. 15. A method as in claim 7, further comprising: measuring the subject's heart rate;computing both the mean and the standard deviation of the input signal;computing, as the cardiovascular parameter, a pressure decay constant a function that is proportional to the mean and inversely proportion to both the standard deviation and the heart rate. 16. A method as in claim 15, further comprising: computing an arterial compliance value; andcomputing, as the cardiovascular parameter, a vascular resistan... | Edwards Lifesciences Corporation | A61B005/02
US07856326 | B2 | 2010-12-21 | Methods and systems for providing a nutraceutical program specific to an individual animal | 2008-06-06 | The present invention relates to methods and systems for providing an animal with a nutraceutical program that is specifically tailored to address the deficiencies and/or needs of the particular animal. Blood is drawn from the animal and analyzed at a lab to obtain blood test results. The blood test results are scored to obtain at least one blood test score for at least one corresponding blood parameter. If the at least one blood test score falls within a normal range but outside of an optimal range, one or more nutraceuticals needed to bring the animal within the optimal range for the at least one corresponding blood parameter are identified. A prescribed dosage amount is then calculated for at least one of the one or more identified nutraceuticals for the animal. The calculation of the prescribed dosage amount is based on at least the blood test score for the corresponding blood parameter and a deviation of the blood test score from the optimal range. | 1. A system for providing a nutraceutical program specific to an individual animal, comprising: a blood analyzer for analyzing a blood sample of an animal to obtain blood test results;a database containing historical blood test scores for a plurality of animals, said historical blood test scores defining normal ranges for corresponding blood parameters and defining corresponding optimal ranges within said normal ranges for the corresponding blood parameters; anda processor in communication with said blood analyzer and said database, said processor adapted to: score the blood test results to obtain at least one blood test score for at least one corresponding blood parameter;determine if the at least one blood test score falls outside the normal range defined for the corresponding blood parameter;if said at least one blood test score falls within said normal range but outside of the optimal range: identify one or more nutraceuticals needed to bring the animal within said optimal range for the at least one corresponding blood parameter; andcalculate a prescribed dosage amount for at least one of said one or more identified nutraceuticals for the animal based on at least said blood test score for said corresponding blood parameter and a deviation of the blood test score from the optimal range;wherein if said at least one blood test score falls outside of said normal range, standard nutrients or medications are prescribed for the animal in order to bring the at least one blood test score within said normal range. 2. A system in accordance with claim 1, wherein said processor identifies said one or more nutraceuticals based on at least one of animal type, animal physiology, animal biochemistry, data describing a physical condition of the animal, and data describing a medical diagnosis of the animal. 3. A system in accordance with claim 1, wherein said processor calculates the prescribed dosage amount for a particular identified nutraceutical by: calculating a pre-cap dosage amount based on said blood test score, said deviation of the at least one blood test score from the optimal range, animal type, animal weight, dosage strength information, and dosage limit information; andcomparing said calculated pre-cap dosage amount to a low-cap dosage amount and a high-cap dosage amount for the particular animal;wherein:if the pre-cap dosage amount is between the low-cap and high-cap dosage amounts, the prescribed dosage amount is equal to the pre-cap dosage amount;if the pre-cap dosage amount is greater than the high-cap dosage amount, the prescribed dosage amount is equal to the hi-cap dosage amount; andif the pre-cap dosage amount is less than the low-cap dosage amount, the prescribed dosage amount is equal to the low-cap dosage amount. 4. A system in accordance with claim 3, wherein: the low-cap dosage amount is the minimum amount of an identified nutraceutical prescribed for an animal regardless of the animal weight or the deviation of the at least one b... | Bionutritional, LLC | G06F007/00
US08208982 | B2 | 2012-06-26 | Evanescent catheter system | 2005-12-01 | An evanescent catheter system having a tip portion of an optical fiber capable of securely measuring fluorescence intensity or fluorescence spectrum of a substance around a catheter tip, without being hindered by optical absorption of hemoglobin in erythrocytes in blood, and being capable of calculating an existing amount of a fluorescence-emitting substance existing in an evanescent light generating part. A protecting layer 22 and a cladding 23 are removed from an upper semicircular part of a core 24 at a tip of a catheter 21 to expose the core 24. When excitation light 29 is incident, evanescent light 25 is generated on the upper semicircular part of the core 24 and a fluorescent substance generates emits fluorescence 30. An interference filter 26 filters out the excitation light and only the fluorescence 30 reaches the photodiode 28, permitting a measuring device to measure an existing amount of the fluorescent substance. | 1. An evanescent catheter system adapted to be placed in blood comprising: an optical fiber obtained by removing a cladding layer to expose a core layer, to generate evanescent light, to detect fluorescence from a fluorescent substance excited by the evanescent light, and to measure an existing amount of the substance in the blood on the basis of a fluorescence intensity,wherein substantially a half of the cladding covering the core is removed at a columnar tip of the optical fiber whereby the evanescent light is generated on an upper columnar surface of the core, andan optical filter for transmitting the fluorescence in a predetermined wavelength region located on an inner lower columnar surface of the core exposed by removing the cladding layer, wherein the optical filter comprises an upper side and a lower side, wherein the upper side of the optical filter is more proximal to the exposed core than the lower side, and a light receiving element is located on the lower side of the optical filter, wherein the light receiving element detects the fluorescence from the fluorescent substance. 2. The catheter system of claim 1 wherein the light receiving element is a photodiode. 3. The catheter system of claim 1 further comprising a detector amplifier able to receive output from the light receiving element. 4. The catheter system of claim 3 further comprising a controller able to receive output from the detector amplifier and calculate fluorescence intensity using the output from the amplifier. 5. The catheter system of claim 4 wherein said controller is able to calculate the existing amount of the fluorescent substance using the calculated fluorescence intensity. 6. The catheter system of claim 1 wherein the optical filter is an interference filter. | National University Corporation Hamamatsu University School of Medicine | A61B005/1455
US07957791 | B2 | 2011-06-07 | Multi-beat integration for cardiac mapping | 2008-06-13 | A non-contact cardiac mapping method is disclosed that includes: (i) inserting a catheter into a heart cavity having an endocardium surface, the catheter including multiple, spatially distributed electrodes; (ii) measuring signals at the catheter electrodes in response to electrical activity in the heart cavity with the catheter spaced from the endocardium surface; and (iii) determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface. Related systems and computer programs are also disclosed. | 1. A method comprising: inserting a catheter into a heart cavity having an endocardium surface, the catheter comprising multiple, spatially distributed electrodes;measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;integrating information derived from the signals for common phases of the electrical heart beat cycles;determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface; anddisplaying at least a portion of the determined physiological information;wherein the signals are measured during multiple different portions of a heart beat cycle and wherein the physiological information is determined, at least in part, by accounting for changes in the location of the catheter electrodes with respect to the endocardium surface during the multiple different portions of a heart beat cycle. 2. The method of claim 1, further comprising: for each of the multiple different portions of a heart beat cycle, determining a position of the catheter electrodes with respect to the endocardium surface and measuring signals at the catheter electrodes in response to electrical activity in the heart cavity,wherein the determination of the physiological information at the multiple locations of the endocardium surface is based further on the positions of the catheter electrodes and the measured signals at the different catheter positions. 3. The method of claim 1, wherein accounting for changes in the location of the catheter electrodes with respect to the endocardium surface during the multiple different portions of a heart beat cycle comprises accounting for changes in the shape of the endocardium surface during the multiple different portions of a heart beat cycle. 4. The method of claim 1, wherein integrating information derived from the signals for common phases of the electrical heart beat cycles comprises separately integrating information derived from the signals for systole and end diastole phases of the heart beat cycle. 5. The method of claim 1, wherein the determined physiological information includes at least one of: electrical potential values at the multiple locations of endocardium surface at each of one or more different phases of the heart beat cycle,a maximum voltage amplitude for one or more heart beat cycles at different ones of the endocardium surface locations,a frequency representation of electrical activity at multiple locations of the endocardium surface during the heart beat cycle, wherein the frequency representation is computed based on electrical potential values at the multiple locations of endocardium surface at different phases of the heart beat cycle, andan activation time for each of different locations of the endocardium surface. 6. The method of claim 1, wherein displaying at least a portion of the determined physiological information comprises displaying se... | Rhythmin Medical, Inc. | A61B005/02
US08971992 | B2 | 2015-03-03 | Magnetic resonance imaging apparatus and control method thereof | 2008-05-05 | A magnetic resonance imaging apparatus includes a collection unit which applies a uniform static magnetic field to a subject and also applies a radio-frequency magnetic field and a gradient magnetic field to the subject in accordance with a predetermined pulse sequence to collect a magnetic resonance signal from the subject, a imaging unit which images the subject based on the magnetic resonance signal collected by the collection unit, a detection unit which detects a respiratory level of the subject, an informing unit which informs the subject of whether the detected respiratory level falls within an allowable range, and a unit which controls the collection unit and the imaging unit in such a manner that the magnetic resonance signal for imaging is collected and the subject is imaged based on the thus collected magnetic resonance signal for imaging when the detected respiratory level falls within the allowable range. | 1. A magnetic resonance imaging (MRI) apparatus comprising: MRI gantry components including static and gradient magnetic field generators, at least one radio frequency (RF) coil coupled to an imaging volume, at least one RF transmitter, at least one RF receiver and at least one control computer having a processor and memory configured to control said gantry components so as to effect operations including:applying a uniform static magnetic field to a subject;applying a radio-frequency magnetic field and a gradient magnetic field to the subject in accordance with a predetermined pulse sequence and collecting from the subject within one heart beat period of the subject (a) magnetic resonance signals for reconstruction acquired using phase encoding magnetic gradient fields, and (b) a magnetic resonance signal for monitoring and at least one magnetic resonance signal for display both acquired without applying phase encoding gradient magnetic fields;detecting a first respiratory level for the subject using the magnetic resonance signal for monitoring;detecting a second respiratory level for the subject using the at least one magnetic resonance signal for display;generating a display image of the detected second respiratory level, an upper threshold of an allowable range and a lower threshold of the allowable range;causing a display system to display the generated display image to the subject; andreconstructing an image of the subject using the collected magnetic resonance signals for reconstruction when the detected first respiratory level falls within the allowable range. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the at least one control computer being further configured to effect: collecting the magnetic resonance signals for monitoring and display for an excitation slice or an excitation slab disposed in a target part that is present in the subject; andacquiring positional information of the target part based on the collected magnetic resonance signals for monitoring and display; anddetecting the respiratory levels based on the acquired positional information. 3. The magnetic resonance imaging apparatus according to claim 2, wherein the at least one control computer is configured to determine a diaphragm as the target part. 4. The magnetic resonance imaging apparatus according to claim 1, wherein the at least one control computer is further configured to effect collecting the magnetic resonance signals for monitoring and display without applying the phase encoding gradient magnetic field immediately before collecting the magnetic resonance signals for reconstruction with application of the phase encoding gradient magnetic field. 5. The magnetic resonance imaging apparatus according to claim 1, wherein the at least one control computer is further configured to effect detecting the second respiratory level at each of plural predetermined timings, and generating the display image showing a relationship between the la... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | A61B005/05 A61B005/055 G01R033/563 G01R033/567 A61B005/00 G01R033/28 G01R033/565 A61B006/00
US08617069 | B2 | 2013-12-31 | Health monitor | 2008-06-20 | Methods and devices to detect analyte in body fluid are provided. Embodiments include enhanced analyte monitoring devices and systems. | 1. An analyte monitoring system, comprising: an analyte sensor for transcutaneous positioning under a skin layer;a data processing device operatively coupled to the analyte sensor, the data processing device comprising: a control unit;a memory operatively coupled to the control unit and configured to store a plurality of data associated with a monitored analyte level received from the analyte sensor; anda communication unit operatively coupled to the control unit; anda receiver device configured for signal communication with the data processing device;wherein when the control unit of the data processing device detects a communication link with the receiver device, the control unit is further configured to retrieve the stored plurality of data from the memory and to communicate the retrieved data to the receiver device only upon receipt of a request from the receiver device when the receiver device is placed within a predetermined distance from the data processing device;wherein the retrieved stored plurality of data correspond to analyte data collected over a predetermined time period; and furtherwherein the analyte sensor is factory calibrated. 2. The system of claim 1 wherein the data processing device does not include a user output component. 3. The system of claim 1 wherein the communication unit is not configured to support real time continuous data communication. 4. The system of claim 1 wherein the predetermined time period includes one of about several weeks or at least one month. 5. The system of claim 1 wherein the communication unit is configured to communicate with the receiver device using one or more of a wired connection, a USB cable connection, a serial cable connection, an RF communication protocol, an infrared communication protocol, a short-range communication protocol, or an 802.11x communication protocol. 6. The system of claim 1 wherein the control unit detects the communication link with the receiver device based on detection of a wired connection to the receiver device. 7. The system of claim 1 wherein the analyte sensor does not require user calibration. 8. The system of claim 1 wherein the analyte sensor is a glucose sensor. 9. The system of claim 1 wherein the receiver device is configured to receive the retrieved data transmitted by the control unit, and wherein the receiver device includes an output unit configured to output one or more of the received data. 10. The system of claim 9 wherein the output unit includes a display unit operatively coupled to a housing of the receiver device. 11. The system of claim 9 wherein the output of one or more of the received data includes a graphical output, a numerical output, or a text output. 12. The system of claim 9 wherein the receiver device includes a storage unit configured to store the received data. 13. The system of claim 1 including a holster device for receiving the receiver device. 14. The system of claim 13 wherein the data processing unit is integrated... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/01 A61B005/04 A61B005/05
US07986823 | B2 | 2011-07-26 | System and method for consistent detection of mid-sagittal planes for magnetic resonance brain scans | 2008-05-07 | A method for finding consistent mid-sagittal planes (MSPs) in a pair of 3D image head scans of a same patient includes, for each image, selecting a 2D transverse localizer image from a middle slice along a transverse view of the 3D image, fitting an ellipse to the transverse localizer image to locate a head position (x, y), where the y-coordinate is indicative of a position of a 2D coronal localizer image in the 3D image, fitting an ellipse to the coronal localizer image to locate a head position (x, y) in the coronal localizer image, calculating a middle line in the transverse localizer image and a middle line in the coronal localizer image, calculating an MSP from the middle lines, and determining a new set of slope and intercept parameters for one MSP that maximizes a similarity measure between the one MSP and the other MSP. | 1. A method of finding consistent mid-sagittal planes (MSPs) in different image head scans of a same patient, said method comprising the steps of: providing a pair of images acquired from different head scans of a same patient, each said image comprising a plurality of intensities associated with a 3-dimensional (3D) grid of points;for each image, selecting a 2-dimensional (2D) transverse localizer image from a middle slice along a transverse view of the 3D image;fitting an ellipse to said transverse localizer image to locate a head position (x, y) in said transverse localizer image, wherein said y-coordinate, a position along a vertical axis, is indicative of a position of a 2D coronal localizer image in the 3D image;fitting an ellipse to said coronal localizer image to locate a head position (x, y) in said coronal localizer image, wherein said x-coordinate, a position along a horizontal axis, is indicative of a position of said 2D coronal localizer image in the 3D image;calculating a middle line in said transverse localizer image and a middle line in said coronal localizer image;calculating a mid-sagittal plane (MSP) from the middle lines of said transverse localizer image and said coronal localizer image; anddetermining a new set of slope and intercept parameters within a proximity of one MSP that maximizes a similarity measure between the one MSP and the other MSP. 2. The method of claim 1, wherein if the middle lines of said transverse localizer image and said coronal localizer image intersect each other, calculating said mid-sagittal plane from the plane defined by the intersection of said middle lines. 3. The method of claim 1, wherein if the middle lines of said transverse localizer image and said coronal localizer image do not intersect each other, calculating said mid-sagittal plane further comprises finding points p1, p2 on each respective middle lines that are closest to each other, finding a mid-pint pm on a line segment connecting points p1, p2, and defining said MSP as a plane passing through point pm that whose normal is determined by a cross product of vectors corresponding to said middle lines. 4. The method of claim 1, further comprising, for each 3D image, updating a column resolution and row resolution of said corresponding MSP with information from each selected transverse and coronal localizer image. 5. The method of claim 1, wherein said slope and intercept parameters are the slope and intercept of the middle lines of the transverse localizer image and the coronal localizer image associated with the MSP. 6. The method of claim 1, wherein said similarity measure is a mutual information measure proportional to-∑x⁢∑y⁢p⁡(x,y)⁢log⁡(p⁡(x,y)p⁡(x)⁢p⁡(y)),wherein x, y represent respectively intensity levels of corresponding points in each of the two MSPs, p(x) is a probability of intensity level x, and p(x, y) is a joint probability of intensity levels x, y. 7. The method of claim 1, wherein said similarity measure is maximiz... | Siemens Aktiengesellschaft | G06K009/00
US08157732 | B2 | 2012-04-17 | Method and apparatus for measuring autonomic-nervous index and apparatus for detecting biological information | 2008-02-25 | A detection-rate calculating unit calculates a detection rate representing a ratio of number of interval data representing a time interval of one cycle of a waveform of at least one of a pulse and a heart rate, generated within a predetermined reference time to a sum of the number of the interval data and number of error data representing a content of an error occurred while the interval data is generated. When the detection rate is larger than the first threshold, an index calculating unit calculates an autonomic-nervous index representing an autonomic-nervous activity state, based on the interval data generated within the reference time. | 1. An apparatus for measuring an autonomic-nervous index, comprising: a detection-rate calculating unit that calculates a detection rate representing a ratio of number of interval data generated within a predetermined reference time to a sum of the number of the interval data and number of error data generated within the reference time, the interval data representing a time interval of one cycle of a waveform of at least one of a pulse and a heart rate, the error data representing a content of an error that occurred within the reference time; andan index calculating unit that calculates an autonomic-nervous index representing an autonomic-nervous activity state, based on the interval data generated within the reference time, when the detection rate is larger than a first threshold. 2. The apparatus according to claim 1, further comprising a receiving unit that receives the interval data and the error data from a detecting device, the detecting device detecting the waveform of at least one of the pulse and the heart rate of a user and generating the interval data and the error data based on the waveform of the at least one of the pulse and the heart rate, wherein the detection-rate calculating unit calculates the number of the interval data and the number of the error data received within the reference time and calculates a ratio of calculated number of the interval data to a sum of the calculated number of the interval data and calculated number of the error data as the detection rate. 3. The apparatus according to claim 1, further comprising: a receiving unit that receives the waveform of at least one of the pulse and the heart rate of a user detected by a detecting device, the detecting device detecting the waveform of at least one of the pulse and the heart rate of the user; andan analyzing unit that analyzes the waveform of the at least one of the pulse and the heart rate received by the receiving unit for generating the interval data and the error data. 4. The apparatus according to claim 1, wherein the error data includes an out-of-range error representing either that the interval data is smaller than a predetermined second threshold or that the interval data is larger than a predetermined third threshold that is larger than the second threshold. 5. The apparatus according to claim 1, wherein the error data includes a body-movement detected error representing that a body movement of a user is detected when the interval data is generated. 6. The apparatus according to claim 1, wherein the index calculating unit further compares number of body movements of a user detected within the reference time with a predetermined fourth threshold, and calculates the autonomic-nervous index when the number of the body movements is smaller than the fourth threshold. 7. The apparatus according to claim 1, wherein the index calculating unit further compares an average value of an amplitude of the waveform of the interval data in the reference time with... | Kabushiki Kaisha Toshiba | A61B005/00 A61B005/0402
US08249685 | B2 | 2012-08-21 | Method and apparatus for mapping and/or ablation of cardiac tissue | 2005-05-17 | An apparatus for mapping and/or ablating tissue includes a braided conductive member that may be inverted to provide a ring-shaped surface. When a distal tip of the braided conductive member is retracted within the braided conductive member, the lack of protrusion allows the ring-shaped surface to contact a tissue wall such as a cardiac wall. In an undeployed configuration, the braided conductive member is longitudinally extended, and in a deployed configuration, the distal end of the braided conductive member is retracted to invert the braided conductive member. | 1. A medical device for electrophysiology procedures, comprising: a catheter shaft having a proximal end and a distal end; anda braided conductive member located at the catheter shaft distal end;the braided conductive member having an undeployed configuration in which the braided conductive member has a distal end and is in a longitudinally extended configuration;the braided conductive member further having a deployed configuration in which the braided conductive member distal end is positioned proximally relative to portions of the braided conductive member; whereinin the deployed configuration, the braided conductive member forms a distally-facing ring-shaped surface and includes ablation electrodes configured to form an annular lesion in tissue when the distally-facing surface is contacted to the tissue and the ablation electrodes are activated; andfurther comprising a braided conductive member adjustment element constructed and arranged to move the braided conductive member distal end, wherein the braided conductive member adjustment element is a cable attached to the braided conductive member distal end, the cable being constructed and arranged to pull the braided conductive member distal end in a proximal direction. 2. A medical device according to claim 1, wherein the braided conductive member comprises electrically conductive filaments. 3. A medical device according to claim 1, wherein the braided conductive member comprises mapping filaments. 4. A medical device according to claim 1, wherein the ablation electrodes comprise ablation filaments. 5. A medical device according to claim 4, wherein the braided conductive member further comprises mapping filaments. 6. A medical device as in claim 1, wherein, in the deployed configuration, nothing protrudes distally beyond the distally-facing surface. 7. A medical device as in claim 1, wherein, in the deployed configuration, the distally-facing ring-shaped surface is substantially perpendicular to a longitudinal direction of the braided conductive member. 8. A medical device as in claim 1, further comprising a sheath that is advanceable over and retractable from the braided conductive member. 9. A medical device as in claim 1, wherein the distal end of the catheter shaft is steerable. 10. A medical device as in claim 1, further comprising support elements within a proximal portion of the braided conductive member. 11. A medical device as in claim 1, wherein the braided conductive member is radially asymmetrically-shaped relative to a longitudinal axis of the braided conductive member. 12. A medical device as in claim 1, wherein the braided conductive member is longitudinally asymmetrically-shaped relative to a longitudinal axis of the braided conductive member. 13. A medical device as in claim 1, further comprising a controller operatively connected to the braided conductive member. 14. A medical device as in claim 1, further comprising an irrigation system. 15. A medical device ac... | C.R. Bard, Inc. | A61B005/0402 A61B018/14
US08989842 | B2 | 2015-03-24 | System and method to register a tracking system with intracardiac echocardiography (ICE) imaging system | 2007-09-25 | A system to navigate an object traveling in an imaged subject is provided. The system includes a four-dimensional (4D) imaging system to create a first three-dimensional model of the acquired imaged data correlated relative to a time reference and defined relative to a first image coordinate system. A second imaging system creates a second three-dimensional model of the imaged subject defined relative to a second image coordinate system. A tracking system tracks movement and orientation of the object relative to a tracking coordinate system. A controller includes program instructions in combination with the processor operable to register the first and second image coordinate systems and the tracking coordinate system relative to a world coordinate system, and to combine the three-dimensional model created by the 4D imaging system with the three-dimensional model created by the second imaging system with a tracked location of the object relative to the world coordinate system. | 1. A system to navigate an image-guided object comprising a tracking element and a transducer array configured to travel in an area of interest of an imaged subject in relation to an acquired image of the imaged subject, comprising: a four-dimensional (4D) imaging system comprising the image-guided object and configured to acquire first image data, the 4D imaging system further configured to use the first image data to create a 4D model representative of a time-dependent three-dimensional spatial model correlated to a time interval and defined in spatial relation and orientation relative to a first image coordinate system defined by a predetermined spatial relation and orientation of the first image data relative to the transducer array, wherein the time interval is one of a cardiac cycle or a respiratory cycle;a tracking system comprising sensors, the tracking system configured to track movement and orientation of the image-guided object through the imaged subject relative to a tracking coordinate system defined by the spatial relations and orientations of the sensors to the tracking element and to time; anda second imaging system different than the image-guided object, the second imaging system configured to acquire second image data, the second imaging system configured to use the second image data to create a three-dimensional (3D) model of the imaged subject correlated to a second image coordinate system defined by a spatial relation and orientation of the second image data correlated to a phase of the cardiac cycle or a phase of the respiratory cycle;a controller electrically connected in communication with the 4D imaging system, the second imaging system, and the tracking system, the controller having a processor configured to: correlate the 4D model to a world coordinate system, the step of correlating further comprising: performing a first registration of first image coordinate system with the tracking coordinate system; andperforming a second registration of the tracking coordinate system with a world coordinate system, wherein the world coordinate system defines a spatial orientation of image pixels or voxels within a display;register the 4D model correlated with the world coordinate system to the 3D model, the step of registering further comprising: allowing a user to select a series of anatomical landmarks within the 3D model displayed on a graphical user interface (GUI); andallowing the user to select a series of anatomical landmarks within a portion of the 4D model acquired in the phase of cardiac or respiratory cycle that corresponds to the phase of cardiac or respiratory cycle of the 3D model based on a time reference, the GUI configured to allow the user to select the series of anatomical landmarks within the 4D model; andcombine the first image data of the 4D model and the second image data of the 3D model according to the series of anatomical landmarks selected within the 3D and 4D models. 2. The system of claim 1, wherein th... | General Electric Company | A61B005/05 A61B008/12 A61B005/06 A61B008/08 A61B008/00
US08734355 | B2 | 2014-05-27 | Motion management in a fast blood pressure measurement device | 2008-07-22 | The invention is in one embodiment an apparatus and a method for measuring the blood pressure of a vertebrate subject. The apparatus uses an inflatable chamber with a sensor to detect signals indicative of a blood pressure of the subject during an inflation interval of the inflatable chamber. If secondary motion or artifact signals are detected, the apparatus and method determine whether the secondary signals are below a predetermined value, and if so, complete the measurement. If the signals indicate that the measurement is not accurate, the system and method immediately institute a measurement of the blood pressure during a deflation interval, which can include a step deflation interval. In the event that the subject is identified as a neonate, the system and method immediately measure using the deflation interval procedure. | 1. A blood pressure measurement apparatus configured to operate according to an oscillometric method of measuring blood pressure, said apparatus comprising: an inflatable chamber configured to be inflated during an inflation interval and deflated during a deflation interval;a sensor coupled to the inflatable chamber, the sensor configured to measure a signal according to the oscillometric method of measuring blood pressure, the signal comprising information indicative of a blood pressure of a vertebrate;a control module configured to receive as input at least a portion of the signal from the sensor, and configured to generate as output a control signal having a selected one of a plurality of values responsive to the input;a first analysis module configured to analyze the signal before the inflatable chamber is fully inflated to extract from the signal a systolic blood pressure and a diastolic blood pressure of the vertebrate according to the oscillometric method of measuring blood pressure;a second analysis module selectively operative in response to one of the plurality of values of the control signal, the second analysis module configured to analyze the signal during the deflation interval of the inflatable chamber to extract from the signal the blood pressure of the vertebrate according to the oscillometric method of measuring blood pressure; anda neonate sensor module configured to represent whether the vertebrate is a neonate vertebrate by determining whether the vertebrate is 28 days or less in terms of age,wherein in response to a positive determination that the vertebrate is a neonate vertebrate, the apparatus completes a blood pressure measurement of the neonate vertebrate using the second analysis module,whereby the apparatus completes the measurement of the blood pressure of the vertebrate using at least one of the first analysis module and the second analysis module. 2. The apparatus of claim 1, wherein the first analysis module configured to analyze the signal during the inflation interval of the inflatable chamber at a rate greater than 3 mmHg per second. 3. The apparatus of claim 1, wherein the first analysis module is configured to determine if the systolic blood pressure and the diastolic blood pressure of the vertebrate are obtained with a fast inflation rate, the inflatable chamber being deflated upon a successful determination by the first analysis module. 4. The apparatus of claim 1, wherein the deflation interval comprises at least one step deflation interval. 5. The apparatus of claim 1, wherein the plurality of control signal values comprises a first value that inhibits operation of the second analysis module and a second value that activates operation of the second analysis module. 6. The apparatus of claim 1, further comprising a reporting module configured to report at least one of the systolic blood pressure and the diastolic blood pressure. 7. The apparatus of claim 1, further comprising: a motion detection modu... | Welch Allyn, Inc. | A61B005/08
US08801609 | B2 | 2014-08-12 | Apparatus and system for predictive health monitoring | 2005-03-24 | An apparatus and system are provided for monitoring an individual, a group or a community of individuals for respiratory infections. Predictive health data, such as waking peak expiratory flow rate (“WPF”) and basal metabolic temperature (“BMT”) are measured and charted to provide an indication of early stage respiratory infection. Data may be collected for target populations and reported to health agencies for use in combating the spread of infection. An apparatus is provided which can simultaneously measure WPF and BMT and other health-condition-related values and automatically transmit readings to a data collection and analysis program on a remote computer. | 1. An apparatus for simultaneously measuring basal metabolic temperature, waking peak flow, plus one or more of heart rate, blood oxygen percentage and saliva acidity; the apparatus comprising: a fluid channel for passing exhaled fluid and having an impellor disposed therein for measuring peak volumetric flow of fluid through the fluid channel by sensing a rotating magnet within the impellor with a field-effect transistor counting the numbers of rotations per second of the impellor;a mouthpiece positioned in line with the fluid channel for insertion into an individual's mouth and forming a seal with the mouth and directing exhaled fluid into the fluid channel;a rapid thermometer probe with a metallic tip disposed in the mouthpiece for measuring basal metabolic temperature at a sublingual location when the mouthpiece is inserted into the mouth;a second probe with a metallic tip disposed in the mouthpiece for measuring saliva acidity at a sublingual location when the mouthpiece is inserted into the mouth; and,a digital memory for storing measured values for basal metabolic temperature, waking peak flow, and saliva acidity (Ph). 2. The apparatus of claim 1 further comprising a connector for transmitting the measured values to a remote location. 3. The apparatus of claim 1 further comprising a wireless means for transmitting the measured values to a remote location. 4. The apparatus of claim 1 wherein said fluid channel has at least one continuous wall extending from the mouthpiece to an exit. 5. A measuring apparatus comprising: a mouthpiece configured to be placed in a user's mouth and forming a fluid tight seal therewith during exhalation to direct exhaled fluid into the apparatus;a fluid channel extending in a plane defined by a straight path from the mouthpiece to an exit;an impellor disposed in the fluid channel;a rotating magnet coupled to the impellor;a field-effect transistor sensing the rotating magnet and counting the number of rotations per second of the impellor;a rapid thermometer probe with a metallic tip disposed in the mouthpiece and measuring basal metabolic temperature in the user's mouth; and,a second probe with a metallic tip disposed in the mouthpiece and measuring saliva acidity together with the first probe in the user's mouth;both probes being located in the mouthpiece to accurately position them in opposing lateral sublingual positions in the user's mouth. 6. The apparatus of claim 5 further comprising a memory storing measured values of basal metabolic temperature, peak flow and saliva acidity. 7. The apparatus of claim 5 further comprising an electrical connector. 8. The apparatus of claim 5 further comprising a sensing module having a light source, the sensing module being disposed on the apparatus to accommodate one of the user's fingers. 9. The apparatus of claim 8 wherein the sensing module measures heart rate. 10. The apparatus of claim 8 wherein the sensing module measures percentage of oxygen saturation in ... | Predictive, Inc. | A61B005/00 A61B005/09 A61B010/00 A61B005/01
US09402558 | B2 | 2016-08-02 | System and method for pain detection and computation of a pain quantification index | 2007-04-05 | A method of detecting pain in a subject, comprises the steps of generating brain wave data based on brain wave activity of the subject and comparing the brain wave data to reference data to generate result data, the reference data corresponding to at least one of (i) population normative data indicative of brain wave activity of a first plurality of individuals in an absence of pain, (ii) population reference data indicative of brain wave activity of a second plurality of individuals generated in response to pain events inflicted on the second plurality of individuals, (iii) subjective population reference data indicative of brain wave activity of a third plurality of individuals reporting a sensation of pain, and (iv) population of reference data indicative of brain wave activity of a fourth population of individuals following an intervention which has changed a subjective report of pain in combination with determining a presence of pain experienced by the subject as a function of the result data. | 1. A method of detecting pain in a subject, comprising the steps of: comparing, using a processor, brain wave data generated based on brain wave activity of the subject to at least one source of reference data to generate comparison data, the reference data corresponding to at least one of (i) population normative data indicative of brain wave activity of a first plurality of individuals in an absence of pain, (ii) population reference data indicative of brain wave activity of a second plurality of individuals generated in response to pain events inflicted on the second plurality of individuals, (iii) subjective population reference data indicative of brain wave activity of a third plurality of individuals reporting a sensation of pain, and (iv) population of reference data indicative of brain wave activity of a fourth population of individuals following an intervention which has changed a subjective report of pain;inputting, using the processor, the comparison data into at least one of a first classifier function A indicating a probability score that pain is being sensed and a second classifier function B indicating a probability score that pain is not being sensed;determining, using the processor, a presence of pain experienced by the subject as a function of the at least one probability score, wherein the pain is not elicited by a sensory stimulus, the determining step comprising comparing the probability score reflecting the presence of pain to the probability score reflecting that pain is not being sensed to calibrate a pain quantification index including values corresponding to one of a plurality of levels of pain and a plurality of types of pain, wherein the pain is not elicited by a sensory stimulus, and the types of pain include chronic and acute;quantifying, using the processor, a level of pain experienced by the subject as a function of the comparison data; andregulating an amount of analgesic administered to the subject by a patient controlled anesthetic apparatus as a function of the comparison data and the level of pain. 2. The method according to claim 1, further comprising: computing, using the processor, the probability score P that the subject is experiencing pain by comparing a value based on eA to a value based on 1−eB; andcomputing, using the processor, one of a predicted intensity and a predicted type of the pain using a logistic regression when the probability is at least a threshold level. 3. The method according to claim 2, wherein the threshold level is at least 0.95. 4. The method according to claim 2, wherein the threshold level is at least 0.99. 5. The method according to claim 1, wherein the brain wave data is analyzed using a neural network to compute the probability score that pain is being sensed and one of a predicted intensity of sensed pain and a predicted type of sensed pain. 6. The method according to claim 1, wherein one of the brain wave data and the comparison data is presented to the subject as a “ta... | NEW YORK UNIVERSITY | A61B005/0476 A61B005/00
US09406134 | B2 | 2016-08-02 | Image system for supporting the navigation of interventional tools | 2008-03-20 | The invention relates to an image system and a method for supporting the navigation of interventional tools when performing imaging controlled minimally invasive interventions within the body of a patient in a predetermined intervention plane, which serves to improve the precision and reliability of interventional accesses that are required for performing a histological tissue sample removal under CT or MRT based image-oriented monitoring, or in the context of tumor therapy or pain therapy. Directional deviations, away from the intervention plane, of the actual course of an interventional tool from a predeterminable desired course are captured and presented for display by registering the shortened represented total length or a shortened represented partial length of this interventional tool in the 2D projection representation of a fluoroscopic radioscopy recording that was recorded in a 2D projection direction running normally relative to the intervention plane. | 1. An image system for supporting a navigation of an interventional tool when performing an imaging controlled invasive intervention within a body of a patient in an intervention plane, comprising: a recording system that records a 2D projection recording representing a current position of the interventional tool from two different projection planes running normally to the intervention plane;an imaging processing system that processes the 2D projection recording; andan input/output interface that receives the recorded 2D projection recording from the recording system and supplies the recorded 2D projection recording to the imaging processing system,wherein the imaging processing system is configured to: register a shortened length of the interventional tool represented in the 2D projection recording;capture a directional deviation of an actual course of the interventional tool from a predetermined desired course that lies within the intervention plane from the registration;generate a 3D data record of a target zone of the patient;define a spatial coordinate of a puncture target in the 3D data record with respect to a 3D Cartesian coordinate system;define a spatial coordinate of an insertion point in the 3D Cartesian coordinate system;identify a position of a tip of the interventional tool in the 2D projection recording;calculate a location coordinate of the position in the 3D Cartesian coordinate system based on the identified position in the 2D projection recording;select a further spatial point of the intervention tool in the 3D Cartesian coordinate system;estimate a position of the further spatial point of the interventional tool in the 2D projection recording;calculate a 3D location coordinate of the 2D estimated position of the further spatial point in the 3D Cartesian coordinate system;adjust the further spatial point until a deviation between a direction of a line connecting the identified position of the tip and the estimated position of the further spatial point in the 2D projection recording and a projected direction of the predetermined desired course in the 2D projection recording being within a predetermined deviation value; andconvert the directional deviation of the interventional tool into an actuating variable for triggering an actuator system to compensate the directional deviation by a reverse control,wherein the intervention plane is defined by the position of the tip of the interventional tool, the further spatial point of the interventional tool, and the puncture target when the intervention tool is properly aligned with a desired path of the intervention tool, andwherein the image system further comprises a robot that guides the interventional tool introduced from the insertion point to the puncture target based on the actuating variable. 2. The image system as claimed in claim 1, wherein the interventional tool is a puncture needle to remove a histological tissue sample or to perform a tumor or pain therapy. 3. The image... | SIEMENS HEALTHCARE GMBH | A61B006/12 G06T007/00 A61B005/055 A61B006/03 A61B010/02
US20080242967 | A1 | 2008-10-02 | Medical imaging and therapy utilizing a scanned beam system operating at multiple wavelengths | 2007-03-27 | An apparatus, for medically treating a patient, includes a scanned beam system including a radiation beam source assembly, a scanner, and a controller. The assembly is adapted to emit a radiation beam (such as, without limitation, a laser beam) at different wavelengths. The controller is operatively connected to the assembly and the scanner. The controller is adapted: to display an image of an area of the patient obtained by the scanned beam system; to receive from a user viewing the image an identification of an in-treatment region; and to control the assembly and the scanner to medically treat the in-treatment region using a therapeutic wavelength of the radiation beam which provides an increased scanned-beam-system-image contrast between the in-treatment region and an out-of-treatment region over another wavelength of the radiation beam available for therapy. A method for medically treating a patient using a scanned beam system is also disclosed. | 1. An apparatus for medically treating a patient comprising a scanned beam system, wherein the scanned beam system includes a radiation beam source assembly, a scanner, and a controller, wherein the radiation beam source assembly is adapted to emit a radiation beam at different wavelengths, wherein the controller is operatively connected to the radiation beam source assembly and the scanner, and wherein the controller is adapted: to display an image of an area of the patient obtained by the scanned beam system;to receive from a user viewing the image an identification of an in-treatment region; andto control the radiation beam source assembly and the scanner to medically treat the in-treatment region using a therapeutic wavelength of the radiation beam which provides an increased scanned-beam-system-image contrast between the in-treatment region and an out-of-treatment region over another wavelength of the radiation beam available for therapy. 2. The apparatus of claim 1, wherein the therapeutic wavelength of the radiation beam substantially maximizes the contrast between the in-treatment region and the out-of-treatment region over any other wavelength of the radiation beam available for therapy. 3. The apparatus of claim 2, wherein the controller is adapted to determine the maximum contrast based on a survey of the area with the scanned beam system using the radiation beam source assembly at different wavelengths. 4. The apparatus of claim 1, wherein the controller is adapted to receive a therapeutic wavelength selection from the user based on the user viewing images of the area obtained by the scanned beam system at different wavelengths. 5. The apparatus of claim 1, wherein the in-treatment region is a calculus of the patient. 6. The apparatus of claim 1, wherein the in-treatment region is tissue of the patient. 7. A method for medically treating a patient comprising the following steps: a) viewing an image of an area of the patient obtained by a scanned beam system;b) identifying an in-treatment region from viewing the image;c) selecting a therapeutic wavelength of the scanned beam system which provides an increased scanned-beam-system-image contrast between the in-treatment region and an out-of-treatment region over another wavelength of the radiation beam available for therapy; andd) medically treating the in-treatment region with the scanned beam system using the therapeutic wavelength. 8. The method of claim 7, wherein step b) includes identifying the in-treatment region from viewing the image which provides the greatest contrast between the in-treatment region and the out-of-treatment region from a plurality of viewed images obtained at different imaging wavelengths and combinations thereof. 9. The method of claim 7, wherein step c) includes selecting a therapeutic wavelength of the scanned beam system which substantially maximizes the contrast between the in-treatment region and the out-of-treatment region over any other wavelen... | Ethicon Endo-Surgery, Inc. | A61B005/00
US20080306353 | A1 | 2008-12-11 | CALCULATION DEVICE FOR METABOLIC CONTROL OF CRITICALLY ILL AND/OR DIABETIC PATIENTS | 2007-10-09 | A method of providing blood glucose therapy for a critically ill patient includes calculating a baseline nutrition feed requirement based on an algorithm that incorporates at least one of age, gender, and body size of the patient: determining a first blood glucose level; determining a second blood glucose level after a preselected time interval: determining a first body temperature reading: comparing the blood glucose levels: and administering either nutrition or insulin. The amount of nutrition administered to the patient is based on a first change in blood glucose level, the current body temperature reading, and a predetermined feed algorithm based on the second blood glucose level as well as the baseline nutritional feed requirement. The amount of insulin administered is based on a second change in blood glucose level, body temperature, and a predetermined insulin algorithm that incorporates at least one of the patient's body frame size, age, and gender. | 1. A method of providing a blood glucose therapy for a critically ill patient, said method comprising the steps of: calculating a baseline nutritional feed requirement for said patient based on an algorithm that incorporates at least a weight of said patient;determining a first body temperature of said patient;determining a first blood glucose level of said patient;determining at least a second body temperature of said patient after a preselected time period;determining at least a second blood glucose level of said patient after a preselected time period; andcomparing said at least second blood glucose level to said first blood glucose level;administering one of an amount of nutritional feed to said patient and an amount of insulin to said patient;wherein said amount of nutritional feed administered to said patient is based on a first decrease in blood glucose level and a predetermined feed algorithm that incorporates said second blood glucose level and said patient's baseline nutritional feed requirement; andwherein said amount of insulin administered to said patient is based on a second decrease in blood glucose level that is less than said first decrease in blood glucose level and a predetermined insulin algorithm that incorporates said at least second blood glucose level and said patient's weight. 2. The method of claim 1, wherein said step of administering one of an amount of nutritional feed to said patient and an amount of insulin to said patient comprises maintaining said blood glucose level of said patient in the range of 4.0 mmol/L to 7.75 mmol/L. 3. The method of claim 1, wherein said preselected time period is 2 hours or less. 4. The method of claim 1, wherein said method is used to reduce the occurrence of ventilator-induced pneumonia by controlling a patient's blood glucose level. 5. The method of claim 1, wherein said patient was not diagnosed with diabetes prior to becoming critically ill and in which said patient is not otherwise considered to be either a type I or type II diabetic. 6. The method of claim 1, wherein said nutritional feed and said insulin administered to said patient is determined using a linear slide rule calculator or a circular slide rule calculator. 7. The method of claim 1, wherein said nutritional feed and said insulin administered to said patient is determined using an electronic computing device. 8. The method of claim 1, wherein blood glucose levels of said patient are determined repetitively. 9. The method of claim 1, wherein said predetermined feed algorithm incorporates said patient's one or more metabolic markers selected from the group of body temperature, renal function, urine output, catecholamine dosage, blood pressure and one or more patient status variables selected from the group consisting of age, weight and height, gender, and body frame size. 10. The method of claim 1, wherein said step of administering one of an amount of nutritional feed to said patient and an amount of insulin to... | - | A61B005/145 A61B005/01 A61B005/02
US07468032 | B2 | 2008-12-23 | Advanced patient management for identifying, displaying and assisting with correlating health-related data | 2002-12-18 | Systems, devices and methods for identifying, displaying and assisting in correlating health-related data. One aspect is a programmable device having machine executable instructions for performing a method for use in correlating data to manage a patient's health. In various embodiments, data related to the patient's health is acquired. The acquired data includes one or more trended health-related parameters, and at least one of one or more predetermined health-related events and one or more health-related alerts. Correlations are identified. The one or more trended health-related parameters and the at least one of the one or more predetermined health-related events and the one or more health-related alerts establish a causal relationship between the acquired data and the patient's health. | 1. A programmable device comprising a computer-readable medium for storing a computer-executable process for use in correlating data to manage a patient's health, the process comprising: acquiring data related to the patient's health, including: acquiring at least one parameter data series for one or more trended health-related parameters, wherein acquiring at least one parameter data series for one or more trended health-related parameters includes acquiring a parameter data series for use in determining atrial fibrillation percentage changes and a parameter data series for use in determining activity log changes; andacquiring at least one alert data series for one or more health-related alerts, wherein each of the one or more health-related alerts represents an alert condition for directing an alert to a patient or a clinician when a predetermined alert definition is satisfied;determining a correlation among the data, including determining a correlation coefficient for the at least one parameter data series and the at least one alert data series, wherein determining a correlation among the data includes automatically identifying a causal relationship with respect to the patient's health, wherein automatically identifying the causal relationship with respect to the patient's health and displaying the causal relationship that has been automatically identified includes automatically searching for a correlation between atrial fibrillation percentage and activity log changes;performing an action using the determined correlation coefficient, wherein the action is useful in managing the patient's health; anddisplaying the causal relationship that has been automatically identified. 2. The device of claim 1, wherein determining a correlation among the data includes displaying the acquired data in a manner useful to visually identify a correlative relationship. 3. The device of claim 2, wherein displaying the acquired data in a manner useful to visually identify the correlative relationship includes displaying the acquired data using a portable device. 4. The device of claim 2, wherein displaying the acquired data in a manner useful to visually identify the correlative relationship includes displaying the acquired data using a wellness monitoring device (WMD). 5. The device of claim 2, wherein displaying the acquired data in a manner useful to visually identify the correlative relationship includes displaying the acquired data using a programmer device. 6. The device of claim 2, wherein displaying the acquired data in a manner useful to visually identify the correlative relationship includes displaying a cursor to visually correlate the acquired data. 7. The device of claim 1, wherein acquiring at least one parameter data series for one or more trended parameters includes acquiring at least one mental/emotional parameter. 8. The device of claim 7, wherein acquiring at least one mental/emotional parameter includes acquiring at least one parameter r... | Cardiac Pacemakers, Inc. | A61B005/00
US07766831 | B2 | 2010-08-03 | System, tools, devices and a program for diabetes care | 2008-03-07 | A method for diabetes care, the method (which also may be referred to as, involve or incorporate at least one of a tool, device or program) allowing for the characterization of the relevance of errors of parameters affecting glucose concentration on a postprandial glucose concentration outcome for a person with diabetes mellitusm, wherein the method involves at least one of sensing, determining, calculating, predicting, describing and communicating the effects of potential errors of parameters affecting glucose concentration on postprandial glucose concentration values within a clinically relevant glucose range. | 1. A system for determining postprandial glucose concentration, said system comprising: a processor adapted to execute computer implemented instructions to calculate postprandial glucose values based on a therapeutic action scheme and each of: a preprandial glucose concentration measurement by self-monitoring of glucose;an effect of carbohydrate-portion on maximum glucose increase;an estimate of carbohydrate amount in a meal;an effect of prandial insulin on maximum glucose decrease; andan insulin dosage;wherein the processor is further adapted to execute computer implemented instructions to account for a margin of error for the self-monitored preprandial glucose concentration measurement in the calculation;wherein the processor is further adapted to execute computer implemented instructions to account for an error associated with at least one of: the effect of carbohydrate-portion on maximum glucose concentration increase;the estimate of carbohydrate amount in a meal;the effect of prandial insulin on maximum glucose concentration decrease; andthe insulin dosage; andwherein a trend of a continuous blood glucose monitoring is considered as follows:Glucose (mg/dl)241-301-Trend&lt;6161-8081-120121-160161-200201-240300360UU0Y+1Y+2Y+3Y+4Y+5Y+6YU0−1YY+1Y+2Y+3Y+4Y+5Y=0−1YY+1Y+2Y+3Y+4Y+5YD0−2Y−1YY+1Y+2Y+3Y+4YDD0−3Y−2Y−1Y Y+1Y+2Y+3Yand wherein the trends are defined as follows:Very Fast glucose increase>+2mg/dl/minUUFast+ (1-2)mg/dl/minUSlow changes&lt;±1mg/dl/min=Fast decrease− (1-2)mg/dl/minDVery Fast decrease>−2mg/dl/mimDD. 2. The system according to claim 1, wherein the processor is further adapted to execute computer implemented instructions to determine postprandial glucose concentration for different ranges of preprandial glucose concentration values according to the therapeutic scheme. 3. The system according to claim 2, wherein the system further comprises a display operatively coupled to the processor, and wherein the processor is further adapted to execute computer implemented instructions to display postprandial blood glucose as a function of preprandial blood glucose on the display. 4. The system according to claim 3, wherein the processor is further adapted to execute computer implemented instructions to determine whether a critical point is reached by exceeding a lower limit for glucose concentration or by exceeding an upper limit for glucose concentration. 5. The system according to claim 4, wherein the therapeutic scheme includes a carbohydrate self-adjustment in relation to preprandial glucose concentration according to the relation:Carbohydrate Self-adjustment in relation to pre-prandial BG:121-161-201-241-301-BG (mg/dl)&lt;4040-6061-120160200240300360CARB-P(n)X + 2X + 1XX − 1X − 2X − 3X − 4X − 5wherein X equals the number of carbohydrate portions (X=1, 2, 3, 4 or 5) for the blood glucose range of 61-120 mg/dl. 6. The system according to claim 4, wherein the therapeutic scheme includes a pre-prandial insulin dose self-adjustment ac... | Roche Diagnostics International AG | A61B005/00
US20080275357 | A1 | 2008-11-06 | Oral/nasal cannula | 2008-05-01 | There is provided an oral nasal cannula for sampling breath of a subject comprising a main body portion, comprising a suction port which is adapted to be connected to a suction device for side sampling of exhaled breath of said subject, wherein said main body portion comprises a recess adapted to support a medical tube, wherein the tube adapted to be inserted through a first nostril of said subject and a nasal prong and adapted to collect nasally exhaled breath from a second nostril of said subject, and in fluid flow communication with a nasal breath collection bore. | 1. An oral nasal cannula for sampling breath of a subject, comprising: a main body portion, comprising a suction port which is adapted to be connected to a suction device for side sampling of exhaled breath of said subject, wherein said main body portion comprises a recess adapted to support a medical tube, wherein the tube is intended to be inserted through a first nostril of said subject; anda nasal prong and adapted to collect nasally exhaled breath from a second nostril of said subject, and in fluid flow communication with a nasal breath collection bore. 2. The oral nasal cannula according to claim 1, further comprising an oral scoop adapted to collect orally exhaled breath of said subject, and terminating in an oral breath collection bore. 3. The oral nasal cannula according to claim 1, wherein said recess is adapted to clip the tube to secure the tube to the recess. 4. The oral nasal cannula according to claim 1, further comprising a clip adapted to secure the tube to the recess. 5. The oral nasal cannula according to claim 2, wherein said main body portion is formed with at least one of an oral oxygen delivery port and a nasal oxygen delivery port. 6. The oral nasal cannula according to claim 5, wherein said nasal oxygen delivery port comprises a plurality of oxygen delivery holes formed in said main body portion. 7. The oral nasal cannula according to claim 5, wherein said nasal oxygen delivery port comprises an oxygen delivery prong which is shorter than said nasal prong and is adapted to be inserted into a nostril of said subject. 8. The oral nasal cannula according to claim 5 and also comprising a nasal spacer, adapted to distance said at least one nasal oxygen delivery port from the nose of said subject when said oral nasal cannula is placed on the face of said subject. 9. An oral nasal cannula according to claim 5 and wherein said at least one oral oxygen delivery port is formed over said oral scoop. 10. An oral nasal cannula according to claim 6 and wherein said at least one oral oxygen delivery port is directed sideways, such that delivered oxygen is directed around said oral scoop. 11. An oral nasal cannula according to claim 2, and wherein said oral scoop comprises an inner surface which is configured to direct breath, exhaled from the mouth of said subject in substantially any direction, toward said suction port. 12. An oral nasal cannula according to claim 2, and wherein said oral scoop covers substantially a majority of the mouth of said subject when said oral nasal cannula is placed on the face of said subject. 13. An oral nasal cannula according to claim 12, and wherein said oral scoop comprises an inner surface which is configured to direct breath, exhaled from the mouth of said subject in substantially any direction, toward said suction port. 14. An oral nasal cannula according to claim 2, and being formed with an angle between said at least one nasal prong and said oral scoop. 15. An oral nasal cannula accor... | - | A61B005/097
US20080262369 | A1 | 2008-10-23 | Capnographic Sampling Catheter | 2005-12-28 | The present invention seeks to provide a system for analyzing the exhaled breath of a subject, including a gas analyzer, sensitive to at least one component of the exhaled breath and a gas sampling catheter having a first end and a second end, the first end being in fluid flow communication with the gas analyzer, wherein the gas sampling catheter has a plurality of exhaled breath sampling holes formed therein. | 1. A system for analyzing the exhaled breath of a subject, comprising: a gas analyzer, sensitive to at least one component of said exhaled breath; and a gas sampling catheter having a first end and a second end, said first end being in fluid flow communication with said gas analyzer, said gas sampling catheter having a plurality of exhaled breath sampling holes formed therein in the vicinity of said second end. 2. A system according to claim 1, wherein different ones of said plurality of sampling holes are disposed at different longitudinal positions along said gas sampling catheter. 3. A system according to claim 1, wherein different ones of said plurality of sampling holes are disposed at different circumferential positions around said gas sampling catheter. 4. A system according to claim 1 and also comprising a suction element, configured to remove fluid from said gas sampling catheter. 5. A system according to claim 1 and wherein said plurality of sampling holes are disposed in a generally helical pattern. 6. A system according to claim 1 and wherein said second end of said gas sampling catheter is sealed. 7. A system according to claim 1 and wherein said first end of said gas sampling catheter is directly connected to said gas analyzer, thereby forming said fluid flow communication therebetween. 8. A system according to claim 1 and wherein said first end of said gas sampling catheter is connected to said gas analyzer by means of a breath conduit. 9. A system according to claim 4 and wherein said gas sampling catheter is fitted with a suction port, said suction port being adapted to connect to said suction element. 10. A system according to claim 1 and wherein said gas sampling catheter has at least one circumferential protrusion formed therearound 11. A system according to claim 1 and also comprising: an oxygen delivery tube mounted onto said gas sampling catheter, thereby forming a multi-lumen tube therewith; andan oxygen source connected to said oxygen delivery tube for supply of oxygen thereto. 12. A gas sampling catheter comprising a catheter tube being formed with a plurality of exhaled breath sampling holes. 13. A gas sampling catheter according to claim 12, wherein different ones of said plurality of sampling holes are disposed at different longitudinal positions along said catheter tube. 14. A gas sampling catheter according to claim 12 wherein different ones of said plurality of sampling holes are disposed at different circumferential positions around said catheter tube. 15. A gas sampling catheter according to claim 12 and wherein said plurality of sampling holes are disposed in a generally helical pattern. 16. A gas sampling catheter according to claim 12 and wherein one end of said catheter tube is sealed. 17. A gas sampling catheter according to claim 12 and wherein said catheter tube has at least one circumferential protrusion formed therearound. 18. A gas sampling catheter according to claim 13 wherein differe... | - | A61B005/08
US20080287822 | A1 | 2008-11-20 | Method and system for representing brain activities | 2008-07-09 | A system for representing brain activities in the form of visual patterns has been described. The system comprises a signal frequency modulator (frequenzer) and a resonating chamber. The signal frequenzer receives a signal corresponding to a brain activity and generates an audio signal corresponding to the received signal. The resonating chamber receives the audio signal and generates a visual pattern corresponding to the received audio signal. The visual pattern may be recorded by using a camera and displayed on a display device. | 1. A system for representing brain activities in the form of visual patterns, the system comprising: a a signal frequency modulator (frequenzer), the signal frequenzer receiving a signal corresponding to a brain activity and generating an audio signal corresponding to the received signal; andb a resonating chamber, the resonating chamber generating a visual pattern corresponding to the audio signal. 2. The system of claim 1 further comprising a signal capturing module, the signal capturing module capturing the signal corresponding to the brain activity. 3. The system of claim 2, wherein the signal capturing module is an Electroencephalogram (EEG). 4. The system of claim 1, wherein the resonating chamber comprises: a loudspeaker, the loudspeaker comprising: a diaphragm; anda magnet, the magnet inducing the diaphragm to vibrate in response to the audio signal;a bowl, the bowl being filled with a liquid; anda first cold-light-electrode, the first cold-light-electrode emitting light,wherein the first cold-light-electrode is placed between the diaphragm and the bowl. 5. The system of claim 4, wherein the loudspeaker further comprises one or more siphons, the one or more siphons forcing the air through the resonating chamber at atmospheric pressure. 6. The system of claim 4, wherein the resonating chamber further comprises a styrofoam plate, the styrofoam plate providing a base for the bowl. 7. The system of claim 4, wherein the resonating chamber further comprises a black fabric plate, the black fabric plate containing the light within the resonating chamber. 8. The system of claim 4, wherein the first cold-light-electrode is a fluorescent lamp. 9. The system of claim 1 further comprising a camera, the camera recording the visual pattern. 10. The system of claim 9 further comprising a light box, wherein the camera is placed in the light box. a a diffusor plate;b a second cold-light electrode, the second cold-light-electrode emitting light; andc a tube, the tube transferring the visual pattern from the bowl to the camera. 12. The system of claim 11, wherein the light box further comprises a cover plate. 13. The system of claim 9 further comprises a display device, the display device displaying the recorded visual pattern. 14. The system of claim 1, wherein the resonating chamber comprises: a magnetic coil;a bowl, the bowl being filled with a liquid, wherein the magnetic coil induces the bowl to vibrate in response to the audio signal; anda third cold-light-electrode, the third cold-light-electrode emitting light,wherein the third cold-light-electrode is placed above the bowl. 15. The system of claim 14, wherein the third cold-light electrode is a plurality of LEDs. 16. The system of claim 1 further comprising an amplifier, the amplifier amplifying the audio signal. 17. A system for representing brain activities in the form of visual patterns, the system comprising: a. a signal capturing module, the signal capturing module capturing a sig... | - | A61B005/0476
US20080300501 | A1 | 2008-12-04 | Fluid Analyser | 2006-07-06 | A gas analyser (12) comprises a transistor (1) that has a cavity (7) between its gate (2) and its organic semiconductor (6) based conducting channel. In operation a component from a gas sample introduced into the cavity (7) may absorb onto an exposed absorption sensitive surface portion of the organic semiconductor (6). A detector (13) detects a change in the threshold voltage of the transistor caused by the component absorbing on the exposed surface portion. In response to detecting this change, the detector generates a measurement signal indicative of a concentration of the component in the sample. | 1. Fluid analyser comprising: a transistor comprising, a gate and a semiconductor conducting channel, wherein the transistor defines a cavity between the gate and the semiconductor conducting channel such that in use, a component from a fluid sample introduced into the cavity may absorb onto an exposed surface portion of the semiconductor; anda detector for detecting a change in a property of the transistor caused by the component absorbing on the exposed surface of the semiconductor and in response thereto generating a measurement signal indicative of a concentration of the component in the sample. 2. Fluid analyser according to claim 1, wherein the semiconductor conducting channel is an organic semiconductor. 3. Fluid analyser according to claim 2 wherein the organic semiconductor comprises polyarylamine. 4. Fluid analyser according to claim 1 wherein the transistor further comprises an insulating layer formed on a surface of the gate, the insulating layer leaving exposed a surface portion of the gate and wherein the gate, the semiconductor conducting channel and the insulating layer together substantially define the cavity with the exposed surface portion of the semiconductor facing the exposed surface portion of the gate. 5. Fluid analyser according to claim 4 wherein the transistor further comprises a source formed between a first portion of the insulator layer and the semiconductor and a drain formed between a second portion of the insulator layer and the semiconductor. 6. Fluid analyser according to claim 1 wherein the transistor further comprises a protective layer formed on an upper surface of the semiconductor. 7. Fluid analyser according to claim 1, wherein the transistor property is its threshold voltage. 8. Fluid analyser according to claim 1 wherein the fluid sample is a breath sample, the analyser further comprising a mouth piece for delivering the breath sample to the cavity. 9. Fluid analyser according to claim 1, wherein the component is one of: nitrogen monoxide, acetone, ethanol, carbon monoxide and isoprene. 10. A method of analyzing a fluid sample, the method comprising; receiving a fluid sample into a cavity defined in a transistor between a gate of the transistor and a semiconductor layer that forms a conducting channel of the transistor,such that a component of the fluid sample can absorb on an exposed surface portion of the semiconductor layer;detecting a change in a characteristic of the transistor induced by the component absorbing on the exposed surface portion and in response thereto generating a signal indicating a concentration of the component in the sample. | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61B005/08 H01L051/05
US20080281169 | A1 | 2008-11-13 | Ear-Thermometer With Ear Identification | 2006-09-14 | A physiological measurement device (10) that concurrently measures a physiological state of an individual and determines an identity of the individual includes at least a measurement component (12), an identification component (14), and processing component (16). The measurement component (12) measures the physiological state of the individual, and the identification component (14) concurrently determines the identity of the individual. The processing component (16) associates the physiological state of the individual with the identity of the individual. | 1. A physiological measurement device that concurrently measures a physiological state of an individual and determines an identity of the individual, comprising: a measurement component that measures the physiological state of the individual;an identification component that concurrently determines the identity of the individual; anda processing component that associates the physiological state of the individual with the identity of the individual. 2. The physiological measurement device as set forth in claim 1, wherein the measurement component includes an electronic ear thermometer that measures a temperature of the individual from an ear of the individual. 3. The physiological measurement device as set forth in claim 1, wherein the identification component determines the identity of the individual from the same information used by the measurement component to determine the physiological state of the individual. 4. The physiological measurement device as set forth in claim 1, wherein the identification component uses one of a reflected sound and an otoacoustic technique to determine the identity of the individual. 5. The physiological measurement device as set forth in claim 1, wherein the measurement component measures at least one of temperature, blood pressure, eye pressure, heart rate, heart electrical activity, pulse, and blood oxygen level of the individual. 6. The physiological measurement device as set forth in claim 1, wherein the identification component determines the identity of the individual through biometric identification including at least one of a unique physical characteristic of an ear, a fingerprint, and an iris map. 7. The physiological measurement device as set forth in claim 1, further including: a storage component for storing multiple temperatures for a plurality of individuals, wherein each temperature is associated with a corresponding individual based on the identity of the individual as determined by the identification component. 8. The physiological measurement device as set forth in claim 1, further including: a display component for displaying the temperature and the identity of the individual. 9. The physiological measurement device as set forth in claim 1, further including: portable storage for transferring stored temperatures and identity information from the system to devices that reads the portable storage. 10. The physiological measurement device as set forth in claim 1, further including: a network component for conveying the temperature and the identity of the individual to other devices over a network. 11. The physiological measurement device as set forth in claim 1, further including: a shaft that houses the measurement component and the identification component, the shaft is designed to be inserted within an ear of the individual. 12. The physiological measurement device as set forth in claim 1, the physiological measurement system uses a local body area network to convey the physiological ... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/00 G01K005/22
US20080312550 | A1 | 2008-12-18 | Multi-Functional and Modular Urine Collection System | 2006-06-29 | Multi-functional urine collection devices, embodiments of which can include a self-expanding container having a receptacle for receiving urine from the tubing, a pump for moving urine through the tubing and into a receptacle, extendable tubing that may be shortened and/or lengthened, and/or one or more meters for monitoring, measuring, transmitting or storing a characteristic from the urine. Embodiments can also include a display kit having a transmitting device in the form of a cable or wireless transmitter and a display module. The display module may include an LCD and may be configured for attachment to a urine receptacle. | 1-29. (canceled) 30. A display kit for a urinary catheter system, comprising: a transmitting device connected to a transducer, the transducer connected to a proximal port of a urinary catheter having a distal end disposed in a bladder; anda display module configured to display a characteristic of urine from the catheter, the display module attached to a urine receptacle. 31. The display kit according to claim 30, wherein the display module includes a liquid crystal display. 32. The display kit according to claim 30, wherein the transmitting device includes a cable connected to the display module. 33. The display kit according to claim 30, further comprising a receiver connected to the display module, wherein the transmitting device includes a wireless transmitter. 34. The display kit according to claim 30, wherein the characteristic of urine is the temperature thereof. 35. The display kit according to claim 30, wherein the urine receptacle is connected to a urine meter including a measurement region and a plurality of closeable ports between the receptacle and the measurement region of the meter for emptying the measurement region into the receptacle. 36. The display kit according to claim 30, wherein the receptacle comprises a self-expanding collection container having a contracted position and an expanded position and a bias to convert the container from the contracted position to the expanded position. 37. The display kit according to claim 30, further comprising a plurality of transmitting devices and display modules in respective communication. | - | A61B005/20 A61M025/00
US07471978 | B2 | 2008-12-30 | Brain function scan system | 2002-01-10 | A portable EEG (electroencephalograph) instrument, especially for use in emergencies and brain assessments in physicians' offices, detects and amplifies brain waves and converts them into digital data for analysis by comparison with data from normal groups. In one embodiment, the EEG electrodes are in a headband which broadcasts the data, by radio or cellular phone, to a local receiver for re-transmission and/or analysis. In another embodiment, the subject is stimulated in two modes, i.e., aural and sensory, at two different frequencies to provide the subject's EPs (Evoked Potentials), assessing transmission through the brainstem and thalamus. | 1. A medical system to analyze brain waves of a subject comprising: (a) an active EEG (electroencephalograph) electrode detecting a subject's brain waves;(b) a stimulus generator providing to the subject concurrent sense stimuli in a plurality of stimulus modes, the stimuli in a first one of the modes being at a frequency F1 and the stimuli in a second one of the modes being at a frequency F2;(c) an amplifier amplifying and digitizing the detected brain waves;(d) a processor receiving the digitized brain waves from (c) and producing subject brain wave F ratio data as a function of a power of brain responses at each of F1 and F2 in the presence and absence of stimulation;(e) a computer comparing the subject brain wave F ratio data from (d) with one of brain wave F ratio data generated as a function of one of (i) data from a control group of patients and (ii) data generated by the subject in the presence and absence of stimulation;(f) an output coupled to the computer and producing a warming when the comparison of (e) indicates one of injury to and dysfunction of one of the subject's spinal cord, brain stem and brain; and(g) a modulator modulating a carrier wave and the amplified brain waves to generate an audio signal therefrom. 2. The medical system according to claim 1, wherein a statistical evaluation of computed measures from a subject is determined by computing a Z-score, where Z =(M−P)/6, wherein M is a mean value of a normative distribution, P is a current measure from the subject and 6 is a standard deviation of a control age- matched population. 3. The medical system according to claim 1, further comprising: a radio broadcast transmitter; anda headband situating thereon the electrode, the amplifier and the radio broadcast transmitter. | New York University | A61B005/00
US07467008 | B2 | 2008-12-16 | Ectography multimodality imaging system for diagnosis and treatment | 2006-09-25 | A system and method of integrating a nuclear medicine imaging device such as a gamma camera with a patient handling system where the nuclear medicine imaging system does not prevent full access to the patient and does not interfere with the field of view of accompanying modalities accommodated by use of the patient handling system. | 1. A multimodality imaging system comprising: an anatomical imaging device having a field of view (FOV);a Patient Handling System (PHS), wherein said PHS has a bed and a support for said bed, for positioning a patient to be imaged by said anatomical imaging device;a nuclear medicine imaging device, having a radiation detector and a collimator being integrated and movable with said PHS such that imaging of a patient on said bed can be conducted without said nuclear medicine imaging device interfering with the field of view of said anatomical imaging device. 2. The multimodality imaging system of claim 1, wherein said bed is movable in the longitudinal axis of the bed so that a target patient contained on said bed may be translated with respect to said anatomical imaging device. 3. The multimodality imaging system of claim 2, wherein the nuclear medicine imaging device is integrated with said PHS such that scanning of a target patient can be conducted without impeding full access to the target patient by medical personnel during scanning. 4. The multimodality imaging system of claim 3, wherein the target patient is scanned by the nuclear medicine device and thereafter translated into the FOV of one or more accompanying imaging modalities for scanning or treatment by said modalities. 5. The multimodality imaging system of claim 3, wherein the nuclear medicine imaging device is enclosed in a protective sealed casing. 6. The multimodality imaging system of claim 5, wherein said casing is a movable shield, and comprised of low attenuation, low Z material; and wherein said movable shield closes during use of accompanying modalities. 7. The multimodality imaging system of claim 5, wherein the casing is constructed of lead or copper. 8. The multimodality imaging system of claim 3, wherein the target subject is translated into the FOV of one or more other accompanying modalities after target subject is scanned by said nuclear medicine imaging device. 9. The multimodality imaging system of claim 2, wherein the target patient is scanned by the nuclear medicine imaging device while the target patient is translated on the bed for imaging by said anatomical imaging device. 10. The multimodality imaging system of claim 1, wherein the collimator is a rotating acceptance angle collimator. 11. The multimodality imaging system of claim 1, wherein the collimator is a rotating slat collimator. 12. The multimodality imaging system of claim 1, wherein the nuclear medicine imaging device is located beneath the bed. 13. The multimodality imaging system of claim 1, wherein a substantial amount of the nuclear medicine imaging device is located beneath the bed. 14. The multimodality imaging system of claim 1, wherein the nuclear medicine imaging device is not fixed or attached to a gantry. 15. The multimodality imaging system of claim 1, wherein the target patient is scanned while the target subject is held stationary on the bed. 16. The multimodality imaging ... | Siemens Medical Solutions USA, Inc. | A61B005/00
US08282550 | B2 | 2012-10-09 | Integrated receiver for continuous analyte sensor | 2008-07-29 | A system is provided for monitoring glucose in a host, including a continuous glucose sensor that produces a data stream indicative of a host's glucose concentration and an integrated receiver that receives the data stream from the continuous glucose sensor and calibrates the data stream using a single point glucose monitor that is integral with the integrated receiver. The integrated receiver obtains a glucose value from the single point glucose monitor, calibrates the sensor data stream received from the continuous glucose sensor, and displays one or both of the single point glucose measurement values and the calibrated continuous glucose sensor values on the user interface. | 1. An integrated system comprising: a continuous glucose sensor configured to continuously measure a glucose concentration in a host, wherein the continuous glucose sensor is configured to generate continuous glucose sensor data associated with the glucose concentration in the host;a single point glucose monitor, wherein the single point glucose monitor is configured to generate single point glucose monitor data associated with a glucose concentration of a biological sample from the host; anda processor module configured to receive and process the continuous glucose sensor data and the single point glucose monitor data, wherein the processor module is configured to check a clinical acceptability of a deviation between the single point glucose monitor data and corresponding continuous glucose sensor data and a risk associated with the deviation based on the glucose concentration indicated by the single point monitor data or the glucose concentration indicated by the corresponding continuous glucose sensor data. 2. The system of claim 1, wherein the processor module is configured to automatically receive the single point monitor data without user interaction. 3. The system of claim 1, wherein the processor module is configured for data communication with an insulin delivery device, and wherein the processor module is further configured to process and transmit insulin administration information to the insulin delivery device. 4. The system of claim 3, further comprising an insulin pump, and wherein the data communication with the insulin delivery device comprises data communication with the insulin pump. 5. The system of claim 1, wherein a sensing membrane of each of the continuous glucose sensor and the test strip comprises an enzyme domain, and wherein the enzyme domain of the sensing membrane of the continuous glucose sensor and the enzyme domain of the sensing membrane of the test strip are substantially the same. 6. The system of claim 1, wherein the processor module is configured to control time and/or amount of administration of insulin to the host. 7. The system of claim 6, wherein the processor module is configured to control a time and/or an amount of administration of insulin to the host by controlling administration of insulin to the host based at least in part on a rate of change of the continuous glucose sensor data or a directional trend of the continuous glucose sensor data. 8. The system of claim 6, wherein the processor module is configured to control a time and/or an amount of administration of insulin to the host by controlling administration of insulin to the host based at least in part on the glucose concentration indicated by the single point monitor data and/or the glucose concentration indicated by the corresponding continuous glucose sensor data. 9. The system of claim 1, wherein the processor module is configured to calculate a rate of change of the continuous glucose sensor data or a directional trend of the conti... | DexCom, Inc. | A61B005/00
US08328735 | B2 | 2012-12-11 | Analytical system for detecting an analyte in a body fluid and disposable integrated puncturing and analyzing element | 2008-04-18 | An analytical system for detecting an analyte in a body fluid comprises a disposable integrated puncturing and analyzing element and a reusable analytical system. The integrated element includes a puncturing element having a tip for generating a puncture wound in a body part and an elongated flat test strip having a top side and a bottom side. The integrated element is held in a holder such that, in a sample application position of the test strip, the top side of the test strip is facing upward. A sample application area for manual application of a drop exiting from the puncture wound of the body part is provided on the top side of the test strip such that, in the sample application position, the drop is applied to the sample application area from directly above. | 1. An analytical system for detecting an analyte in a body fluid comprising a disposable integrated puncturing and analyzing element including a puncturing element having a tip for generating a puncture wound in a body part and including an elongated flat test strip having a top side and a bottom side, wherein the puncturing element defines a sample transfer orifice, anda reusable analyzing instrument including a housing which includes a holder for holding the integrated puncturing and analyzing element, a drive, and a coupling unit for coupling the puncturing element to said drive, said drive being adapted for driving a puncturing movement of the puncturing element during which the puncturing element is moved on a movement path, said puncturing movement including a forward phase in which the puncturing element is moved in a puncturing direction and a retraction phase during which the puncturing element is moved, after having reached a reversal point, opposite to the puncturing direction,whereinthe puncturing element is arranged parallel to the test strip and the puncturing element is movable on a movement path relative to the test strip,the holder adapted for holding the integrated puncturing and analyzing element, in a sample application position of the test strip, in which the top side of the test strip faces upward, the holder and the drive adapted and arranged such that, in a puncturing position of the test strip, the tip of the puncturing element is moved beyond the front end of the test strip in the puncturing direction, in order to generate a puncture wound in a body part, wherein the puncturing element covers the sample application area when at the puncturing position,the sample application area located at the top side of the test strip such that, in the sample application position a drop of sample liquid from the puncture wound can be directly applied manually from above to the sample application area, wherein the sample transfer orifice is positioned directly above the sample application area when at the sample application position, andthe housing being shaped to allow access to the sample application area of a test strip positioned in the sample application position. 2. The analytical system according to claim 1, characterized in that the front end of the test strip in a puncturing direction comprises a skin contact surface and serves as a puncturing depth reference element, whereby the puncturing depth is determined by a distance in the puncturing direction between the skin contact surface and a position of the tip of the puncturing element at the reversal point of the puncturing movement. 3. The analytical system according to claim 1, characterized in that the coupling unit comprises a coupling element for coupling to the puncturing element, said coupling element connecting to the test strip from the bottom side of the test strip. 4. The analytical system according to claim 1, wherein the puncturing element is at least partly cov... | Roche Diagnostics Operations, Inc. | A61B005/00 A61B017/14 A61B017/32 B65D081/00
US08348839 | B2 | 2013-01-08 | Systems and methods for active listening/observing and event detection | 2007-04-10 | Embodiments of the present invention provide a method of monitoring an environment comprising: monitoring at least one data stream wherein the data stream is a data stream in the environment; detecting a specified event from the data stream; and triggering a response to the specified event. Embodiments of the present invention provide a system for monitoring an environment comprising: a receiver adapted to receive at least one input data stream wherein the input data stream is a data stream in the environment; an active listener/observer system adapted to monitor the data stream; and an interface adapted to express at least one output stream. Embodiments of the present invention provide a computer-readable medium having instructions comprising: an active listener/observer routine configured to monitor at least one data stream; a detection routine configured to find specified events in the data stream; and an output routine configured to express a response event. | 1. A method of monitoring a healthcare environment for purposes of early event detection comprising: actively monitoring a healthcare environment for speech input using a microphone to input speech from the healthcare environment;activating a speech recognition program in real time upon detection of the speech input to interpret the speech;using a computer processor to compare the interpreted speech to keywords in a database; andif the interpreted speech matches a keyword in the database, triggering a response event for the actively monitored healthcare environment that includes outputting an alert and activating an event log that includes a listing of at least one of resuscitation activities and voice identifications,wherein the speech input at the microphone is actively monitored in substantially real-time using the speech recognition program to interpret the speech and the computer processor to compare the interpreted speech to keywords in the database. 2. The method of claim 1 wherein the steps are performed sequentially. 3. The method of claim 1 wherein the alert includes at least one of contacting an individual, contacting an agency, initiating a page and initiating a software message pop-up. 4. The method of claim 1 further comprising: differentiating between speech of a plurality of speakers in the healthcare environment. 5. The method of claim 1 further comprising: using an input device to input data, wherein the input device includes at least one of a keyboard, a touchscreen, a joystick, a mouse and a touchpad; andoutputting an alert based on the input data. 6. The method of claim 1 further comprising: triggering an enterprise event if the interpreted speech matches a keyword in the database, wherein the enterprise event includes at least one of locking a door and activating a sprinkler system. 7. The method of claim 1 wherein multiple alerts are triggered. 8. The method of claim 1 further comprising: using a physiological monitoring device to input patient data; andoutputting an alert based on the patient data. 9. The method of claim 1 further comprising: inputting information from a healthcare information system; andoutputting an alert based on the information from the healthcare information system. 10. A healthcare environment monitoring system comprising: a microphone configured to input speech from a healthcare environment in order to actively monitor the healthcare environment, wherein the speech is interpreted in real time upon detection of the speech input using a speech recognition program; anda computer processor configured to compare the interpreted speech to keywords in a database, the processor configured such that if the interpreted speech matches a keyword in the database, the processor triggers a response event for the actively monitored healthcare environment, the response event including outputting an alert and activating an event log that includes a listing of at least one of resuscitation activities and voic... | General Electric Company | A61B005/00 G10L011/00
US08320995 | B2 | 2012-11-27 | Fiducial marker with rings | 2007-07-19 | An interstitial marker for the location of one of an organ, tumor and tumor bed within a mammalian body includes a plurality of rings having an inner diameter and a ribbon positioned within the inner diameters of the plurality of rings. The plurality of rings may be of the same material or different materials. The materials are visible under different imaging modalities. The ribbon similarly may be of a material visible under various imaging modalities. Tips are attached to the ribbon to hold the rings in place. Alternatively, one or more of the rings may be attached to the ribbon. A helical coil may also be used in connection with a central ribbon for improved imaging. A plurality of coils, of the same material or different materials, may be positioned on a ribbon for a marker. | 1. An interstitial marker for the location of one of an organ, tumor and tumor bed within a mammalian body, the marker comprising: a plurality of adjacent rings, each ring of the plurality of adjacent rings being configured as a closed band defining a circular curve and having an inner diameter;a ribbon positioned within the inner diameters of the adjacent rings, each ring of the plurality of adjacent rings and the ribbon comprising a material visible under a common imaging modality; andat least one tip connected to an end of the ribbon;wherein the at least one tip is formed of a material visible under at least one imaging modality; andwherein the at least one tip includes:a first tip having a D-shaped cross-section, the first tip secured to a first end of the ribbon,a second tip having a D-shaped cross-section the second tip secured to a second end of the ribbon, the second end opposing the first end,the first tip and the second tip each having an outer diameter that is greater than the inner diameter of each ring of the plurality of adjacent rings, andthe first tip and the second tip being configured to capture the plurality of adjacent rings along the ribbon there between, each ring of the plurality of adjacent rings being configured to translate along a longitudinal axis of the ribbon between the first tip and the second tip. 2. The interstitial marker according to claim 1 wherein each of the plurality of rings and the ribbon is comprised of a dense, precious metal. 3. The interstitial marker according to claim 1 wherein the plurality of adjacent rings include a first set of rings having a first set of dimensions and a second set of rings having a second set of dimensions. 4. The interstitial marker according to claim 1 wherein each of the plurality of rings is movable with respect to the ribbon. 5. The interstitial marker according to claim 1 wherein the ribbon is attached to at least one of the rings so that the at least one ring does not move with respect to the ribbon. 6. The interstitial marker according to claim 1 wherein at least one component of the marker, from among the ribbon and the plurality of rings, comprises a radiopaque material. 7. The interstitial marker according to claim 1 wherein at least one component of the marker, from among the ribbon and the plurality of rings, comprises a material visible under ultrasound imaging. 8. The interstitial marker according to claim 1 wherein at least one component of the marker, from among the ribbon and the plurality of rings, comprises a material visible under diagnostic x-ray. 9. The interstitial marker according to claim 1 wherein at least one component of the marker, from among the ribbon and the plurality of rings, comprises a material visible under therapeutic x-ray. 10. The interstitial marker according to claim 1 wherein the ribbon and the rings comprise biocompatible materials for implantation. 11. The interstitial marker according to claim 1 wherein: a first one of t... |  | A61B005/00
US08282564 | B2 | 2012-10-09 | Systems and methods for model-based estimation of cardiac output and total peripheral resistance | 2008-05-15 | The methods and systems for estimating cardiac output and total peripheral resistance include observing arterial blood pressure waveforms to determine intra-beat and inter-beat variability in arterial blood pressure and estimating from the variability a time constant for a lumped parameter beat-to-beat averaged Windkessel model of the arterial tree. Uncalibrated cardiac output and uncalibrated total peripheral resistance may then be calculated from the time constant. Calibrated cardiac output and calibrated total peripheral resistance may be computed using calibration data, assuming an arterial compliance that is either constant or dependent on mean arterial blood pressure. The parameters of the arterial compliance may be estimated in a least-squares manner. | 1. A method for estimating beat-by-beat cardiovascular parameters and variables, comprising: processing two or more cycles of arterial blood pressure to determine a mean arterial blood pressure over each cycle beginning at a beat onset time and an inter-beat variability in blood pressure between the two or more cycles; andcomputing estimates of one or more cardiovascular parameters and variables from the mean arterial blood pressures, the inter-beat variability, and a beat-to-beat averaged Windkessel model of an arterial tree. 2. The method of claim 1, wherein the one or more cardiovascular parameters and variables include a beat-by-beat time constant of the arterial tree. 3. The method of claim 2, wherein the time constant is estimated over a data window. 4. The method of claim 2, wherein the time constant is estimated through optimization of an error criterion. 5. The method of claim 4, wherein the error criterion is least-squared error. 6. The method of claim 1, wherein the one or more cardiovascular parameters and variables include an uncalibrated beat-by-beat cardiac output. 7. The method of claim 6, further comprising computing calibrated beat-by-beat cardiac output from the uncalibrated beat-by-beat cardiac output using a calibration factor. 8. The method of claim 7, wherein the calibration factor is computed for each of the cycles. 9. The method of claim 7, wherein the calibration factor represents a lumped arterial compliance. 10. The method of claim 9, wherein the lumped arterial compliance is modeled as a function of mean arterial blood pressure. 11. The method of claim 9, wherein the lumped arterial compliance is modeled as a parameterized function of mean arterial blood pressure. 12. The method of claim 9, wherein the lumped arterial compliance is modeled as a two-parameter function of mean arterial blood pressure. 13. The method of claim 9, wherein the lumped arterial compliance is modeled as a constant. 14. The method of claim 11, 12 or 13, wherein a parameter of the lumped arterial compliance is estimated through optimization of an error criterion. 15. The method of claim 14, wherein the error criterion is least-squared error. 16. The method of claim 6, wherein the one or more cardiovascular parameters and variables include an uncalibrated beat-by-beat total peripheral resistance. 17. The method of claim 16, further comprising computing calibrated beat-by-beat total peripheral resistance from a ratio of the time constant to a lumped arterial compliance. 18. The method of claim 16, further comprising computing calibrated beat-by-beat total peripheral resistance from a ratio of mean arterial blood pressure to calibrated cardiac output. 19. The method of claim 16, further comprising computing calibrated beat-by-beat total peripheral resistance from a ratio of mean pressure to a systemic blood flow. 20. The method of claim 19, wherein the systemic blood flow is calculated as the calibrated cardiac output minus the p... | Massachusetts Institute of Technology | A61B005/02
US08353288 | B2 | 2013-01-15 | Patient care unit with a reclining surface | 2008-05-29 | A patient care unit (1) with a reclining surface has one or more docking devices designed as inductive coupling sites (7) for medical technical devices (10, 11, 12), which consume electrical energy and can be connected to the patient care unit. The inductive coupling sites (7) also supply the patient care unit (1) with electrical energy by means of an external line connection, which is likewise coupled inductively with the patient care unit (1). | 1. A patient system comprising: a patient care unit with a reclining surface and a docking device connected to said reclining surface, said docking device including an inductive coupling site for inductive coupling;a medical technical device which consumes electric energy, said medical device including an inductive supply tap for inductively coupling said medical technical device to the patient care unit via said inductive coupling site;an external line connection for supplying said patient care unit with electrical energy; anda flexible breathing tube wherein said external line connection for the electrical energy supply for said patient care unit comprises wire lines extending in a longitudinal direction of said flexible breathing tube. 2. A patient system in accordance with claim 1, further comprising a respirator wherein said flexible breathing tube is in a breathing line connection with said respirator. 3. A patient system in accordance with claim 1, wherein said docking device includes an inductive feeding coupling site and said flexible breathing tube is equipped with a tube inductive supply tap for inductively coupling to said inductive feeding coupling site as an inductive feeding site for feeding energy into said inductive coupling site at said patient care unit. 4. A patient system in accordance with claim 3, wherein said patient care unit further comprises a supply bus provided in said patient care unit and connected to each of said inductive coupling site and said inductive feeding coupling site for supplying said inductive coupling site with electrical energy via said feeding coupling site. 5. A patient system in accordance with claim 1, wherein said medical technical device comprises one of a patient monitor, an infusion pump, and a portable monitor. 6. A patient system in accordance with claim 4, further comprising an information network with a hospital information system and with an application server, said patient care unit being connected to said application server via said information network. 7. A patient system in accordance with claim 6, further comprising a portable monitor provided with a radio interface for a first data path into said information network via a wireless access node of said information network as well as with a second line-bound data path via a connection line with said feeding coupling site, said flexible breathing tube, a respirator and a line from said respirator into said information network. 8. A patient system in accordance claim 7, wherein a transmission path for the energy supply also forms a data transmission path. 9. A patient system in accordance with claim 8, wherein said portable monitor and/or a infusion pump are remotely controllable by means of said respirator or by means of a stationary monitor. 10. A patient system comprising: a patient care unit with a reclining surface, a supply bus fixed to said patient care unit and a docking device connected to said reclining surface, said d... | Dräger Medical GmbH | A61M015/00 A61M016/00 A62B007/00 A62B009/00
US09983282 | B2 | 2018-05-29 | Stimulus induced rotary saturation for magnetic resonance functional imaging | 2007-11-21 | Spin-lock T1ρ-weighted images of a subject are acquired and processed to produce an image that is indicative of bioelectromagnetic activity in the subject. A spin-lock RF field B1ρ is produced such that the Larmor frequency in the rotating frame corresponds to the expected frequency of the bioelectromagnetic activity. As a result, the magnetic field fluctuations generated by the bioelectromagnetic currents shorten the T1ρ of the surrounding tissue, creating a contrast mechanism that is seen as a reduced MR signal in the T1ρ-weighted image that is produced. | 1. A method for producing an image of bioelectromagnetic activity in a subject with a magnetic resonance imaging (MRI) system, the steps comprising: a) acquiring T1ρ-weighted image data from the subject using a pulse sequence that directs the MRI system to perform a spin-lock preparatory pulse sequence prior to acquiring nuclear magnetic resonance (NMR) image data, wherein the spin-lock preparatory pulse sequence includes applying a spin-lock radio frequency (RF) field, B1ρ, that establishes a spin-lock condition in transverse spin magnetization such that transverse spin magnetization in the spin-lock condition is rotated by bioelectromagnetic fields directly produced by neuronal activity in the subject;b) reconstructing an image with the acquired T1ρ-weighted image data; andc) analyzing image signals in the T1ρ-weighted image to detect locations of neuronal activity in the subject. 2. The method as recited in claim 1 in which a magnitude of the spin-lock RF field, B1ρ, is set to establish a Larmor frequency in a rotating frame of transverse magnetization that is substantially similar to a frequency of magnetic field fluctuations produced by the bioelectromagnetic field produced by neuronal activity in the subject. 3. The method as recited in claim 2 in which the Larmor frequency in the rotating frame is less than 100 Hz. 4. The method as recited in claim 1 in which the spin-lock preparatory pulse sequence includes: a first RF excitation pulse that tips longitudinal spin magnetization to a transverse plane prior to establishing the spin-lock condition with the spin-lock RF field, B1ρ; anda second RF excitation pulse that tips the transverse spin magnetization back to longitudinal spin magnetization after the established spin-lock condition. 5. The method as recited in claim 4 in which an imaging pulse sequence is performed by the MRI system after the second RF excitation pulse. 6. The method as recited in claim 5 in which the imaging pulse sequence is an echo planar imaging (EPI) pulse sequence. 7. The method as recited in claim 5 in which the spin-lock preparatory pulse sequence also includes a gradient pulse that dephases transverse magnetization that remains after application of the second RF excitation pulse. 8. The method as recited in claim 1 in which the spin-lock RF field, B1ρ, is applied for a time period sufficiently long to enable the bioelectromagnetic fields produced by neuronal activity to alter the transverse spin magnetization in the spin-lock condition, the altered transverse spin magnetization providing a contrast mechanism for detecting the neuronal activity. 9. The method as recited in claim 1 in which the spin-lock preparatory pulse sequence includes: a first RF excitation pulse that produces transverse spin magnetization by tipping longitudinal spin magnetization to a transverse plane prior to establishing the spin-lock condition with the spin-lock RF field, B1ρ; andan imaging pulse sequence that reads out an NMR sig... | The General Hospital Corporation | A61B005/05 G01R033/48 G01R033/56
US09820671 | B2 | 2017-11-21 | Devices and methods for assessing motor point electromyogram as a biomarker | 2008-05-16 | Described herein are methods of detecting the presence or progression of neuromuscular disorders (including ALS) based at least in part on diaphragm EMG signals taken at different points in time from the same implanted electrode(s). Signals may be used to diagnose or to monitor progression of a disorder or to track a treatment. Thus, diaphragm EMG signal may be used as a marker (e.g., a “biomarker”) for the detection or progression of a neurological disorder such as ALS. The characteristics or parameters of diaphragm EMG signals may be used to create an activity index, which may be output or compared and further analyzed. These signals may also be analyzed to show a difference between the left and right sides of the diaphragm. | 1. A method of diagnosing or tracking the progression or treatment of a neurological disorder that affects the function of a diaphragm, the method comprising: measuring a first EMG signal during a volitional breathing maneuver at an initial recording time before the start of therapy from one or more electrodes chronically implanted in a diaphragm;measuring a second EMG signal from the same one or more electrodes at a later recording time during a volitional breathing maneuver after the start of the therapy; andcomparing one or more characteristics of the first EMG signal and the second EMG signal and characterizing the presence, progression or treatment of a neurological disorder based on the comparison. 2. The method of claim 1, wherein the disorder comprises any of a centrally-mediated breathing disorder or a neurodegenerative disease. 3. The method of claim 1, wherein the comparing step comprises comparing one or more of amplitude, frequency, power spectrum, time-to-peak rise time, relaxation time, compound action potential rise time, compound action potential fall time, fast twitch muscle contribution, and slow-twitch muscle contribution. 4. The method of claim 1, further comprising: mapping the diaphragm to determine an active region of the diaphragm; andimplanting the one or more electrodes at an active region of the diaphragm. 5. The method of claim 1, further comprising: mapping the diaphragm to determine a phrenic nerve motor point of the diaphragm; andimplanting the one or more electrodes at the phrenic nerve motor point. 6. The method of claim 5, wherein the mapping step includes applying electrical stimuli to an area of diaphragm surface and identifying a site within the area that provokes a maximal physiological response to the stimuli. 7. The method of claim 6, wherein the mapping step includes observing a physiological response to the stimuli including any of: abdominal pressure or airway pressure, and diaphragm contraction. 8. The method of claim 1, wherein measuring the first EMG signal comprises: selecting a duration and/or period of time for collection, wherein the duration and/or period of time for collection for the first EMG signal is the same for the second EMG signal; andquantifying at least one characteristic of the EMG signal over the selected time. 9. The method of claim 8, wherein the time duration of collection of the EMG signal is selected from the group consisting of: a single breath and a particular number of breaths. 10. The method of claim 8, wherein the characteristic of the EMG signal quantified is selected from the group consisting of: amplitude, frequency, power spectrum, time-to-peak rise time, relaxation time, compound action potential rise time, compound action potential fall time, fast twitch muscle contribution, and slow-twitch muscle contribution. 11. The method of claim 1, wherein the steps of measuring the first EMG signal and measuring the second EMG signal comprises filtering the EMG signa... | SYNAPSE BIOMEDICAL, INC. | A61B010/00 A61B005/0488 A61N001/36
US20080249420 | A1 | 2008-10-09 | Vascular Catheter Device and Related Methods of Using the Same | 2005-10-14 | A catheter device fur diagnostic vascular treatment and/or therapeutic vascular treatment of a subject's vasculature having a blunt, atraumatic (non-traumatic) distal tip and orifice edge to the catheter. This will eliminate or mitigate catastrophic complications caused by contemporary catheters such as vascular dissection, thrombosis, distal embolization and vessel occlusion. The catheter tip can be shaped in various ways with solid material (compressible or non-compressible) or balloon inflation devices to create a blunt, atraumatic (non-traumatic) tip and orifice, and may include a set-back region as well. | 1. A catheter device for diagnostic vascular treatment and/or therapeutic vascular treatment of a subject's vasculature, said device comprising: a catheter shaft having a proximal portion and a distal portion; anda distal tip disposed on said distal portion, said distal tip having a blunt shape adapted to avoid or mitigate trauma with an ostium of the vasculature. 2. The catheter device of claim 1, wherein said distal tip is adapted to be compressible to fit through a sheath or other conduit. 3. The catheter device of claim 1, wherein said distal tip is inflatable. 4. The catheter device of claim 1, wherein said distal tip has at least one the following shapes: olive, bulbous, rounded, spherical, hemispherical, conical, oval, tapered, beveled, chamfered, graduated and/or multi-faceted, or any combination thereof. 5. The device of claim 4, wherein said distal tip comprises a set-back extension located on the distal end of said distal tip. 6. The device of claim 5, wherein said set-back extension having a blunt shape adapted to avoid or mitigate trauma with an ostium of the vasculature. 7. The catheter device of claim 6, wherein said set-back extension has at least one the following shapes: olive, bulbous, rounded, spherical, hemispherical, conical, oval, tapered, beveled, chamfered, graduated and/or multi-faceted, or any combination thereof. 8. The device of claim 5, wherein said set-back extension has at least one the following shapes: semi-elliptical, semi-spherical, hemispherical, semi-oval, partly rounded or partly olive, or any combination thereof. 9. The device of claim 5, wherein said set-back extension is adapted to be manipulated along the entire geometric spectrum of potential shapes to create non-traumatic tip. 10. The catheter device of claim 5, further comprising: a set-back balloon disposed on said set-back extension that is inflatable. 11. The catheter device of claim 10, wherein said set-back balloon has a pre-formed shape for inflation. 12. The device of claim 1, wherein said distal tip comprises a set-back off extension located on the distal end of said distal tip. 13. The device of claim 12, wherein said set-back extension having a blunt shape adapted to avoid or mitigate trauma with an ostium of the vasculature. 14. The catheter device of claim 13, wherein said set-back extension has at least one the following shapes: olive, bulbous, rounded, spherical, hemispherical, conical, oval, tapered, beveled, chamfered, graduated and/or multi-faceted, or any combination thereof. 15. The device of claim 12, wherein said set-back extension has at least one the following shapes: semi-elliptical, semi-spherical, hemispherical, semi-oval, partly rounded or partly olive, or any combination thereof. 16. The device of claim 12, wherein said set-back extension is adapted to be manipulated along the entire geometric spectrum of potential shapes to create non-traumatic tip. 17. The catheter device of claim 12, further comprising: a... | - | A61M025/10 A61B005/02 A61M025/01
US20080319281 | A1 | 2008-12-25 | Device for Detecting and Warning of Medical Condition | 2006-12-19 | A portable device detects a medical condition, such as an epileptic seizure. The device is implemented on a wrist band, possibly together with a watch, or in a helmet. The device may use heart rate detection to identify characteristic patterns associated with epileptic seizure. The device optionally combines more than one measurement to eliminate false positives. In the case of epileptic seizures, heart rate related measurements may be combined with body motion related measurements to ensure greater accuracy. | 1. A portable device for predicting epileptic seizures comprising: at least one sensor (203, 204, 206-1, 206-2, 206-3, 207) for sensing a physical phenomenon in a patient's body, which phenomenon is known to be able to predict epileptic seizures, and for supplying signals characteristic of that phenomenon;at least one processor (201), coupled with the sensor (203, 204, 206-1, 206-2, 206-3, 207), for performing operations, the operations comprising processing the signals to determine whether they meet at least one criterion characteristic of epileptic seizures;at least one output device (205, 501, 903, 904, 905) adapted to supply an alarm indication, when the criterion is met; andat least one fastening apparatus for affixing the device to a patient's body. 2. The device of claim 1, wherein the fastening apparatus comprises a wrist band (101) carrying at least the processor (CE) and the output device (CE). 3. The device of claim 1, wherein the means for affixing comprises headgear (401) carrying the sensor, the processor and the output device. 4. The device of claim 1, wherein the phenomenon comprises a heart beat pattern (FIG. 3);the at least one criterion comprises at least one stored heartbeat pattern known to be predictive of epileptic seizures; andthe processing means uses an artificial intelligence algorithm to determine whether the signals match the stored pattern. 5. The device of claim 4, wherein the sensor further comprises at least one motion detection device (204, 206-1, 206-2, 206-3) suitable for detecting movement artifacts, acceleration, or both;the signals supplied by the sensor comprise signals relating to heartbeat pattern and one or more of movement artifacts and acceleration in at least one direction;the means for processing analyzes the signals in accordance with at least two criteria, one for each of the types of signals supplied by the sensors to create a combined analysis result; andthe alarm indication is supplied or not supplied responsive to the combined result. 6. The device of claim 1, wherein the sensor comprises a plurality of sensors including a heart beat detector 204, at least three accelerometers (206-1, 206-2, 206-3) for detecting motion in three dimensions, and an EEG device (207);the criterion comprises a plurality of criteria, including: at least one heart beat pattern associated with epileptic seizure;at least one first motion criterion designed to distinguish normal motion from seizure motion;at least one second motion criterion designed to correct heart beat measurements for movement artifacts; andat least one EEG criterion associated with epileptic seizure; andthe operations comprise considering signals from the plurality of sensors in view of the plurality of criteria in order to determine the presence or absence of epileptic seizure. 7. The device of claim 1 wherein the processing device is adapted to use the signals from the sensor for a second purpose separate from detecting seizures. 8. The dev... | KONINKLIJLE PHILIPS ELECTRONICS, N.V. | A61B005/00
US20080306395 | A1 | 2008-12-11 | Portable Electrocardiograph | 2007-02-08 | A portable electrocardiograph comprising a housing and a second electrode which is provided on one end of the housing, characterized in that a clamping cover which tends to close automatically is provided on the other end of the housing; an oval finger hole is provided between the clamping cover and the housing; the lower half of the inner wall of the finger hole is defined on the housing; the lower half of the inner wall is provided with a first electrode; and the upper half of the inner wall of the finger hole is provided on the clamping cover, such that the present invention provides the portable electrocardiograph of simple structure, convenient use, and high reliability for heart disease sufferers. | 1. A portable electrocardiograph comprising a housing (1) and a second electrode (10) which is provided on one end of the housing (1), characterized in that a clamping cover (11) which automatically tends to close is provided on the other end of the housing (1); an oval finger hole (6) is provided between the clamping cover (11) and the housing (1); a lower half (14) of the inner wall of the finger hole (6) is defined on the housing (1); the lower half (14) of the inner wall is provided with a first electrode (7); and an upper half (15) of the inner wall of the finger hole (6) is defined on the clamping cover (11). 2. The portable electrocardiograph according to claim 1, characterized in that a photoelectric receiving cell (8a) is provided in the inner wall of the lower half (14) of the finger hole (6); and a light emitting diode (8) is provided at the corresponding position of the inner wall of the upper half (15) of the finger hole (6). 3. The portable electrocardiograph according to claim 2, characterized in that the clamping cover (11) is closed with the housing (1) by means of a torsion spring (13). 4. The portable electrocardiograph according to claim 1, characterized in that the torsion spring (13) is sleeved on a shaft (12) which is connected with the housing (1). 5. The portable electrocardiograph according to claim 1, characterized in that the clamping cover (11) is closed with the housing (1) by means of a U-shaped spring sheet. 6. The portable electrocardiograph according to claim 1, characterized in that the first electrode (7) on the lower half (14) of the inner wall is of a half-ring shape. 7. The portable electrocardiograph according to claim 1, characterized in that a half-ring shaped soft material layer (9) is provided to cover the inner wall of the upper half (15). 8. The portable electrocardiograph according to claim 1, characterized in that the top face of the photoelectric receiving cell (8a) is provided to be at the same surface as the top face of the first electrode (7); and the top face of the light emitting diode (8) is slightly indented into the surface of the soft material layer (9). 9. The portable electrocardiograph according to claim 1, characterized in that the side face having the display of the housing is of a rectangle, trapezoid, or drum shape. 10. The portable electrocardiograph according to claim 1, characterized in that the second electrode is of straight line or curvilinear shape closely fit with human body. | Beijing Choice Eletronic Technology Co., Ltd. | A61B005/0404
US20080294041 | A1 | 2008-11-27 | Devices, Systems and Methods for Determining Sizes of Vessels | 2007-01-25 | Devices, systems and methods are disclosed for determining the cross sectional area of a vessel. Through a combination of fluid injection with different conductivities and measurement of the resultant conductances, parallel tissue conductance measure is obtained that assists in determining the cross sectional area, taking into account the presence of a stent. | 1. A device for determining a cross sectional size of a vessel, the device comprising: an elongated body having a longitudinal axis extending from a proximal end to a distal end, the body having a lumen along the longitudinal axis and enabling introduction of the distal end into a lumen of a vessel;a first excitation electrode and a second excitation electrode along the longitudinal axis, both located in respective grooves near the distal end; anda first detection electrode and a second detection electrode located in respective grooves along the longitudinal axis and in between the first and second excitation electrodes;wherein at least one of the first and second excitation electrodes is in communication with a current source, thereby enabling a supply of electrical current to the vessel, thereby enabling measurement of two or more conductance values in the blood vessel by the detection electrodes, and thereby enabling calculation of parallel tissue conductance in the vessel, whereby tissue conductance is the inverse of resistance to current flow, which depends on the cross sectional area of the blood vessel. 2. The device of claim 1, wherein the conductance measurement takes into account an offset created in measured conductance by the presence of a stent in the cross sectional area that is being measured. 3. The device of claim 1, further comprising: a data acquisition and processing system that receives conductance data from the detection electrodes and determines the conductance of the lumen. 4. The device of claim 1, further comprising: a suction/infusion port located near the distal end, wherein said suction/infusion port is in communication with said lumen, thereby enabling injection of two or more solutions into the lumen. 5. The device of claim 4, wherein the solution comprises an NaCl solution. 6. The device of claim 4, wherein the lumen is in communication with a source of a solution to be injected therethrough and through the suction/infusion port into the lumen. 7. A device for determining a cross sectional area of a vessel, the device comprising: an elongated body having a lumen therethrough along its longitudinal length;a pair of excitation electrodes located in respective grooves on the elongated body; anda pair of detection electrodes located in respective grooves located in between the pair of excitation electrodes such that a distance between one detection electrode and its adjacent excitation electrode is equal to the distance between the other detection electrode and its adjacent excitation electrode;wherein at least one excitation electrode is in communication with a current source, thereby enabling a supply of electrical current to a lumen of a vessel, and enabling measurement of two or more conductance values at the lumen by the detection electrodes, resulting in an assessment of the cross sectional area of the blood vessel. 8. A catheter for determining a cross sectional area of a vessel, the device comprising: an... | - | A61B005/053 A61B006/00
US20080319278 | A1 | 2008-12-25 | Sensor | 2005-09-08 | A physiological sensing device comprises, in combination a sensor (4) for the measurement of the partial pressure of carbon dioxide (pCO2), a body temperature sensor (5) and a heart rate and oxygen saturation sensor (54). The sensor device can be used to continuously monitor the vital signs of a patient. | 1. A physiological sensing device comprising in combination: a sensor for the measurement of the partial pressure of carbon dioxide (pCO2);a body temperature sensor;a heart rate sensor; and;an oxygen saturation sensor. 2. A sensing device as claimed in claim 1, wherein the pCO2 sensor is configured for insertion through a patient's skin. 3. A sensing device as claimed in claim 1, wherein the temperature sensor is configured for insertion through a patient's skin. 4. A sensing device as claimed in claim 1, wherein the temperature sensor and the pCO2 sensor are provided by a sensor unit for insertion through a patient's skin. 5. A sensing device as claimed in claim 2 wherein the device comprises a sharp tip for puncturing a patient's skin on insertion of the pCO2 sensor. 6. A physiological sensing device comprising a pCO2 sensor configured for insertion through a patient's skin anda sharp tip for puncturing a patient's skin on insertion of the pCO2 sensor. 7. A sensing device as claimed in claim 6, wherein the sharp tip is provided by a removable hollow needle in which the pCO2 sensor is located for insertion through a patient's skin. 8. A sensing device as claimed in claim 1, wherein the oxygen saturation sensor is configured for application to the surface of a patient's skin. 9. A sensing device as claimed in claim 8, wherein the heart rate sensor and the oxygen saturation sensor are provided by a pulse oxymetry sensor. 10. A sensing device as claimed in claim 1 comprising an adhesive patch for adhering the device to a patient's skin. 11. A physiological sensing device comprising a pCO2 sensor configured for insertion through a patient's skin andan ablest patch for adhering the device to a patient's skin to retain the inserted pCO2 sensor in position. 12. A sensing device as claimed in claim 1, wherein the pCO2 sensor comprises a chamber bounded, at least in part, by a carbon dioxide permeable membrane and containing a substantially electrolyte-free liquid and at least two electrodes. 13. A sensing device as claimed in claim 6, wherein the pCO2 sensor comprises a chamber bounded, at least in part, by a carbon dioxide permeable membrane and containing a substantially electrolyte-free liquid and at least two electrodes. 14. A sensing device as claimed in claim 11, wherein the pCO2 sensor comprises a chamber bounded, at least in part, by a carbon dioxide permeable membrane and containing a substantially electrolyte-free liquid and at least two electrodes. 15. A sensing device as claimed in claim 5, wherein the sharp tip is provided by a removable hollow needle in which the pCO2 sensor is located for insertion through a patient's skin. 16. A sensing device as claimed in claim 6 comprising an adhesive patch for adhering the device to a patient's skin. | ALERTIS MEDICAL AS | A61B005/02
US20080262385 | A1 | 2008-10-23 | Backflow Preventing Structure of a Blood Sampler, Luer Adapter, Blood Sampling Needle and Blood Sampling Holder | 2005-05-16 | A backflow preventing structure of a blood sampler capable of preventing backflow of blood positively without causing a significant increase in the number of components and cost. The backflow preventing structure of a blood sampler comprises a hub (3) as a tubular body having an internal channel (3a) extending from a first end part (3b) toward the second end part (3c) side, and a bottomed tubular backflow preventing member (6) which is composed of a material having rubber elasticity and contained in the internal channel (3a), has an opening opened toward the first end part (3b), and is closed on the second end part (3c) side. Outer circumferential wall of the backflow preventing member (6) is secured liquidtightly to the inner wall of the internal channel (3a), and a cut (6c) is made in the backflow preventing member (6) at a part closer to the second end part side than the secured part. | 1. A backflow preventing structure of a blood sampler, for preventing backflow of blood in the blood sampler, the structure comprising: a tubular body having an internal channel extending from a first end part toward a second end part through which blood flows; anda backflow preventing member contained in the internal channel of the tubular body, made of a material having rubber elasticity, the backflow preventing member being a bottomed tube closed at on the second end part side and having an opening that opens toward the first end part,wherein an outer circumferential wall of the backflow preventing member is secured liquid-tightly to an inner wall of the internal channel, anda cut is made in the backflow preventing member at a part closer to the second end part side than the secured part. 2. The backflow preventing structure of a blood sampler according to claim 1, wherein the cut is provided in at least two positions. 3. The backflow preventing structure of a blood sampler according to claim 1, wherein the cut formed in the backflow preventing member extends in a direction that intersects with a direction connecting the first end part and the second end part. 4. The backflow preventing structure of a blood sampler according to claim 3, wherein the cut formed in the backflow preventing member extends in a direction inclined at an angle ranging from 15 to 165 degrees with respect to the direction connecting the first end part and the second end part. 5. The backflow preventing structure of a blood sampler according to claim 3, wherein the cut formed in the backflow preventing member extends in a direction perpendicular to the direction connecting the first end part and the second end part. 6. The backflow preventing structure of a blood sampler according to claim 1, wherein in the backflow preventing member, a part that serves as a hinge for opening/closing achieved by the cut has a relatively smaller thickness than the remaining part. 7. A luer adapter comprising: a hollow needle used for piercing a plug member of a blood sampling tube in blood sampling, anda hub secured to one end of the hollow needle,wherein the hub has the backflow preventing structure of a blood sampler according to claim 1, and the hollow needle is in communication with the second end part side of the internal channel of the backflow preventing structure of a blood sampler. 8. A blood sampling needle comprising: a hollow needle for collecting blood; anda hub secured to one end of the hollow needle,wherein the hub has the backflow preventing structure of a blood sampler according to claim 1, and the hollow needle is in communication with the first end part side of the internal channel of the backflow preventing structure of a blood sampler. 9. A blood sampling holder comprising: a tubular holder main body formed with an opening for inserting a blood sampling tube at its one end, and closed at the other end;a hub provided on the other end side of the holder main bod... | - | A61B005/153
US20080255433 | A1 | 2008-10-16 | OPTOACOUSTIC MONITORING OF MULTIPLE PARAMETERS | 2008-04-11 | An optoacoustic technique for absolute, accurate, continuous, and real-time measurement of a variety of important diagnostic variables is disclosed, where the variables include: (1) noninvasive measurements of circulating blood volume (BV) and cardiac output (CO); (2) calculation from the measured variables of cardiac index (CI) and systemic oxygen delivery (DO2); (3) concentrations of hemoglobin derivatives (e.g., carboxyhemoglobin [HbCO], reduced hemoglobin [Hb], oxygenated hemoglobin [HbOxy], and methemoglobin [HbMet]), total hemoglobin concentration [THb], concentrations of lactate, myoglobin, cholesterol, body pigments, and exogenous dyes; (4) content in tissues of water, fat, protein, calcium, and blood; as well as density of hard and soft tissues; and (5) accurate noninvasive measurement of blood pressure (or vascular pressure) in arteries, arterioles, veins, capillaries, using occlusion-induced changes in optoacoustic signal induced in blood circulating in the vessels. The optoacoustic technique can be used for single measurement, continuous measurement, or continuous monitoring of these variables. | 1. An optoacoustic system for absolute or relative, accurate, and real-time or near real-time measuring or monitoring of a variety of at least one diagnostic variable comprising: a pulsed light source,a light source delivery subsystem comprising one optical fiber or a plurality of optical fibers or a light delivery subsystem comprising conventional (non-fiber) optical elements,a probe including a transducer for detection of acoustic waves generated in a site or tissue of a patient due to the optical pulses generated by the light source,a registration subsystem to acquire, analyze, and process data from the transducer and to display or send data containing at least one absolute or relative value of the diagnostic variable or changes in the value of the diagnostic variable. 2. The system of claim 1, wherein the registration system comprises: a low-noise amplifier is connected to the transducer via a signal conduit,a digitizer is connected to the amplifier via a connection,a laptop computer is connected to the digitizer via a connection and to the light source via a connection, where the computer controls the light source and acquires, analyzes and processes data from the transducer via and amplifier and digitizer. 3. The system of claim 1, wherein the light source is capable of generating at least one pulse with duration in the range between about 1 femtosecond and about 10 microseconds and with at least one wavelength in the range between about 200 nm and about 20,000 nm; and preferably with duration in the range between about 0.1 nanoseconds and about 200 nanoseconds and with wavelength in the range between about 600 nm and 2,000 nm. 4. The system of claim 1, wherein the fiber comprises or fibers comprise multi-mode optical fibers. 5. The system of claim 1, wherein the fiber has or fibers have diameters in the range between about 10 microns to about 2 mm. 6. The system of claim 1, wherein the fiber has or fibers have diameters in the range between about 100 microns and about 1 mm. 7. The system of claim 1, wherein an array of the transducers is used for more accurate monitoring. 8. The system of claim 1, wherein the transducer is a broad-band sensitive focused or unfocused transducer containing at least one conventional piezo-element or at least one peizo-element with low planar coupling coefficient or containing a subsystem for optical detection of acoustic waves. 9. The system of claim 1, wherein the transducer has a bandwidth in the range between about 0.01 MHz to about 100 MHz; and preferably in the range between about 0.05 MHz and about 10 MHz. 10. The system of claim 1, where in the probe is scanned over tissue surface in at least one direction or in at least one angular direction, or in mixtures or combinations thereof to provide more accurate monitoring. 11. The system of claim 1, where in the probe is inserted in a hollow organ to better access blood vessel or tissue of interest. 12. The system of claim 1, wherein the variable... | The Board of Regents of The University of Texas Syatem | A61B005/145
US20080275318 | A1 | 2008-11-06 | BIOSENSORS FOR MEASURING ANALYTES IN THE INTERSTITIAL FLUID | 2008-04-21 | The invention relates to methods and devices for measuring blood glucose levels in a subject, where the methods and devices are designed for exposing a sensing mechanism to interstitial fluid in the subject. | 1. A method for determining the blood levels of glucose in a subject, the method comprising a) inserting in the skin of the subject to a depth of up to about 2 mm from the surface of the skin, a sensor that produces a detectable signal in the presence of glucose in the interstitial fluid of the subject,b) maintaining said sensor in the skin at said depth such that at least a portion of said sensor is exposed to the interstitial fluid in the skin,c) measuring the detectable signal of the sensor produced by said sensor upon exposure to said interstitial fluid, andd) correlating the signal produced by the sensor to the level of blood glucose, thereby determining the levels of blood glucose in the subject. 2. The method of claim 1, wherein the sensor comprises a glucose-galactose binding protein (GGBP). 3. The method of claim 2, wherein the GGBP is labeled with a signaling moiety that emits an optical signal. 4. The method of claim 1, wherein a needle houses the sensor, said needle comprising a body with a proximal end and a distal end, wherein the proximal end of the needle comprises a tip, said needle being a 31 gauge needle or smaller. 5. The method of claim 1, wherein the sensor is inserted into the subject at a depth of between about 0.1 mm and 2 mm below the surface the skin. 6. The method of claim 5, wherein the sensor is inserted into the subject at a depth of between about 0.1 mm and 1 mm below the surface the skin. 7. The method of claim 6, wherein the sensor is inserted into the subject at a depth of between about 0.1 mm and 0.8 mm below the surface the skin. 8. The method of claim 1, wherein the sensor remains inserted in the subject for at least one day. 9. The method of claim 1, wherein the sensor is a continuous sensor. 10. The method of claim 1, wherein the sensor produces an optical signal in the presence of glucose for at least 4 hours. 11. The method of claim 1, wherein the subject is under a physiological stress that results in decreased peripheral blood flow. 12. A biosensor for detecting blood glucose levels in a subject, the biosensor comprising a) a sensing mechanism that produces an optical signal in the presence of glucose said sensing mechanism comprising a fluorescently labeled glucose-galactose binding protein (GGBP) entrapped within a hydrogel matrix,b) a shallow-depth housing that encases the sensing mechanism, wherein the shallow-depth housing comprises a hollow needle, wherein said hollow needle is a 31 gauge needle or smaller and penetrates the skin of a subject to a depth of up to about 2 mm from the surface of the skin such that said sensing mechanism is exposed to interstitial fluid, andc) an optical conduit, wherein said optical conduit is in optical connectivity with said sensing mechanism and the needle houses at least a portion of said conduitwherein the optical signal of the sensing mechanism generated from exposure to the interstitial fluid of the subject is correlated to blood glucose levels. 13... | Becton, Dickinson and Company | A61B005/1459 A61B005/145
US20080300478 | A1 | 2008-12-04 | SYSTEM AND METHOD FOR DISPLAYING REAL-TIME STATE OF IMAGED ANATOMY DURING A SURGICAL PROCEDURE | 2007-05-30 | A system and method for displaying surgical instruments accurately, in both time and space, within both slice and volumetric medical images of organs that may deform in a predictable manner over time due to an ongoing body activity. The system and method comprising using a gating signal associated with an ongoing body activity to determine which image in a sequence of acquired images best represents the state of the imaged anatomy at a given point in time and accurately displaying navigated surgical instruments within that image at that point in time. | 1. A method of displaying real-time state of imaged anatomy during a surgical procedure comprising the steps of: (a) attaching a measurement device to a patient to measure and ongoing body activity and generating a gating signal corresponding to the ongoing body activity;(b) establishing a navigation reference frame around the patient in a surgical field of interest;(c) acquiring a time series of images of a patient's anatomy in the surgical field of interest in sync with the gating signal;(d) determining an image position for each image in the time series in relation to the navigation reference frame;(e) determining a position for at least one navigated surgical instrument used during the surgical procedure in relation to the navigation reference frame; and(f) displaying the time series of images along with the at least one navigated surgical instrument used during the surgical procedure superimposed on the time series of images showing the real-time state of the imaged anatomy and the accurate positions of the at least one navigated surgical instrument within the real-time state of imaged anatomy. 2. The method of claim 1, wherein the patient's anatomy deforms in a predictable manner over time due to the ongoing body activity. 3. The method of claim 1, wherein the gating signal is a respiration cycle measurement signal of the patient. 4. The method of claim 1, wherein the gating signal is an EKG measurement signal of the patient. 5. The method of claim 1, wherein the step of acquiring a time series of images includes acquiring pre-operative or intraoperative images. 6. The method of claim 1, further comprising the step of continuously updating the display of the time series of images along with the navigated surgical instruments superimposed on the time series of images in sync with the gating signal. 7. A method of displaying real-time state of imaged anatomy during a surgical procedure comprising the steps of: (a) attaching a measurement device to a patient to measure and ongoing body activity and generating a gating signal corresponding to the ongoing body activity;(b) acquiring a time series of pre-operative 3D images of a patient's anatomy in a surgical field of interest in sync with the gating signal;(c) establishing a navigation reference frame around the patient in the surgical field of interest;(d) acquiring a time series of intraoperative 2D images of the patient's anatomy in the surgical field of interest in sync with the gating signal;(e) reconstructing a time series of intraoperative 3D images from the time series of intraoperative 2D images;(f) registering the time series of pre-operative 3D images with the time series of intraoperative 3D images; and(g) displaying the time series of pre-operative 3D images along with at least one navigated surgical instrument used during the surgical procedure superimposed on the time series of pre-operative 3D images showing the real-time state of the imaged anatomy and the accurate positi... | GENERAL ELECTRIC COMPANY | A61B005/00
US20080281187 | A1 | 2008-11-13 | Ferromagnetic threat detection method apparatus | 2007-10-17 | A method and apparatus for detecting ferromagnetic threat objects on a recumbent patient, including magnetizing/sensing stations having mutually orthogonal magnetizing axes. Two magnetizing/sensing stations can be used, but three or more are preferred. The magnetizing/sensing stations are arranged on a table-like mounting structure providing a path for the patient to roll beneath the magnetizing/sensing stations on a gurney. Three additional magnetizing/sensing stations can be provided on either side, or on each side, of the gurney path, at the same height as the patient. | 1. An apparatus for detecting a ferromagnetic threat object, comprising: a mounting structure, said mounting structure having a gurney path along which a recumbent patient on a gurney can pass;a plurality of detection stations on said mounting structure, said detection stations being arranged sequentially along said gurney path;a magnetic field source in each said detection station; andat least one magnetic sensor in each said detection station, said at least one sensor being adapted to sense the magnetization of a ferromagnetic threat object in the field established by an associated said magnetic field source in said detection station;wherein the primary magnetic axis of each said magnetic field source is orthogonal to the primary magnetic axis of at least one other said magnetic field source. 2. The apparatus recited in claim 1, wherein said plurality of detection stations are limited to only two of said detection stations. 3. The apparatus recited in claim 1, further comprising three of said detection stations, each said detection station having the primary magnetic axis of its said magnetic field source oriented orthogonal to the primary magnetic axes of said magnetic field sources at each of the other said detection stations. 4. The apparatus recited in claim 1, wherein said detection stations are arranged above the path of said patient on said gurney passing along said gurney path. 5. The apparatus recited in claim 4, further comprising a second plurality of detection stations on said mounting structure, said second plurality of detection stations being arranged sequentially alongside said gurney path, said second plurality of detection stations being positioned at substantially the same height as said patient on said gurney. 6. The apparatus recited in claim 5, wherein said second plurality of detection stations are mounted along one side of said gurney path. 7. The apparatus recited in claim 5, wherein said second plurality of detection stations are mounted along both sides of said gurney path. 8. A method for detecting a ferromagnetic threat object, comprising: providing a plurality of detection stations on a mounting structure, each said detection station having a magnetic field source and an associated magnetic sensor, wherein the primary magnetic axis of each said magnetic field source is orthogonal to the primary magnetic axis of at least one other said magnetic field source;passing a recumbent patient on a gurney along a path, sequentially bringing said patient into close proximity with each of said detection stations;magnetizing any ferromagnetic threat object present on the patient with a magnetic field established by one of said magnetic field sources; andsensing, with said magnetic sensor associated with said one magnetic field source, the magnetization of said ferromagnetic threat object in said magnetic field. 9. The method recited in claim 8, further comprising magnetizing said ferromagnetic threat object with an off-... | MedNovus, Inc. | A61B005/05
US20080281167 | A1 | 2008-11-13 | DIAGNOSTIC INSTRUMENT WORKSTATION | 2008-05-22 | An integrated medical workstation for use in patient clinical encounters includes an input device such as a bar code scanner that is interconnected to a computing device. At least one device capable of obtaining at least one physiological parameter is either attached directly to the workstation or is in communication therewith. Preferably, the input scanning device controls at least substantial overall operation of the medical workstation that can be placed, for example, into a network. | 1. An integrated apparatus for use in a patient encounter, said apparatus comprising: an input device having means for reading machine-readable information;a computing device connected to at least one medical instrument and said input device; andcontrol means for controlling the operation of said at least one medical instrument, said input device, and said computing device wherein said apparatus is substantially and automatically controlled using said input device by scanning appropriate machine-readable information portions, said portions including instructions that are interpreted by and executed by said computing device to control operation of said apparatus. 2. An integrated apparatus according to claim 1, further including a miniature imaging device for selectively capturing images during a patient encounter. 3. An integrated apparatus according to claim 2, wherein said miniature imaging device is provided in said input device. 4. An integrated apparatus according to claim 1, wherein said input device is a bar-code scanner. 5. An integrated apparatus according to claim 1, including a printer connected to said computing device. 6. An integrated apparatus according to claim 1, wherein said control means further includes user-actuable controls for controlling the operation of the apparatus. 7. An integrated apparatus according to claim 2, wherein said computing device includes data storage means for selectively storing at least one image captured by said miniature imaging device and data from said at least one of said input device and said at least one medical instrument. 8. An integrated apparatus according to claim 2, wherein said input device includes means for identifying at least one machine readable portion in at least one image that is captured by said miniature imager device. 9. An integrated apparatus according to claim 8, wherein said computing device includes means for decoding said at least one machine readable portion, if said at least one portion is identified in a captured image. 10. An integrated apparatus according to claim 1, including encryption means for preventing unauthorized operation of said apparatus. 11. An integrated apparatus according to claim 1, wherein scanning of an appropriate machine-readable portion by said input device automatically causes activation of said at least one medical instrument. 12. An integrated apparatus according to claim 1, wherein at least one said machine-readable information portion is presented on a display of said apparatus. 13. An integrated apparatus according to claim 1, wherein at least one said machine-readable information portion is on a patient record sheet. 14. An integrated apparatus according to claim 13, wherein said patient record sheet further includes at least one machine-readable portion that includes patient-related data, wherein the patient-related data can be uploaded into the data storage means of said computing device and in which additional patient data c... | Welch Allyn, Inc. | A61B005/00
US20080298614 | A1 | 2008-12-04 | System for and Method of Offering an Optimized Sound Service to Individuals within a Place of Business | 2005-06-09 | The present invention relates to optimizing sound for a person based on his or her location and hearing profile. More particularly, the present invention relates to personalizing the sound of an event or performance by examining the technical specifications of an individual's hearing aid (which is provided by a third party) and the acoustics of an establishment. Using this information, a place of business alters the settings of each hearing aid device to optimize the sound for each individual. | 1. A method for optimizing sound service to an individual customized to a specific location, comprising the steps of: providing acoustic characterization data corresponding to a location,providing access to a database of hearing loss profiles associated with a respective plurality of individuals using hearing aids, wherein the hearing loss profiles include respective digital sound processor (“DSP”) correction factors for use by a DSP of a hearing aid; andcomputing an adjusted DSP correction factor for a user based on the acoustic characterization data corresponding to the location. 2. The method of claim 1, further comprising transmitting the adjusted DSP correction factor as a DSP programming signal for programming a DSP of a hearing aid, wherein the DSP of the hearing aid modifies input audio signals using the adjusted DSP correction factors. 3. The method of claim 1, wherein the location comprises a plurality of regions having distinct acoustic properties and the acoustic characterization data comprises regional data corresponding thereto, and further comprising computing the adjusted DSP correction factor based on characterization data associated with a particular region within the location. 4. The method of claim 1, wherein at least one of the acoustic characterization data and the hearing loss profile database is accessible over a communications network. 5. The method of claim 3, further comprising determining a favored acoustic region for an individual within a location based on the regional data and the hearing loss profile of the individual. 6. The method of claim 1, wherein the acoustic characterization data and the hearing loss profile database are accessible over a communications network, and wherein the location comprises a plurality of regions having distinct acoustic properties and the acoustic characterization data comprises regional data corresponding thereto, the location further comprising a connection interface at each of said regions for allowing a data connection between the hearing aid and the communications network, and further comprising connecting the hearing aid to the communications network when the hearing aid is located within a particular region within the location, computing the adjusted DSP correction factor based on characterization data associated with that region and transmitting the adjusted DSP correction factor as a DSP programming signal for programming the DSP of the hearing aid, wherein the DSP of the hearing aid modifies input audio signals using the adjusted DSP correction factors. 7. The method of claim 1, further comprising the step of processing a user fee prior to prior to performing one or more of the steps. 8. A system for optimizing sound service to an individual customized to a specific location, comprising: acoustic characterization data corresponding to a location,a database of hearing loss profiles associated with a respective plurality of individuals using hearing aids, wherein the he... | Johnson & Johnson Consumer Companies, Inc. | H04R025/00 A61B005/00 H04R029/00
US07451002 | B2 | 2008-11-11 | Automated generation of transfer functions based upon machine data | 2006-01-06 | A system and method for servicing a medical device, which provides for generation of a transfer function that correlates historical machine data with the health of the medical device. The transfer function may be validated and stored. The transfer function is automatically updated based on current machine data. | 1. A method of maintaining a medical device, comprising: generating a transfer function correlating historical machine data with a health of the medical device;validating and storing the transfer function;receiving current machine data of the medical device in substantially real time; andautomatically updating the transfer function based on the current machine data. 2. The method according to claim 1, wherein automatically updating comprises updating the transfer function substantially contemporaneously with receipt of the current machine data. 3. The method according to claim 1, wherein automatically updating comprising updating the transfer function during monitoring and diagnosis of the medical device. 4. The method according to claim 1, comprising detecting an anomaly of the medical device via the transfer function. 5. The method according to claim 4, comprising implementing corrective action in response to the detection of the anomaly. 6. A method of monitoring a medical system, comprising: storing machine parameter data of the medical system;generating a transfer function that correlates a health of the medical system with the stored machine parameter data;storing the transfer function; andautomatically updating the transfer function based on current machine parameter data. 7. The method according to claim 6, comprising storing and updating the transfer function at a location remote from the medical system. 8. The method according to claim 6, comprising storing or updating the transfer function on a portable computer system. 9. The method according to claim 6, comprising servicing the medical system in response to diagnosis of the medical system via the updated transfer function. 10. A method of servicing a medical device, comprising: storing a model having a transfer function that correlates a health of the medical device with one or more machine parameters of the medical device;receiving machine data of the medical device; andautomatically updating the model based on the machine data. 11. The method according to claim 10, wherein storing comprises storing the model on the medical device or in a remote service center, or a combination thereof. 12. The method according to claim 10, comprising diagnosing the health of the medical device via the updated model. 13. The method according to claim 12, wherein diagnosing comprises diagnosing the health of the device locally or remotely, or a combination thereof. 14. The method according to claim 12, comprising implementing corrective action in response to the diagnosis. 15. The method according to claim 12 wherein the medical device comprises a medical imaging system. 16. The method according to claim 12, wherein the medical device comprises a magnetic resonance imaging system (MRI) and the one or more machine parameters comprise magnet center frequency. 17. The method according to claim 16, wherein the device health comprises a gradient of magnet shim. 18. The method according to... | GE Medical Systems Global Technology Company, LLC | G05B011/01 G05B009/02 G01D003/00 G01M019/00 G01P021/00 G01R035/00 G06F019/00 G06F011/00 G01C025/00 A61B005/00 A61B019/00 G08B029/00 A61N001/00
US07441950 | B2 | 2008-10-28 | Probe for electronic clinical thermometer | 2006-03-15 | Providing a probe for an electronic clinical thermometer, which can measure a surface temperature and a deep body temperature of a human body securely in short time, the probe includes a cylindrical housing, a bottomed metal pipe fitted into a top end of the cylindrical housing, and a substrate having two thin-film heat-sensitive elements arranged perpendicularly on an inner wall of the bottomed metal pipe so as to fix one side edge of the substrate tightly on the inner wall. By measuring heat flux between the two thin-film heat-sensitive elements, the surface and the deep body temperatures of the human body can be predicted by applying the measured values into a heat conduction equation without waiting the sensor to reach in heat balance. | 1. A probe for an electronic clinical thermometer, comprising: a cylindrical housing;a metal pipe having a closed end and being fitted with an end of said cylindrical housing; anda temperature sensor being inserted and fixed in said metal pipe for measuring heat flux according to heat flow between a human body and the temperature sensor to predict a body temperature;wherein said temperature sensor includes a substrate and two thin-film heat-sensitive elements being formed on the substrate along the heat flow from a side edge of the substrate,wherein said substrate is arranged perpendicularly on an inner wall of the metal pipe so as to fix the side edge of the substrate tightly on the inner wall for thermally coupling the inner wall of the metal pipe and the substrate. 2. The probe for an electronic clinical thermometer according to claim 1, wherein the side edge of said substrate is fixed on the inner wall of the metal pipe with one of resin and adhesive. 3. The probe for an electronic clinical thermometer according to claim 1, further comprising a long strip printed wiring board having a wiring pattern to form lands at an end thereof and lead terminals, said printed wiring board being inserted in the metal pipe, wherein the temperature sensor is electrically connected with the lands at the end of the printed wiring board so as to fix the temperature sensor tightly on the inner wall of the metal pipe. 4. The probe for an electronic clinical thermometer according to claim 3, wherein the side edge of said substrate is fixed on the inner wall of the metal pipe with one of resin and adhesive. 5. The probe for an electronic clinical thermometer according to claim 1, wherein the two thin-film heat-sensitive elements formed on the substrate, which structures the temperature sensor, fixed on the inner wall of the metal pipe are arranged respectively at an upstream side and a downstream side of the heat flow. 6. The probe for an electronic clinical thermometer according to claim 5, further comprising a long strip printed wiring board having a wiring pattern to form lands at an end thereof and lead terminals, said printed wiring board being inserted in the metal pipe, wherein the temperature sensor is electrically connected with the lands at the end of the printed wiring board so as to fix the temperature sensor tightly on the inner wall of the metal pipe. 7. The probe for an electronic clinical thermometer according to claim 5, wherein the side edge of said substrate is fixed on the inner wall of the metal pipe with one of resin and adhesive. | Ishizuka Electronics, Corp. | G01K007/00 A61B005/01
US07450724 | B1 | 2008-11-11 | Hearing aid adjustment device | 1999-08-27 | A hearing device fitting device that includes a computing device (3), connected on an input side with a connection (E3) for data entry and on an output side with a connection (A3) for a hearing device. The hearing device fitting device further includes an audio storage medium play-back unit having a control input (E9) connected to a computing device (3) output and having an audio output (A3) connectable to a loud speaker unit (11) input. The device supports a method a method for fitting the hearing device in situ by applying the hearing device to an individual; subjecting the individual to an audio test signal; having the individual appraise said audio test signal; and automatically selecting, in dependency of said appraising, a subsequent audio test signal. | 1. An apparatus for fitting a hearing device which is worn by an individual, said apparatus comprising: a data entry device;a computing device connected on an input side with a connection for connecting to said data entry device and on an output side with a connection for a hearing device adjusting input,an audio storage medium play-back unit storing a plurality of audio tracks with audio test signals and having a control input connected to another output of said computing device and having an audio output connectable to a loud speaker unit input, anda storage device for storing a plurality of individual assessment data and audio test signals experienced by the individual, whereinsaid individual assessment data is entered into said data entry device based on perceptions of said individual wearing said hearing device and listening to one of said audio tracks, said individual assessment data and said audio test signals experienced by the individual being stored in said storage device, and further whereinsaid computing device computes a control signal based on said stored plurality of individual assessment data and said stored audio test signals experienced by the individual, wherein said control signal is applied to said other output of said computer device, and said control signal is used for automatically selecting another one of said audio tracks. 2. The apparatus for fitting a hearing device according to claim 1, wherein said playback unit contains at least one audio storage chip. 3. The apparatus for fitting a hearing device according to claim 1, wherein said playback unit is a CD playback unit. 4. The apparatus for fitting a hearing device according to claim 1, further comprising a comparer unit that tests the audio storage medium in said playback unit for a predetermined identification and which disables said playback unit on non-recognition of said predetermined identification. 5. The apparatus for fitting a hearing device according to claim 1, further comprising a decoding unit, wherein said playback unit is an audio CD playback unit generating a specification of an extent of at least one of the segments on the audio storage medium in said playback unit, and wherein said specification is fed from an output of said playback unit to said decoding unit which then generates a control signal for the operation of said playback unit. 6. The apparatus for fitting a hearing device according to claim 1, wherein said fitting device further comprises a set-level comparative unit having an output operationally connected to a level control input of said playback unit for controlling said audio output, wherein the hearing device is connected to said hearing device output, the hearing device having a level detector which is connected to an acoustical/electrical converter of the hearing device, such that said computing device generates, on a level detector control output, a level detector control signal for controlling an operational connection betwee... | Phonak AG | H04R029/00 H04R025/00 A61B005/00 A61B005/12 A61N001/00
US07470235 | B2 | 2008-12-30 | Pulse wave detecting device and method therefor | 2005-09-23 | A pulse wave measuring apparatus can determine an optimum combination of a pair of light-emitting element and light-receiving element out of plural light-emitting elements and plural light-receiving elements on a real time basis according to a difference of signal intensity of a pulse wave signal from the light-receiving element and always measure an stable pulse wave of a wrist artery robustly. | 1. A measuring apparatus for detecting a pulse wave signal indicating a change in a blood flow in the blood vessel of a patient by using light, comprising: a sensor module including a plurality of light-emitting elements to irradiate the blood vessel, and a plurality of light-receiving elements to receive reflected light as a pulse wave signal from the blood vessel, each element being configured to be attached on the surface of the patient;a processor for light emission that causes the plurality of light-emitting elements to emit light one after another;an autocorrelation value calculating processor that calculates autocorrelation values of respective pulse wave signals corresponding to respective combinations of the light-emitting elements, which have emitted light, and the light-receiving elements, which have received light, respectively; andan optimum position identifying processor that selects a combination of a light-emitting element and a light-receiving element, which has outputted a pulse wave signal with a highest autocorrelation value among the respective autocorrelation values, as an optimum combination,wherein the autocorrelation value calculating processor slices waveforms of pulse wave signals in an arbitrary period from the pulse wave signals, respectively, and obtains an autocorrelation between the sliced waveforms and the pulse wave signals to thereby calculate autocorrelation values of the respective pulse wave signals. 2. A pulse wave measuring apparatus according to claim 1, wherein the arbitrary period includes a pulse wave of at least one beat. 3. A measuring apparatus for detecting a pulse wave signal indicating a change in a blood flow in the blood vessel of a patient by using light, comprising a sensor module including a plurality of light-emitting elements to irradiate the blood vessel, and a plurality of light-receiving elements to receive reflected light as a pulse wave signal from the blood vessel, each element being configured to be attached on a surface of the patient;a processor for light emission that causes the plurality of light-emitting elements to emit light one after another;an autocorrelation value calculating processor that calculates autocorrelation values of respective pulse wave signals corresponding to respective combinations of the lite-emitting elements, which have emitted light, and the light-receiving elements, which have received light, respectively; andan optimum position identifying processor that selects a combination for a light-emitting element and a light-receiving element, which has outputted a pulse wave signal with a highest autocorrelation value among the respective autocorrelation values, an optimum combination,wherein, in selecting an optimum combination again after an optimum combination of a light-emitting element and a light-receiving element is determined by the optimum position identifying processor, the processor for light emission causes a light-emitting element near the light-... | Kabushiki Kaisha Toshiba | A61B005/02
US07811236 | B2 | 2010-10-12 | Puncturing system for collecting body fluid sample | 2008-05-19 | A microsampler puncturing system for collecting a body fluid from a body part is described, comprising a sample collection unit having a piercing element, and a puncturing instrument having a drive, by which a sample collection unit is movable on a movement path for piercing the piercing element into the skin of the body part and withdrawing it in a piercing and retraction movement, a setting device for setting the puncturing depth of the piercing wound to be generated, and a control device for controlling the piercing and retraction movement having the following sequentially executed movement phases: a forward phase, a retraction phase, and a collecting phase (S). The setting device is settable independently of the chronological mean of the defined residual puncturing depth (dr), with reference to the collecting phase. | 1. A microsampler puncturing system for collecting a body fluid sample from a body part, comprising: a sample collection unit having a piercing element, anda puncturing instrument having: a drive, wherein the sample collection unit is movable on a movement path in a piercing and retraction movement for piercing the piercing element into skin of the body part, creating a piercing wound and withdrawing the piercing element thereafter,a setting device configured to set a defined puncturing depth of the piercing wound, anda control device configured to control the piercing and retraction movement, the piercing and retraction movement including the following sequentially executed movement phases: a forward phase, in which the piercing element is moved in a puncturing direction and pierced to the defined puncturing depth,a retraction phase, in which the piercing element is partially withdrawn by a retraction distance and is decelerated toward an end of the retraction phase, in such a manner that the piercing element projects into the skin to a residual puncturing depth, anda collecting phase, in which a body fluid sample is withdrawn by the sample collection unit, via the piercing element projecting into the skin and being stationary, such that the residual puncturing depth is constant, or moving opposite to the puncturing direction,wherein a mean value of the residual puncturing depth corresponds to the time integral of the residual puncturing depth from the beginning of the collecting phase to the end of the collecting phase divided by the duration of the collecting phase,the setting device is adapted to set the defined puncturing depth independently of the mean value of the residual puncturing depth andthe setting device is adapted to set the mean value of the residual puncturing depth independently of the defined puncturing depth. 2. The puncturing system according claim 1, wherein the control device is adapted to control, as a further movement phase, a further retraction phase, in which the piercing element is accelerated again and withdrawn completely from the skin. 3. The puncturing system according to claim 1, wherein the puncturing instrument comprises a retraction stop, wherein the piercing element is stopped at the residual puncturing depth at the end of the retraction phase. 4. The puncturing system according to claim 3, wherein the retraction stop comprises a spring element, wherein an abutting element of the retraction stop is moved into a blockade position to stop the piercing element. 5. The puncturing system according to claim 1, further comprising a reference element which is adapted for contacting the skin surface to ensure a reproducible defined puncturing depth during piercing, the reference element being movable in relation to the piercing element and in relation to a device housing. 6. The puncturing system according to claim 5, wherein the piercing element is moved together with the reference element in the puncturing direc... | Roche Diagnostics Operations, Inc. | A61B005/00 A61B005/15 A61B005/157 A61B017/32
US07685728 | B2 | 2010-03-30 | Method and apparatus for determining comfortable footbed shapes | 2008-03-28 | A method and apparatus for determining the comfortable footbed shape for a person. The apparatus comprises a rear platform for supporting the rear of a foot, a front platform for supporting a forefoot, the height of at least one of said front and rear platforms being adjustable relative to the height of the other one of said platforms and fixable at a selected height relative to the other platform; and a supporting surface for accommodating the plantar surface of a foot, the supporting surface comprising a length of flexible material fixable mounted to and extending between said front and rear platforms. The apparatus can adjust the aforementioned components to various settings with a view to determining the plantar foot shape and the perceived comfort in any chosen posture. | 1. An apparatus comprising; a rear platform for supporting the rear of a foot;a front platform for supporting a forefoot;the height of at least one of said front and rear platforms being adjustable relative to the height of the other one of said platforms and fixable at a selected height relative to the other platform;and a supporting surface for accommodating the plantar surface of a foot, the supporting surface comprising a length of flexible material mounted to and extending between said front and rear platforms. 2. The apparatus of claim 1 further comprising an adjustable middle foot supporting member, said middle foot supporting member being arranged to contact a portion of the supporting surface. 3. The apparatus of claim 2 wherein the height of the middle foot supporting member is adjustable. 4. The apparatus of claim 2 wherein an angle of the middle foot supporting member is adjustable. 5. The apparatus of claim 2 wherein the lateral distance between the middle foot supporting member and the front or rear platform is adjustable. 6. The apparatus of claim 1 wherein an angle of inclination of the rear platform is adjustable. 7. The apparatus of claim 6 wherein the rear platform is rotatable about an axis and its angle of inclination may be varied by one or more adjustment members which support the platform. 8. The apparatus of claim 6 wherein the rear platform has a replaceable foot supporting portion, which can be replaced by one or more other foot supporting portions having different angles of inclination. 9. The apparatus of claim 1 wherein an angle of inclination of the front platform is adjustable. 10. The apparatus of claim 1 wherein the relative height of the front and rear platforms is adjustable continuously. 11. The apparatus of claim 1 wherein the relative height of the front and rear platforms is adjustable by one or more electrical actuators. 12. The apparatus of claim 1 having a plurality of replaceable supporting surfaces, each of which is mountable to the front and rear platforms and each of which has different characteristics. 13. The apparatus of claim 1 wherein the supporting surface is formed of a plurality of portions having different characteristics from each other; said characteristics being selected from the group comprising stiffnesses and resilience properties. 14. The apparatus of claim 1 further comprising an optical device or touch probe for obtaining a shape of a foot supported by the support surface. 15. A method of determining the comfortable plantar shape of a foot comprising the steps of: a. placing the rear of a foot on a rear platform, the forefoot on a front platform and resting the middle of the foot on a supporting surface comprising a strip of material extending between and mounted to the front and rear platformsb. adjusting the height and/or an angle of inclination of at least one of the front and rear platformsc. determining when the foot is comfortable. 16. An apparatus for measuring... | The Hong Kong University of Science and Technology | A61B005/103
US07445602 | B2 | 2008-11-04 | Carbon dioxide sensor and airway adapter incorporated in the same | 2006-06-29 | In a sensor for detecting a carbon dioxide gas in an expiration gas of a living body, an airway case is adapted to be disposed below nostrils of the living body, and formed with an airway passage extending across an optical axis of a light beam emitted from a light emitter of the sensor. A mouth guide is adapted to be disposed in front of a mouth of the living body so as to define a space communicated with the airway passage. The mouth guide is pivotably supported on the airway case. A retainer is adapted to retain an oxygen supply tube on the airway adapter body in such an attitude that an oxygen gas supplied from prongs of the oxygen supply tube is not directly injected into the nostrils. | 1. An airway adapter of a sensor for detecting a carbon dioxide gas in an expiration gas of a living body, the airway adapter comprising: an airway case, adapted to be disposed below nostrils of the living body, and formed with an airway passage extending across an optical axis of a light beam emitted from a light emitter of the sensor, the airway case having a front side and a back side that opposes the front side, the back side positioned adjacent a face of the living body when the airway case is positioned on the face of the living body;an oxygen supply tube; anda retainer that retains the oxygen supply tube on the front side of the airway case such that an exit of the oxygen supply tube is adjacent a side of the airway case other than the back side of the airway case, when the back side of the airway case is adjacent the face of the living body, whereby an oxygen gas supplied from the exit of the oxygen supply tube is not directly injected into the nostrils. 2. The airway adapter as set forth in claim 1, wherein the oxygen supply tube is retained at such a position that a gap is defined between the prongs and the nostrils. 3. The airway adapter as set forth in claim 1, further comprising a mouth guide, adapted to be disposed in front of a mouth of the living body so as to define a space communicated with the airway passage, the mouth guide being pivotably supported on the airway case. 4. The airway adapter as set forth in claim 1, further comprising an inlet member, adapted to be inserted into at least one of the nostrils having a passage for guiding a nasal expiration gas to the airway passage, the inlet member being formed with a vent hole communicating the passage and an exterior of the inlet member. 5. The airway adapter as set forth in claim 4, wherein: the passage of the inlet member is defined by a pair of tube members adapted to be inserted into the nostrils and a junction at which the tube members are merged; andthe vent hole is formed at the junction. 6. The airway adapter as set forth in claim 5, wherein the vent hole is arranged such that a flow of a gas discharged from the vent hole is not substantially interfered by the living body. 7. The airway adapter as set forth in claim 6, wherein the vent hole is arranged so as not to oppose to a face of the living body. 8. A sensor for detecting a carbon dioxide gas in an expiration gas of a living body, comprising: an oxygen supply tube;a photo emitter;a photo receiver; andan airway adapter, which supports the photo emitter and the photo receiver such that a light beam emitted from the photo emitter is received by the photo receiver, the airway adapter comprising:an airway case, adapted to be disposed below nostrils of the living body, and formed with an airway passage extending across an optical axis of the light beam, the airway case having a front side and a back side that opposes the front side, the back side positioned adjacent a face of the living body when the airway case ... | Nihon Kohden Corporation | A61B005/08
US07458942 | B2 | 2008-12-02 | Systems and methods for collecting, testing and transporting liquid biological specimens | 2004-09-22 | A collection system for collecting, testing and transporting liquid biological specimens that limits exposure of a tester to the liquid sample and keeps the collected samples pristine even after exposing some of the sample to test strips. The system includes a collection body for receiving and holding a liquid sample, a lid that is coupled to the body, and a test strip housing. The lid includes a transfer conduit while the test strip housing includes an opening defined within the test strip housing. The transfer conduit and the opening are aligned when the test strip housing is coupled to the lid. When the system is moved or tilted, some of the liquid sample moves through the transfer conduit and the opening into the test strip housing. | 1. A collection system for collecting a liquid sample, the collection system comprising: a body for receiving and holding the liquid sample, the body including a coupling structure defined at a top portion of the body for receiving a lid;the lid including a mating structure for mating with the coupling structure, the lid further comprising a test strip housing receiving structure on a top portion of the lid, and around a transfer conduit defined therein; anda test strip housing removably coupled to the lid through the test strip housing receiving structure, the test strip housing comprising a test strip receiving area, outside of the lid, having at least one test strip reservoir, and further comprising an opening defined within the test strip housing for alignment with the transfer conduit within the lid; andsurrounding the opening, a housing mating structure for mating with the test strip housing receiving structure on the top portion of the lid, thereby coupling the test strip housing to the lid. 2. A collection system in accordance with claim 1, further comprising a vent located adjacent the test strip receiving area. 3. A collection system in accordance with claim 2 wherein the vent is defined within a cover over the test strip receiving area and is located at a position relative to the opening whereby during use of the system, fluid level of liquid sample within the test strip receiving area is defined by positioning of the vent. 4. A collection system in accordance with claim 3 wherein the cover over the test strip receiving area is clear. 5. A collection system in accordance with claim 1, wherein the test strip housing comprises a trough adjacent the at least one test strip reservoir. 6. A collection system in accordance with claim 5 wherein a manifold pad is positioned within the trough. 7. A collection system in accordance with claim 1, wherein a test strip is positioned within the at least one test strip reservoir. 8. A collection system in accordance with claim 1 wherein the transfer conduit is positioned approximately one-half inch from an edge of the lid. 9. A collection system in accordance with claim 1 wherein the transfer conduit is defined within a conical protrusion that extends such that a top of the conical protrusion is substantially flush with a top surface of the test strip housing receiving structure. 10. A collection system in accordance with claim 9 wherein the transfer conduit comprises an inlet and an outlet, and wherein the inlet is larger than the outlet. 11. A collection system in accordance with claim 1, wherein the test strip receiving area comprises a plurality of test strip reservoirs. 12. A collection system in accordance with claim 11, wherein a test strip is positioned within at least one of the test strip reservoirs. 13. A collection system in accordance with claim 2, wherein the transfer conduit is defined within a conical protrusion that extends such that a top of the conical protrusion is subs... | MedTox | A61B005/00 B65D081/00 A61M001/00 A61F005/44
US08185179 | B2 | 2012-05-22 | Fingertip oximeter and a method for observing a measurement result thereon | 2007-03-12 | There is provided a fingertip oximeter that has a plurality of display modes which are presented sequentially in a circulating way, allowing users to easily observe a measurement result from any of surrounding directions. The present invention makes users to be able to observe a measurement result of the fingertip oximeter from any of surrounding directions, without the need of bending his/her finger. Thus, any partial occlusion of the arterial blood capillary can be avoided, so that strength of the pulse will not decrease, and strength of the signal will not be affected. As a result, the precision of the measurement is improved. | 1. A fingertip oximeter, comprising: a display update unit and display for displaying a measurement result to a user using a plurality of display modes which are configured to be presented one by one in a circulating way, headings and parameters of the measurement result being displayed in different directions for the user to observe from any of surrounding directions without changing the direction of the fingertip oximeter relative to the user; anda manually pressing button for updating the display mode in use,wherein each display mode divides a screen of the display into a heading area and a parameter area, in which characters in the heading area are presented in an upright standing way or in an upside-down standing way, while the parameter area is provided to show a measurement parameter, or a measurement parameter and a pulse column, or a measurement parameter and a waveform in an upright standing way, in a right laying way, in an upside-down standing way, or in a left laying way,wherein the display update unit is configured to generate at least six display modes (Mode 0-Mode 5), wherein:in Mode 0, the heading is displayed in a landscape upright standing way, measurement parameters are displayed in a landscape upright standing way, and a pulse column for indicating the strength of pulse is displayed in a landscape left standing way;in Mode 1, the heading, measurement parameters, and the pulse column are respectively displayed in a portrait upright standing way, a portrait right laying way, and a centered standing way;in Mode 2, the heading, measurement parameters, and the pulse column are respectively displayed in a landscape upside-down standing way, a landscape upside-down standing way, and a right standing way;in Mode 3, the heading is displayed in a landscape upright standing way, measurement parameters are displayed in a portrait left laying way, and the pulse column is displayed in a centered standing way;in Mode 4, the heading is displayed in a landscape upright standing way, measurement parameters are displayed in a landscape upright standing way, and a landscape solid waveform is displayed; andin Mode 5, the heading is displayed in a landscape upright standing way, measurement parameters are displayed in a landscape upright standing way, and a landscape outlined waveform is displayed. 2. A fingertip oximeter according to claim 1, wherein the manually pressing button is configured to receive an instruction for updating the display mode of the fingertip oximeter in use, the fingertip oximeter further comprising: a central processing unit which is adapted to determine whether any button is pressed down, upon such an operation being detected, set a new display mode and update the display mode in use with the new one, and transmit a signal regarding the new display mode to the display;a signal drive unit which is adapted to make a light emitter emit a beam of light under control of the central processing unit;a signal acquisition and amp... | Beijing Choice Electronic Technology, Co., Ltd. | A61B005/1455
US08009913 | B2 | 2011-08-30 | System, method, apparatus and computer program product for capturing human-readable text displayed on a unit dose package | 2007-05-29 | A system, method, apparatus and computer program product are provided for capturing human-readable text displayed on a unit dose package. In particular, identification information associated with a unit dose package may be used to determine a location and/or a format in which human-readable text, such as an expiration date or lot number associated with the corresponding unit dose medication, is displayed on the unit dose package. Once the location and/or format of the human-readable text has been determined, this information can be used in order to scan the unit dose medication and translate the human-readable text into machine-readable text using, for example, optical character recognition. | 1. A method of capturing human-readable text displayed on a unit dose package, said method comprising: capturing identification information associated with a unit dose package;determining, based at least in part on the identification information, a location on the unit dose package at which human-readable text is displayed; andelectronically capturing the human-readable text at the determined location,wherein determining, based at least in part on the identification information, a location on the unit dose package at which human-readable text is displayed comprises accessing a mapping of the identification information associated with one or more unit dose packages to information describing the location on the corresponding unit dose package at which the human-readable text is displayed. 2. The method of claim 1, wherein capturing identification information comprises reading an identification code displayed on the unit dose package. 3. The method of claim 2, wherein reading the identification code comprises: locating an edge of the unit dose package. 4. The method of claim 3, wherein reading the identification code further comprises: capturing an image of the unit dose package; anddecoding one or more pixels of the captured image on a row-by-row basis in a predefined direction relative to the edge, wherein the predefined direction is selected from a group consisting of substantially parallel to the edge and substantially perpendicular to the edge. 5. The method of claim 4 further comprising: continuing to decode one or more pixels of the captured image on a row-by-row basis in the predefined direction relative to the edge until the identification code is located or substantially all of the pixels of the captured image have been decoded; andwhere substantially all of the pixels of the captured image have been decoded and the identification code has not been located, decoding one or more pixels on a row-by-row basis in a direction substantially perpendicular to the predefined direction relative to the edge until the identification code is located. 6. The method of claim 3, wherein reading the identification code further comprises: scanning one or more optical signals across the unit dose package in a predefined direction relative to the edge, wherein the predefined direction is selected from a group consisting of substantially parallel to the edge and substantially perpendicular to the edge. 7. The method of claim 6 further comprising: continuing to scan the optical signals across the unit dose package in the predefined direction relative to the edge until the identification code is located or substantially the entire surface of the unit dose package has been scanned; andwhere substantially the entire surface of the unit dose package has been scanned and the identification code has not been located, scanning one or more optical signals across the unit dose package in a direction substantially perpendicular to the predefined direction relative ... | McKesson Automation, Inc. | G06K009/00 G06K005/04
US07904153 | B2 | 2011-03-08 | Method and apparatus for subcutaneous ECG vector acceptability and selection | 2007-04-27 | An implantable medical device system and associated method select subcutaneous sensing electrodes for use in monitoring the heart rhythm. A subset of multiple sensing vectors is selected from a number of electrodes positioned at a first subcutaneous location and one electrode positioned at a second subcutaneous location. The subset of sensing vectors includes one vector that includes the electrode positioned at the second location. A signal quality parameter corresponding to each of the sensing vectors of the subset is determined. | 1. An implantable medical device system, comprising: a plurality of electrodes positioned at a first subcutaneous location;at least one electrode positioned at a second subcutaneous location, wherein the plurality of electrodes at the first location and the electrode at the second location form multiple sensing vectors;a sensing module for receiving ECG signals from the multiple sensing vectors; anda control module configured to select a subset of the multiple sensing vectors, determine a signal quality parameter corresponding to each of the sensing vectors of the subset, determine whether each of the sensing vectors of the subset is acceptable in response to the determined signal quality parameter, and select at least one sensing vector for heart rhythm monitoring from the multiple sensing vectors in response to at least one of the subset sensing vectors being acceptable;wherein the subset of sensing vectors includes only one vector that includes the electrode positioned at the second location;wherein determining whether each of the sensing vectors of the subset is acceptable comprises comparing the signal quality parameter to a threshold requirement,wherein selecting at least one sensing vector for heart rhythm monitoring comprises: computing an acceptability index as a weighted function of a first signal quality parameter and a second signal quality parameter for each of the sensing vectors determined to be acceptable;comparing the computed acceptability indices; andselecting the at least one sensing vector for heart rhythm monitoring in response to the comparison of the computed acceptability indices. 2. The system of claim 1 wherein determining the signal quality parameter includes determining an R-wave peak amplitude. 3. The system of claim 1 wherein determining the signal quality parameter includes determining a low slope content. 4. The system of claim 1 wherein a first of the selected two sensing vectors is a vector including the electrode positioned at the second location and one of the plurality of electrodes positioned at the first location in response to the only one vector of the subset that includes the electrode positioned at the second location being acceptable. 5. The system of claim 4 wherein a second of the selected two sensing vectors includes two electrodes selected from the plurality of electrodes at the first location that are different electrodes than the one of the plurality of electrodes included in the first selected sensing vector. 6. The system of claim 4 wherein the first selected sensing vector is the only one vector of the subset that includes the electrode positioned at the second location in response to only two of the subset sensing vectors being acceptable. 7. The system of claim 1 wherein the first location corresponds to an anterior location. 8. The system of claim 1 wherein the second location corresponds to a posterior location. 9. The system of claim 1 further comprising a housing enclosing the se... | Medtronic, Inc. | A61N001/39
US08235910 | B2 | 2012-08-07 | Systems and methods for model-based estimation of cardiac ejection fraction, cardiac contractility, and ventricular end-diastolic volume | 2008-05-16 | The methods and systems for estimating cardiac ejection fraction, cardiac contractility, and ventricular end-diastolic volume on a beat-by-beat basis include observing arterial blood pressure waveforms to determine ventricular compliances for a pressure-volume loop in the ventricle. Uncalibrated or calibrated cardiac ejection fraction may be calculated from estimates of stroke volume and the ventricular compliances. Cardiac contractility may be calculated from estimates of a ventricular compliance. Uncalibrated or calibrated ventricular end-diastolic volume may also be calculated from estimates of stroke volume and the ventricular compliances. A set of calibration parameters for calibrating cardiac ejection fraction or ventricular end-diastolic volume may be estimated in a least-squares manner. | 1. A method for estimating beat-by-beat cardiovascular parameters and variables, comprising: processing one or more cycles of arterial blood pressure information to determine intra-beat and inter-beat variability in blood pressure; andcomputing estimates of one or more cardiovascular parameters and variables from the intra-beat variability, the inter-beat variability, beat-by-beat estimates of stroke volume, and a ventricular pressure-volume model;wherein the cardiovascular parameters and variables include at least one of a beat-by-beat pre-ejection ventricular compliance, a beat-by-beat end-diastolic compliance, a beat-by-beat end-systolic ventricular compliance, and a beat-by-beat peak-systolic compliance, determined for at least one of the left and right ventricle and computing estimates comprises optimization of an error criterion. 2. The method of claim 1, further comprising using information from phonocardiographic measurements. 3. The method of claim 2, wherein the phonocardiographic information includes parameters relating to the beat-by-beat timing of cardiac cycle events. 4. The method of claim 3, wherein the parameters relating to the beat-by-beat timing of cardiac cycle events include a duration of ventricular ejection. 5. The method of claim 1, wherein the at least one of a beat-by-beat pre-ejection ventricular compliance, a beat-by-beat end-systolic ventricular compliance, and a beat-by-beat peak-systolic compliance is estimated over a data window. 6. The method of claim 1, wherein the error criterion is least-squared error. 7. A method for estimating beat-by-beat cardiovascular parameters and variables, comprising: processing one or more cycles of arterial blood pressure information to determine intra-beat and inter-beat variability in blood pressure; andcomputing estimates of one or more cardiovascular parameters and variables from the intra-beat variability, the inter-beat variability, beat-by-beat estimates of stroke volume, and a ventricular pressure-volume model, wherein the cardiovascular parameters and variables include an uncalibrated beat-by-beat ejection fraction; andcomputing calibrated beat-by-beat ejection fraction from the uncalibrated beat-by-beat ejection fraction using one or more calibration parameters, wherein the calibration parameters are computed using an optimization criterion. 8. The method of claim 7, wherein the calibration parameters are computed for each of the beats. 9. The method of claim 7, wherein the calibration parameters include ventricular dead volume. 10. The method of claim 7, wherein the calibration parameters include intra-thoracic pressure. 11. The method of claim 7, wherein the cardiovascular parameters and variables include uncalibrated beat-by-beat end-diastolic volume. 12. The method of claim 11, further comprising computing calibrated beat-by-beat end-diastolic volume from the uncalibrated beat-by-beat end-diastolic volume using a second set of calibration parameters. 13. Th... |  | A61B005/02
US08099250 | B2 | 2012-01-17 | Impedance parameter values | 2006-07-27 | A method of determining parameter values used in impedance analysis of a subject. The method includes using a processing system to determine a number of impedance measurements at a corresponding number of frequencies. The impedance measurements are used to determine estimates of parameter values, with the estimates being used to determine theoretical impedance values based on the parameter value estimates. The theoretical impedance values are compared to the measured impedance values to allow a modification direction to be determined for at least one of the parameter value estimates. This is then used to modify at least one parameter value estimate, with the parameter values being determined at least in part from one or more modified parameter value estimates. | 1. A method of determining parameter values used in impedance analysis of a subject, the method comprising: a) determining a number of impedance measurements at a corresponding number of frequencies;b) determining estimates of the parameter values;c) determining theoretical impedance values based on the parameter value estimates;d) comparing, in a processor of a processing system, the theoretical impedance values to the measured impedance values;e) determining, in the processor, at least one modification direction for at least one of the parameter value estimates in accordance with the results of the comparison; and,f) modifying, in the processor, at least one parameter value estimate in accordance with the determined direction, the parameter values being determined at least in part from one or more modified parameter value estimates. 2. A method according to claim 1, wherein the method includes, in the processing system, iteratively modifying the parameter value estimates to thereby determine the parameter values. 3. A method according to claim 1, wherein the method includes, in the processing system, comparing the theoretical impedance values to the measured impedance values: a) determining a theoretical impedance locus using the parameter value estimates;b) determining a measured impedance locus associated with each measured impedance; and,c) determining an intersection between each measured impedance locus and the theoretical impedance locus. 4. A method according to claim 3, wherein the method includes, in the processing system: a) determining a point associated with each intersection;b) determining a linear regression based at least partially on the points; and,c) determining the modification direction based at least partially on at least one of: i) the relative position of the regression line and at least one of the points; and,ii) the gradient of the regression line. 5. A method according to claim 4, wherein the method includes, in the processing system: a) determining a pair of chords associated with each point of intersection, the chords representing the distance to locations at which the theoretical impedance locus represents predetermined impedance values; and,b) determining a point associated with each pair of chords; and,c) determining the linear regression in accordance with the value of the points. 6. A method according to claim 1, wherein the method includes, in the processing system: a) determining four impedance measurements, each impedance measurement being made at a respective frequency;b) comparing each measured impedance value to a corresponding theoretical impedance value;c) determining the modification direction based at least in part on the results of the comparison. 7. A method according to claim 6, wherein the method includes, in the processing system: a) determining a point based on the comparison of each measured impedance value with a corresponding theoretical impedance value;b) determining a linear regression... | Impedimed Limited | G01R027/00
US07972266 | B2 | 2011-07-05 | Image data normalization for a monitoring system | 2007-05-22 | A method for providing normalized physiological monitoring data of an individual with a measure of the quality of the normalization, including providing a reference feature based on a physiological attribute of the individual for which an attribute value can be determined; unobtrusively capturing physiological monitoring data for the individual during a series of capture events and determining the capture conditions present during each capture event; detecting the presence of the reference feature in the series of captured physiological monitoring data and determining associated attribute values for each capture event; normalizing the captured physiological monitoring data from each capture event according to differences in the attribute values associated with each event and previously calculated attribute values; and calculating normalization confidence values for the individual at the series of capture events, based on capture conditions, normalization transforms, or semantic data, wherein the confidence values statistically measure the quality of the normalization. | 1. A method for providing normalized physiological monitoring data of an individual with a measure of the quality of the normalization, comprising; providing a reference feature based on a physiological attribute of the individual for which an attribute value can be determined;unobtrusively capturing physiological monitoring data for the individual during a series of capture events and determining the capture conditions present during each capture event;detecting the presence of the reference feature in the series of captured physiological monitoring data and determining associated attribute values for each capture event;normalizing the captured physiological monitoring data from each capture event according to differences in the attribute values associated with each event and previously calculated attribute values; andcalculating normalization confidence values for the individual at the series of capture events, based on capture conditions, normalization transforms, or semantic data, wherein the confidence values statistically measure the quality of the normalization. 2. The method as in claim 1 wherein the semantic data related to the individual, that can affect the quality of the normalization, includes the time of day of a capture event, the use of cosmetics, or the tiredness of the individual. 3. The method as in claim 1 wherein the normalization confidence values are tracked over multiple capture events to monitor changes in normalization quality over time. 4. The method as in claim 1 further comprising determining if valid wellness parameters can be derived based on the normalization confidence values, wherein the wellness parameters are associated with the physiological attributes being monitored. 5. The method as in claim 1 wherein the attribute values or normalization transforms are obtained by RGB or spectral color measurements of: the sclera, the skin, the tongue, the ambient light, or the capture light. 6. The method as in claim 1 wherein the physiological monitoring data is obtained by image capture, audio capture or multi-spectral capture. 7. The method as in claim 6 wherein the capture conditions are derived as capture parameters from one or more captured images acquired during a capture event. 8. The method as in claim 6 wherein the face of the individual in the captured images is normalized in each image according to color attributes of one or more facial reference features of the individual. 9. The method as in claim 6 wherein the face of the individual in the captured images is normalized based on the spatial disparity between two or more facial reference features of the individual. 10. The method as in claim 9 wherein the captured images of the individual are normalized based on the spatial disparity between corresponding features of both eyes of the individual. 11. The method as in claim 6 wherein the captured images of the individual are normalized to provide changes in the size or orientation or combinations ther... | Eastman Kodak Company | A61B005/00 G06K009/80 H04N005/243
US07861397 | B2 | 2011-01-04 | Method of making an electrochemical sensor | 2007-10-30 | A process for the manufacture of small sensors with reproducible surfaces, including electrochemical sensors. One process includes forming channels in the surface of a substrate and disposing a conductive material in the channels to form an electrode. The conductive material can also be formed on the substrate by other impact and non-impact methods. In a preferred embodiment, the method includes cutting the substrate to form a sensor having a connector portion and a transcutaneous portion, the two portions having edges that define one continuous straight line. | 1. A method of manufacturing an in vitro sensor for the determination of the presence and/or level of glucose in a sample, the sensor comprising: forming at least one conductive trace on a substrate by non impact printing, including applying a conductive material on the substrate to provide one or more traces including the at least one conductive trace to form one or more electrodes, at least one trace having a width of about 25 microns to 250 microns, and the substrate comprising a first portion that is narrower than a second portion; wherein the in vitro sensor is configured to determine the presence and/or level of glucose in a volume sample of about 1 microliter or less. 2. The method of claim 1, wherein the conductive material forms a working electrode. 3. The method of claim 1, further comprising forming a sample chamber. 4. The method of claim 1, wherein the at least one trace has a width of about 100 microns or less. 5. The method of claim 4, wherein the at least one trace has a width of about 75 microns or less. 6. The method of claim 5, wherein at least one trace has a width of about 50 microns or less. 7. The method of claim 6, wherein at least one trace has a width of about 25 microns or less. 8. The method of claim 1, wherein the sensor is configured to determine the presence and/or level of glucose in a volume sample of about 0.5 microliters or less. 9. The method of claim 1 wherein the non impact printing to form the at least one conductive trace on the substrate includes one or more of ink jet printing or piezo jet printing. 10. The method of claim 1 wherein the volume sample is a blood sample. 11. The method of claim 1 including: depositing a liquid precursor over at least a portion of the substrate; andcuring the liquid precursor with a photomask having an image of the at least one conductive trace. 12. The method of claim 11 including directing a light through the photomask. 13. A method of manufacturing an in vitro sensor for the determination of the presence and/or level of glucose in a sample, the sensor comprising: forming at least one conductive trace on a substrate by non impact printing, including applying a conductive material on the substrate to provide one or more traces including the at least one conductive trace to form one or more electrodes, at least one trace having a width of about 25 microns to 250 microns; andproviding an indentation in the substrate; wherein the in vitro sensor is configured to determine the presence and/or level of glucose in a volume sample of about 1 microliter or less. 14. The method of claim 13, wherein the sensor is configured to determine the presence and/or level of glucose in a volume sample of about 0.5 microliter or less. 15. The method of claim 13, wherein the volume sample is a blood sample. 16. The method of claim 13 including: depositing a liquid precursor over at least a portion of the substrate; andcuring the liquid precursor with a photomask having an image of... | Abbott Diabetes Care Inc. | G01R003/00
US08897417 | B2 | 2014-11-25 | Method and device for planning a treatment | 2006-11-23 | A device for planning an irradiation is provided. The device includes an evaluation module with an input for receiving input data, a memory and an output for outputting determined output data. The evaluation module is designed for using the input data that includes the type and number of the imaging units present, variables characterizing the tumor and/or variables characterizing the patient, in order to determine the output data that includes the type of the imaging unit, the frequency of use of the imaging unit and/or the parameters for setting the imaging unit to be used in the form of an imaging plan with the aid of a functional relationship, which is based on experience, stored in the memory. | 1. A method for planning a treatment of a tumor of a patient in the case of particle, radiation, or particle and radiation therapy, the method comprising: determining, in an evaluation module, output data based on input data using a functional relationship based on experience, the functional relationship being stored in a memory, the input data including a type of imaging unit present and a number of the imaging units present, variables characterizing the tumor, variables characterizing the patient, or combinations thereof, the output data being for an imaging plan to take images of the tumor during treatment and comprising the type of the imaging unit, a frequency of use of the imaging unit, parameters for setting the imaging unit, or combinations thereof; anddetermining the imaging plan as part of the therapy based on the determined output data; andobtaining an image data record by the imaging unit to be used, the image data record being used to change treatment parameters. 2. The method as claimed in claim 1, further comprising changing the functional relationship in the evaluation module. 3. The method as claimed in claim 2, wherein the change in the functional relationship is determined based on results of images obtained with the output data. 4. The method as claimed in claim 1, wherein the treatment parameters are calculated by comparing the image data record and an original basic image data record. 5. The method as claimed in claim 1, wherein the input data includes a type and number of irradiation devices present, and output data for a treatment plan is determined in the evaluation module. 6. The method as claimed in claim 5, wherein the evaluation module exchanges data with the imaging units, the irradiation devices, or the imaging units and the irradiation devices via a network. 7. The method as claimed in claim 1, further comprising simulating, by the evaluation module, a dry run of the imaging unit to be used. 8. A device for planning a treatment of a tumor of a patient in the case of particle, radiation, or particle and radiation therapy, the device comprising: an evaluation module with an input operable to receive input data, a memory and an output operable to output determined output data,wherein the evaluation module is operable to determine the output data using the input data, the input data comprising a type and number of imaging units present, variables characterizing the tumor, variables characterizing the patient, or combinations thereof, the output data comprising the type of the imaging unit, a frequency of use of the imaging unit, parameters for setting the imaging unit, or combinations thereof to be used in an imaging plan to take images of the tumor, the evaluation module operable to use a functional relationship based on experience and stored in the memory to determine the imaging plan as part of the therapy, andwherein the evaluation module is operable to process a type and number of irradiation devices includ... | Siemens Aktiengesellschaft | A61N005/10 G06F019/00
US09398863 | B2 | 2016-07-26 | Detection of anomalies in measurement of level of hypnosis | 2007-06-20 | The invention relates to detection of anomalies related to a measurement of the hypnotic level of a subject. In order to detect when a measure indicative of the hypnotic level of a subject is anomalous either due to a medical reason or due to an interference, the measure is monitored and state information indicative of the activity of the autonomous nervous system of the subject and drug effect information indicative of the hypnotic drug effect in the subject are employed to check whether the measure fulfills a predetermined consistency condition requiring that the measure changes consistently with at least one of the drug effect information and the state information. An anomaly, indicative of an abnormal change in the measure, is detected when the measure fails to fulfill the predetermined consistency condition. | 1. A method for detecting an anomaly in a measurement of the hypnotic level of a subject, the method comprising: monitoring using a patient monitor a change in a measure indicative of the hypnotic level in the subject;receiving at a processing unit state information indicative of the activity of the autonomous nervous system of the subject;receiving at the processing unit drug effect information indicative of the hypnotic drug effect in the subject;comparing the change in the measure to one of the drug effect information and the state information;checking with the processing unit whether the change in the measure is a consistent change with at least one of the drug effect information and the state information; anddetecting with the processing unit an anomaly, wherein an anomaly is indicative that the change in the measure is not a consistent change with at least one of the drug effect information and the state information, and wherein the anomaly is indicative of an alteration in the normal relation between the measure and the at least one of the drug effect information and the state information. 2. A method according to claim 1, wherein the checking includes: defining an estimate for the measure based on a model that depends on the state information and on the drug effect information; anddetermining a variable indicative of the difference of the measure and the estimate;wherein the detecting includes detecting the anomaly when the variable fails to remain within a predetermined value range. 3. A method according to claim 1, wherein: the checking includes checking whether the change in the measure is a consistent change with changes in the drug effect information and the state information; andthe detecting includes detecting the anomaly when the checking indicates that the change in the measure is not a consistent change with both the drug effect information and the state information. 4. A method according to claim 1, further comprising: deriving a rate of change in the measure; andcomparing the rate of change to a predetermined threshold value. 5. A method according to claim 4, wherein the comparing indicates that the rate of change exceeds the predetermined threshold value and then the checking is triggered. 6. A method according to claim 5, wherein the checking is aborted if it is detected that the change in the measure is a consistent change with one of the drug effect information and the state information. 7. A method according to claim 1, further comprising: administering a hypnotic drug to the subject; andcontrolling the administering in response to the detecting. 8. A method according to claim 7, wherein the controlling includes freezing the administering to its current state. 9. A method according to claim 1, further comprising generating, in response to the detecting, an indication informing a user about the anomaly. 10. A method according to claim 1, further comprising deriving the measure from EEG signal data obtained from th... | General Electric Company | A61B005/04 A61B005/0476 A61B005/00 A61M005/172 A61M016/00 A61M016/18 G06F019/00 G09B023/28
US09801565 | B2 | 2017-10-31 | Methods and workflow design for improving specificity of breast MR | 2007-01-08 | In a breast imaging method, magnetic resonance data are acquired of one or both breasts of a subject (16) using a breast coil (20, 20′) coupled with said one or both breasts. The acquiring employs at least four independent channels (41, 42, 43, 44, 45, 46; 61, 62, 63, 64, 65, 66) per breast. The acquired magnetic resonance data are processed to generate at least one of (i) an image, (ii) a spectrum, and (iii) elasticity data. The at least four independent channels per breast are suitably embodied as a dual breast coil (20) including left breast conductors (31, 32, 33, 34, 35, 36) defining at least four independent left breast acquisition channels (41, 42, 43, 44, 45, 46) and right breast conductors (51, 52, 53, 54, 55, 56) defining at least four independent right breast acquisition channels (61, 62, 63, 64, 65, 66). | 1. A breast imaging method comprising: acquiring magnetic resonance data of one or both breasts of a subject using a dual breast coil coupled with said one or both breasts, the acquiring employing at least four independent channels per breast, the dual breast coil including at least four independent acquisition channels per breast wherein the at least four independent acquisition channels for each breast include acquisition channels defined by (i) at least one conductor positioned anterior of the breast, (ii) at least one conductor positioned posterior of the breast, (iii) at least one conductor positioned above the breast, (iv) at least one conductor positioned below the breast, (v) at least one conductor positioned lateral of the breast, and (vi) at least one conductor positioned medial of the breast, wherein at least one of the four independent acquistion channels is defined by an electrical or inductive coupling between two of the conductors; andprocessing the acquired magnetic resonance data to generate elasticity data. 2. The breast imaging method as set forth in claim 1, wherein the acquiring includes (i) acquiring at least some magnetic resonance data as undersampled magnetic resonance data using the at least four independent acquisition channels, and the processing includes (ii) performing image reconstruction including combining the undersampled magnetic resonance data from the at least four independent acquisition channels using SENSE to generate an image. 3. The breast imaging method as set forth in claim 2, wherein the at least four independent acquisition channels for each breast include at least six independent acquisition channels per breast, and the acquiring (i) and combining (ii) operate with a SENSE factor of at least four. 4. The breast imaging method as set forth in claim 1, wherein the processing includes: generating at least two inherently aligned images including at least two of (i) a T1 weighted image, (ii) a T2 weighted image, (iii) a diffusion weighted image, and (iv) a contrast dynamic image; andcombining the at least two generated images to produce a fused image. 5. The breast imaging method as set forth in claim 4, wherein the generating includes one of: generating at least the T2 weighted image, the diffusion weighted image, and the contrast dynamic image, andgenerating the T1 weighted image, the T2 weighted image, and the diffusion weighted image, but not the contrast dynamic image. 6. The breast imaging method as set forth in claim 4, wherein the acquiring includes: acquiring magnetic resonance data for the at least two of (i) the T1 weighted image, (ii) the T2 weighted image, (iii) the diffusion weighted image, and (iv) the contrast dynamic image over the same imaged volume, the generating producing images spanning the same volume, the combining operating on a voxel-by-voxel basis. 7. The breast imaging method as set forth in claim 4, wherein the acquiring further includes: acquiring magnetic resonance ela... | Koninklijke Philips N.V. | A61B005/055 G01R033/3415 G01R033/561
US09320452 | B2 | 2016-04-26 | Magnetic resonance imaging of amyloid plaque in the brain | 2005-06-28 | Amyloid plaque in the brain of a subject is imaged in an MRI system with or without the use of a contrast agent. Contrast is achieved using a spin-echo pulse sequence that is both respiratory gated and cardiac gated to reduce motion artifacts at the very high image resolution required to see plaque. A preparatory pulse sequence is used to insure longitudinal magnetization remains constant for all the acquired views even if the effective TR changes during the scan due to irregular breathing. | 1. A method for acquiring an in vivo, high resolution magnetic resonance image from a subject, the steps comprising: a) producing a trigger signal when the subject is in a preselected phase of the subject's respiratory cycle;b) performing a pulse sequence to acquire an NMR signal when the trigger signal is produced, the pulse sequence including: b)i) performing a preparatory pulse sequence to create uniform view-to-view longitudinal magnetization in the presence of non-uniform view-to-view trigger times in which the longitudinal magnetization in a region of interest in the subject is set to a preset amount by applying an rf nulling pulse with a substantially 90° flip angle followed by a preset null recovery time (TN);b)ii) producing transverse magnetization in the region of interest by applying an rf excitation pulse after step b)i);b)iii) producing an NMR echo signal at a preselected time (TE) after application of the rf excitation pulse;b)iv) acquiring the NMR echo signal in the presence of a readout magnetic field gradient;c) changing a magnetic field gradient in the pulse sequence and repeating step b) to acquire another NMR signal; andd) repeating step c) until sufficient NMR echo signals are acquired to reconstruct an image. 2. The method as recited in claim 1 in which the region of interest is in the subject's brain. 3. The method as recited in claim 2 in which the region of interest is in a portion of the subject's brain containing amyloid plaque and the NMR echo signal is produced in step b)iii) by producing an rf refocusing pulse. 4. The method as recited in claim 1 in which the rf nulling and rf excitation pulses are adiabatic half-passage RF pulses that flip substantially all of the longitudinal magnetization into the transverse plane. 5. The method as recited in claim 1 in which the pulse sequence performed in step b) further includes: producing a first selective RF refocusing pulse in the presence of a first gradient after step b)ii) to refocus the transverse magnetization in a first slice through the region of interest; andproducing a second selective RF refocusing pulse in the presence of a second gradient after step b)ii) to refocus the transverse magnetization in a second slice through the region of interest. 6. The method as recited in claim 5 in which the selective RF refocusing pulses are hyperbolic scant type pulses. 7. A method for acquiring an in vivo, high resolution magnetic resonance image from a subject, the steps comprising: a) producing a first trigger signal when the subject is in a preselected phase of the subject's respiratory cycle;b) performing a preparatory pulse sequence when the first trigger signal is produced by applying an rf nulling pulse followed by a preset null recovery time (TN);c) producing a second trigger signal when the subject is in a preselected phase of the subject's cardiac cycle;d) performing an imaging pulse sequence to acquire an NMR echo signal from a selected region of interest in t... | MAYO FOUNDATION FOR MEDICAL EDUCATION AND RESEARCH REGENTS OF THE UNIVERSITY OF MINNESOTA | A61B005/055 A61B005/00 G01R033/567
US20080242961 | A1 | 2008-10-02 | TRANSCUTANEOUS ANALYTE SENSOR | 2008-06-11 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A transcutaneous analyte sensor assembly, the assembly comprising: a housing adapted for mounting on a skin of a host; anda sensor configured to measure an analyte concentration in the host, wherein the sensor comprises: an in vivo portion configured for insertion through the host's skin at an exit-site;an ex vivo portion extending through the housing and adapted to remain above the host's skin after insertion of the sensor; anda transition portion located between the in vivo portion and the ex vivo portion and configured to extend through the exit-site after sensor insertion,wherein the sensor possesses an antimicrobial property. 2. The assembly of claim 1, wherein the sensor is configured to passively inhibit growth of a microorganism. 3. The assembly of claim 2, wherein the sensor comprises a silver material that confers the antimicrobial property to the sensor. 4. The assembly of claim 1, wherein the sensor is configured to actively inhibit growth of a microorganism. 5. The assembly of claim 4, wherein the sensor comprises an auxiliary electrode configured to pass a current that inhibits growth of a microorganism. 6. The assembly of claim 1, wherein the transition portion possesses the antimicrobial property. | DexCom, Inc. | A61B005/05
US20080281171 | A1 | 2008-11-13 | ANALYTE MONITORING SYSTEM AND METHODS | 2008-05-08 | Methods and systems for providing data communication in medical systems are disclosed. | 1. An analyte monitoring device, comprising: a data transmission section;a data receiving section; anda data processing section coupled to the data receiving section;wherein the data receiving section is configured to receive a predefined command, and further, wherein the data processing section is configured to control the data transmission section to transmit one or more data in response to the received predefined command. 2. The device of claim 1 wherein the data receiving section includes a close proximity receiver. 3. The device of claim 1 wherein the data transmission section includes an RF transmitter. 4. The device of claim 1 including an antenna coupled to the data transmission section and the data receiving section for wireless data communication. 5. The device of claim 1 wherein the data processing section includes a state machine. 6. The device of claim 1 wherein the transmitted one or more data includes a device identification information. 7. The device of claim 1 wherein the transmitted one or more data includes a response data packet. 8. The device of claim 7 wherein the response data packet includes one or more of a communication mode information, a device status information, a device type information, a close proximity command information, or a sensor count information. 9. The device of claim 1 including a memory unit for storing data. 10. A method, comprising: receiving a command to initiate communication with an analyte sensor;retrieving an identification information;transmitting the retrieved identification information; andreceiving a communication key associated with the transmitted identification information. 11. The method of claim 10 wherein the received command includes a close proximity command. 12. The method of claim 10 wherein the retrieved identification information is wirelessly transmitted over an RF communication link. 13. The method of claim 10 wherein the identification information includes a data processing device identification information. 14. The method of claim 13 wherein the data processing device identification information includes a transmitter identification information. 15. The method of claim 10 wherein the communication key includes an 8 bit data associated with the identification information. 16. The method of claim 10 including transmitting a data packet including the communication key. 17. The method of claim 10 including storing the communication key. 18. The method of claim 10 wherein the analyte sensor includes a glucose sensor. 19. The method of claim 10 including receiving one or more signals related to a monitored analyte level. 20. The method of claim 19 wherein the analyte level includes glucose level. | Abbott Diabetes Care, Inc. | A61B005/145
US20080306348 | A1 | 2008-12-11 | MINIATURE WIRELESS APPARATUS FOR RECORDING PHYSIOLOGICAL SIGNALS OF HUMANS AND USE THEREOF | 2007-06-06 | A miniature wireless apparatus for recording and remote monitoring physiological signals and use thereof are described. The miniature wireless apparatus comprises a signal collecting module, a signal processing module, a wireless module, and a recording device. The methods for recording and remote monitoring physiological signals comprising collecting physiological signals by a signal collecting module, processing the said physiological signals by a signal processing module and recording the processed signals by a recording device, modulating the said processed signals and synchronously transmitting and receiving the modulated signals by a wireless module to integrate the physiological signals, receiving the signals by a remote receiver, and recording, monitoring, and analyzing the physiological signals by the remote receiver and further transmitting the physiological signals to the remote monitoring station through a network. Methods for sleep monitoring and evaluating and for monitoring and evaluating of autonomic nerve function using the miniature wireless apparatus are further described. | 1. A miniature wireless apparatus for recording and remote monitoring physiological signals, comprising: (a) a signal collecting module for collecting one or more physiological signals;(b) a signal processing module coupled to the signal collecting module for processing the said physiological signals;(c) a wireless module for modulating the signals from the signal processing module, transmitting the modulated signals, and receiving external wireless signals; and(d) a recording device. 2. The miniature wireless apparatus of claim 1, wherein the signal collecting module comprises one or more sensors for detecting physiological signals. 3. The miniature wireless apparatus of claim 2, wherein the sensors comprises wireless or non-wireless sensors. 4. The miniature wireless apparatus of claim 3, wherein the sensors comprises an electroencephalogram sensor, an electrooculogram sensor, an electrocardiogram sensor, an electromyogram sensor, an electrogastrography sensor, an temperature sensor, and an acceleration sensor. 5. The miniature wireless apparatus of claim 4, wherein the acceleration sensor is a miniature motion sensor for detecting 3-dimensional acceleration and the measured signals of acceleration of motion are converted to analog voltage signals. 6. The miniature wireless apparatus of claim 1, wherein the signal processing module comprises an amplifier for amplifying the electrical physiological signals and a microcontroller for performing analog-to-digital conversion and data compression for an amplified physiological signal generated by the amplifier. 7. The miniature wireless apparatus of claim 6, wherein the amplifier is an operational amplifier for processing single-point signals or differential signals. 8. The miniature wireless apparatus of claim 1, wherein the wireless module comprises an RF transceiver and an antenna for modulating the digital signals from the signal processing module to form modulated radio frequency signals, transmitting the modulated radio frequency signals to an internal and a remote receiver, and receiving external wireless signals. 9. The miniature wireless apparatus of claim 8, wherein the internal receiver comprises the receiver of the apparatus transmitting the signals and a receiver of other miniature wireless apparatuses used simultaneously on the body of a user. 10. The miniature wireless apparatus of claim 8, wherein the remote receiver is a fixed or portable microcomputer system. 11. The miniature wireless apparatus of claim 10, wherein the physiological signals received by the remote receiver are recorded, monitored, and analyzed by the microcomputer system and further transferred to remote monitoring stations through a network. 12. The miniature wireless apparatus of claim 8, wherein the external wireless signals comprises the signals of remote manipulation of the apparatus and the physiological signals from other miniature wireless apparatuses used simultaneously on the body of the user. ... | NATIONAL YANG-MING UNIVERSITY | A61B005/00
US20080249393 | A1 | 2008-10-09 | METHOD AND APPARATUS FOR ENHANCEMENT AND QUALITY IMPROVEMENT OF ANALYTE MEASUREMENT SIGNALS | 2007-04-04 | While measuring analytes in an object, movements between the object and a device that is taking the measurement may adversely effect the accuracy of the measurements. A probe that engages an object includes a sticky surface that comes into contact with the object to alleviate the relative movement between the probe and the object. The sticky substance can be applied in a variety of manners and embodiments and in general, helps to reduce the relative movement. In addition, the probe may include the ability to apply pressure to the object. The pressure can range from a slight pressure to help reduce relative movement to a pressure that exceeds systolic pressure, thereby constricting the flow of blood through the object. | 1. A probe for application to an object for non-invasive analyte measurement, said probe comprising: a) a body with at least one sticky surface, andb) a source of radiation and a detector built into said body. 2. The probe according to claim 1, wherein said source of radiation and detector are operable to measure a blood and tissue analyte. 3. The probe according to claim 1, wherein said body comprises one of a pressure development article, a multilayer resilient structure, or a combination of both of them. 4. The probe according to claim 1, wherein said sticky surface engages said object. 5. The probe according to claim 1, wherein said sticky surface restricts relative movements between said probe and said object. 6. The probe according to claim 1, wherein said sticky surface is one of a silicone layer having a hardness of less than 15 Shorr surface, a coating, a plurality of micron size columns, or a temporarily sprayed coating. 7. The probe according to claim 1, wherein said sticky surface is a cleanable surface. 8. The probe according to claim 1, wherein said source of radiation is a source of red and infrared radiation. 9. The probe according to claim 1, wherein the resilient structure of said body is a multilayer structure, and wherein at least one of said layers develops and supports application of pressure to said object. 10. The probe according to claim 1, wherein the resilient structure of said body is a multilayer structure, and wherein at least one of said layers has at least a section of a surface possessing sticky properties. 11. The probe according to claim 1, wherein at least one of the layers of the multilayer structure of said body is armored by one of external or internal armor. 12. The probe according to claim 1, wherein said body is a lightproof thimble-like or sleeve-like body. 13. The probe according to claim 1, wherein operation of the pressure development article of said probe includes operation in occlusion-release mode, and wherein said occlusion mode develops pressure higher than systolic pressure. 14. An apparatus for determination of a physiological parameter of a subject, said apparatus comprising: a) a probe for receiving a measurement object, wherein the surface of said probe engaging the object is a sticky surface, andb) a control and processing unit. 15. The apparatus according to claim 14, wherein said sticky surface restricts relative movements between said probe and the object. 16. The apparatus according to claim 14, wherein said probe is one of a resilient structure, a pressure development article, or a combination of both. 17. The apparatus according to claim 14, wherein a source of red and infrared radiation and a detector are built into said probe. 18. The apparatus according to claim 14, wherein said probe includes a pressure development article and operation of said article includes operation in occlusion-release mode. 19. The apparatus according to claim 14, wherein the pressure deve... | - | A61B005/1455
US20080269596 | A1 | 2008-10-30 | Orthpaedic Monitoring Systems, Methods, Implants and Instruments | 2005-03-10 | Orthopaedic operating systems, operating rooms, computer aided surgical methods, trackable implants, instruments and surgical planning and IGS software applications are described. The integrated surgical system is for use in an orthopaedic operating room to enable a surgeon to carry out a computer aided surgical procedure on a subject. A subject support, wireless magnetic tracking system, registration system, at least a first display device, a control system and a surgeon interface operable by the surgeon to control operation of the plurality of parts of the integrated surgical system are provided. The method comprises determining the position of at least a first marker being wirelessly tracked by a wireless magnetic tracking system, registering the position of the body part of the subject with an image of the body part of the subject, displaying a registered image of the body part of the subject and at least an image representative of an implant, and receiving a command from a surgeon interface operable by the surgeon. | 1. An integrated surgical system for use in an orthopaedic: operating room to enable a surgeon to carry out a computer aided surgical procedure on a subject, the integrated surgical system comprising: a subject support on which the subject can be positioned;a first wireless magnetic tracking system, the tracking system generating a magnetic field defining a working volume of the tracking system, the subject support being located at least partially within the working volume, and the tracking system including a tracking control system configured to track the position of a marker detectable by the tracking system within the working volume and generate a signal indicative of the position of the marker within a reference frame of the tracking system;a registration system configured to register the position of the body part of the subject with an image of the body part of the subject within the reference frame of the tracking system;at least a first display device configured to display a registered image of the body part of the subject and at least an image representative of a trackable implant during the computer aided surgical procedure;a control system configured to integrate the functionalities of a plurality of the parts of the surgical system; anda surgeon interface operable by the surgeon to control operation of the plurality of parts of the integrated surgical system. 2. The system of claim 1, further comprising a second wireless tracking system, the second wireless tracking system being an infrared wireless tracking system and being in communication with the control system and configured to generate a signal indicative of the position of a tracked element in the reference frame of the second wireless tracking system. 3. The system claim 1, wherein the first display device is a touch sensitive display and comprises a part of the surgeon interface. 4. The system of claim 1, wherein the surgeon interface includes an orientation sensitive device operable by a surgeon to enter control commands. 5. The system of claim 1, wherein the surgeon interface includes heads up display wearable by the surgeon and configured to display at least a one of the images selected from the group comprising: a captured image of the body part; an image of a model of the body part; a registered image of the body part; a video image of the body part; a representation of an implant; a representation of an instrument; an indication of the planned position of an implant, instrument or incision; and any combination of the preceding. 6. The system of claim 1, further comprising a wall display unit, the wall display unit being configured to provide a plurality of image regions, each image region being capable of displaying a different image. 7. The system of claim 6, wherein the different images are selected from the group comprising: a captured image of the body part; an image of a model of the body part; a registered image of the body part; a video image of the body par... | - | A61B005/00 G06Q050/00 A61B019/00
US20080242945 | A1 | 2008-10-02 | Facility management system | 2007-03-30 | The invention provides systems and methods for a acquiring, manipulating, and monitoring data associated with one or more assisted living facilities. The invention includes utilizing one or more monitoring devices to gather data regarding at least one assisted living facility. The data is then transmitted across a network and saved in a database. An enterprise management application then operates on the data to manipulate and display the data according to parameters and instructions specified by one or more users. The data may be used by the one or more users to monitor and manage the one or more assisted living facilities. | 1. A system for acquiring and monitoring operational data in assisted living facilities, the system comprising: a plurality of event monitoring devices that gather data associated with at least one assisted living facility;a database that receives, over a network, the data associated with the at least one assisted living facility from the plurality of event monitoring devices;an enterprise management application that generates one or more reports regarding the at least one assisted living facility using the data associated with the at least one assisted living facility; anda graphical user interface that displays the one or more reports to one or more user devices. 2. The system of claim 1, wherein the data associated with the at least one assisted living facility includes one or more of patient distress event data, patient distress event response data, employee time data, security system data, automatic entryway data, and inventory data. 3. The system of claim 1, wherein the plurality of event monitoring devices include one or more of patient medical alert necklaces, medical monitoring devices, staff event response devices, employee time entry systems, facility security systems, and automatic entryway systems. 4. The system of claim 1, wherein the network includes the Internet. 5. The system of claim 4, wherein the graphical user interface is accessible over the network via a web portal. 6. The system of claim 1, wherein the network includes a wireless telecommunications network. 7. The system of claim 1, wherein the graphical user interface is an alert sent to the one or more user devices over the network. 8. The system of claim 7, wherein the alert includes a task to be performed. 9. The system of claim 7, wherein the alert is automatically generated. 10. The system of claim 7 wherein the alert is generated on demand. 11. The system of claim 1, wherein the one or more user devices include one or more of a server, a desktop computer, a laptop computer, a dumb terminal, a handheld computer, a personal digital assistant, and a cell phone. 12. A method for acquiring and monitoring operational data in assisted living facilities, the method comprising: receiving, over a network, data associated with at least one assisted living facility from a plurality of event monitoring devices;storing the data associated with the at least one assisted living facility in a database;generating one or more reports regarding the at least one assisted living facility from the stored data associated with the at least one assisted living facility;presenting at least one of the one or more reports to one or more user devices via a graphical user interface. 13. The method of claim 12, wherein the data associated with the at least one assisted living facility includes one or more of patient distress event data, patient distress event response data, employee time data, security system data, automatic entryway data, and inventory data. 14. The method of claim 1... | Stanley Security Solutions, Inc. | A61B005/00
US20080269632 | A1 | 2008-10-30 | Systems and Methods for Analyzing and Assessing Attention Deficit Hyperactivity Disorder | 2007-10-23 | Embodiments of the invention can provide systems and methods for analyzing and assessing attention deficit hyperactivity disorder (ADHD) by integrating the use of electroencephalography (EEG), and ADHD diagnostic and assessment tools, such as an ADHD rating scale. Embodiments of the invention can provide some or all of the following improvements over conventional systems and methods, including: (1) Increased sensitivity, specificity, and overall accuracy; (2) Improved detection of ADHD; and (3) Distinguishing subjects with ADHD from subjects with a different disorder but having ADHD-like symptoms, such as attention and behavior symptoms similar to ADHD. | 1. A method for assessing attention deficit hyperactivity disorder in a subject, the method comprising: receiving EEG data associated with a subject;selecting a portion of the EEG data based at least in part on the number of artifacts in the EEG data;based at least in part on the selected EEG data, determining at least one theta-beta ratio; andstandardizing the theta-beta ratio;receiving ADHD rating scale data associated with the subject;based at least in part on some or all of the ADHD rating scale data, determining at least one ADHD rating scale score;standardizing the at least one ADHD rating scale score; andbased at least in part on both the standardized theta-beta ratio and standardized ADHD rating scale score, determining a probability the subject has ADHD. 2. The method of claim 1, wherein at least a portion of the EEG data is collected from the subject's body from about the CZ site location using the International 10-20 system of electrode placement. 3. The method of claim 1, wherein at least a portion of the EEG data is collected while the subject's eyes are opened in a fixed gaze. 4. The method of claim 1, wherein the number of artifacts is a predefined threshold. 5. The method of claim 1, wherein the theta-beta ratio is determined by dividing the percent power of a theta band associated with the selected EEG data by the percent power of a beta1 band associated with the selected EEG data, wherein the theta band is about 4 Hz to about 8 Hz, and the beta1 band is about 13 Hz to about 21 Hz. 6. The method of claim 1, wherein the theta-beta ratio is standardized to a Z-score using a normative or clinical database by at least age. 7. The method of claim 1, wherein the ADHD rating scale data can comprise data from at least one of the following: an inattentive score or total score of a ADHD-IV rating scale, or an ADHD index of a CRS-R rating scale. 8. The method of claim 1, wherein the at least one ADHD rating scale score is standardized to a normative or clinical database by at least age and gender. 9. The method of claim 1, wherein the probability determination is facilitated with a logistic regression model. 10. The method of claim 1, wherein a receiver operating characteristic (ROC) curve can be used to output the probability determination. 11. A system for assessing attention deficit hyperactivity disorder in a subject, the system comprising: a data collection module operable to: receive EEG data associated with a subject;receive ADHD rating scale data associated with the subject;a processing module operable to: determine at least one theta-beta ratio based at least in part on some or all of the EEG data;determine at least one ADHD rating scale score based at least in part on the ADHD rating scale data;standardize the at least one theta-beta ratio and the at least one ADHD rating scale score; anddetermine a probability the subject has ADHD based at least in part on both the standardized theta-beta ratio and standardized ADHD rat... | LEXICOR MEDICAL TECHNOLOGY, LLC | A61B005/0476
US20080249381 | A1 | 2008-10-09 | Combined Apparatus For Measuring the Blood Glucose Level From an Ocular Fluid | 2005-06-13 | A combined apparatus for measuring the blood glucose level from an ocular fluid comprises a hand-held fluorescence photometer (D1) for determining the glucose level of the ocular fluid and a device (D2) for the amperometric or photometric determination of the blood glucose level. The fluorescence photometer (D1) comprises: first irradiating means (7,8,10) for providing a pilot beam (1) for irradiation of the eye (3) in order to excite a first fluorescence comprising a first wavelength band; first detecting means (8,9,12,13) for measuring the intensity of the of the pupil fluorescence within the first wavelength band; second irradiating means (17,18,21) for providing a measurement beam (5) for irradiation of the eye in order to excite an ocular sensor (4) to emit a second fluorescence comprising a second wavelength band; second detecting means (18,19,20,22,23,24,25) for measuring the intensity of the second fluorescence within the second wavelength band. The first and second irradiating means are arranged such that the measurement beam irradiates the eye under a predetermined angle (a) between 0° and 90° and at a predetermined distance from the pilot beam. | 1-13. (canceled) 14. Combined apparatus for measuring the blood glucose level from an ocular fluid of a user, comprising a hand-held fluorescence photometer for determining the glucose level of the ocular fluid of the usera device for the amperometric or photometric determination of the blood glucose level of the user for calibration purpose,wherein the hand-held fluorescence photometer comprises: (a) first irradiating means for providing a pilot beam for irradiation of the eye of the user in order to excite a pupil fluorescence, i.e. a first fluorescence, comprising a first wavelength band, said pupil fluorescence traveling along a first optical path;(b) first detecting means arranged in the first optical path for measuring the intensity of the of the pupil fluorescence within the first wavelength band;(c) second irradiating means for providing a measurement beam for irradiation of the eye of the user in order to excite an ocular sensor which is in contact with the ocular fluid of the user and which—upon irradiation with said measurement beam—emits a second fluorescence comprising a second wavelength band, said second fluorescence traveling along a second optical path;(d) second detecting means arranged in the second optical path for measuring the intensity of the second fluorescence within the second wavelength band;wherein the first and second irradiating means are arranged such that the measurement beam irradiates the eye of the user under a predetermined angle (α) and at a predetermined distance from the pilot beam, said angle (α) being between 0° and 90°. 15. Apparatus according to claim 14, wherein the predetermined angle (α) and distance are chosen such, that once the fluorescence photometer is arranged at a proper distance and orientation in front of the eye of the user, the measuring beam impinges on the iris while the pilot beam impinges on the pupil. 16. Apparatus according claim 14, wherein the first detecting means comprise a first detector for detecting the pupil fluorescence or first fluorescence and at least one wavelength selective element arranged in the first optical path for preventing light having a wavelength outside the first wavelength band from impinging on the first detector. 17. Apparatus according to claim 14, wherein the second detecting means comprise at least a second detector for detecting the second fluorescence emitted by the ocular sensor and at least one wavelength selective element for preventing light having a wavelength outside the second wavelength band from impinging on the second detector. 18. Apparatus according to claim 17, wherein the second detecting means further comprise a third detector for detecting a third fluorescence emitted by the ocular sensor and at least one wavelength selective element for preventing light having a wavelength outside the third wavelength band from impinging on the third detector. 19. Apparatus according to claim 14, wherein the hand-held fluorescence photometer and ... | Eyesense AG | A61B005/00
US20080255560 | A1 | 2008-10-16 | Fracture Fixation and Site Stabilization System | 2005-05-20 | A system for percutaneous fixation and stabilization of a fracture with a spanning, expandable structural frame placed in the intramedullary canal of the bone, comprising a column of surgical fluid such as bone cement, within which the structural frame acts as a reinforcing cage, and a sheath positioned at the fracture site and at least partially surrounding the frame and fluid. The fluid may be supported by a restrictor and may be agitated with a vibrating probe to remove entrapped air. The structural frame may be self-expanding or opened by an internal force, and it may be retrievable. The frame, fluid, or sheath may contain antibiotics, pharmaceuticals, or other therapeutic compounds to be delivered to the fracture site. | 1. In a system for treating a fracture site in a bone having an intramedullary canal, a structural frame comprising: an elongate size and shape suitable for insertion into said canal;a length sufficient to span said fracture site; anda contour adapted to engage an internal surface of said canal. 2. The structural frame of claim 1, wherein said structural frame is expandable from a first cross-sectional size to a second cross-sectional size which is larger than said first size. 3. The structural frame of claim 2, wherein said structural frame is biased toward expansion to said second size, such that said structural frame requires a retainer in order to remain at said first size. 4. The structural frame of claim 2, wherein said structural frame is configured to receive an expandable member which, upon expansion, expands said structural frame from said first size to said size. 5. The structural frame of claim 2, wherein, after expansion, said structural frame is collapsible from said second size to said first size. 6. The structural frame of claim 2, wherein, after expansion, said structural frame is collapsible from said second size to said first size and retrievable from said canal. 7. The structural frame of claim 2, wherein said structural frame is biased toward remaining at said second size once expanded. 8. The structural frame of claim 2, wherein said structural frame resists collapsing forces once expanded to said second size. 9. The structural frame of claim 2, further comprising a locking member configured to maintain said structural frame at said second size and resist collapsing forces. 10. The structural frame of claim 1, wherein said frame has a generally hollow, substantially tubular shape. 11. The structural frame of claim 1, wherein said structural frame is made of a biocompatible material. 12. The structural frames of claim 1, wherein said structural frame is made of a bioabsorbable material. 13. The structural frame of claim 1, wherein said structural frame carries a therapeutic agent. 14. The structural frame of claim 1, wherein said frame has a plurality of discrete components nested together. 15. The structural frame of claim 1, further comprising a plurality of elongate members connected to one another by one or more linking members. 16. The structural frame of claim 1, further comprising: a generally central elongate shaft; anda plurality of other elongate members, each connected to said shaft by one or more linking members. 17. The structural frame of claim 1, further comprising: a generally central elongate shaft sized and shaped to receive a guide wire having near its distal end one or more fins disposed thereon;a plurality of other elongate members, each connected to said shaft by one or more linking members,said one or more linking members positioned such that distal motion against said one or more fins urges said plurality of other elongate members away from said shaft, thereby expanding said structural ... | MYERS SURGICAL SOLUTIONS, LLC | A61B017/56 A61B005/103 A61N001/00
US20080255477 | A1 | 2008-10-16 | Method of Neuromusculoskeletal Proprioceptive Re-education and Development of a Living Body and Apparatus Therefor | 2008-06-20 | In combination, vertically percussive vibration and non-symmetrical forces are applied to a patient. The non-symmetrical forces may be generated by a variety of unevenly applied weights that are supported at appropriate body locations in a manner that upsets normal body load symmetry. The non-symmetrical forces in combination with vertically percussive vibrations induce neuromusculoskeletal proprioceptive re-education and development within a living body to correct musculo-skeletal disorders including but not limited to scoliosis. A special posture modifying chair and head-orienting glasses may be used instead of or in addition to unevenly applied weights. An additional force concentrating treatment apparatus is also presented. | 1. A method of neuromusculoskeletal proprioceptive re-education and development of a living body, comprising the steps of: identifying a neuromusculoskeletal proprioceptive disorder;determining at least one unbalanced external force magnitude and orientation which is unbalanced about a body axis of symmetry to upset a typical body orientation in a corrective manner;applying at least one unbalanced external force at said at least one unbalanced external force load magnitude and orientation unbalanced about said axis of symmetry; andinducing vibratory energy in said body during said applying step to induce said neuromusculoskeletal proprioceptive re-education and development. 2. A method of using a vibrating device to provide localized vibratory excitation to a joint in a direction generally perpendicular to adjacent bones, localized to a specific location of a living body adjacent said joint, and having direction and magnitude of vibratory excitation to correct a musculo-skeletal disorder within said living body, comprising the steps of: providing an elongate vibrating head having a longitudinal axis parallel to a direction of elongation and a vibratory motion transverse thereto, and further having an operatively incompressible dimension transverse to said longitudinal axis of a substantially similar order of magnitude to a dimension of said joint between said adjacent bones;placing said elongate vibrating head generally perpendicular to said adjacent bones and generally adjacent said joint;producing said vibratory motion; andcoupling said vibratory motion from said vibrating head to said joint with sufficient magnitude to correct a neuromusculoskeletal disorder within said living body. 3. The method of using a vibrating device to provide localized vibratory excitation to a joint of claim 2, further comprising the steps of: supporting said living body upon a generally fixed support; andseparating a limited portion of said living body from said support to form a gap prior to said placing step;said placing step further comprising the steps of: positioning said elongate vibrating head within said gap; andallowing said body to rest under at least a gravitational force against said vibrating device. 4. The method of using a vibrating device to provide localized vibratory excitation to a joint of claim 3, wherein said supporting step further comprises positioning said living body horizontally extensive upon said generally fixed support. 5. The method of using a vibrating device to provide localized vibratory excitation to a joint of claim 2, further comprising the step of applying additional massive load to at least a portion of said specific location of said living body to upset a typical body orientation about said joint in a manner corrective of said musculo-skeletal disorder within said living body. 6. The method of using a vibrating device to provide localized vibratory excitation to a joint of claim 2, wherein said joint and adjacent bones fu... | - | A61B005/103 A61H001/00
US20080312514 | A1 | 2008-12-18 | Serum Patterns Predictive of Breast Cancer | 2006-05-12 | Models for classifying a biological sample are developed from samples taken from a mammalian subject into one of at least two possible biological states related to breast cancer. Samples may be processed by mass spectral and other high-throughput analytical techniques. | 1. A model for classifying a biological sample taken from a mammalian subject into one of at least two possible biological states related to breast cancer using a data stream that is obtained by performing a mass spectral analysis of the biological sample, the data stream including magnitude values for a range of mass-to-charge values, comprising: at least one classifying hypervolume associated with one of the at least two biological states related to breast cancer and disposed within a vector space having n dimensions, each dimension corresponding to a different mass-to-charge value;wherein n is at least three and at least a first of the dimensions corresponds to a mass-to-charge value in a range of m/z values selected from the m/z ranges consisting of between 200 to 300, 300 to 400, 400 to 500, 500 to 600, 600 to 700, and 700 to 900. 2. The model of claim 1, wherein n is at least 5. 3. The model of claim 1, wherein n is between 5 and 25. 4. The model of claim 1, wherein at least a second of the dimensions corresponds to a mass-to-charge value of between 500 and 1100. 5. The model of claim 1, wherein at least a second of the dimensions corresponds to a mass-to-charge value of between 500 and 900. 6. The model of claim 1, wherein at least a second of the dimensions corresponds to a mass-to-charge value of between 700 and 900. 7. The model of claim 1, the at least one classifying hypervolume being a first classifying hypervolume, further comprising: a second classifying hypervolume disposed within the vector space;the first classifying hypervolume being associated with a presence of breast cancer, the second classifying hypervolume being associated with an absence of breast cancer. 8. The model of claim 1, the at least one classifying hypervolume being a first classifying hypervolume, further comprising: a second classifying hypervolume disposed within the vector space;the first classifying hypervolume and the second classifying hypervolume being associated with a presence of breast cancer. 9. The model of claim 1, the at least one classifying hypervolume being a first classifying hypervolume, further comprising: a second classifying hypervolume disposed within the vector space;the first classifying hypervolume and the second classifying hypervolume being associated with an absence of breast cancer. 10. The model of claim 1, wherein the classifying hypervolume is associated with a presence of breast cancer. 11. The model of claim 10, wherein the classifying hypervolume is associated with a presence of in situ breast cancer. 12. The model of claim 10, wherein the classifying hypervolume is associated with a presence of invasive breast cancer. 13. The model of claim 10, wherein the classifying hypervolume is associated with a likelihood of metastasis of the invasive breast cancer. 14. The model of claim 1, wherein the classifying hypervolume is associated with an absence of breast cancer. 15. The model of claim 14, wherein the classify... | - | A61B005/00
US20080243008 | A1 | 2008-10-02 | Blood Pressure Measuring Apparatus and Blood Pressure Measuring Method | 2005-10-03 | To provide a technique which makes it possible to obtain a proper pulse wave signal for high-accuracy blood pressure measurements, thereby saving the trouble of taking measurements repeatedly and reducing a physical burden imposed on the user by cuff pressure. A blood pressure measuring apparatus includes a cuff which is attached to and around an external ear; a first pulse wave detector and a second pulse wave detector which detect a pulse wave in a part squeezed by the cuff and which are affected differently from each other by a characteristic of body movements; a body movement detecting means which detects the characteristic of body movements; a pulse wave selecting means which selects a pulse wave detected by one of the first pulse wave detector and the second pulse wave detector based on the characteristic of body movements detected by the body movement detecting means; and a blood pressure value deriving means which derives a blood pressure value based on the pulse wave selected by the pulse wave selecting means. | 1. A blood pressure measuring apparatus comprising: a cuff which is attached to and around an external ear;a first pulse wave detector and a second pulse wave detector which detect a pulse wave in a part squeezed by said cuff and which are affected differently from each other by a characteristic of body movements;body movement detecting means which detects the characteristic of body movements;pulse wave selecting means which selects a pulse wave detected by one of said first pulse wave detector and said second pulse wave detector based on the characteristic of body movements detected by said body movement detecting means; andblood pressure value deriving means which derives a blood pressure value based on the pulse wave selected by said pulse wave selecting means. 2. The blood pressure measuring apparatus according to claim 1, wherein: said body movement detecting means comprises level detecting means which detects magnitude of the body movements; andsaid pulse wave selecting means selects a pulse wave for use to derive blood pressure based on the magnitude of body movements detected by said level detecting means. 3. The blood pressure measuring apparatus according to claim 2, wherein: said body movement detecting means further comprises a period detecting means which detects a period of the body movements; andsaid pulse wave selecting means selects a pulse wave for use to derive blood pressure based on the magnitude of body movements detected by said level detecting means and the period of the body movements detected by said period detecting means. 4. A blood pressure measuring apparatus comprising: a first cuff which is attached to and around an external ear;a first pulse wave detector and a second pulse wave detector which detect a pulse wave in a part squeezed by said first cuff and which are affected differently from each other by a characteristic of body movements;body movement detecting means which detects the characteristic of body movements;first pulse wave selecting means which selects a pulse wave detected by one of said first pulse wave detector and said second pulse wave detector based on the characteristic of body movements detected by said body movement detecting means;first blood pressure value deriving means which derives a blood pressure value based on the pulse wave selected by said first pulse wave selecting means;a second cuff mounted in a different location from said first cuff; andblood pressure determining means which determines blood pressure by detecting a pulse wave in a part squeezed by said second cuff; andpressurization control means which synchronizes pressurization of said first cuff and said second cuff. 5. The blood pressure measuring apparatus according to claim 4, wherein said blood pressure determining means comprises: a third pulse wave detector and a fourth pulse wave detector which detect a pulse wave in the part squeezed by said second cuff and which are affected differently from each other by a charac... | TERUMO KABUSHIKI KAISHA NIPPON TELEGRAPH AND TELEPHONE CORPORATION | A61B005/02
US20080294020 | A1 | 2008-11-27 | SYSTEM AND METHOD FOR PHYSLOLOGICAL DATA READINGS, TRANSMISSION AND PRESENTATION | 2008-01-25 | Systems, methods, and computer program products for facilitating the reading, transmission and presentation of physiological data within a wireless body area network are disclosed. The remote collection and monitoring of a person's (e.g., patient's) physiological data and activity levels for the purposes of determining the well-being of the person, as well as making additional health status determinations based on the historical information and trends of the collected data are provided. The systems, methods, and computer program products disclosed herein, in varying embodiments, readily lend themselves to incremental component and functionality modifications, which allow for increased sensor data sources, accuracy, reliability and utility of the collected information, further solidifying the uniqueness and desirability of the systems methods and computer program products. | 1. A system for facilitating the reading, transmission and presentation of physiological data, comprising: (a) a patch capable of being attached to the body of a person, wherein the patch comprises a plurality of sensors, wherein each of the plurality of sensors is capable of measuring a physiological parameter indicative of the health status of the person;(b) a gateway device wearable by the person, wherein said gateway device is in short-range wireless communications with the patch in order to receive a plurality of physiological parameters from the plurality of sensors; and(c) a data center remotely located from said gateway device and capable of long-range wireless communications with the gateway device in order to receive the plurality of physiological parameters from the plurality of sensors;wherein a user is able to remotely monitor the health status of the person by utilizing a processing device capable of wireless communications with the data center. 2. The system of claim 1, wherein the processing device capable of wireless communications with the data center is one selected from a group consisting of a desktop computer, a laptop computer, a palmtop computer, a workstation, a set-top box, a mobile telephone, and a personal data assistant. 3. The system of claim 1, wherein the gateway device comprises: (d) a short-range transceiver in wireless communications with the patch; and(e) a long-range transceiver in wireless communications with the data center. 4. The system of claim 3, wherein the gateway device further comprises: (f) a plurality of gateway sensors, wherein each of the plurality of gateway sensors is capable of measuring a physiological parameter indicative of the health status of the person; and wherein each of the plurality of gateway sensors measure a physiological parameter different than the plurality of sensors located on the patch. 5. The system of claim 1, further comprising: (d) a call center remotely located from the gateway device and capable of long-range, two-way wireless voice communications with the gateway device in order to receive indications of an alert condition. 6. The system of claim 1, wherein the data center comprises at least one server capable of analyzing at least one of the plurality of physiological parameters received from the gateway device. 7. The system of claim 6, wherein the data center further comprises at least one server capable of presenting at least one of the plurality of physiological parameters to the user utilizing the processing device. 8. The system of claim 1, wherein the gateway device is in short-range wireless communications with the patch via a wireless communications protocol selected from a group consisting of: ZigBee, Wi-Fi, ANT, and Bluetooth. 9. The system of claim 1, wherein the gateway device is in long-range wireless communications with the data center via a cellular communications protocol. 10. The system of claim 1, wherein the data center comprises an identi... | - | A61B005/02
US20080317313 | A1 | 2008-12-25 | SYSTEM AND METHOD FOR TRACKING MOTION FOR GENERATING MOTION CORRECTED TOMOGRAPHIC IMAGES | 2008-05-08 | A method and related system for generating motion corrected tomographic images includes the steps of illuminating a region of interest (ROI) to be imaged being part of an unrestrained live subject and having at least three spaced apart optical markers thereon. At least one camera is used to obtain images of the markers. Motion data comprising 3D position and orientation of the markers relative to an initial reference position is then calculated. The at least three spaced apart optical markers and the at least one camera are sufficient in quantity and position to avoid multiple epipolar solutions. Motion corrected tomographic data obtained from the ROI using the motion data is then obtained, where motion corrected tomographic images obtained therefrom. | 1. A method for generating motion corrected tomographic images, comprising the steps of: illuminating a region of interest (ROI) to be tomographically imaged, wherein said ROI is part of an unrestrained live subject having at least three spaced apart retro-reflective optical markers attached thereto, wherein said markers are proximate said ROI and each marker is either polarizing or depolarizing for an illuminating wavelength;acquiring filtered optical images of said markers from at least one filtered camera, wherein a polarization filter on the at least one filtered camera enables selective detection of illumination reflected by the at least three optical markers;calculating motion data comprising 3D position and orientation of said markers relative to an initial reference position, wherein the at least three spaced apart retro-reflective optical markers and the at least one filtered camera are sufficient in quantity and position to avoid multiple epipolar solutions; andmotion correcting tomographic data of said ROI obtained simultaneously with said motion data using said motion data, wherein motion corrected tomographic images are obtained. 2. The method of claim 1, wherein during said method said animal is disposed in a confinement volume which is optically transparent to said illumination wavelength. 3. The method of claim 1, wherein said tomographic images are single photon emission computed tomography (SPECT) images. 4. The method of claim 1, wherein said illumination is aligned to be approximately coaxial with an optical axis of said first camera. 5. The method of claim 1, wherein said illuminating is strobed illuminating. 6. The method of claim 5, wherein said calculating motion data step comprises processing said images using a combination of segmentation, object feature extraction and digital filtering. 7. The method of claim 1, further comprising at least a fourth spaced apart retro-reflective optical marker, wherein at least four of the markers are arranged to eliminate multiple epipolar solutions. 8. The method of claim 1, comprising a first and a second filtered camera, wherein the acquiring step comprises acquiring simultaneous images from the first and second filtered cameras. 9. The method of claim 8, wherein said calculating motion data step comprises processing said simultaneous images using a combination of segmentation, object feature extraction and digital filtering. 10. The method of claim 1, comprising a first, a second, and a third filtered camera, wherein the acquiring step comprises acquiring simultaneous images of at least three of the spaced apart retro-reflective optical markers from at least two of the first, second, and third filtered cameras. 11. The method of claim 1, wherein said illumination is polarized. 12. The method of claim 1, wherein at least two of said markers are polarized and have different polarization characteristics. 13. The method of claim 1, wherein at least one of the at least one fi... | UT-Battelle, LLC | A61B005/055
US20080262340 | A1 | 2008-10-23 | Moving Table Mri | 2005-05-12 | The invention relates to a MRI system and to a method for producing an image with such an system. In order to provide a MR imaging technique with a high efficient MR signal acquisition, which provides a high level of comfort to a patient, a MRI system and method are suggested, where image data from an object are acquired while the object is moving with variable speed relative to the MRI system, and where the image data are combined and an image of the object is reconstructed. | 1. A magnetic resonance imaging (MRI) system, the system comprising an object carrier movable within the MRI system,a control unit for controlling the MRI system, the control unit being adaptedto control the movement of the object carrier,to derive the position of the object carrier,to acquire image data from the object, while the object is moving with variable speed relative to the MRI system, anda processing unit being adaptedto combine the image data acquired andto reconstruct an image of the object. 2. The MRI system as claimed in claim 1, wherein the control unit is adapted to drive the object carrier according to the part of k-space being acquired. 3. The MRI system as claimed in claim 1, wherein the control unit is adapted to acquire image data from an object, while the object is stationary relative to the MRI system, andto acquire image data from the object, while the object is moving relative to the MRI system. 4. The MRI system as claimed in claim 1, wherein the control unit is adapted to acquire image data of a well-defined portion of the object, whereby part of that data is acquired during displacement of that portion of the object with respect to a default position or a stationary position. 5. The MRI system as claimed in claim 1, wherein the control unit is adapted to acquire image data while the object shows a small displacement from the default position or the stationary position. 6. The MRI system as claimed in claim 1, wherein the control unit is adapted to shift the field of view in accordance to the displacement. 7. The MRI system as claimed in claim 1, wherein the control unit is adapted to acquire low k-values for the stationary part and high k-values for displacement. 8. MRI system as claimed in claim 1, wherein the control unit is adapted to select k-space parts according to the position of the object carrier. 9. The MRI system as claimed in claim 1, wherein the control unit is adapted to keep the field of view fixed with respect to a magnet of the MRI system andto perform frequency-encoding parallel to the motion direction, and wherein the processing unit is adaptedto perform a fourier-transformation in motion-direction of each profile, andto fill the data in (x,ky,kz)-space. 10. The MRI system as claimed in claim 1, wherein the control unit is adapted to acquire high k-values during movement with high velocity and to acquire low k-values during movement with low velocity. 11. The MRI system as claimed in claim 1, wherein the control unit is adapted to acquire high-resolution image data of part of the object while the object is in a stationary position and to acquire image data of other part of the object with lower resolution while the object is moving. 12. The MRI system as claimed in claim 1, wherein the control unit is adapted to acquire very high k-space values while the object is in a stationary position. 13. The MRI system as claimed in claim 1, wherein multi-element RF receive coils are used for the ob... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/055 G01R033/563
US07470236 | B1 | 2008-12-30 | Electromyography system | 2000-11-24 | A method for detecting the presence of a nerve adjacent the distal end of at least one probe, comprising: determining relative neuro-muscular response onset values for a plurality of spinal nerves; emitting a stimulus pulse from a probe or surgical tool; detecting neuro-muscular responses to the stimulus pulse in each of the plurality of spinal nerves; and concluding that the electrode disposed on the distal end of the at least one probe is positioned adjacent to a first spinal nerve when the neuro-muscular response detected in the first spinal nerve is detected as a current intensity level less than or equal to a corresponding neuro-muscular response onset value of the first spinal nerve. | 1. A method for assessing the proximity of a spinal nerve relative to a distal end of at least one probe or surgical tool being introduced towards at least one of a lumbar region and thoracic region of a patient's spine, said lumbar region and said thoracic region of said spine having a ventral column and a dorsal column, said ventral column including a plurality of vertebral bodies and a plurality of intervertebral discs disposed in between said vertebral bodies, said vertebral bodies and intervertebral discs each having an anterior aspect, a posterior aspect opposite from said anterior aspect, and a lateral aspect extending between said anterior and posterior aspects, said dorsal column including a plurality of bone elements extending from said vertebral bodies to form a spinal canal that contains and protects the spinal chord, said spinal nerve exiting from said spinal canal and disposed generally parallel to a longitudinal axis of said spine along said lateral aspect, the method comprising: (a) emitting a stimulus signal from an electrode disposed on a probe or surgical tool as said probe or tool is introduced towards a lateral aspect of at least one of a vertebral body and an intervertebral disc of at least one of a lumbar region and thoracic region of a patient's spine;(b) electromyographically monitoring muscles coupled to said spinal nerve to determine if a predetermined neuro-muscular response is elicited by the stimulus signal;(c) increasing the intensity level of said stimulus signal until said predetermined neuro-muscular response is elicited by said stimulus pulse and stopping the emission of said stimulus signal immediately after said predetermined neuro-muscular response is detected; and(d) communicating to an operator said intensity level of said stimulus signal required to elicit said predetermined neuro-muscular response, wherein said intensity level required to elicit said predetermined neuro-muscular response represents the proximity of said spinal nerve to said probe or surgical tool. 2. The method of claim 1, wherein the stimulus signal is emitted from an electrode disposed on the distal end of the at least one probe or surgical tool. 3. The method of claim 1, wherein detecting neuro-muscular responses involves detecting the neuro-muscular responses at a plurality of distally spaced apart myotome locations corresponding to each of a plurality of spinal nerves. 4. The method of claim 1, further comprising: repeating the method of claim 1 while the intensity level of the electrical stimulus signal is varied over time. 5. The method of claim 4, wherein the intensity level of the stimulus signal is varied incrementally. 6. The method of at least one of claims 4 and 5, wherein the intensity level of the stimulus signal is increased over time. 7. The method of claim 1, wherein said spinal nerve is one of a plurality of spinal nerves exiting from successive vertebrae. 8. The method of claim 1, wherein the method of claim 1 i... | NuVasive, Inc. | A61B005/05 A61N001/00
US07742801 | B2 | 2010-06-22 | Planning method and system for free-form implant modification | 2007-07-25 | A method for planning a bone implant ascertains a condition of a bone to be treated, and identifies a dysfunctional part of the bone. A free-form area then can be ascertained and registered, wherein the free-form area lies below the dysfunctional part. A control data set then is produced that forms the basis for ablating the bone on the ascertained free-form area. | 1. A method for planning a bone implant, comprising: obtaining image data corresponding to a bone to be treated;identifying in the image data a dysfunctional part of the bone and a functional part of the bone;ascertaining a contour of an interface between the dysfunctional part of the bone and the functional part of the bone;registering the ascertained contour in a medical navigation system, wherein registering the contour includes registering the contour a predetermined distance below the interface;using a processor to produce a control data set based on the registered contour of the interface, said control data set forming the basis for ablating the bone; andoutputting the control data set. 2. The method according to claim 1, wherein identifying the dysfunctional part and the functional part of the bone includes using an imaging method to ascertain the dysfunctional part and the functional part of the bone. 3. The method according to claim 2, wherein using the imaging method includes using an x-ray method, a computer tomography method or a magnetic resonance tomography method. 4. The method according to claim 2, wherein identifying includes using graphical segmentation of image data obtained from the imaging method to identify the dysfunctional part of the bone. 5. The method according to claim 1, further comprising using a navigation-assisted ablating device to remove the dysfunctional part of the bone based on the control data set. 6. The method according to claim 1, wherein producing the control data set includes producing a control data set that includes information for guiding an ablating device. 7. The method according to claim 6, further comprising: positionally detecting and tracking the ablating device; andbased on the control data set and the detected and tracked position of the ablating device, providing information for guiding the ablating device via free hand. 8. The method according to claim 7, wherein positionally detecting and tracking the ablating device includes using a medical tracking system to detect and track the ablating device, and wherein the information for guiding the ablating device is provided by a medical navigation system communicatively coupled to said tracking system. 9. The method according to claim 6, further comprising: positionally detecting and tracking the ablating device, wherein the ablating device is mechanically mounted to a support structure; andbased on the control data set and the detected and tracked position of the ablating device, providing information for guiding the ablating device by hand. 10. The method according to claim 9, wherein positionally detecting and tracking the ablating device includes using a medical tracking system to detect and track the ablating device, and wherein the information for guiding the ablating device is provided by a medical navigation system communicatively coupled to said tracking system. 11. The method according to claim 9, wherein providing information... | BrainLAB AG | A61B005/05 G06G007/48 A61B006/00
US07749170 | B2 | 2010-07-06 | Multiple configurable electronic thermometer | 2007-05-22 | A plurality of thermometer models are manufactured from a universal clinical thermometer cure. The thermometer core includes a housing, a controller storing a software program, a power source, a probe, a probe cord, and a display. A plurality of model selection devices are produced, each corresponding to a different thermometer model. One of the model selection devices is selected and connected to the controller. The controller operates the thermometer as a function of the connected model selection device such that the thermometer is configured with the features and default settings of the thermometer model corresponding to the connected model selection device. | 1. A configurable thermometer comprising: a controller for operating the thermometer, said controller having a memory for storing one or more software instructions implementing a plurality of functions of the thermometer; anda model selection device identifying a selected one of a plurality of operating configurations of the thermometer, each of said configurations of the thermometer being defined by said model selection device being connected to the controller, said controller being responsive to the model selection device connected thereto for operating the thermometer in the selected configuration according to software instructions selected from the one or more software instructions, said selected software instructions corresponding to the model selection device, wherein the model selection device comprises one of a plurality of interchangeable user input devices, wherein each of said user input devices comprises a keypad overlay comprising one or more buttons, said overlay, when connected to the controller, selectively enabling execution of the selected software instructions by a user of the thermometer, each of said user input devices and said controller having mating connectors for connecting a selected one of the user input devices to the controller to configure the one or more software instructions within the thermometer according to the selected configuration. 2. The thermometer of claim 1 wherein the mating connectors comprise a set of pins and a corresponding set of pin-receiving contacts, at least two pins of the set of pins being electrically connected to the corresponding pin-receiving contacts when the selected user input device is connected to the controller, and wherein the controller selectively enables the one or more functions of the thermometer defining the selected configuration based on which pins of the set of pins are electrically connected to the corresponding pin-receiving contacts. 3. The thermometer of claim 1 wherein the controller is configured to selectively enable the one or more functions of the thermometer defining the selected configuration based on the model selector device connected thereto for operating the thermometer in the selected configuration. 4. The thermometer of claim 1 wherein the controller is configured to selectively disable at least one of the functions of the thermometer not defining the selected configuration based on the model selector device connected thereto for operating the thermometer in the selected configuration. 5. The thermometer of claim 1 further comprising a plurality of software profiles stored in the controller memory, said software profiles each being representative of a pre-defined set of thermometer functions corresponding to one of the configurations of the thermometer, and wherein the controller is configured to load the software profile corresponding to the selected configuration identified by the model selection device when the thermometer initializes. 6. The thermo... | Tyco Healthcare Group LP | A61B005/00 G01K001/08 G01K001/14 G01K013/00 G01K001/00
US07468036 | B1 | 2008-12-23 | Monitoring device, method and system | 2007-06-13 | A monitoring device (20) for monitoring the vital signs of a user is disclosed herein. The monitoring device (20) is preferably an article (25) having an optical sensor (30) and a circuitry assembly (35). The optical sensor (30) preferably comprises a photodetector (130) and a plurality of light emitting diodes (135). The monitoring device (20) preferably provides for the display of the following information about the user: pulse rate; blood oxygenation levels; calories expended by the user of a pre-set time period; target zones of activity; time; distance traveled; and/or dynamic blood pressure. | 1. A method of monitoring a user's vital signs, the method comprising: deactivating a processor for a deactivation period of 128 to 640 microseconds;activating an optical sensor during the deactivation period to generate a signal corresponding to the flow of blood through at least one artery of the user, the optical sensor comprising a photodetector and a plurality of light-emitting diodes;reactivating the processor to process the signal to obtain an average pulse value of a first set of time periods;processing the signal to obtain an average pulse value of a second set of time periods, the second set of time periods less than the first set of time periods;subtracting the average pulse value of the second set of time periods from the average pulse value of the first set of time periods to generate a first filtered pulse data value;processing the first filtered pulse data value to obtain a first heart rate value;using the first heart rate value to filter the subsequent average pulse values generated from the digital signal to obtain a real-time heart rate value of the user;processing the real-time heart rate data of the user for analysis of real-time calories expended by the user and for real-time display of a plurality of the user's vital signs; anddisplaying the plurality of user's vital signs on a display member, the display of the plurality of user's vital signs controlled by the user using a control component. 2. The method according to claim 1 further comprising filtering motion noise from the signal utilizing a motion sensor. 3. The method according to claim 1 wherein the plurality of light emitting diodes comprises a plurality of pulsed infrared light emitting diodes. 4. The method according to claim 1 wherein the plurality of light emitting diodes comprises a pulsed infrared light emitting diode and a green light emitting diode. 5. The method according to claim 1 wherein the real-time heart rate data is based on 60 to 180 sample readings and each of the sample readings is based on 4 to 15 sub-sample readings. 6. The method according to claim 1 wherein the processor, the optical sensor and the display member are disposed on a monitoring device. | Impact Sports Technology, Inc. | A61B005/00
US07452365 | B2 | 2008-11-18 | Load-controlled device for a patterned skin incision of constant depth | 2006-09-13 | A skin incision device has a casing with a bottom surface and a slot formed therein, a cover positioned on the casing and slidable in a direction toward the bottom surface, a blade positioned in the casing adjacent the slot, an actuator cooperative to a position between the cover and interior of the casing, and a carriage. The actuator engages the blade by its horizontal displacement triggered by slidable movement of the cover toward the bottom surface of the casing. The carriage element guides movement of the blade between a pre-actuated position and a post-actuated position. The blade extends outwardly of the bottom surface with a fixed depth of cut during a horizontal movement of the blade through the slot. | 1. A skin incision device comprising: a casing having a slot formed at a bottom surface thereof;a cover positioned on said casing, said cover being slidable in a direction transverse to a plane of said bottom surface of said casing;a blade pivotally positioned within said casing generally adjacent said slot, said blade having a pre-actuated position and a post-actuated position;an actuator means anchored to said casing, said actuator means being actuatable by slidable movement of said cover toward said bottom surface, said actuator means for converting transverse movement of said cover into horizontal movement of said actuator means, said actuator means engaging said blade at said pre-actuated position through said post-actuated position such that at least a portion of said blade extends outwardly of said slot during the movement between said pre-actuated position and said post-actuated position; anda carriage means anchored to said casing and pivotally attached to said blade, said carriage means for guiding movement of said blade from said pre-actuated position through said post-actuated position, said carriage means comprising: an arcuate member;a cam formed at one end of said arcuate member, said cam being pivotally affixed to said blade; andan anchor end formed at an opposite end of said arcuate member from said cam, said anchor end being affixed to said casing, said blade comprising: a razor member having a cutting edge;a retaining hole positioned adjacent an end of said blade, said cam being pivotally affixed to said retaining hole; andan elongated hole positioned adjacent an opposite end of said blade, said elongated hole being rotatably interconnected to said casing, said casing having a blade retainer peg formed therein adjacent said slot of said casing, said elongated hole of said blade positioned over said blade retainer peg, said elongated hole being an obround aligned with a longitudinal axis of said blade. 2. A skin incision device comprising: a casing having a slot formed at a bottom surface thereof;a cover positioned on said casing, said cover being slidable in a direction transverse to a plane of said bottom surface of said casing;a blade pivotally positioned within said casing generally adjacent said slot, said blade having a pre-actuated position and a post-actuated position;an actuator means anchored to said casing, said actuator means being actuatable by slidable movement of said cover toward said bottom surface, said actuator means for converting transverse movement of said cover into horizontal movement of said actuator means, said actuator means engaging said blade at said pre-actuated position through said post-actuated position such that at least a portion of said blade extends outwardly of said slot during the movement between said pre-actuated position and said post-actuated position; anda carriage means anchored to said casing and pivotally attached to said blade, said carriage means for guiding movement of said blad... | Helena Laboratories | A61B017/32 A61B005/151
US07852082 | B2 | 2010-12-14 | Remote monitoring system and method | 2008-07-22 | A remote control system for a modulatable device is provided. The remote control system comprises a receiver system coupled to the modulatable device and configured to obtain an output characteristic of the modulatable device, the receiver system being located remotely with respect to the modulatable device. The system further comprises a command signal setting system coupled to the receiver system and configured to use the output characteristic to generate a drive command signal and a bias system coupled to the command signal setting system and configured to receive the drive command signal and set a bias point of the modulatable device based on the drive command signal. The bias system is located locally with respect to the modulatable device. The command signal setting system and the bias system are coupled via a first optical conduit. | 1. A magnetic resonance (MR) system including a remote control system for a modulatable device, the magnetic resonance system comprising: a modulatable device located within a magnetic field of the magnetic resonance system and configured to provide an output characteristic of the modulatable device to an output characteristic receive and command signal setting assembly situated outside the magnetic field;a bias system optically coupled to the output characteristic receive and command signal setting assembly and coupled to the modulatable device, the bias system configured to receive a drive command signal from the output characteristic receive and command signal setting assembly and to set a bias point of the modulatable device based on the drive command signal;wherein the output characteristic comprises an output power and wherein the output characteristic receive and command signal setting assembly is configured to generate the drive command signal based on a maximum output power and a minimum output power generated by the modulatable device. 2. The MR system of claim 1, wherein the modulatable device is coupled to output characteristic receive and command signal setting assembly via an optical conduit. 3. The MR system of claim 1, wherein the modulatable device comprises a device selected from the group consisting of optical modulators, RF modulators, and laser diodes. 4. The MR system of claim 1, wherein the bias system is located within the magnetic field. 5. The MR system of claim 1, wherein the output characteristic receive and command signal setting assembly is configured to generate the drive command signal by mapping a transfer function and obtaining an optimum position for the drive command signal. 6. The MR system of claim 5, wherein the modulatable device comprises an optical modulator. 7. The MR system of claim 1, wherein the output characteristic receive and command signal setting assembly is configured to generate the drive command signal by taking an average of a first drive signal corresponding to the maximum output power and a second drive signal corresponding to the minimum output power. 8. The MR system of claim 1, wherein the output characteristic receive and command signal setting assembly is configured to generate the drive command signal based on output characteristic signal distortion. 9. The MR system of claim 8, wherein the output characteristic receive and command signal setting assembly further comprises a pilot tone generator and a band pass filter centered about an expected second order output of the pilot tone generator, and wherein the output characteristic receive and command signal setting assembly is further configured to generate a drive command signal corresponding to a minimum output signal of the band pass filter. 10. The MR system of claim 9, wherein the pilot tone generator is configured to provide a pilot tone to the modulatable device. 11. The MR system of claim 9, wherein the pilot tone is ... | General Electric Company | G01V003/00
US07465273 | B2 | 2008-12-16 | Method for monitoring pre-eclamptic patients | 2007-11-30 | A system and method for monitoring and estimating the blood pressure of a pregnant patient that modifies the blood pressure estimating algorithm when the patient is pre-eclamptic. The level of carbon monoxide within a patient's bloodstream or exhaled breath can be analyzed to determine whether a pregnant patient is pre-eclamptic. After the patient has been diagnosed as pre-eclamptic, the NIBP monitoring system adjusts its algorithm for estimating the patient's blood pressure to compensate for the physical changes that occur in the patient during pre-eclampsia. The adjusted blood pressure estimates calculated by the NIBP monitoring system can be calculated using different adjustment techniques and methods and are displayed on the NIBP monitor. | 1. A method of enhancing the operation of a non-invasive blood pressure (NIBP) monitor having a blood pressure cuff selectively inflatable and deflatable to determine at least a systolic pressure, a mean arterial pressure and a diastolic pressure for a patient, the method comprising the steps of: receiving an indication that the patient is pregnant at the NIBP monitor;receiving an indication that the patient is pre-eclamptic at the NIBP monitor;adjusting at least the determined systolic pressure and the determined diastolic pressure to compensate for the effects of pre-eclampsia; anddisplaying the adjusted blood pressure estimates. 2. The method of claim 1 further comprising the steps of: receiving an indication of the carbon monoxide level in the blood of the patient at the NIBP monitor; anddetermining whether the patient is pre-eclamptic based upon the carbon monoxide level. 3. The method of claim 2 further comprising the steps of: positioning a carbon monoxide detection device on the patient to detect the level of carbon monoxide in the patient's blood;generating a signal related to the carbon monoxide level; andreceiving the signal at the NIBP monitor. 4. The method of claim 3 wherein the carbon monoxide detection device is a pulse oximeter that detects the carboxyhemoglobin (COHb) level in the patient's blood. 5. The method of claim 3 wherein the carbon monoxide detection device detects the level of carbon monoxide in the patient's expired breath. 6. The method of claim 1 wherein the step of adjusting the blood pressure estimates includes: increasing the estimated systolic pressure by a first offset value when the patient is pre-eclamptic; andincreasing the estimated diastolic pressure by a second offset value when the patient is pre-eclamptic. 7. The method of claim 1 wherein the step of adjusting the blood pressure estimates includes: adjusting the ratio of the systolic pressure from an estimated mean arterial pressure when the patient is pre-eclamptic; andadjusting the ratio of the diastolic pressure from the mean arterial pressure when the patient is pre-eclamptic. 8. A method of enhancing the operation of a non-invasive blood pressure (NIBP) monitor having a blood pressure cuff selectively inflatable and deflatable to determine at least a systolic pressure, a mean arterial pressure and a diastolic pressure for a patient, the method comprising the steps of: receiving an indication that the patient is pre-eclamptic at the NIBP monitor;adjusting at least the determined systolic pressure and the determined diastolic pressure to compensate for the effects of pre-eclampsia; anddisplaying the adjusted systolic and diastolic pressure estimates. 9. The method of claim 8 wherein the step of adjusting the blood pressure estimates includes: increasing the estimated systolic pressure by a first offset value when the patient is pre-eclamptic; andincreasing the estimated diastolic pressure by a second offset value when the patient is pre-eclam... | The General Electric Company | A61B005/02 A61B005/08 A61B005/00
US07853327 | B2 | 2010-12-14 | Heart sound tracking system and method | 2007-04-17 | A system and method provide heart sound tracking, including an input circuit, configured to receive heart sound information, and a heart sound recognition circuit. The heart sound recognition circuit can be coupled to the input circuit and can be configured to recognize, within a particular heart sound of a particular heart sound waveform, a first intra heart sound energy indication and a corresponding first intra heart sound time indication using the heart sound information from the particular heart sound waveform and the heart sound information from at least one other heart sound waveform. The particular heart sound can include at least a portion of one of S1, S2, S3, and S4. Further, the first intra heart sound energy indication and the corresponding first intra heart sound time indication can correspond to the at least a portion of one of S1, S2, S3, and S4, respectively. | 1. A system, comprising: an input circuit, configured to receive heart sound information;a heart sound recognition circuit, coupled to the input circuit, configured to recognize, within a particular heart sound of a particular heart sound waveform, a first intra heart sound energy indication and a corresponding first intra heart sound time indication using the heart sound information from the particular heart sound waveform and the heart sound information from at least one other heart sound waveform;wherein the particular heart sound of the particular heart sound waveform includes an ensemble averaged heart sound over multiple physiological cycles;wherein the particular heart sound includes at least a portion of one of S1, S2, S3, and S4; andwherein the first intra heart sound energy indication and the corresponding first intra heart sound time indication corresponds to the at least a portion of one of S1, S2, S3, and S4, respectively. 2. The system of claim 1, wherein the input circuit is configured to receive S2 information; and wherein the heart sound recognition circuit is configured to recognize, within the S2 of the particular heart sound waveform, an intra S2 energy indication and a corresponding intra S2 time indication using the S2 information from the particular heart sound waveform and the S2 information from at least one other heart sound waveform. 3. The system of claim 1, wherein each physiological cycle of the multiple physiological cycles includes a cardiac cycle. 4. The system of claim 1, wherein the at least one other heart sound waveform includes an immediately preceding heart sound waveform to the particular heart sound waveform. 5. The system of claim 1, wherein the heart sound recognition circuit is configured to recognize the first intra heart sound energy indication and the corresponding first intra heart sound time indication using a heart sound template. 6. The system of claim 5, wherein the heart sound template includes at least a portion of the heart sound information from the at least one other heart sound waveform. 7. The system of claim 6, wherein the heart sound recognition circuit is configured to recognize the first intra heart sound energy indication and the corresponding first intra heart sound time indication using a cross correlation between at least a portion of the heart sound information from the particular heart sound waveform and at least a portion of the heart sound template. 8. The system of claim 1, including: an implantable medical device, the implantable medical device including: a heart sound sensor, configured to sense a heart sound signal; anda heart sound feature detector, coupled to the heart sound sensor, the heart sound feature detector configured to detect at least one heart sound feature using the heart sound signal; andwherein the heart sound information includes the at least one heart sound feature. 9. The system of claim 1, wherein the heart sound recognition circuit includes a d... | Cardiac Pacemakers, Inc. | A61N001/365
US07460898 | B2 | 2008-12-02 | Dual electrode system for a continuous analyte sensor | 2006-10-04 | Disclosed herein are systems and methods for a continuous analyte sensor, such as a continuous glucose sensor. One such system utilizes first and second working electrodes to measure additional analyte or non-analyte related signal. Such measurements may provide a background and/or sensitivity measurement(s) for use in processing sensor data and may be used to trigger events such as digital filtering of data or suspending display of data. | 1. An analyte sensor configured for insertion into a host for measuring an analyte in the host, the sensor comprising: a first working electrode comprising a first electroactive surface disposed beneath an active enzymatic portion of a sensor membrane, wherein the first working electrode is configured to generate a first signal; anda second working electrode comprising a second electroactive surface disposed beneath an inactive-enzymatic or a non-enzymatic portion of a sensor membrane, wherein the second working electrode is configured to generate a second signal, wherein the first electroactive surface and the second electroactive surface are configured and arranged to be equivalently influenced by in vivo environmental factors when implanted in a host such that a signal component associated with the in vivo environmental factors can be removed from the first signal, wherein the first and second electrodes are configured and arranged around a circumference of the sensor along an overlapping length of the sensor. 2. The sensor of claim 1, wherein the first working electrode and the second working electrode are coaxial. 3. The sensor of claim 1, wherein at least one of the first working electrode and the second working electrode is twisted or helically wound to integrally form at least a portion of the sensor. 4. The sensor of claim 3, wherein the first working electrode and the second working electrode are twisted together to integrally form an in vivo portion of the sensor. 5. The sensor of claim 1, wherein the first working electrode and the second working electrode each comprise a first end and a second end, wherein the first ends are configured for insertion in the host, and wherein the second ends are configured for electrical connection to sensor electronics. 6. The sensor of claim 5, wherein the second ends are coaxial. 7. The sensor of claim 5, wherein the second ends are stepped. 8. The sensor of claim 1, further comprising at least one additional electrode selected from the group consisting of a reference electrode and a counter electrode. 9. The sensor of claim 8, wherein the additional electrode, together with the first working electrode and the second working electrode, integrally form at least a portion of the sensor. 10. The sensor of claim 8, wherein the additional electrode is located at a position remote from the first and second working electrodes. 11. The sensor of claim 8, wherein a surface area of the additional electrode is at least six times a surface area of at least one of the first electroactive surface and the second electroactive surface. 12. The sensor of claim 1, wherein the sensor is configured for implantation into the host. 13. The sensor of claim 12, wherein the sensor is configured for subcutaneous implantation in a tissue of a host. 14. The sensor of claim 12, wherein the sensor is configured for indwelling in a blood stream of a host. 15. The sensor of claim 1, wherein the sensor substantially c... | DexCom, Inc. | A61B005/00 A61B005/05
US07435226 | B2 | 2008-10-14 | Method and apparatus for displaying information obtained by electrical impedance tomography data | 2002-12-06 | The invention refers to a method and an apparatus for displaying information obtained by electrical impedance tomography (EIT) data from a part of a patient's body. In order to provide a method for displaying information which can be immediately used in order to determine a specific pathological condition of said part, a method according to the invention performs te following steps: processing the EIT data with a plurality of predetermined processing modes; determining specific pathological conditions of said part in accordance with the results of the predetermined processing modes; selecting a screen mode from a plurality of predetermined screen modes in accordance with the specific pathological conditions; and displaying the EIT data in accordance with the selected screen modes. | 1. Method for displaying information obtained by electrical impedance tomography (EIT) data from a part of a patient's body on an EIT monitor, comprising the steps of: implementing on said EIT monitor a plurality of predetermined processing modes and a plurality of predetermined screen modes,providing the predetermined screen modes with functional images having an online imaging corresponding to one of the processing modes and with screen modes yielding further data reduction compared with the functional images,processing the EIT data with said plurality of predetermined processing modes,determining specific pathological conditions of said part in accordance with the results of the predetermined processing modes, selecting a functional image as a screen mode which corresponds to a respective processing mode, if said processing mode leads to the determination of a pathological condition, and selecting a screen mode which provides a further data reduction, if not a pathological condition is determined, anddisplaying the EIT data in accordance with the selected screen modes. 2. Method according to claim 1, wherein the screen mode is updated on demand of a user. 3. Method according to claim 1, wherein the displayed EIT data are updated in real time. 4. Method according to claim 1, wherein the EIT data are calculated on the basis of a running data window. 5. Method according to claim 1, wherein an estimate of a pathological condition is input by a user and wherein processing of the EIT data is carried out with a subcombination of the predetermined processing modes in accordance with the estimated pathological condition. 6. Method according to claim 1, wherein a screen mode consists of graphic elements which are pre-stored in a data store as graphic patterns and which are fetched from the data store as soon as the corresponding screen mode is selected. 7. Method according to claim 1, wherein said part to be examined is a patient's lung or selected regions thereof. 8. Method according to claim 7, wherein a relative mode is implemented as a predetermined processing mode which yields as a result the mean variation of impedance changes due to ventilation. 9. Method according to claim 8, wherein a perfusion mode is implemented as a predetermined processing mode which yields as a result the mean variation of impedance changes due to lung perfusion, and a perfusion ratio mode is implemented which calculates the ratio of the results of the relative mode and the perfusion mode and which yields as a result the ventilation/perfusion ratio. 10. Method according to claim 8, wherein a perfusion mode is implemented as a predetermined processing mode which yields as a result the mean variation of impedance changes due to lung perfusion, an absolute mode is implemented as a predetermined processing mode which yields as a result the mean absolute impedance, andthe lung condition of a pulmonary infarct is determined, if the relative mode yields a low result, the... | Stephan Bohm Marcelo B. P. Amato | A61B005/08 A61B005/05
US07434991 | B2 | 2008-10-14 | Thermal tympanic thermometer | 2006-05-19 | A tympanic thermometer includes a thermally conductive nozzle extending from a distal end of the thermometer. A base of a sensor can is thermally connected to the nozzle to define a path of conductive heat transfer from the nozzle to the base of the can thereby minimizing a thermal gradient between proximal and distal ends of the sensor can when temperature is measured in the ear. | 1. A tympanic thermometer having a proximal end and a distal end, the thermometer comprising: a thermally conductive nozzle extending from the distal end of the thermometer, the nozzle being located and shaped to receive a probe cover thereon,a sensor can housing temperature sensing electronics for sensing temperature, the sensor can including a base thermally connected to the nozzle to define a path of conductive heat transfer from the nozzle to the base of the can, the sensor can being otherwise free of thermal connection with a thermally conductive structure of greater mass than the sensor can,the thermometer in combination with the probe cover thermally connected to the nozzle as part of the path of conductive heat transfer, wherein heat from outside the thermometer is transferred from the probe cover through the nozzle to the base of the sensor can via the path of conductive heat transfer, and further wherein the probe cover is thermally connected to the nozzle solely at a distal end margin of the nozzle. 2. A tympanic thermometer as set forth in claim 1 wherein the base of the sensor can is secured directly to the nozzle. 3. A tympanic thermometer as set forth in claim 2 wherein the base of the sensor can is secured to an internal ridge of the nozzle. 4. A tympanic thermometer as set forth in claim 3 further comprising a non-thermally conductive retainer secured to the nozzle to sandwich the base of the sensor can between the internal ridge and the retainer. 5. A tympanic thermometer as set forth in claim 3 wherein the base of the sensor can is secured to the nozzle using a thermally-conductive adhesive. 6. A tympanic thermometer as set forth in claim 5 wherein a distal surface of the base of the sensor can is secured to the internal ridge of the nozzle and an opposite proximal surface of the base is free from contact with any structure. 7. A tympanic thermometer as set forth in claim 1 wherein the nozzle defines an air space adjacent a proximal side of the sensor can. 8. A tympanic thermometer as set forth in claim 7 wherein the thermometer is free of structure in thermal connection with the sensor can other than the nozzle. 9. A tympanic thermometer as set forth in claim 1 wherein the nozzle and sensor can are constructed and arranged to define a heat transfer path from the nozzle to the sensor can and to inhibit heat transfer away from the sensor can thereby minimizing a thermal gradient between proximal and distal ends of the sensor can when temperature is measured in the ear. 10. A tympanic thermometer as set forth in claim 1 wherein the nozzle is made of metal and the sensor can is free of thermal connection with a metallic object other than the nozzle. 11. A tympanic thermometer having a proximal end and a distal end, the thermometer comprising: a thermally conductive nozzle extending from the distal end of the thermometer, the nozzle being located and shaped to receive a probe cover thereon,a sensor can housing temperature... | Covidien AG | G01J005/06 G01K001/16 A61B005/01
US07603168 | B2 | 2009-10-13 | Method and apparatus for the estimation of anesthesia depth | 2008-04-17 | A method and apparatus to monitor and control the anesthesia state of a patient undergoing general anesthesia is provided. Previous automated systems to monitor the anesthesia state of a patient undergoing general anesthesia involve a significant time delay between the patient's true anesthesia state and the computed indices. The present invention reduces this time delay by using a novel analysis technique applied to spontaneous electrophysiological activity. A transformation and statistical analysis of the observed electrophysiological activity is conducted and compared to reference data to provide numerical indicators. In addition, these indicators are consistent with the levels of depression of patients CNS and ANS states. This method is illustrated in detail by CNS monitoring of electroencephalogram signals. | 1. A method for extracting information from an observed electrophysiological signal from a subject in order to evaluate the level of depression of the ANS (Autonomous Nervous System) of said subject with a processor, said method comprising: a) acquiring at least two reference signals, said at least two reference signals corresponding to two distinct ANS states obtained from a one or more reference subject or subjects;b) selecting a transformation function which, when applied to one of said at least two reference signals, yields a set of coefficients;c) selecting a statistical function which, when applied to said set of coefficients derived from one of said at least two reference signals, or a subset of said set of coefficients, yields a reference data set which characterizes the distinct ANS state corresponding to said one of at least two reference signals;d) applying said transformation and statistical function to said at least two reference signals to produce reference data sets which distinguish between the distinct ANS state(s) corresponding to said reference signals;e) observing electrophysiological activity of said subject to produce said observed electrophysiological signal;f) applying said transformation and statistical function to said observed electrophysiological signal to produce an observed data set;g) comparing the observed data set to one or more of said reference data sets; andh) computing with the processor a numerical value or values representative of said level of depression of the ANS of said subject which results from said comparison;wherein said at least two reference signals correspond to distinct ANS states which are two extreme states. 2. The method of claim 1 wherein said observed electrophysiological signal and reference signals representing measured cardiovascular activity of the subject. 3. The method of claim 2 wherein said observed electrophysiological signal and reference signals representing measured cardiovascular activity of a subject comprises one or more electrocardiograms. 4. The method of claim 1 wherein said reference subject or subjects are not the same individual as said subject. 5. The method of claim 1 wherein said statistical function is selected from the group histogram, probability density function, standard deviation, or variance. 6. The method of claim 1 used to measure electrophysiological activity in said subject to ascertain the level of analgesia of said subject. 7. The method of claim 1 used to measure electrophysiological activity in said subject to ascertain the effects of anesthesia and analgesia agents on the ANS of said subject. 8. The method of claim 1 used to measure electrophysiological activity in said subject to ascertain titration and dosage profiles of anesthesia and analgesia compounds end medicaments. 9. The method of claim 1 wherein said transformation function is a wavelet transform. 10. The method of claim 1 wherein said transformation function is any transform with j... | BioNova Technologies Inc. | A61B005/04
US08204597 | B2 | 2012-06-19 | Evaluating patient incontinence | 2007-05-30 | Systems and methods for determining a number of interruptions in a sleep state of a patient attributable to an urge or need to void during a sleep event include monitoring an activity level of the patient. The urge or need to void may be attributable to fecal or urinary incontinence. The patient activity level may be determined with one or more sensors that detect motion and/or one or more sensors that monitor a physiological parameter of the patient that varies as a function of patient activity. In one embodiment, a clinician selects a therapy parameter set for the patient based on the severity of the patient's incontinence. In another embodiment, the number of interruptions in the sleep state is associated with a therapy parameter set that was implemented during the sleep state in order to evaluate the efficacy of the therapy parameter set. | 1. A system comprising: an implantable activity sensor that generates a signal that varies as a function of activity of a patient; andat least one processor that is configured to receive the signal from the activity sensor, detect a sleep state of the patient, determine a first number of interruptions in the sleep state of the patient based on the signal, and determine a second number of sleep event voiding events based on the number of interruptions in the sleep state of the patient. 2. The system of claim 1, wherein the at least one processor is configured to adjust a therapy parameter to treat sleep event voiding based on the number of interruptions in the sleep state. 3. The system of claim 2, further comprising a memory, wherein the at least one processor is configured to compare the number of interruptions in the sleep state to a quality of life threshold stored in the memory and adjust the therapy parameter if the number of interruptions exceeds the quality of life threshold. 4. The system of claim 2, wherein the therapy parameter comprises at least one of voltage or current amplitude of electrical stimulation, pulse width of electrical stimulation, pulse frequency of electrical stimulation, a drug dosage, or a frequency of a delivery of a drug. 5. The system of claim 2, wherein the at least one processor is configured to adjust a therapy parameter by at least selecting a therapy parameter set to treat sleep event voiding based on the number of interruptions in the sleep state. 6. The system of claim 1, wherein the at least one processor is configured to select a therapy parameter set to treat sleep event voiding based on the number of interruptions in the sleep state. 7. The system of claim 1, wherein the at least one processor is configured to detect the sleep state of the patient by receiving an indication from the patient that the patient is attempting to sleep. 8. The system of claim 1, wherein the at least one processor is configured to detect the sleep state of the patient based on the signal from the activity sensor. 9. The system of claim 1, further comprising a memory, wherein the at least one processor is configured to determine an activity level of the patient based on the signal, determine the number of interruptions in the sleep state of the patient by comparing the activity level to an activity level threshold stored in the memory, and determine when the activity level exceeds the activity level threshold. 10. The system of claim 9, wherein the at least one processor is configured to generate an awake count when the activity level exceeds the activity level threshold and a subsequent activity level falls below the activity level threshold within a predetermined range of time, wherein each awake count represents an interruption in the sleep state. 11. The system of claim 9, wherein the implantable activity sensor senses a physiological parameter to generate the signal. 12. The system of claim 11, wherein the signal... | Medtronic, Inc. | A61B005/20
US08100123 | B2 | 2012-01-24 | Respiratory system for inducing therapeutic hypothermia | 2005-01-24 | The present invention provides a method and apparatus for controlling a patient's body temperature and in particular for inducing therapeutic hypothermia. Various embodiments of the system are described. The system includes: a source of breathing gas, which may be in the form of a compressed breathing gas mixture; a heat exchanger or other heating and/or cooling device; and a breathing interface, such as a breathing mask or tracheal tube. Optionally, the system may include additional features, such as a mechanical respirator, a nebulizer for introducing medication into the breathing gas, a body temperature probe and a feedback controller. The system can use air or a specialized breathing gas mixture, such as He/O2 or SF6/O2 to increase the heat transfer rate. In addition, the system may include an ice particle generator for introducing fine ice particles into the flow of breathing gas to further increase the heat transfer rate. | 1. Apparatus for inducing therapeutic hypothermia, comprising: a source of breathing gas;a heat exchanger configured to exchange heat with a flow of breathing gas from the source;an ice particle generator configured to introduce fine ice particles into the flow of breathing gas; anda breathing interface configured to deliver the flow of breathing gas containing the fine ice particles to a patient. 2. The apparatus of claim 1, wherein the source of breathing gas comprises a mechanical ventilator. 3. The apparatus of claim 1, wherein the source of breathing gas comprises a mechanical ventilator configured to deliver ambient air to the apparatus. 4. The apparatus of claim 1, wherein the source of breathing gas comprises a supply of compressed breathing gas. 5. The apparatus of claim 4, wherein the compressed breathing gas comprises a mixture containing oxygen and helium. 6. The apparatus of claim 4, wherein the compressed breathing gas comprises a mixture containing oxygen and sulfur hexafluoride. 7. The apparatus of claim 1, wherein the ice particle generator includes a source of water and an atomizer for introducing fine droplets of water into the flow of breathing gas. 8. The apparatus of claim 1, wherein the ice particle generator includes a source of saline solution and an atomizer for introducing fine droplets of saline solution into the flow of breathing gas. 9. The apparatus of claim 7, wherein the atomizer is an ultrasonic atomizer. 10. The apparatus of claim 7, wherein the ice particle generator includes a chiller for cooling the water prior to atomization. 11. The apparatus of claim 1, wherein the heat exchanger is configured to selectively cool or heat the flow of breathing gas from the source. 12. The apparatus of claim 1, further comprising a temperature sensor for measuring body temperature of the patient. 13. The apparatus of claim 11, further comprising a feedback controller for controlling operation of the apparatus based on the measured body temperature of the patient. | Thermocure, Inc. | A61M015/00 A61M016/00 A62B007/00 F24F005/00
US08224420 | B2 | 2012-07-17 | MRI guided ultrasound therapy apparatus | 2005-08-31 | This invention provides a MRI guided ultrasound therapy apparatus. It comprises a static field magnet adapted to apply a static magnetic field in an magnetic resonance volume at a predetermined disposition; at least one ultrasound energy applicator adapted to apply energy within an energy application zone at a predetermined disposition; and The mechanical positioning means for moving said ultrasound energy applicator to position the applicator so that the energy application zone intersects said magnetic resonance volume within said region of the subject. In that apparatus, the static field magnet is open at both ends or open at side. The sided open is upward or downward and the mechanical positioning means of this ultrasound energy applicator is close to and is located outside of this sided open. This invention reduces the space limitation for the mechanical positioning means of the ultrasound transducer. Meanwhile, the non-magnetic requirement on the mechanical positioning means become less greatly and particularly the problem of interference from magnetic field produced by working current of the transducer power cord to MRI system can be solved. | 1. An MRI guided high-intensity focused ultrasound therapy apparatus, comprising, a) a static field magnet adapted to apply a static magnetic field in a magnetic resonance volume at a predetermined disposition;b) at least one ultrasound energy applicator adapted to apply energy within an energy application zone at a predetermined disposition; andc) a mechanical positioning means for moving said ultrasound energy applicator to position the applicator so that the energy application zone intersects said magnetic resonance volume within a region of a patient to be treated,wherein said static field magnet is open at both ends and open at one side, and, during treatment, the patient to be treated is fixed and positioned in a magnetic gap of the static field magnet,the open at side is either upward or downward, and the mechanical positioning means of said ultrasound energy applicator is close to and outside of said open at side, andsaid ultrasound energy applicator is also located outside of the open at side. 2. The apparatus as claimed in claim 1, when said static field magnet has a downward open side, further comprises an open flexible fluid container, which makes said patient to be treated contact said fluid directly. 3. The apparatus as claimed in claim 2, further comprising a treatment bed for fixing said patient to be treated, said treatment bed has an open side, and said flexible fluid container couples to the patient through this open side, wherein said fluid is degassed water. 4. The apparatus as claimed in claim 3 wherein, when said static field magnet has an upward open side, said flexible fluid container is closed by a flexible acoustic membrane and the flexible acoustic membrane is used as a contacting surface with the patient. 5. The apparatus as claimed in claim 4, wherein the closed side of said magnet may be used as a treatment bed for fixing the patient. 6. The apparatus as claimed in claim 1, wherein an MRI image is used to locate the region of the patient to be treated, check ultrasound beams transmission path and make treatment procedures. 7. The apparatus as claimed in claim 1, wherein a real-time MRI fast image is used to monitor acoustic energy application to the region of the patient to be treated and/or check an ultrasound beam transmission path. 8. The apparatus as claimed in claim 1, further comprising a receiving and processing means to obtain local temperature information on the region of the patient to be treated and/or an ultrasound beams transmission path within the magnetic resonance volume from an MRI system. 9. The apparatus as claimed in claim 8, wherein the temperature images are used for monitoring treatment. | Chongqing Ronghai Medical Ultrasound Industry Ltd. | A61B005/055
US08204599 | B2 | 2012-06-19 | System for anchoring an implantable sensor in a vessel | 2008-04-16 | A system and a method of disposing a second sensor module overlying a first sensor module system is described. A first assembly including an expandable anchor and a sensor module is at least partially overlapped by a second assembly including an expandable anchor and a sensor module. If necessary or desired, the functions of the second sensor module can replace the functions of the first sensor module. The sensor module may include a blood pressure sensor. | 1. A method of overlapping a first assembly including a first expandable anchor coupled to a first sensor module with a second assembly including a second expandable anchor coupled to a second sensor module within an artery or blood vessel, each of the first and second expandable anchors being configured for transition between a collapsed configuration and an expanded configuration, the method comprising: locating a position of the first assembly at a target location within the artery or blood vessel;delivering the second assembly with the second expandable anchor in its collapsed configuration to a position proximate the target location within the artery or blood vessel;positioning the second assembly within the first assembly such that the first expandable anchor at least partially overlaps the second expandable anchor; andexpanding the second expandable anchor to its expanded configuration. 2. The method of claim 1, further comprising aligning a first alignment feature located on the first anchor with a second alignment feature located on the second anchor. 3. The method of claim 1, further comprising adjusting the position of the second sensor module assembly. 4. The method of claim 1, further comprising deactivating the first sensor module. 5. The method of claim 1, further comprising activating the second sensor module. 6. A method of overlapping a first assembly including an expandable first anchor coupled to a first sensor module with a second assembly including a second expandable anchor coupled to a second sensor module within an artery or blood vessel, each of the first and second expandable anchors being configured for transition between a collapsed configuration and an expanded configuration, the method comprising: locating and evaluating functionality of the first sensor module, the expandable first anchor having been previously expanded to its expanded configuration within the artery or blood vessel;delivering the second assembly with the second expandable anchor in its collapsed configuration to a location proximate the first assembly within the artery or blood vessel;positioning the second assembly within the first assembly within the artery or blood vessel such that the first expandable anchor at least partially overlaps the second expandable anchor; andexpanding the second expandable anchor to its expanded configuration. 7. The method of claim 6, further comprising aligning a first alignment feature located on the first anchor with a second alignment feature located on the second anchor. 8. The method of claim 6, further comprising adjusting the position of the second sensor module assembly. 9. The method of claim 6, further comprising deactivating the first sensor module. 10. The method of claim 6, further comprising activating the second sensor module. | Cardiac Pacemakers, Inc. | A61N001/00
US08271063 | B2 | 2012-09-18 | System and method for a non-invasive medical sensor | 2007-06-15 | A sensor is disclosed having one or more light sources, and one or more light detectors. At least one of the detectors produces an output in response to received light. Also, electrical circuitry in response to the output generates a pulse waveform proportional to the arterial and venous pulse of a human body, wherein the pulse waveform is used to synchronize an arterial-pulse measurement of the absorption or reflectance of one or more coherent light sources. | 1. An apparatus comprising: a sensor having one or more coherent light sources, one broadband light source and one or more wavelength detectors wherein the sources and detectors are mounted opposite each other, and at least one of the detectors produces an output value corresponding to a ratio and indicating the centroid wavelength of received light. 2. The apparatus of claim 1, wherein the apparatus is capable of being placed on a finger or other extremity of the human body and is capable of shining light from the one or more light sources through the extremity. 3. The apparatus of claim 1, wherein the sensor is capable of reflecting light off skin from any part of the human body. 4. The apparatus of claim 1, wherein at least one of the one or more wavelength detectors is a dual photodiode or other type that produces the output indicating a centroid wavelength of received light. 5. The apparatus of claim 1, wherein at least one of the one or more light sources emits light of a wide light spectrum from 200 nm to 1100 nm or any portion thereof. 6. The apparatus of claim 5, wherein at least one of the one or more light sources emits light of a wide light spectrum is a white LED or any other broad emission device. 7. The apparatus of claim 1, wherein the sensor is connected to a processing system that drives the light sources, including a broad emission device and also receives signals from photodiodes and the wavelength detector. 8. The apparatus of claim 7, wherein the processing system includes computer media, memory, and signal processing circuitry to calculate various parameters of a blood stream through comparison of the signals received by the photodiodes when the output value corresponding to an arterial pulse is detected. 9. The apparatus of claim 8, wherein the processing system generates light from the broad emission light source which is capable of illuminating tissue. 10. The apparatus of claim 9, wherein the wavelength detector generates a signal proportional to a strongest transmission wavelength of the light emitted from the broad emission light source, after it either passed through a body part or is reflected off of skin of a human. 11. The apparatus of claim 10, wherein the signal derived from the wavelength detector is processed by circuitry and computer processing to obtain a pulse waveform proportional to the arterial and venous pulse of the human body. 12. The apparatus of claim 11, wherein the time domain and amplitude properties of both the arterial and venous pulse are stored and further used to synchronize an arterial-pulse synchronized measurement of the absorption or reflectance of the one or more coherent light-sources through a portion of the human body. 13. The apparatus of claim 12, wherein the arterial and venous pulse are used to calculate oxygen-saturation, carbon-monoxide concentration or any other parameter of the blood. 14. The apparatus of claim 13, wherein the arterial and venous pulse coinciding... | Medtor LLC | A61B005/1455
US08945009 | B2 | 2015-02-03 | Remote health monitoring system | 2008-06-12 | The invention relates generally to remote health monitoring systems, as applied to the field of public health surveillance. In particular, it relates to a multi-user remote health monitoring system that is capable of reliably identifying and collecting data from frontline healthcare providers, laboratory and hospital information systems, patients and healthy individuals in a number of ways, with a view to aid in the field of public health. The system can also be used to query and collect additional information regarding specifics pertaining to the health of the individuals, and for patient tracking, monitoring, and the collection of individual data. | 1. A computer implemented method for monitoring a plurality of individuals located across a regional, state, or national geographic area using a server executing instructions stored on a computer readable medium, the instructions comprising: (A) receiving, with the server, health-related data corresponding to one or more health conditions of said plurality of individuals from a respective device of a plurality of devices associated with said plurality of individuals and configured to produce said health-related data by (i) measuring one or more physiological parameters related to said one or more health conditions and (ii) receiving responses to one or more queries presented to said plurality of individuals, wherein said plurality of devices are (i) remotely located from said server, (ii) in communication with said server via a communications network, (iii) configured to generate data corresponding to a geographical location of each respective device of said plurality of devices, and (iv) configured to transmit said health-related data and said data corresponding to a geographical location of each respective device of said plurality of devices to said server;(B) highlighting (i) a population and (ii) a geographic region at risk of exposure to a disease by analyzing, with the server, said health-related data and said data corresponding to said geographical location of each respective device of said plurality of devices;(C) generating, with the server, one or more reports based on said health-related data and said data corresponding to said geographical location of each respective device of said plurality of devices; and(D) scaling, with the server, a level of surveillance among said plurality of individuals using the plurality of devices in response to a change in a threat level of said disease. 2. The method according to claim 1, further comprising the step of: generating, with the server, a visualization of said health-related data in response to said data corresponding to said geographical location of each respective device of said plurality of devices received from said devices. 3. The method according to claim 2, further comprising the steps of: generating, with the server, one or more surveys for one or more symptoms in response to said health-related data suggesting one or more outbreaks of said disease; andtransmitting, with the server, said surveys only to said devices in proximity to one or more of said locations associated with said outbreaks based on said data corresponding to said geographical location of each respective device of said plurality of devices. 4. The method according to claim 1, further comprising the steps of: storing one or more programs in said server; andtransmitting said programs to one or more of said plurality of devices to convey one or more messages to one or more of said plurality of individuals. 5. The method according to claim 1, wherein said health-related data and said data corresponding to said geograp... | Robert Bosch Heathcare Systems, Inc. | A61B005/00 G06F019/00
US08852127 | B2 | 2014-10-07 | System and method for monitoring information related to sleep | 2008-06-02 | A system configured to monitor ambient illumination experienced by a subject. In one embodiment, the system comprises an illumination sensor, a timer, and a storage module. The illumination sensor is configured to monitor an intensity of ambient illumination within two or more wavelength ranges by generating one or more output signals that convey information related to the intensity of ambient illumination within the two or more wavelength ranges. The timer is configured to indicate the passage of periods of time. The storage module is configured to store information related to the intensity of ambient illumination within the two or more wavelength ranges, as conveyed by the one or more output signals, for individual periods of time. The system is portable to be carried by the subject. | 1. A system configured to monitor ambient illumination experienced by a subject, the system comprising a wearable piece, the wearable piece comprising: an illumination sensor configured to monitor an intensity of ambient illumination within two or more wavelength ranges by generating one or more output signals that convey information related to the intensity of ambient illumination within the two or more wavelength ranges;a timer configured to indicate the passage of periods of time;a motion sensor configured to monitor movement of the subject by generating one or more output signals that convey information related to movement of the subject;a storage module configured to: store information related to the intensity of ambient illumination within the two or more wavelength ranges, as conveyed by the one or more output signals, in separate sets of data that correspond to individual periods of time; andstore information related to movement of the subject, as conveyed by the one or more output signals generated by the motion sensor, during the individual periods of time; anda communication interface configured to facilitate communication of the sets of data and the information related to movement of the subject to an external processor;wherein, at least the illumination sensor, the timer, the motion sensor, the storage module, and/or the communication interface is portable to be carried by the subject. 2. The system of claim 1, further comprising a strap that engages an extremity of the subject, wherein the illumination sensor, the timer, and the storage module are carried on the strap. 3. The system of claim 1, further comprising a proximity sensor configured to monitor whether or not the subject is wearing the system by generating an output signal that indicates whether or not the subject is proximate to the proximity sensor. 4. The system of claim 1, wherein the two or more wavelength ranges comprise three wavelength ranges. 5. The system of claim 4, wherein the three wavelength ranges include a first wavelength range corresponding to red, a second the wavelength range corresponding to green, and a third wavelength range corresponding to blue. 6. The system of claim 1, wherein the information stored by the storage module enables a determination of photon flux for individual ones of the two or more wavelength ranges during individual periods. 7. The system of claim 1, wherein the illumination sensor comprises one or both of (i) an array of photodiodes, the photodiodes included in the array comprise, for a given one of the two or more wavelength ranges, one or more photodiodes capable of monitoring the intensity of ambient illumination in the given one of the two or more wavelength ranges, or (ii) a spectrophotometer. 8. The system of claim 1, wherein the external processor comprises an illumination score module configured to determine, from the intensities of ambient illumination within the two or more wavelength ranges, a metric indicating ... | Ric Investments, LLC | A61B005/103 A61B005/117 A61B005/04 G08B023/00 G01J001/42 G01J003/51 A61B005/00 A61B005/11 G01J001/02 G01J001/04
US08352005 | B2 | 2013-01-08 | Noninvasive blood analysis by optical probing of the veins under the tongue | 2008-04-24 | A method for analysis of blood components or parameters is disclosed where a probe having an excitation outlet and a response inlet is placed in proximity to or in contact with a tissue of an underside of a patient's tongue over a big vein in the tongue so that an excitation signal exits the outlet, produces a response which enters the inlet for detection and analysis. | 1. A method for noninvasive analysis of at least one blood component in venous blood comprising the steps of: placing a probe in contact with a surface of a tissue directly under a big vein associated with an underside of a patient's tongue, where the probe includes a tip having a radiation outlet and a response inlet so that both the radiation outlet and the response inlet are simultaneously in contact with the surface and where the outlet is connected to a light source via a light source conduit and the inlet is connected to a detector via a response conduit;irradiating venous blood in the big vein with radiation having at least one frequency or wavelength;detecting a response from the venous blood irradiated in the irradiating step; andcalculating a concentration of a blood component, a value of a blood parameter or a mixture or combination thereof in the venous blood from the response. 2. The method of claim 1, further comprising the step of: displaying the response, the concentration and/or the value from the calculating step. 3. The method of claim 1, wherein the detecting step comprises the step of: utilizing one or a combination of techniques selected from the group consisting of reflectance technique, confocal technique, scanning confocal technique, polarization techniques, interferometry, optoacoustics, low coherence interferometry and reflectometry, techniques based on speckle measurements, fluorescence, Raman scattering, and a two or multi-photon technique. 4. The method of claim 1, wherein the wavelength of the radiation is from about 200 nanometers to about 20 microns. 5. The method of claim 1, wherein the response corresponds to a concentration of hemoglobin in the blood and the wavelength of the radiation is selected from the group consisting of 548 nm, 568 nm, 587 nm, and 805 nm, a wavelength in a range from about 400 nm to about 640 nm, and a wavelength in a range from about 520 nm to about to about 1700. 6. The method of claim 1, wherein the radiation has a plurality of wavelengths or frequencies. 7. The method of claim 1, wherein the response corresponds to a concentration of hemoglobin in the blood and the wavelength of the radiation is in a range from 250 to 2500 nm. 8. The method of claim 1, wherein the blood component and/or value of a parameter of the blood is selected from the group consisting of hematocrit, hemoglobin, glycosylated hemoglobin, glucose, cholesterol, oxy-hemoglobin, deoxy-hemoglobin, carboxy-hemoglobin, an exogenous substance, and combinations thereof. 9. The method of claim 8, wherein the exogenous substance is selected from the group consisting of one or more of: a drug; a dye or other reporter in a molecular state; particles made of liquid, gas or solid materials and combinations thereof; and a layered structure. 10. The method of claim 9, wherein the exogenous substance is a dye selected from the group consisting of indocyanine green and Evans blue. 11. The method of claim 9, wherein the exo... | Board of Regents, The University of Texas System | A61B005/1455
US08107698 | B2 | 2012-01-31 | Image processing apparatus and computer program product | 2008-04-17 | An image processing apparatus includes an organ determination unit that determines a type of observation target which appears in a target image among a sequence of observation images, an imaging distance estimation unit that estimates an imaging distance at a time of image pickup of the observation image, an abnormal region detection unit that detects an abnormal region, which is a specific region, from the target image using an abnormality detection parameter, and an image processing control unit that sets as the abnormality detection parameter, a parameter value corresponding to a result of determination by the organ determination unit, and causes the abnormal region detection unit to detect an abnormal region using the abnormality detection parameter. | 1. An image processing apparatus for detecting a feature image region, which is a specific region, from a sequence of observation images in which an image of at least one type of plural types of observation targets is picked up individually, comprising: a target determination unit that determines a type of an observation target whose image is picked up in a target image among the sequence of observation images;a region detection unit that detects the feature image region from the target image using a process parameter for detecting the feature image region;a distance estimation unit that estimates an imaging distance at a time of image pickup of the target image; anda setting control unit that sets a parameter value corresponding to a result of determination by the target determination unit and a result of estimation by the distance estimation unit, as the process parameter, and causes the region detection unit to detect the feature image region using the process parameter. 2. The image processing apparatus according to claim 1, wherein the target determination unit determines the type of the observation target whose image is picked up in the target image based on at least one type of information among compression information, frequency component information, an amount of image information, and texture information of the target image. 3. The image processing apparatus according to claim 1, wherein the parameter value represents one of a value defining a feature of the feature image region, and a processing unit pixel block size of the target image processed by the region detection unit. 4. The image processing apparatus according to claim 3, wherein the value defining a feature is a region size or color. 5. The image processing apparatus according to claim 1, wherein the distance estimation unit estimates the imaging distance at the time of image pickup of the target image based at least on brightness information or gain information of the target image. 6. The image processing apparatus according to claim 1, further comprising a parameter storage unit that stores therein the parameter value corresponding to the type of observation target based on the imaging distance,the setting control unit sets a parameter value corresponding to the type of observation target determined by the target determination unit and the imaging distance estimated by the distance estimation unit among the parameter values stored in the parameter storage unit, as the process parameter. 7. The image processing apparatus according to claim 1, wherein the target determination unit sequentially switches the target image from one image to another among the sequence of observation images, and determines the observation target for each target image,the region detection unit sequentially switches the target image from one image to another among the sequence of observation images and detects the feature image region of each target image. 8. The image processing apparatus acc... | Olympus Corporation | G06K009/00 H04N007/18
US08784331 | B2 | 2014-07-22 | Method and device for determining the hydrodynamics of the cerebrospinal fluid system | 2006-02-27 | The present invention relates to a method and a device for determining the hydro-dynamic properties of the fluid system surrounding the brain and the spinal cord, the so called cerebrospinal fluid system, whereby the method comprises continuous pressure measurement through a fluid contact passage and active infusion of artificial cerebrospinal fluid through another fluid contact passage to a number of pressure-flow levels, and analysis of the connections between the measured pressures and flows. The method uses an adaptive way of procedure which on each pressure-flow level takes account for the measure time and the patient's fluctuations in physiologic signals for calculating in real time when the relation between the measure time and measure accuracy in pressure and flow on the actual pressure-flow level is sufficient and at an optimum, whereupon the investigation or examination is automatically proceeding to the next level according to a predetermined protocol. In this way, the method gives rise to an investigation with measurements having sufficient pressure and flow information on each level, which forms a base for determination in a correct manner, with an uncertainty estimate, of the patient's hydrodynamic parameters. The device comprises a hose pump (1) for infusion of artificial cerebrospinal fluid in a bottle or bag (2), a standardized hose set including a pump hose (3) and pressure transducers (4) for continuous registration of the intracranial pressure, an invasive contact object (5) for creating fluid contact with the cerebrospinal fluid system and consisting of needles or a catheter, as well as a computer (12) with software for computerized collection and analysis as well as control of pump speed. The device uses a computerized implementation of the above-mentioned adaptive method in order to systematically and safely carry through a standardized protocol which generates pressure and flow information which the software of the measure system then uses for determining in real time and with an uncertainty estimate, the hydrodynamic parameters of the system. | 1. A device for automatically determining hydrodynamic characteristics of a patient cerebrospinal fluid system, comprising: (a) a hose pump for infusion of artificial cerebrospinal fluid from a container into or out of the patient cerebrospinal fluid system for continuous pressure measurement through a first fluid contact passage and active infusion of artificial cerebrospinal fluid through a second fluid contact passage; (b) a pump hose through which the artificial cerebrospinal fluid is fed from the container;(c) pressure transducers for continuous measurement of intracranial pressure;(d) an invasive contact object, connected to the pump hose, with two lumens for creating fluid contact with the cerebrospinal fluid system; and(e) a computer for control of the hose pump, the computer having software for computerized collection and analysis, and programmed according to a predetermined protocol that generates or provides pressure and flow information and uses the information for determining, with an uncertainty estimate, the hydrodynamic characteristics of the patient, wherein the computer is programmed to initially measure each level of pressure and flow for predetermined time intervals, and to automatically proceed to a next level of pressure and flow based upon a determination by the computer that a sufficient accuracy or precision of the hydrodynamic properties has been obtained. 2. A device for automatically determining hydrodynamic characteristics of a patient cerebrospinal fluid system, comprising: (a) a hose pump for infusion of artificial cerebrospinal fluid from a container into or out of the patient cerebrospinal fluid system for continuous pressure measurement through a first fluid contact passage and active infusion of artificial cerebrospinal fluid through a second fluid contact passage; (b) a pump hose through which the artificial cerebrospinal fluid is fed from the container;(c) pressure transducers for continuous measurement of intracranial pressure;(d) an invasive contact object, connected to the pump hose, with two lumens for creating fluid contact with the cerebrospinal fluid system; and(e) a computer for control of the hose pump, the computer having software for computerized collection and analysis, and programmed according to a predetermined protocol that generates or provides pressure and flow information and uses the information for determining, with an uncertainty estimate, the hydrodynamic characteristics of the patient, wherein the computer is programmed to:(i) make a determination that a patient has large fluctuations in physiological signal if the computer determines physiological signal fluctuations meet predetermined conditions, and based upon that determination, measure pressure and flow for a first time period of a first length; and(ii) make a determination that a patient has a small fluctuation if the computer determines physiological signal fluctuations do not meet the predetermined conditions, and based upon tha... | Likvor AB | A61B005/03 A61M005/00
US08442606 | B2 | 2013-05-14 | Optical sensor apparatus and method of using same | 2008-05-12 | An apparatus for, and method of, sensing characteristics of a vessel and a fluid conveyed therein. | 1. A sensing device for acquiring signals and computing measurements comprising: a sensor assembly including a plurality of emitters and a plurality of detectors for generating a plurality of signals;a computing device operating the plurality of emitters and detectors and processing the plurality of signals to obtain measurement values;an energy storage device configured for powering the computing device and the sensor assembly; anda housing enclosing the sensor assembly, the energy storage device, and the computing device,the sensing device configured for implantation in a patient wherein the emitters and the detectors are arranged in the housing to face one side of a vessel when implanted, and said processing includes computing, after implantation, a distance from the sensor assembly to the vessel, a diameter of the vessel, and an oxygen saturation value of blood flowing in the vessel. 2. The sensing device of claim 1, wherein the plurality of emitters are arranged in a matrix. 3. The sensing device of claim 2, wherein the plurality of detectors are arranged in a matrix. 4. The sensing device of claim 3, wherein the emitter matrix includes four rows of four emitters and the detector matrix includes four rows of four detectors. 5. The sensing device of claim 1, wherein each detector is operationally paired with a different emitter. 6. The sensing device of claim 1, wherein the number of emitters in the plurality of emitters is different from the number of detectors in the plurality of detectors. 7. The sensing device of claim 1, wherein the sensing device is dimensioned about the same as two stacked quarter dollar coins. 8. The sensing device of claim 1, wherein the sensor assembly and the computing device are integrated in a single piece. 9. The sensing device of claim 1, further including one or more communication devices for transmitting and receiving communication signals. 10. The sensing device of claim 9, wherein the communication device transmits and receives wireless communication signals. 11. The sensing device of claim 9, wherein the communication device includes a connector adapted to operably couple to one or more of a docking station and a second sensing device. 12. The sensing device of claim 9, wherein the sensor assembly, the computing device, and the one or more communication devices are integrated in a single piece. 13. The sensing device of claim 1, wherein the energy storage device includes an energy coupler for receiving energy to recharge the energy storage device. 14. The sensing device of claim 1, said processing further including computing the diameter of the vessel by activating the plurality of emitters to avoid beam overlays, and computing the oxygen saturation value by concurrently activating at least selected emitters from the plurality of emitters. 15. The sensing device of claim 14, the selected emitters corresponding to emitters selected, based on the diameter of the vessel, to emit beams that impin... | Cardio Art Technologies Ltd. | A61B005/1455
US08355552 | B2 | 2013-01-15 | Automated determination of lymph nodes in scanned images | 2008-06-19 | Techniques include automatically detecting a lymph node in a scanned image of a body without human intervention, using one or more of three approaches. First, a subset of scanned images is determined, which belongs to one anatomical domain. A search region for lymph tissue is in a particular spatial relationship outside an anatomical object in the domain. Second, scanned images are segmented without human intervention to determine a boundary of a particular lymph node. The scanned images and outline data are received. Some of these embodiments automatically segment by determining an external marker, based on the outline data, and an internal marker, based on a geometric center of the outline data or thresholds determined automatically inside detected edges, or both, for a marker-controlled watershed algorithm. Third, based on lymph node data at a particular time, a second scanned image at a different time is segmented automatically, without human intervention. | 1. A method comprising automatically detecting a lymph node in a computer tomography X-ray image of internal tissues of a body without human intervention based on applying a dot-enhancement Hessian matrix to voxels in the computer tomography X-ray image. 2. The method as recited in claim 1, further comprising: receiving a single set of a plurality of computer tomography X-ray images representing corresponding different slices of the body between a neck of the body and a bottom of a pelvis of the body; anddetermining a subset of the plurality of the computer tomography X-ray images wherein the subset of the computer tomography X-ray images belongs to one anatomical domain of a plurality of anatomical domains between the neck and the bottom of the pelvis. 3. The method as recited in claim 2, wherein: the method further comprises determining voxels in the subset associated with an anatomical object that is not a lymph node and which is associated with the anatomical domain; andautomatically detecting the lymph node comprises determining voxels in a search region of a particular image of the subset in a particular spatial relationship outside the voxels associated with the anatomical object. 4. The method as recited in claim 3, wherein: applying the dot-enhancement Hessian matrix further comprises applying dot-enhancement Hessian matrixes at a plurality of scales to the voxels in the search region to produce a corresponding plurality of transformed images; andautomatically detecting the lymph node further comprises: determining a set of one or more candidate nodes on the transformed images, wherein each candidate node includes voxels with values above a lymph threshold value which form a contiguous region of more than a particular number of voxels; andselecting one candidate node of the set as the lymph node. 5. The method as recited in claim 4, wherein selecting one candidate node of the set further comprises: determining whether a candidate node of the set is a false positive based on known statistics of lymph nodes; andif the candidate node is a false positive, then removing the candidate node from the set before selecting the one candidate node of the set as the lymph node. 6. A method for segmenting a lymph node comprising the steps of: receiving a single set of one or more computer tomography X-ray images representing internal tissues of a body at a particular time;receiving outline data that indicates a region of interest that encloses voxels that represent a lymph node, wherein the region of interest is in the single set of one or more computer tomography X-ray images; andautomatically, without human intervention, segmenting the single set of one or more computer tomography X-ray images to identify a lymph node boundary between voxels that represent the lymph node and voxels that do not, based at least in part on applying a marker-controlled watershed algorithm. 7. The method as recited in claim 6, wherein automatically segmenting the s... | The Trustees of Columbia University in the City of New York | G06K009/00
US09943234 | B2 | 2018-04-17 | Using respiration distress manifestations for heart failure detection | 2007-04-17 | Systems and methods for diagnosing one or more respiration distress manifestations by implantably recognizing their occurrence and evaluating information about the same to provide an indication of present or impending worsening heart failure are discussed. Using information produced by an implantable respiration sensor circuit and an implantable physiological sensor circuit, such as at least one of a physical activity sensor circuit or a posture sensor circuit, an implantable or external processor circuit may detect a respiration disturbance and an associated subsequent arousal from stable state occurrence and thereafter evaluate over time arousal from stable state occurrences to provide the indication of present or impending worsening heart failure. In one example, information about a fluid level within a subject is used in determining the indication of worsening heart failure. In various examples, a regimen is initiated or adjusted in response to the indication of present or impending worsening heart failure. | 1. A system comprising: an implantable medical device including, a respiration sensor circuit configured for producing a respiration signal indicative of a respiration or a respiration related parameter of a subject; anda physiological sensor circuit that includes at least one of a physical activity sensor circuit configured for producing a physical activity signal indicative of a physical activity level of the subject or a posture sensor circuit configured for producing a posture signal indicative of a posture of the subject; andan implantable or external processor circuit including an input to receive information about the respiration signal and at least one of the physical activity signal or the posture signal, the processor circuit configured for detecting a paroxysmal nocturnal dyspnea (PND) event by detecting a respiration disturbance using the respiration signal and detecting an associated subsequent arousal from stable state occurrence caused by the respiration disturbance, from at least one of the physical activity signal or the posture signal, the processor circuit further configured for evaluating over time, at least in part, the respiration disturbances and the associated subsequent arousal from stable state occurrences caused by the respiration disturbance by comparing a trend or count of the respiration disturbances and the associated arousal from stable state occurrences to a specified threshold, providing an indication of present or impending worsening heart failure. 2. The system of claim 1, comprising a clock circuit configured for producing at least one timestamp associated with at least one of the respiration signal, the physical activity signal, or the posture signal. 3. The system of claim 2, comprising a memory configured for storing information about the timestamp and at least one of the respiration signal, the physical activity signal, or the posture signal with which the timestamp is associated. 4. The system of claim 1, comprising a regimen control circuit configured for adjusting a regimen provided to the subject using, at least in part, information about at least one of the detected respiration disturbance, the detected arousal from stable state occurrence, or the indication of present or impending worsening heart failure. 5. The system of claim 1, comprising an external user-interface device communicatively coupled to the implantable medical device and including a user-detectable indication of an evaluation over time of at least one of the respiration disturbance, the arousal from stable state occurrence, or the indication of present or impending worsening heart failure. 6. The system of claim 5, wherein the external user-interface device includes a user input device configured for receiving programming information from a user and communicating the programming information to the implantable medical device. 7. The system of claim 1, comprising a fluid detector circuit configured for producing an indication of a ... | Cardiac Pacemakers, Inc. | A61B005/08 A61B005/0205 A61B005/11 A61B005/00 A61N001/362
US10178965 | B2 | 2019-01-15 | Activity monitoring system for pregnant women | 2007-06-22 | An activity feedback system for an individual is disclosed. The activity feedback system can be wearable by the individual and provide activity information and possibly other information to the individual. The individual can, for example, be a pregnant woman. The activity information can, for example, be used by a pregnant woman to monitor the amount of activity by the woman during as well as after her pregnancy. Besides activity information, the activity feedback system can also monitor, record or provide other information that is beneficial for individuals. | 1. A wearable electronic device, comprising: a housing;an activity monitor configured to monitor activity of a user of said wearable electronic device;a memory device configured to store information related to the activity of the user and a due date associated with the user;a processor operatively connected to said activity monitor and said memory device, said processor being configured to compare information related to the activity of the user with at least one threshold level, wherein the processor is configured to determine the at least one threshold level being dependent on (i) the due date associated with the user, (ii) a user characteristic pertaining to the user, and (iii) calibration information associated with the user, the calibration information being determined based on data acquired from monitoring a predetermined activity of the user while wearing said wearable electronic device; andan output device operatively connected to said processor, said output device being configured to output at least activity output information for the user of said wearable electronic device, the activity output information being based at least in part on the comparing of the information related to the activity of the user with the at least one threshold level,wherein at least said processor and said memory device are included within said housing, andwherein said housing, said activity monitor and said output device of said wearable electronic device are wearable. 2. A wearable electronic device as recited in claim 1, wherein said memory device stores a plurality of threshold levels. 3. A wearable electronic device as recited in claim 2, wherein at least one of the plurality of threshold levels is for use prior to the due date, and wherein at least one of the plurality of threshold levels is for use after the due date. 4. A wearable electronic device as recited in claim 2, wherein the predetermined activity of the user is walking. 5. A wearable electronic device as recited in claim 2, wherein said processor is further configured to determine the one of the threshold levels to be used for comparison with the activity of the user based on a doctor's input associated with the user. 6. A wearable electronic device as recited in claim 2, wherein said memory device is further configured to store at least one weight of the user. 7. A wearable electronic device as recited in claim 1, wherein said processor is configured to determine a number of days until the due date, and causes said output device to output the number of days until the due date. 8. A wearable electronic device as recited in claim 1, wherein said output device is a display. 9. A wearable electronic device as recited in claim 1, wherein said wearable electronic device further comprises: a battery for providing power to said wearable electronic device. 10. A wearable electronic device as recited in claim 1, wherein said memory device is further configured to store weight progression of the ... | IpVenture, Inc. | A61B005/00 A61B005/11
US10198958 | B2 | 2019-02-05 | Method and apparatus for training a team by employing brainwave monitoring and synchronized attention levels of team trainees | 2008-04-30 | Training methods and apparatus wherein a training environment is activated only when a trainee is in a focused attention state. A brainwave monitor is employed for determining level of attention. The training environment is activated when the level of attention of the trainee is at or above a predetermined attention threshold. Activation of the training environment provides feedback to the trainee that he or she is in a focused attention state, and at the same time provides the incentive to remain in the focused attention state. A focused attention state is important if not required for training/learning success. The method and apparatus may be employed for training individual trainees, as well as for training team member trainees. | 1. An apparatus for training a team including at least two team member trainees, comprising: a training environment for teaching a skill beyond merely maintaining a particular brainwave state, said training environment having not-activated and activated states, the trainee not being allowed to proceed in said training environment when said training environment is in the not-activated state, and the trainee being allowed to actively participate in said training environment when said training environment is in the activated state;a brainwave monitor for each trainee for monitoring electrical activity within the brain of each trainee, and determining each trainee's level of attention; andan activation device connected to said brainwave monitors and to said training environment, and operable to activate said training environment to the activated state when the levels of attention of all team member trainees are simultaneously at or above a predetermined attention threshold. 2. An apparatus for training a team including at least two team member trainees, comprising: a training environment for teaching a skill beyond merely maintaining a particular brainwave state, said training environment having not-activated and activated states, the trainee not being allowed to proceed in said training environment when said training environment is in the not-activated state, and the trainee being allowed to actively participate in said training environment when said training environment is in the activated state;a brainwave monitor for each trainee for monitoring electrical activity within the brain of each trainee, and determining each trainee's level of attention; andan activation device connected to said brainwave monitors and to said training environment, and operable to activate said training environment to the activated state when the level of attention of each team member trainee is simultaneously at or above a predetermined attention threshold for the particular team member trainee, which is the attention baseline for the particular team member trainee. 3. A training method for training a team including at least two team member trainees, comprising: providing an apparatus for training a team including at least two team member trainees according to claim 1;employing brainwave monitors of the apparatus for determining the level of attention for each team member trainee;providing a training environment from the apparatus for teaching a skill beyond merely maintaining a particular brainwave state, said training environment having not-activated and activated states, the trainee not being allowed to proceed in said training environment when said training environment is in the not-activated state, and the trainee being allowed to actively participate in said training environment when said training environment is in the activated state; andmonitoring the levels of attention of the team member trainees and activating the training environment of the apparatus to th... | Freer Logic | A61B005/0476 A61B005/0482 G09B019/00 G09B023/28 G06F015/18 G06F017/00 G09B025/00 G09B007/02 A61B005/11 A61M005/172 A61N001/36 A61B005/16 A61B005/00 A61B005/04 G06F019/00 A61N001/05 A61B005/18
US09532848 | B2 | 2017-01-03 | Computer-assisted surgery system and method | 2008-06-16 | There is described a system and a method for assisting a user manipulating an object during a surgery, the method comprising: tracking the object in a sterile field in which surgery is being performed at a location using a tracking device which generates tracking data; processing the tracking data using a processing device located outside the sterile field to generate position and orientation information related to the object; and sending the position and orientation information related to the object to a displaying device positioned in the sterile field adjacent to the location at which the surgery is being performed, for display to the user. | 1. A system for assisting a user manipulating an object during a surgery, the system comprising: a tracking device for tracking the object in a sterile field in which surgery is being performed, and for generating tracking data for the object;a processing device positioned outside the sterile field, for generating tracking information related to the object by processing the tracking data received from the tracking device; anda sterilized displaying device located in a volume within the sterile field defined as being above a plane of and delimited by an operating table and below the shoulders of an operator standing next to a patient lying on the operating table, the displaying device for displaying the tracking information received wirelessly from the processing device to the user during the surgery, the displaying device having a touch-screen for interaction with the user through said processing device, the displaying device being supported directly by the operating table via a support extending between the displaying device and the operating table. 2. The system as in claim 1, wherein the displaying device is positioned within a same field of view as a location at which the surgery is being performed from a surgeon's viewpoint, for display of the information to the surgeon during manipulation of the object. 3. The system as in claim 1, further comprising an imaging device for acquiring images of at least one of the object and a location at which the surgery is being performed, the displaying device displaying the images. 4. The system as in claim 3, wherein the processing device comprises a coordinate translating device for translating the object position and orientation information into a coordinate system relative to the images. 5. The system as in claim 1, further comprising a server operatively connected to the processing device, for allowing a communication between the processing device and the displaying device. 6. The system as in claim 5, wherein the server comprises a wireless server for allowing wireless communication between the displaying device and the processing device. 7. The system as in claim 6, wherein the displaying device comprises a portable handheld computing device in the sterile field in close proximity to the location of the surgery. 8. The system as in claim 1, wherein the user comprises multiple users, and the displaying device comprises multiple displaying devices each comprising a dedicated user interface, wherein each of the displaying devices are dedicated to one of the users, and each of the users executes a specific role. 9. The system as in claim 8, further comprising a communications interface for assigning a communication priority between one of the displaying devices and the processing device, the communication priority being in accordance with the dedicated user interface involved in the communication. 10. The system as in claim 1, further comprising a database operatively connected to the processi... | OTHOSOFT, INC. | A61B006/00 A61B017/00
US20080243021 | A1 | 2008-10-02 | Signal Common Mode Cancellation For Handheld Low Voltage Testing Device | 2007-03-30 | An apparatus for monitoring bioelectric signals of a patient which includes a processing system and an interface for receiving external electrical signals representative of a condition of the patient. The interface is configured to convey a representation of the received external signals to the processing system, and includes a common mode cancellation amplifier circuit which is adapted to reduce common mode signal noise present in the external signals. | 1. Apparatus for monitoring bioelectric signals of a patient, comprising: a processing system;an interface coupled to said processing system for receiving external electrical signals representative of a condition of the patient and for conveying a representation of said received external signals to said processing system, said interface including a common mode cancellation amplifier circuit configured to reduce common mode signal noise present in said external signals, and at least one amplifier circuit configured to alter an amplitude of said external signals; andwherein said processing system is configured for processing representations of said received external signals with at least one software algorithm to generate at least one index value representative of a patient condition. 2. The apparatus of claim 1 wherein said interface includes a first stage gain amplifier for amplifying a received electrical signal, said first stage gain amplifier operatively coupled to a gain setting network; and wherein said gain setting network is a resistor network configured to provide a gain setting for said first stage gain amplifier and to provide a common-mode signal to said common mode cancellation amplifier circuit. 3. The apparatus of claim 2 wherein said interface further includes an input radio-frequency interference filter circuit coupled between a source of said received external electrical signals and said first stage gain amplifier, said input radio-frequency interference filter circuit including at least one resistive component and at least one capacitive component configured to reduce differential and common-mode radio frequency interference present in said received electrical signals. 4. The apparatus of claim 1 wherein said interface further includes a second amplifier circuit coupled between said first amplifier circuit and said processing system, said second amplifier circuit configured to provide a bandpass filter and gain function for said representation of said electrical signals. 5. The apparatus of claim 4 wherein said second amplifier circuit includes a first amplification stage and a second amplification stage, said first stage having a gain of 300× and said second stage having a gain of 2×. 6. The apparatus of claim 4 wherein said second amplifier circuit is configured to include passband corners of 2 Hz and 1000 Hz. 7. The apparatus of claim 2 wherein said common mode cancellation amplifier circuit includes a high-pass filter circuit for receiving said common mode signal, a unity gain buffer operatively coupled to receive an output of said high-pass filter circuit, and a low-pass filter and gain stage operatively coupled to said unity gain buffer, said low-pass filter and gain stage configured to output a reference signal for cancelling common-mode voltage present at a source of said external electrical signals. 8. The apparatus of claim 1 wherein external signals include at least one signal selected from a set of external si... | EVEREST BIOMEDICAL INSTRUMENTS CO. | A61B005/04
US20080319277 | A1 | 2008-12-25 | SLEEP DISORDER MONITORING AND DIAGNOSTIC SYSTEM | 2006-06-12 | A portable or wearable system for monitoring and diagnosing sleep disorders, such as sleep apnea, and an associated method of monitoring and diagnosis. The device which can be used for the detection, assessment, diagnosis and pre-diagnosis (screening) of sleep apnea, as well as other sleep-related disorders associated with sleep apnea, such as hypopnea, snoring and abnormal cardiac rhythms. The device preferably samples, stores and records sound at a frequency of 1000 Hz and higher to allow for an accurate analysis of the subject's condition to be carried out. Memory is provided in the device to store at least six hours of continuous data. Data collected by the device can be downloaded to an external computing device for later use and analysis by a medical professional. | 1. A sleep disorder monitoring and diagnostic device, comprising: a processing and recording unit to be worn by a subject for monitoring and diagnosis of a sleep disorder in the subject, the processing and recording unit having:a plurality of connectors to permit reconfigurable attachment of various physiological sensors for sensing physiological conditions of the subject;processing means for sampling and processing signals from the physiological sensors, at least one of the signals being sampled at a rate of at least about 1000 Hz; andstorage means for recording the sampled and processed signals. 2. The device of claim 1, wherein the plurality of connectors include two or more different connector types. 3. The device of claim 2, wherein the two or more different connector types are selected from the group consisting of leur lock connectors, auxiliary connectors, and pin keyed connectors. 4. The device of claim 1, wherein the physiological sensors include a microphone. 5. The device of claim 4, wherein a sound sample rate for the signal obtained from the microphone is at least 1000 Hz. 6. The device of claim 1, wherein the physiological sensors are selected from the group consisting of oxyhemoglobin sensors, pulse rate sensors, electrocardiogram (ECG) sensors, electroencephalography (EEG) sensors, electromyography (EMG) sensors and respiratory effort sensors. 7. The device of claim 1, further including an airflow pressure sensor for use with a nasal cannula. 8. The device of claim 1, further including a body position detector. 9. A sleep disorder and diagnostic device kit, comprising: a plurality of physiological sensors for sensing physiological conditions of a subject; anda processing and recording unit to be worn by a subject for monitoring and diagnosis of a sleep disorder in the subject, the processing and recording unit having a plurality of connectors to permit reconfigurable attachment of the plurality of physiological sensors, a processing means for sampling and processing signals from the physiological sensors, at least one of the signals being sampled at a rate of at least about 1000 Hz, and a storage means for recording the sampled and processed signals. 10. The kit of claim 8, wherein the kit is a single-use kit. 11. The kit of claim 8, wherein the plurality of connectors are selected from the group consisting of leur lock connectors, auxiliary connectors, and pin keyed connectors. 12. The kit of claim 8, wherein the plurality of physiological sensors include a microphone. 13. The device of claim 12, wherein a sound sample rate for the signal obtained from the microphone is at least 1000 Hz. 14. The kit of claim 8, wherein the plurality of physiological sensors are selected from the group consisting of oxyhemoglobin sensor, pulse rate sensors, electrocardiogram (ECG) sensors, electroencephalography (EEG) sensors, electromyography (EMG) sensors and respiratory effort sensors. 15. The kit of claim 8, wherein the processin... | BRAEBON MEDICAL CORPORATION | A61B005/00
US20080255464 | A1 | 2008-10-16 | SYSTEM AND METHOD FOR DIAGNOSING AND TREATING LONG QT SYNDROME | 2007-04-10 | A system for diagnosing Long QT Syndrome (LQTS) derives a QT/QS2 ratio from an electrical systole (QT) and a mechanical systole (QS2) to detect a prolonged QT interval in a patient's cardiac cycle. A processor acquires the systoles from a microphone and chest electrodes, calculates the QT/QS2 ratio, and outputs the result to a display. The processor may compare the QT/QS2 ratio to a threshold value stored in memory for diagnosing LQTS in the patient. A user interface provides for programming, set-up, and customizing the display. A mode selector allows the system to operate alternatively as a phonocardiograph, a 12 lead electrocardiograph, or a machine for diagnosing LQTS. A related method for diagnosing cardiac disorders such as LQTS includes measuring QT and QS2 during a same cardiac cycle, calculating a QT/QS2 ratio, and comparing the result to a threshold value derived from empirical data. The method may include measuring systoles both at rest and during exercise, and may be used for drug efficacy, dosage optimization, and acquired LQTS causality tests. | 1. A system for diagnosing LQTS in a patient, comprising: electrodes for detecting an electrical systole, QT, in the patient;at least one microphone for detecting a mechanical systole, QS2, in the patient;a processing module receiving the detected electrical systole and the detected mechanical systole, calculating a QT/QS2 ratio, and outputting the ratio as a diagnostic result. 2. The system of claim 1 further comprising a display unit coupled to the processing module for displaying the QT/QS2 ratio. 3. The system of claim 2 wherein the display unit displays a first waveform representing the electrical systole and a second waveform representing the mechanical systole. 4. The system of claim 1 further comprising a mode selector for selecting between first and second operating modes of the processing module, the processing module in the first operating mode operating as an electrocardiograph, the processing module in the second operating mode calculating the QT/QS2 ratio. 5. The system of claim 1 wherein the electrodes comprise a minimum of ECG leads II, V2 and V3. 6. The system of claim 1 further comprising a user interface coupled to the processing module. 7. The system of claim 1 wherein the processing module further comprises an analog-to-digital converter receiving analog input from the electrodes and the at least one microphone. 8. The system of claim 7 wherein the processing module further comprises a signal conditioning module receiving digital output from the analog-to-digital converter. 9. The system of claim 8 wherein the processing module further comprises a microprocessor coupled to memory. 10. The system of claim 9 wherein the microprocessor calculates the QT/QS2 ratio based on output from the signal conditioning module. 11. The system of claim 10 wherein the microprocessor calculates the QT/QS2 ratio by executing an algorithm stored in the memory. 12. The system of claim 10 wherein the microprocessor compares the QT/QS2 ratio to a threshold value stored in the memory. 13. The system of claim 12 wherein if the QT/QS2 ratio exceeds the threshold value, the microprocessor outputs an alarm. 14. The system of claim 13 further comprising a display unit, wherein the microprocessor outputs the alarm to the display unit. 15. The system of claim 1 wherein the detected electrical systole and the detected mechanical systole received by the processor for calculating the QT/QS2 ratio occur during a same cardiac cycle. 16. The system of claim 1 wherein the electrical systole comprises a time interval during one period of a cardiac cycle from onset of electrical depolarization of heart ventricles to end of electrical repolarization of the heart ventricles. 17. The system of claim 1 wherein the mechanical systole comprises a time interval during one period of a cardiac cycle from onset of electrical depolarization of heart ventricles to occurrence of second heart sound S2. 18. A module for diagnosing LQTS in a patient, comprising: el... | - | A61B005/0452
US20080312523 | A1 | 2008-12-18 | Apparatus and Method for Reducing Interference | 2005-06-07 | An electronic apparatus for reducing interference in a desired signal, the apparatus comprising: (a) a plurality of measurement signal lines, each connected to a respective measurement signal electrode; and (b) one or more reference signal lines, each connected to respective one or more reference electrodes; each of said measurement signal lines or a respective group of said measurement signal lines being associated by being in close physical proximity with a respective one of said reference signal lines for a substantial part of their lengths, so that each measurement signal line or signal line group with its corresponding reference signal line forms a measurement signal line or measurement signal line group/reference signal line pair, said electronic apparatus further comprising subtraction means for subtracting an interference on each reference signal line from an interference signal on the associated measurement signal line or from each measurement signal line in the measurement signal line group in that measurement signal line or measurement signal line group/reference signal line pair, wherein at least one of the measurement signal electrodes is arranged to be in direct electrical connection with a subject and at least one of the reference signal electrodes is arranged to be in close physical proximity but not in direct electrical contact with the subject. | 1. An electronic apparatus for reducing interference in a desired signal, the apparatus comprising:— (a) a plurality of measurement signal lines, each connected to a respective measurement signal electrode; and(b) one or more reference signal lines, each connected to a respective one or more reference electrodes;each of said measurement signal lines or a respective group of said measurement signal lines being associated by being in close physical proximity with a respective one of said reference signal lines for a substantial part of their lengths, so that each measurement signal line or signal line group with its corresponding reference signal line forms a measurement signal line or measurement signal line group/reference signal line pair, said electronic apparatus further comprising subtraction means for subtracting an interference signal on each reference signal line from an interference signal on the associated measurement signal line or from each measurement signal line in the measurement signal line group in that measurement signal line or measurement signal line group/reference signal line pair;wherein at least one of the measurement signal electrodes is arranged to be in direct electrical connection with a subject and at least one of the reference signal electrodes is arranged to be in close physical proximity but not in direct electrical contact with the subject. 2. An electronic apparatus according to claim 1 further comprising an electrically conductive mesh comprising one or more of said reference electrodes. 3. An apparatus according to claim 2 wherein an insulating layer is provided for insulating the conductive mesh from a subject. 4. An apparatus according to any one of claim 2 or 3 wherein said conductive mesh comprises a continuous laminar member. 5. An apparatus according to any one of claim 2 or 3 wherein said conductive mesh comprises a matrix of discrete members respectively comprising said reference electrodes. 6. An apparatus according to any of claims 2 to 5 further comprising an electrode support structure for supporting said electrodes and said conductive mesh. 7. An apparatus according to claim 6, wherein said electrode support structure comprises a flexible cap. 8. An apparatus according to claim 6 wherein said electrode support structure comprises a rigid cap for supporting said electrodes, said conductive mesh being flexible. 9. An apparatus according to any one of claims 6 to 8, wherein said electrode support structure is arranged to effect an EPM. 10. An apparatus according to any one of claims 6 to 9, wherein the electrode support structure apparatus further comprises an electrode support having supported thereon, an array of said measurement signal electrodes presented for contacting the skin of a subject, first connection means being provided for independent electrical connection to each of said measurement signal electrodes, the electrically conductive mesh further having second connection means for i... | - | A61B005/055 A61B005/053
US20080269590 | A1 | 2008-10-30 | MEDICAL INSTRUMENT FOR PERFORMING A MEDICAL INTERVENTION | 2008-04-18 | A medical instrument for performing a medical intervention on a body region is provided. The medical instrument includes a sensor for recording tissue parameters of a body region to be treated using measuring techniques. | 1. A medical instrument for performing medical interventions on a body region, the medical instrument comprising: at least one sensor for recording tissue parameters of a body region to be treated. 2. The medical instrument as claimed in claim 1, wherein the sensor is coupled with an evaluation unit. 3. The medical instrument as claimed in claim 2, wherein the evaluation unit comprises an output unit for outputting optical and/or acoustic signals. 4. The medical instrument as claimed in claim 1, wherein the evaluation unit comprises a signal processing unit for subdividing the examined body tissue into vulnerable body tissue and non-vulnerable body tissue. 5. The medical instrument as claimed in claim 1, wherein the evaluation unit includes a signal processing unit for identifying a classified tissue type. 6. The medical instrument as claimed in claim 1, wherein the evaluation unit is bidirectionally coupled. 7. The medical instrument as claimed in claim 1, wherein the evaluation unit is coupled with at least one further processing device. 8. The medical instrument as claimed in claim 1, wherein the sensor is an optical and/or an electrical sensor. 9. The medical instrument as claimed in claim 1, wherein the sensor comprises at least one transmitter and at least one receiver. 10. The medical instrument as claimed in claim 1 wherein the medical instrument is an endoscopic instrument, a scalpel, an HF scalpel, a medical robot, or a surgical laser. 11. The medical instrument as claimed in claim 5, wherein the classified tissue type includes malignant tissue. 12. A method for controlling a medical instrument during medical intervention, the method comprising: identifying a tissue type with a sensor; andcontrolling a function of the medical instrument as a function of the identified tissue type. 13. The method as claimed in claim 12, wherein the tissue type is healthy tissue or diseased tissue. 14. The method as claimed in claim 13, wherein controlling the function of the medical insturment includes stopping the fucntion of the medical instrument when healthy tissue is identified. 15. The method as claimed in claim 13, wherein controlling the function of the medical insturment includes starting the function of the medical instrument when diseased tissue is identified. | - | A61B005/00 A61B017/00
US20080312516 | A1 | 2008-12-18 | Living body information detection and display method | 2008-06-05 | A living body information display apparatus, together with a sleeping posture and position detection apparatus, is used to correctly and suitably capture an abnormal condition in a living body while the living body is sleeping/lying on a bed or on the apparatus. Living body information and sleeping posture/position are captured by pressure sensors placed under a sleeper. The living body information includes, for example, respiratory information, body movement information, and sleeping posture. The information of the sleeper is represented in color-coded time chart, in writing and in a rough image based on the sensor signals. By using this apparatus, a user can readily understands the position, posture and respiratory information of the sleeper. | 1. A method for displaying living body information comprising: arranging a plurality of sensors under a living body in rows and columns approximately perpendicular and parallel to a longitudinal direction of the living body, and positioning the sensors in predetermined intervals;outputting signals from the plurality of sensors, the signals corresponding to at least one of a pressure and variation of pressure of the living body;detecting and outputting the living body information of the living body;obtaining the signals from the plurality of sensors, processing the signals to display a difference of the at least one of the pressure and variation of pressure in a color coded manner, and outputting processed signals providing a posture of the living body in the color coded manner;outputting the processed signals at an end of each body movement over a period of time, and outputting the living body information over the same period of time as the posture of the living body; anddisplaying the posture of the living body and the living body information. 2. The method according to claim 1, wherein the detecting and outputting the living body information includes:obtaining the signals from the sensors;converting the signals to the living body information; andoutputting converted signals. 3. The method according to claim 2, wherein the living body information includes at least one or more of a respiration, a body movement, a pulse, a thoracoabdominal movement, and a position of the living body. 4. The method according to claim 1, wherein in the processing the signals, the signals are processed in order to display only a portion of the living body as the posture including an upper body of the living body, andthe outputting the processed signals providing only the portion of the living body in the color coded manner. 5. The method according to claim 1, wherein the processing the signals is performed only when a notable change or an abnormality occurs in the living body information so that the displaying the posture of the living body is performed only when the notable change or the abnormality occurs in the living body information. 6. A method for displaying living body information comprising: arranging a plurality of sensors under a living body in rows and columns approximately perpendicular and parallel to a longitudinal direction of the living body, and positioning the sensors in predetermined intervals;outputting signals from the plurality of sensors, the signals corresponding to at least one of a pressure and variation of pressure of the living body;obtaining the signals from the plurality of sensors, processing the signals to display a difference of the at least one of the pressure and variation of pressure in a color coded manner, and outputting processed signals providing a posture of the living body in the color coded manner;determining whether the living body is in one of a supine/prone position and a sideways position based on the signals and a... | DENSO CORPORATION | A61B005/00
US20080319330 | A1 | 2008-12-25 | TRANSMITTER AND RECEIVER FOR OBSERVING PERIODICAL EVENTS | 2008-06-25 | A mobile transmitter for observing events and transmitting data on the observed events to a receiver and a device for receiving data messages from the mobile transmitter. The transmitter and receiver are designed so as to provide efficient and reliable transmission of data from the mobile transmitter to the receiver device. | 1. A mobile transmitter for observing events and transmitting data on the observed events to a receiver, the mobile transmitter comprising: a detector for detecting occurrences of the observed events and responsive to the detection of each of said occurrences to generate a detection signal,a timer for providing time references, anda memory for recording time references,wherein the mobile transmitter is responsive to each detection signal to obtain a time reference from the timer and record the obtained time reference in the memory, and wherein the mobile transmitter is also operative to produce data messages containing a predetermined number of time references obtained from the memory and further to transmit the produced data messages to the receiver. 2. A mobile transmitter for observing events and transmitting data on the observed events to a receiver, the mobile transmitter comprising: a detector for detecting occurrences of the observed events and responsive to the detection of each of said occurrences to generate a detection signal,a counter for counting detection signals,a timer for providing time references, anda memory for recording time references,wherein the mobile transmitter is adapted to count detection signals in the counter up to a predetermined count, and as response to reaching said predetermined count, adapted to obtain a time reference from the timer and record the obtained time reference in the memory and further to immediately start a new count of detection signals up to the predetermined count, and wherein the mobile transmitter is also operative to produce data messages containing a predetermined number of time references obtained from the memory and further to transmit the produced data messages to the receiver. 3. The mobile transmitter of claim 1 or 2, wherein the mobile transmitter is operative to produce and transmit the data messages at predetermined time intervals. 4. The mobile transmitter of claim 3, wherein each of the predetermined time intervals have a common constant value. 5. The mobile transmitter of claim 4, wherein the common constant value of the predetermined time intervals is in the range of 50 to 1000 ms and the predetermined number of time references in the data messages is from 1 to 10. 6. The mobile transmitter of claim 4, wherein the common constant value of the predetermined time intervals is in the range of 150 to 400 ms. 7. The mobile transmitter of claim 6, wherein the predetermined number of time references in the data messages is from 2 to 4. 8. The mobile transmitter of claim 1 or 2, wherein the mobile transmitter is operative to produce the data messages such that each of the data messages contains the predetermined number of the latest time references in the memory. 9. The mobile transmitter of claim 8, wherein the mobile transmitter is operative to produce the data messages periodically such that a predetermined interval of time lapses between each of the successive data messages. ... | SUUNTO OY | A61B005/0404 H04B007/00
US07468043 | B2 | 2008-12-23 | Method and apparatus for noninvasive intraductal fluid diagnositc screen | 2005-04-05 | Disclosed is a method and apparatus for accomplishing a noninvasive screen for breast disease markers, including breast cancer markers and cytologically abnormal cells. Intraductal fluid is noninvasively aspirated using compression, heating and suction cycles. The removed sample is thereafter assayed for the presence of cytologically abnormal cells and/or one or more breast disease markers. In addition to the diagnostic applications of the disclosed method and apparatus, a therapeutic application is also contemplated wherein intraductal fluids are noninvasively aspirated using compression, heating and suction cycles to remove toxic buildups within the fluid. | 1. An intraductal fluid withdrawal system, comprising: a control unit;a patient interface unit, including an adjustable support having a plurality of petals which are adjustable throughout a range of motion and a bladder having a plurality of inflatable lobes for applying pressure to a breast;a control line extending between the control unit and the patient interface unit; anda closed fluid circulation loop, having a first component removably carried by the control unit and the bladder carried by the patient interface unit. 2. An intraductal fluid withdrawal system as in claim 1, wherein the first component comprises a reservoir. 3. An intraductal fluid withdrawal system as in claim 2, wherein the circulation loop comprises a movable wall. 4. An intraductal fluid withdrawal system as in claim 3, wherein the circulation loop comprises a compressible tube. 5. An intraductal fluid withdrawal system as in claim 1, wherein the bladder comprises at least 3 inflatable lobes. 6. An intraductal fluid withdrawal system as in claim 5, comprising at least 6 inflatable lobes. 7. An intraductal fluid withdrawal system as in claim 5, wherein each lobe has an inflated width of no more than about 3 inches and an inflated length of no more than about 4 inches. 8. An intraductal fluid withdrawal system as in claim 7, wherein each lobe has an inflated width of no more than about 2 inches and an inflated length of no more than about 3 inches. 9. An intraductal fluid withdrawal system as in claim 5, wherein each lobe has an inflated thickness of no more than about 1 inch. 10. An intraductal fluid withdrawal system as in claim 9, wherein each lobe has an inflated thickness of no more than about 0.5 inches. 11. An intraductal fluid withdrawal system as in claim 1, further comprising a heat exchange fluid contained within the closed loop. 12. An intraductal fluid withdrawal system as in claim 1, further comprising a heat source in the control unit. 13. An intraductal fluid withdrawal system as in claim 12, wherein the first component is removably in thermal communication with the heat source. 14. An intraductal fluid withdrawal system as in claim 1, further comprising a pump in the control unit. 15. An intraductal fluid withdrawal system as in claim 14, wherein the fluid circulation loop is positioned such that the pump causes fluid circulation through the loop. 16. An intraductal fluid withdrawal system as in claim 1, further comprising a disposable patient interface carried by the patient interface unit, for contacting the patient. 17. An intraductal fluid withdrawal system as in claim 1, wherein the disposable patient interface comprises a flexible polymeric membrane. 18. An intraductal fluid withdrawal system, comprising: a patient interface unit, including an adjustable support having a plurality of petals which are adjustable throughout a range of motion and a bladder having a plurality of inflatable lobes for applying pressure to a breast;a contro... | Neomatrix, LLC | A61B005/00 B65D081/00
US20080300494 | A1 | 2008-12-04 | REAL-TIME MEASUREMENT OF VENTRICULAR STROKE VOLUME VARIATIONS BY CONTINUOUS ARTERIAL PULSE CONTOUR ANALYSIS | 2008-08-12 | Ventricular stroke volume variation (SVV) is estimated as a function of the standard deviation of arterial blood pressure value measured over each of at least two cardiac cycles, preferably over each of the cardiac cycles in a computation interval covering a full respiratory cycle. In one embodiment, maximum and minimum standard deviation values are determined over the computation interval. SVV is then estimated proportional to the ratio of the difference between the maximum and minimum standard deviation values and the mean of the standard deviation values. In another embodiment, SVV is then estimated proportional to the ratio of the standard deviation of the standard deviation values and the mean standard deviation over the entire computation interval. A pre-processing arrangement for improving reliability of estimates of more general cardiac or hemodynamic parameters is also disclosed and involves smoothing with an approximating function, and sampling and low-pass filtering at an adjustable rate. | 1. A method for determining a cardiac or hemodynamic output value comprising: inputting a waveform data set corresponding to measurements of a measurement parameter;generating a series of measurement values from the waveform data set;computing an approximating function that best matches the measurement values according to a predetermined metric;over each of at least one computation interval, creating a set of sampled, approximating values by sampling the approximating function at an interval-specific sampling rate;low-pass filtering the sampled, approximating values; andcalculating an estimate of the output value as a function of the low-pass filtered, sampled approximating values. 2. A method as in claim 1, in which the output value displays respiration-induced variation, further comprising: identifying respiratory cycles;detecting cardiac cycles;setting the computation interval to be a respiratory cycle; andfor each computation interval, setting the interval-specific sampling rate as a function of the number of cardiac cycles in each respective respiratory cycle. 3. A method as in claim 1, in which the step of generating the series of measurement values includes pre-processing the waveform data set. 4. A method as in claim 1, in which the waveform data set corresponds to blood pressure, further comprising computing each measurement value to be a function of the standard deviation of blood pressure over each respective computation interval. 5. A method as in claim 1, in which the output value is systolic pressure variation. 6. A method as in claim 1, in which the output value is pulse pressure variation. 7. A system for determining a cardiac or hemodynamic output value comprising: an arrangement for determining and inputting a waveform data set corresponding to measurements of a measurement parameter;a low-pass filter;a processing system including computer-executable code for generating a series of measurement values from the waveform data set; for computing an approximating function that best matches the measurement values according to a predetermined metric;over each of at least one computation interval, for creating a set of sampled, approximating values by sampling the approximating function at an interval-specific sampling rate;calculating an estimate of the output value as a function of the sampled approximating values after filtering by the low-pass filter. 8. A system as in claim 7, in which the output value displays respiration-induced variation, further comprising: an arrangement for identifying respiratory cycles;an arrangement for detecting cardiac cycles;in which the processing system further includes additional computer-executable code for setting the computation interval to be a respiratory cycle; andfor each computation interval, setting the interval-specific sampling rate as a function of the number of cardiac cycles in each respective respiratory cycle. 9. A system as in claim 7, further including a pre-processing modul... | Edwards Lifesciences Corporation | A61B005/029 A61B005/021
US20080243018 | A1 | 2008-10-02 | SYSTEM AND METHOD TO TRACK A RESPIRATORY CYCLE OF A SUBJECT | 2007-03-30 | A system operable to track a respiratory cycle of a subject is provided. The system includes at least a first sensor positioned on the subject, and at least a second sensor located at a reference relative to a change in position of the first sensor associated with respiration of the subject. The system also includes a respiratory cycle measurement device coupled to receive the position data of the first sensor relative to the reference. The respiratory cycle measurement device is configured to translate the position data of the first sensor relative to time into a respiratory signal representative of a respiratory cycle of the subject. | 1. A system to track a respiratory cycle of a subject, the system comprising: at least a first sensor positioned on the subject;at least a second sensor located at a reference relative to a change in position of the first sensor associated with respiration of the subject; anda respiratory cycle measurement device coupled to receive the position data of the first sensor relative to the reference, wherein the respiratory cycle measurement device is configured to translate the position data of the first sensor in relation to the reference relative to time into a respiratory signal representative of a respiratory cycle of the subject. 2. The system of claim 1, wherein the respiratory cycle measurement device includes a display device coupled in communication to receive and illustrate a display of the respiratory signal. 3. The system of claim 2, wherein the respiratory cycle measurement device includes a processor operable to generate a gate signal in based on a comparison of the position of the first sensor, wherein the gate signal is in an ON state when the position of the first sensor is within a first threshold of a peak position and a second threshold of a minimum position relative to a position of the second sensor, and wherein the gate signal is in an OFF state when the position of the first sensor is detected not within the first and second thresholds relative to the second sensor. 4. A system to acquire an image data of an imaged subject, the system comprising: an imaging system operable to selectively acquire the image data of the imaged subject; anda respiratory cycle measurement device coupled to receive a position data of a first sensor at the imaged subject in relation to a second sensor at a reference, wherein the respiratory cycle measurement device translates the position data relative to time into a respiratory signal representative of a respiratory cycle of the imaged subject. 5. The system of claim 4, wherein the reference comprises a positioning assembly configured to support the imaged subject or the imaging system, and wherein the tracking system includes at least a first sensor positioned at the imaged subject and at least a second sensor located at a reference relative to a position of the imaged subject, wherein the second sensor is configured to detect position data of the first sensor relative the second sensor. 6. The system of claim 4, wherein the respiratory cycle measurement device includes a processor operable to generate a gate signal correlated to the respiratory signal. 7. The system of claim 6, wherein the processor synchronizes the gate signal with a clock output signal. 8. The system of claim 7, wherein the processor compares the position data of the first sensor relative to a threshold, and wherein the gate signal is in an ON state when the position data of the first sensor is within the threshold, and wherein the gate signal is in an OFF state when the position data of the first sensor is detected not w... | GENERAL ELECTRIC COMPANY | A61B005/08
US20080242944 | A1 | 2008-10-02 | METHOD AND SYSTEM FOR FACILITATING ERROR FREE SCAN ADMINISTRATION | 2007-03-30 | A method and system for facilitating error free scan is disclosed herein. The system for preventing wrong scan administration to a patient comprises: an identifier associated with a patient for identifying at least one patient parameter. The patient parameter may be configured to indicate patient information such as pregnancy or the presence of metallic implants. An interrogator is provided for automatically interrogating the identifier to obtain the at least one patient parameter during one or more predefined stages when the patient is in proximity of a scanner. A controller is operatively coupled to the interrogator for generating a wrong scan signal based on the communicated information. | 1. A system for preventing wrong scan administration to a patient, comprising: (a) an identifier associated with a patient for identifying at least one patient parameter;(b) an interrogator for automatically interrogating the identifier to obtain the at least one patient parameter when the patient is in proximity of a scanner; and(c) a controller operatively coupled to the interrogator, wherein the controller is configured to generate a wrong scan signal based on the at least one patient parameter. 2. A system as in claim 1, wherein the identifier is configured to identify a patient parameter indicative of whether the patient is pregnant. 3. A system as in claim 1, wherein the identifier is configured to identify a patient parameter indicative of whether the patient has a metallic implant. 4. A system as in claim 1, wherein the identifier is an RFID tag. 5. A system as in claim 1, wherein the identifier is selected from a group consisting of a barcode, a magnetic strip, and a wireless transponder. 6. A system as in claim 1, wherein the interrogator is associated with a scanner. 7. A system as in claim 1, wherein the interrogator is associated with a scanning room. 8. A system as in claim 4, wherein the interrogator is an RFID reader. 9. A system as in claim 5, wherein the interrogator is selected from a group consisting of a barcode reader, a magnetic strip reader and a wireless receiver. 10. A system as in claim 1, wherein the interrogator is configured to obtain the at least one patient parameter upon the patient entering into a scanning room. 11. A system as in claim 1, wherein the interrogator is configured to obtain the at least one patient parameter upon the patient being positioned in a scanner. 12. A system as in claim 1, wherein the interrogator includes a first interrogator and a second interrogator, the first interrogator being associated with a scanning room and the second interrogator being associated with a scanner. 13. A system as in claim 12, wherein the controller includes a first controller and a second controller, the first controller being associated with the first interrogator and the second controller being associated with the second interrogator. 14. A system as in claim 13, wherein the controller is configured to generate the wrong scan signal when the patient parameter indicates the patient is pregnant. 15. A system as in claim 13, wherein the controller is configured to generate the wrong scan signal when the patient parameter indicates the patient has a metallic implant. 16. A system as in claim 1, wherein the wrong scan signal is a scan inhibit signal for inhibiting a scan of the patient when applied to a scanner. 17. A system as in claim 1, wherein the system further comprises an alarm device, wherein the alarm device is configured to generate an operator-perceptible alarm in response to the wrong scan signal. 18. A method of facilitating error free scanning of a patient, comprising: (a) providing a p... | GENERAL ELECTRIC COMPANY | A61B005/00
US20080275351 | A1 | 2008-11-06 | Model-based pulse wave velocity measurement method | 2007-11-02 | A method for measuring pulse wave velocity through an artery. The method includes: obtaining an image of the artery; generating from the image a model of the artery; selecting two reference points along the artery, such two reference points being separated a predetermined distance, D; observing a change in the model at a first one of the two reference points and a corresponding change in the model at a second one of the two reference points along with a time difference T between the observed changes; and estimating the pulse wave velocity, PWV, in accordance with PWV=D/T. | 1. A method for measuring pulse wave velocity through an artery, comprising: obtaining an image of the artery;generating from the image a model of the artery;selecting two reference points along the artery, such two reference points being separated a predetermined distance, D;observing a change in the model at a first one of the two reference points and a corresponding change in the model at a second one of the two reference points along with a time difference T between the observed changes;estimating the pulse wave velocity, PWV, in accordance with PWV=D/T. 2. The method recited in claim 1 wherein the change is a sudden expansion of the artery along a path between the two reference points. 3. The method recited in claim 2 wherein the expansion is observed as an upslope in an artery diameter curve at each of the two reference points. 4. The method recited in claim 3 wherein the reference points are disposed along diameters of the artery. 5. The method recited in claim 4 wherein the reference points are disposed along centerlines of the artery. 6. The method recited in claim 5 wherein the onset of the radius increase in the artery is tracked using a quadratic sub-frame interpolation scheme. 7. The method recited in claim 1 wherein the distance is derived directly from the model by integrating the centerline length of the centerline of the artery between the two reference points. | Siemens Corporate Research, Inc. | A61B005/02
US20080269628 | A1 | 2008-10-30 | Denoising and Artifact Rejection for Cardiac Signal in a Sensis System | 2007-07-31 | A method and apparatus for denoising and rejecting artifacts from cardiac signals, having the steps of accepting a cardiac signal from a patient, subjecting the cardiac signal from the patient to a frequency band width controllable choke to separate the cardiac signal into predefined frequencies, filtering each of the predefined frequencies to remove dynamic common noise, joining each of the predefined frequencies into a cardiac signal without the dynamic common noise, and providing a feedback control to the filtering of each of the predefined frequencies. | 1. A method for denoising and rejecting artifacts from cardiac signals, comprising: accepting a cardiac signal from a patient;separating the cardiac signal into predefined frequencies;filtering each of the predefined frequencies to remove dynamic noise;joining each of the predefined frequencies into a cardiac signal without the dynamic noise; andproviding a feedback control to the filtering of each of the predefined frequencies. 2. The method according to claim 1, wherein the filtering of each of the predefined frequencies to remove dynamic noise is accomplished by identification of interference signals from medical instruments and elimination of noise related to the medical instruments. 3. The method according to claim 1, wherein the dynamic noise removed is a non-linear signal. 4. The method according to claim 1, wherein the signals from the patient are controlled through a computer programmable selection that decreases common mode noise. 5. The method according to claim 1, wherein the filtering of each of the predefined frequencies is through empirical mode decomposition processing. 6. A method for denoising and rejecting artifacts from cardiac signals, comprising: accepting a cardiac signal from a patient;separating the cardiac signal from the patient into predefined frequencies;filtering each of the predefined frequencies to remove dynamic noise; andjoining each of the predefined frequencies into a cardiac signal without the dynamic noise. 7. The method according to claim 6, wherein the filtering of each of the predefined frequencies is through empirical mode decomposition processing. 8. The method according to claim 6, wherein the filtering of each of the predefined frequencies to remove dynamic noise is accomplished by identification of interference signals from medical instruments and elimination of noise related to the medical instruments. 9. The method according to claim 6, wherein the dynamic noise removed is a non-linear signal. 10. A program storage device readable by machine, tangibly embodying a program of instructions executable by the machine to perform method steps for denoising and rejecting artifacts from cardiac signals, comprising: accepting a cardiac signal from a patient;separating the cardiac signal from the patient into predefined frequencies;filtering each of the predefined frequencies to remove dynamic noise;joining each of the predefined frequencies into a cardiac signal without the dynamic noise; andproviding a feedback control to the filtering of each of the predefined frequencies. 11. The device according to claim 10, wherein the filtering of each of the predefined frequencies is through empirical mode decomposition processing. 12. The device according to claim 10, wherein the filtering of each of the predefined frequencies to remove dynamic noise is accomplished by identification of interference signals from medical instruments and elimination of noise related to the medical instruments. 13. The device... | Siemens Medical Solutions USA, Inc. | A61B005/0402
US07457657 | B2 | 2008-11-25 | Magnetic resonance method and apparatus for determining the position and/or orientation of the image plane of slice image exposures of a vessel region in a contrast agent bolus examination | 2005-01-21 | In a method and a magnetic resonance system for determining the position and/or orientation of the image plane of slice image exposures of a vessel region in a contrast agent bolus examination, in particular in the region of the aorta bifurcation or the aorta arch, wherein the image plane traverses the vessel region in the longitudinal section, initially a group of slice image exposures disposed essentially orthogonal to the vessel is automatically acquired from different planes of the vessel region lying parallel to one another, and the image regions showing the vessel region are automatically determined within the orthogonal slice image exposures in an image-processing device, and using the determined image regions the spatial position of the vessel region is determined. The position and/or orientation of the image plane disposed essentially orthogonal to the planes of the slice image exposures and traversing the vessel region in the longitudinal section is determined automatically using the position information, such that the image plane lies in the center of the vessel region. | 1. A method for determining at least one of a position and orientation of an image plane of a plurality of slice image exposures of a vessel region in a magnetic resonance contrast agent bolus examination, comprising the steps of: automatically acquiring a plurality of slice image exposures in respectively different planes of a vessel region encompassing a vessel into which a contrast agent bolus is yet to be administered, said slice image exposures being disposed parallel to each other and substantially orthogonally to the vessel;automatically electronically determining respective image regions, showing said vessel region, in said orthogonal slice image exposures in an image processor;using the determined image regions, automatically electronically determining a spatial position of said vessel region in said image processor;in said image processor automatically electronically determining at least one of a position and orientation of an image plane, disposed substantially orthogonal to the respective planes of the slice image exposures and traversing said vessel region in longitudinal section, using said spatial position of said vessel region, said image plane being disposed in a center of said vessel region;subsequently administering said contrast agent bolus and monitoring progress of said contrast agent bolus through said vessel using an image of said vessel region obtained in said image plane. 2. A method as claimed in claim 1 comprising obtaining said slice image exposures of a region containing the aorta arch, as said vessel region. 3. A method as claimed in claim 1 comprising obtaining said slice image exposures of a region containing the aorta bifurcation, as said vessel region. 4. A method as claimed in claim 1 comprising automatically electronically determining the respective image regions in said plurality of slice image exposures using a segmentation algorithm. 5. A method as claimed in claim 1 comprising using said at least one of said position and orientation of said image plane for automatically controlling said magnetic resonance system. 6. A method as claimed in claim 1 comprising acquiring at least one localizer slice image before acquiring said slice image exposures, and using said localizer slice image in said image processor to determine at least one of a coarse position and a coarse orientation of said vessel region and the respective positions of the planes of said slice image exposures. 7. A method as claimed in claim 6 comprising obtaining a plurality of localizer slice images in succession. 8. A magnetic resonance apparatus comprising: a magnetic scanner that automatically acquires a plurality of slice image exposures in respectively different planes of a region encompassing a vessel into which a contrast agent bolus is yet to be administered, said slice image exposures being disposed parallel to each other and substantially orthogonally to the vessel;an image processor supplied with said slice image exposures, sa... | Siemens Aktiengesellschaft | A61B005/05
US07612560 | B2 | 2009-11-03 | Selectively adjustable phantom that is compatible with a magnetic resonance imaging system and environment | 2007-04-24 | A phantom for use in an MRI or MRS system includes an array of subresolvable compartments that are each connected to receive fluid from one of a plurality of fluid reservoirs. The array of compartments is divided into sub-arrays which differ from each other by the mix of compartments in each sub-array receiving fluids from the different reservoirs. Studies can be performed with the phantom by filling the compartments with selected fluids from the reservoirs, disconnecting the reservoirs and placing the phantom in the system bore. The phantom can be reused in other studies by replacing the fluids with different fluids. | 1. A selectively adjustable phantom system configured for use in a medical imaging system that is compatible with a magnetic resonance imaging environment, the selectively adjustable phantom system comprising: A phantom housing having formed therein an array of compartments divided into sub-arrays that are distinguished from each other by a ratio of the amount of fluid resent within the compartments;a plurality of fluid circuits, each fluid circuit connected to at least one of said compartments in a sub-array;a plurality of detachable fluid reservoirs external to the phantom housing and selectively coupled through the plurality of fluid circuits into portions of the sub-arrays within the phantom housing; anda control system configured to allow user selection of the selectively coupled portions of the subarrays connecting with the plurality of detachable fluid reservoirs in order to allow the sub-arrays to be selectively filled to a desired level through the fluid circuits by the fluid reservoirs. 2. The selectively adjustable phantom recited in claim 1 in which each fluid reservoir contains a fluid comprised of various metabolites. 3. The selectively adjustable phantom as recited in claim 1 in which each sub-array has a subresolvable size with respect to the medical imaging system in which it is used. 4. The selectively adjustable phantom as recited in claim 3 in which the sub-arrays are arranged in a three-dimensional configuration within the phantom housing. 5. The selectively adjustable phantom as recited in claim 1 in which the compartments in each sub-array are contiguous to one another. 6. The selectively adjustable phantom as recited in claim 1 in which the medical imaging system is a magnetic resonance spectroscopy system and the reservoirs are filled with fluids having different resonance responses. 7. The selectively adjustable phantom as recited in claim 1 in which fluids within the compartments in the sub-arrays are selectively recirculated back into the fluid reservoirs. 8. The selectively adjustable phantom as recited in claim 7 in which the selected amount of fluids from the separate fluid reservoirs differ in each sub-array. | The General Hospital Corporation | G01V003/00 A61B005/055
US07447540 | B1 | 2008-11-04 | Systems and methods for detection of VT and VF from remote sensing electrodes | 2004-11-24 | Methods and systems are provided for performing ventricular arrhythmia monitoring using at least two sensing channels that are each associated with different sensing vectors, for example by different pairs of extracardiac remote sensing electrodes. Myopotential associated with each of the sensing channels in monitored, and a ventricular arrhythmia monitoring mode is selected based thereon (e.g., based on determined myopotential levels). Ventricular arrhythmia monitoring is then performed using the selected monitoring mode. This description is not intended to be a complete description of, or limit the scope of, the invention. Other features, aspects, and objects of the invention can be obtained from a review of the specification, the figures, and the claims. | 1. In an implantable system, a method for monitoring myopotential associated with a sensing vector comprising: (a) sensing a cardiac signal using a pair of extracardiac remote sensing electrodes corresponding to the sensing vector;(b) sampling the cardiac signal to produce a plurality of samples that are representative of a window of the cardiac signal; and(c) producing a histogram, for the window of the cardiac signal, based on the plurality of samples; fitting a line to the histogram distribution; determining a slope of the line fit to the histogram distribution; and comparing the slope of the line to at least one threshold to determine the level of myopotential. 2. The method of claim 1, further comprising filtering the cardiac signal sensed at step (a) and the samples produced at step (b) in order to filter out at least some of the frequencies that are not of interest, prior to step (c). 3. The method of claim 1, further comprising processing the cardiac signal sensed at step (a) and the samples produced at step (b) such that all the samples upon which monitoring is based at step (c) have a commom polarity. 4. An implantable system for detecting myopotential associated with a sensing vector, comprising: a sensing channel to sense a cardiac signal using a pair of extracardiac remote sensing electrodes corresponding to the sensing vector;a sampler to sample the cardiac signal to produce a plurality of samples that are representative of a window of the cardiac signal; anda myopotential detector to monitor myopotential based on the plurality of samples, wherein the myopotential detector is configured to: produce a histogram, for the window of the cardiac signal, based on the plurality of samples; fit a line to the histogram distribution; determine a slope of the line fit to the histogram distribution; and compare the slope of the line to at least one threshold to determine the level of myopotential. 5. In an implantable system, a method for updating at least one parameter that is used for monitoring myopotential associated with a sensing vector corresponding to a pair of extracardiac remote sensing electrodes implanted in a patient, the method comprising: determining whether the patient is active;if the patient is not active then determining whether a heart rate of the patient is within an acceptable range; andif the heart rate of the patient is within the acceptable range then updating the at least one parameter. 6. The method of claim 5, wherein the at least one parameter defines at least one of the following: (i) a ratio threshold;(ii) a slope threshold; and(iii) a highest bin edge and a lowest bin edge. 7. A method for updating at least one parameter that is used for monitoring myopotential associated with a sensing vector corresponding to a pair of subcutaneous extracardiac electrodes implanted in a patient, the method comprising: (a) monitoring whether an R-wave value has changed by at least a specific amount since the at least one pa... | Pacesetter, Inc. | A61B005/0452
US20080269639 | A1 | 2008-10-30 | Device for obtaining bodily fluids for analysis | 2008-03-14 | A lancing device for piercing skin to obtain a bodily fluid, such as blood, for testing has an electrically activated actuator, such as a shape memory alloy or electro-active polymer, that changes shape upon the application of electrical energy to advance and withdraw the lancing element. The lancing device operates when the actuator changes shape to advance a lancing element to pierce the skin when the actuator changes shape, delaying the lancing element tip near the advanced position to allow fluid to flow into the lancing tip, withdrawing the lancing element to remove the lancing element form the skin when the actuator again changes shape. | 1. A device for obtaining bodily fluids for analysis, comprising: a lancing element configured for insertion into a body;a movement converter coupled to the lancing element to control movement of the lancing element;a drive coupled to the lancing element and the movement converter to advance and withdraw the lancing element, the drive comprising, an electrical source coupled to the drive to selectively apply electrical energy to the drive,at least one actuator coupled to the electrical source, the actuator changes shape upon the application of electrical energy to advance the lancing element, and the actuator changes shape upon the removal of electrical energy to withdraw the lancing element. 2. The device as in claim 1, wherein the actuator is formed from a shape memory alloy that changes shape by contracting upon the application of electrical energy causing the shape memory alloy reaches a transformational temperature. 3. The device as in claim 1, wherein the actuator is formed from an electro-active polymer that changes shape by deformation upon the application of the electrical energy causing electrostatic attraction. 4. The device as in claim 1, further comprising, a dampener coupled to the drive to maintain the lancing element in an advanced position briefly to allow for fluid flow into a lancing tip before the lancing element moves to a withdraw position. 5. The device as in claim 1, further comprising, a return element coupled to the drive to aid in moving the lancing element from an advance position to a withdraw position. 6. The device as in claim 1, further comprising, a trigger switch coupled to the electrical source to provide a current pulse to the actuator causing the actuator to change shape. 7. A device for obtaining bodily fluids for analysis, comprising: a lancing element configured for insertion into a body part;a movement converter coupled to the lancing element to control movement of the lancing element;a drive coupled to the lancing element and the movement converter to advance and withdraw, the drive comprising, at least one actuator wire composed of a shape memory alloy coupled to the movement converter,a heating element thermally coupled to the actuator wire to contract the actuator wire when a transformational temperature is reached to advance the lancing element and expand when the transformational temperature is no longer reached to withdraw the lancing element. 8. The device as in claim 7, wherein the heating element is electrical energy selectively applied by an energy source. 9. The device as in claim 7, further comprising, a trigger switch coupled to a capacitor to provide a current pulse as the heating element to the actuator wire to cause the actuator wire to reach the transformational temperature. 10. The device as in claim 9, further comprising, movable electrical contact points that selectively engage the actuator wire to apply electrical energy to heat only the a selected portion of the actuator wir... | - | A61B005/151
US20080269575 | A1 | 2008-10-30 | Method and Apparatus for Monitoring Bodily Analytes | 2006-02-14 | The present invention relates mainly to a method and apparatus for monitoring bodily analytes, such as glucose, by use of an analyzer (100) configured for emitting radiation oriented to impinge on a selected area of the skin (S), for collecting returned imprinted radiation exiting the skin, for processing and analysis of the imprinted radiation, and for display of analysis results. At least one retroreflector (RR, RI 1, R1 & 2) is inserted subcutaneously to the selected area (SA) for receiving radiation and returning retroreflected radiation, as returned imprinted radiation. The retroreflector (RR, R1 1, R1 & 2) has at least one level of retroreflection (RR1L, RR2L). | 1-80. (canceled) 81. A method for monitoring bodily analytes by use of an analyzer (100) configured for emitting radiation oriented to impinge on a selected area (SA) of the skin (S), for collecting returned imprinted radiation exiting the skin, for processing and analysis of the imprinted radiation, and for display of analysis results, the method being characterized by the steps of: inserting at least one retroreflector (RR, R1, R2) subcutaneously to the selected area (SA) for receiving radiation and returning retroreflected radiation, andcollecting retroreflected radiation exiting from the selected area as returned imprinted radiation. 82. The method according to claim 81, wherein: imprinted radiation is collected separately from each at least one retroreflector (RR, R1, R2). 83. The method according to claim 81, wherein: the at least one retroreflector (RR, R1, R2) has at least one level of retroreflection (RR1L, RR2L). 84. The method according to claim 81, wherein: the at least one retroreflector (RR, R1, R2) is disposed appropriately and configured to return retroreflected radiation as a plurality of distinct separate imprinted beams collected for either one of both separate and combined processing and analysis. 85. The method according to claim 81, wherein: the at least one retroreflector (RR, R1, R2) is configured and disposed in selected position and orientation to receive radiation and return radiation crossing at least the skin (S) and tissue (T). 86. The method according to claim 81, wherein: the received and returned radiation have at least a first portion and a second portion of radiation crossing bodily matter via a separate, respectively, first path and a second path, andthe first path and the second path cross different bodily matter. 87. The method according to claim 81, wherein: subcutaneous insertion into a selected location (TL) is provided by ejection of the at least one retroreflector (RR, R1, R2) out of a hypodermic needle (205). 88. The method according to claim 87, wherein: subcutaneous insertion into a selected location (TL) is provided by ejection out of an open-end extremity (217) of a hypodermic needle and by operation in association with an imaging apparatus. 89. The method according to claim 88, wherein: the imaging apparatus is an ultrasound imaging apparatus. 90. The method according to claim 81, wherein: subcutaneous insertion into a selected location (TL) is provided by coupling the at least one retroreflector (RR, R1, R2) to at least one surgical thread attached to a surgical needle (205), and pulling the at least one surgical thread trough body tissue (T) until the at least one retroreflector (RR, R1, R2) reaches a selected location and orientation. 91. The method according to claim 81, wherein: the apparatus has a sampler (200) configured for emitting radiation and for collecting returned imprinted radiation, a processing unit (300) for processing imprinted radiation and for analysis, and a display... | - | A61B005/00
US07468033 | B2 | 2008-12-23 | Blood contacting sensor | 2004-09-08 | The invention provides methods and apparatus for detecting an analyte in blood. The apparatus is particularly suited for bringing a sensor into direct contact with blood in vivo. The apparatus comprises a sensor that detects the presence of an analyte and an assembly means. The assembly means has a sensor end, wherein the sensor end of the assembly means is affixed to the sensor, and the assembly means is adapted for coupling with a venous flow device. By coupling with a venous flow device, the assembly means brings the sensor into direct contact with blood flowing through the venous flow device. Examples of venous flow devices that bring the sensor into direct contact with the blood of a subject include, but are not limited to, intravenous catheters and external blood loops, such as are used in extra corporeal membrane oxygenation or hemodialysis. | 1. A method of introducing a sensor into the circulating blood of a subject comprising contacting a venous flow device with the circulating blood of the subject and introducing an apparatus for detecting an analyte in blood into the venous flow device, the apparatus comprising: a sensor that detects the presence of an analyte; andan assembly means having a sensor end and an exterior face, wherein the sensor is affixed to the sensor end of the assembly means, and the assembly means is adapted for coupling with the venous flow device,wherein the sensor contacts the circulating blood of the subject as the blood flows through the venous flow device; and the venous flow device comprises a blood loop having a port adapted to receive the apparatus, wherein the blood loop further comprises a second port adapted to receive an alignment means, and wherein the method further comprises introducing the alignment means into the blood loop from a side opposing the port adapted to receive the apparatus prior to introducing the apparatus. 2. A method of introducing a sensor into the circulating blood of a subject comprising contacting a venous flow device with the circulating blood of the subject and introducing an apparatus for detecting an analyte in blood into the venous flow device, wherein the venous flow device comprises a blood loop, the apparatus comprising: a sensor that detects the presence of an analyte;an assembly means having a sensor end and an exterior face, wherein the sensor is affixed to the sensor end of the assembly means, and the assembly means is adapted for coupling with a venous flow device; andan alignment means;wherein:the sensor contacts the circulating blood of the subject as the blood flows through the venous flow device; and the venous flow device comprises a port adapted to receive the apparatus; andthe method comprises introducing the alignment means into the port prior to or simultaneously with introduction of the apparatus into the port. 3. The method of claim 2, wherein the alignment means is removed with the introduction of the apparatus into the venous flow device. | Medtronic MiniMed, Inc. | A61B005/00 A61M037/00
US07467007 | B2 | 2008-12-16 | Respiratory gated image fusion of computed tomography 3D images and live fluoroscopy images | 2006-05-16 | A system and method is provided directed to improved real-time image guidance by uniquely combining the capabilities of two well-known imaging modalities, anatomical 3D data from computer tomography (CT) and real time data from live fluoroscopy images provided by an angiography/fluoroscopy system to ensure a more efficient and safer intervention in treating punctures, drainages and biopsies in the abdominal and thoracic areas of the human body. | 1. A method for providing real-time image guidance in an image-guided medical procedure, the method comprising: (a) constructing, in a pre-operative stage of said procedure, a fused image comprising a single image of a computer tomography (CT) generated anatomical 4D data set of a patient's region of operative interest with a static fluoroscopic image of said patient's region of operative interest;(b) acquiring, in an operative stage of said procedure, a live fluoroscopic image of said patient's region of operative interest;(c) acquiring, in said operative stage, an operative respiratory gating signal of the patient, substantially coincident with said step (b); and(d) triggering, in said operative stage, the fused image, obtained at said step (a), with the live fluoroscopic image, obtained at said step (b) by the real-time respiratory gating signal, obtained at said step (c) to generate a real-time fused, and triggered image of said patient's region of operative interest;wherein the fused and triggered image provides real-time image guidance during said image-guided medical procedure by displaying a medical device in real time to ensure a safer and more efficient interventionwherein the constructing step (a) further comprisesacquiring a pre-operative respiratory gating signal of the patient's respiratory movement, at substantially a point in time when the monitored respiratory signal reaches a preset phase of respiration;acquiring an anatomical 3D data set of a patient's region of operative interest;constructing an anatomical 4D data set from the acquired anatomical 3D data set;acquiring a static fluoroscopic image of said patient's region of operative interest;selecting a single frame from the anatomical 4D data set as an optimal image: andstatically fusing the selected single frame from the anatomical 4D data set with the acquired static fluoroscopic image. 2. The method of claim 1, wherein the fused image comprises one of a landmark based fused image, surface based fused image, intensity based fused image, and point based fused image. 3. The method of claim 1, wherein the operative respiratory gating signal is acquired at substantially a point in time when the monitored respiratory signal reaches a preset phase of respiration. 4. The method of claim 1, wherein the step of identifying a single frame from the anatomical 4D data set as an optimal image further comprises: simultaneously displaying the anatomical 4D data set and the static fluoroscopic image on a display device; andreviewing each frame of the 4D data set and the static fluoroscopic image to identify said single frame as a frame having the highest degree of correlation with the displayed static fluoroscopic image based on a set of common anatomical landmarks. | Siemens Medical Solutions USA, Inc. | A61B005/05
US07457753 | B2 | 2008-11-25 | Telephone pathology assessment | 2005-06-29 | A system for remote assessment of a user is disclosed. The system comprises application software resident on a server and arranged to interact across a network with a user operating a client device to obtain one or more sample signals of the user's speech. A datastore is arranged to store the user speech samples in association with details of the user. A feature extraction engine is arranged to extract one or more first features from respective speech samples. A comparator is arranged to compare the first features extracted from a speech sample with second features extracted from one or more reference samples and to provide a measure of any differences between the first and second features for assessment of the user. | 1. A system for remote assessment of a user comprising: application software resident on a server and arranged to interact across a network with a user operating a client device to obtain one or more sample signals of the user's speech;a datastore arranged to store said one or more user speech samples in association with details of the user;a feature extraction engine arranged to: extract one or more first features from respective speech samples; extract one or more second features from one or more reference samples, said reference samples comprising a database of speech samples, each sample having a pathology associated therewith, said pathologies comprising: normal, neuromuscular pathologic, physical pathologic and mixed pathologic; and store said second features in association with respective pathologies; anda comparator arranged to compare said first features extracted from a speech sample with second features extracted from said reference samples and to provide a measure of any differences between said first and second features for assessment of said user. 2. A system as claimed in claim 1 wherein said client is a cellular phone, said network includes the Global System for Mobile Communications (GSM) network, wherein said application software comprises an interactive voice recognition (IVR) application and wherein said server includes a voice mark-up language (VML) gateway. 3. A system as claimed in claim 1 wherein said client is a telephone handset, said network includes the public switched telephone network (PSTN), wherein said application software comprises an interactive voice recognition (IVR) application and wherein said server includes a voice mark-up language (VML) gateway. 4. A system as claimed in claim 1 wherein said client is a computing device, said network includes a packet switched network, wherein said application software comprises one or more web pages and wherein said server includes a web server. 5. A system as claimed in claim 1 wherein said first and second features comprise one or more of pitch perturbation, amplitude perturbation and harmonic-to-noise ratio features. 6. A system as claimed in claim 5 wherein said pitch perturbation features include a mean frequency measure for a sample signal. 7. A system for remote assessment of a user comprising: application software resident on a server and arranged to interact across a network with a user operating a client device to obtain one or more sample signals of the user's speech, said sample signals comprising a sustained phonation of the vowel sound /a/;a datastore arranged to store said one or more user speech samples in association with details of the user;a feature extraction engine arranged to: extract one or more first features from respective speech samples; extract one or more second features from one or more reference samples, said reference samples comprising a database of speech samples, each sample having a pathology associated therewith; and store said ... | University College Dublin National University of Ireland | A61B005/08 G06F003/048 G10L011/04 G10L015/00 G10L017/00 G10L021/00
US07844325 | B2 | 2010-11-30 | Bioelectricity impedance measuring device, a malnutrition measurement system, a malnutrition measurement method | 2008-06-04 | A bioelectricity impedance measuring device includes a contact surface capable of being placed on a part of the body, a gripper which is formed to be easily grasped by a single hand, and electrodes on the contact surface. In addition, the bioelectricity impedance measuring device further includes a malnutrition measuring device which measures malnutrition. An operation and control device measures Phase angle theta and/or an RcXc ratio, and measures a nutrient state. The malnutrition measuring device measures the state of malnutrition from a value of Phase angle theta. A display shows the results of measurement with a bar classified by display of a color. In addition, the operation and control device further measures muscular volume. The operation and control device, by being capable of being provided data or regression, can also measure muscular volume for a child aged 5 or younger or an adult aged 60 or older. | 1. A bioelectricity impedance measuring device, which comprises: a housing with a contact surface capable of being placed on a part of the body;a handheld gripper configured to be easily grasped;electrodes facing the same direction from the contact surface of the housing;an operation and control device configured to measure a state of malnutrition from a phase angle theta and/or an RcXc ratio eta by a malnutrition measurement program; anda display which shows the measured state of malnutrition,wherein the operation and control device is configured to store values of Phase angle theta and/or an RcXc ratio eta in a data table as threshold values of Phase angle theta and/or values of RcXc ratio eta to measure a state of malnutrition from Phase angle theta and/or an RcXc ratio eta,wherein the values of Phase angle theta and/or RcXc ratio eta are equivalent to values of an upper arm circumference corresponding to several states of malnutrition determined by the Mid-Upper Arm Circumference (MUAC) method, andwherein the operation and control device thereby is configured to measure a state of malnutrition by comparing values of Phase angle theta and/or RcXc ratio eta calculated from measurement with the threshold values of Phase angle theta and/or RcXc ratio eta stored in the data table. 2. The bioelectricity impedance measuring device according to claim 1, wherein the gripper forms a concave surface, andthe electrodes facing from the contact surface are located at both ends of the concave surface of the gripper. 3. The bioelectricity impedance measuring device according to claim 1, wherein the display shows a color-coded indication of a judged result of the state of malnutrition with a bar classified by color, wherein the display is configured to provide a numerical value of the state of malnutrition along with the color-coded indication of the state of malnutrition, andwherein the display is configured to include a further notification to warn of serious malnutrition and thus a danger of death when an acute state of malnutrition has been measured. 4. The bioelectricity impedance measuring device according to claim 1, wherein the operation and control device is configured to further measure muscular volume by a muscular volume measuring program. 5. The bioelectricity impedance measuring device according to claim 4, wherein the operation and control device is configured to measure muscular volume for a child aged 5 or younger or an adult aged 60 or older, the operation and control device being provided with at least one of data from the data table or regression analysis. 6. The bioelectricity impedance measuring device according to claim 4, wherein the operation and control device is configured to measure a state of malnutrition and/or measure a muscular volume by individual parameters with sex, height and age. 7. The bioelectricity impedance measuring device according to claim 1, wherein the operation and control device is configured to measure th... | Tanita Corporation | A61B005/06
US07751907 | B2 | 2010-07-06 | Expert system for insulin pump therapy | 2007-05-24 | An apparatus comprising a controller. The controller includes an input/output (I/O) module and a rule module. The I/O module is configured to present a question for a patient when communicatively coupled to a user interface and receive patient information in response to the question via the user interface. The rule module is configured to apply a rule to the patient information and generate a suggested insulin pump setting from application of the rule. Other devices, systems, and methods are disclosed. | 1. A device comprising: a controller including: an input/output (I/O) module configured to: present a question for a patient when communicatively coupled to a user interface; andreceive patient information in response to the question via the user interface;a rule module configured to: apply a rule to the patient information; andgenerate a device determined insulin pump setting from application of the rule; anda rule development module configured to receive modifications to the rule via the I/O module. 2. The device of claim 1, including: a memory communicatively coupled to the controller, wherein the memory is to store parameters associated with an insulin pump initial setup, andwherein the rule module is configured to apply the rule to match a subset of the parameters to at least one of received patient health information and received patient lifestyle information to generate a suggested insulin pump initial setup. 3. The device of claim 1, including a communication port communicatively coupled to the controller, and wherein the I/O module is configured to: present the suggested insulin pump setting using the device via the user interface;present a request for confirmation of the insulin pump setting; andreceive a confirmation, andcommunicate the suggested insulin pump setting to an insulin pump via the communication port when a confirmation is received into the device. 4. The device of claim 1, wherein rule module is configured to generate at least one of a basal rate pattern, a correction factor, and a carbohydrate ratio. 5. The device of claim 1, wherein rule module is configured to generate a target blood glucose level. 6. The device of claim 1, including: a memory communicatively coupled to the controller, wherein the memory is to store a database of food options in association with a known amount of nutrient content, andwherein the rule module is configured to generate, using the patient information, a suggested database of food options for storage in an insulin pump. 7. The device of claim 6, wherein the patient information includes at least one of patient health information, and patient lifestyle information. 8. The device of claim 1, wherein the controller includes a rule development module configured to receive an edit to the rule via the I/O module. 9. The device of claim 8, wherein the rule development module is configured to: display a representation of the rule when the I/O module is communicatively coupled to a user interface that includes a display; andconvert a manipulation of the representation via the user interface into the edit to the rule. 10. The device of claim 1, wherein the I/O module is configured to: present an action for the user; andreceive blood glucose information, and wherein the rule module is configured to:apply the rule to the blood glucose data; andgenerate a suggested insulin pump setting from application of the rule. 11. The device of claim 10, wherein the device includes a communication port comm... | Smiths Medical ASD, Inc. | G05B011/01 G05B015/00 G05B019/42 A61M031/00 A61M037/00 A61M001/00
US08332024 | B2 | 2012-12-11 | Low-power analog architecture for brain-machine interfaces | 2008-05-27 | An ultra-low-power circuit for wireless neural recording and stimulation is provided. The circuit includes a neural amplifier with adaptive power biasing for use in multi-electrode arrays and a decoding and/or learning architecture. An impedance-modulation telemetry system provides low-power data telemetry. Also, the circuit includes a wireless link for efficient power transfer, and at least one circuit for wireless stimulation of neurons. | 1. An ultra-low-power microelectronic system for wireless neural recording capable of being completely implanted in the brain, comprising: a plurality of neural amplifiers with adaptive power biasing used in multi-electrode arrays, wherein said neural amplifiers comprise a gain stage implemented using an all-subthreshold and folded cascode architecture to allow low-voltage operation and a bandwidth limiting stage to maintain a constant overall bandwidth and to vary a bias current of said gain stage so as to adjust a noise per unit bandwidth of said gain stage;an adaptive learning architecture for neural signal decoding that uses a continuous-time analog gradient descent algorithm to learn a decoding filter via supervised learning or other learning techniques that is implantable in a brain;an impedance-modulation telemetry system for low-power data telemetry used by a plurality of brain-machine interfaces to transmit RF data from the brain flexibly and efficiently so as to permit power consumption to occur outside a body rather than within the brain;a wireless link for efficient power transfer. 2. The system of claim 1 further comprising at least one circuit for wireless stimulation of neurons. 3. The system of claim 1 further comprising an impedance-modulation wireless link in conjunction with a brain-machine interface. 4. The system of claim 3, wherein the link uses near-field inductive coupling between the implanted system and an external transceiver. 5. The system of claim 4, wherein the link is designed to minimize power consumption in the brain-implanted system or support high data rates from the implanted to the external system. 6. The system of claim 1, wherein the analog decoding and learning architecture includes a data compression stage. 7. An ultra-low-power circuit for wireless neural recording and stimulation comprising: a neural amplifier with adaptive power biasing for use in multi-electrode arrays, wherein said neural amplifier comprises a gain stage implemented using an all-subthreshold and folded cascode architecture to allow low-voltage operation and a bandwidth limiting stage to maintain a constant overall bandwidth and to vary a bias current of said gain stage so as to adjust a noise per unit bandwidth of said gain stage;an adaptive learning architecture for neural signal decoding that uses a continuous-time analog gradient descent algorithm to learn a decoding filter via supervised learning or other learning techniques that is implantable in a brain;an impedance-modulation telemetry system for low-power data telemetry used by a plurality of brain-machine interfaces to transmit RF data from the brain flexibly and efficiently so as to permit power consumption to occur outside a body rather than within the brain;a wireless link for efficient power transfer; andat least one circuit for wireless stimulation of neurons. 8. The ultra-low-power circuit of claim 7, wherein said impedance-modulation telemetry system further com... | Massachusetts Institute of Technology | A61B005/04 H03F003/45
US08202224 | B2 | 2012-06-19 | System and method for calibrating cardiac pressure measurements derived from signals detected by an implantable medical device | 2008-04-24 | Various techniques are provided for calibrating and estimating left atrial pressure (LAP) using an implantable medical device, based on impedance, admittance or conductance parameters measured within a patient. In one example, default conversion factors are exploited for converting the measured parameters to estimates of LAP. The default conversion factors are derived from populations of patients. In another example, a correlation between individual conversion factors is exploited to allow for more efficient calibration. In yet another example, differences in thoracic fluid states are exploited during calibration. In still yet another example, a multiple stage calibration procedure is described, wherein both invasive and noninvasive calibration techniques are exploited. In a still further example, a therapy control procedure is provided, which exploits day time and night time impedance/admittance measurements. | 1. A method for estimating cardiac pressure within a patient using an implantable medical device wherein the estimation exploits conversion factors representative of a correlation between a measured electrical parameter within the patient and the cardiac pressure of the patient, the method comprising: determining a set of conversion factors for converting measured electrical parameters to estimates of cardiac pressure, wherein the conversion factors are determined, at least in part, based on a physiological state of the patient and wherein the physiological state is either an acute state wherein an acute change of fluid volume within the intra-vascular fluid compartment of the patient does not substantially affect the fluid volume within the interstitial fluid compartment of the patient or a chronic state wherein fluid volume between the intra-vascular and interstitial fluid compartments of the patient has substantially equilibrated; andestimating cardiac pressure within the patient by applying a first set of conversion factors to measured electrical parameters when the patient is in the acute state and a second set of conversion factors to measured electrical parameters when the patient is in the chronic state. 2. The method of claim 1 wherein determining the set of conversion factors includes: inputting a first set of conversion factors appropriate for converting electrical parameters measured while the patient is in the acute state;inputting an adjustment factor for use in adjusting the first set of conversion factors to yield a second set of conversion factors appropriate for use with electrical parameters measured while the patient is in the chronic state; andapplying the adjustment factor to the first set of conversion factors to generate the second set of conversion factors. 3. The method of claim 2 wherein the first and second sets of conversion factors include slope and baseline values representative of a linear relationship between the measured parameter and cardiac pressure and wherein the adjustment factor is applied only to the slope value. 4. The method of claim 3 wherein the slope of the first set of conversion factors is adjusted by dividing the slope by adjustment factor in the range of 3 to 5 to yield the slope of the second set of conversion factors. 5. The method of claim 2 wherein adjustment factor is specific to the particular patient in which the device is implanted. 6. The method of claim 1 wherein determining the set of conversion factors includes: determining the current patient physiological state of the patient; andretrieving and applying the appropriate set of conversion factors for the current physiological state. 7. The method of claim 6 wherein the implantable device stores a first set of conversion factors appropriate for use in the acute state and a second set of conversion factors appropriate for use in the chronic state. 8. A method for estimating cardiac pressure within a patient using an implantable me... | Pacesetter, Inc. | A61B005/0205 A61B005/021 A61B005/02
US07965881 | B2 | 2011-06-21 | Methods for obtaining MRI frequency maps | 2005-05-16 | A magnetic resonance imaging (MRI) method includes generating first and second spin populations within each of a plurality of pixels of a slice of an object. These spin populations have different phase histories. For each pixel, MRI signals from those spin populations are acquired at different times These MRI signals can then be used to determine a field strength difference between different pixels of the slice. | 1. A magnetic resonance imaging (MRI) method comprising: performing a pulse sequence including excitation pulses and magnetic gradients designed to excite first and second spin populations within a slice of an object, to have different phase histories;acquiring MRI signals from the first and second spin populations at different times;using the MRI signals to determine a phase difference between the MRI signals from the first and second spin populations;using the determined phase difference to determine a magnetic field strength difference between different portions of the slice; andusing the determined magnetic field strength difference between different portions of the slice, preparing a frequency map indicative of differences in a precessional frequency of the spin populations in the slice. 2. The MRI method of claim 1, wherein the pulse sequence comprises: subjecting the slice to a first radio frequency (RF) pulse;subjecting the slice to a second RF pulse; andat a time between the first and second RF pulses, subjecting the slice to a dephasing gradient. 3. The MRI method of claim 1, wherein the pulse sequence further comprises subjecting the first and second populations to a radio frequency (RF) focusing pulse. 4. The MRI method of claim 2, wherein the steps of subjecting the slice to a first RF pulse and subjecting the slice to a second RF pulse are performed within 1 ms of one another. 5. The MRI method of claim 2, wherein the steps of subjecting the slice to a first RF pulse and subjecting the slice to a second RF pulse are performed within 500 ps of one another. 6. The MRI method of claim 2, wherein the first and second RF pulses have a total tip angle of 90° or less. 7. The MRI method of claim 1, further comprising shimming a magnetic resonance imaging system based on the field strength differences and using the frequency map, wherein the magnetic resonance imaging system was used to generate the first and second spin populations. 8. A magnetic resonance imaging (MRI) method comprising: subjecting an object to first and second radio frequency (RF) excitation pulses spaced apart by less than 1 ms;receiving a set of MRI signals, wherein the set of MRI signals comprises a plurality of subsets of MRI signals, each subset of MRI signals comprising a first echo signal from the first RF excitation pulse and a second echo signal from the second RF excitation pulse, and wherein the MRI signals of a given subset arise from the same portion of the object, and the MRI signals of different subsets of MRI signals arise from different portions of the object; andusing the subsets of MRI signals to determine a field strength difference between different portions of the object. 9. The method of claim 8, further comprising subjecting the object to a dephasing gradient at a time between subjecting the object to the first and second RF excitation pulses. 10. The method of claim 9, further comprising subjecting the object to a rephasing RF pulse after... | The General Hospital Corporation | G06K009/00
US08311611 | B2 | 2012-11-13 | Method for performing multiple registrations in a navigated procedure | 2007-04-24 | A system can be used to navigate or guide an instrument or device into an anatomy of a patient. The navigation can occur with the use of image data acquired of the patient. The image data can be registered to the patient space for navigation. Registration can occur more than once to increase efficiency and speed of the procedure. | 1. A system to register image space to physical space of a patient for a surgical navigation procedure, comprising: a first dynamic reference frame operable to be attached relative to the patient in a first manner;a second dynamic reference frame operable to be attached to the patient in a second manner;a tracked device operable to be used to determine at least one fiducial point on the patient;a processor operable to correlate the fiducial point on the patient to an image fiducial point in image data of the patient; anda tracking system operable to track at least one of the tracked device, the first dynamic reference frame, the second dynamic reference frame, or combinations thereof;wherein the processor registers the image space and physical space with the first dynamic reference frame attached to the patient in the first manner in a first registration with a first measurable accuracy;wherein the processor registers the image space and physical space with the second dynamic reference frame attached to the patient in the second manner in a second registration with a second measurable accuracy;wherein the first and second dynamic reference frames each include a tracking device configured to track real-time movement of the patient in physical space relative to the image space; andwherein the second measurable accuracy is greater than the first measurable accuracy. 2. The system of claim 1, further comprising: a guide device selectively fixed to the patient;wherein the second dynamic reference frame is fixed to the guide device. 3. The system of claim 2, wherein the guide device is secured to a boney structure of the patient in a substantially invasive manner for use during a surgical procedure; wherein the second dynamic reference frame is fixed to a selected portion of the guide device during the procedure. 4. The system of claim 1, wherein the first dynamic reference frame is fixed to a surface of the patient with a surface contact member. 5. The system of claim 1, wherein the second dynamic reference frame is directly or indirectly fixed to a boney structure portion of the patient with a bone fixation member. 6. The system of claim 1, further comprising: at least one fiducial marker operable to be fixed to the patient and imaged with the imaging system;wherein the tracked device is operable to be moved relative to the fiducial marker to determine a position of the fiducial marker. 7. The system of claim 1, further comprising: an imaging system operable to obtain image data of the patient. 8. The system of claim 1, wherein the tracked device includes: a contour determining system operable to be used to determine a contour of the patient in physical space;wherein the contour of the patient in physical space is operable to be used as a physical fiducial point to be correlated with an image fiducial point in the image data with the processor. 9. The system of claim 8 wherein the contour determining system includes a laser emitting device op... | Medtronic, Inc. | A61B005/05
US08190228 | B2 | 2012-05-29 | Doppler velocimetry of retinal vessels and application to retinal vessel oximetry | 2007-12-20 | A new model based on ray tracing is developed to estimate power spectral properties in laser Doppler velocimetry of retinal vessels and to predict the effects of laser beam size and eccentricity as well as absorption of laser light by oxygenated and reduced hemoglobin. There is described the model and show that it correctly converges to the traditional rectangular shape of the Doppler shift power spectrum, given the same assumptions, and that reduced beam size and eccentric alignment cause marked alterations in this shape. The changes in the detected total power of the Doppler-shifted light due to light scattering and absorption by blood can also be quantified with this model and may be used to determine the oxygen saturation in retinal arteries and veins. The potential of this approach is that it uses direct measurements of Doppler signals originating only from moving red blood cells. The invention opens new avenues for retinal vessel oximetry. | 1. A method for determining the oxyhemoglobin content in a blood vessel of a mammalian subject comprising: directing a first incident beam of laser light into an ocular blood vessel having an ocular blood vessel outer wall and an ocular blood vessel inner wall such that said first incident beam of laser light has a radius that is less than the radius of the ocular blood vessel inner wall;measuring the Doppler shift power spectrum of the scattered laser light from said first incident beam with a detector;directing a second incident beam of laser light at a different wavelength than said first incident beam into the same ocular blood vessel having an ocular blood vessel outer wall and an ocular blood vessel inner wall such that said second incident beam of laser light has a radius that is less than the radius of the ocular blood vessel inner wall;measuring the Doppler shift power spectrum of the scattered laser light from said second incident beam with a detector;comparing the total Doppler shift power of said scattered beam of laser light from said first incident beam with the total Doppler shift power of said scattered laser light from said second incident beam to obtain a differential between the total Doppler shift power of said two scattered beams of laser light of different wavelength;choosing the wavelength of the first said laser light to be at an isobestic point and the wavelength of the second said laser light to be at a point where the extinction coefficient of oxy- and deoxyhemoglobin is maximally different orchoosing the wavelength of the first said laser light to be at a point where the extinction coefficient of oxy- and deoxyhemoglobin is maximally different and the wavelength of the second said laser light to be at a point where the extinction coefficient of oxy- and deoxyhemoglobin is maximally different in the opposite direction; and,correlating said differential to a reference value of oxyhemoglobin content, determined from a theoretical model of a blood vessel of a similar size, with that of the measured ocular blood vessel; and,determining the oxyhemoglobin content in a blood vessel based on results of the correlating step. 2. The method according to claim 1 wherein the wavelength of the incident beam of laser light is between 500 nm and 1500 nm. 3. The method according to claim 1 wherein ocular blood vessel is a retinal vein or venule. 4. The method according to claim 1 wherein ocular blood vessel is a retinal artery or arteriole. 5. The method according to claim 1 wherein ocular blood vessel is a retinal capillary. 6. The method according to claim 1 wherein the reference value of oxyhemoglobin content for a blood vessel is determined by other non-invasive or invasive means. 7. A method for determining the flow of oxygenated blood through a blood vessel of a mammalian subject comprising: directing a first incident beam of laser light into an ocular blood vessel having an ocular blood vessel outer wall and an ocular blood... |  | A61B005/1455
US09250227 | B2 | 2016-02-02 | System and method for determining a time when the blood alcohol concentration has passed a threshold level | 2006-11-29 | A system operable to determine a future point of time, at which the alcohol concentration in blood of a user is less than a threshold value. The system includes an alcohol sensor operable to detect a current alcohol concentration in blood of the user, and a portable apparatus. The alcohol sensor and the portable apparatus comprises each a wireless communication apparatus. The wireless communication apparatus in the alcohol sensor is operable to wirelessly transmit the current alcohol concentration to the wireless communication apparatus in the portable apparatus. | 1. A system operable to determine a future point of time when a user can operate a vehicle, at which the alcohol concentration in blood of a user is less than a threshold value, the system comprises an alcohol sensor operable to detect a current alcohol concentration in blood of the user, and a portable apparatus, wherein the alcohol sensor, and the portable apparatus each independently comprise a wireless communication device, wherein a first wireless communication device in the alcohol sensor is operable to wirelessly transmit the current alcohol concentration to a second wireless communication device in the portable apparatus, andwherein the portable apparatus further comprises a controller, a clock, and a memory, wherein the controller is connected to the second wireless communication device, the clock, and the memory, and the controller is operable to store the current alcohol concentration and a time stamp from the clock in the memory, andwherein the controller is connected to a calculation device operable to calculate the future point of time in dependence of the current alcohol concentration detected by the alcohol sensor and of a known alcohol burn rate of the user stored in the memory means, or an alcohol burn rate calculated by the calculation means in dependence of characteristics of the user which have an influence on the alcohol burn rate, wherein said characteristics of the user are selected from the group consisting of sex and weight and combinations thereof, andwherein the alcohol sensor further comprises a device configured to receive exhalation air from the user, and a fuel cell, andwherein the calculation device is operable to be self-calibrated during use through software comprising formulas that are parameterized to the user. 2. A system operable to determine a future point of time when a user can operate a vehicle, at which the alcohol concentration in blood of a user is less than a threshold value according to claim 1, wherein the controller is further connected to an input with the aid of which the user is able to input the characteristics, wherein the characteristics comprise sex and weight. 3. A system operable to determine a future point of time when a user can operate a vehicle, at which the alcohol concentration in blood of a user is less than a threshold value according to claim 1, wherein the first and second wireless communication device independently comprise a radio transmitting and receiving device. 4. A system operable to determine a future point of time when a user can operate a vehicle, at which the alcohol concentration in blood of a user is less than a threshold value according to claim 1, wherein the portable apparatus is a mobile telephone, a mobile computer, a personal digital assistant, a handheld computer, a navigation equipment, or a portable computer. 5. A system operable to determine a future point of time when a user can operate a vehicle, at which the alcohol concentration in blood of a user ... | Alco Systems Sweden AB | G01N033/497
US09883828 | B2 | 2018-02-06 | Test element and test system for examining a body fluid | 2008-04-12 | The invention relates to a test element as a single-use article for examining a body fluid, comprising a lancing member that can be inserted into a body part, a collecting area configured thereon for body fluid obtained by the lancing and at least one light guide for an optical measurement in the collecting area. The collecting area is formed by a collection recess of the lancing member which extends in the lancing direction and the light guide is integrated into the lancing member such that it is secure against displacement and the distal end thereof is arranged in a proximal measurement zone of the collection recess. | 1. A test system for examining a body fluid, comprising: a plurality of test elements for single use that are successively positionable in an active position for use, each test element comprising: a lancing member movable in a lancing direction and having a tip configured for insertion into a body part;a collecting area disposed in the lancing member and configured to collect body fluid obtained by lancing, the collecting area being shaped as an open longitudinal slot extending substantially in the lancing direction;a light guide integrated into the lancing member, the light guide having an end thereof protruding into the collecting area for an optical measurement; anda test field arranged at the end of the light guide, the test field aligned substantially transversely to the longitudinal direction of the slot, wherein the test field is configured to receive body fluid substantially head-on from the slot; anda lancet drive configured to drive the lancing member of each test element when positioned in the active position;wherein the light guide of each one of the plurality of test elements is detachably connectable to an optical transmitter or receiver. 2. The test element of claim 1, wherein the end of the light guide is configured for detecting body fluid. 3. The test element of claim 1, further comprising at least one electrode arranged about the collecting area. 4. The test element of claim 3, wherein the light guide or the at least one electrode is configured to generate a measurement signal for controlling the movement of the lancing member as a function of detected body fluid contact. 5. The test element of claim 1, wherein the light guide is configured to generate a measurement signal for controlling the movement of the lancing member as a function of detected body fluid contact. 6. The test element of claim 1, wherein the test field comprises a reagent layer for an analyte in the body fluid. 7. The test element of claim 1, wherein the light guide comprises two separate light guides integrated into the lancing member to provide a parallel measurement, the ends of the two separate light guides being spaced apart. 8. The test element of claim 7, further comprising a carrier foil that optically couples the distal ends of the light guides to a common test field, the common test field being formed by a reagent layer applied to a side of the carrier foil that faces away from the light guides. 9. The test element of claim 8, wherein the thickness of the carrier foil is about 0.5 to 1.5 times the distance between the midlines of the two separate light guides. 10. The test element of claim 9, wherein the light guides, the carrier foil and the reagent layer have substantially identical refractive indices. 11. The test element of claim 1, wherein a reagent layer of the test field contains reaction-neutral auxiliary substances configured to, when dissolved in the body fluid, adjust the refractive index of the reagent layer to substantially t... | Roche Diabetes Care, Inc. | A61B005/00 A61B005/157 A61B005/145 A61B005/1459 G01N021/77 A61B005/15 A61B005/151
US20080295837 | A1 | 2008-12-04 | METHOD TO LIMIT LEAK COMPENSATION BASED ON A BREATHING CIRCUIT LEAK ALARM | 2007-05-29 | A method and system for ventilating a patient that compensates for leaks occurring within the patient breathing circuit while limiting the volume of breathing gas delivered to the patient. During the operation of a ventilator to supply breathing gases to a patient, the ventilator monitors for leak volumes occurring during the inspiratory phase and expiratory phase of the breathing cycle. Based upon the leak volumes sensed, the volume of breathing gas delivered by the ventilator is increased such that the tidal volume delivered to the patient is the desired tidal volume set by a clinician. The ventilator operates to generate a leak alarm when the leak volume exceeds an alarm threshold. If the leak alarm is generated, the tidal volume delivered to the patient is limited to the tidal volume being delivered prior to generation of the leak alarm. During compensation of the breathing gases delivered to the patient, the system and method determines whether the compensated tidal volume exceeds a maximum tidal volume threshold and limits the compensated tidal volume to the maximum tidal volume threshold. | 1. A method of operating a ventilator for ventilating a patient through a breathing circuit having a patient limb, an inspiratory limb and an expiratory limb, comprising the steps of: operating the ventilator to generate a ventilation delivery level to the patient as set by a user;sensing a parameter associated with the operation of the ventilator;modifying the ventilation delivery level based upon the sensed parameter such that the ventilation delivery level is different than the ventilation delivery level set by the user;comparing the sensed parameter to an alarm threshold;generating an alarm when the sensed parameter exceeds the alarm threshold; andlimiting the ventilation delivery level to the modification present when the sensed parameter exceeds the alarm threshold. 2. The method of claim 1 wherein the ventilation delivery level is a tidal volume of gas delivered by the ventilator during inspiration of the patient. 3. The method of claim 2 wherein the step of sensing a parameter associated with the delivery of gas to the patient includes: sensing an inspired volume of gas delivered to the patient;sensing an expired volume of gas exhaled from the patient; andcalculating an inspiratory leak volume based upon the difference between the inspired volume and the expired volume. 4. The method of claim 3 wherein the tidal volume of gas set by the user is modified to a compensated tidal volume when the difference between the sensed inspired volume and the sensed expired volume exceeds the alarm threshold. 5. The method of claim 4 wherein the compensated tidal volume is limited to the compensated tidal volume being delivered to the patient when the alarm is generated. 6. The method of claim 4 further comprising the steps of: comparing the compensated tidal volume to a maximum tidal volume; andlimiting the compensated tidal volume to the maximum tidal volume. 7. The method of claim 6 wherein the maximum tidal volume is based on the tidal volume set by the user. 8. The method of claim 7 wherein the maximum tidal volume is a percentage of the tidal volume set by the user. 9. The method of claim 1 further comprising the step of manually setting the alarm threshold in the ventilator. 10. A method of operating a ventilator for ventilating a patient through a breathing circuit having a patient limb, an inspiratory limb and an expiratory limb, comprising the steps of: selecting a desired tidal volume of gas to be delivered to the patient;operating the ventilator to generate the desired tidal volume for each inspiratory phase of a breath cycle;sensing an inspired volume of gas delivered to the patient during the inspiratory phase;sensing an expired volume of gas from the patient during the expiratory phase of the breath cycle;comparing the sensed inspired volume and the sensed expired volume;calculating an inspiratory leak volume based upon the difference between the expired tidal volume and the inspired volume;increasing the delivered volume of gas ... | - | A61M016/00 A61B005/08
US20080319280 | A1 | 2008-12-25 | Implantable Biotelemetry Device | 2008-03-27 | An implantable device for in vivo monitoring of biotelemetry data includes: a waterproof housing completely encasing the implantable device, the waterproof housing constructed from a material with chemical and fatigue resistance plus thermal stability for placement in a living being; a radio frequency modem located inside the housing and operable at a low radio frequency not exceeding one megahertz; an antenna located within the housing and operatively coupled with the radio frequency modem; a fully programmable microprocessor located within the housing and operatively coupled with the modem; at least one sensor located within the housing for detecting the biotelemetric data; a memory; and a connector for connecting to a power source to power the programmable microprocessor. | 1. An implantable device for in vivo monitoring of biotelemetry data, the implantable device comprising: a waterproof housing completely encasing the implantable device, the waterproof housing constructed from a material with chemical and fatigue resistance plus thermal stability that is insoluble to all common solvents for placement in a living being;a radio frequency modem operable at a low radio frequency not exceeding one megahertz, said radio frequency modem comprising a full duplex transmitter and receiver;an antenna located operatively coupled with the radio frequency modem;a fully programmable microprocessor operatively coupled with the radio frequency modem, and operable to transmit data received from the at least one sensor;at least one sensor operatively coupled with the microprocessor for detecting the biotelemetric data;a memory located operatively coupled with the programmable microprocessor, said memory comprising identification data for identifying the implantable device; anda connector configured for connecting to a power source to power the programmable microprocessor,wherein the radio frequency modem, the antenna, the microprocessor, the at least one sensor, the memory and the connector are all located within the housing. 2. The implantable device of claim 1 wherein the waterproof housing is approximately one inch by one and one-half inches in size. 3. The implantable device of claim 1 wherein the waterproof housing is ovoid in shape. 4. The implantable device of claim 1 wherein the material for the waterproof housing is an engineered thermoplastic. 5. The implantable device of claim 4 wherein the engineered thermoplastic comprises polyetheretherketones. 6. The implantable device of claim 1 wherein the radio frequency modem is configured to transmit and receive at a low radio frequency less than 450 kilohertz. 7. The implantable device of claim 6 wherein the radio frequency modem is configured to transmit and receive at a low radio frequency less than 312 kilohertz. 8. The implantable device of claim 1 wherein the identification data further comprises identifying data about the living being. 9. The implantable device of claim 1 wherein the living being is a human. 10. The implantable device of claim 1 wherein the at least one sensor is a physiological parameter sensor that detects biotelemetry data selected from a group consisting of: blood pressure, blood flow, body temperature, PH, glucose, respiration rate, blood oxygen saturation level, length of Q-T interval, length of a P-R interval, thoracic impedance changes, nerve activity, and biochemical concentrations such as enzymes and glucose. 11. The implantable device of claim 10 wherein the at least one sensor is a pulse oximetry sensor that provides differential measurements of arterial blood flow. 12. The implantable device of claim 1 wherein the at least one sensor is an activity parameter sensor that detects motion and acceleration. 13. The implantable device o... | Visible Assets Inc. | A61B005/00
US20080249387 | A1 | 2008-10-09 | Enhanced monitor compliance | 2008-04-02 | A monitoring system that enables enhanced compliance with a bio-sign monitor includes a method of transferring data to a consumer device, such as a cell phone or MP3 player and incorporating information related to the bio-sign data into the normal communication channel used by the consumer device. The combination of monitor and consumer device enables bio-sign related information to be made available in convenient and discreet manner that enhances the likelihood of compliance with the counter measures indicated by the bio-sign monitor. | 1. A method of enabling compliance with a bio-sign monitor, comprising: transferring data from said monitor to at least one consumer device;incorporating said data as information in a conventional information format;transmitting said information by means of a conventional communication channel; andmaking said information available to a capable entity to enable compliance with said bio-sign monitor. 2. The method of claim 1, wherein said monitor is a glucose concentration level monitor. 3. The method of claim 1, wherein said data is related to glucose concentration level. 4. The method of claim 1, wherein said consumer device is an iPod. 5. The method of claim 1, wherein said consumer device is a cell phone. 6. The method of claim 1, wherein said consumer device is a computer. 7. The method of claim 1, wherein said consumer device is an MP3 player. 8. The method of claim 1, wherein said conventional information format is audio. 9. The method of claim 1, wherein said conventional information format is text. 10. The method of claim 1, wherein said conventional information format is video. 11. The method of claim 1, wherein said conventional communication channel is an audio channel. 12. The method of claim 1, wherein said conventional communication channel is a wireless channel. 13. The method of claim 1, wherein said conventional communication channel is an Internet channel. 14. The method of claim 1, wherein said capable entity is the diabetic associated with said bio-sign monitor. 15. The method of claim 1, wherein said capable entity is an informed care giver. 16. The method of claim 1, wherein said capable entity is a medical emergency service. 17. A system for enabling compliance with a bio-sign monitor, said system comprising: a module operable to transfer data from said monitor to at least one consumer device;a module operable to incorporate said data as information in a conventional information format; anda module operable to transmit said information by means of a conventional communication channel. 18. An apparatus for enabling compliance with a bio-sign monitor, said system comprising: means for transferring data from said monitor to at least one consumer device;means for incorporating said data as information in a conventional information format;means for transmitting said information by means of a conventional communication channel; andmeans for making said information available to a capable entity to enable compliance with said bio-sign monitor. 19. The apparatus of claim 18, wherein said monitor is a glucose concentration level monitor. 20. The apparatus of claim 18, wherein said data is related to glucose concentration level. 21. The apparatus of claim 18, wherein said consumer device is an iPod. 22. The apparatus of claim 18, wherein said consumer device is a cell phone. 23. The apparatus of claim 18, wherein said consumer device is a computer. 24. The apparatus of claim 18, wherein said consumer device is an... | - | A61B005/00
US20080262371 | A1 | 2008-10-23 | Method for Adaptive Complex Wavelet Based Filtering of Eeg Signals | 2005-09-16 | A method for adaptive filtering of EEG signals in the wavelet domain using a nearly shift-invariant complex wavelet transform. EEG signal data is segmented into a set of K “trials” or “light averages” of M-frames of data each. These trials are overlapped by a number of frames P, where P | 1. A method for adaptive filtering of EEG signal data to extract at least one evoked potential response, comprising: segmenting the EEG signal data into a plurality of sets, each set including a plurality of frames of data;overlapping each of said plurality of sets by a predetermined number of data frames;computing a complex wavelet transform for each of said sets to identify associated normalized wavelet coefficients;computing a phase variance of each associated normalized wavelet coefficient;selectively scaling a magnitude of each associated normalized wavelet coefficient; andreconstructing the at least one evoked potential response from said selectively scaled wavelet coefficients. 2. The method of claim 1 where said at least one evoked potential response is an auditory brainstem response. 3. The method of claim 1 where said step of selectively scaling a magnitude of each associated normalized wavelet coefficient is responsive to said phase variance. 4. The method of claim 1 wherein said predetermined number of data frames in said overlapping step is less than said plurality of frames of data in each set. 5. The method of claim 1 wherein said phase variance is computed from:Fij=(1K)∑k=1Kwijk-wij2where wikj is the normalized spectral component calculated according to:wij=WijkWijkwhere Wijk is the ith complex wavelet coefficient at wavelet scale j of the kth trial, andwhere wij is the mean normalized component calculated according to:wij=(1K)∑k=1Kwijk. 6. The method of claim 1 wherein said step of computing a complex wavelet transform includes computing a dual-tree complex wavelet transform for each of said sets to identify associated normalized wavelet coefficients. 7. The method of claim 1 wherein said step of scaling said magnitude of each associated normalized wavelet coefficient wi,j includes computing: wij=αi,j·Ai,jejθi,j where Ai,j and θi,j are respectively the magnitude and phase of the unprocessed complex ith wavelet coefficient at the jth scale; andwhere:αi,j=exp(-0.75·(FijTmax)4)where Fij is the phase variance of coefficient wi,j across said sets, and the parameter Tmax is a decreasing function. 8. The method of claim 1 wherein said step of scaling said magnitude of each associated normalized wavelet coefficient wi,j includes computing: wij=αi,j·Ai,jejθi,j where Ai,j and θi,j are respectively the magnitude and phase of the unprocessed complex ith wavelet coefficient at the jth scale; andwhere: αi,j1 if FijTmax; αi,j=0where Fij is the phase variance of coefficient wi,j across said sets, and the parameter Tmax is a decreasing function. | - | A61B005/0476
US07433727 | B2 | 2008-10-07 | Implantable biosensor | 2004-09-24 | The invention consists of a sensor with multiple indicating (sensing) electrodes covered with a selectively permeable membrane for monitoring fluid concentrations in a biological environment. The indicating electrodes respond to changes in certain analytes, such as glucose, through an enzyme-mediated reaction. The currents generated from the enzyme-mediated reactions are transmitted through radio signals to an external receiver where the information is processed and recorded. Through the use of various biomaterials and biochemicals associated with the sensor, the monitoring accuracy is improved and the overall viability is prolonged. The process of foreign body fibrosis (formation of a scar capsule around the implanted sensor) eventually limits the functional life of the device. We teach methods of delivery of certain biochemicals that can increase the functional life of the sensor by inhibiting the formation of the foreign body capsule or by stimulating the growth of capillaries into the capsule. | 1. A system for monitoring analyte levels in an individual comprising an implantable biosensor that uses an oxidoreductase enzyme and has multiple electrodes, anda processor that receives signals from the electrodes, excludes outlying signals based on a statistical median analysis, and calculates a mean or median from the remaining signals, wherein the statistical median analysis uses a Z-statistic, Median of Absolute Deviations. 2. A system for monitoring analyte levels in an individual comprising an implantable biosensor that uses an oxidoreductase enzyme and has multiple electrodes, anda processor that receives signals from the electrodes, excludes outlying signals based on a statistical median analysis, and calculates a mean or median from the remaining signals, wherein each electrode generates a sensor glucose value (SG), outlier exclusion being determined based on a Z score calculated for each sensor at a given time, the Z score being calculated according to the following formula:Z⁢⁢score=Median⁢⁢SGall⁢⁢units-SGunit⁢⁢XMedian⁢⁢of⁢⁢all⁢⁢deviations*1.483. 3. The system of claim 2, wherein the processor calculates a mean from the remaining SG values after excluding outliers. 4. The system of claim 2, wherein an SG value is excluded as an outlier if it has a Z score greater than 1. 5. A system for monitoring analyte levels in an individual comprising an implantable biosensor that uses an oxidoreductase enzyme and has multiple electrodes, anda processor that receives signals from the electrodes, excludes outlying signals based on a statistical median analysis, and calculates a mean or median from the remaining signals, wherein the statistical median analysis uses a Z-score. 6. The system of claim 5, wherein the Z-score is based on the difference between a given SG value and a median SG value of all units compared to a scaled SG value of all units. 7. A system for monitoring analyte levels in an individual comprising an array of implantable glucose sensors, each sensor using an electrode pair and glucose oxidase to generate real time glucose sensor values (SG), anda processor programmed to exclude outlying SG values based on the difference between a given SG value and the median SG value of all units, wherein the processor is programmed to exclude outlying SG values based on the difference between a given SG value and the median SG value of all units divided by a scaled SG value of all units. 8. The system of claim 7, wherein the scaled SG value of all units is a median of all absolute deviations value. 9. A system for monitoring analyte levels in an individual comprising an array of implantable glucose sensors, each sensor using an electrode pair and glucose oxidase to generate real time glucose sensor values (SG), anda processor programmed to exclude outlying SG values based on a Z-score. 10. The system of claim 9, wherein the Z-score is the difference between a given SG value and a location SG value of all units compared to a scale... | Legacy Good Samaritan Hospital and Medical Center | A61B005/05 A61B005/00
US07439856 | B2 | 2008-10-21 | Health care patient status event processing and reporting | 2006-07-27 | The invention provides a method and system for communicating, processing and responding to information reporting the status of health care patients. In one aspect, the invention provides a procedure for allocating the attention of health care personnel to issues and events that are reported in association with health care patients located within a health care facility. In some embodiments, assignment of health care personnel to health care patients can be configured according to a set of directives. The directives can direct assignment of personnel based upon many criteria including, but not limited to, the identity and/or location of a health care patient, an issue or event associated with the status of the health care patient, the relative availability of other health care personnel to respond to the issue or event, or a combination thereof. | 1. A method for responding to a communication regarding the status of a health care patient, comprising the steps of: receiving a first communication that includes information regarding the status of an associated health care patient;determining an identity of the patient and whether an event of sufficient priority has occurred in association with said patient based upon one or more characteristics of said first communication;determining an identity of each member of a set of one or more first level respondents, said first level respondents having an association with said first communication;determining whether at least one of said first level respondents is available to receive a second communication including information regarding said first communication, and if no first level respondents are available to receive said second communication, performing a set of escalation actions, or elsetransmitting said second communication to a respondent device associated with said at least one of said first level respondents that are available to receive said second communication; anddetermining if at least a first one of said first level respondents transmits a first response communication that includes an indication of an acceptance of responsibility to respond to said first communication within a pre-determined period of time, and if not true, performing said set of escalation actions, and else if true, performing a step of assignment of said respondent to said first communication. 2. The method of claim 1 where if said step of assignment is performed, then performing a step of transmitting a third communication to each of said first level respondents as a notification of said step of assignment. 3. The method of claim 1 where at least one of said initial respondents transmits a second response communication that includes an indication of a refusal to respond to said second communication. 4. The method of claim 1 where if said first one of said first level respondents transmits a third response communication including an indication of a revocation of said acceptance of responsibility to respond, then a fourth communication to notify of said revocation is transmitted to said set of first level respondents, and said steps of claim 1 that are performed in association with said second communication are performed in association with said fourth communication. 5. The method of claim 4 where if no other first level respondent is available to accept responsibility, then performing said escalation actions. 6. The method of claim 1 where said set of escalation actions includes performing a step of raising a mode of escalation from a lowest first escalation level to a higher second escalation level and transmitting said second communication to a set of second escalation level respondents. 7. The method of claim 6 including a plurality of escalation levels and where at least one of said escalation levels includes the step of broadcast transmitting said second ... | Welch Allyn, Inc. | G08B001/08 G08B023/00 A61B005/00
US07433726 | B2 | 2008-10-07 | Soc-Ox | 2005-09-27 | This invention consists of a sock, a sending unit, a receiving unit, a sensor light, and a line that can be connected to a pulse-oximeter cord. A medical professional places the sock over a patient's foot and adjusts its position until the equipment is properly aligned. The cord from the pulse-oximeter can then be attached to the line from the sending unit. The sending unit activates the sensor light that emits radiation at a minimum of two wavelengths. The receiving unit detects the radiation after it passes through the skin and produces an electrical signal in response to the radiation that can be decoded by the pulse-oximeter. | 1. An oxygen and pulse monitoring device comprising: (a) a signal-carrying line protruding from a sock with a plug that can be connected to a pulse-oximeter, wherein the signal-carrying line includes an exterior portion extending through a first side portion of the sock and an interior portion within the sock, wherein the interior portion of the signal-carrying line laterally extends at least partially around a toe-region of the sock in a manner whereby opposing segments of the signal-carrying line extend laterally in an opposing fashion above and below the toe region and a spanning segment of the signal-carrying line extends along a second side portion of the sock generally opposite the first side portion of the sock to interconnect said opposing segments of the signal-carrying line; (b) a sending unit attached to a first one of said opposing segments of the signal-carrying line and located in the toe region of the sock immediately adjacent to the first one of said opposing segments; (c) a receiving unit attached to a second one of said opposing segments of the signal-carrying line and located in the toe region of the sock immediately adjacent to the second one of said opposing segments; and (d) a sensor light located in the toe region of the sock opposite the receiving unit. 2. The oxygen and pulse monitoring device of claim 1 further comprising an elasticized cuff located on the sock. 3. The oxygen and pulse monitoring device of claim 1 wherein the sock is made of a cotton and elastic blend textile. 4. An oxygen and pulse monitoring device comprising: (a) a signal-carrying line protruding from a cotton and elastic blend sock with a plug that can be connected to a pulse-oximeter, wherein the signal-carrying line includes an exterior portion extending through a first side portion of the cotton and elastic blend sock and an interior portion within the cotton and elastic blend sock, wherein the interior portion of the signal-carrying line laterally extends at least partially around a toe-region of the cotton and elastic blend sock in a manner whereby opposing segments of the signal-carrying line extends laterally in an opposing fashion above and below the toe region and a spanning segment of the signal-carrying line extends along a second side portion of the cotton and elastic blend sock generally opposite the first side portion of the cotton and elastic blend sock to interconnect said opposing segments of the signal-carrying line; (b) a sending unit attached to a first one of said opposing segments of the signal-carrying line and located in the toe region of the cotton and elastic blend sock immediately adjacent to the first one of said opposing segments; (c) a receiving unit attached to a second one of said opposing segments of the signal-carrying line and located in the toe region of the cotton and elastic blend sock immediately adjacent to the second one of said opposing segments; and (d) a sensor light located in the toe region of the cott... | - | A61B005/1455
US07458939 | B2 | 2008-12-02 | Procedure for extracting information from a heart sound signal | 2003-10-09 | A procedure for extracting features from phonocardiographic signals without the use of synchronizing information from electrocardiographic signals. The features extracted are the timing and value of first and second heart sounds and various combinations of timing and value of signal components constituting heart murmur. Such combinations are directly related to various heart conditions, which are more easily diagnosable by a medically trained person when assisted by the signal extraction. The features are extracted by a novel combination of energy/time relationships for the heart signal and various novel classification schemes. | 1. A procedure used for extracting murmur information, characterised in that it comprises the following steps: obtaining a digital representation of heart sound for a predetermined number of seconds,identifying the time of occurrence of the first and second heart sounds in each cycle,windowing the parts of heart sounds falling between the first and second heart sounds, and second and first heart sounds, respectivelydecomposition of the signals into a predetermined first number n1 of frequency bands, each band being decomposed into a predetermined second number n2 of time-slicesobtaining a systole (SP) and a diastole (DP) power vector consisting of the sum of n1 powers measured in each of the n2 time slicesfor each combination of a frequency band and a time slice, the power values from the different systoles are compared, and the median value is chosen to be the standard value for a power vectorobtaining a systole (SMF) and a diastole (DMF) mean frequency vector by weighting the power value for each of n1 frequency bands with the mean frequency of the corresponding band, summing the results and dividing the sum by the corresponding element in the respective systole or diastole power vectorwhile using the time of occurrence of the intensity vectors of the various classes for classifying the time distribution of murmurs. 2. A procedure for extracting murmur information according to claim 1, characterised in that a step preceding the step of obtaining systole and diastole murmur intensity vectors SI and DI consists of refining the windowing by setting the values of SP, DP, SMF, and DMF of the first or last elements equal to the second or last-but-one values, respectively, if the values of the first or last elements of the corresponding vectors fulfil predetermined deviation criteria. 3. A procedure according to claim 1, characterised in that further steps are included in the procedure, comprising subjecting the signal to double differentiation before decompositionobtaining a systole (SI) and diastole (DI) murmur intensity vector, respectively, by taking the logarithm of the corresponding SP and DP vectors,classifying the obtained logarithmic vectors into murmur intensity classescomparing the energy distributions to a catalogue of distributions related to known medical conditions. 4. An apparatus for performing the procedure of extracting information from a phonocardiographic signals obtained from a transducer and subjected to signal processing including identification of characteristic signal components, comprising the steps of: detecting first and second heart sounds and placing them correctly on a time axis by performing the steps of:extracting the first and second heart sounds by classification according to energy levels,then performing instantaneous power and lowpass filtering of the original phonocardiographic signal by performing spectral analysis and feature enhancement to obtain the energy content present in areas of a time-frequency repre... | Bang & Olufsen Medicom A/S | A61B005/02
US20080300477 | A1 | 2008-12-04 | SYSTEM AND METHOD FOR CORRECTION OF AUTOMATED IMAGE REGISTRATION | 2007-05-30 | An image guided surgical system and method for correction of automated image registration via user interaction. The system and method comprising at least one imaging apparatus adapted to acquire a first image and a second image of a region of interest of a subject, a registration component adapted to perform a registration of the second image to a dataset of the first image, at least one display for displaying a visualization of the registration of the second image to a dataset of the first image as it is occurring, and a user interface for manipulating the visualization of the registration to correct any misalignments between the first image and the second image in the registration. | 1. A medical navigation system comprising: at least one imaging apparatus adapted to acquire a first image and a second image of a region of interest of a subject;a registration component adapted to perform a registration of the second image to a dataset of the first image;at least one display for displaying a visualization of the registration of the second image to a dataset of the first image as it is occurring; anda user interface for manipulating the visualization of the registration to correct any misalignments between the first image and the second image in the registration. 2. The system of claim 1, wherein the first image is acquired by a first imaging apparatus. 3. The system of claim 1, wherein the first image is acquired prior to a medical procedure and transferred and stored on a storage device of the medical navigation system. 4. The system of claim 3, wherein the second image is acquired by a second imaging apparatus. 5. The system of claim 4, wherein the second image is acquired during the medical procedure. 6. The system of claim 1, wherein the registration component includes a feedback mechanism for user interaction with the registration process. 7. The system of claim 6, wherein the feedback mechanism provides visualization of the second image and the dataset of the first image on the display. 8. The system of claim 1, wherein the acquired first image data is selected from the group consisting of computed tomography data, magnetic resonance data, positron emission tomography data, ultrasound data, and X-ray data and any combinations thereof. 9. The system of claim 1, wherein the acquired second image data is selected from the group consisting of computed tomography data, magnetic resonance data, positron emission tomography data, ultrasound data, and X-ray data and any combinations thereof. 10. The system of claim 1, wherein the at least one display is a touch screen display with graphical inputs. 11. The system of claim 1, wherein the user interface includes standard input tools selected from a group consisting of a mouse, a keyboard, a joystick, a plurality of pushbuttons, and a touch screen display. 12. A method for performing image registration comprising: acquiring a first image and a second image of a region of interest of a patient;performing a registration of the second image to a dataset of the first image;viewing a visualization of the registration on at least one display as the registration is occurring; andmanipulating the visualization of the registration to correct any misalignments between the first image and the second image in the registration using a user interface. 13. The method of claim 12, further comprising the step of terminating the registration and re-starting the registration with a new set of images. 14. The method of claim 12, wherein the first image is acquired prior to a medical procedure and transferred and stored on a storage device of the medical navigation system. 15. The method o... | GENERAL ELECTRIC COMPANY | A61B005/05
US20080255615 | A1 | 2008-10-16 | Treatments for Correcting Spinal Deformities | 2007-03-27 | The present application is directed to methods of treating a spinal deformity. The methods may begin by initially testing the spinal deformity. An implant may be chosen to treat the deformity based on the results of the testing. The implant may then be attached to vertebral members to begin treating the deformity. The implant and vertebral members may be monitored afterwards to determine whether the deformity is being corrected, and that the vertebral members remain healthy. If the monitoring warrants, the implant may be adjusted to better correct the deformity and/or prevent damage to the vertebral members. | 1. A method of treating a spinal deformity comprising: testing the spinal deformity;determining an implant to treat the spinal deformity based on testing results;attaching the implant to vertebral members and applying a corrective force to the spinal deformity;monitoring the implant after attachment to the vertebral members; andadjusting the corrective force applied by the implant based on monitoring results. 2. The method of claim 1, wherein the step of testing the spinal deformity includes performing measurement testing on concave and convex sides of the vertebral members. 3. The method of claim 1, wherein the step of testing the spinal deformity includes performing biochemical testing on concave and convex sides of the vertebral members. 4. The method of claim 1, wherein the step of monitoring the implant after attachment to the vertebral members comprises testing an intervertebral disc. 5. The method of claim 1, further comprising obtaining a baseline of the spinal deformity prior to implanting the implant, and the step of monitoring the implant after attachment to the vertebral members comprises comparing the baseline to a monitored status of the vertebral members. 6. The method of claim 1, wherein the step of attaching the implant to the vertebral members comprises attaching one of a staple system and a tether system. 7. The method of claim 1, wherein the step of monitoring the implant after attachment to the vertebral members includes monitoring convex and concave sides of the vertebral members. 8. The method of claim 1, wherein the step of attaching the implant to the vertebral members comprises percutaneously attaching the implant to the vertebral members. 9. The method of claim 1, wherein the step of monitoring the implant includes performing one of measurement testing and biochemical testing. 10. The method of claim 1, wherein the step of adjusting the implant based on the monitoring results includes decreasing a force applied by the implant to the vertebral members. 11. The method of claim 1, wherein the step of adjusting the implant based on the monitoring results includes releasing a force applied by the implant to the vertebral members. 12. The method of claim 1, further comprising monitoring the vertebral members after the implant is attached and adjusting the implant based on a health of the vertebral members. 13. The method of claim 1, wherein the step of attaching the implant to vertebral members and applying the corrective force to the spinal deformity applies a distractive force to a concave side of the vertebral members. 14. The method of claim 1, wherein the step of attaching the implant to vertebral members comprises attaching the implant to a posterior side of the vertebral members. 15. A method of treating a spinal deformity comprising: determining an implant to treat the spinal deformity based on testing results;determining a baseline measurement of the spinal deformity;attaching the implant to a convex si... | Warsaw Orthopedic, Inc. | A61B017/70 A61B017/58 A61B005/00 A61B019/00
US20080312511 | A1 | 2008-12-18 | ANIMAL HEALTH MONITORING SYSTEM AND METHOD | 2007-06-14 | The invention provides a real-time method and computer-implemented system of monitoring animal health comprising sensing at least one physical characteristic by means of an active physical sensor attached to the animal, and sensing at least one activity of the animal by means of an activity sensor attached to the animal; positioning an active activity signal generator in the environment, such that the activity signal generator is associated with an activity, gathering physical characteristic data and activity data, and wirelessly transmitting all such data to a network receiver/converter, in real-time; converting all such data if necessary, and transmitting all such data over a computer network to one or more users, in real-time. | 1. A system for monitoring at least one physical characteristic and at least one activity of an animal within an environment, said system comprising: (a) at least one active physical sensor capable of directly detecting at least one physical characteristic or at least one environmental characteristic, or both a physical characteristic and an environmental characteristic,(b) at least one active activity sensor for detecting an activity signal generator associated with an activity;(c) a transceiver operatively connected to the at least one physical activity sensor and the at least one activity sensor, for receiving data from either or both the at least one physical sensor and the at least one activity sensor and transmitting the data using a wireless protocol, in real-time;d) a receiver/converter for receiving data received from the transceiver, and for transmitting the data over a computer network, in real-time. 2. The system of claim 1 wherein the at least one active physical sensor detects body temperature of the animal. 3. The system of claim 2 wherein the at least one active physical sensor also detects ambient temperature of the environment. 4. The system of claim 1 wherein the at least one active activity sensor comprises a first activity sensor which detects a first transmitter associated with feeding, and a second activity sensor which detects a second transmitter associated with watering. 5. The system of claim 4 wherein the first and second activity sensors comprise a single coil and circuit responsive to an electromagnetic field created by the activity signal generator. 6. The system of claim 1 wherein the at least one active physical sensor is contained within a first module, and the at least one activity sensor and the transceiver are contained within a second module. 7. The system of claim 1 wherein the at least one active physical sensor and the at least one activity sensor and the transceiver is contained within a unitary module. 8. The system of claim 6 wherein the first module is a tag comprising means to be attached to an appendage of the animal. 9. The system of claim 8 wherein the first module is an ear tag. 10. The system of claim 6 wherein the second module is attached to a neck collar. 11. The system of claim 1 wherein the transceiver is a non-protocol based RF transmitter, a wireless USB transmitter, or a wireless personal area network transmitter. 12. (canceled) 13. The system of claim 11 wherein the converter is a wireless personal area network to Ethernet converter. 14. The system of claim 1 further comprising an accelerometer for detecting movement of the animal, operatively connected to the transceiver. 15. A method for monitoring at least one physical characteristic and at least one activity of an animal within an environment, said method comprising the steps of: a) sensing at least one physical characteristic by means of an active physical sensor attached to the animal, and sensing at least one activit... | ALBERTA RESEARCH COUNCIL INC. | A61B005/01 A61B005/00
US20080281165 | A1 | 2008-11-13 | SYSTEM AND METHOD FOR ACQUIRING AND TRANSFERRING DATA TO A REMOTE SERVER | 2008-04-29 | Systems and methods are provided for transferring data from a data acquisition device to a remote server in real time facilitated by wireless storage device operative to transmit data across a wireless network and to a portable electronic device, such as a cell phone, PDA, laptop, wireless headset or MP3 Player. Such systems and methods may be employed to transform traditional processes that require continuous analysis and monitoring to be integrated into readily available informational systems. An exemplary embodiment of the present invention is designed to streamline and facilitate the monitoring of physiological parameters so that healthcare providers may continuously monitor the physiological parameters of patients in real time. | 1. A system for acquiring and transferring data to a remote server, comprising: a data acquisition device for acquiring and transmitting the data, the data acquisition device further including: an input device for acquiring the data;a computer processor having memory to store and actively maintain the data and being operative to run a data management software program which is stored in the memory;an electrical power source for providing power to the data acquisition device;a wireless transmitter being operative to transmit the data to a wireless network;a wireless storage device including: a wireless transmitter for receiving data from said wireless network;memory for storing the data; anda connection terminal for transferring said data to a portable electronic device;said portable electronic device further including: memory to store said data;an Internet connectivity means; andan input connection terminal for receiving the data from said wireless storage device;a remote server that is configured for receiving the data from the portable electronic device via the Internet and having memory for storing the data. 2. The system for acquiring and transferring data to a remote server of claim 1, wherein said wireless storage device is a wireless micro secure digital card, wherein a wireless micro secure card comprises: a package having the dimensions of a memory card, the package further including: a wireless transmitter for receiving data from said wireless network;a memory device incorporated in the package for storing and maintaining the data;a connection terminal for transferring the data to said portable electronic device; andan internal power source disposed in the package for providing power to the wireless transmitter. 3. The wireless micro secure digital card of claim 2, wherein the memory device is partitioned such that the memory device can actively store the data in a secure storage area. 4. The wireless micro secure digital card of claim 2, wherein the wireless transmitter of the wireless storage device and the wireless transmitter of the data acquisition device is a Bluetooth transmitter and the wireless network is a Bluetooth network. 5. The wireless micro secure digital card of claim 2, wherein the wireless transmitter of the wireless storage device and the wireless transmitter of the data acquisition device is a radio frequency identification tag and the wireless network is a radio frequency identification network. 6. The wireless micro secure digital card of claim 5, wherein the wireless transmitter of the wireless storage device and the wireless transmitter of the data acquisition device is a passive radio frequency identification tag. 7. The wireless secure micro digital card of claim 2, wherein the connection terminal is a standard connection terminal employed by a conventional micro SD card. 8. The system for acquiring and transferring data to a remote server of claim 1, wherein the wireless storage device is a wireless USB... | - | G06F015/16 A61B005/00
US20080292160 | A1 | 2008-11-27 | METHOD FOR ESTIMATING THE PHYSIOLOGICAL PARAMETERS DEFINING THE EDEMA INDUCED UPON INFUSION OF FLUID FROM AN INTRAPARENCHYMALLY PLACED CATHETER | 2008-05-15 | A method for estimating the physiological parameters defining the edema induced upon infusion of fluid from an intraparenchymally placed catheter including; a) acquisition of patient-specific medical data; b) estimation of pertinent tissue microstructure based on the patient-specific medical data and/or generalized information derived or drawn from one or more of the following: experience, literature, modeling, studies, research, analysis; c) acquisition of information about delivery parameters, and/or delivery device geometry, and/or fluid properties, such as: delivery device trajectory, flow rate, pressure, catheter diameter, catheter profile, fluid viscosity, fluid molecular size; and d) computing a field of values of predicted extracellular volume fraction over the tissue region using the information obtained in (b) and (c). According to a further aspect, a method of infusing or planning and/or monitoring an infusion of a contrast agent such that the distribution of such agent can be detected by means of medical data acquisition (e.g. MRI, CT, x-ray, ultrasound, SPECT, PET) and observing and/or measuring the backflow length along the catheter track. | 1. A method for estimating the physiological parameters defining the edema induced upon infusion of fluid from an intraparenchymally placed catheter including; a) acquisition of patient-specific medical data;b) estimation of pertinent tissue microstructure based on the patient-specific medical data and/or generalized information derived or drawn from one or more of experience, literature, modeling, studies, research, and/or analysis;c) acquisition of information about delivery parameters, and/or delivery device geometry, and/or fluid properties; andd) computing a field of values of predicted extracellular volume fraction over the tissue region using the information obtained in (b) and (c). 2. The method of claim 1, wherein the patient-specific medical data includes one or more of MRI, CT, PET, SPECT, and/or x-ray data. 3. The method of claim 1, wherein the tissue microstructure includes information about crossing and intra-voxel directionality of fibers in white matter. 4. The method of claim 1, wherein patient-specific medical data includes one or more of diffusivity of water molecules, capillary permeability, blood flow, and/or blood volume. 5. The method of claim 1, wherein any of the computations include one or more of the following parameters: pore fraction, intracellular volume fraction, extracellular volume fraction, and/or hydraulic conductivity. 6. The method of claim 1, wherein the diffusivity of water molecules within the measurement volume is computed. 7. The method of claim 1, wherein the directionality of extracellular matrix scaffolding with respect to the cellular fiber directions is estimated. 8. The method of claim 7, wherein the tautness or slackness of the extracellular matrix scaffolding is estimated. 9. The method of claim 1, wherein patient-specific medical data and/or generalized information is used to extract regions and/or structures of special relevance to the procedure, including one or more of surfaces, functional areas, nerve fiber tracks, cavities, and/or intracranial structures that influence fluid flow and/or distribution. 10. The method of claim 9 wherein the extraction of regions and/or structures is performed automatically or semi-automatically using models and/or algorithms for surface detection and delineation, atlases, anatomical information, and computerized versions thereof. 11. The method of claim 1, wherein a derivation of reference values refer to literature values and/or properties about the target volume and/or the entire measurement volume of the patient specific medical information. 12. The method of claim 1, wherein a derivation of reference values refer to a model for fluid distribution within tissue. 13. The method of claim 1, wherein a derivation of reference values refer to generalized values derived from a database. 14. The method of claim 1, wherein an extracellular volume is measured by using medical images. 15. The method of claim 14, wherein the extracellular volume is measur... | - | A61B005/00
US20080281221 | A1 | 2008-11-13 | Method and Apparatus for Encouraging Physiological Change Through Physiological Control of Video Devices | 2008-05-10 | The present invention concerns a method and system for encouraging physiologic change through physiological display control of video devices. The system includes interfacing an output from a physiologic acquisition device with a junction box that is capable of controlling the output of the physiologic acquisition device which can be an EEG, and whose output can be analog or digital, visually modulated output of the EEG signal or can be an audio feedback output of the EEG training signal. In one embodiment of the present invention, the PC is no longer required to process the audio/visual changes affected on the video device. The junction box can include means for changing the video feed characteristics such as brightness or darkness of the video image and also volume control of the video feed in response to changes in the signals of brain activity. | 1. An apparatus for encouraging physiologic change through physiological display control of video devices consisting of the following parts in combination: a physiologic acquisition device;a junction box;means for interfacing an output from a physiologic acquisition device with the junction box; andwherein said junction box is capable of controlling the output of the physiologic acquisition device. 2. The apparatus of claim 1 wherein the physiologic acquisition device is an EEG whose output can be analog or digital;visually modulated output of the EEG signal; oran audio feedback output of the EEG training signal. 3. The apparatus of claim 1 wherein multiple signal outputs acquired from the physiologic acquisition device are sent to the junction box;the acquired signals are sent utilizing one or more interfaces;the junction box interfaces with a playback device;the output signals from the physiologic acquisition device can modulate multiple signals of brain activity. 4. The apparatus of claim 3 wherein the interface may be wireless or wired. 5. The apparatus of claim 3 wherein the playback device is a DVD player, Television set, computer, or other personal electronic device. 6. The apparatus of claim 3 wherein the playback device supports either audio, video, or a combined audio and visual information. 7. The apparatus of claim 3 wherein the output signals from the physiologic acquisition device can modulate multiple signals of brain activity, including Sensory Motor Rhythm (SMR);Delta waves;Theta waves;Alpha waves;Beta waves; andGamma wavesthat are sent to the junction box. 8. The apparatus of claim 3 wherein the junction box can includes means for changing the video feed characteristics of a video image, andvolume control of a video feedin response to changes in the signals of brain activity. 9. The apparatus of claim 3 wherein the means for changing the video feed characteristics can be potentiometers. 10. The apparatus of claim 3 wherein standalone video devices are utilized to modulate the video feed characteristics. 11. The apparatus of claim 10 wherein, a standalone DVD player is used in conjunction with a TV. 12. The apparatus of claim 10 wherein, the signal that is being sent to a TV can be directly modulated by a microchip utilized to modulate multiple voltage outputs; and said multiple voltage outputs act as a variable resistor or potentiometer, in a standalone video device. 13. The apparatus of claim 12 wherein the microchip is placed in a standalone project box and can derive power from an internal battery of a standalone video device. 14. The apparatus of claim 1 wherein the junction box is configured to interface between a plurality of standalone video device devices. 15. The apparatus of claim 1 wherein the interface from the junction box modulates the brightness and contrast control of the device's video signal while another output would be used to modulate the device's volume control of the signal based upon real ti... | - | A61B005/0482
US20080236596 | A1 | 2008-10-02 | CUFF SHIELD FOR A PRESSURE CUFF | 2007-04-02 | A cuff shield for a pressure cuff comprises a liquid impervious layer and an absorbent layer. The liquid impervious layer includes an outer surface and an inner surface. The outer surface of the liquid impervious layer has an adhesive portion for removably securing the cuff shield to the pressure cuff. The absorbent layer is secured to the liquid impervious layer and includes an outer surface, an inner surface, and a release strip on the outer surface of the absorbent layer. The release strip on the absorbent layer is configured to serve as a protective cover for an adhesive portion on a second cuff shield. This design of the release strip enables delivering the cuff shields on a roll | 1. A cuff shield for a pressure cuff comprising: a liquid impervious layer having an outer surface and an inner surface, the outer surface of the liquid impervious layer having an adhesive portion for removably securing the cuff shield to the pressure cuff; andan absorbent layer coupled to the liquid impervious layer and having an outer surface, an inner surface, and a release strip on the outer surface of the absorbent layer, wherein the release strip is configured to serve as a protective cover for an adhesive portion on a second cuff shield. 2. The cuff shield of claim 1, wherein the liquid impervious layer is formed from a polymer film. 3. The cuff shield of claim 1, wherein the absorbent layer is formed from a non-woven web. 4. The cuff shield of claim 1, further comprising edge markers on the outer surface of the absorbent layer for indicating proper alignment of the cuff shield with the pressure cuff. 5. The cuff shield of claim 1, further comprising a sensor marker on the outer surface of the absorbent layer for indicating the desired placement of a sensor with respect to the pressure cuff. 6. The cuff shield of claim 1, wherein the adhesive portion extends substantially between a first end and a second end of the liquid impervious layer. 7. The cuff shield of claim 6, wherein the release strip extends substantially between a first end and a second end of the absorbent layer. 8. The cuff shield of claim 1, wherein the release strip includes an outer surface and an inner surface, the inner surface of the release strip having an adhesive material for bonding the inner surface of the release strip to the absorbent layer. 9. The cuff shield of claim 8, further comprising one or more openings extending through the absorbent layer and the liquid impervious layer, wherein the adhesive portion is formed from the adhesive material exposed through the one or more openings in the absorbent and liquid impervious layers. 10. A cuff shield for a pressure cuff comprising: a first layer having an outer surface and an inner surface, the outer surface of the first layer having an adhesive portion for removably securing the cuff shield to the pressure cuff at a desired cuff location; anda second layer secured to the first layer and having an inner surface, an outer surface, and a release strip on the outer surface of the second layer configured to serve as a protective cover for an adhesive portion on a second cuff shield, the outer surface of the second layer comprising at least one position marker for properly positioning the cuff shield at the desired cuff location. 11. The cuff shield of claim 10, wherein the first layer is a liquid impervious layer. 12. The cuff shield of claim 11, wherein the first layer is formed from a polymer film. 13. The cuff shield of claim 10, wherein the second layer is an absorbent layer. 14. The cuff shield of claim 13, wherein the second layer is formed from a non-woven web. 15. The cuff shield of claim 10, wh... | OPTICAL SENSORS INCORPORATED | A61B005/02 A61F005/37
US07450984 | B2 | 2008-11-11 | Local coil arrangement for a magnetic resonance system | 2003-05-15 | A local coil arrangement for a magnetic resonance system has an antenna arrangement that is arranged in a retainer structure. The retainer structure has a waistband and a perineum element. A person to be examined dons the local coil arrangement like trousers. | 1. A local coil arrangement for a magnetic resonance system comprising: a flexible non-rigid retainer structure having a waistband;a flexible antenna arrangement built into said retainer structure, said antenna arrangement including a perineum element having a middle section;said retainer structure with said antenna arrangement built therein being configured to be donned by an examination subject as non-rigid apparel to bring said middle section of said perineum element adjacent the perineum of the subject with said perineum element extending from a front side of the perineum to a back side of the perineum; andsaid antenna arrangement comprising a conductor configuration selected from the group consisting of (a) a conductor forming a half loop comprising said perineum element and extending from said front side to said back side of the perineum, said half loop having two half-loop ends, and a single connecting ring attached to said ends and (b) a ring conductor connected to said perineum element configured to proceed around the subject substantially parallel to the waistband. 2. A local coil arrangement as claimed in claim 1 wherein said antenna arrangement comprises a perineum ring conductor extending from a first point on said waistband to said middle section of said perineum element at one side of said middle section. 3. A local coil arrangement as claimed in claim 2 wherein said antenna arrangement comprises a further ring conductor extending from said second point on said waistband opposite said first point, to an opposite side of said middle section, opposite to said one side of said middle section. 4. A local coil arrangement as claimed in claim 3 wherein said ring conductor and said auxiliary ring conductor form a single conductor loop which crosses itself in said middle section of said perineum element. 5. A local coil arrangement as claimed in claim 1 wherein said waistband has a variable circumference. 6. A local coil arrangement as claimed in claim 5 wherein said waistband is composed of elastic material. 7. A local coil arrangement as claimed in claim 5 wherein said waistband is manually adjustable in circumference. 8. A local coil arrangement as claimed in claim 7 wherein said waistband has a separating location bridged with a bridge element that is releasably connectable to said waistband. 9. A local coil arrangement as claimed in claim 1 wherein said perineum element of said antenna arrangement is enveloped with a liquid-tight material. 10. A local coil arrangement as claimed in claim 1 wherein said antenna arrangement is enveloped with a liquid-tight material. 11. A local coil arrangement as claimed in claim 1 wherein said antenna arrangement comprises a conductor configuration and circuit components connected to said conductor configuration which emit an antenna arrangement output signal, said conductor configuration and said circuit component being configured to acquire, as said output signal, a magnetic resonance sign... | Siemens Aktiengesellschaft | A61B005/055 G01V003/00
US07446528 | B2 | 2008-11-04 | Coil array for magnetic resonance imaging | 2003-08-29 | A radio frequency (RF) coil array is used in resonance imaging and/or analysis of a subject located within a cylindrical space in which a magnetic field is operatively applied in an axial direction (z). The coil array comprises a plurality of coil elements (10, 11, 12, 13) angled relative to each other about the axis of the cylindrical space, each coil element having a pair of main conductors extending generally parallel to the direction of the magnetic field and located on diametrically opposite sides of the cylindrical space, and a pair of connection conductors connected between respective ends of the main conductors. Each coil element has its maximum sensitivity near the centre of the cylindrical space, so that the subject under study is located in a region of maximum sensitivity. | 1. A radio frequency (RF) coil array configured for use in magnetic resonance imaging and/or magnetic resonance analysis of a subject located within a space in which a homogeneous static magnetic field is operatively applied in a first direction, the RF coil array comprising: a plurality of individual coil elements angled relative to each other, with each of the individual coil elements (i) being electrically separate from each of the other individual coil elements, and (ii) having a pair of main conductors extending generally parallel to the first direction and located on opposite sides of the space in a single plane, said single plane passing through the middle of the space; anda connection conductor connected between respective ends of the main conductors, said connection conductor being generally transverse to the first direction. 2. An RF coil array as claimed in claim 1, wherein the homogeneous space is a cylindrical space. 3. An RF coil array as claimed in claim 2 wherein the individual coil elements are equi-angularly spaced about the axis of the cylindrical space, the angle between adjacent coil elements being 360/N, where N is the number of coil elements in the array. 4. An RF coil array as claimed in claim 2, wherein the connection conductor extends around the periphery of the cylindrical space at an axial end thereof in order to thereby permit access to the cylindrical space through that end. 5. An RF coil array as claimed in claim 2, wherein the individual coil elements are arranged in one or more orthogonal pairs with respect to one another. 6. A magnetic resonance imaging apparatus comprising: a magnet configured for applying a homogeneous static magnetic field in a first direction within an imaging space configured for receiving a subject to be imaged;a radio frequency (RF) coil array comprising a plurality of angularly spaced individual coil elements, each individual coil element being electrically separate from each of the other individual coil elements;each individual coil element comprising: a pair of main conductors extending generally parallel to the first direction and located on opposite sides of the imaging space in a single plane, said single plane passing through the middle of the imaging space; anda connection conductor connected between respective ends of the main conductors, said connection conductor being generally transverse to the first direction. 7. A magnetic resonance imaging apparatus as claimed in claim 6, wherein the imaging space is a cylindrical space. 8. A magnetic resonance imaging apparatus as claimed in claim 7, wherein the individual coil elements are equi-angularly spaced about the axis of the cylindrical space, the angle between adjacent individual coil elements being 360/N, where N is the number of individual coil elements in the array. 9. A magnetic resonance imaging apparatus as claimed in claim 7, wherein the connection conductor extends around the periphery of the cylindrical space at a... | The University of Queensland | G01V003/00 A61B005/055
US07438893 | B2 | 2008-10-21 | Polybiotin compounds for magnetic resonance imaging and drug delivery | 2004-10-01 | The invention relates generally to biotin-containing compounds that are useful as imaging agents and drug-delivery agents. Another aspect of the invention relates to the aforementioned compounds chelated to a metal atom. In a preferred embodiment, the metal atom is a gadolinium. Another aspect of the invention relates to a compound comprising three biotin moieties and a pharmaceutical agent covalently bound to a heterocyclic core. In certain embodiments, the pharmaceutical agent is an antibiotic, antiviral, or radionuclide. Another aspect of the present invention relates to a method of treating disease involving administering the compounds of the invention to a mammal. Another aspect of the present invention relates to a method of acquiring a magnetic resonance image using the compounds of the invention. | 1. A compound represented by formula I:wherein R represents independently for each occurrence H or alkyl;Y represents independently for each occurrence —C(O)— or —S(O)—;n represents independently for each occurrence 1, 2, 3, or 4; andX represents independently for each occurrence 2. The compound of claim 1, wherein n is 2. 3. The compound of claim 1, wherein n is 2; and R is hydrogen. 4. The compound of claim 1, wherein n is 2; R is hydrogen; and Y is —C(O)—. 5. A formulation, comprising a compound of formula I and a pharmaceutically acceptable excipient; wherein a compound represented by formula I is:wherein R represents independently for each occurrence H or alkyl;Y represents independently for each occurrence or —C(O)— or —S(O)—;n represents independently for each occurrence 1, 2, 3, or 4; andrepresents independently for each occurrence 6. The formulation of claim 5, wherein n is 2. 7. The formulation of claim 5, wherein n is 2; and R is hydrogen. 8. The formulation of claim 5, wherein n is 2; R is hydrogen; and Y is —C(O)—. | Massachusetts General Hospital | A61B005/055 C07F005/00 C07D235/02
US07553285 | B2 | 2009-06-30 | Diagnosing peripheral vascular disease by monitoring oxygen saturation changes during an accumulation phase | 2008-07-09 | Peripheral vascular disease is diagnosed through measurements of oxygen saturation. In a specific implementation, peripheral vascular disease is diagnosed based on changes in oxygen saturation in tissue. Ischemia is induced, and measurements of changes in oxygen saturation in tissue are made. Based on changes in oxygen saturation during an accumulation phase, a diagnosis is provided of whether a patient has or does not have peripheral vascular disease. | 1. A method of diagnosing peripheral vascular disease comprising: measuring oxygen saturation in tissue of a patient during an accumulation phase prior to ischemia;analyzing the oxygen saturation in the tissue during the accumulation phase; anddiagnosing the patient as having peripheral vascular disease according to changes in the oxygen saturation in the tissue during the accumulation phase. 2. The method of claim 1 wherein the analyzing the oxygen saturation comprises calculating a rate of change of oxygen saturation during the accumulation phase. 3. The method of claim 2 wherein the rate of change is calculated at a midpoint of the accumulation phase. 4. The method of claim 2 wherein the diagnosing the patient comprises determining if the rate of change crosses a threshold. 5. The method of claim 1 wherein the analyzing the oxygen saturation comprises calculating a time for oxygen saturation to reach a specified percentage during the accumulation phase. 6. The method of claim 5 wherein the diagnosing the patient comprises determining if the time for oxygen saturation to reach a specified percentage crosses a threshold. 7. The method of claim 1 wherein the analyzing the oxygen saturation comprises: calculating a rate of change of oxygen saturation during the accumulation phase; andcalculating a time for oxygen saturation to reach a specified percentage during the accumulation phase. 8. The method of claim 7 wherein the diagnosing the patient comprises determining if the rate of change and the time for oxygen saturation during the accumulation phase cross thresholds. 9. A method comprising: measuring oxygen saturation of the tissue of the patient during an accumulation phase prior to ischemia; anddiagnosing the patient as having peripheral vascular disease based on changes in the oxygen saturation of the tissue during an accumulation phase before ischemia. 10. The method of claim 9 comprising: providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured rate of change in the oxygen saturation during an accumulation phase crosses a threshold value. 11. The method of claim 9 comprising providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured time for the oxygen saturation to cross a percentage value during the accumulation phase exceeds a specific time value. 12. The method of claim 9 comprising: providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured time, from a starting time value to an ending time value, for the oxygen saturation to cross a percentage value exceeds a specific time value, wherein the ending time is within the accumulation phase. 13. The method of claim 9 comprising: providing a visual indication of a positive diagnosis of peripheral vascular disease on a display when a measured time, from a starting time value to an ending time value, for the oxygen s... | ViOptix Inc. | A61B005/00 A61B005/02
US07828742 | B2 | 2010-11-09 | Method and system of monitoring respiratory signal by radio | 2007-07-19 | A method and system is provided for monitoring a respiratory signal by radio. The method includes the steps of: converting a change in electric resistance, which is caused by a change in abdominal circumference measured through a rubber waistband that is made of conductive rubber and is mounted on a lower garment of a testee during respiration, into a voltage signal, performing A/D conversion on the voltage signal, and transmitting the converted digital signal to a short distance by radio using a wireless communication protocol for ZigBee; and receiving the respiratory signal transmitted by radio, transmitting it to a computer unit by wire through an RS-232 port that is a serial communication port, and enabling a tester to monitor the respiratory signal through a screen. | 1. A method of monitoring a respiratory signal by radio, the method comprising: attaching a single rubber waistband that is made of a conductive rubber to a lower garment of a testee so that the waistband completely surrounds the testee's body;a respiratory signal measuring and transmitting step of converting a change in electric resistance, which is caused by a change in abdominal circumference measured through the rubber waistband during respiration, into a voltage signal, performing analog/digital (A/D) conversion on the voltage signal, and transmitting the converted digital signal to a short distance by radio using a wireless communication protocol for ZigBee; anda respiratory signal receiving and monitoring step of receiving the respiratory signal transmitted by radio, transmitting it to a computer unit by wire through an RS-232 port that is a serial communication port, and enabling a tester to monitor the respiratory signal through a screen. 2. The method as claimed in claim 1, wherein the rubber waistband is produced by mixing conductive particles into rubber so as to have the electric resistance from tens of Ω to several kΩ. 3. The method as claimed in claim 1, wherein the rubber waistband in the respiratory signal measuring and transmitting step is operated by one of four resistors constituting a Wheatstone bridge circuit, and undergoes the conversion of the electric resistance change thereof into the voltage signal. 4. The method as claimed in claim 1, wherein the voltage signal converted by the electric resistance change of the rubber waistband corresponds to the respiratory signal that is amplified through a differential amplifier circuit and minimizes high-band noise through a low pass filter circuit. 5. The method as claimed in claim 1, wherein, prior to the respiratory signal receiving and monitoring step, a calibration procedure of performing a CO2 inhalation test on each testee once, analyzing correlation between an abdominal respiratory signal and a tidal volume of the testee to calculate a regression line, and substituting the abdominal respiratory signal, measured in the respiratory signal measuring and transmitting step, into the regression line to measure the lung volume. | Chungbuk National University Industry Academic Cooperation Foundation | A61B005/08
US07468035 | B2 | 2008-12-23 | Automated method for discriminating the cardiac beat | 2004-03-11 | An automated method for discriminating the cardiac beat, on the basis of a blood pressure sampled signal, having a starting point Pstart, operating according to a finite state machine for determining at least the diastolic point Pdia, the systolic point Psys, and the dicrotic point Pdic of the pressure signal, the method being apt to iteratively repeat on subsequent sections of the pressure signal; and instruments necessary to perform the automated method and the apparatus performing the same. | 1. Automated method for discriminating the cardiac beat, on the basis of a blood pressure sampled signal, having a starting point Pstart, characterised in that it operates according to a finite state machine, comprising: a first state (1), wherein the method searches for: a pressure absolute minimum value Pmin, by scanning pressure values included within a first time interval not exceeding the interval going from the starting point Pstart up to the point distant from the determined minimum value Pmin by a first time threshold DTMIN_SYS,a pressure absolute maximum value Pmax, by scanning the pressure values included within a second time interval not exceeding the interval going from the starting point Pstart up to the point distant from the determined minimum value Pmin by a second time threshold DTMAX_SYS, anda pressure signal first derivative maximum value Y1max_postdia included within a third time threshold not exceeding the interval going from the starting point Pstart up to the point distant from the determined minimum value Pmin by a period equal to the second time threshold DTMAX_SYS,the method assuming the point Pmin as diastolic point Pdia and the point Pmax as systolic point Psys, and passing to a following second state (2);the second state (2), wherein the method searches for a pressure signal inflection point Pinflection following the systolic point Psys in a fifth time interval not exceeding the interval starting from the systolic point Psys and of duration equal to a third time threshold DTMAX_MINY1_SYS, the method then passing to a following third state (3);the third state (3), wherein the method verifies whether, in a sixth time interval not exceeding the interval starting from the inflection point Pinflection and of duration equal to a fourth time threshold DTMAX_SYS2Y1DIC, the pressure signal presents a hump with downward concavity, so that: if the outcome of the verification is positive, the method searches, in a seventh time interval not exceeding the interval starting from the inflection point Pinflection and of duration equal to the fourth time threshold DTMAX_SYS2Y1DIC, for the first pressure curve relative minimum, and it assumes the latter as dicrotic point Pdic, whereasif the outcome of the verification is negative, the method searches in said seventh time interval the instant wherein the pressure signal second derivative assumes the maximum value Y2max_postinflection, and it assumes the related pressure signal point as dicrotic point Pdic,the method then passing to a following fourth state (4);the fourth state (4), wherein the method searches for a maximum value Y1max_postdic of the pressure signal first derivative in an eighth interval not exceeding the interval starting from the dicrotic point Pdic and of duration equal to a fifth time threshold DPOSTDIC, the method verifying that the maximum value Y1max_postdia determined in the first state (1) is not less than the value Y1max_postdic, so that: if the outcome of the v... | - | A61B005/02
US08403865 | B2 | 2013-03-26 | Prediction and monitoring of clinical episodes | 2007-07-25 | A method is provided for predicting an onset of an asthma attack. The method includes sensing at least one parameter of a subject without contacting or viewing the subject or clothes the subject is wearing, and predicting the onset of the asthma attack at least in part responsively to the sensed parameter. Also provided is a method for predicting an onset of an episode associated with congestive heart failure (CHF), including sensing at least one parameter of a subject without contacting or viewing the subject or clothes the subject is wearing, and predicting the onset of the episode at least in part responsively to the sensed parameter. Other embodiments are also described. | 1. A method for predicting an onset of an asthma attack, comprising: sensing, using a single sensor, at least one parameter of a subject without contacting the subject or clothes the subject is wearing, and without performing viewing; andusing a pattern-analysis module, predicting the onset of the asthma attack at least in part responsively to the sensed parameter. 2. The method according to claim 1, wherein the at least one parameter includes at least one motion-related parameter of the subject, and wherein sensing the at least one parameter comprises sensing the at least one motion-related parameter. 3. The method according to claim 1, wherein predicting the onset comprises predicting the onset before the subject or a caretaker of the subject becomes aware of the onset. 4. The method according to claim 1, comprising determining a measure of coughing of the subject, wherein predicting the onset comprises predicting the onset at least in part responsively to the sensed parameter and the measure of coughing. 5. The method according to claim 1, wherein sensing the at least one parameter comprises sensing the at least one parameter without requiring human compliance. 6. The method according to claim 1, wherein sensing the at least one parameter comprises sensing the at least one parameter while the subject is sleeping. 7. The method according to claim 1, wherein the at least one parameter includes at least one breathing-related parameter of the subject, and wherein sensing the at least one parameter comprises sensing the at least one breathing-related parameter. 8. The method according to claim 7, wherein predicting the onset comprises: determining at least one breathing pattern of the subject responsively to the sensed breathing-related parameter;comparing the breathing pattern with a baseline breathing pattern; andpredicting the onset at least in part responsively to the comparison. 9. The method according to claim 1, wherein sensing the at least one parameter comprises measuring a pressure in, on, or under a reclining surface upon which the subject lies. 10. The method according to claim 9, wherein sensing the at least one parameter comprises measuring the pressure using exactly one sensor placed in, on, or under the reclining surface. 11. A method for predicting an onset of an episode associated with congestive heart failure (CHF), comprising: sensing, using a single sensor, at least one parameter of a subject without contacting the subject or clothes the subject is wearing, and without performing viewing; andusing a pattern-analysis module, predicting the onset of the episode at least in part responsively to the sensed parameter. 12. The method according to claim 11, wherein the at least one parameter includes at least one motion-related parameter of the subject, and wherein sensing the at least one parameter comprises sensing the at least one motion-related parameter. 13. The method according to claim 11, wherein sensing the at lea... | Earlysense Ltd. | A61B005/08
US08142375 | B2 | 2012-03-27 | Automated infant massager | 2008-03-20 | An automated infant massager has a base for supporting the infant and a plurality of first movable massage elements structured to underlie a portion of the infant. At least one motor is employed for effecting reciprocating movement of the first movable massage element. A plurality of second movable massage elements is provided for massaging the arms and legs of the infant. A plurality of third movable elements is provided within a head support for massaging portions of the infant's head, and a plurality of fourth movable massage elements are provided for massaging the infant's shoulders. The infant massager may, in one embodiment, be structured to receive and massage a premature infant, and in another embodiment, may be structured to receive and massage a full-term infant. | 1. An automated infant massager comprising: a base for supporting said infant,a plurality of first movable massage elements structured to underlie at least a portion of said infant,at least one motor for effecting reciprocating movement of said first movable massage elements with respect to said infant,said first movable massage elements being structured to massage at least a portion of said infant's back,a plurality of second movable massage elements being structured to massage at other portions of said infant,said automated infant massager structured to automatically monitor physical and behavioral changes of said infant, andsaid automated infant massager having a microprocessor for receiving information from monitoring sensors of said infant massager regarding said monitored physical and behavioral changes of said infant, processing the same, and emitting signals with respect thereto to said infant massager when it is performing said massage to adjust or shut off said infant massager. 2. The infant massager of claim 1 including said plurality of second movable massage elements being structured to massage at least portions of said infant's arms and legs. 3. The infant massager of claim 2 including said second movable massage elements having one group of said movable massage elements for massaging a first arm and a first leg and another group of said movable massage elements for massaging a second arm and a second leg. 4. The infant massager of claim 1 including said infant massager having an infant's head support. 5. The infant massager of claim 4 including a plurality of third movable massage elements structured to be subjected to translational movement and being secured to said head support and structured to massage at least a portion of said infant's head. 6. The infant massager of claim 1 including said infant massager structured to accommodate a premature infant. 7. The infant massager of claim 1 including said infant massager structured to accommodate a full-term infant. 8. The infant massager of claim 1 including a majority of said first movable massage elements being structured to reciprocate in a direction generally parallel to the other said first movable massage elements. 9. The infant massager of claim 8 including said first movable massage elements reciprocating in a path generally parallel to the spinal column of said infant. 10. The infant massager of claim 1 including said at least one motor energized by at least one battery. 11. The infant massager of claim 1 including said at least one motor energized from an electrical outlet. 12. The infant massager of claim 1 including employing a plurality of motors to reciprocate said first movable massage elements. 13. The infant massager of claim 1 including a plurality of fourth movable massage elements structured to underlie the shoulders of said infant and establish reciprocating movement thereof in a path generally perpendicular to the orientation of the spinal column of... |  | A61B005/00
US08131043 | B2 | 2012-03-06 | Method and apparatus for detecting interventricular dyssynchrony | 2006-09-18 | Several embodiments of a method and apparatus for detecting interventricular dyssynchrony are provided. In one embodiment, the apparatus includes an image segment processor and an image classifier. The image segment processor may segment left and right ventricles in source images of a subject's heart to form left and right ventricle segments in segmented images. Each source image may include cross sections of the right and left ventricles in common spatial relation along the short axis plane and in common temporal relation in reference to a cardiac cycle. The source images and corresponding segmented images may be temporally-spaced in relation to the cardiac cycle. The image classifier may determine first and second cross-sectional areas associated with the left and right ventricle segments for each segmented image, compare the first and second cross-sectional areas, and classify the subject's heart in an interventricular dyssynchronous class or a non-dyssynchronous class. | 1. An apparatus for detecting interventricular dyssynchrony, including: an image segment processor segmenting a left ventricle and a right ventricle in each of a plurality of source images of a heart to form left and right ventricle segments in a corresponding plurality of segmented images, each source image including at least a portion of a cross section of the heart along a short axis plane, each source image including cross sections of the right and left ventricles of the heart in common spatial relation along the short axis plane and in common temporal relation in reference to a cardiac cycle associated with the heart, the plurality of source images and corresponding plurality of segmented images being temporally-spaced in relation to the cardiac cycle; andan image classifier, in operative communication with the image segment processor, determining a first cross-sectional area associated with the left ventricle segment and a second cross-sectional area associated with the right ventricle segment for each segmented image, comparing, at least in part based on use of a linear regression model, the first and second cross-sectional areas for each segmented image, and classifying the heart in an interventricular dyssynchronous class or a non-dyssynchronous class, based at least in part on the comparing, by classifying the heart in the non-dyssynchronous class if the linear regression model attains statistical significance and indicates a linear relation between the first and second cross-sectional areas over the cardiac cycle and by classifying the heart in the interventricular dyssynchronous class if the linear regression model indicates there is not a linear relation between the first and second cross-sectional areas over the cardiac cycle. 2. The apparatus of claim 1, further including: an image sequence processor, in operative communication with the image segment processor, generating the plurality of source images of the heart from image data resulting from detecting characteristics of the heart. 3. The apparatus of claim 2, further including: a medical imaging device, in operative communication with the image sequence processor, generating the image data by detecting the characteristics of the heart. 4. The apparatus of claim 3 wherein the medical imaging device includes at least one of a medical radiography device, a medical tomography device, a medical sonography device, and a nuclear medicine device. 5. The apparatus of claim 4 wherein the medical imaging device includes a magnetic resonance imaging device, an echocardiography device, a multi-detector computed tomography device, a single photon emission computed tomography device, or a positron emission tomography device. 6. The apparatus of claim 1, further including: a network interface, in operative communication with the image segment processor, adapted to receive the plurality of source images from a medical imaging system via a communication network. 7. A method of detecting in... | The Ohio State University | G06K009/00
US07903020 | B2 | 2011-03-08 | System and methods for remote sensing using double-sideband signals | 2006-03-31 | A sensing system is provided that includes a transceiver. The transceiver includes a transmitter chain that transmits a double-sideband signal having first and second frequency components, and a receiving chain that receives the double-sideband signal after it is reflected by a target. The system further includes a baseband circuit for extracting information content from the received double-sideband signal. A separation between a first frequency of the first frequency component and a second frequency of the second frequency component causes a spike in a signal response generated by one sideband of the received double-sideband signal to substantially overlap a null point of a signal response generated by the other sideband of the received double-sideband signal. | 1. A method of remote sensing, the method comprising: transmitting a double-sideband signal comprising a first frequency component and a second frequency component; andreceiving the double-sideband signal after it is reflected by a subject;wherein a separation between a first frequency of the first frequency component and a second frequency of the second frequency components causes a spike in a signal response corresponding to one sideband of the received double-sideband signal to substantially overlap a null point of a signal response corresponding to the other sideband of the received double-sideband signal. 2. The method of claim 1, wherein the spike is an optimum point. 3. The method of claim 1, further comprising determining a separation between the first frequency and the second frequency that causes the spike to substantially overlap the null point. 4. The method of claim 1, further comprising adjusting at least one of the first frequency and the second frequency to thereby effect the separation that causes the spike to substantially overlap the null point. 5. The method of claim 4, wherein adjusting the at least one of the first frequency and the second frequency to thereby effect the separation comprises sampling a plurality of detection sensitivity measurements, each of the plurality of detection sensitivity measurements corresponding to a different separation between the first frequency and the second frequency, and selecting from among the different separations a separation corresponding to a greatest detection sensitivity from the plurality of detection sensitivity measurements. 6. The method of claim 1, further comprising generating the double-sideband signal by multiplying a first signal having a frequency equal to the first frequency and a second signal having a frequency equal to the second frequency. 7. The method of claim 1, wherein the first frequency and the second frequency are within a Ka-band of an electromagnetic spectrum. 8. A sensing system, comprising: a transceiver including a transmitter chain that transmits a double-sideband signal having a first frequency component and a second frequency component, anda receiving chain that receives the double-sideband signal after it is reflected by a target; anda baseband circuit for extracting information content from the received double sideband signal;wherein a separation between a first frequency of the first frequency component and a second frequency of the second frequency component causes a spike in a signal response generated by one sideband of the received double-sideband signal to substantially overlap a null point of a signal response generated by the other sideband of the received double-sideband signal. 9. The system as in claim 8, wherein the spike is an optimum point. 10. The system as in claim 8, wherein the receiving chain is configured according to a two-step indirect conversion receiver architecture. 11. The system as in claim 8, wherein the target is... | University of Florida Research Foundation, Inc. | G01S013/58 A61B005/02
US08099158 | B2 | 2012-01-17 | System and method for selecting end of diastole and end of systole frames | 2007-08-31 | A method and apparatus for automatically selecting an end of diastole image frame and an end of systole image frame for cardiac analysis. In the method and apparatus, a plurality of image frames of a heart having a contrast medium injected into the heart and an electrocardiogram (ECG) curve of the heart are obtained. Candidate image frames are identified in the plurality of image frames that correspond to a first predetermined point on the ECG curve of the heart associated with the end of diastole. An end of diastole image frame is selected from the candidate image frames. An end of systole image frame is then identified and selected from the remaining one of the plurality of image frames. | 1. A method for automatically selecting an end of diastole image frame and an end of systole image frame for cardiac analysis, the method comprising the steps of: obtaining a plurality of image frames of a heart having a contrast medium injected therein;obtaining an electrocardiogram (ECG) curve of the heart while performing the image frames obtaining step;identifying candidate image frames in the plurality of image frames that correspond to a first predetermined point on the ECG curve of the heart between predetermined different ECG heart cycle wave segments, the first predetermined point on the ECG curve being associated with the end of diastole;selecting from the candidate image frames an image frame that shows the greatest amount of the contrast medium injected into the heart, the selected image frame comprising the end of diastole image frame; andidentifying and selecting from the remaining image frames of the plurality of image frames an image frame that corresponds to the selected end of diastole image frame and a second predetermined point on the ECG curve of the heart within at least one predetermined ECG heart cycle wave segment, the second predetermined point on the ECG curve being associated with the end of systole and the identified and selected image frame comprising the end of systole image frame. 2. The method according to claim 1, wherein the ECG curve includes a T wave and a P wave following the T wave, the first predetermined point on the ECG curve comprising about 70% of a distance starting at a falling edge of the T wave and ending at a leading edge of the P wave. 3. The method according to claim 2, wherein the second predetermined point on the ECG curve comprises about 30% of a distance starting at a maximum amplitude of the T wave and ending at a falling edge of the T wave. 4. The method according to claim 1, wherein the ECG curve includes a T wave, the second predetermined point on the ECG curve comprises about 30% of a distance starting at a maximum amplitude of the T wave and ending at a falling edge of the T wave. 5. The method according to claim 1, wherein the selected end of systole image frame follows the selected end of diastole image frame. 6. The method according to claim 1, wherein the plurality of image frames comprise x-ray image frames. 7. The method according to claim 1, wherein the image frames obtaining step is performed by angiography. 8. The method according to claim 1, displaying the selected end of diastole and systole image frames. 9. The method according to claim 1, further comprising the step of transmitting the selected end of diastole and systole image frames over a network. 10. The method according to claim 1, wherein the step of selecting from the candidate image frames the image frame that shows the greatest amount of the contrast medium injected into the heart is performed by generating a histogram of each of the candidate image frames and comparing the histograms to one another. 11. ... | Siemens Medical Solutions USA, Inc. | A61B005/04
US10327701 | B2 | 2019-06-25 | Apparatuses and methods for electrophysiological signal delivery and recording during MRI | 2005-11-21 | Methods, systems and arrangements are provided for obtaining electroencephalograph (“EEG”) EEG signals from a patient e.g., during a concurrent EEG/MRI examination of the patient. The methods, systems and arrangements include a cap made of conductive inks with sensor positions for attaching a plurality of sensors to the patient's head. The sensors can include electrodes as well as motion sensors for improving EEG signal quality and MRI image quality in the presence of motion noise and other artifacts within the MRI environment. The electrodes may be composed of conductive inks, and can be used in high magnetic fields due to a weak interaction with the RF fields generated by the MRI scanner. The exemplary methods, systems and arrangements can achieve lower SAR and lower temperature increase, as compared to conventional electrodes. | 1. A system for transmitting electrical signals to or from the head of a subject, the system comprising: an electrode carrier;a plurality of electrodes configured to transmit electrical signals and connected to the carrier, the carrier configured and arranged for placement on the head of the subject and, upon placement thereon, arranged to position the electrodes for transmitting electrical signals to or from the head;a conductive trace configured to be electrically connected to the subject to receive or deliver the electrical signals; andwherein each electrode includes a conductive lead electrically extending from the conductive trace to communicate the electrical signals, the conductive lead configured to be electrically isolated from the subject and extending to a connector, and the conductive lead constructed of a predetermined amount of a conductive ink forming a conductive ink layer, the predetermined amount of the conductive ink being substantially dispersed through the extent of the conductive lead, the predetermined amount of the conductive ink having a resistivity in a range of 1 kilo-ohm·meter to 100 kilo-ohm·meters, the resistivity of the predetermined amount of the conductive ink selected to reduce exposure from RF pulses used during magnetic resonance imaging. 2. The system of claim 1 further comprising a motion sensor configured to at least detect ballistocardiogram motion. 3. The system of claim 2 wherein the motion sensor includes a piezo-electric sensor. 4. The system of claim 2 wherein the motion sensor is connected to the electrode carrier. 5. The system of claim 1 wherein the conductive trace has a resistivity of at least 25 kilo-ohm·meters. 6. The system of claim 5 wherein the conductive trace and the conductive lead include at least one of carbon, silver, and silver chloride. 7. The system of claim 1 wherein the conductive trace and the conductive lead include a polymer thick film. 8. The system of claim 1 wherein the conductive trace forms a partial circle and the conductive lead is electrically connected to the conductive trace between two endpoints of the partial circle so as to reduce eddy currents induced by an applied RF signal. 9. The system of claim 1 wherein the conductive lead includes a plurality of layers including at least one of a foam layer, a polyester layer, a conductive ink layer, a dielectric layer, and a silicone layer. 10. The system of claim 1 wherein the carrier forms a stretchable cap. 11. The system of claim 10 wherein the stretchable cap includes a polyester substrate and having connected thereto a plurality of microstrips of conductive ink forming the conductive leads. 12. The system of claim 11 wherein the conductive ink is coated with dielectric ink. 13. The system of claim 12 wherein the carrier includes a vibration dampener including at least one of an elastic mesh dressing and a pillow. 14. The system of claim 11 wherein the conductive trace is connected to the polyester substrat... | The General Hospital Corporation | A61B005/0478 A61B005/00 A61B005/11
US20080319292 | A1 | 2008-12-25 | ANALYTE MONITORING DEVICE AND METHODS OF USE | 2008-08-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. A sensor control unit comprising: a housing adapted for placement on skin and adapted to receive a portion of an electrochemical sensor having a plurality of contact pads;a plurality of conductive contacts disposed on the housing and configured for coupling to the plurality of contact pads on the sensor; anda transmitter disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. 2-120. (canceled) 121. A method of calibrating a glucose measuring device in a host, the method comprising: generating a signal from a continuous glucose measuring device indicative of a glucose concentration in a host;calculating a rate of change of the signal;determining a reference value from reference data obtained using a reference glucose measuring device; andcalibrating the signal to obtain a calibrated signal, wherein the step of calibrating comprises using the reference value only when the rate of change of the signal is less than a predetermined threshold. 122. The method of claim 121, wherein the step of determining a reference value comprises determining a reference value using a single point glucose monitor. 123. The method of claim 122, wherein the single point glucose monitor is built-in with at least a portion of the continuous glucose measuring device. 124. The method of claim 121, wherein the reference glucose measuring device is a continuous glucose measuring device. 125. The method of claim 121, wherein the step of calibrating is performed prospectively. 126. The method of claim 121, wherein the step of calibrating is performed during initialization of the continuous glucose measuring device. 127. The method of claim 121, wherein the step of calibrating is performed after initialization of the continuous glucose measuring device. 128. The method of claim 121, wherein the step of calibrating is performed using a single reference value. 129. The method of claim 121, wherein the step of calibrating is performed using more than one reference value. 130. The method of claim 121, wherein the predetermined threshold is a rate of change of glucose concentration of at least about 1.5 mg/dL/min. 131. The method of claim 121, wherein the signal comprises a raw signal, and wherein the step of calculating a rate of change of the signal comprises calculating a rate of change of the raw signal. 132. The method of claim 121, further comprising filtering the signal to obtain a filtered signal, wherein the step of calculating a rate of change of the signal comprises calculating a rate of change of the filtered signal. 133. The method of claim 121, wherein the step of calculating a rate of change of the signal comprises calculating a rate of change of the calibrated signal. 134. The method of claim 121, further comprising a step of predicting a future value indicative of a glucose concentration in the host for at least one of a future time point or a future time period. 135. The method o... | Abbott Diabetes Care, Inc. | A61B005/055
US20080294018 | A1 | 2008-11-27 | PRIVACY MANAGEMENT FOR WELL-BEING MONITORING | 2007-05-22 | A method for monitoring physiological conditions of a particular individual, while providing unobtrusively implemented controls for managing the privacy of the particular individual and their physiological data, including providing a set of personal data regarding the particular individual, the personal data including privacy settings data, wellness parameter data, or semantic data; determining initial values for the wellness parameters for the particular individual; unobtrusively acquiring physiological data of the particular individual from an image captured during an image capture event, wherein the image capture and image data management is in accordance with a privacy setting; analyzing the physiological data using wellness parameters to assess physiological conditions of the particular individual; and providing a report of the physiological conditions of the particular individual, where the reporting access is limited by privacy controls based on the privacy setting. | 1. A method for monitoring physiological conditions of a particular individual, while providing unobtrusively implemented controls for managing the privacy of the particular individual and their physiological data, comprising; providing a set of personal data regarding the particular individual, the personal data including privacy settings data, wellness parameter data, or semantic data;determining initial values for the wellness parameters for the particular individual;unobtrusively acquiring physiological data of the particular individual from an image captured during an image capture event, wherein the image capture and image data management is in accordance with a privacy setting;analyzing the physiological data using wellness parameters to assess physiological conditions of the particular individual; andproviding a report of the physiological conditions of the particular individual, where the reporting access is limited by privacy controls based on the privacy setting. 2. The method according to claim 1, wherein privileges to change the privacy settings reside in an individual. 3. The method according to claim 1, wherein the privacy setting limits access to the personal data. 4. The method according to claim 3 wherein the access limits for the personal data includes processing, altering, viewing, printing, copying, or e-mailing data. 5. The method according to claim 1 wherein the method is applied to a plurality of individuals in a family and wherein the personal data includes family-related data, such as genetic data, health and disease histories, or ethnic origin. 6. The method according to claim 1, wherein during image capture events only information for the particular individual is collected. 7. The method according to claim 1, wherein the privacy settings determine what images are captured during an image capture event, relative to pose and field of view of the particular individual. 8. The method according to claim 1, wherein the privacy settings determine how image-altering effects, such as cropping or blurring are applied to captured images. 9. The method according to claim 1 wherein the privacy settings define portions of the particular individual for which image data should not be retained. 10. The method according to claim 1, wherein the privacy settings are maintained in a privacy settings database. 11. The method according to claim 1, wherein the privacy controls include one or more passwords, or other biometric identification means, that control access to the privacy settings or the personal data. 12. The method according to claim 1 wherein the privacy settings define the privacy management of the personal data relative to the encryption of the data for data storage or data transmission. 13. The method according to claim 1 wherein the unobtrusive monitoring occurs on an ongoing basis, without need for intervention or preparation by the particular individual. 14. The method according to claim 1 wherein the acquisiti... | - | A61B005/00
US20080312557 | A1 | 2008-12-18 | Method and system for measuring urinary flow rate | 2007-06-18 | Disclosed are a method and an apparatus capable of efficiently and precisely measuring a urinary flow rate by using a general purpose pressure sensor with a low price, instead of a conventional expensive weight sensor, even without a conventional funnel. According to the method and the apparatus, water pressure of urine accumulated in a vessel is transferred to an external pressure sensor through a pressure transfer pipe, the received water pressure is converted into electrical signals through the pressure sensor, the electrical signals are amplified through a bridge amplifier, the amplified analog signals into digital signals (water pressure signals), the water pressure signals are converted into volume signals through an interface with an analyzer, and urinary flow rate signals are acquired by differentiating the volume signals, in which the electrical signals are voltage signals corresponding to volume signals.;An apparatus capable of measuring a urinary flow rate. According to the method and the apparatus, water pressure of urine accumulated in a vessel is transferred to an external pressure sensor through a pressure transfer pipe, the received water pressure is converted into electrical signals through the pressure sensor, the electrical signals are amplified through a bridge amplifier, the amplified analog signals into digital signals (water pressure signals), the water pressure signals are converted into volume signals through an interface with an analyzer, and urinary flow rate signals are acquired by differentiating the volume signals, in which the electrical signals are voltage signals corresponding to volume signals. | 1. A method for measuring a urinary flow rate in diagnosis of prostatomegaly, the method comprising the steps of: transferring water pressure of urine accumulated in a vessel to an external pressure sensor through a pressure transfer pipe;converting the received water pressure into electrical signals through the pressure sensor, in which the electrical signals are voltage signals corresponding to volume signals;amplifying the electrical signals through a bridge amplifier;converting the amplified analog signals into digital signals (water pressure signals);converting the water pressure signals into volume signals through an interface with an analyzer; andacquiring urinary flow rate signals by differentiating the volume signals. 2. The method as claimed in claim 1, wherein the pressure transfer pipe has a diameter of 0.1 to 2 mm and the pressure sensor comprises a Wheatstone bridge sensor including four resistors. 3. The method as claimed in claim 1, wherein the interface with the analyzer is carried out through an RS-232C serial communication port. 4. The method as claimed in claim 1, wherein the analyzer comprises a computer or a dedicated analyzer. 5. A system for performing the method as claimed in claim 1, the system comprising: a vessel in which urine is accumulated;a pressure transfer pipe inserted inside the vessel and connected to an external pressure sensor;an amplifier for extracting and amplifying pressure signals;a converter for converting the amplified pressure signals into digital signals; andan analyzer interfaced with the converter. 6. The system as claimed in claim 5, wherein the pressure transfer pipe has a diameter of 0.5 to 2 mm and the pressure sensor comprises a bridge amplification circuit including four resistors. 7. The system as claimed in claim 5, wherein the interface is carried out through an RS-232C serial communication port. 8. The system as claimed in claim 5, wherein the analyzer comprises a computer or a dedicated analyzer. 9. A system for performing the method as claimed in claim 2, the system comprising: a vessel in which urine is accumulated;a pressure transfer pipe inserted inside the vessel and connected to an external pressure sensor;an amplifier for extracting and amplifying pressure signals;a converter for converting the amplified pressure signals into digital signals; andan analyzer interfaced with the converter. 10. A system for performing the method as claimed in claim 3, the system comprising: a vessel in which urine is accumulated;a pressure transfer pipe inserted inside the vessel and connected to an external pressure sensor;an amplifier for extracting and amplifying pressure signals;a converter for converting the amplified pressure signals into digital signals; andan analyzer interfaced with the converter. 11. A system for performing the method as claimed in claim 4, the system comprising: a vessel in which urine is accumulated;a pressure transfer pipe inserted inside the vessel and connec... | - | A61B005/20
US20080262326 | A1 | 2008-10-23 | Signal Processing Method and Apparatus for Processing a Physiologic Signal such as a Photoplethysmography Signal | 2008-04-21 | A signal processing method of processing a physiologic signal, such as a Photoplethysmography Signal having at least some cardiac components and/or respirator components in the physiologic signal, the processing including the steps of: Identifying a potential cardiac and or respiratory components of a physiologic signal wherein the potential cardiac and or respiratory components have a series of peaks and valleys; Calculating a comparison of the durations of a peak to valley sub-component and a valley to peak sub component of the potential cardiac and or respiratory components; and Utilizing the calculated comparison to evaluate the potential cardiac and or respiratory components. | 1. A signal processing method of processing a physiologic signal having at least some cardiac components in the physiologic signal, the processing including the steps of: Identifying a potential cardiac component of a physiologic signal wherein the potential cardiac component has a series of peaks and valleys;Calculating a comparison of the durations of a peak to valley sub-component and a valley to peak sub component of the potential cardiac component; andUtilizing the calculated comparison to evaluate the potential cardiac component. 2. The signal processing method according to claim 1 wherein the signal includes at least some respiratory components. 3. The signal processing method according to claim 2 wherein the signal is a Photoplethysmography Signal. 4. The signal processing method according to claim 3 wherein the calculated comparison is a ratio of the durations of a peak to valley sub-component and a valley to peak sub component of the potential cardiac component. 5. The signal processing method according to claim 4 wherein the evaluation of the potential cardiac component includes determining whether the calculated ratio is above or below a preset threshold. 6. The signal processing method according to claim 4 wherein the evaluation of the potential cardiac component includes flagging the potential cardiac component when the calculated ratio fails to satisfy a preset threshold. 7. The signal processing method according to claim 4 wherein the calculated comparison includes a calculation of at least a portion of the slopes of the sub-components. 8. A signal processing method of processing a physiologic signal having at least some respiratory and some cardiac components in the physiologic signal, the processing including the steps of: Identifying a potential respiratory component of a physiologic signal wherein the potential respiratory component has a series of peaks and valleys;Calculating a comparison of the durations of a peak to valley sub-component and a valley to peak sub component of the potential respiratory component; andUtilizing the calculated comparison to evaluate the potential respiratory component. 9. The signal processing method according to claim 8 wherein the signal is a Photoplethysmography Signal. 10. The signal processing method according to claim 9 wherein the calculated comparison is a ratio of the durations of a peak to valley sub-component and a valley to peak sub component of the potential respiratory component. 11. The signal processing method according to claim 10 wherein the evaluation of the potential respiratory component includes determining whether the calculated ratio is above or below a preset threshold. 12. The signal processing method according to claim 10 wherein the evaluation of the potential respiratory component includes flagging the potential respiratory component when the calculated ratio fails to satisfy a preset threshold. 13. The signal processing method according to claim 10 where... | STARR LIFE SCIENCES CORP. | A61B005/1455
US20080306349 | A1 | 2008-12-11 | Visual Working Memory Assessment Using Abstract Symbols and Associated Methods | 2007-06-08 | A neuropsychological assessment of visual working memory includes a first display having a first set of distinct abstract elements positioned in a first order. A second display includes a second set of distinct abstract symbols positioned in a second order. The second set includes all the abstract elements of the first set plus at least one additional distinct abstract element not present in the first set. The first display is shown to an examinee. The examinee is asked to view the first display and to note the abstract elements in the first set and the first order thereof. The examinee is instructed to view the second display and to indicate, with reference to the second set, a location of each of the abstract elements present in the first set, in the first order. | 1. A neuropsychological assessment of visual working memory comprising: a first display comprising a first set of distinct abstract elements positioned in a first order;a second display comprising a second set of distinct abstract symbols positioned in a second order, the second set including all the abstract elements of the first set plus at least one additional distinct abstract element not present in the first set;means for displaying the first display to an examinee;means for asking the examinee to view the first display and to remember the abstract elements in the first set and the first order thereof; andmeans for instructing the examinee to view the second display and to indicate, with reference to the second set, a location of each of the abstract elements present in the first set, in the first order. 2. The assessment recited in claim 1, wherein the second display further comprises an indicium adjacent each of the abstract elements of the second set. 3. The assessment recited in claim 2, wherein the instructing means comprises means for instructing the examinee to indicate the location of each of the abstract elements present in the first set, in the first order, by indicating the indicia associated with the abstract elements. 4. The assessment recited in claim 1, wherein the displaying means comprises means for displaying the first display for a predetermined period of time. 5. The assessment recited in claim 4, wherein the predetermined time comprises approximately 5 seconds. 6. The assessment recited in claim 1, wherein the means for instructing the examinee to indicate the location of each of the abstract elements present in the first set, in the first order, comprises means for instructing the examinee to point at each of the abstract elements present in the first set, in the first order. 7. The assessment recited in claim 1, further comprising means for tracking the examinee indications and for scoring the examinee indications. 8. A method of assessing visual working memory of an examinee comprising the steps of: displaying to an examinee a first display comprising a first set of distinct abstract elements positioned in a first order;displaying to the examinee a second display comprising a second set of distinct abstract symbols positioned in a second order, the second set including all the abstract elements of the first set plus at least one additional distinct abstract element not present in the first set;asking the examinee to view the first display and to note the abstract elements in the first set and the first order thereof; andinstructing the examinee to view the second display and to indicate, with reference to the second set, a location of each of the abstract elements present in the first set, in the first order. 9. The method recited in claim 8, wherein the second display further comprises an indicium adjacent each of the abstract elements of the second set. 10. The method recited in claim 9, wherein the instruc... | HARCOURT ASSESSMENT, INC. | A61B005/00
US20080306402 | A1 | 2008-12-11 | METHOD AND SYSTEM FOR DETERMINING VITALITY, HEALING AND CONDITION OF TISSUE OR ORGAN FOR SURGERY | 2008-08-19 | A method of organ and tissue vitality assessment for surgery, including subjecting the organ or tissue to bioelectrical impedance analysis;;taking initial and serial measurements of resistance, reactance, capacitance, phase angle, impedance or any value derived therefrom; and;tracking the initial and serial measurements to establish vitality, healing, and condition of the organ or tissue for surgery. | 1. A method of organ and tissue vitality assessment for surgery, comprising the steps of: subjecting the organ or tissue to bioelectrical impedance analysis;taking initial and serial measurements of resistance, reactance, capacitance, phase angle, impedance or any value derived therefrom; andtracking said initial and serial measurements to establish vitality, healing, and condition of said organ or tissue for surgery. 2. The method according to claim 1, including the step of: utilizing a modified bioelectrical impedance analysis for composition analysis to assess the health of cells of said organ and tissue by the measured reactance thereof. 3. A method according to claim 1, wherein upon harvesting, processing, preserving, treating or transporting said organ or tissue from the donor or source, including the steps of: placing signal introduction electrodes on opposite lateral peripheral borders of said organ or tissue;placing signal detection electrodes at superior and inferior borders of said organ or tissue for a first part of an initial measurement;measuring and recording first values of resistance and reactance and calculating the phase angle of said organ or tissue in said initial measurement;then placing said signal introduction electrodes on said superior and said inferior borders of said organ or tissue;placing said signal detection electrodes on said opposite lateral borders of said organ or tissue; measuring and recording second values of said resistance and said reactance and calculating the phase angle of said organ or tissue; andcomparing said first and second values to normal values to assess vitality. 4. A method according to claim 2, wherein upon arrival of said organ or tissue at the location of the recipient including the steps of: placing signal introduction electrodes on opposite lateral peripheral borders of said organ or tissue;placing signal detection electrodes at superior and inferior borders of said organ or tissue for a first part of an initial measurement of said organ or tissue;measuring and recording first values of resistance and reactance and calculating the phase angle of said organ or tissue in said initial measurement;then placing said signal introduction electrodes on said superior and said inferior borders of said organ or tissue;placing said signal detection electrodes on said opposite lateral borders of said organ or tissue;measuring and recording second values of said resistance and said reactance and calculating the phase angle of said organ or tissue; andcomparing said first and second values to normal values to assess vitality of said organ or tissue. 5. A method according to claim 3, wherein prior to implantation of said organ or tissue into the recipient including the following additional steps: again placing said signal introduction electrodes on said opposite lateral peripheral borders of said organ or tissue;again placing said signal detection electrodes at said superior and said inferior borders... | - | A61B005/053
US07458938 | B2 | 2008-12-02 | Peripherally inserted central catheter with continuous central venous oximetry and proximal high flow port | 2004-10-22 | A peripherally inserted central catheter includes three lumens that communicate with its proximal end. A large lumen terminates short of the distal end of the catheter and is used for the infusion of fluids into the venous system. A second lumen terminates at the distal end and is suitable for measuring blood pressure in the central venous system and infusion of fluids into the central venous system. The third lumen houses a pair of optical fibers which form part of a central venous oxygen saturation (SvO2) monitoring system. | 1. A peripherally inserted central venous system catheter which comprises: a sheath having a length and a diameter suitable for extending from a peripheral vein insertion point on a patient to the superior vena cava of the patient, the sheath being constructed of a material suitable for insertion into the venous system of a subject through a peripheral vein and forming two lumens therein which extend from its proximal end at the insertion point to its distal end in the superior vena cava;a connector fastened to the proximal end of one of said two lumens for connecting the lumen to a medical instrument,a third lumen formed in the sheath and extending from its proximal end to a point intermediate the ends of the sheath where it communicates with blood surrounding the sheath through a port formed in the sheath; anda third connector fastened to the proximal end of the third lumen for connecting the third lumen to a fluid infusion device;wherein the third lumen is larger than said two lumens such that fluids may be infused into the venous system of the patient through said port; andwherein the sheath has a length from 35 to 65 centimeters. 2. The catheter of claim 1 in which the medical instrument is a blood pressure monitor which measures blood pressure at the distal end of the sheath. 3. The catheter as recited in claim 1 in which the sheath has a diameter less than 5.5 French. 4. The catheter as recited in claim 1 in which the sheath is formed from a polyurethane base polymer. 5. The catheter as recited in claim 1 in which the port has a valve. 6. The catheter as recited in claim 1 in which the port has a slit valve. 7. The catheter as recited in claim 1 in which the third lumen is not present distal the port. | MAYO Foundation for Medical Education & Research | A61B005/0215 A61M025/00
US20080255461 | A1 | 2008-10-16 | Real-time optical monitoring system and method for thermal therapy treatment | 2008-03-26 | Multiple site information of light intensity is obtained by application of a multiple-fibre probe in a real-time optical monitoring system. The multiple-fibre probe includes a plurality of optical fibres distributed along the length of the probe. Each optical fibre may be is switchable between the mode for transmitting optical signal into the malignant tissue and the mode for collecting the optical signal from the same tissue. Thus the numbers of the probes can be minimized for collecting multiple site light information and the irritation to the tissue is reduced. A method of using such a probe to determine coagulated boundary in thermal or other treatment is also described. | 1. An optical monitoring system for differentiating coagulated tissue from non-coagulated tissue in an organ in tissue treatment, the system comprising: at least one probe comprising a housing, and at least two optical elements housed within the probe and distributed along the length of the probe,each of said optical elements being switchable between emitting optical output into target tissue and collecting at least some of the optical output affected by the targeted tissue. 2. The system of claim 1, further comprising a control for configuring whether each of said elements for transmitting optical signal into the targeted tissue or for collecting the optical signal from the same targeted tissue. 3. The system of claim 2 wherein the control can enable one of the at least two optical elements to emit optical output and another of the at least two optical elements to collect the at least some optical output affected by the targeted tissue. 4. The system of claim 3 wherein the affecting by the targeted tissue is on an optical property selected from the group consisting of optical density, reflectivity, radiance, fluence and light dispersion. 5. The system of claim 1, further comprising a light source; a detector for detecting the optical output affected by the targeted tissue collected by the collecting optical elements; each said optical element having a first end proximal to the targeted tissue and a second end distal from the targeted tissue, said control having a first set of terminals connected to a light source and the detectors and a second sets of terminals connected to the second ends of said optical elements to switch individually at least two optical elements to effect transmittal of the optical signal from the light source to the targeted tissue or to the detector for collecting the at least some optical output affected by the targeted tissue. 6. The system of claim 5, further comprising an analyzer for analyzing and interpreting signals detected by the detector, and an electronic visual display for displaying result of analysis. 7. The system of claim 2, wherein a light source is present in the system and the light source is selected from coherent and incoherent light sources, 8. The system of claim 2, wherein the light source is a continuous single wave light source. 9. The system of claim 2, wherein the light source is an amplitude modulated light source. 10. The system of claim 2, wherein the light source is a pulsed light source. 11. The system of claim 2, wherein the optical elements comprise optical fibres. 12. The system of claim 11, wherein a micro-prism is mounted on the tip of each optical fibre to selectively accept optical signals. 13. The system of claim 11, wherein each of said optical fibres is an angle-cleaved or bevelled fibre. 14. The system of claim 11, wherein said target tissue is prostate tissue. 15. The system of claim 14, wherein each of said optical fibres is so configured that the optical signal e... | - | A61B005/00 A61B018/18
US20080243016 | A1 | 2008-10-02 | Pulmonary Artery Pressure Signals And Methods of Using | 2007-03-28 | Embodiments of the invention are related to methods and systems for using a pulmonary artery pressure signal to detect and/or monitor physiological parameters, physiological status, and aspects of disorders and diseases, amongst other things. In an embodiment, the invention includes a method of measuring a pulmonary function parameter of a patient. In an embodiment, the invention includes a method of detecting a sleeping disorder. In an embodiment, the invention includes a method of tracking sleep characteristics of a patient. Other aspects and embodiments are provided herein. | 1. A method of measuring a pulmonary function parameter of a patient comprising: chronically implanting a pulmonary artery pressure sensor;obtaining a pulmonary artery pressure signal from the pressure sensor; andprocessing the signal to obtain the pulmonary function parameter. 2. The method of claim 1, comprising transmitting the pulmonary artery pressure signal to an implanted device. 3. The method of claim 2, the implanted device comprising a pulse generator. 4. The method of claim 2, comprising wirelessly transmitting the signal. 5. The method of claim 1, the pulmonary function parameter selected from the group consisting of respiratory rate, respiratory excursion, breath interval, inspiration slope, expiration slope, tidal volume, relative tidal volume, minute ventilation, relative minute ventilation, pulmonary vascular resistance, relative pulmonary vascular resistance, forced expiration volume in one second (FEV1), relative FEV1, forced vital capacity (FVC), relative FVC, total lung capacity (TLC) and relative (TLC). 6. The method of claim 1, wherein processing the signal to obtain the pulmonary function parameter includes the step of converting the pulmonary artery pressure signal into a respiration signal. 7. The method of claim 6, further comprising deriving the frequency of the respiration signal. 8. The method of claim 6, further comprising deriving the amplitude of the respiration signal. 9. The method of claim 1, further comprising comparing a real-time value of the pulmonary function parameter with a stored value of the pulmonary function parameter. 10. The method of claim 1, further comprising reporting the pulmonary function parameter to a care provider. 11. A method of detecting a sleeping disorder comprising: measuring a pulmonary artery pressure signal with a pressure sensor; andmonitoring the pulmonary artery pressure signal to identify a breathing pattern indicative of a sleeping disorder. 12. The method of claim 11, wherein the pressure sensor is chronically implanted. 13. The method of claim 11, the pressure sensor disposed in the pulmonary artery of a patient. 14. The method of claim 11, the breathing pattern indicative of a central sleep apnea. 15. The method of claim 11, the breathing pattern indicative of an obstructive sleep apnea. 16. The method of claim 11, the breathing pattern characterized by a plurality of apneas. 17. The method of claim 11, the breathing pattern characterized by five or more obstructed breathing events per hour, the obstructed breathing event selected from the group consisting of an apnea and a hypopnea. 18. The method of claim 11, the breathing pattern characterized by a plurality of hypopneas. 19. The method of claim 11, the breathing pattern characterized by a plurality of cyclical increases and decreases in breathing amplitude. 20. A method of tracking sleep characteristics of a patient comprising: monitoring a pulmonary artery pressure signal for a change to a cardiopulm... | CARDIAC PACEMAKERS, INC. | A61B005/00 A61B005/08
US20080312543 | A1 | 2008-12-18 | MEASUREMENT OF PULMONARY HYPERTENSION FROM WITHIN THE AIRWAYS | 2008-06-18 | This is directed to methods and devices suited for airway based measurements of pressure in a pulmonary artery. A device is advanced into an airway and in the vicinity of the pulmonary artery. Physical properties of the pulmonary artery are observed through the airway wall using one or more minimally invasive modalities. In a variation, a bronchial balloon catheter measures pressure of the pulmonary artery. | 1. A method for observing a pressure in a pulmonary artery, the method comprising: advancing an expandable device within an airway of a patient;pressurizing the expandable device against an airway wall to compress a pulmonary artery adjacent to the airway wall in a sufficient amount to interfere with a flow of blood in the pulmonary artery;observing a pressure of the expandable device; andcorrelating the interference of the flow of blood in the pulmonary artery with the pressure in the expandable device to determine the pressure of the pulmonary artery. 2. The method of claim 1 wherein the expandable device is advanced within the airway and to a location in a vicinity of a ligamentum arteriosum 3. The method of claim 1, where pressuring the expandable device comprises stopping the flow of blood in the pulmonary artery. 4. The method of claim 3, where observing the pressure of the expandable device comprises observing the pressure when blood flow stops. 5. The method of claim 4, further comprising subsequently depressurizing the expandable device to re-establish blood flow, where observing the pressure of the expandable device comprises observing the pressure when blood flow is re-established. 6. A method for observing a pressure in a pulmonary artery, the method comprising: advancing a first device into an airway;measuring a first flow rate of blood flow in the pulmonary artery with the first device;pressurizing an expandable member within an airway to interfere with the blood flow in the pulmonary artery;measuring a second flow rate of blood flow in the pulmonary artery; andcomparing the first flow rate to the second flow rate to determine the pressure in the pulmonary artery. 7. A method for observing a pressure in a pulmonary artery, the method comprising: locating the pulmonary artery along a surface of the airway;advancing a device into an airway of a lung; inserting a pressure measuring device through the airway wall directly into the pulmonary artery;measuring the pressure of the pulmonary artery with the pressure measuring device;removing the pressure measuring device from the pulmonary artery. 8. The method of claim 1, further comprising sealing an opening in the pulmonary artery created by the advancement of the measuring device. 9. The method of claim 8, wherein sealing is carried out with a tamponade. 10. A method for observing a pressure in a pulmonary artery, the method comprising: locating the pulmonary artery along a surface of the airway;placing a motion sensing device against the surface of the airway near a location of the pulmonary artery;measuring a change in a physical parameter of a wall of the pulmonary artery resulting from an increase in pressure in the pulmonary artery through the airway wall;determining a radius of the pulmonary artery; andcalculating an increase in pressure of the pulmonary artery using the amount of movement of the airway wall and the radius of the pulmonary artery. 11. The method of claim ... | BRONCUS TECHNOLOGIES, INC. | A61B005/02
US20080262363 | A1 | 2008-10-23 | MOBILE HEMODYNAMIC AND ELECTROPHYSIOLOGICAL INTERFACE TO PHYSIOLOGICAL MONITORS AND METHOD OF USE | 2007-04-20 | Certain embodiments provide systems and methods for patient monitoring using a mobile workstation and one or more patient monitors. Certain embodiments include selecting a patient monitor via a mobile physiological monitoring workstation; acquiring physiological data from the patient monitor; and displaying the physiological data at the mobile physiological monitoring workstation. The data is transmitted over a network from the patient monitor to the mobile physiological monitoring workstation without a dedicated connection between the mobile physiological monitoring workstation and the patient monitor. Certain embodiments include at least one patient monitor capable of obtaining and transmitting physiological data for a patient; a mobile physiological monitoring workstation receiving physiological data; and a network enabling exchange of physiological data between the at least one patient monitor and the workstation. The workstation may select one or more of the at least one patient monitor from which to receive physiological data for the patient. | 1. A method for monitoring patient physiological data, said method comprising: selecting a patient monitor via a mobile physiological monitoring workstation;acquiring physiological data from said patient monitor, said data transmitted over a network from said patient monitor to said mobile physiological monitoring workstation without a dedicated connection between said mobile physiological monitoring workstation and said patient monitor; anddisplaying said physiological data at said mobile physiological monitoring workstation. 2. The method of claim 1, wherein said mobile physiological monitoring workstation comprises a wireless mobile physiological monitoring workstation. 3. The method of claim 1, further comprising correlating said physiological data from a plurality of patient monitors. 4. The method of claim 1, wherein said physiological data comprises patient vital sign data. 5. The method of claim 1, wherein said physiological data comprises waveform signal data. 6. The method of claim 5, wherein said waveform signal data comprises at least electrocardiogram waveform data. 7. The method of claim 1, wherein said selecting step further comprises selecting among a plurality of patient monitors, wherein said plurality of patient monitors comprise wired and wireless patient monitors. 8. The method of claim 7, wherein a wireless access point is used for wireless communication with a wire-based network. 9. The method of claim 1, wherein said acquiring step further comprises acquiring data from said monitor before, during and after a medical procedure on a patient. 10. A patient monitoring system, said system comprising: a plurality of patient monitors, said plurality of patient monitors communicating via at least one of wired and wireless communication links, each of said plurality of patient monitors capable of obtaining physiological data for a patient;a wireless mobile physiological monitoring workstation receiving physiological data; anda network enabling exchange of physiological data from said plurality of patient monitors and said workstation,wherein said workstation may select one or more of said plurality of patient monitors from which to receive physiological data for the patient. 11. A patient monitoring system, said system comprising: at least one patient monitor capable of obtaining and transmitting physiological data for a patient;a mobile physiological monitoring workstation receiving physiological data; anda network enabling exchange of physiological data between said at least one patient monitor and said workstation,wherein said workstation may select one or more of said at least one patient monitor from which to receive physiological data for the patient. 12. The system of claim 11, wherein said mobile physiological monitoring workstation comprises a wireless mobile physiological monitoring workstation. 13. The system of claim 11, further comprising a plurality of patient monitors, said plurality of patient monitors c... | General Electric Company | A61B005/02 A61B005/00
US20080287760 | A1 | 2008-11-20 | METHOD OF MAKING AN ELECTROCHEMICAL SENSOR | 2007-10-30 | A process for the manufacture of small sensors with reproducible surfaces, including electrochemical sensors. One process includes forming channels in the surface of a substrate and disposing a conductive material in the channels to form an electrode. The conductive material can also be formed on the substrate by other impact and non-impact methods. In a preferred embodiment, the method includes cutting the substrate to form a sensor having a connector portion and a transcutaneous portion, the two portions having edges that define one continuous straight line. | 1. An in vivo glucose sensor, the sensor comprising: a working electrode having an end; anda counter electrode having an end that is staggered with respect to the end of the working electrode; andsensor chemistry comprising a catalyst,wherein the sensor is capable of operation for at least three days. 2. The method of claim 1, wherein each produced sensor is capable of operation for a period of at least one week. | THERASENSE, INC. | A61B005/055
US07450981 | B2 | 2008-11-11 | Apparatus and method for measuring blood component using light trans-reflectance | 2005-08-19 | An apparatus for measuring a blood component using the trans-reflectance of light irradiated to a transflective member disposed behind a test site. The apparatus includes an optical source unit irradiating light to the test site of a test subject; a transflective unit transflecting light transmitted through the test site; a detection unit detecting the transflected light; and an analysis unit analyzing the transflected light detected in the detection unit. A related method includes irradiating the light to the test site; transflecting light through the test site, detecting the transflected light, and analyzing the detected light. Blood component measurement can be easily performed without collecting blood by analyzing the transflected light with a high optical power containing much information about a blood component. | 1. An apparatus for measuring a blood component using light trans-reflectance, the apparatus comprising: an optical source unit to irradiate light to a test site of a test subject;a transflective unit to transflect light transmitted through the test site among the light irradiated from the optical source unit;a detection unit to detect the transflected light; andan analysis unit to analyze the transflected light detected in the detection unit,wherein the test subject is a human body and the test site is an eyelid, andwherein the transflective unit is a contact lens configured to transflect light transmitted through the eyelid. 2. The apparatus of claim 1, wherein the optical source unit includes an optical source that irradiates a 1,600-1,800 nm or 2,100-2350 nm wavelength light beam. 3. The apparatus of claim 1, wherein the optical source unit includes an optical source that irradiates different wavelength light beams, and a spectrometer to select a specific wavelength light beam from the different wavelength light beams. 4. The apparatus of claim 1, wherein the optical source unit irradiates near-infrared light to the test site. 5. The apparatus of claim 1, wherein the transflecting light is infrared light which is in the ranges of 1600-1800 nm or 2100-2350 nm. 6. The apparatus of claim 1, wherein the optical source unit irradiates light reacting with glucose to the test site. 7. The apparatus of claim 1, wherein the optical source unit includes an optical fiber to transmit light from the optical source unit to the test site. 8. The apparatus of claim 1, wherein the transflective unit is a contact lens transflecting light irradiated from the optical source unit and the contact lens is adapted to be attached to the test site. 9. The apparatus of claim 1, wherein the contact lens is adapted to be disposed on the cornea. 10. The apparatus of claim 1, wherein the blood component is selected from the group consisting of glucose, cholesterol, hemoglobin, albumin, triglyceride, alcohol and gamma globulin. 11. The apparatus of claim 1, wherein the transflective unit is configured to be spaced from a cornea. 12. A method of measuring a blood component using light trans-reflectance, the method comprising: irradiating light to a test site;transflecting light transmitted through the test site among the light irradiated to the test site;detecting the transflecting light; andanalyzing the detected light by calculating a blood component level based on the transflected light,wherein the operation of irradiating the light to the test site includes irradiating light reacting with the blood component to an eyelid, andwherein the operation of transflecting the light transmitted through the test site includes transflecting light transmitted through the eyelid. 13. The method of claim 12, wherein the blood component is glucose and the operation of analyzing the light includes calculating a glucose level in the blood based on the transflecting light. | Samsung Electronics Co., Ltd. | A61B005/1455
US07469158 | B2 | 2008-12-23 | Fetal oximetry system and sensor | 2004-10-01 | An optical measuring device having multiple optical paths between one or more light emitters and one or more light detectors and/or providing at least two sets of wavelength of light along at least one path, with a final measurement being produced as a combination of measurements of the sets of wavelengths of light taken along one or more of the optical paths. Features that contribute to increased safety and ease of use include providing (1) a receiving cavity in a proximal end of an insertion rod that holds a free end of a circuit connector to keep it from becoming tangled or snagged, (2) a mechanism to keep the sensor within an introducer tube during storage and insertion and to expose a portion of the sensor only when the sensor is applied to the unborn baby, (3) a tab on the insertion rod to prevent the circuit connector from becoming tangled or snagged within the introducer tube, (4) a rotating feature whereby if a torque applied on the sensor exceeds a first predetermined amount, the sensor rotates, and a disengaging feature whereby the sensor detaches from insertion rod if a pull-off force exceeds a second predetermined amount, the rotating and disengaging features being independent of one another, (5) a circuit connector that includes at least one of the following features: (a) a stiffening member provided at the proximal end to minimize bending, (b) a shielding layer, and (c) at least one slit to increase the flexibility of the circuit connector, and (6) an interface that includes an identification element that is detected by an external circuit only if the circuit connector is connected to the interface. The present invention also pertains to a method of manufacturing a needle that is used in an invasive sensor, and preferably for fetal monitoring, that provides features not heretofore available in conventional sensors. | 1. A physiological condition measuring device comprising: an insertion rod;a sensor selectively attached to a distal end of said insertion rod;a circuit connector having a distal end coupled to said sensor and a proximal end that selectively couples said sensor to an external circuit;an introducer tube selectively positioned relative to said insertion rod so as to house at least a portion of said insertion rod therein with said circuit connector extending along at least a portion of a length of said insertion rod between said insertion rod and said introducer tube, said introducer tube being moveable in an axial direction relative to said insertion rod;a mechanism that selectively couples said insertion rod and said introducer tube to maintain said introducer tube in a first position relative to said insertion rod wherein said sensor is located entirely within said introducer tube and that permits said introducer tube to be moved to a second position relative to said insertion rod wherein at least a portion of said sensor is located outside said introducer tube;wherein said mechanism is a deflectable tab disposed on said insertion rod, said tab being deflectable from a non-deflected position to a deflected position, in said deflected position, said tab provides a configuration that permits a first portion of said tab to insert between a portion of said introducer tube and said insertion rod, with said tab being biased into an engaged relation with said introducer tube to maintain said introducer tube in said first position, and wherein removing said tab from within said introducer tube causes said tab to move to its non-deflected position out of a path of movement of said introducer tube relative to said insertion rod to permit said introducer tube to be moved to said second position; andfurther comprising a first tab provided on a side of said insertion rod opposite a side on which said deflectable tab is located, said first tab being sized and configured to provide a force on said introducer tube opposite a force provided thereon by said deflectable tab responsive to said deflectable tab being biased into said engaged relation with said introducer tube. 2. A physiological condition measuring device comprising: an insertion rod;a sensor selectively attached to a distal end of said insertion rod, said sensor and an associated portion of said insertion rod being sized and configured such that said sensor rotates relative to said insertion rod if a torque applied on said sensor by said insertion rod exceeds a first predetermined amount and said sensor disconnects from said distal end of said insertion rod if a pull-off force exerted on said sensor by said insertion rod exceeds a second predetermined amount, wherein said first predetermined amount of torque necessary to cause rotation is independent of said second predetermined amount of force necessary to cause said sensor to disconnect from said insertion rod;a circuit connector having a distal e... | RIC Investments, LLC | A61B005/1455
US07440904 | B2 | 2008-10-21 | Method and system for generating personal/individual health records | 2006-07-28 | A system and method for generating and/or updating a personal/individual health record. Inputs of data to the system may come from diverse sources including, but not limited to, patient questionnaires, insurance company (or other payor) claims data, hospitals, clinics and other institutional providers, and individual physicians and physicians' offices. | 1. A method for generating a personal/individual health record of a patient enrolled in a health plan of a payor, said method comprising a computer performing the steps: (a) enrolling a patient in a health care plan of a payor, wherein the patient undergoes encounters with a plurality of health care providers regarding a health related issue of the patient, wherein the health care providers request payment for at least a portion of the cost for the encounters from the payor, wherein the health care providers submit health related information of the patient to the payor to support the cost of the encounters with the patient;(b) initially establishing a communication link with between a personal health record (“PHR) system and the payor, wherein there is no communication between PHR system and the health care providers and between the PHR system and the patient;(c) receiving to the PHR system server a payor data element indicative of a health related parameter for the patient from the payor via the communication link, wherein said payor data element includes payor claims data;(d) extracting at the PHR system server encounter data comprising at least one of outpatient encounter history, hospital outpatient encounter history, and hospital admissions data from said payor data element;(e) determining at the PHR system server at least one universal health care concept code corresponding to the encounter data;(f) creating at the PHR system server a personal/individual health record of said patient by storing one or more entries into said personal/individual health record based on said encounter data that was extracted from said payor data element and associating the universal health care concept code with each entry;(g) establishing from the PHR system server a communication link with at least one of the health care providers and the patient;(h) supplementing to the PHR system server said personal/individual health record with supplemental data received from the health care providers and the patient;(i) associating at the PHR system server an access list with each user capable of accessing the personal/individual health record, wherein said access list categorizes said personal/individual health record into a restricted set of entries and an accessible set of entries based on the universal health care concept code associated with each entry in the personal/individual health record;(j) denying at the PHR system server access to a user requesting an entry in the personal/individual health record if the requested entry is in the restricted set of entries;(k) allowing at the PHR system server access to a user requesting an entry in the personal/individual health record if the requested entry is in the accessible set of entries; andwherein steps (b), (c), (d), (e) and (f) occur prior to steps (g) and (h). 2. The method of claim 1, wherein said payor claims data includes at least one health insurance claim code. 3. The method of claim 1, wherein said payor d... | Malik M. Hanson | A61B005/00
US07746070 | B2 | 2010-06-29 | Method for generating high resolution MRI image and recording medium thereof | 2007-07-24 | The present invention relates to a high-resolution magnetic resonance imaging (MRI) image generating method using a new MRI image method called a generalized series parallel imaging technique, and a recording medium thereof. In the high-resolution MRI image generating method using the generalized series parallel imaging technique according to the present invention, selecting only a predetermined low frequency band of all frequency bands, and sampling the predetermined low frequency band of all frequency bands at a lower rate than a Nyquist rate to acquire magnetic resonance data. | 1. A method for generating a high-resolution magnetic resonance imaging (MRI) image on a display device by using a generalized series parallel imaging method, the method comprising the steps of: selecting only a predetermine low frequency band of all frequency bands;sampling the predetermine low frequency band of all frequency bands at a lower rate than a Nyquist rate to acquire magnetic resonance data; andproducing on the display device a final MRI image that is generated using the magnetic resonance data. 2. The method of claim 1, wherein the method of sampling is a SPACE RIP (sensitivity profiles from an array of coils for encoding and reconstruction in parallel) sampling method. 3. The method of claim 1, wherein the step of producing the final MRI image comprises: generating an intermediate MRI image by reconstructing the magnetic resonance data with a reconstruction method of a generalized series imaging technique; andgenerating the final MRI image, in which the magnetic resonance data has been reconstructed, by reconstructing the intermediate MRI image with a reconstruction method of a parallel imaging technique. 4. The method of claim 3, wherein the step of generating the intermediate MRI image comprises a regularization method. 5. The method of claim 3, wherein the step of generating the final MRI image comprises a regularization method. 6. The method of claim 3, wherein the reconstruction method of the parallel imaging technique is a SPACE RIP reconstruction method. 7. A computer readable recording medium implemented with a program for executing a high-resolution MRI generating method that produces a high-resolution MRI image on a display device by using a generalized series parallel imaging method, the high-resolution MRI image generating method comprising the steps of: selecting only a predetermined low frequency band of all frequency bands; andsampling the predetermined low frequency band of all frequency bands at a lower rate than a Nyquist rate to acquire magnetic resonance data. 8. The compute readable recording medium of claim 7, wherein the method of sampling is a SPACE RIP (sensitivity profiles from an array of coils for encoding and reconstruction in parallel) sampling method. 9. The computer readable recording medium of claim 7, further comprising the steps of: generating an intermediate MRI image by reconstructing the magnetic resonance data with a reconstruction method of a generalized series imaging technique; andgenerating a final MRI image, in which the magnetic resonance data has been reconstructed, by reconstructing the intermediate MRI image with a reconstruction method of a parallel imaging technique. 10. The computer readable recording medium of claim 9, wherein the step of generating the intermediate MRI image comprises a regularization method. 11. The computer readable recording medium of claim 9, wherein the step of generating the final MRI image comprises a regularization method. 12. The computer reada... | Korea Advanced Institute of Science & Technology | G01V003/00
US07433733 | B2 | 2008-10-07 | Motion measurement method, motion measurement system, and motion measurement program | 2005-06-16 | A method, system, and program enabling a variable varying with an animal's motions to be accurately measured. First, a reference relational expression representing a relation between myoelectric potentials and motion is set on the basis of the myoelectric potentials and the variable measured by using an inverse dynamics model. Then, the motion variable is measured on the basis of the measured myoelectric potentials and according to the reference relational expression. A reference relational expression is set on the basis of myoelectric potentials measured in a first motion and a motion variable corresponding to the first motion. Then, a motion variable corresponding to a second motion is measured on the basis of the myoelectric potentials measured in the second motion and according to the reference relational expression. | 1. A method of measuring motions of an animal made by activities of muscle fibers, comprising the steps of: a first myoelectric potential measurement step for measuring, by electrodes adapted to be attached to an animal's body, first myoelectric potentials generated in the animal's body;a first variable measurement step for measuring, based on outputs from sensors adapted to be attached to the animal's body, a motion variable varying with the animal's motions;a reference relational expression setting step for setting a reference relational expression representing a relation between the myoelectric potentials and the motion variable based on the first myoelectric potentials measured in the first myoelectric potential measurement step and the motion variable measured in the first variable measurement step;a second myoelectric potential measurement step for measuring, by said electrodes attached to the animal's body, second myoelectric potentials generated in the animal's body, following said reference relational expression setting step; anda second variable measurement step for newly measuring said motion variable based on the myoelectric potentials measured in the second myoelectric potential measurement step and according to the reference relational expression set in the reference relational expression setting step. 2. The motion measurement method according to claim 1, wherein the first variable measurement step includes measuring the motion variable according to an inverse dynamics model representing a behavior of the animal. 3. The motion measurement method according to claim 1, further comprising a motion condition determination step of determining a motion condition of the animal based on the motion variable varying with the animal's motions, wherein the first variable measurement step includes measuring the motion variable based on the motion condition determined in the motion condition determination step. 4. The motion measurement method according to claim 1, further comprising the steps of: a third variable measurement step of measuring, based on the outputs from the sensors attached to the animal's body, the motion variable intended to be newly measured in the second variable measurement step simultaneously with and separately from the second variable measurement step; anda deviation determination step of determining whether a deviation between the motion variable newly measured in the second variable measurement step and the motion variable measured in the third variable measurement step exceeds a threshold,wherein the reference relational expression setting step includes setting a new reference relational expression based on the myoelectric potentials measured in the myoelectric potential measurement step and the variable measured in the third variable measurement step if the deviation is determined to exceed the threshold in the deviation determination step. 5. The motion measurement method according to claim 1, wherein the refere... | Honda Motor Co., Ltd. | A61B005/04 A61B005/117 A61B005/103
US07801593 | B2 | 2010-09-21 | System, software, and method for detection of sleep-disordered breathing using an electrocardiogram | 2008-02-13 | A system to form and store an electrocardiogram (ECG) signal derived from a cardiac electrical signal that includes an apparatus having a pair of the electrodes to connect to a patient to detect the cardiac electrical signal. A signal sampler samples the cardiac electrical signal to form the ECG signal. A data storage device stores the ECG signal. A computer communicates with the data storage device to retrieve the ECG signal for analysis by software stored in the memory of the computer. The software analyzes a morphology of the amplitude of a plurality of R-wave peaks contained within the ECG signal and/or analyzes a morphology of the area of a plurality of QRS complex pulses contained within the ECG signal. | 1. A method for detecting sleep-disordered breathing, the method comprising: detecting a plurality of QRS complex pulses within an electrocardiogram signal from only a single lead;determining a plurality of QRS complex pulse area values for a preselected time band, variations in the area of each of the plurality of QRS complex pulse areas defining QRS complex pulse area morphology for the plurality of QRS complex pulses;analyzing the QRS complex pulse area morphology for the plurality of QRS complex pulses; anddetecting sleep disordered breathing, when existing, responsive to the analysis of the change in value between the area of each of the plurality of QRS complex pulse areas. 2. A method as defined in claim 1, further comprising the step of: forming a QRS complex pulse area value sampled series by interpolating between QRS complex pulse area values of an electrocardiogram signal to form a QRS complex pulse area value trend and by sampling the formed QRS complex pulse area value trend; and wherein the step of analyzing the QRS complex pulse area morphology of the plurality of QRS complex pulses includes the steps of: determining a power spectral density of the QRS complex pulse area value sampled series, anddetermining a power level in a preselected frequency band of the power spectral density. 3. A method as defined in claim 2, further comprising the step of comparing the determined power level to a predetermined threshold level to determine if the predetermined threshold level is exceeded. 4. A method as defined in claim 2, wherein the preselected frequency band is a first frequency band having a first power level, the method further comprising performing at least one of the following steps: comparing the first power level to a first predetermined threshold level to determine if the first predetermined threshold level is exceeded, the preselected frequency band approximately between 0.0 to 0.1 Hz; andcomparing a second power level of a second frequency band to a second predetermined threshold level to determine if the second power level fails to exceed the second predetermined threshold level, the second preselected frequency band approximately between 0.2 to 0.3 Hz. 5. A method as defined in claim 2, wherein the preselected time band is less than 100 milliseconds wide. 6. A method as defined in claim 1, further comprising the steps of: dividing the electrocardiogram signal into a first plurality of epochs;forming an QRS complex pulse area value sampled series for each epoch in a subset of the first plurality of epochs defining a second plurality of epochs by interpolating between QRS complex pulse area values of each epoch in the second plurality of epochs to form an QRS complex pulse area value trend for each epoch in the second plurality of epochs and by sampling the formed QRS complex pulse area value trend for each of the second plurality of epochs; anddetermining a power spectral density of the QRS complex pulse area value sampled... | Board of Regents, The University of Texas System John R. Burk Edgar A. Lucas | A61B005/04
US07471977 | B2 | 2008-12-30 | Method and system for detecting pace pulses | 2004-06-30 | The present technique provides for the detection of pace pulses in electrocardiogram data. The technique provides for processing one or more sets of electrocardiogram data via a non-linear algorithm. Furthermore, the technique provides for detecting one or more pace pulses in the one or more sets of electrocardiogram data via a non-linear detection algorithm. Systems and computer programs that afford functionality of the type defined by this method are also provided by the present technique. | 1. A method for detecting pace pulses, the method comprising the steps of: processing one or more sets of digital electrocardiogram data via a non-linear enhancement algorithm;detecting one or more pulses in the one or more sets of digital electrocardiogram data via a non-linear detection algorithm; anddisplaying the pace pulses on a display unit. 2. The method of claim 1, wherein processing one or more sets of digital electrocardiogram data comprises filtering the one or more sets of electrocardiogram data. 3. The method of claim 2, wherein filtering the one or more sets of digital electrocardiogram data comprises applying a median filter. 4. The method of claim 1, wherein processing one or more sets of digital electrocardiogram data comprises applying a non-linear pulse signature enhancement to the one or more sets of digital electrocardiogram data. 5. The method of claim 4, comprising processing the one or more sets of digital electrocardiogram data with a differentiator. 6. The method of claim 4, wherein applying the non-linear pulse signature enhancement comprises enhancing a pulse signature of differentiated pace pulses. 7. The method of claim 1, comprising the step of acquiring at least one respective set of electrocardiogram data from at least one electrode. 8. The method of claim 7, comprising the step of digitizing the each set of electrocardiogram data to generate each set of digitized electrocardiogram data. 9. The method of claim 1, comprising the step of combining two or more sets of digital electrocardiogram data via a cross-multiplicative combiner to generate a combined set of digital electrocardiogram data. 10. The method of claim 1, wherein detecting one or more pulses in the one or more sets of digital electrocardiogram data comprises detecting one or more pace pulses in the one or more sets of electrocardiogram data based upon a signal to noise ratio. 11. The method of claim 10, wherein detecting one or more pulses in the one or more sets of digital electrocardiogram data based upon a signal to noise ratio comprises detecting a signal to noise ratio associated with the combined set of electrocardiogram data. 12. The method of claim 1, wherein displaying the pace pulses comprises superimposing the pace pulses on an electrocardiogram trace. 13. A computer program comprising one or more tangible media, wherein the one or more tangible media comprise: code adapted to process one or more sets of digital electrocardiogram data via a non-linear enhancement algorithm;code adapted to detect one or more pulses in the one or more sets of digital electrocardiogram data via a non-linear detection algorithm; andcode adapted to display the pace pulses on a display unit. 14. The computer program, as recited in claim 13, wherein the code adapted to process one or more sets of digital electrocardiogram data comprises: code adapted to filter one or more sets of electrocardiogram data;code adapted to process the one or more sets of e... | General Electric Company | A61B005/04
US07806832 | B2 | 2010-10-05 | False positive reduction in SPO2 atrial fibrillation detection using average heart rate and NIBP | 2007-04-30 | The method and system includes detecting atrial fibrillation in a patient by monitoring the blood oxygen saturation level over a period of time. The method and system produces a plethysmographic waveform from the monitored blood oxygen saturation level and analyzes the plethysmographic waveform and detected intervals and determines whether the patient is in atrial fibrillation. The method and system is preferably implemented in a software application and may be configured to report to the user on the current state of atrial fibrillation (AFIB) and a current trend. | 1. A method of detecting atrial fibrillation in a human or animal subject, the method comprising: acquiring a plethysmographic waveform of a patient from a monitoring device over a period of time;extracting a pulse from the plethysmographic waveform;measuring a set of pulse intervals from the plethysmographic waveform;measuring a set of features from the plethysmographic waveform;analyzing the set of pulse intervals and the set of pulse features from the plethysmographic waveform to determine an atrial fibrillation state of the patient;analyzing the pulse from the plethysmographic waveform, wherein analyzing the pulse from the plethysmographic waveform includes analyzing an average heart rate increase and analyzing an average heart rate variability;acquiring and analyzing an NIBP pulse of the patient;acquiring and analyzing a set of medical records of the patient; andcombining results from the pulse analysis, NIBP pulse analysis and the medical records analysis and confirming the atrial fibrillation state based on said results. 2. The method as claimed in claim 1, wherein the monitoring device is a pulse oximeter sensor. 3. The method as claimed in claim 1, wherein one of the set of features is a set of pulse amplitudes. 4. The method as claimed in claim 1, further comprising analyzing the set of pulse intervals and the set of pulse features from the plethysmographic waveform to detect an atrial fibrillation trend of the patient. 5. The method as claimed in claim 4, further comprising reporting the atrial fibrillation trend. 6. The method as claimed in claim 1, further comprising reporting the atrial fibrillation state. 7. The method as claimed in claim 1, wherein analyzing the set of pulse intervals includes: conducting a hidden Markov model analysis on the set of pulse intervals; and conducting a contextual analysis on the set of pulse intervals. 8. The method as claimed in claim 7, wherein the hidden Markov model analysis determines a probability of the patient having an irregular heart rhythm. 9. The method as claimed in claim 7, wherein the contextual analysis rules out atrial fibrillation as a source of irregular heart rhythm. 10. The method as claimed in claim 1, wherein analyzing the set of pulse features includes: calculating amplitude variation between each consecutive one of the set of pulse amplitudes. 11. The method as claimed in claim 10, wherein the atrial fibrillation state determination is dependent on an analysis of pulse amplitude variation. 12. The method as claimed in claim 1, wherein analyzing the set of pulse amplitudes includes determining a set of filling time dynamics of the set of pulse amplitudes. 13. The method as claimed in claim 12, wherein the steps of analyzing the set of pulse amplitudes and determining the set of filling time dynamics provides an indication of the degree of atrial augmentation present. 14. The method as claimed in claim 1, wherein the pulse analyzing step contributes to determining ... | The General Electric Company | A61B005/02
US07471814 | B2 | 2008-12-30 | Curved-slab maximum intensity projections | 2004-12-02 | A method to define a curved slab region of interest that includes vessels while maximally excluding surrounding soft tissue and bone is provided. The thickness of the curved slab is automatically adapted to the thickness of the vessel and follows the tortuous vessel(s) so that an increase in tortuousity does not result in a disproportionate increase in the region of interest required to enclose the vessel. A plurality of boundary pairs is determined in the view plane to define a vessel. Vessel-intensities are determined for each one of the boundary pairs. The boundary pairs with associated intensities define the view of the vessel in the projection plane. Context-intensity could be defined in the area surrounding the boundary pairs in the projection and/or transverse plane. The method also includes several steps that will result in a better isolation and removal of non-vessel structures and view of the vessel(s) and its(their) branches. | 1. A method for generating a curved-slab region of interest including a vessel, comprising the steps of: (a) identifying a vessel in the projection plane of an image;(b) projecting, for each pixel in said projecting plane, a line in the view plane of said image, said projected line in said view plane intersecting with said vessel;(c) identifying, along each of said lines in the view plane of said image, a first set of boundary points which approximate the intersection points of the vessel boundaries with surrounding tissue or bone along said line;(d) determining, for each of said first set of boundary pairs, a vessel-intensity value representing a combined-intensity value of the voxels within said boundary points and along said line, and determining a context-intensity representing a combined-intensity value for the voxels outside said boundary points and along said line; and(e) assigning said boundary pairs and corresponding vessel intensities to the respective pixels resulting in said curved slab region of interest of said vessel. 2. The method as set forth in claim 1, further comprising removing the corresponding voxels having said context-intensity from said image. 3. The method as set forth in claim 1, further comprising fine-tuning each of said first set of boundary points and corresponding vessel- and context-intensity values, said fine-tuned boundary points define new intersection points of the vessel boundaries with surrounding tissue or bone along said lines, and assigning said fine-tuned boundary points and corresponding vessel intensities to the respective pixels resulting in said curved slab region of interest of said vessel. 4. The method as set forth in claim 3, further comprising removing the corresponding voxels having said fine-tuned context-intensity from said image. | The Board of Trustees of the Leland Stanford Junior University | G06T017/00 G06K009/00 A61B005/05
US07435229 | B2 | 2008-10-14 | System for transcutaneous monitoring of intracranial pressure (ICP) using near infrared (NIR) telemetry | 2005-02-24 | A system for measuring and converting to an observer intelligible form an internal physiological parameter of a medical patient. The invention allows transcutaneous telemetry of the measured information intracranial pressure via a system which includes a patient implanted sensor module and a processing and display module which is external of the patient and optically coupled to the sensor module via an external coupling module. A sensor within the implanted module transduces the measured information and a near infrared (NIR) emitter transmits this telemetry information when interrogated by the complementary external coupling module. Power for the sensor module is derived inductively through rectification of a transcutaneously-applied high-frequency alternating electromagnetic field which is generated by a power source within the external coupling module, in concept much like a conventional electrical transformer. A computer within the processing and display module calculates the parameter value from the NIR telemetry signal and represents this data either in numerical, graphical, or analog format. | 1. A system for measuring and converting to an observer intelligible form an internal physiological pressure of a medical patient, comprising: (a) a patient-implantable sensor module;(b) an external coupling module positionable externally of the medical patient; and(c) a processing and display module for receiving an output from the coupling module and generating the observer intelligible form indicative of the internal physiological pressure;(d) the patient-implantable sensor module including: (1) means for obtaining a measured physiological pressure and providing a digital signal indicative of the measured physiological pressure;(2) a source of NIR wavelength light;(3) means for modulating Near infrared (NIR) wavelength light from the source of Near infrared (NIR) wavelength light to incorporate the digital signal and thereby producing a NIR telemetry signal indicative of the measured physiological parameter;(4) means for directing the NIR telemetry signal through the medical patient's tissues to a point external to the medical patient for coupling with the processing and display module;(5) a chamber in communication with a bodily fluid of the medical patient, and a deformable diaphragm mounted in the chamber and being deformable as a function of the pressure of the bodily fluid; and means for converting a deformation state of the diaphragm to the digital signal;(6)wherein the diaphragm comprises a transducer crystal which exhibits a change in oscillation frequency as a function of its deformation;(7) a monolithic circuit including means for modulating the output of the source of NIR wavelength light as a function of the pressure on the transducer crystal;(8) wherein the transducer crystal comprises an x-cut quartz crystal, means for resonating the transducer crystal at its fundamental frequency, and means for determining a change in oscillation frequency and utilizing the change in oscillation frequency to convert the deformation state to the digital signal;(9) wherein the transducer crystal is paired with a nominally identical reference crystal which is not in contact with the bodily fluid, a set of output signals from the transducer crystal and the nominally identical reference crystal, respectively, controlling a transducer oscillator and a reference oscillator; means for detecting and measuring a difference frequency between the transducer oscillator and the reference oscillator; and means for using the difference frequency to compute the pressure applied to the transducer crystal;(10) wherein the monolithic circuit is positioned between the transducer crystal and the nominally identical reference crystal; and,(11) wherein the source of NIR wavelength light is a vertical cavity surface emitting laser positioned to transmit NIR radiation through a NIR portal in the transducer crystal; and including means for turning the source of NIR wavelength light on and off at the difference frequency;(e) the coupling module including means for receivin... | - | A61B005/00
US07815576 | B2 | 2010-10-19 | Method for determining haemodynamic parameters | 2005-02-16 | A method for determining haemodynamic parameters from patient pressure signals includes receiving patient pressure signals from a measurement receiver to which vessel pressures measured in the body of a patient are supplied, selecting and determining a patient pressure signal as reference pressure signal from the patient pressure signals sent from the measurement receiver, dividing the patient pressure signals and reference pressure signal into segments, and examining these segments for artifacts with methods in the time range by using the first derivative with respect to time of the pressure signals wherein artifact afflicted segments are discarded. The method further includes determining the start and end of a heart beat by the reference pressure signal and using the non-discarded segments of the patient pressure signals for calculating haemodynamic parameters. | 1. A method for determining haemodynamic parameters from patient pressure signals, the method comprising: a) receiving patient pressure signals from a measurement receiver to which vessel pressures measured in the body of a patient are supplied;b) selecting and determining a patient pressure signal as a reference pressure signal from the patient pressure signals sent from the measurement receiver;c) dividing the patient pressure signals and the reference pressure signal into segments;d) examining the patient pressure signal segments and the reference pressure signal segments in a time range for artifacts with time range methods by using a first derivative with respect to time of the patient pressure signals and the reference pressure signal wherein artifact-afflicted segments are discarded;e) determining a start and end of a heart beat on at least one of the patient pressure signals using the reference pressure signal; andf) using non-discarded segments of the patient pressure signals for calculating haemodynamic parameters. 2. The method according to claim 1, wherein the examination of the segments for artifacts is carried out with time range methods and frequency range methods. 3. The method according to claim 1, wherein an analysis of the first derivative with respect to time is carried out for determining signal type. 4. The method according to claim 1, wherein for determining signal type, a discriminant analysis or logistic regression is carried out based on beat-wise analysis of the patient pressure signals in time and frequency ranges. 5. The method according to claim 1, wherein to calculate beat volume as haemodynamic parameters, a beat-wise multiple regression based on parameters from standardized Fourier transformation and analysis of the patient pressure signals and derivatives of the parameters is undertaken in the time range. 6. The method according to claim 5, wherein at least one of size, weight, sex, and age is used for calibrating the beat volume as a haemodynamic parameter. 7. The method according to claim 1, wherein a signal-type evaluation is carried out to determine a type of measured pressure signal. 8. The method according to claim 1, wherein a low frequency analysis is carried out on the pressure signals. 9. The method according to claim 1, wherein the calculated haemodynamic parameters are stored electronically or displayed on an output unit. 10. The method according to claim 9, wherein the haemodynamic parameters are displayed on a monitor. | Deutsches Herzzentrum Berlin | A61B005/00
US08172764 | B2 | 2012-05-08 | Method and system for high-resolution extraction of quasi-periodic signals | 2006-10-26 | A quasi-periodic signal with high signal-to-noise ratio containing signal features that exhibit poor temporal localization is processed to identify waveform temporal reference points that are used to provide the temporal reference points for extracting a representative waveform of a signal feature having high temporal localization in a second, related, quasi-periodic signal that has low signal-to-noise ratio. The resulting representative waveform exhibits much improved signal-to-noise ratio while preserving the temporal detail contained in the second, related, quasi-periodic signal. | 1. A method for determining a plurality of temporal reference points within a quasi-periodic signal comprising the steps of: acquiring (304) a first quasi-periodic signal, the first quasi-periodic signal having a low temporal resolution, and a high signal-to-noise ratio;acquiring a second quasi-periodic signal, the second quasi-periodic signal being related to the first quasi-periodic signal, and having a high temporal resolution, and a low signal-to-noise ratio;determining (310) a nominal interval of the first quasi-periodic signal by determining a most common frequency component of the first quasi-periodic signal;constructing and providing (312) a shape filter for the first quasi-periodic signal using the nominal interval for determining (320) slope crossings;forming (324) a template for the first quasi-periodic signal;identifying (328) reference time points for the first quasi-periodic signal by comparing the template with the first quasi-periodic signal;providing (330) an envelope for the second quasi-periodic signal using the reference times;estimating (336) reference points for the second quasi-periodic signal within the envelope;forming (342) a second template for the second quasi-periodic signal;cross-correlating (346) the second template with the second quasi-periodic signal around the reference points resulting in a highest correlation peak (350); andsetting (352) the reference points to the time of the highest correlation peak. 2. The method according to claim 1, wherein the first quasi-periodic signal is a first heart-beat waveform, and the second quasi-periodic signal is a second heart-beat waveform. 3. The method according to claim 2, wherein the second heart beat waveform is a phonocardiogram. 4. The method according to claim 2, wherein the first heart beat waveform is a plethysmogram having a plurality of blood-density pulses. 5. The method according to claim 1, wherein the step of determining a nominal beat rate further comprises the steps of: low-pass-filtering (306) the first quasi-periodic signal; anddetermining (310) the largest frequency component in a frequency domain after a Fourier transform of the low-pass-filtered first quasi-periodic signal. 6. The method according to claim 1, wherein the shape filter for identifying signal events of interest is a morphological opening. 7. The method according to claim 6, wherein the step of providing a threshold further comprises the steps of: subtracting (312) the morphological opening of the first quasi-periodic signal from the first quasi-periodic signal;determining (314) a maximum for the subtracted first quasi-periodic signal;providing a plurality of sub-thresholds between a baseline and the maximum;determining (316) slope-crossings for each of the plurality of sub-thresholds resulting in a group of sub-thresholds, each of the group of sub-thresholds having identical number of crossings;determining (318) a midpoint of the sub-thresholds, andproviding the midpoint as the thr... | Coeurmetrics Inc. | A61B005/02
US08260405 | B2 | 2012-09-04 | Monitoring apparatus for monitoring a user's heart rate and/or heart rate variation; wristwatch comprising such a monitoring apparatus | 2006-12-11 | A monitoring apparatus (4) monitors a user's heart rate and/or heart rate variation. The apparatus includes a capacitor (22) which is positionable on or near a body part of a person, for example a person's limb, for example an arm (3), such that an electrical capacitance of the capacitor (22) is influenced by blood pressure pulses of blood running through the body part. The apparatus (4) also includes a data processor (26) for determining the heart rate and/or heart rate variation from changes in the capacitance of the capacitor (22). | 1. A monitoring apparatus for monitoring a user's heart rate and/or heart rate variation, the apparatus comprising: an oscillator which generates an oscillating signal;first and second capacitor electrodes, the first electrode is connected to said oscillator, the capacitor electrodes being arranged to be mounted to capacitively couple with different surface parts of a blood carrying body part of a person, such that an electrical capacitance between the electrodes across the body part between the electrodes changes in response to blood pressure pulses of blood running through the capacitively coupled body part;a data processor connected to said oscillator and said second electrode, the processor being programmed to determine a heart rate or heart rate variation from changes in the capacitance of the capacitively coupled electrodes and body part; andan output device connected to the processor to provide an auditory and/or visual output indicative of the heart rate or the heart rate variation. 2. The apparatus according to claim 1, wherein the oscillator is arranged to generate a binary signal. 3. The apparatus according to claim 1, wherein the data processor comprises a frequency detector arranged to determine frequency changes. 4. The apparatus according to claim 1, wherein the data processor comprises a one shot unit which generates a pulse train, the pulse train being formed by pulses generated by the one shot unit in response to an up going flank or a down going flank in the oscillating signal. 5. The apparatus according to claim 4, wherein the data processor comprises an averaging unit for generating a moving average of the pulse train, the moving average being indicative of a frequency changes of the oscillating signal. 6. The apparatus according to claim 1, the different surface parts being located at substantially opposite sides of the person's body part, particularly such that the body part extends between the electrodes during use and changes in blood volume in the body part during each heart beat changes the capacitive coupling of the electrodes. 7. The apparatus according to claim 1, wherein the apparatus is provided with a wristband, the wristband supporting the first and second capacitor electrodes in a spaced relationship, and supporting the data processor. 8. A wristband or wrist watch comprising a monitoring apparatus according to claim 1. 9. The apparatus according to claim 1, wherein the heart rate and/or the heart rate variations are used as a parameter in computer games. 10. The apparatus according to claim 1, wherein the heart rate and/or the heart rate variations are used in a training scheme. 11. A method for monitoring a heart of a patient using the apparatus according to claim 1, wherein the heart rate and/or heart rate variations are used to determine the occurrence of a heart failure, wherein, upon the determination of a heart failure, the output device generates an alarm signal. 12. The method according to cl... | Koninklijke Philips Electronics N.V. | A61N005/04 A61N005/02
US08082021 | B2 | 2011-12-20 | Markers for interventional devices in magnetic resonant image (MRI) systems | 2008-08-18 | Methods and apparatuses for locating a catheter, or other medical international devices are described herein. In one aspect of the invention, an exemplary method includes providing a pressure sensitive device attached to a body of a catheter, providing a resonant circuit coupled to the pressure sensitive device in the catheter, the resonant circuit having an oscillating frequency, generating a magnetic field, the magnetic field inducing and causing the resonant circuit to oscillate, applying a pressure on the pressure sensitive device, the pressure applied on the pressure sensitive device changing the frequency of the resonant circuit, detecting the oscillation of the resonant circuit, and determining location of the catheter based on the oscillation of the resonant circuit. Other methods and apparatuses are also described. | 1. A method for locating a catheter, the method comprising: providing a pressure sensitive device and a piezoelectric device attached to a body of a catheter;providing a resonant circuit coupled to the pressure sensitive device and the piezoelectric device of the catheter, the resonant circuit having a resonant oscillating frequency;applying a first pressure on the piezoelectric device, the first pressure applied on the piezoelectric device causing the piezoelectric device to generate a current in and cause the resonant circuit to oscillate;applying a second pressure on the pressure sensitive device, the second pressure applied on the pressure sensitive device changing the resonant frequency of the resonant circuit;detecting the oscillation of the resonant circuit; anddetermining location of the catheter based on the oscillation of the resonant circuit. 2. The method of claim 1, wherein the pressure sensitive device and the piezoelectric device are attached to an outer wall of the catheter. 3. The method of claim 1, wherein the pressure sensitive device and the piezoelectric device are incorporated into a wall of the catheter. 4. The method of claim 1, wherein the resonant circuit is located at a distal tip region of the catheter. 5. The method of claim 1, wherein the resonant circuit is an LC circuit incorporated into a tip structure of the catheter. 6. The method of claim 5, wherein the LC circuit comprises: a capacitor;an inductor coil coupled to the capacitor in parallel or in series; andwherein the inductor coil is wound around the catheter in a way such that the resonant circuit contains a maximum Q factor. 7. The method of claim 6, wherein the LC circuit comprises multiple inductor coils. 8. The method of claim 7, wherein the multiple inductor coils are separated by capacitors coupled in series or in parallel. 9. The method of claim 1, wherein the first pressure applied on the pressure sensitive device is received from an internal part of a patient's body. 10. The method of claim 9, wherein the first pressure applied on the pressure sensitive device is received from a blood pressure of the patient's body. 11. The method of claim 1, wherein the first pressure applied on the pressure sensitive device is received from an external pressure source through a pressure controller. 12. The method of claim 11, further comprising adjusting the first pressure applied on the pressure sensitive device through the pressure controller, such that the resonant circuit is operating in and out of a resonant state. 13. The method of claim 1, wherein the pressure sensitive device is a pressure sensitive resistor, wherein the resistance of the pressure sensitive resistor changes based on the pressure applied on the pressure sensitive resistor. 14. The method of claim 1, wherein the second pressure applied on the piezoelectric device is received from an internal part of a patient's body. 15. The method of claim 14, wherein the second pressure applie... | Abbott Cardiovascular Systems Inc. | A61B005/05 A61F002/00
US08887733 | B2 | 2014-11-18 | Endovascular treatment apparatus and method | 2008-06-12 | An endovascular sheath device for use with a thermal treatment apparatus is provided. The device includes a longitudinal tube which is designed to receive a thermal treatment device and is designed to be inserted into a blood vessel. An ultrasonically visible reinforcement element is disposed along the length of the longitudinal tube. The reinforcement element such as a braided wire provides several functions including increased visibility under ultrasound, clearer identification of sheath tip, and increased durability to protect the fiber from needle punctures during tumescent injections into the perivenous space. The wire reinforcement also increases shaft torquability and kink resistance during sheath insertion and withdrawal. | 1. A method of using an endovascular thermal treatment delivery device comprising the steps of: inserting into a blood vessel a longitudinal tube having an ultrasonically visible reinforcement element disposed along a wall of the longitudinal tube;positioning a tip of the longitudinal tube at a target site using the ultrasonically visible reinforcement element in the longitudinal tube as an ultrasonic guide; andinserting a thermal treatment device through a lumen of the longitudinal tube;wherein the thermal treatment device is an optical fiber, said method including the step of visually checking an aiming beam to locate an optical fiber tip while the optical fiber is being inserted through the longitudinal tube. 2. The method according to claim 1, wherein at least one end marker is disposed near the tip of the longitudinal tube, and further comprising: using the at least one end marker as a visual check that the tip of the longitudinal tube is near an exit site. 3. The method according to claim 1, wherein at least one end marker is disposed near the tip of the longitudinal tube, and further comprising preparing to turn off the thermal treatment device when the end marker becomes visible. 4. The method according to claim 1, further comprising: sliding an adjustable depth stop along the longitudinal tube to a position that approximately indicates the length of the longitudinal tube that should be inserted into the blood vessel to reach the target site. | AngioDynamics, Inc. | A61B019/00 A61B018/24
US08442610 | B2 | 2013-05-14 | System and methods for processing analyte sensor data | 2008-08-21 | Systems and methods for processing sensor analyte data, including initiating calibration, updating calibration, evaluating clinical acceptability of reference and sensor analyte data, and evaluating the quality of sensor calibration. During initial calibration, the analyte sensor data is evaluated over a period of time to determine stability of the sensor. The sensor may be calibrated using a calibration set of one or more matched sensor and reference analyte data pairs. The calibration may be updated after evaluating the calibration set for best calibration based on inclusion criteria with newly received reference analyte data. Fail-safe mechanisms are provided based on clinical acceptability of reference and analyte data and quality of sensor calibration. Algorithms provide for optimized prospective and retrospective analysis of estimated blood analyte data from an analyte sensor. | 1. A method for calibrating a substantially continuous analyte sensor, the method comprising: receiving, by a communication module, a data stream comprising one or more sensor data points from a substantially continuous analyte sensor;receiving, by the communication module, reference data comprising one or more reference data points from a reference analyte monitor;forming, by a processor module, at least one matched data pair by matching the one or more reference data points to the one or more sensor data points;forming, by the processor module, a calibration data set using the one or more sensor data points and the one or more reference points using the at least one matched data pair;determining, by the processor module, a level of confidence based on data that forms the calibration data set by measuring an association of the calibration data set to a curve using a statistical analysis; andproviding one of a plurality of different display modes on a display screen based on the level of confidence. 2. The method of claim 1, wherein providing one of a plurality of different display modes comprises displaying at least one of error bars or ±values. 3. The method of claim 1, wherein providing one of a plurality of different display modes comprises initiating a normal display mode based on the level of confidence. 4. The method of claim 1, wherein providing one of a plurality of different display modes comprises initiating an initial fail-safe display mode based on the level of confidence. 5. The method of claim 4, wherein initiating the initial fail-safe mode comprises altering a user interface so that sensor data is not displayed. 6. The method of claim 4, wherein initiating the initial fail-safe mode comprises prompting a user to provide an updated reference analyte value. 7. The method of claim 6, further comprising re-determining a level of confidence based on the updated reference analyte value. 8. The method of claim 7, wherein providing one of a plurality of different display modes comprises initiating a full fail-safe mode based on the re-determined level of confidence. 9. The method of claim 1, wherein determining the level of confidence further comprises comparing the measure of the association with a threshold and wherein the level of confidence is determined to be sufficient based on a result of the comparison. 10. The method of claim 1, wherein determining the level of confidence comprises evaluating the calibration data set for clinical acceptability. 11. The method of claim 1, wherein determining the level of confidence comprises measuring an association of the calibration data set using statistical analysis and evaluating the calibration data set for clinical acceptability. 12. The method of claim 1, wherein the processor module is incorporated into sensor electronics, receiver electronics or both. 13. A system for calibrating a substantially continuous analyte sensor, the system comprising: a communication module configu... | DexCom, Inc. | A61B005/05
US08750957 | B2 | 2014-06-10 | Microfabricated neural probes and methods of making same | 2005-06-01 | An embodiment of the invention provides a neural probe containing a plurality of nanoscale recording electrodes. The recording electrodes have a width of 1 micron or less and a distance between adjacent recording electrodes is 10 microns or less. Another embodiment of the invention provides a neural probe comprising a plurality of microfabricated recording electrodes located on a polymer base material, such as a flexible polymer cantilever. | 1. A neural probe comprising: a polymer handle portion made of a first polymer material;metal contact pads exposed in the polymer handle portion and presenting an external flat surface;a flexible polymer cantilever made from a second polymer material extending from the polymer handle portion;a plurality of metal recording pads exposed in the flexible polymer cantilever and presenting a flat external surface, wherein the recording pads are electrically connected to the contact pads via metal conductive interconnections buried in the polymer of the flexible polymer cantilever;a polymer layer made of a third polymer material and presenting a flat external surface, wherein the polymer layer is between the recording pads and the contact pads and the metal conductive interconnections are behind the polymer layer away from the flat external surface of the polymer layer;wherein the contact pad flat surfaces, the recording pad flat surfaces, and the polymer layer flat external surface are flush with each other; anda plurality of elongated metal recording electrodes extending normally from the exposed flat surfaces of the recording pads and separated by a distance, wherein the recording electrodes have a width of 1 micron or less and a height of six microns or less, and wherein a distance between adjacent recording electrodes is 10 microns or less. 2. The probe of claim 1, wherein the distance between adjacent recording electrodes is 1 to 4 microns. 3. The probe of claim 1, wherein the height of the electrodes is in a range of 2 to 5 microns. 4. The probe of claim 1, wherein the width of the electrodes is about 0.1 to about 0.8 microns. 5. The probe of claim 1, wherein the width of the electrodes is about 0.3 to about 0.5 microns. 6. The probe of claim 1, wherein the probe has a probe electrode density of 1/10 nm squared to about 1/50 nm squared. 7. The probe of claim 1, wherein the electrodes have a round cross section. 8. The probe of claim 1, wherein the electrodes comprise sharpened tips. 9. The probe of claim 1, wherein an insulation layer is formed over the electrodes and then the insulation layer is etched back at the electrode tip. 10. The probe of claim 1, wherein the polymer of the flexible polymer cantilever comprises SU-8. 11. The probe of claim 1, wherein the polymer of the flexible polymer cantilever comprises parylene. 12. The probe of claim 1, wherein the polymer handle portion is 75 microns to 500 microns thick. 13. The probe of claim 1, wherein the flexible polymer cantilever is five to 100 microns thick. 14. The probe of claim 1, wherein: a circuit board is electrically connected to the probe;a voltage or current measurement device which is adapted to measure a change in voltage or current due to neuron action potentials detected by the probe is electrically connected to the circuit board; anda voltage or current source adapted to provide a voltage or current to the electrodes to stimulate neuron activity in a living organis... | California Institute of Technology | A61B005/04 A61N001/00
US08371298 | B2 | 2013-02-12 | Method and apparatus for lung volume estimation | 2005-12-06 | In a method, ventilator and ventilator control unit for determining an end-expertorial lung volume (EELV) for a mechanically ventilated patient, a breathing gas is provided to the patient that has a first fixed N2/O2 gas composition, at least until the N2/O2 gas composition in air expired from the patient is constant. At least once, at a first predetermined point in time, the N2/O2 gas composition in the breathing gas is changed to a second fixed composition. The change in the N2/O2 gas composition exhaled by the patient for each breath is measured until a second point in time at which the level of expired O2 in at least two subsequent breadths is substantially stable. The measurement is made downstream of the expiratory tube of the ventilator. The total gas volume is determined for each breath, and the EELV of the patient's lungs is determined based on the change in O2 level between the first and second points in time. | 1. A ventilator for ventilation of a patient, comprising an inhalation gas circuit that provides breathing air comprising a mixture of N2 and O2 to the patient and an exhalation gas circuit that receives exhaled air from the patient, and a gas outlet for said exhaled air, and a control unit configured to calculate the end expiratory lung volume, EELV, of a patient being ventilated by the ventilator, based on the expired volume per breathing cycle, and a composition measuring unit configured to obtain a measurement of a quantity that varies linearly with the N2/O2 composition in the expired gas, said measuring unit being arranged near the gas outlet, the control unit being configured to determine the change in the N2/O2 composition in the expired gas between a first point in time (t1) at which the N2/O2 composition of the breathing gas is changed from a first fixed composition to a second fixed composition, and a second point in time (t2) at which the N2/O2 composition of the expired gas in at least two subsequent breaths is substantially stable, and said control unit being further configured to calculate, said EELV based on the measurement of the quantity that varies linearly with the N2/O2 composition in the expired gas, and the change in the N2/O2 composition exhaled and the total gas volume, for each breath between the first and the second points in time t1 and t2. 2. A ventilator according to claim 1, wherein the control unit is configured to calculate the end expiratory lung volume, EELV, according toEELV=∑n=1N⁢⁢VCen·(xn-xafter)xbefore-xafter-Vsystemwhere VCen is the volume of expired gas at breath number n,x is the quantity that varies linearly with the expired N2/O2 composition,xn is the value of the quantity at breath number n in each breath between the first and second points in time,xbefore is the value of the quantity before the first point in time and,xafter is the value of the quantity at or just after the second point in time. 3. A ventilator according to claim 1, wherein the control unit is arranged to calculate the end expiratory lung volume, EELV, according toEELV=∑n=1N⁢⁢VCen·(O⁢⁢2⁢mixn-O⁢⁢2⁢mixafter)O⁢⁢2⁢mixbefore-O⁢⁢2⁢mixafter-Vsystemwhere VCen is the volume of expired gas at breath number n,O2mixn is the level of O2 in the expired gas at breath number n between the first and second points in time,O2mixafter is the level of O2 in the expired gas at a breath just after the second point in time andO2mixbefore is the level of O2 in the expired gas before the first point in time. 4. A ventilator according to claim 1, wherein the control unit is configured to calculate the EELV according toEELV=∑n=1N⁢⁢VCen·(ρn-ρafter)ρbefore-ρafter-Vsystemwhere ρn is the density of the gas at breath number n between the first and second points in time,ρbefore is the density of the gas at the first point in time, andρafter is the density of the gas at or just after the second point in time. 5. A ventilator according to claim 4, wherein the contro... | Maquet Critical Care AB | F16K031/02
US08897887 | B2 | 2014-11-25 | Band stop filter employing a capacitor and an inductor tank circuit to enhance MRI compatibility of active medical devices | 2008-07-10 | A band stop filter is provided for a lead wire of an active medical device (AMD). The band stop filter includes a capacitor in parallel with an inductor. The parallel capacitor and inductor are placed in series with the lead wire of the AMD, wherein values of capacitance and inductance are selected such that the band stop filter is resonant at a selected frequency. The Q of the inductor may be relatively maximized and the Q of the capacitor may be relatively minimized to reduce the overall Q of the band stop filter to attenuate current flow through the lead wire along a range of selected frequencies. In a preferred form, the band stop filter is integrated into a TIP and/or RING electrode for an active implantable medical device. | 1. An implantable lead, comprising: a) at least one implantable lead wire extending from a proximal lead wire end to a distal lead wire portion having a distal lead wire end, wherein the proximal lead wire end is electrically connectable to electronic circuits of an implantable medical device;b) at least one electrode electrically connected to the distal lead wire portion or the distal lead wire end, wherein the electrode is configured to be placed in contact with biological cells inside a body; andc) a wound portion of the lead wire providing a self-resonant inductor disposed physically and electrically in series with the at least one electrode,d) wherein, along with a capacitive series resistance (Rc) and an inductor series resistance (RL), the self-resonant inductor has sufficient parasitic capacitance between adjacent turns of the wound lead wire so that at low frequencies of about 10 Hz to about 1 kHz, an inductive reactance and the inductor series resistance (RL) permit passage of biological signals from the electrode to the proximal lead wire end τ while a capacitive reactance and the capacitive series resistance (Rc) substantially act as an open circuit to the same biological signals that the inductive reactance and the inductor series resistance (RL) permit to pass, ande) wherein the capacitive series resistance (Rc) and the inductor series resistance (RL) provide a self-resonant inductor Q having a 3-dB bandwidth that is at least 100 kHz. 2. The implantable lead of claim 1, wherein the self-resonant inductor is disposed at or proximally adjacent to the at least one electrode. 3. The implantable lead of claim 1 wherein the at least one electrode is a TIP electrode or a RING electrode, and the self-resonant inductor is integrated at least partially into the electrode. 4. The implantable lead of claim 1 wherein the at least one electrode is selected from the group consisting of a TIP electrode, a RING electrode, and a PAD electrode. 5. The implantable lead of claim 1 wherein the self-resonant inductor has an equivalent L-C resonant electrical circuit comprising a total capacitance in series with the capacitive series resistance (Rc) which are both in parallel with a total inductance in series with a total inductance series resistance. 6. An implantable lead, comprising: a) an implantable lead wire extending from a proximal lead wire end to a distal lead wire portion having a distal lead wire end, wherein the proximal lead wire end is electrically connectable to electronic circuits of an implantable medical device;b) an electrode configured to be placed in contact with biological cells inside a body;c) a wound wire providing a self-resonant coil, wherein the self-resonant coil has a proximal coil end physically and electrically connected to the distal lead wire end, and wherein the self-resonant coil has a distal coil end physically and electrically connected to the electrode so that the self-resonant coil is physically and electricall... | Greatbatch Ltd. | A61B005/055 G01R033/28 A61N001/08 A61B018/14 A61B019/00 A61B018/00
US09114028 | B2 | 2015-08-25 | Arm prosthetic device | 2008-02-06 | A prosthetic arm apparatus comprising a plurality of segments that provide a user of the prosthetic arm apparatus with substantially the same movement capability and function as a human arm. The segments are connectable to one another and connectable to a harness mount that may be adorned by the user. Each segment of the plurality of segments provides a portion of the movement capability, enabling the plurality of connected segments connected to the harness mount to provide substantially the same movement capability as that lacking in the user. | 1. A prosthetic arm apparatus comprising: a plurality of interconnected segments arranged adjacent to one another from a first segment to a final segment, each segment of the plurality of interconnected segments having: an input interface;an output interface; andat least one actuator disposed within the segment between the input interface and the output interface, the at least one actuator driving the output interface through a drive system to cause relative movement in at least one degree of freedom between the input interface and the output interface when actuated such that each segment of the plurality of segments operates independently of each of the other segments of the plurality of interconnected segments;wherein the input interface of the first segment is a harness mount adapted to connect with a prosthetic harness and the output interface of the first segment is adapted to connect to the input interface of an adjacent segment;wherein the input interface of the final segment is adapted to connect to the output interface of an adjacent segment and the output interface of the final segment is adapted to allow a user to manipulate an object; andwherein each segment of the plurality of interconnected segments between the first segment and the final segment is arranged with the input interface engaging the output interface of a first adjacent segment and with the output interface engaging the input interface of a second adjacent segment such that the plurality of interconnected segments provides the user of the prosthetic arm apparatus with movement capability in a plurality of degrees of freedom;wherein each segment of the plurality of segments is adapted to be connected and disconnected from the adjacent segments at the input interface and output interface so that a number of segments in the plurality of segments may be selected depending upon a degree of amputation of the user, thereby providing the user with substantially the same movement capabilities as a human arm; andwherein at least one segment of the plurality of segments includes a non-backdriving clutch disposed in the drive system between the at least one actuator and the output interface, the non-backdriving clutch comprising: an output hex connected to the output interface of the at least one segment;an input cage connected to the at least one actuator;a clutch race constraining the input cage and the output hex; anda plurality of roller bearings disposed within the clutch race around the periphery of the output hex at an interface between the output hex and the input cage;wherein the non-backdriving clutch transfers power from the at least one actuator to the output interface of the at least one segment when the actuator is actuated to generate movement of the output interface in either of two opposing directions and prevents backward transfer of power from the output interface to the at least one actuator of the at least one segment in both of the two opposing directions by an... | DEKA Products Limited Partnership | A61F002/54 A61F002/68 A61F002/58 A61F002/30 A61F002/50 A61F002/70 A61F002/76 A61F002/78
US07983751 | B2 | 2011-07-19 | Measuring conduction velocity using one or more satellite devices | 2006-08-09 | A method for measuring the conduction velocity of a depolarization wave in a tissue employs a first satellite located within the tissue and a second that satellite is located within the tissue a distance away from the first satellite, e.g., by using the time of depolarization wave as reported from each satellite and the distance to determine velocity of the wave. Also provided are systems and kits that find use in accordance with the invention. | 1. A method of determining conduction velocity of a depolarization wave in a tissue, the method comprising: (a) providing a first satellite coupled to a central controller, the first satellite having a first satellite circuit and at least two individually and separately addressable first satellite electrodes, the first satellite circuit adapted to simultaneously and separately couple and decouple each of the at least two first satellite electrodes as directed by the central controller;(b) providing a second satellite coupled to the central controller, the second satellite having a second satellite circuit and at least two individually and separately addressable second satellite electrodes, the second satellite circuit adapted to simultaneously and separately couple and decouple each of the at least two second satellite electrodes as directed by the central controller;(c) determining a time interval between a first satellite's reporting of a depolarization wave in the tissue and a second satellite's reporting of a depolarization wave in the tissue, wherein the first satellite and the second satellite are formed on a lead, and wherein the second satellite is located a fixed distance away from the first satellite; and(d) calculating the conduction velocity using the time interval and the fixed distance between the first satellite and the second satellite. 2. The method according to claim 1, wherein the first and second satellites are coupled to at least one conductive element within the lead. 3. The method according to claim 2, wherein the lead comprises two wires and the first and second satellites are each coupled to the two wires. 4. The method according to claim 2, wherein the lead comprises a single wire and the first and second satellites are each coupled to the single wire. 5. The method according to claim 2, wherein the at least one conductive element is coupled to the central controller. 6. The method according to claim 1, wherein each of the first and second satellites comprises a circuit and at least one effector coupled to the circuit. 7. The method according to claim 5, wherein the method further comprises configuring the first and second satellites to report a presence of the depolarization wave to the central controller. 8. The method according to claim 7, wherein the configuring the first and second satellites are performed sequentially. 9. The method according to claim 7, wherein each of the first and second satellites comprises an effector, and wherein the method further comprises coupling the effector to the at least one conductive element. 10. The method according to claim 9, wherein the coupling of the first and second satellites is performed sequentially. 11. The method according to claim 1, wherein the method further comprises applying a stimulus to the tissue to produce the depolarization wave. 12. The method according to claim 11, wherein the stimulus is a pharmacological stimulus. 13. The method according to cla... | Proteus Biomedical, Inc. | A61N001/00
US08007437 | B2 | 2011-08-30 | Method and apparatus for interactive 4-dimensional (4D) virtual endoscopy | 2007-10-19 | A method of performing 4-dimensional virtual endoscopy includes acquiring a sequence of cardiac images of a patient's heart from an imaging device; acquiring an electrocardiographic signal exhibiting cardiac cycles of the heart; and utilizing the electrocardiographic signal to gate the images for deriving respective 3-dimensional views in succession at corresponding cardiac cycles, from respective cardiac images at a selected phase point common to each of the corresponding cardiac cycles. | 1. A method of performing 4-dimensional (4D) virtual endoscopy, comprising: acquiring a sequence of cardiac images of a patient's heart from an imaging device;acquiring an electrocardiographic signal exhibiting cardiac cycles of said heart;utilizing said electrocardiographic signal to gate said images for deriving respective 3-dimensional (3D) views in succession at corresponding cardiac cycles, from respective cardiac images at a selected phase point common to each of said corresponding cardiac cycles so as to provide a 3 dimensional visualization of said patient's heart as a function of time for said 4D virtual endoscopy;utilizing a virtual camera; andvirtually attaching said virtual camera to a point selected on a wall or surface in said sequence of cardiac images of said patient's heart, such that said virtual camera moves together with said point and will appear static in said respective 3D views in succession so as to provide said 3-dimensional visualization of said patient's heart as a function of time for said 4D virtual endoscopy, so as to produce a substantially steady image of said heart. 2. A method in accordance with claim 1, wherein said step for deriving respective 3D views includes placing said cardiac images in a 3D space and reconstructing a volume. 3. A method in accordance with claim 1, wherein said step of acquiring an electrocardiographic signal includes a step of detecting said cardiac cycles and a step of interactively selecting said phase point. 4. A method in accordance with claim 1, wherein said step for deriving respective 3D views utilizes one of (a) a ray casting algorithm and (b) a mesh extraction algorithm. 5. A method in accordance with claim 1, wherein said step for deriving respective 3D views in succession includes amending a 3D view so as to produce the next 3D view in said succession. 6. A method in accordance with claim 1, wherein said step for deriving respective 3D views includes a step of sensing breathing of said patient and excluding ones of said images affected by said breathing. 7. A method in accordance with claim 1, wherein said step for deriving respective 3D views includes a step of sensing breathing and respiratory cycles of said patient and excluding images not in the same phase of a respiratory cycle. 8. A method in accordance with claim 1, wherein said step for deriving respective 3D views includes a step of sensing sudden catheter movements in said patient and excluding ones of said images affected by said sudden catheter movements. 9. A method of performing 4-dimensional (4D) virtual endoscopy, comprising: acquiring a motion signal of a patient's organ exhibiting motion cycles of said organ;acquiring a sequence of 3-dimensional (3D) images of said organ from an imaging device;utilizing said motion signal to gate said images for deriving respective 3D views in succession at corresponding motion cycles, from respective organ images at a selected phase point common to each of said corre... | Siemens Aktiengesellschaft | A61B008/00
US08229544 | B2 | 2012-07-24 | Detecting temperature and protein denaturation during thermal therapy | 2008-05-05 | In one aspect, in general, a method is provided for detecting temperature and protein denaturation of a tissue during thermal therapy. The method includes generating a plurality of MR pulse sequences that include a first group of pulse sequences and a second group of pulse sequences, and receiving a plurality of response signals that include a first and second group of response signals in response to the first and second groups of pulse sequences, respectively. A first information associated with a degree of protein denaturation of the tissue is determined based on the first and second groups of response signals. A second information associated with a temperature of the tissue is determined based on at least some of the plurality of response signals. | 1. A method for detecting temperature and protein denaturation of a tissue, said method comprising: generating a plurality of MR pulse sequences, including a first group of pulse sequences lacking MT pulses and a second group of pulse sequences having MT pulses, wherein the first and second groups of pulse sequences are interleaved;delivering the plurality of MR pulse sequences to the tissue;receiving a plurality of response signals, including a first group of response signals in response to the first group of pulse sequences, and a second group of response signals in response to the second group of pulse sequences;determining a first information associated with a degree of protein denaturation of the tissue based on the first and second groups of response signals to thereby detect protein denaturation of the tissue; anddetermining a second information associated with a change in temperature of the tissue based on at least some of the plurality of response signals to thereby detect temperature of the tissue,wherein the first information and the second information are determined simultaneously. 2. The method of claim 1, wherein the plurality of MR pulse sequences are configured to be suitable for inducing response signals that are sufficient for determining the first and second information. 3. The method of claim 1, wherein the plurality of MR pulse sequences include dual gradient echo sequences. 4. The method of claim 3, wherein each one of the plurality of response signals includes a first echo component and a second echo component. 5. The method of claim 4, wherein said determining the first information includes comparing at least some of the first echo components of the first group of response signals with some of the first echo components of the second group of response signals. 6. The method of claim 4, wherein said determining the second information includes comparing at least some of the second echo components of the first group of response signals with some of the second echo components of the second group of response signals. 7. The method of claim 1, wherein each one of the first group of pulse sequences includes a RF excitation pulse, and each one of the second group of pulse sequences includes the RF excitation pulse and a MT pulse. 8. The method of claim 7, wherein the MT pulse includes an off-resonance pulse. 9. The method of claim 7, wherein the MT pulse is configured to cause a change in the response signals induced by the RF excitation pulse alone. 10. The method of claim 1, wherein said determining the first information includes comparing the first group of response signals with the second group of response signals. 11. The method of claim 1, wherein the first information includes a magnetization transfer ratio, and the second information includes a phase information. 12. The method of claim 1, further comprising reconstructing an image and forming a temperature map of the tissue based on the plurality of response sig... | National Health Research Institutes | A61B005/05 G01V003/00
US08333719 | B2 | 2012-12-18 | Load apparatus | 2008-03-31 | An oscillation arm moves in a seesaw fashion about an oscillation axis which is supported by a support rod and which serves as a fulcrum. A rotation arm and a pressurizer are provided on one end of the oscillation arm. One end of the rotation arm is rotatably attached to the oscillation arm, and a load weight is placed on the other end of the rotation arm. By rotating the rotation arm, the magnitude of a force to be applied by the load weight to the pressurizer is adjusted. More specifically, by displacing the load weight to a predetermined position by rotation and then causing the load weight to stop at this position, a static load corresponding to the rotation angle can be applied to the pressurizer. Further, by causing the load weight to perform rotational motion, the magnitude of the force to be applied to the pressurizer is periodically varied, so that a load which varies periodically can be applied to hard tissue. | 1. A load apparatus which applies a load to hard tissue, comprising: a support member;an oscillation arm which moves in a seesaw fashion about an axis serving as a fulcrum, the axis being supported by the support member;a pressure mechanism which is provided on one side of the oscillation arm with respect to the fulcrum; anda balancer which is provided on the other side of the oscillation arm with respect to the fulcrum,whereinthe pressure mechanism includes:a load weight disposed on the one side of the oscillation arm with respect to the fulcrum;a pressurizer for applying to the hard tissue a load corresponding to a force supplied by the load weight; andan adjustment unit disposed on the one side of the oscillation arm with respect to the fulcrum, wherein the adjustment unit rotates the load weight to adjust a magnitude of the force to be applied to the pressurizer by the load weight. 2. The load apparatus according to claim 1, wherein the adjustment unit includes a rotation arm for rotating the load weight, one side of the rotation arm being rotatably attached to the oscillation arm, and the load weight being attached to the other side of the rotation arm. 3. The load apparatus according to claim 2, wherein the balancer adjusts weight balance between one side and the other side of the oscillation arm in a state in which the load weight is moved by the rotation arm to a position immediately above the fulcrum. 4. The load apparatus according to claim 3, wherein the balancer moves an adjustment weight linearly along the longitudinal direction of the oscillation arm to fine-tune the weight balance. 5. The load apparatus according to claim 2, wherein the rotation arm causes the load weight to perform rotational motion to allow periodic variation of the magnitude of a force to be applied to the pressurizer, for applying to the hard tissue a load which varies periodically. 6. The load apparatus according to claim 5, wherein the rotation arm rotates about an axis which is attached to the oscillation arm on one side of the rotation arm, to cause the load weight which is attached on the other side of the rotation arm to perform rotational motion between the position immediately above the fulcrum and a position immediately above the pressurizer. 7. The load apparatus according to claim 6, wherein the rotation arm moves the load weight along a circumference of a perfect circle to allow application to the hard tissue of a periodical load having a sinusoidal waveform. 8. The load apparatus according to claim 7, wherein the balancer adjusts weight balance between one side and the other side of the oscillation arm in a state in which the load weight is moved by the rotation arm to a position immediately above the fulcrum. 9. The load apparatus according to claim 8, wherein the balancer moves an adjustment weight linearly along the longitudinal direction of the oscillation arm to fine-tune the weight balance. 10. The load apparatus according to claim 2... | Hitachi Aloka Medical, Ltd. | A61B005/103 A61B005/117 A61B005/00 A61B019/00 B23Q017/20 G01N003/00 G01N003/48 G01N003/42
US08228060 | B2 | 2012-07-24 | Method and apparatus for generating a flip angle schedule for a spin echo train pulse sequence | 2007-06-25 | A method for calculating a flip angle schedule for a train of refocusing radio frequency (RF) pulses with reduced flip angles allows control of RF power deposition and use of a longer echo train. A target signal is defined for each echo in the echo train, and flip angles are then calculated from the target signals. The target signal schedule includes two phases. In the first phase, the target signals drop asymptotically to efficiently establish a pseudo-steady state at a pre-defined minimum signal level, Smin. In the second phase, the target signal is increased monotonically for the remainder of the train to a pre-defined maximum signal level, Smax. By increasing the target signal, the effect of relaxation may be reduced, decreasing blurring and ringing artifacts. Flip angles are then calculated from the target signal schedule, using a simplified method that requires no information about the tissues' relaxation properties. | 1. A method for calculating a flip angle schedule for a spin echo train with reduced refocusing flip angles, usable in magnetic resonance imaging, the spin echo train having a plurality of spin echoes, the method comprising: defining a target signal schedule including target signals for each spin echo in the spin echo train, the target signal schedule comprising a first phase configured to establish a pseudo-steady state and a different, second phase subsequent to the first phase where each of the target signals increase from a pre-defined minimum signal value to a pre-defined maximum signal value; andcalculating a corresponding flip angle schedule from the target signal schedule using only non-relaxation properties wherein the calculated corresponding flip angle schedule is utilized in performing magnetic resonance imaging. 2. A method according to claim 1, wherein the first phase of the target signal schedule comprises target signals that decrease asymptotically to the pre-defined minimum signal value. 3. A method according to claim 1, wherein each of the target signals of the second phase of the target signal schedule increase from the pre-defined minimum signal value to the pre-defined maximum signal value. 4. A method for acquiring magnetic resonance images using a spin echo train with reduced refocusing flip angles, usable in magnetic resonance imaging, the spin echo train having a plurality of spin echoes, the method comprising: applying a magnetic field to an imaging subject comprising tissues to cause a net longitudinal magnetization in the tissues;defining a target signal schedule including target signals for each spin echo in the spin echo train, the target signal schedule comprising a first phase configured to establish a pseudo-steady state and a different, second phase subsequent to the first phase where each of the target signals increase from a pre-defined minimum signal value to a pre-defined maximum signal value;calculating a corresponding flip angle schedule from the target signal schedule using only non-relaxation properties;generating an excitation radio frequency pulse;generating a plurality of refocusing radio frequency pulses, each refocusing radio frequency pulse having a flip angle determined based on the flip angle schedule; andacquiring magnetic resonance imaging data from which magnetic resonance images are created. 5. A method according to claim 4, further comprising: calculating an equivalent TE for each spin echo in the spin echo train. 6. A method according to claim 5, further comprising: selecting a spin echo to fill a center line of k-space based on the equivalent TE of the spin echo. 7. A method according to claim 4, wherein the plurality of refocusing radio frequency pulses is generated in accordance with a single-shot fast spin echo (SSFSE) pulse sequence. 8. A method according to claim 4, wherein the plurality of refocusing radio frequency pulses is generated in accordance with a gradient and spin ech... | General Electric Company | G01V003/00
US08038614 | B2 | 2011-10-18 | Establishing baseline data for physiological monitoring system | 2007-05-22 | A method for establishing baseline data for normalization in a physiological monitoring system which captures images of an individual in order to monitor the well being of the individual, including establishing the identity of an individual; assembling a set of personal data associated with the individual, including semantic data; using the semantic data to identify at least a portion of the wellness parameters to be monitored; acquiring images of the individual during separate capture events to establish baseline image-based wellness parameter data and baseline capture condition data for the individual and; using the baseline wellness parameter data and baseline capture condition data to monitor the physiological condition of the individual. | 1. A method for establishing baseline data for a physiological monitoring system which captures images of an individual in order to monitor the well being of the individual, comprising; (a) operating a physiological monitoring system, including an image processor, a memory, and an image capture device, to automatically assess physiological conditions of a subject individual, where the physiological monitoring system is operated in a fixed capture area and the image capture device automatically acquires color images;(b) automatically acquiring color images of the individual using the image capture device during multiple image capture events;(c) automatically analyzing the color images using the image processor to identify an individual as a subject individual being monitored by the system, based on known personal data for the individual;(d) automatically analyzing the color images using the image processor to derive physiological data and image capture condition data, wherein the image capture condition data includes ambient lighting condition data;(e) automatically deriving baseline wellness parameter data for the subject individual, using the acquired physiological data and the image capture condition data;(f) automatically deriving target criteria for image capture, relative to the capture conditions, that apply during subsequent automatic image captures of the subject individual by the physiological monitoring system; and(g) automatically associating the derived baseline wellness parameter data, baseline capture condition data, and derived target criteria, for the subject individual with the individual's personal data, whereby baseline data is established for the physiological monitoring system for the subject individual. 2. The method of claim 1 in which the derived target criteria for capture conditions define ranges of variation in which the captured images are acceptable. 3. The method of claim 1, in which a subsequent set-up process is initiated based upon detected changes in the image capture conditions prevalent in the capture area. 4. The method of claim 1, further including determining baseline normalization transforms. 5. The method of claim 4, further including determining baseline normalization confidence values, associated with the baseline normalization transforms. 6. The method of claim 1, wherein the wellness parameters are quantitative metrics that characterize physiological traits or conditions, as relate to the health, wellness, or physical attributes of the subject individual. 7. The method of claim 1 wherein the baseline wellness parameter data includes data regarding reference features for the individual. 8. The method of claim 7 wherein the reference features include the whiteness of the sclera or the eye-to-eye distance of the individual. 9. The method of claim 1 wherein the baseline wellness parameter data or the baseline capture data is averaged over multiple capture events. 10. The method of claim 1 wherein ... | Eastman Kodak Company | A61B005/00
US07967810 | B2 | 2011-06-28 | Sub-atmospheric wound-care system | 2007-10-19 | Methods and systems are provided for a sub-atmospheric wound-care (SAWS) system for treating an open wound. The SAWS system includes a regulated vacuum source for developing a negative pressure, a flow rate meter configured to measure a flow rate of liquid removed from the wound, a primary pressure regulating sensor located proximate the wound for directly measuring the negative pressure at the wound, a backup pressure regulating sensor located vacuum tube, a porous dressing suitable to be sealed airtight which is positioned within a wound interface chamber, a collection canister configured to collect said liquid removed from the wound, and an adapter configured to use wall suction a primary regulated vacuum source. | 1. A sub-atmospheric wound-care system for treating an open wound comprising: a connector configured to be connected to wall suction as a regulated vacuum source for developing a negative pressure;a flow rate meter configured to measure a flow rate of liquid removed from the wound;a pressure regulating sensor to be placed at a surface of the wound for directly measuring the negative pressure at the wound;a collection canister with volume level markings configured to collect said liquid removed from the wound; anda porous dressing positioned over the wound and over the pressure regulating sensor placed at the surface of the wound. 2. The sub-atmospheric wound-care system of claim 1, wherein the wall suction is a first regulated vacuum source, the system further comprising: a vacuum motor configured for use as a second regulated vacuum source; anda valve configured to be switched between the first regulated vacuum source and the second regulated vacuum source. 3. The sub-atmospheric wound-care system of claim 1, wherein the vacuum motor is a brushless vacuum motor. 4. The sub-atmospheric wound-care system of claim 1, further comprising: a wound interface chamber connected via a vacuum path to the regulated vacuum source; anda tube connected to the wound interface chamber;wherein the porous dressing is disposed between the wound interface chamber and the wound. 5. The sub-atmospheric wound-care system of claim 4, further comprising: an electronic-programmable vacuum regulator configured to access the vacuum path. 6. The sub-atmospheric wound-care system of claim 4, wherein the pressure regulating sensor is a first pressure regulating sensor, the system further comprising: a second pressure regulating sensor disposed in said tube for taking indirect measurements of the negative pressure at the wound. 7. The sub-atmospheric wound-care system of claim 1, wherein the porous dressing is configured to be sealed airtight over the wound. 8. The sub-atmospheric wound-care system of claim 1, wherein the porosity of the porous dressing is greater than 60%. 9. The sub-atmospheric wound-care system of claim 1, wherein the pressure regulating sensor is located between the porous dressing and the wound. 10. The sub-atmospheric wound-care system of claim 4, wherein the pressure regulating sensor is located between the wound interface chamber and the wound. 11. The sub-atmospheric wound-care system of claim 2, wherein the vacuum motor is battery operated. 12. The sub-atmospheric wound-care system of claim 1, further comprising: a first alarm configured to indicate that a predefined flow rate has been reached as detected by said flow rate meter;fluid level detection means for detecting a fluid level in the collection canister; anda second alarm configured to indicate that a predefined fluid level has been reached in the collection canister fluid level detection means. 13. The sub-atmospheric wound-care system of claim 1, wherein the pressure regulating se... | Mary Beth Kelley | A61M027/00
US08738112 | B2 | 2014-05-27 | Electrocardiograph monitoring device and connector | 2008-05-07 | The present invention relates to electrocardiography and to electrode arrangements used in electrocardiographic monitoring devices, and is more particularly related to a pad or patch containing said electrodes which may be used to passively and non-invasively monitor electrical activity generated by a patient's heart from the surface of that patient's chest, and to a connector which allows for fast and simple connection between the pad containing said electrodes and the devices and equipment typically used to monitor and view electrocardiographic information. | 1. An EKG patch, comprising: a base layer; andfour electrodes held in a fixed configuration with respect to said base layer, said four electrodes being capable of receiving an electrical signal generated by the myocardium of a human heart,wherein said base layer is configured to contact the skin of a human patient and attach thereto,wherein said four electrodes are also configured to contact the skin of a human patient concurrently with said base layer, andwherein said four electrodes are arranged in a square configuration and are equidistantly spaced about 5.5 cm apart from one another along the periphery of the square configuration. 2. An EKG patch according to claim 1, further comprising a cable attachment configured to electrically connect said four electrodes to a single cable. 3. An EKG patch according to claim 1, wherein dimensions of said EKG patch are within a range of about 5.5 cm×5.5 cm square to 16 cm×16 cm square. 4. An EKG patch according to claim 1, wherein said base layer comprises an adhesive material on a surface thereof for attaching said EKG patch to a patient. 5. An EKG patch according to claim 1, wherein said plurality of electrodes comprise a layer of electrode gel thereon. 6. An EKG patch according to claim 1, wherein said base layer has a shape suitable to facilitate at least one of positioning, attaching and removing said EKG patch from a patient. 7. The EKG patch of claim 1, wherein said EKG patch is constructed from nonmagnetic materials such that said EKG patch is suitable for use in conjunction with magnetic fields in a MRI system. 8. The EKG patch of claim 1, wherein said EKG patch comprises perforations to allow at least one of said plurality of electrodes to be ripped apart from a remaining portion of said EKG patch. 9. An EKG patch according to claim 1, further comprising a second layer coupled to said base layer. 10. An EKG patch according to claim 9, wherein said second layer comprises an adhesive layer. 11. The EKG patch of claim 1, wherein the base layer comprises a flexible cloth material capable of conforming to the contours of a human patient. 12. The EKG patch of claim 1, wherein the base layer is configured to be placed along the midline of the sternum of the human patient approximately 2 to 4 cm below the suprasternal notch. 13. An EKG patch, consisting essentially of: a base layer comprising a flexible cloth material capable of conforming to the contours of a human patient;an adhesive layer coupled to said base layer, wherein said adhesive layer is configured to contact the skin of a human patient and attach thereto; andfour electrodes held in a fixed configuration with respect to said base layer, said four electrodes being capable of receiving an electrical signal generated by the myocardium of a human heart,wherein said four electrodes are also configured to contact the skin of a human patient concurrently with said adhesive layer, andwherein said four electrodes are arranged in a square ... | Cardiac Lead Technologies, LLC | A61B005/0408 A61B005/0416
US08463344 | B2 | 2013-06-11 | Antigen monitoring system | 2007-06-22 | A method for detecting cancer in a subject includes administering polysilicon mirrors to the subject, transmitting near infrared light through subject's skin, receiving light which is reflected from the polysilicon mirrors though the subject's skin, converting received light into a digital signal and calculating a level of CEA in the subject's blood from the digital signal. | 1. A method for detecting or monitoring cancer in a subject, comprising: administering polysilicon mirrors to random locations within the bloodstream of a subject orally or by injection;transmitting near infrared light through the subject's skin;detecting light which is reflected from the polysilicon mirrors though the subject's skin;converting received light into a digital signal; andcalculating a level of CEA in the subject's blood from the digital signal. 2. The method of claim 1 wherein differences in the level of CEA in a subject's blood are determined based on previously stored measurements and currently taken measurements. 3. The method of claim 1 wherein the level of CEA is stored. 4. The method of claim 1 wherein the level of CEA is transmitted to a third party. 5. The method of claim 4 wherein the level of CEA is transmitted to the third party via an Internet. 6. The method of claim 1 wherein an intensity of light reflected from said polysilicon mirrors varies according to an amount of antigen markers in the subject's blood. | - | A61B005/1455
US08444559 | B2 | 2013-05-21 | Skin impedance matching system and method for skin/electrode interface | 2008-05-02 | The present invention relates to a system for measuring the input impedance of a skin/electrode interface and selectively modifying the input impedance of the monitoring circuit to match the measured input impedance. More particularly, a simplified method for correcting for input impedance mismatch between electronic monitoring circuitry and the skin/electrode interface. In accordance with one embodiment of the invention, an input impedance measuring circuit will interface with a microprocessor and a reconfigurable switch network to select the input impedance of the electronic monitoring circuitry, thus eliminating the impedance sensitivity of EMG or EKG instruments. | 1. A system for matching skin impedance of a patient, comprising: a first module adapted to determine the skin impedance of the patient by sensing an input impedance from the patient through electrodes, and amplifying and digitizing the input impedance;a resistor ladder network communicably coupled to the first module;a microprocessor configured to analyze the input impedance and generate a series of control signals to direct the resistor ladder network to match the skin impedance of the patient; anda second module adapted to sense electrical signals from the patient through electrodes coupled to the patient's skin, wherein the electrical signals are sensed with the resistor ladder network matching the skin impedance of the patient. 2. The system as set forth in claim 1, wherein the first and second modules use identical electrodes. 3. The system as set forth in claim 1, wherein the first module comprises a sensing resistor, an amplifier, and an analog to digital converter. 4. The system as set forth in claim 1, further comprising a pair of switches controlled by the microprocessor for switching between the first and second modules. 5. The system as set forth in claim 1, wherein the resistor ladder network is digitally configured by the microprocessor through the control signal. 6. The system as set forth in claim 1, wherein the second module provides the electrical signals to an external medical device. 7. The system as set forth in claim 6, wherein the electrical signals are calibrated by the second module using the resistor ladder network. 8. The system as set forth in claim 6, wherein the external medical device includes an electromyography (EMG) monitor. 9. A method for matching skin impedance of a patient, comprising: applying at least one pair of electrodes to a skin of the patient, and coupling the at least one pair of electrodes to an impedance measurement module;switching microcontroller activated switches in the impedance measurement module to a matching mode;determining the skin impedance of the patient by measuring a voltage drop between the at least one pair of electrodes;switching a resistor ladder network in the impedance measurement module to match the determined skin impedance of the patient;switching the microcontroller activated switches to a sensing mode;measuring voltage inputs from the skin of the patient; andcalibrating the voltage inputs using the resistor ladder network to obtain calibrated inputs. 10. The method as set forth in claim 9, wherein the determining and measuring steps are performed using identical electrodes. 11. The method as set forth in claim 9, further comprising providing the calibrated inputs to an external medical device. 12. The method as set forth in claim 11, wherein the external medical device includes an electromyography (EMG) monitor. 13. The method as set forth in claim 11, wherein the steps of switching microcontroller activated switches to a matching mode and determining the skin ... | Reproductive Research Technologies, LP | A61B005/053
US08600681 | B2 | 2013-12-03 | Method and apparatus for providing data processing and control in a medical communication system | 2008-05-14 | Methods and apparatus for providing data processing and control for use in a medical communication system are provided. | 1. A computer-implemented method to determine a composite sensitivity of an in vivo analyte sensor, comprising: receiving, at one or more processors, a calibration parameter to calibrate the in vivo analyte sensor;determining, using the one or more processors, a sensitivity value associated with the received calibration parameter;retrieving, using the one or more processors, a prior sensitivity value associated with the in vivo analyte sensor;comparing, using the one or more processors, the determined sensitivity value and the retrieved prior sensitivity value;determining, using the one or more processors, if the comparison of the determined sensitivity value and the retrieved prior sensitivity value is within a predetermined range; anddetermining, using the one or more processors, a composite sensitivity for the in vivo analyte sensor based on one or more of the received calibration parameter, the determined sensitivity value and the retrieved prior sensitivity value, if the comparison of the determined sensitivity value and the retrieved prior sensitivity value is within the predetermined range. 2. The method of claim 1 wherein the calibration parameter includes a blood glucose value. 3. The method of claim 1 wherein the retrieved prior sensitivity value is associated with a prior calibration parameter used to calibrate the in vivo analyte sensor. 4. The method of claim 3 wherein the prior calibration parameter includes a blood glucose value. 5. The method of claim 1 wherein determining the composite sensitivity includes applying a first weighted parameter to the determined sensitivity value and applying a second weighted parameter to the retrieved prior sensitivity value. 6. The method of claim 5 wherein the first weighted parameter and the second weighted parameter are different. 7. The method of claim 5 wherein the first weighted parameter and the second weighted parameter are substantially the same. 8. The method of claim 5 wherein the first weighted parameter and the second weighted parameter are time based. 9. The method of claim 8 wherein the prior sensitivity value associated with the in vivo analyte sensor is based on a prior calibration parameter used to calibrate the in vivo analyte sensor prior to a predetermined time period of receiving the calibration parameter. 10. The method of claim 8 wherein the first weighted parameter is associated with a current in vivo analyte sensor calibration event, and the second weighted parameter is associated with a prior in vivo analyte sensor calibration event. 11. The method of claim 1 wherein the in vivo analyte sensor comprises at least one working electrode including an analyte-responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode, and further wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer. 12. An apparatus, co... | Abbott Diabetes Care Inc. | G01N033/48 A61B005/00
US08805508 | B2 | 2014-08-12 | Collecting activity data for evaluation of patient incontinence | 2007-05-30 | Systems and methods for determining whether an involuntary voiding event was attributable to stress or urge incontinence include determining an activity level of a patient that coincides with the occurrence of the involuntary voiding event or the activity level within a certain time range of the involuntary voiding event. Patient activity data may be collected via a signal that varies as a function of patient activity. The signal may be generated with one or more sensors that detect motion, such as an accelerometer or a piezoelectric crystal, and/or one or more sensors that monitor a physiological parameter of the patient that varies as a function of patient activity, such as heart rate, respiratory rate, electrocardiogram morphology, respiration rate, respiratory volume, core temperature, muscular activity level or subcutaneous temperature of the patient. | 1. A method comprising: generating, with a sensor, a signal that varies as a function of a physical activity engaged in by a patient, wherein the physical activity engaged in by the patient applies pressure to a bladder of the patient;receiving an indication of an occurrence of an involuntary voiding event; anddetermining, with a processor, whether the involuntary voiding event was attributable to stress incontinence or urge incontinence based on the signal,wherein the indication is separate from the signal used by the processor to determine whether the involuntary voiding event was attributable to stress incontinence or urge incontinence. 2. The method of claim 1, wherein the signal comprises a signal generated by at least one of an accelerometer, a bonded piezoelectric crystal, a mercury switch or a gyro. 3. The method of claim 1, wherein the signal is indicative of a physiological parameter measurement of the patient. 4. The method of claim 3, wherein the physiological parameter comprises at least one of a heart rate, respiratory rate, electrocardiogram morphology, respiration rate, respiratory volume, core temperature, muscular activity level or subcutaneous temperature of the patient. 5. The method of claim 3, further comprising determining an activity level of the patient by comparing the physiological parameter measurement to at least one threshold. 6. The method of claim 3, wherein the physiological parameter comprises at least one of a heart rate, respiratory rate, electrocardiogram morphology, respiration rate, respiratory volume, core temperature, or subcutaneous temperature of the patient. 7. The method of claim 5, wherein determining the activity level of the patient by comparing the physiological parameter measurement to the at least one threshold comprises determining whether the activity level of the patient was relatively high or relatively low based on comparing the activity level to the threshold, the method further comprising associating the determined activity level with the involuntary voiding event, wherein determining whether the involuntary voiding event was attributable to stress incontinence or urge incontinence based on the signal comprises determining whether the involuntary voiding event was attributable to stress incontinence or urge incontinence based on whether the activity level of the patient associated with the involuntary voiding event was relatively high or relatively low. 8. The method of claim 7, wherein determining whether the involuntary voiding event was attributable to stress incontinence or urge incontinence comprises determining the involuntary voiding event was attributable to stress incontinence in response to determining a relatively high activity level is associated with the involuntary voiding event and determining the involuntary voiding event was attributable to urge incontinence in response to determining a relatively low activity level is associated with the involuntary voiding event. ... | Medtronic, Inc. | A61N001/00
US08932217 | B2 | 2015-01-13 | Vital signs monitor | 2006-01-13 | A multi-parametric vital signs monitoring device configured for use as an ambulatory and a bedside monitor wherein the device can be patient-wearable and is battery powered. The monitoring device can be used with a charging cradle to provide power to the device in lieu of the battery as a power source for bedside applications, in which the cradle further serves as an intermediary device to enable a data link with a PC or other peripheral device. The monitoring device can include a wireless radio to enable bi-directional transfer of patient-related data to a separate remote station. | 1. A multi-parametric patient monitoring device comprising: a device housing that retains a portable power supply, a CPU, and a display for permitting the monitoring device to operate in a first stand-alone mode wherein a plurality of physiological sensors are attached to said device housing and in which parameter data received from said sensors is processed by said CPU within said device housing;a wireless transceiver enabling said monitoring device to operate in a second mode in which patient-related data is transmitted to a remote station and in which said monitoring device is further configured to be separately operated in a separate charging mode wherein said device can still operate simultaneously in the second remote wireless transmitting mode while in said charging mode, said device further including a wired transceiver to permit patient-related data transfer remotely in said charging mode, and in which said device simultaneously transfers data over each of said wired and said wireless transceivers,wherein a plurality of physiological parameters are continuously monitored by said monitoring device, said monitoring device further comprising a pulse oximeter assembly in which said monitoring device is programmed to continuously monitor at least one of said plurality of physiological parameters while said pulse oximeter assembly is enabled to selectively operate in each of a continuous monitoring mode and a spot-check monitoring mode and in which the at least one of said plurality of physiological parameters is continuously monitored independent of the monitoring mode selected for the pulse oximeter assembly. 2. A device according to claim 1, in which the monitoring device is placed in a charging cradle in the charging mode, the charging cradle including a data port to permit communication between said patient monitoring device and at least one peripheral device. 3. A device as recited in claim 1, wherein said monitoring device is patient wearable in said first and second modes. 4. A device as recited in claim 2, in which the peripheral device is a large display wherein patient-related data is further transmitted in real time to said display in said charging mode using said data port. 5. A device as recited in claim 2, in which the peripheral device is a computing device. 6. A device as recited in claim 5, in which the computing device sends a configuration file to the patient monitoring device providing default settings for said monitoring device. 7. A device as recited in claim 5, in which said computing device automatically receives stored data from said patient monitoring device when said monitoring device is attached to said charging cradle. 8. A device as recited in claim 2, wherein said CPU includes memory for storing and trending patient-related data, said device further including a user interface that permits snapshots of data to be selectively taken, wherein all trended data and snapshot data is automatically transmitted whe... | Welch Allyn, Inc. | G06F019/00 A61B005/00 A61B005/024 A61B005/0205 A61B005/0424 A61B005/1455
US08657752 | B2 | 2014-02-25 | Biologic information detecting apparatus | 2005-10-04 | In measurement requiring application of pressure to a tissue of a living body such as blood pressure measurement, noise due to vibration tends to occur. It is difficult to accurately measure a pulse wave and a blood pressure value. It is also difficult to measure blood pressure in life activities or to measure blood pressure at intervals or continuously where a tonometer is always attached. There is consequently a problem of holding a biologic information detecting apparatus. The present invention solves the problems by providing an easy-to-wear biologic information detecting apparatus for stably detecting biologic information. The biologic information detecting apparatus includes a sensor for detecting biologic information in a pair of arms connected via a spindle, and the sensor is tightly attached to a projecting part in a living body, particularly, a tragus of an auricle. | 1. A biologic information detecting apparatus comprising: a pair of arms that face each other;a spindle connecting the pair of arms at one end of each of the pair of arms;a distance varying mechanism provided for the spindle and adjusting an interval between other ends of the pair of arms; anda cuff configured to press a living body, which is attached to the other end of at least one of the pair of arms and attached on facing sides of the pair of arms, characterized in thata cuff rotating mechanism is provided for adjusting an angle of the cuff relative to the arm to which the cuff is attached in order to vary a direction in which the cuff is configured to come into contact with the living body;an axis of rotation of the cuff rotating mechanism extends in a same direction as an axis of rotation of the distance varying mechanism;a rotating mechanism is obtained by breaking a tip of the spindle and moves the spindle in a rotating direction; andthe pair of arms is configured to sandwich a tragus as a projecting part of an auricle in the living body by an arm configured to be disposed on one side face side of the tragus and an arm configured to be disposed on an other side face side of the tragus. 2. The biologic information detecting apparatus according to claim 1, further comprising a sensor that includes a light emitting element for emitting output light to the living body, and a light receiving element for receiving the output light from the light emitting element, which is scattered by the living body, or passed through the living body. 3. The biologic information detecting apparatus according to claim 1, further comprising an ear attachment configured to be placed around the base of the auricle. 4. The biologic information detecting apparatus according to claim 1, further comprising: a bridge connecting the pair of arms configured to be attached to one of right and left auricles of the living body and the other auricle to which the pair of arms is not attached; anda power supply unit disposed in a midpoint of the bridge and driving a sensor. 5. The biologic information detecting apparatus according to claim 1, further comprising: an ear attachment configured to be hung from the base of the auricle; and a power supply unit disposed in the ear attachment. | Nippon Telegraph and Telephone Corporation | A61B005/02 A61B005/00
US07890162 | B2 | 2011-02-15 | Method and apparatus for sensing improvement using pressure data | 2007-04-30 | Method and apparatus for sensing improvement using pressure data. The method and apparatus may be used in an implantable medical device to confirm that an EGM event signifies a true mechanical cardiac activity and not just electrical oversensing. The mechanical activity may be used to create a mechanical marker channel in the implantable medical device. | 1. A method of confirming that an EGM deflection represents an R-wave comprising: sensing an EGM via an implantable medical device (IMD);sensing an intracardiac pressure via the IMD;determining a dP/dt of the sensed intracardiac pressure;detecting an electrical event using the EGM;opening a time window after detecting the electrical event;determining a maximum positive dP/dt of the intracardiac pressure sensed during the open time window;detecting the presence of a mechanical event responsive to the maximum positive dP/dt being greater than a first threshold; anddetermining that the electrical event includes an R-wave responsive to the presence of the mechanical event;determining a minimum negative dP/dt of the intracardiac pressure sensed during the open time window; anddetecting the presence of an oversensed T-wave or noise using the minimum negative dP/dt, the oversensed T-wave or noise being deemed present when the mechanical event is deemed absent and an absolute value of the minimum negative dP/dt is greater than a second threshold. 2. A method according to claim 1, further comprising withholding defibrillation when the EGM indicates a tachyarrhythmia and the oversensed T-wave or noise is deemed present. 3. A method according to claim 1, further comprising detecting the presence of an R-wave with no corresponding mechanical event, the R-wave without a corresponding mechanical event being deemed present when the mechanical event is deemed absent and an absolute value of the minimum dP/dt is not greater than a second threshold. 4. A method according to claim 1, further comprising recording the presence of the mechanical event by the device on a marker channel. 5. A method according to claim 1, wherein the first threshold is approximately 100 mmHg/s. 6. A method according to claim 1, wherein the time window is between approximately 400 and approximately 500 milliseconds or equal to the duration between two sequentially detected R-waves, which ever is smallest. | Medtronic, Inc. | A61B005/0452
US08311619 | B2 | 2012-11-13 | Detection of drastic blood pressure changes | 2006-08-10 | A cardiac-activity based prediction of a rapid drop in a patient's blood pressure during extracorporeal blood treatment is disclosed. A proposed alarm apparatus includes a primary beat morphology analysis unit bank of secondary analysis units and an alarm generating unit. The primary beat morphology analysis unit discriminates heart beats in a received basic electrocardiogram signal, classifies each beat into one out of at least two different beat categories, and associates each segment of the signal with relevant event-type data. The event-type data and the basic electrocardiogram signal together form an enhanced electrocardiogram signal, based upon which the primary beat morphology analysis unit determines whether one or more secondary signal analyses should be performed. Depending on the enhanced electrocardiogram signal's properties, the bank of secondary analysis units performs none, one or more of up to at least two different types of secondary analyses, and for each analysis performed produces a respective test signal. The alarm generating unit receives the test signals, and triggers an alarm signal indicative of an estimated rapid blood pressure decrease, if at least one alarm criterion is fulfilled. | 1. An alarm apparatus for predicting a rapid blood pressure decrease in a patient undergoing extracorporeal blood treatment, the apparatus comprising: an input interface configured to receive a basic electrocardiogram signal of the patient;a primary beat morphology analysis unit configured to: discriminate heartbeats in the basic electrocardiogram signal, said basic electrocardiogram signal having a plurality of signal segments,classify each discriminated heartbeat into one out of at least two different beat categories of which one beat category represents a beat whose morphology is typical for the patient,associate each of the plurality of signal segments of the basic electrocardiogram signal with event-type data reflecting one of said at least two different beat categories represented by the signal during that segment, the event-type data and the basic electrocardiogram signal together forming an enhanced electrocardiogram signal, anddetermine a number of secondary signal analyses to be performed in respect of the basic electrocardiogram signal;a bank of secondary analysis units configured to: receive and use the enhanced electrocardiogram signal to perform up to at least two different types of said secondary analyses in response to control instructions from the primary beat morphology analysis unit, and for each analysis performed, produce a respective test signal,wherein the bank of secondary analysis units comprises at least two different secondary analysis units selected from a group of: a heart-rate-variability analysis unit configured to produce a first test signal describing a ratio between a low-frequency band and a high-frequency band of a power spectrum representation of the basic electrocardiogram signal,an ectopic-beat count unit configured to produce a second test signal describing an intensity of ectopic beats, anda heart-rate-turbulence analysis unit configured to produce at least one third test signal describing at least one heart-rate-turbulence parameter; andan alarm generating unit configured to: receive the test signals,investigate whether any of the test signals fulfill at least one alarm criterion, and if so,trigger an alarm signal indicative of an estimated rapid blood pressure decrease, wherein the alarm generating unit is configured to trigger the alarm signal: if the first test signal fulfills a first alarm criterion, orif the second test signal fulfills a second alarm criterion, orif the at least one third test signal fulfills at least one third alarm criterion, wherein the at least one third test signal comprises at least one of: a first parameter expressing a turbulence-onset measure reflecting a relative change in RR-intervals of the basic electrocardiogram signal during a period following a particular ectopic beat,a second parameter expressing a turbulence-slope measure reflecting a rise rate of the RR-intervals during a period following a particular ectopic beat, anda pair of third quantities expressing geometric... | Gambro Lundia AB | A61B005/04 A61B005/02
US08249698 | B2 | 2012-08-21 | General diagnostic and real-time applications of discrete hermite functions to digital data | 2008-04-04 | General diagnostic and real-time application of digital Hermite functions allows features to be extracted from a measured signal through expansion of the measured signal. Specifically, the digital Hermite functions represent the shape of the measured signal in a set of vectors derived from a symmetrical tridiagonal matrix. This allows for efficient computation of the Hermite expansion coefficients, in real-time, to represent the expanded signal. The signal expansion also allows any artifacts, such as noise, to be isolated and removed, allowing the underlying signal of interest to be revealed. | 1. A method for identifying a biomedically notable digital signal of interest comprising the steps of: (a) obtaining a digital electronic signal; (b) isolating an interval of interest including the notable digital signal; (c) applying a dilated digital Hermite transform to the interval of interest to represent the shape of the signal in the interval of interest in a set of orthogonal vectors; and (d) modifying the transform to identify an underlying digital signal which is the notable digital signal of interest. 2. The method of claim 1 wherein the set of vectors are derived from a symmetrical tridiagonal matrix. 3. The method according to claim 2 wherein the notable digital signal of interest is in the form of at least one of an EEG, an ECG, an EMG, and an EOG. 4. The method according to claim 1 wherein the modification of step (d) is based on shape content of the underlying signal as revealed by the digital Hermite transform. 5. The method according to claim 1 wherein the steps are performed in real-time. 6. The method according to claim 1 wherein the modification of step (d) includes the removal of noise to isolate the notable digital signal of interest. 7. The method according to claim 1 wherein the notable digital signal of interest is part of an EEG and the modification of step (d) includes the removal of BCG artifacts. 8. The method according to claim 1 further comprising the steps of: providing a computer system having at least one power source, at least one input device, at least one display, and at least one memory device;calculating the dilated digital Hermite transform using the computer system; andmodifying the notable signal using the dilated digital Hermite transform calculated at the calculating step. 9. A method for monitoring/detecting abnormalities in an ECG, the method comprising the steps of: (a) gathering at least one ECG; (b) subjecting the at least one ECG to a QRS detection algorithm in order to scan for R-peak location; (c) calculating a digital Hermite transform corresponding to the individual ECG complexes from each individual ECG to represent the shape of the ECG in a set of orthogonal vectors; and (d) subjecting the Hermite transform coefficients to a Neural Network in order to determine the present and/or absence of ECG abnormalities. 10. The method according to claim 9 wherein the digital Hermite transform coefficients of step (c) are calculated by a simple dot product. 11. The method according to claim 9 wherein the steps are performed in real-time. 12. The method according to claim 9 wherein the ECG abnormalities being monitored are associated with cardiac ischemia. 13. The method according to claim 12 wherein the Neural Network outputs the presence or absence of ST segment changes, T wave changes, beat classification, ischemia, or a combination of one or more thereof. 14. The method according to claim 9 further comprising the steps of: providing a computer system having at least one power source, at... | The University of Akron | A61B005/0402 A61B005/0476
US08714153 | B2 | 2014-05-06 | Method for selecting a device adapted to treat disordered breathing | 2008-03-25 | A method for selecting a device adapted to treat disordered breathing includes determining a number of characteristics associated with at least a portion of the airway of a patient, comparing each of at least some of the number of characteristics to a corresponding threshold, and selecting such a device in response to the comparison. An alternative method for selecting a device adapted to treat disordered breathing includes delivering a flow of breathing gas to the airway of a patient, determining a minimum effective therapeutic pressure associated with such patient, and selecting, relative to the minimum effective therapeutic pressure, an oral appliance or one of a number of patient interface devices. | 1. A method for selecting a device adapted to treat sleep disordered breathing, the method comprising: fitting a patient with a first device, the first device being adapted to measure one or more resistance characteristics associated with at least a portion of an airway of the patient;determining from the first device a number of the resistance characteristics associated with at least the portion of the airway of the patient;comparing each of at least some of the number of the resistance characteristics to a corresponding threshold; andselecting a second device in response to the comparing step, the second device being adapted to treat sleep disordered breathing,wherein selecting the second device comprises choosing one device from among a group consisting of at least two of a mandibular advancement device, nasal prongs, a nasal mask, a full nasal/oral mask structured to cover the patient's mouth and nose, and a total face mask,wherein each device of the group is chosen over another device in the group based at least in part on the comparison of the resistance characteristics to the threshold. 2. The method of claim 1, wherein determining the number of resistance characteristics associated with at least of portion of the airway comprises determining at least one pressure/flow attribute associated with a portion of the airway between the nares and the laryngopharynx of the patient. 3. The method of claim 2, wherein the first device is a nasal mask, and wherein determining the at least one pressure/flow attribute associated with the portion of the airway between the nares and the laryngopharynx of the patient comprises: supplying a flow of gas to the patient via the nasal mask; andmeasuring the pressure/flow attributes of the flow of gas entering the patient's nares and exiting the patient's mouth. 4. The method of claim 1, wherein determining the number of resistance characteristics associated with at least of the portion of the airway comprises determining a nasal resistance associated with a portion of the airway between the nares and the laryngopharynx of the patient. 5. The method of claim 4, wherein determining the nasal resistance comprises: providing a breathing gas at a number of different pressures;measuring a flow rate of the breathing gas at each of at least some of the number of different pressures;generating a pressure-flow curve from the measured flow rates and the number of different pressures; andcorrelating the pressure-flow curve to the nasal resistance associated with the portion of the airway. 6. The method of claim 4, wherein determining the nasal resistance comprises: inserting a manometer into a closed mouth of the patient; andmeasuring the difference between the pressure present at the patient's nares and the pressure present at the patient's mouth when the patient breathes though their nose. 7. The method of claim 1, wherein comparing each of at least some of the number of resistance characteristics to the correspond... | RIC Investments, LLC | A62B007/00 A61M016/00
US08784338 | B2 | 2014-07-22 | Electrical means to normalize ablational energy transmission to a luminal tissue surface of varying size | 2008-06-20 | Methods and devices for measuring the size of a body lumen and a method for ablating tissue that uses the measurement to normalize delivery of ablational energy from an expandable operative element to a luminal target of varying circumference are provided. The method includes inserting into the lumen an expandable operative element having circuitry with resistivity or inductance that varies according to the circumference of the operative element, varying the expansion of the operative element with an expansion medium, measuring the resistivity of the circuitry, and relating the resistivity or inductance to a value for the circumference of the operative element. In some embodiments the sizing circuit includes a conductive elastomer wrapped around the operative element. Other embodiments of the method apply to operative elements that include an overlapping energy delivery element support in which the overlap varies inversely with respect to the state of expansion, and which is configured with sizing electrodes that sense the amount of the overlap. | 1. A device for measuring the size of a body lumen comprising: an expandable operative element including an inflatable balloon and an ablation energy delivery element support arranged around the inflatable balloon, the support having a longitudinal first edge and a second longitudinal edge that overlap each other to form an area of mutual overlap; anda sensing circuitry whose resistivity varies according to the area of mutual overlap. 2. The device of claim 1, wherein the expandable operative element includes an expandable balloon. 3. The device of claim 1, wherein the operative element further includes an ablative energy delivery element. 4. The device of claim 3, wherein the ablative energy delivery element includes any of a radiofrequency electrode, a radio frequency electrode array, or solid-state circuitry. 5. The device of claim 1, wherein the circuitry comprises a portion of a band of conductive elastomer wrapped around the operative element such that when the operative element is contracted a length of the conductive elastomer band is also contracted, and when the operative element is expanded, the length of the conductive elastomeric band is stretched. 6. The device of claim 1, wherein the support is circumferentially expandable by the balloon such that the area of mutual overlap is inversely related to an amount of expansion of the balloon. 7. The device of claim 1, wherein the circuitry includes sensing elements on both longitudinal edges of the ablational energy delivery element support in the area of mutual overlap, such elements configured to connect across the area of mutual overlap to form a circuit with a particular resistivity, the elements further configured such that the connection formed between sensing elements varies according to an amount of mutual overlap of the two longitudinal edges. 8. The device of claim 7, wherein the first edge includes a single sensing element and the second edge includes a plurality of spaced-apart sensing elements, the particular element among the plurality of elements on the second edge that connects to the element on the first edge varying according to an amount of mutual overlap of the two edges, the resistivity of circuit thus formed varying according to the particular element on the second edge that is included in the circuit. 9. The device of claim 7, wherein the first edge includes a single sensing element and the second edge includes a plurality of closely-spaced sensing elements, the elements configured such that the element on the first edge connects with one of the plurality of the elements on the second edge or with a pair of adjacent elements on the second edge, the resistivity of circuit formed varying according to which of the one or which adjacent pair of the elements on the second edge is included in the circuit. 10. The device of claim 7, wherein the first edge includes a single sensing element and the second edge includes an elongated sensing element, the elements conf... | Covidien LP | A61B005/00 A61B005/103 A61B005/107
US08287467 | B2 | 2012-10-16 | Puncturing system | 2008-03-28 | The invention relates to a puncturing system for generating a puncture wound for obtaining a sample of a body fluid, comprising a press-on part to be pressed onto a body part in which a puncture wound is to be generated, a triggering means, by the actuation of which a user can trigger a puncturing motion of a puncturing element after the press-on part is pressed on, and a testing facility for determining at least one test parameter on which a sample-obtaining probability depends. A securing facility is also provided that, in a locked state, locks the triggering means such that no puncturing motion can be triggered, and, in a triggering state, releases the triggering means such that a puncturing motion can be triggered by actuation thereof, whereby the securing facility is transitioned from the locked state to the triggering state by the testing facility when the test parameter determined by the testing facility meets defined minimum requirements. A signaling facility is also provided for signaling a transition of the securing facility to the triggering state and/or for signaling that the securing facility is in the triggering state. The invention also relates to a method for preparing a puncturing system for generating a puncture wound. | 1. A puncturing system for generating a puncture wound for obtaining a sample of a body fluid, the puncturing system comprising: a housing which contains components of the puncturing system;a puncturing element movable relative to the housing in a puncturing direction to generate a puncture wound in a body part;a press-on part configured to be pressed onto the body part and having an opening for the lancet tip to puncture the body part;a trigger, a manual actuation of which causes the puncturing element to move in the puncturing direction;a testing facility for determining at least one test parameter on which a sample-obtaining probability depends;a securing facility having a locked state and a triggering state, wherein in the locked state the securing facility prevents the trigger from being actuated, and in the triggering state, the securing facility releases the trigger to allow the manual actuation, the securing facility being transitioned from the locked state to the triggering state when the at least one test parameter determined by the testing facility meets defined minimum requirements, wherein, the puncturing element is movable in the puncturing direction to generate a puncture wound only if (i) the testing facility determines a test parameter that indicates a successful puncture can be executed and if (ii) the trigger is subsequently actuated; anda signaling facility that signals a transition of the securing facility from the locked state to the triggering state or signals that the securing facility is in the triggering state, wherein the signaling facility comprises a device which produces an audible signal or comprises a visual indicator disposed on or visible from the exterior of the housing. 2. The puncturing system of claim 1, wherein the testing facility comprises a sensor configured to determine the blood supply status of a body part touching against the press-on part. 3. The puncturing system of claim 1, wherein the testing facility comprises a pressure sensor configured to determine whether a pressure that exceeds a predetermined minimum pressure bears on the press-on part. 4. The puncturing system of claim 3, wherein the pressure sensor comprises a snap-action element that snaps from a first configuration to a second configuration when a critical minimum pressure is applied thereto. 5. The puncturing system of claim 4, wherein the snap-action element comprises a metal sheet. 6. The puncturing system of claim 3, wherein the pressure sensor comprises a spring element having a spring force against which the press-on part can be shifted with respect to a housing part, further wherein, the press-on-part reaches an end position relative to the housing part when at least the minimum pressure is applied. 7. The puncturing system of claim 6, wherein the press-on part is rigidly connected to a puncturing element drive. 8. The puncturing system of claim 7, further comprising a marker that indicates when the press-on part is in the... | Roche Diagnostics Operations, Inc. | A61B005/00 A61B017/14 A61B017/32
US08414503 | B2 | 2013-04-09 | Methods and apparatus for lancet actuation | 2007-03-16 | A lancet driver is provided wherein the driver exerts a driving force on a lancet during a lancing cycle and is used on a tissue site. The driver comprises of a drive force generator for advancing the lancet and a processor coupled to the drive force generator capable of changing the direction and magnitude of force exerted on the lancet during the lancing cycle. The driver further includes a human interface on the housing providing at least one output for communicating with the patient. | 1. A body fluid sampling device for use on a tissue site, comprising: a housing;an electronic drive force generator configured to be coupled to a penetrating member;a multiple penetrating member cartridge;a processor coupled to the drive force generator capable of changing a direction and magnitude of force exerted on a penetrating member during a lancing cycle;a feedback control coupled to the processor, the feedback control providing information to the processor based on at least one of penetrating member, displacement velocity and acceleration;a human interface on the housing;a human interface processor; andwherein the housing and all elements therein have a combined weight of less than about 0.5 lbs and the electronic drive force generator uses electricity to create a controllable electromagnetic field for actuating a penetrating member. 2. The device of claim 1, wherein the electronic drive force generator actuates a penetrating member to penetrate to a depth in the tissue site and pause for a controlled dwell time while in the tissue site, the dwell time sufficient to draw body fluid toward a wound channel created by the penetrating member. 3. The device of claim 1, wherein the electronic drive force generator uses electricity and is configured to hold a penetrating member in the tissue site at a fixed position when electric current is turned off, allowing for unlimited dwell time in the tissue site. 4. The device of claim 1, further including: a penetrating member coupler for removably coupling a penetrating member to the drive force generator. 5. The device of claim 1, wherein the human interface is selected from: an LED, an LED digit display, an LCD display, a sound generator, a buzzer, or a vibrating device. 6. The device of claim 1, further comprising an input device on a housing, the input device selected from: one or more pushbuttons, a touch pad independent of the display device, or a touch sensitive screen on the human interface. 7. The device of claim 1, further comprising a data exchange device for coupling the lancing device to support equipment. 8. The device of claim 1, further comprising a data exchange device for coupling the lancing device to support equipment selected from one of the following: personal computer, modem, PDA or a computer network. 9. The device of claim 1, wherein the electronic drive force generator is a linear voice coil drive force generator. 10. A body fluid sampling device for use on a tissue site, comprising: a housing;an electronic drive force generator configured to be coupled to a penetrating member;a multiple penetrating member cartridge;a processor coupled to the drive force generator capable of changing a direction and magnitude of force exerted on a penetrating member during a lancing cycle;a feedback control coupled to the processor, the feedback control providing information to the processor based on at least one of penetrating member, displacement velocity and acceleration;a human i... | Sanofi-Aventis Deutschland GmbH | A61B005/00 A61B017/14 A61B017/32
US08589177 | B2 | 2013-11-19 | Virtual clinic for medical practice | 2008-07-08 | A method is provided for a virtual clinic to establish communications between a user (e.g., a patient) and a specialist (e.g., a physician) in order to facilitate a diagnosis. In various embodiments, a user provides a request for the specialist to a virtual clinic. The request may be a request for medical services identifying medical needs. The user may then receive specialist information from the virtual clinic. The virtual clinic may establish real-time communication between the user and a specialist. The user may receive instructions from the specialist via the virtual clinic to perform a test. The user may then report a test result from the test to the specialist via the virtual clinic. The virtual clinic may receive and store a diagnosis from the specialist. | 1. A method comprising: establishing a virtual clinic on a network, the virtual clinic having working relationships with specialists;establishing a real-time communication between a user and a particular one of the specialists via the virtual clinic;providing user information to the particular one of the specialists via the virtual clinic;providing instructions from the particular one of the specialists to the user via the virtual clinic to perform a test using diagnostic test equipment operatively coupled to a diagnostic center at a remote location away from the user's home and away from a dedicated treatment facility, and the diagnostic center being administered by an administrator who is not a licensed physician;providing a test result to the particular one of the specialists via the virtual clinic; andreceiving and storing a diagnosis from the specialist. 2. The method of claim 1, further comprising storing locations of a plurality of diagnostic centers, storing a list of diagnostic testing equipment present at each different diagnostic center, generating a list of diagnostic centers from the plurality of diagnostic centers based at least in part on proximity to the user, and providing the list of diagnostic centers to the user. 3. The method of claim 2, wherein the list of diagnostic centers is also based at least in part on the request for the particular one of the specialists. 4. The method of claim 1, further comprising establishing a real-time communication between the particular one of the specialists and a consultant via the virtual clinic during the real-time communication between the user and the particular one of the specialists. 5. The method of claim 4, wherein the consultant does not have access to an identity of the user. 6. The method of claim 1, further comprising providing access to the real-time communication between the user and the particular one of the specialists to a third-party. 7. The method of claim 6, wherein the third-party is a trainee. 8. The method of claim 6, wherein the third-party is a peer to the particular one of the specialists. 9. The method of claim 1, wherein the particular one of the specialists is a physician and the user is a patient. 10. The method of claim 1, further comprising receiving the user information and determining the particular one of the specialists based on the user information and specialist information. 11. The method of claim 1, further comprising establishing a real-time communication between the particular one of the specialists and a supervisor via the virtual clinic during the real-time communication between the user and the particular one of the specialists. 12. The method of claim 1, further comprising storing the communication between the user and the particular one of the specialists and enabling access to the stored communication to a third-party. 13. A system comprising: a specialist device for use by a specialist and configured to receive user information from... | - | G06Q050/00 G06Q010/00
US08613703 | B2 | 2013-12-24 | Insertion devices and methods | 2008-05-29 | Devices and methods relating to analyte sensor insertion devices and methods are provided. | 1. A glucose sensor insertion system, the system comprising: a mounting unit configured to mate with an inserter, the inserter comprising a housing, a structure within the housing configured to carry a glucose sensor, and a driver within the housing to drive the glucose sensor at least partially into a user's skin; wherein when the inserter and the mounting unit are mated, the mounting unit causes the structure to be pushed into an interior of the inserter housing;wherein mating the inserter and the mounting unit together arms the system;wherein the mounting unit comprises a mount base and mount arms and the structure within the housing to carry the glucose sensor comprises carrier arms; andwherein the structure within the housing configured to carry the glucose sensor comprises a hook that engages a portion of the housing when the driver is armed. 2. A self arming inserter comprising: an inserter housing to be received in a mounting unit;a carrier supported by the inserter housing;a sensor mounted on the carrier; anda driver which drives the carrier; wherein when the inserter and the mounting unit are mated, the mounting unit causes the carrier to be pushed into an interior of the inserter housing;wherein inserting the inserter housing in the mounting unit arms the driver;wherein the mounting unit comprises a mount base and mount arms and the carrier comprises carrier arms; andwherein the carrier comprises a hook that engages a portion of the housing when the driver is armed. 3. The self arming inserter according to claim 2, further comprising an actuator which actuates the driver and wherein actuating the actuator releases the driver. 4. The self arming inserter according to claim 3, wherein the driver comprises a spring and wherein arming the inserter comprises compressing the spring. 5. The self arming inserter according to claim 2, wherein the mounting unit comprises an inserter hole through which the sensor passes when the carrier is driven by the driver. 6. The self arming inserter according to claim 2, wherein the mounting unit is shiftable with respect to the housing after the mounting unit is inserted into the housing. 7. The self arming inserter according to claim 2, further comprising an inserter tip which supports the sensor on the carrier. 8. The self arming inserter according to claim 2, wherein the sensor attaches to the mounting unit after the driver drives the carrier. 9. The self arming inserter according to claim 2, wherein an adhesive is provided on a surface of the mounting unit. 10. The glucose sensor insertion system of claim 1, wherein the mounting unit comprises an adhesive on a surface of the mounting unit. 11. The glucose sensor insertion system of claim 1, wherein the inserter further comprises an actuator, and wherein actuating the actuator releases the driver to drive the glucose sensor at least partially into the user's skin. 12. The glucose sensor insertion system of claim 1, wherein the driver drives t... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/05
US08406838 | B2 | 2013-03-26 | Apparatus for evaluating biological function, a method for evaluating biological function, a living body probe, a living body probe mounting device, a living body probe support device and a living body probe mounting accessory | 2007-01-22 | The apparatus for evaluating biological function of the present invention has living body probes 1, a behavioral information measuring part 2 and an apparatus body 3, and it utilizes near-infrared spectroscopy to evaluate biological function; apparatus body 3 has a controller 8 for calculating (based on light information from living body probes 1) a variety of parameters derived from two-dimensional diagrams showing relationships between changes in oxyhemoglobin and changes in deoxyhemoglobin and two-dimensional diagrams showing relationships between absolute amounts of oxyhemoglobin and absolute amounts of deoxyhemoglobin, a behavioral information input part for entering behavioral information measured by means of behavioral information measuring part 12, and a display part 10 for performing various types of image displays based on various parameters calculated by means of controller 8 and/or behavioral information entered in the behavioral information input part. | 1. An apparatus for evaluating biological function comprising: a plurality of living body probes provided with light-emitting elements for irradiating light to specified sites of a living body and light-receiving elements for receiving and detecting light exiting the living body;an apparatus body for entering light information detected by said living body probe and behavioral information of said living body measured in response to said living body being given a task to perform, and performing calculation, control and memory operations, and utilizing near-infrared spectroscopy to evaluate biological function, said apparatus body comprising: a controller for calculating, based on light information from said living body probe, a variety of parameters derived from two-dimensional diagrams, wherein each diagram shows at least one of: the relationship between amounts of change in oxyhemoglobin and amounts of change in deoxyhemoglobin is inclined 45 degrees to show the relationship between amounts of change in oxygen saturation and amounts of change in total hemoglobin concentration, andthe relationship between absolute amounts of oxyhemoglobin and absolute amounts of deoxyhemoglobin is inclined 45 degrees to show the relationship between absolute oxygen saturation and absolute total hemoglobin concentration;a behavioral information input part for entering behavioral information of said living body measured in response to said living body being given a task to perform; anda display part for performing various types of image displays based on various parameters calculated by means of said controller and/or behavioral information entered in said behavioral information input part. 2. The apparatus for evaluating biological function according to claim 1, wherein said controller calculates a variety of parameters derived from three-dimensional diagrams, in which a time axis is added to said two-dimensional diagrams, and said display part performs three-dimensional displays based on the variety of parameters calculated by said controller. 3. The apparatus for evaluating biological function according to claim 1, wherein said controller calculates a variety of parameters establishing a relationship between hemoglobin and oxygen exchange in a formula of rotational motion or wave dynamics, by considering oxygen exchange in the hemodynamics as rotational motion or wave motion. 4. The apparatus for evaluating biological function according to claim 1, wherein said controller calculates integral values, corresponding to a behavior time, for light information detected by the living body probes, and, based on the calculated integral values, acquires living body site information that is dependent on said behavior time and/or information on network interdependence between sites of the living body; and said display part displays the living body site information. 5. The apparatus for evaluating biological function according to claim 1, wherein said controller calculate... | - | A61B005/1455
US08160836 | B2 | 2012-04-17 | Wireless transmission of temperature data for a geographic area | 2007-04-17 | A user obtains an individual's body temperature data and transmits the data to a medical monitor (e.g., a medical device) for display. Additional data includes a timestamp and location of the body temperature data. Once the data is transmitted, a user may view the medical monitor for a temperature reading. For example, a doctor may take a patient's temperature and the temperature reading is displayed on a medical monitor. The body temperature data of each patient is detected using a preferred temperature detector, such as a temporal artery thermometer using an arterial heat balance approach. After collecting an individual's body temperature data, the body temperature data can be transferred to a processor. By sending body temperature data for many individuals for a geographic region, the processor can identify a pattern (e.g., a pandemic) in the body temperature data. | 1. A method of identifying a pattern in body temperature data, comprising: using a common temperature detector to detect body temperature data for a plurality of individuals;transmitting from the common temperature detector the body temperature data over a wireless communications path to a processor; anddetermining a pattern in the body temperature data from the plurality of individuals using the processor to indicate a disease outbreak in a geographic area based on a change in percentage of fevers for a time period and using statistical analysis of the temperature data to avoid false positives in indication of the disease outbreak. 2. The method of claim 1 wherein using the common temperature detector to detect body temperature data of a plurality of individuals further comprises scanning across the temporal artery for each individual to obtain body temperature data. 3. The method of claim 1 further comprising storing the body temperature data in a database in the processor. 4. The method of claim 1 wherein transmitting the body temperature data uses a cell phone module. 5. The method of claim 4 wherein the cell phone module includes a Subscriber Identity Module (SIM) card. 6. The method of claim 4 wherein the cell phone module includes a General Packet Radio Service (GPRS). 7. The method of claim 4 wherein the cell phone module transmits in the form of a Short Message Service (SMS), text message, or email message. 8. The method of claim 1 further comprising displaying the pattern to a user of the processor. 9. The method of claim 1 wherein the pattern identifies an epidemic. 10. The method of claim 1 wherein the pattern identifies a pandemic. 11. The method of claim 1 wherein the pattern identifies an outbreak in an identified geographic location. 12. The method of claim 1 further comprising transmitting the body temperature data from multiple locations. 13. The method of claim 1 wherein the pattern identifies an outbreak in multiple geographic locations. 14. A system for identifying a pattern in temperature data, comprising: a common temperature detector to detect body temperature data for a plurality of individuals;a transmitter associated with the common temperature detector that transmits body temperature data from the common temperature detector over a wireless communications path to a processing unit; andthe processing unit that determines a pattern in the body temperature data from the plurality of individuals to indicate a disease outbreak in a geographic area based on a change in percentage of fevers for a time period and uses statistical analysis of the temperature data to avoid false positives in indication of the disease outbreak. 15. The system of claim 14 wherein the common temperature detector scans across the temporal artery to obtain body temperature data. 16. The system of claim 14 wherein the processor stores the body temperature data in a database. 17. The system of claim 14 wherein the transmitter is a cell ... | Exergen Corporation | G01K001/02 G01K003/06 G06F003/00 G06F003/01
US08175687 | B2 | 2012-05-08 | Light source for fluorescence diagnosis and photodynamic therapy | 2006-12-05 | Disclosed herein is a light source device for fluorescence diagnosis and photodynamic therapy comprising: a plurality of light source means having a wavelength range different from each other and providing selected light; a light path coupling means coupling light emitted from the light source means to the incident direction of a light guide; a filtering means selectively transmitting the emitted light based on the wavelength; and a switching device that is directed to an opening/closing means, established at a long distance to be controlled by a remote control device and selecting the kinds of light sources of the light source means to change the modes of the light. Accordingly, the device simplifies the configuration by excluding elements causing errors in the process of fluorescence diagnosis in the diagnostic region. Moreover, the light source device for fluorescence diagnosis and photodynamic therapy couples the light emitted from a combined light source to provide a sufficient light strength in the visible light region, thus improving the illumination quality to increase the accuracy of the diagnosis and further improving the therapeutic efficiency. | 1. A light source device for fluorescence diagnosis and photodynamic therapy comprising: a light source comprised of an excitation light source emitting light in a short wavelength range and a long wavelength light source emitting light in a long wavelength range, wherein said short wavelength range is different from said long wavelength range;a light guide adapted to guide light to a diagnostic region;a light path coupler coupling light emitted from the light source to an incident end of said light guide;a first filter for selectively transmitting emitted light based on wavelength;a second filter for photodynamic therapy positionable in a light path between the light path coupler and the light guide; anda switching device which controls emissions of the long wavelength light source so as to achieve any of (i) a white light condition where both said excitation light and said long wavelength light are directed to said diagnostic region, (ii) an excitation condition where only said excitation light is directed to said diagnostic region, and (iii) a photodynamic condition where both said excitation and said long wavelength light are directed to said diagnostic region under a photodynamic therapy condition. 2. The light source device for fluorescence diagnosis and photodynamic therapy as recited in claim 1, wherein the excitation light source is a mercury lamp. 3. The light source device for fluorescence diagnosis and photodynamic therapy as recited in claim 1, wherein the long wavelength light source is selected from the group consisting of a halogen lamp, a light emitting diode, a xenon lamp and a metal-halide lamp. 4. The light source device for fluorescence diagnosis and photodynamic therapy as recited in claim 1, wherein the light path coupler is a dichroic mirror that transmits excitation light emitted from the excitation light source and, at the same time, reflects light emitted from the long wavelength light source so as to couple the excitation light together with the light emitted from the long wavelength light source to the incident end of the light guide. 5. The light source device for fluorescence diagnosis and photodynamic therapy as recited in claim 4, wherein the dichroic mirror that transmits the excitation light emitted from the excitation light source has a high light transmittance in a light spectrum range of less than 460 nm, thereby enabling a detector to observe fluorescence generated from acne. 6. The light source device for fluorescence diagnosis and photodynamic therapy as recited in claim 4, wherein the dichroic mirror that transmits the excitation light emitted from the excitation light source has a high light transmittance in a light spectrum range of less than 560 nm, thereby enabling a detector to observe fluorescence produced from a malignant tumor. 7. The light source device for fluorescence diagnosis and photodynamic therapy as recited in claim 4, wherein the dichroic mirror is angled at a 45° with respect to t... | Korea Electro Technology Research Institute | A61B006/00
US07941308 | B2 | 2011-05-10 | Disease simulation system and method | 2008-07-17 | A system and method for predicting the effect of patient self-care actions on a disease control parameter. A future disease control parameter value X(tj) at time tj is determined from a prior disease control parameter value X(ti) at time ti based on an optimal control parameter value R(tj) at time tj, the difference between the prior disease control parameter value X(ti) and an optimal control parameter value R(ti) at time ti, and a set of differentials between patient self-care parameters having patient self-care values SM(ti) at time ti and optimal self-care parameters having optimal self-care values OM(ti) at time ti. The differentials are multiplied by corresponding scaling factors KM. The system includes an input device for entering the patient self-care values SM(ti). A memory stores the optimal control parameter values R(ti) and R(tj), the prior disease control parameter value X(ti), the optimal self-care values OM(ti), and the scaling factors KM. A processor in communication with the input device and memory calculates the future disease control parameter value X(tj). A display is connected to the processor to display the future disease control parameter value X(tj) to a patient. | 1. A system for diabetes management, comprising: (a) an electronic device for measuring glucose values; and(b) a personal computer for receiving the glucose values measured by the electronic device, wherein the personal computer comprises: (i) a microprocessor;(ii) a memory having digital code embedded therein for programming the microprocessor to calculate diabetes care values and to generate a representation based on (i) one or more optimal glucose values, (ii) insulin dose values of a patient, (iii) said measured glucose values received from said electronic device, and (iv) one or more scaling factors for weighting the impact of changes to said insulin dose values on the measured glucose values, wherein said representation of said diabetes care values is customizable to the patient; and(iii) a display configured to display information according to said diabetes care values including said representation. 2. The system of claim 1, wherein said digital code is further for programming the microprocessor to calculate said diabetes care values based on carbohydrate intake values. 3. The system of claim 1, wherein said display is configured for displaying multiple values in graphical form. 4. The system of claim 1, wherein said display of said representation is in the form of a visual representation including a figure and enables the patient to consider appropriate self care actions. 5. The system of claim 1, wherein said digital code is further for programming the microprocessor to calculate said diabetes care values based on optimal insulin dose values. 6. The system of claim 1, wherein said digital code is further for programming the microprocessor to determine said one or more scaling factors from one or more physiological parameters including body mass, metabolism rate, fitness level or hepatic or peripheral insulin sensitivity, or combinations thereof. 7. The system of claim 1, wherein said digital code is further for programming the microprocessor to determine said one or more scaling factors from insulin sensitivity. 8. The system of claim 1, wherein said digital code is further for programming the microprocessor to calculate at least one future glucose value. 9. The system of claim 1, further comprising a server computer connectable to the said personal computer for communicating data between the server computer and the personal computer. 10. The system of claim 1, wherein said digital code is for programming the microprocessor to generate graphs of the measured glucose values versus time. 11. The system of claim 10, wherein the measured glucose values comprise at least one future glucose value. 12. The system of claim 11, wherein said digital code is further for programming the microprocessor to calculate said at least one future glucose value. 13. The system of claim 1, wherein said digital code is further for programming the microprocessor to calculate said diabetes care values based on food intake values. 14. The system of c... | Health Hero Network, Inc. | G01N033/487 G06F019/00
US07927274 | B2 | 2011-04-19 | Integrated receiver for continuous analyte sensor | 2008-07-29 | A system is provided for monitoring glucose in a host, including a continuous glucose sensor that produces a data stream indicative of a host's glucose concentration and an integrated receiver that receives the data stream from the continuous glucose sensor and calibrates the data stream using a single point glucose monitor that is integral with the integrated receiver. The integrated receiver obtains a glucose value from the single point glucose monitor, calibrates the sensor data stream received from the continuous glucose sensor, and displays one or both of the single point glucose measurement values and the calibrated continuous glucose sensor values on the user interface. | 1. A system for monitoring a glucose concentration in a biological sample in a host, the system comprising: a continuous glucose sensor configured to produce sensor data indicative of a glucose concentration in a host, wherein the glucose sensor comprises a sensing membrane comprising an enzyme domain; anda receiver unit adapted to receive the sensor data from the continuous glucose sensor, wherein the receiver unit comprises a housing, and wherein the receiver unit comprises: a single point glucose monitor physically connected to the housing, wherein the single point glucose monitor is configured to produce a glucose concentration measurement from a biological sample obtained from a host, wherein the glucose monitor comprises a sensing membrane comprising an enzyme domain, wherein the receiver unit is configured to calibrate the sensor data from the continuous glucose sensor using the glucose concentration measurement from the single point glucose monitor, and wherein the enzyme domain of the sensing membrane of the continuous glucose sensor and the enzyme domain of the sensing membrane of the single point glucose monitor are substantially the same. 2. The system of claim 1, wherein the enzyme domain of the sensing membrane of the continuous glucose sensor and the enzyme domain of the sensing membrane of the single point glucose monitor comprise the same materials. 3. The system of claim 1, wherein the single point glucose monitor comprises a single use test strip configured to be releasably received by a port on the single point glucose monitor, and wherein the single use test strip of the stele point glucose monitor comprises the sensing membrane. 4. The system of claim 1, wherein the receiver unit further comprises programming to control administration of a constituent to the host. 5. The system of claim 4, wherein the programming to control administration of a constituent to the host is configured to determine at least one of an amount of insulin and a time of insulin administration. 6. The system of claim 1, wherein the receiver unit further comprises programming to calculate a rate of change of the continuous glucose data or a directional trend of the continuous glucose data. 7. The system of claim 6, wherein the receiver unit further comprises programming to control administration of a constituent to the host based at least in part on the rate of change of the continuous glucose data or the directional trend of the continuous glucose data. 8. The system of claim 1, wherein both the continuous glucose sensor and the single point glucose monitor use the same measurement technology, and wherein the measurement technology is selected from the group consisting of enzymatic, chemical, physical, electrochemical, spectrophotometric, polarimetric, calorimetric, and radiometric. 9. The system of claim 1, wherein the sensing membrane of the single point glucose monitor comprises one or more membrane domains or layers that are the same as one... | DexCom, Inc. | A61B005/00
US08057408 | B2 | 2011-11-15 | Pulsed cavitational ultrasound therapy | 2008-05-15 | Therapy methods using pulsed cavitational ultrasound therapy can include the subprocesses of initiation, maintenance, therapy, and feedback of the histotripsy process, which involves the creation and maintenance of ensembles of microbubbles and the use of feedback in order to optimize the process based on observed spatial-temporal bubble cloud dynamics. The methods provide for the subdivision or erosion of tissue, liquification of tissue, and/or the enhanced delivery of therapeutic agents. Various feedback mechanisms allow variation of ultrasound parameters and provide control over the pulsed cavitational process, permitting the process to be tuned for a number of applications. Such applications can include specific tissue erosion, bulk tissue homogenization, and delivery of therapeutic agents across barriers. | 1. A method for controlled mechanical sub-division of soft tissue comprising: de-initiating a non-treatment portion of the soft tissue to increase a cavitation threshold from a first threshold to a second threshold in said non-treatment portion of the soft tissue, wherein said de-initiating said non-treatment portion of the soft tissue comprises de-initiating said non-treatment portion of the soft tissue using a deleting pulse sequence at a pulse intensity less than said first threshold;using a transducer outputting an initiation pulse sequence at a treatment portion of the soft tissue, wherein said de-initiating a non-treatment portion of the soft tissue and said outputting an initiation pulse sequence are performed simultaneously;detecting initiation of a bubble cloud in said treatment portion of the soft tissue in response to said initiation pulse sequence;outputting a bubble cloud maintenance and sustaining pulse sequence at said treatment portion; andoutputting a therapy pulse sequence which generates and interacts with said bubble cloud to at least partially fractionate the soft tissue in said treatment portion. 2. The method according to claim 1, further comprising: detecting cessation of said bubble cloud. 3. The method according to claim 2, further comprising: re-outputting said initiation pulse sequence in response to said cessation of said bubble cloud. 4. The method according to claim 2, further comprising: re-outputting said bubble cloud sustaining pulse sequence prior to said cessation of said bubble cloud. 5. The method according to claim 1 wherein said outputting an initiation pulse sequence comprises outputting an initiation pulse sequence at a frequency equal to a frequency of at least one of said bubble cloud sustaining pulse sequence and said therapy pulse sequence. 6. The method according to claim 1 wherein said outputting an initiation pulse sequence comprises outputting an initiation pulse sequence operable to create at least one of a vapor cloud or a plasma cloud. 7. The method according to claim 1 wherein said outputting an initiation pulse sequence comprises outputting an initiation pulse sequence causing an ionizing radiation bubble chamber effect. 8. The method according to claim 1 wherein said outputting an initiation pulse sequence comprises mechanically injecting cavitation nuclei. 9. The method according to claim 8 wherein said mechanically injecting cavitation nuclei comprises mechanically injecting cavitation nuclei intravascularly. 10. The method according to claim 8 wherein said mechanically injecting cavitation nuclei comprises mechanically injecting cavitation nuclei into a therapy volume. 11. The method according to claim 1, further comprising: targeting a therapeutic substance at a predetermined location in the soft tissue to enhance enucleation of dissolved gases in the soft tissue. 12. The method according to claim 11 wherein said targeting a therapeutic substance comprises targeting a therapeu... | The Regents of the University of Michigan | A61H001/00
US07970462 | B2 | 2011-06-28 | Implantable medical devices evaluating thorax impedance | 2007-05-29 | Implantable medical device with an impedance determination unit with constant current/voltage source having current feed terminals connected to electrodes for intracorporal placement which generates measuring current pulses having constant current/voltage, for causing a current through a body via intracorporally placed electrodes, a measuring unit for measuring voltage/current strength of voltage/current fed through body, an impedance value determination unit connected to the current/voltage source and adapted to determine an impedance value for each measuring current pulse, and an impedance measuring control and evaluation unit connected to the impedance determination unit which controls the unit and evaluates a sequence of consecutive impedance values, the impedance determination unit further adapted to determine at least intrathoracic and intracardiac impedance values for same period of time, the intrathoracic values sampled with a lower sampling rate than the intracardiac values. | 1. An implantable medical device comprising: an impedance determination unit comprising a constant current source or a constant voltage source having current feed terminals that are connected or can be connected to electrodes for intracorporal placement and that is configured to generate measuring current pulses having a constant current strength or a constant voltage, respectively, to cause a current to be fed through a body via intracorporally placed electrodes;a measuring unit configured to measure a voltage corresponding to a current fed through a body by said constant current source or a current strength of a current fed through a body by said constant voltage source, respectively;an impedance value determination unit that is connected to said constant current source or said constant voltage source and to the measuring unit wherein said impedance value determination unit is configured to determine an impedance value for each measuring current pulse; and,an impedance measuring control and evaluation unit, that is connected to said impedance determination unit and that is configured to control said impedance determination unit and to evaluate a sequence of consecutive impedance values determined by said impedance determination unit; andsaid impedance determination unit further configured to determine at least intrathoracic impedance values and intracardiac impedance values concurrently, wherein the intrathoracic impedance values are sampled with a lower sampling rate than the intracardiac impedance values,determine respiration rate values from said intrathoracic impedance values,determine heart rate values from said intracardiac impedance values,evaluate said sequence of consecutive impedance values based on said respiration rate values and said heart rate values. 2. The implantable medical device according to claim 1, wherein the impedance determination unit is further configured to determine at least two kinds of intrathoracic impedance values comprising: a first kind corresponding to an impedance between an IMD case and at least one electrode to be placed in one ventricle of a heart; and,a second kind corresponding to an impedance between an IMD case at least one intracardiac electrode. 3. The implantable medical device according to claim 1, wherein the impedance determination unit is further configured to determine at least two kinds of intrathoracic impedance values comprising: a first kind corresponding to an impedance between an IMD case and an electrode to be placed in a right ventricle of a heart, anda second kind corresponding to an impedance between an IMD case and an electrode to be placed in a left ventricle of a heart. 4. The implantable medical device according to claim 2 wherein the impedance determination unit is further configured to determine said intrathoracic impedance values wherein: said constant current source is connected to the IMD case and a right ventricular or a left ventricular ring electrode, respectively, suc... | Biotronik CRM Patent AG | A61B005/02 A61B005/08 A61B005/05 A61N001/00
US08096958 | B2 | 2012-01-17 | Blood-collection device for newborn babies and infants | 2005-02-04 | The invention relates to a blood-collection device for newborn babies and infants, comprising a cannula, which is provided with a blood inlet and outlet and is located in a holder (2) that has a grip region (7). Said collection device is equipped with a bow-shaped bridging element (6) that connects the grip region to the holder. | 1. A blood-collecting device for newborn babies and infants, the device comprising: a cannula extending along an axis and having at one axial end a blood inlet and at an opposite axial end a blood outlet;a bow-shaped bridge element having a front end in which the cannula is fixed and a rear end through which the axis passes and defining with the front end a free space, the cannula outlet being exposed between the front and rear ends in the free space and the inlet being outside the free space; anda grip part centered on and substantially rotationally symmetrical about the axis and fixed on the rear end of the bridge element rearward of the rear end and outside the free space. 2. The device defined in claim 1 wherein the bridge element is C-shaped. | Sarstedt AG & Co | A61B005/00 B65D081/00 A61M005/32
US08406492 | B2 | 2013-03-26 | Image detecting system, image detecting method and computer readable medium | 2008-05-19 | There is provided an image detecting system including an image detecting section that detects a moving image of an examination subject into which a radiopaque contrast medium flows through one of an artery and a portal vein, and a change image generating section that generates a change image representing a change in the moving image. Here, the change is caused by movement of the radiopaque contrast medium after a timing at which the radiopaque contrast medium that has flown into one of the artery and the portal vein flows into capillaries. | 1. An image detecting system comprising: an image detecting section configured for detecting a moving image of an examination subject into which a radiopaque contrast medium flows through one of an hepatic artery and a hepatic portal vein; anda change image generating section configured for generating a change image representing a change in the moving image, the change being caused by movement of the radiopaque contrast medium after a timing at which the radiopaque contrast medium that has flown into one of the artery and the portal vein flows into capillaries;whereinthe change image generating section is configured for generating (i) a change image when the radiopaque contrast medium flows into the examination subject through the artery and (ii) a change image when the radiopaque contrast medium flows into the examination subject through the portal vein, andthe image detecting system further comprisesan image processing apparatus that is configured to cause a display section to display the change images generated by the change image generating section, to emphasize a difference between (i) the change image when the radiopaque contrast medium flows into the examination subject through the artery and (ii) the change image when the radiopaque contrast medium flows into the examination subject through the portal vein by displaying a differential image between (i) the change image when the radiopaque contrast medium flows into the examination subject through the artery and (ii) the change image when the radiopaque contrast medium flows into the examination subject through the portal vein generated by the change image generating section. 2. The image detecting system as set forth in claim 1, further comprising a timing identifying section configured for identifying the timing at which the radiopaque contrast medium that has flown into one of the artery and the portal vein flows into the capillaries, with reference to what is shown by the moving image. 3. The image detecting system as set forth in claim 2, wherein the change image generating section is configured for generating the change image represented by a value obtained by integrating a pixel value of the moving image for a predetermined time period after the timing identified by the timing identifying section. 4. The image detecting system as set forth in claim 3, wherein the change image generating section is configured for generating the change image represented by a value obtained by integrating, for the predetermined time period, a difference in pixel value between (i) the moving image detected after the timing identified by the timing identifying section and (ii) an image of the examination subject detected when the radiopaque contrast medium does not flow into the examination subject. 5. The image detecting system as set forth in claim 3, wherein the change image generating section is configured for generating the change image represented by a value obtained by integrating, for the pre... | FUJIFILM Corporation | G06K009/00
US08423105 | B2 | 2013-04-16 | Sensor for eliminating undesired components and measurements method using said sensor | 2006-04-05 | In the present invention a novel method and device for measuring characteristics from a relatively weak signal comprising desired and undesired components is presented. Undesired signals may arise from the nature of the characteristic, from the detector or from the circuitry. The signal is extracted from a first measurement element (1) comprising these desired and undesired components. Using another signal from this first measurement element or from another second measurement element (2) the undesired components can be eliminated. The measurement method is extremely useful when using organic materials for the detectors, electronic circuitry, and measurement elements. These devices can be produced relatively cheap, but less reliable. They can also be combined in a one- or two-dimensional array for measuring characteristics over a larger area. | 1. A sensor comprising at least one sensor unit, each sensor unit comprising: a charge storage element;a first circuitry part connected between said charge storage element and a first measurement element recording a measurement signal, said first circuitry part being provided for periodically applying a first measurement current indicative of a first recorded measurement value to said charge storage element for charging said charge storage element, said first circuitry part comprising a first switching element;a second circuitry part connected between said charge storage element and either said first measurement element or a second measurement element recording said measurement signal, said second circuitry part provided for periodically applying a second measurement current indicative of a second recorded measurement value to the charge storage element for discharging said charge storage element, said second measurement current consisting of components to be eliminated from said first measurement current, said second circuitry part comprising a second switching element;said first and second switching elements being switchable between a conducting state and a non-conducting state;a third circuitry part connected between said charge storage element and an output of said sensor, said third circuitry part comprising a third switching element switchable between a conducting and a non-conducting state; andcontrol circuitry for applying voltages to said first and second switching elements according to a predetermined periodic timing scheme, such that in each period of said timing scheme said first and second switching elements are switched to the conducting state for a given time and that the charge stored on said charge storage element as a result of the components to be eliminated from the first measurement current is substantially equal to said charge which is removed from the charge storage element as a result of the second measurement current;wherein: said charge storage element and said first, second and third circuitry parts are made of organic materials, each of said switching elements in each sensor unit having an uncertain threshold voltage, said conducting and non-conducting states being located on opposite sides of a predetermined voltage range containing said uncertain threshold voltage;said control circuitry is adapted for applying voltages to said switching elements outside said predetermined voltage range; andsaid third switching element is arranged for conducting charge stored on the charge storage element to said output, said third circuitry part being designed for periodically conducting said charge after a predefined number of said periods. 2. A sensor according to claim 1, wherein said first circuitry part consists of said first switching element, said second circuitry part consists of said second switching element and a current mirror, and said third circuitry part consists of said third switching element. 3. A sensor according ... | IMEC | A61B005/00
US08457732 | B2 | 2013-06-04 | Biometric apparatus | 2008-05-09 | The invention provides a biometric apparatus having a display which allows a measured person to recognize biometric data easily and reliably irrespective of whether the measured person is on a platform or not. The biometric apparatus for measuring biometric data of the measured person includes the platform on which the measured person gets on, a sensor unit for acquiring the biometric data of the measured person; a display for displaying the biometric data acquired by the sensor unit; and a control unit for determining whether the measured person is on the platform or not on the basis of a signal from the sensor unit and changes the display mode to display the biometric data on the display according to the determination thereof. | 1. A biometric apparatus for measuring body composition data of a measured person comprising: a platform on which the measured person gets on;a sensor unit for acquiring the body composition data of the measured person and outputting a signal indicating a status of whether the measured person is on or off the platform;a display for displaying the body composition data acquired by the sensor unit; anda control unit for determining, after having acquired the body composition data of the measured person, whether or not the measured person is still on the platform on the basis of the signal from the sensor unit, and displaying the acquired body composition data on the display in one of different manners according to the determination thereof, wherein:when the control unit determines that the measured person is still on the platform, the control unit causes the display to display the acquired body composition data in one of the different manners based on the signal received from the sensor unit, and,when the control unit determines that the measured person is not on the platform, the control unit causes the display to display the acquired body composition data in another one of the different manners based on the signal received from the sensor unit,wherein characters indicating the acquired body composition data displayed on the display are larger in the one of the different manners than characters indicating the acquired body composition data displayed on the display in the another one of the different manners, andwherein the control unit switches between the different manners in response to the signal indicating that said status has changed. 2. The biometric apparatus according to claim 1, wherein: the body composition data includes a plurality of types of body composition data, andthe control unit displays the acquired body composition data in the one of different manners such that at least one, but not all, of the plurality of types of body composition data is displayed on the display when the control unit determines that the measured person is still on the platform, and all of the plurality of types of body composition data are displayed on the display when the control unit determines that the measured person is not on the platform. 3. The biometric apparatus according to claim 1, wherein: the body composition data includes a plurality of types of body composition data, andthe control unit displays the acquired body composition data in the one of different manners such that all of the plurality of types of body composition data are displayed sequentially on the display one by one when the control unit determines that the measured person is on the platform. 4. The biometric apparatus according to claim 1, wherein the body composition data to be displayed on the display includes at least one of the weight, the degree of fatness, the percent body fat, the amount of subcutaneous fat, the amount of offal fat and the body age. 5. The biometric appa... | Tanita Corporation | A61B005/05
US08504128 | B2 | 2013-08-06 | Method and apparatus for coupling a channeled sample probe to tissue | 2008-04-24 | Sampling is controlled in order to enhance analyte concentration estimation derived from noninvasive sampling. More particularly, sampling is controlled using controlled fluid delivery to a region between a tip of a sample probe and a tissue measurement site. The controlled fluid delivery enhances coverage of a skin sample site with the thin layer of fluid. Delivery of contact fluid is controlled in terms of spatial delivery, volume, thickness, distribution, temperature, and/or pressure. | 1. A sample probe configured to communicate with a non-invasive spectroscopic analyzer to determine an analyte property from a sample site of a body part, said sample probe comprising a sample probe tip having a microfluidic fluid delivery channel in a surface of said sample probe tip, said microfluidic fluid delivery channel forming a pathway for dispersion of a coupling fluid across the surface of said sample probe tip, wherein said microfluidic fluid delivery channel comprises a first moat circumferentially surrounding an optic in said sample probe. 2. The sample probe of claim 1, further comprising a fluid delivery aperture in said sample probe tip and proximate to said first moat, wherein said coupling fluid is delivered from a reservoir via said fluid delivery aperture to said first moat. 3. The sample probe of claim 2, further comprising a second moat in the surface of said sample probe tip, said second moat circumferentially surrounding said first moat. 4. The sample probe of claim 2, further comprising means for maintaining temperature of said coupling fluid at about ninety to ninety-two degrees Fahrenheit prior to delivery of said coupling fluid. 5. The sample probe of claim 1, further comprising means for creating a negative pressure in said first moat during use of said analyzer. 6. The sample probe of claim 1, wherein said microfluidic fluid delivery channel further comprises at least one radially extending channel. 7. The sample probe of claim 1, wherein said microfluidic fluid delivery channel further comprises said first moat coupled to at least one radially extending channel. 8. The sample probe of claim 2, further comprising means for delivering said coupling fluid from said reservoir, via said fluid delivery aperature, to said microfluidic fluid delivery channel. 9. The sample probe of claim 8, in combination with said analyzer, wherein said analyzer comprises a processor configured to control said means for delivering. 10. The sample probe of claim 9, wherein said processor is configured to operate in accordance with a control signal from a sensor. 11. The sample probe of claim 10, wherein said sensor comprises any of: an electrical contact sensor configured to sense proximate contact of said sample probe tip with the sample site, wherein upon proximate contact of said sample probe tip with the sample site, an electrical circuit is completed and said electrical contact sensor generates a signal indicative of completion of said electrical circuit; anda pressure sensor configured to sense proximate contact of said sample probe tip with the sample site, wherein said pressure sensor comprises a film layer with air voids internally contained, wherein said pressure sensor is configured to generate a charge and corresponding voltage indicative of different capacitive charges measured as pressure is applied to said film layer. 12. The sample probe of claim 1, in combination with said analyzer, wherein said analyzer comprise... | GLT Acquisition Corp. | A61B005/1455
US08078258 | B2 | 2011-12-13 | Assessment of vascular compartment volume for PET modelling | 2008-05-15 | A method is described for acquiring and analysing the data produced by Positron Emission Tomography (PET), which method provides for an accurate estimation of vascular compartment volume, VB. An MRI scan and the PET scan is performed simultaneously and the results of the former is used to derive a value for VB. The value so derived is then used in pharmacokinetic modelling along with the results of the functional scan. | 1. A method of processing a set of Positron Emission Tomography (PET) scan images comprising the steps of: administering a PET scan contrast agent and a magnetic resonance imaging (MRI) contrast agent to a subject;simultaneously performing a PET scan and an MRI scan on the subject to generate PET scan images and corresponding data from the MRI scan;calculating VB, the fraction of pixel volume occupied by blood vessels, using data acquired from the MRI scan; andperforming pharmacokinetic analysis based on i) the value of VB so derived and ii) the PET scan images. 2. A method according to claim 1, wherein the PET scan contrast agent is 18-F fluoro-2-deoxy-2-glucose (FDG). 3. A method according to claim 1 wherein the PET scan contrast agent is 3′-deoxy-3′-[18-F]-fluorothymidine (FLT). 4. A method according to claim 1 wherein the PET scan contrast agent is 6-[(18)F]fluoro-L-DOPA (FDOPA). 5. Apparatus comprising: a PET scanner;an MRI scanner, said PET scanner and said MRI scanner being configured to perform simultaneous scans of a subject; anda processor;wherein the process is configured, to receive data from each of the PET scanner and the MRI scanner;to calculate VB, the fraction of pixel volume occupied by blood vessels, using the data acquired during the MRI scan; andto perform pharmacokinetic analysis based on i) the value of VB so derived, and ii) the PET scan images. | Siemens Medical Solutions USA, Inc. | A61B005/05
US07986998 | B2 | 2011-07-26 | Methods and devices for detection of context when addressing a medical condition of a patient | 2008-07-10 | Methods and devices detect context related to a patient when monitoring a physiological condition of the patient and/or when applying one or more modes of therapy. The context may be a patient context such as posture or an environmental context such as ambient conditions. The context may be used in various ways in relation to the physiological measurement, such as to control when the physiological measurements are made, to appropriately flag physiological measurements, to be recorded in association with the physiological measurements, and/or to correct the physiological measurements based on a reference context. A device such as a beacon transmitter is used in detecting the context and a measurement device such as an implantable cardiovascular device is used to capture the physiological measurements. | 1. A method of monitoring physiological conditions of an ambulatory patient who is capable of moving to a plurality of non-clinical locations, comprising: measuring at least one physiological condition of the patient with an implanted medical device;associating a non-clinical location with a contextual condition for the patient that is relevant when measuring the at least one physiological condition of the patient, wherein the non-clinical location is one of the plurality of non-clinical locations to which the ambulatory patient is capable of moving; andwhile measuring the at least one physiological condition, utilizing the implantable medical device for sensing an external device in the non-clinical location and determining that the contextual condition associated with the non-clinical location exists when the patient is within a range of the external device and the implantable medical device senses the external device; andflagging with a first indicator one or more measurements resulting from the measuring done while the contextual condition continues to exist. 2. The method of claim 1, wherein the contextual condition is determinative of a patient context. 3. The method of claim 2, wherein detecting whether the at least one contextual condition exists comprises interacting from the external device to the implanted medical device for detecting the patient context. 4. The method of claim 1, wherein sensing that the external device is within the range comprises: emitting a signal from the external device;receiving the signal at the implanted medical device when the implanted medical device is within the range of the external device; andflagging with the first indicator one or more measurements of the at least one physiological condition upon the implanted medical device receiving the signal. 5. The method of claim 1, wherein the at least one contextual condition defines an environmental context. 6. The method of claim 5, wherein the environmental context is temperature present proximal to the patient. 7. The method of claim 1, wherein measuring the at least one physiological condition comprises measuring a cardiovascular value. 8. The method of claim 1, wherein measuring the at least one physiological condition comprises measuring a respiratory value. 9. The method of claim 1, comprising changing a therapy mode of the implantable medical device in association with sensing the external device. 10. A system for monitoring physiological conditions of a patient, comprising: a transmitter external to the patient and in a predetermined location associated with at least one contextual condition;a detector adapted to detect whether the detector is within a range of the transmitter to detect the at least one contextual condition;an implanted physiological sensor adapted to sense a physiological condition of the patient inside and outside of the range; anda processing device configured to communicate with the implanted physiological sensor to meas... | Cardiac Pacemakers, Inc. | A61B005/04
US08260392 | B2 | 2012-09-04 | Analyte monitoring device and methods of use | 2008-06-09 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. An analyte monitoring device, comprising: a housing;an input unit coupled to the housing to provide one or more patient condition parameters, the input unit configured to initiate data communication from the analyte monitoring device to a remotely located sensor control unit operatively coupled to a transcutaneously positioned analyte sensor;a receiver unit coupled to the housing and configured to receive analyte sensor signals detected by the transcutaneously positioned analyte sensor, the receiver unit further configured to receive sensor parameter data that is related to the transcutaneously positioned analyte sensor, wherein the sensor parameter data is determined prior to transcutaneous positioning of the analyte sensor;a reading unit coupled to the housing and configured to generate a reading based upon the analyte sensor signals detected by the transcutaneously positioned analyte sensor and received from the sensor control unit, and the sensor parameter data that is related to the transcutaneously positioned analyte sensor and received from the sensor control unit;a processor coupled to the housing and configured to process one or more analyte related signals detected by the transcutaneously positioned analyte sensor and received from the remotely located sensor control unit and the sensor parameter data that is related to the transcutaneously positioned analyte sensor and received from the remotely located sensor control unit, the one or more analyte related signals modified by the processor based on the reading generated by the reading unit, the processor further configured to process the one or more patient condition parameters, wherein no calibration of the one or more analyte sensor signals detected by the transcutaneously positioned analyte sensor is performed after the analyte sensor is transcutaneously positioned; anda display unit coupled to the housing to display the processed one or more analyte related signals and the one or more patient condition parameters. 2. The device of claim 1, wherein the one or more patient condition parameters includes a meal event time indication. 3. The device of claim 1, wherein the one or more patient condition parameters includes a food intake information. 4. The device of claim 1, wherein the one or more patient condition parameters includes an exercise information. 5. The device of claim 1, wherein the one or more patient condition parameters includes a fasting information. 6. The device of claim 1, wherein the one or more patient condition parameters includes a change in body temperature information. 7. The device of claim 1, wherein the one or more patient condition parameters includes a physiological condition information. 8. The device of claim 1, wherein the one or more patient condition parameters includes a treatment information. 9. The device of claim 1 wherein the one or more analyte related signals is associated with a monitored analyte level, and wherein the display unit i... | Abbott Diabetes Care Inc. | A61B005/04 A61B005/00
US07866212 | B2 | 2011-01-11 | High-sensitivity sensors for sensing various physiological phenomena, particularly useful in anti-snoring apparatus and methods | 2005-06-09 | A mechanical vibration sensor adapted to be brought into contact with an object for sensing mechanical vibrations in the object, includes a body of a soft elastomeric material having high transmissivity and low attenuation properties with respect to a preselected type of energy waves; and a pair of transducers mounted, by mounting members having high attenuation properties with respect to the energy waves, in spaced relationship to each other to define a transmission channel between the transducers. Such sensor is particularly useful in a method and apparatus for controlling snoring by a person, by utilizing a stimulus device effective, when sensing snoring, to immediately produce a response in the person tending to interrupt the person's snoring. | 1. Apparatus for controlling snoring or other predetermined body condition by a person, comprising: a sensor for sensing said predetermined body condition of the person while sleeping; and for outputting a sensor signal corresponding to the sensed condition; a stimulus device effective, when actuated, to immediately produce a response in the person tending to interrupt said sensed body condition; and a processor for processing said outputted sensor signal to determine whether it indicates the presence of said sensed body condition, and if so, for actuating said stimulus device to produce said response tending to interrupt said sensed body condition; characterized in that said sensor is a mechanical vibrations sensor and is included between, and in contact with, a pair of plates of larger size than said sensor for sensing mechanical vibrations in an external part of the person's body and for transmitting said mechanical vibrations to said sensor for generating and outputting said sensor signal. 2. The apparatus according to claim 1, wherein said processor: analyzes said sensor signal for frequency components within a predetermined frequency band;accumulates the energy of the sensor signal within said predetermined frequency band for a preselected time interval;compares said accumulated energy for each predetermined time interval with the average of the accumulated energy in a plurality of the preceding preselected time intervals;and controls the actuation of said stimulus device as a result of such comparison. 3. The apparatus according to claim 2, wherein said processor controls the actuation of said stimulus device in accordance with a determination of whether or not the accumulated energy in the preselected time interval exceeds by a predetermined percentage the average of the accumulated energy in said plurality of preceding preselected time intervals. 4. The apparatus according to claim 3, wherein said processor actuates said stimulus device when the accumulated energy in each of at least two successive preselected time intervals exceeds by said predetermined percentage the average of the accumulated energy in the plurality of preceding preselected time intervals. 5. The apparatus according to claim 3, wherein said apparatus includes a manual control for varying said predetermined percentage, and thereby for varying the sensitivity of the apparatus. 6. The apparatus according to claim 1, wherein said preselected time interval is substantially less than the inhale and exhale half-cycle times of the person's respiration cycle time such that the actuation of the stimulus device is closely time-coordinated with the sensed body condition during the person's respiration cycle. 7. The apparatus according to claim 1, wherein said sensor senses said mechanical vibrations in the part of the person's body by causing said mechanical vibrations sensor to change the transit time of an acoustic wave transmitted from a transmitter to a receiver in an a... | Nexense Ltd. | G01N029/00 A61B005/08
US07973211 | B2 | 2011-07-05 | Animal model for hyperpigmentation | 2008-03-11 | An animal model for hyperpigmentations in which the formation of hyperpigmentations in human skin is faithfully simulated is provided. An animal model for hyperpigmentations, wherein a black person's skin is grafted onto a non-human animal, is provided. | 1. An animal model for evaluating hyperpigmentations, wherein said model is an immunodeficient mouse or an immunodeficient rat onto which a black person's skin has been grafted, wherein within a month after the grafting, the grafted skin comprises (a) darkly pigmented parts and (b) light parts from decolorization after grafting. 2. The animal model according to claim 1, wherein said animal model is said immunodeficient mouse. 3. The animal model according to claim 1, wherein the area of the darkly pigmented part comprises up to 60% of the area of the grafted skin. 4. The animal model according to claim 3, wherein the expression of a gene of at least one cytokine selected from the group consisting of Endothelin-1, SCF and POMC is increased in the darkly pigmented part as compared to the light part. 5. A method of preparing a skin graft for evaluating the formation of hyperpigmentations in said graft, said method comprising grafting a black person's skin onto an immunodeficient mouse or an immunodeficient rat, and raising said mouse or said rat after said grafting, wherein within one month after said grafting, the grafted skin comprises (a) darkly pigmented parts and (b) light parts from decolorization after grafting. 6. A method of evaluating or screening a hyperpigmentation formation suppressant or a hyperpigmentation remover, wherein said method comprises administering said suppressant or said remover to said immunodeficient mouse or said immunodeficient rat animal model of any one of claims 1 to 4, and evaluating the effect of said suppressant or said remover on the grafted skin. 7. The method of claim 6, wherein said method evaluates whether said suppressant suppresses the formation of darkly pigmented parts in the grafted skin. 8. The method of claim 6, wherein said method evaluates the thickness of the epidermis of darkly pigmented parts of the grafted skin. 9. The animal model of claim 2, wherein said immunodeficient mouse is a BALBcA-nu/Scid mouse or a B-17/Icr-Scid mouse. 10. The animal model of claim 1, wherein said animal model is said immunodeficient rat. 11. The animal model of claim 10, wherein said immunodeficient rat is a F344Jca-rnu rat. | Kao Corporation | A01K067/00
US08273227 | B2 | 2012-09-25 | Sensor for in vitro determination of glucose | 2007-10-30 | A process for the manufacture of small sensors with reproducible surfaces, including electrochemical sensors. One process includes forming channels in the surface of a substrate and disposing a conductive material in the channels to form an electrode. The conductive material can also be formed on the substrate by other impact and non-impact methods. In a preferred embodiment, the method includes cutting the substrate to form a sensor having a connector portion and a transcutaneous portion, the two portions having edges that define one continuous straight line. | 1. A sensor for the in vitro determination of glucose level in a sample, the sensor comprising: a substrate comprising a first recessed channel;a first conductive material to provide an electrode disposed in the recessed channel on the substrate to substantially fill the channel such that the substrate and the conductive material form a substantially flat surface,wherein a catalyst is mixed or dispersed in said first conductive material,wherein the catalyst includes an enzyme, wherein the enzyme is mixed in a carbon ink to form a mixture,wherein the sensor generates an electrical signal associated with a level of glucose in a volume sample of about 1 microliter or less. 2. The sensor of claim 1, wherein the catalyst forms a conductive trace associated with a working electrode. 3. The sensor of claim 1, wherein the substrate comprises a second recessed channel and a second conductive material is applied in the second recessed channel. 4. The sensor of claim 3, wherein the first conductive material and the second conductive material are the same. 5. The sensor of claim 3, wherein the first conductive material and the second conductive material are different. 6. The sensor of claim 1, wherein the substrate includes an indentation. 7. The sensor of claim 1, wherein a trace is associated with the electrode, the trace having a width of about 25 microns to 250 microns. 8. The sensor of claim 7, wherein the trace has a width of about 75 microns or less. 9. The sensor of claim 8, wherein the trace has a width of about 50 microns or less. 10. The sensor of claim 9, wherein the trace has a width of about 25 microns or less. 11. The sensor of claim 1, wherein the sensor is coupled to a skin piercer. 12. The sensor of claim 1, comprising a tapered end. 13. The sensor of claim 1, wherein the enzyme is glucose oxidase. 14. The sensor of claim 1, wherein the enzyme is glucose dehydrogenase. 15. The sensor of claim 1, wherein the enzyme is pyrroloquinoline quinone glucose dehydrogenase. | Abbott Diabetes Care Inc. | G01N027/327
US08221328 | B2 | 2012-07-17 | Visual indicating device for bad breath | 2008-06-06 | The invention provides a breath testing device which includes a visual indicating agent which changes color in the presence of an odor associated with bad breath, such as sulfur and ammonia odors. An example of the visual indicating agent is 4,4′-bis(dimethylamino)-benzhydrol (Michler's hydrol or BDMB) and related dyes having a similar chemical structure. The indicating agent is applied to a substrate which is then inserted into a tube or straw, or which covers one end of a straw. When a user with bad breath blows into the tube or straw, the indicating agent will change color. The breath testing devices provide a quick and affordable means for a user to test their breath, and they may be packaged in discreet, pocket-sized dispensers which can be carried in a pocket or purse. | 1. A method of testing for bad breath in a user comprising the steps of: causing the user to blow or breathe onto or into a carrier portion of a breath testing device containing a visual indicating agent that is sensitive to bad breath odors, wherein the visual indicating agent is Michler's hydrol; andobserving if the visual indicating agent changes color to indicate that the user has bad breath. 2. The method of claim 1, the device further comprising nanoparticles. 3. The method of claim 1, the device comprising a dried residue of an applied solution of the visual indicating agent contained on a substrate. 4. The method of claim 3, wherein the substrate is located in a passage of the carrier portion of the device. 5. The method of claim 3, wherein the substrate covers one end of the carrier portion of the device. 6. The method of claim 1, wherein the visual indicating agent is dissolved in a solvent and remains in solution within the device. 7. The method of claim 1, wherein the visual indicating agent is in powder form. | Kimberly-Clark Worldwide, Inc. | A61B005/08
US08050736 | B2 | 2011-11-01 | Hybrid imaging method to monitor medical device delivery and patient support for use in the method | 2008-04-25 | This invention discloses a method and apparatus to deliver medical devices to targeted locations within human tissues using imaging data. The method enables the target location to be obtained from one imaging system, followed by the use of a second imaging system to verify the final position of the device. In particular, the invention discloses a method based on the initial identification of tissue targets using MR imaging, followed by the use of ultrasound imaging to verify and monitor accurate needle positioning. The invention can be used for acquiring biopsy samples to determine the grade and stage of cancer in various tissues including the brain, breast, abdomen, spine, liver, and kidney. The method is also useful for delivery of markers to a specific site to facilitate surgical removal of diseased tissue, or for the targeted delivery of applicators that destroy diseased tissues in-situ. | 1. An apparatus used in conjunction with imaging and intervention for supporting and transporting a patient to an MR imaging system having a bore, the apparatus comprising: a transport stretcher including at least two elevated portions defining an interventional volume therebetween, and a movable bridge which in a first position spans the interventional volume,a patient support structure supported by the elevated portions of the transport stretcher, attached in a horizontally translatable relationship to the transport stretcher, configured to support the patient, and having an opening such that a body part is accessible from below the patient when the opening is aligned with the interventional volume and the bridge is moved out of the way, andwherein when the bridge is in the first position, the bridge provides support for the patient support structure as the patient support structure is moved into and out of the bore. 2. The apparatus of claim 1, wherein the bridge is one of removable, retractable, foldable and pivotable. 3. The apparatus of claim 1, wherein the opening is formed between a superior section and an inferior section of the patient support structure and the sections are connected together by arches which allow medial and lateral and anterior access to the body part. 4. The apparatus of claim 3, wherein the arches connect the two sections together at lateral edges thereof. 5. The apparatus of claim 3, wherein the arches curve inwardly to conform to the shape of the bore. 6. The apparatus of claim 3, wherein the arches are configured to connect together above the patient to form a generally semi-cylindrical structure. 7. The apparatus of claim 1, further including a compression device with compression plates attachable to and removable from the patient support structure when the patient is supported on the patient support structure. 8. The apparatus of claim 7, wherein the compression device includes multiple locking supports movable along linear guides in a lateral direction for locking the compression plates in a desired position. 9. The apparatus of claim 7, wherein a one of the compression plates comprises a fenestrated plate. 10. The apparatus of claim 7, wherein the compression plates are provided with at least one RF coil for MR imaging. 11. The apparatus of claim 10, wherein the compression plates are provided with at least one RF coil having a selected size corresponding to a size of the body part being imaged. 12. The apparatus of claim 1, further including at least one RF coil attachable to and removable from the patient support structure when the patient is supported on the patient support structure. 13. The apparatus of claim 1, wherein the transport stretcher includes wheels for moving the apparatus. 14. The apparatus of claim 1, wherein the transport stretcher is dockable to the MR imaging system. 15. The apparatus of claim 1, further including guides on the transport stretcher to constrain the movement of... | Hologic, Inc. | A61B005/05 A47B013/00 A47B007/00
US07867235 | B2 | 2011-01-11 | System and method for joint restoration by extracapsular means | 2006-06-14 | A system and method for joint restoration by extracapsular means includes an actuator operable to apply a force to a portion of a bone to effect a change in the joint space geometry. One embodiment of the system includes an actuator operable to apply a cyclic loading to subchondral bone of a femur, wherein loads of a predetermined magnitude are alternately applied and released. Between periods of cyclic loading, rest periods are provided where no load is applied. Over time, the femoral joint surface is remodeled in accordance with the location, direction, magnitude, and frequency of the loading. | 1. A method for effecting displacement of bone for restoration of a skeletal joint using a bone restoration system, the joint including a portion of the bone having an articular surface thereon and a mating member configured to cooperate with the articular surface of the bone, the articular surface being disposed at a first position relative to the mating member, the bone restoration system including an actuator operable for cyclic, linear movement, the actuator having an extendable component and a secondary portion, the bone restoration system further including a housing configured for attachment to the bone and for securing the actuator, the method comprising: engaging the housing in the bone at a position generally opposite an underside of the articular surface;securing the actuator to the housing such that the extendable component is proximate the underside of the articular surface;operating the actuator such that the extendable component applies at least one cyclic loading condition to the underside of the articular surface, thereby facilitating a structural change in the bone proximate the articular surface;monitoring the bone for structural changes;moving the secondary portion of the actuator relative to the housing toward the underside of the articular surface when a structural change in the bone has occurred; andremoving application of the at least one cyclic loading condition when the bone has achieved a certain structural change, including a change in the position of the articular surface relative to the mating member. 2. The method of claim 1, wherein the housing includes at least one external thread for engaging the bone, and an aperture therethrough having at least one internal thread for receiving the actuator, the secondary portion of the actuator including at least one external thread for engaging the at least one internal thread in the actuator, the step of engaging the housing in the bone including creating an aperture in the bone and screwing the housing into the aperture to secure the housing to the bone with the at least one external thread on the housing, thereby making accessible an internal portion of the bone generally opposite the articular surface;the step of securing the actuator to the housing including screwing the actuator into the aperture in the housing until the extendable component is at least within a contact distance of the underside of the articular surface; andthe step of moving the secondary portion of the actuator relative to the housing toward the underside of the articular surface when a structural change in the bone has occurred includes screwing the actuator further into the housing, thereby moving the extendable component back to within a contact distance of the articular surface. 3. The method of claim 2, wherein each of the at least one cyclic loading condition is defined by at least a respective force amplitude and a respective frequency of force application. 4. A method for effecting displacem... | - | A61F005/00
US08168051 | B2 | 2012-05-01 | Sensor for determination of glucose | 2007-10-30 | A process for the manufacture of small sensors with reproducible surfaces, including electrochemical sensors. One process includes forming channels in the surface of a substrate and disposing a conductive material in the channels to form an electrode. The conductive material can also be formed on the substrate by other impact and non-impact methods. In a preferred embodiment, the method includes cutting the substrate to form a sensor having a connector portion and a transcutaneous portion, the two portions having edges that define one continuous straight line. | 1. An analyte sensor, comprising: a substrate having a proximal end portion having a width defined by a first side and a second side, and a distal end portion having a width defined by a third side and a fourth side, the width of the distal end portion being less than the width of the proximal end portion, the distal end portion being adjacent to the proximal end portion and extending therefrom;one or more electrodes formed on the substrate, the one or more electrodes including an analyte-response enzyme and a mediator, wherein at least one of the analyte-response enzyme and the mediator is chemically bound to a polymer disposed on the one or more electrodes; andat least two traces separated by a distance of about 25 microns to about 250 microns, wherein at least one trace has a width of about 25 microns to 250 microns; wherein the flexibility of the sensor is varied based on a composition or thickness of the sensor; andwherein the first side and the third side together form a continuous straight edge and the second side and the fourth side do not form a continuous straight edge. 2. The sensor of claim 1, wherein the at least two traces are separated by a distance of about 100 microns or less. 3. The sensor of claim 2, wherein the at least one trace has a width of about 75 microns or less. 4. The sensor of claim 3, wherein the at least one trace has a width of about 50 microns or less. 5. The sensor of claim 1, wherein the sensor is coupled to a skin piercer. 6. The sensor of claim 5, wherein the sensor separates from the skin piercer after positioning. 7. The sensor of claim 1, comprising a varying width. 8. The sensor of claim 1, wherein the one or more electrodes includes two or more electrodes. 9. The sensor of claim 8, wherein two of the two or more electrodes are working electrodes. 10. The sensor of claim 1, wherein the at least two traces are on the same side of the substrate. 11. The sensor of claim 1, wherein the flexibility of the sensor is varied along a continuum of the sensor. 12. The sensor of claim 1, wherein the sensor includes a code applied to the sensor. 13. The sensor of claim 12, wherein the code is printed on a substrate of the sensor. 14. The sensor of claim 12, wherein the code includes one or more of a batch number, a quantity of chemistry applied to the sensor, or calibration data. 15. The sensor of claim 12, wherein the code is applied to the sensor during manufacturing. | Abbott Diabetes Care Inc. | G01N027/327
US07918226 | B2 | 2011-04-05 | Method and system for detecting breathing tube occlusion | 2007-04-10 | A method for identifying an occlusion in a breathing tube of a ventilator system is disclosed herein. The method includes obtaining data related to a respiratory signal, transferring the data to a processor, and implementing the processor to evaluate the data in a manner adapted to automatically identify the presence of an occlusion in the breathing tube. A corresponding ventilator system is also disclosed. | 1. A method for identifying an occlusion in a breathing tube of a ventilator system comprising: obtaining respiratory data related to the respiration of a patient;transferring the respiratory data to a processor; andimplementing the processor to evaluate the respiratory data to automatically identify the presence of an occlusion in the breathing tube, the evaluation comprising: identifying a first data section in the respiratory data that represents an unobstructed respiratory cycle for the patent;identifying a second data section in the respiratory data that represents a respiratory cycle for the patient for which the presence of the occlusion in the breathing tube is to be determined;calculating a first rate of change of at least a portion of the first data section;calculating a second rate of change of at least a portion of the second data section; andcomparing the second rate of change to the first rate of change to determine if the second rate of change differs from the first rate of change by more than a threshold amount. 2. The method of claim 1, wherein the step of obtaining respiratory data includes implementing one or more pressure sensors or flow sensors to obtain data. 3. The method of claim 1, wherein the step of obtaining respiratory data includes obtaining pressure data. 4. The method of claim 3, wherein the step of implementing the processor to evaluate the respiratory data includes comparing pressure data obtained from a first location in the breathing tube with pressure data obtained from a second location in the breathing tube. 5. The method of claim 1, wherein the step of implementing the processor to evaluate the respiratory data further includes identifying relative minimum values and/or relative maximum values in the second data section and comparing the relative minimum and/or relative maximum values in the second data section to the relative minimum and/or relative maximum values in the first data section. 6. The method of claim 1, wherein the step of obtaining respiratory data includes obtaining flow data. 7. The method of claim 1, further comprising sounding an alarm if the processor identifies the presence of an occlusion in the breathing tube. 8. A method for identifying an occlusion in a breathing tube of a ventilator system comprising: obtaining flow data related to the respiration of a patient;transferring the flow data to a processor; andimplementing the processor to evaluate the flow data to automatically identify the presence of an occlusion in the breathing tube, the evaluation comprising: identifying a first data section in the flow data that represents an unobstructed respiratory cycle for the patent;identifying a second data section in the flow data that represents a respiratory cycle for the patient for which the presence of the occlusion in the breathing tube is to be determined; andcomparing the second data section to the first data section to determine if the second data section differs from the f... | General Electric Company | A61M011/00
US08469895 | B2 | 2013-06-25 | Deriving central aortic systolic pressure and analyzing arterial waveform data to derive central aortic systolic pressure values | 2007-09-24 | A method of deriving central aortic systolic pressure comprises (a) creating a set having a predetermined number of blood pressure measurements, the set representative of an arterial waveform; (b) determining an integer interval value; (c) averaging a series of consecutive blood pressure measurement readings in the set equal to the integer interval value commencing from the fth blood pressure measurement in the set; (d) storing the averaged value in a central aortic systolic pressure set; and (e) setting the central aortic systolic pressure as the highest value in the central aortic pressure set. Steps (c) and (d) are repeated with the value of f being incremented by 1 each time until the value of f plus the integer interval value equals the predetermined number of blood pressure measurements in the set. | 1. A system for deriving central aortic systolic pressure comprising: an arterial waveform measuring device; anda processing unit;wherein the arterial waveform measuring device takes a set of blood pressure measurements at predetermined intervals until at least one arterial waveform is represented by the set of blood pressure measurements taken, the set of blood pressure measurements representative of one arterial waveform then being communicated to the processing unit which:a) determines an integer interval value;b) averaging a series of consecutive blood pressure measurement readings in the set equal to the integer interval value commencing from an fth blood pressure measurement in the set;c) stores the averaged value in a central aortic systolic pressure set; andd) sets the central aortic systolic pressure as the highest value in the central aortic pressure set;wherein steps (b) and (c) are repeated with the value of f commencing at 1 and being incremented by 1 each time until the value of f plus the integer interval value equals the number of blood pressure measurements in the set. 2. A computer readable medium that, in use, is non-transitory, the computer readable medium having embodied thereon software comprising computer code, the computer code comprising instructions to derive central aortic systolic pressure, the instructions comprising: (a) creating a set having a predetermined number of blood pressure measurements, the set representative of an arterial waveform;(b) determining an integer interval value;(c) averaging a series of consecutive blood pressure measurement readings in the set equal to the integer interval value commencing from an fth blood pressure measurement in the set;(d) storing the averaged value in a central aortic systolic pressure set; and(e) setting the central aortic systolic pressure as the highest value in the central aortic pressure set;wherein a value of f corresponds with the fuh blood pressure measurement, steps (c) and (d) are repeated with the value of f commencing at 1 and being incremented by 1 each time until the value of f plus the integer interval value equals the predetermined number of blood pressure measurements in the set. 3. A computer readable medium that, in use, is non-transitory, the computer readable medium having embodied thereon software comprising computer code, the computer code comprising instructions to analyze arterial waveform data to derive central aortic systolic pressure values, the instructions comprising: (a) receiving an arterial waveform dataset;dividing each arterial waveform in the arterial waveform dataset into a representative set of blood pressure measurements having a predetermined number;(b) determining an integer interval value for the set being processed;(c) averaging a series of consecutive blood pressure measurement readings in the set equal to the integer interval value commencing from an fth blood pressure measurement in the set;(d) storing the averaged value in a... | Healthstats International Pte Ltd | A61B005/02
US08456301 | B2 | 2013-06-04 | Analyte monitoring system and methods | 2008-05-08 | Methods and systems for providing data communication in medical systems are disclosed. | 1. An analyte monitoring system, comprising: a data processing unit; anda control unit in wireless communication with the data processing unit, wherein the control unit is configured to establish a communication range with the data processing unit when the control unit is placed within a predetermined distance from the data processing unit, wherein the control unit is configured to transmit one or more predefined commands to the data processing unit when the communication range is established, and wherein the data processing unit is configured to perform one or more predefined functions in response to the received one or more predefined commands and further configured to transmit a response data packet based, at least in part, on the one or more predefined commands; and wherein, in response to receiving the one or more predefined commands, the data processing unit is configured to retrieve identification information and to transmit the identification information to the control unit; andwherein the control unit is further configured to generate a communication key and to transmit the generated communication key to the data processing unit upon receipt of the identification information. 2. The system of claim 1 wherein the data processing unit includes a close proximity receiver coupled to an antenna for receiving the one or more predefined commands from the control unit. 3. The system of claim 1 wherein the data processing unit is configured to transmit one or more signals related to a monitored analyte level in response to the received one or more predefined commands. 4. The system of claim 1 wherein the predetermined distance includes a distance of less than one foot. 5. The system of claim 1 including a memory unit. 6. The system of claim 5 wherein the memory unit includes an EEPROM. 7. The system of claim 1 wherein the control unit includes an application specific integrated circuit (ASIC) configuration. 8. The system of claim 1 including an analyte sensor coupled to the data processing unit, wherein the data processing unit receives one or more signals from the analyte sensor. 9. The system of claim 8 wherein the control unit and the data processing unit are in wireless communication using a radio frequency (RF) communication protocol. 10. A method, comprising: positioning a controller unit within a predetermined transmission range from a transmitter unit to establish a communication range between the controller unit and the transmitter unit;transmitting one or more predefined commands from the controller unit to the transmitter unit when the controller unit is placed within the predetermined transmission range;retrieving from a memory identification information upon receipt of the one or more predefined commands, and transmitting the retrieved identification information to the control unit;receiving a response data packet from the transmitter unit in response to the transmitted one or more predefined commands when the controller uni... | Abbott Diabetes Care Inc. | G08B001/08
US08486371 | B2 | 2013-07-16 | Quantitative two-photon flow cytometry | 2005-09-16 | A method and apparatus of multi-dye analysis of particles using flow cytometer. The method includes dying particles to be detected using two or more dyes; urging the particles through a capillary in a non-uniform flow; exciting a first of the particles within the capillary using a multiphoton excitation laser beam causing the two or more dyes each to fluoresce thereby producing a first output signal and a second output signal respectively; and detecting the first output signal and the second output signal. A second of the particles within the capillary being excited using the multiphoton excitation laser beam causing the two or more dyes each to fluoresce thereby producing a third output signal and a forth output signal respectively. The method finally includes comparing a ratio of the first output signal and the second output signal to a ratio of the third output signal and the forth output signal to detect a desired change in the particles. | 1. A method of multi-dye analysis of particles using flow cytometer, said method comprising: labeling particles to be detected using two or more dyes, said two or more dyes comprising a trigger dye and a reporter dye, said trigger dye is selected from carboxyfluorescein succinimidyl ester (CFSE) and fluorescein isothiocyanate (FITC), said reporter dye is DeepRed;suspending said particles within a non-uniform flow in a capillary;exciting a first of said particles within said capillary using a multiphoton excitation laser beam causing said two or more dyes each to fluoresce thereby producing a first output signal and a second output signal respectively;detecting said first output signal and said second output signal;exciting a second of said particles within said capillary using said multiphoton excitation laser beam causing said two or more dyes each to fluoresce thereby producing a third output signal and a fourth output signal respectively; andcomparing a ratio of said first output signal and said second output signal to a ratio of said third output signal and said fourth output signal to detect a desired change in said particles. 2. The method according to claim 1 wherein said first output signal has a first wavelength and said second output signal has a second wavelength. 3. The method according to claim 1 wherein said suspending said particles within said non-uniform flow in said capillary comprises creating said non-uniform flow using a pump. 4. The method according to claim 1 wherein said suspending said particles within said non-uniform flow in said capillary comprises suspending said particles within a non-uniform, pulsatile flow. 5. The method according to claim 1 wherein said suspending said particles within said non-uniform flow in said capillary comprises suspending bacteria within said non-uniform flow. 6. The method according to claim 1 wherein said suspending said particles within said non-uniform flow in said capillary comprises suspending cancer cells within said non-uniform flow. 7. The method according to claim 1 wherein said suspending said particles within said non-uniform flow in said capillary comprises suspending stem cells within said non-uniform flow. 8. The method according to claim 1 wherein said suspending said particles within said non-uniform flow in said capillary comprises suspending a virus within said non-uniform flow. 9. The method according to claim 1 wherein said suspending said particles within said non-uniform flow in said capillary comprises suspending a particle having a diameter of about 1 nanometers within said non-uniform flow. 10. The method according to claim 1 wherein said suspending said particles within said non-uniform flow is performed in vivo. 11. A method of multi-dye analysis of particles using flow cytometer, said method comprising: labeling particles to be detected using two or more dyes, a first of said two or more dyes being selected from carboxyfluorescein succinimidyl ester (C... | The Regents of The University of Michigan | A61K049/00
US08547248 | B2 | 2013-10-01 | Implantable zero-wire communications system | 2006-09-01 | The present invention provides implantable systems that communicate wirelessly with each other using a unique format that enables devices configurations and applications heretofore not possible. Embodiments of the present invention provide communication apparatuses and methods for exchanging information with implantable medical devices. In some embodiments, two implantable devices communicate with each other using quasi-electrostatic signal transmission in a long wavelength/low frequency electromagnetic band, with the patient's body acting as a conductive medium. | 1. A system for communicating information within a body of a patient, the system comprising: (a) a first device comprising a transmitter configured to transmit a quasi-electrostatic signal via a quasi-electrostatic coupling to the body of the patient; and(b) a second device comprising a receiver configured to receive the transmitted quasi-electrostatic signal via a quasi-electrostatic coupling to the body of the patient;wherein the quasi-electrostatic signal is transmitted between the first and second device within the body of the patient and wherein the body of the patient is the conducting medium for the quasi-electrostatic signal. 2. The system according to claim 1, wherein said system includes at least a third device which is configured to at least one of: (i) transmit a quasi-electrostatic signal via a quasi electrostatic coupling to the body of the patient; and(ii) receive the transmitted quasi-electrostatic signal via a quasi electrostatic coupling to the body of the patientwherein the quasi-electrostatic signal is transmitted between the first and second device within the body of the patient and wherein the body of the patient is the conducting medium for the quasi-electrostatic signal;wherein the quasi-electrostatic signal is transmitted and received by the third device within the body of the patient and wherein the body of the patient is the conducting medium for the quasi-electrostatic signal. 3. The system according to claim 1, wherein said quasi-electrostatic signal comprises data produced by said first device. 4. The system according to claim 1, wherein said quasi-electrostatic signal comprises an identifier of said first device. 5. The system according to claim 1, wherein said quasi-electrostatic signal is used by said second device as a power source. 6. The system according to claim 1, wherein said first device is an implantable medical device. 7. The system according to claim 6, wherein said first device includes a sensor. 8. The system according to claim 7, wherein said sensor is a blood flow sensor. 9. The system according to claim 8, wherein said blood flow sensor comprises a resistivity sensor. 10. The system according to claim 8, wherein said blood flow sensor comprises an electromagnetic blood flow sensor. 11. The system according to claim 6, wherein said first device includes an effector. 12. The system according to claim 11, wherein said effector comprises an electrode. 13. The system according to claim 1, wherein the first device comprises an ingestible medical device. 14. The system according to claim 1, wherein the first device is an external device configured to contact a surface of the body of the patient and is capable of detecting and/or generating the quasi-electrostatic signal through the skin of the body of the patient. 15. The system according to claim 1, wherein the second device is an implantable medical device. 16. The system according to claim 15, wherein the second device is further configur... | Proteus Digital Health, Inc. | G08C017/00
US08287470 | B2 | 2012-10-16 | Swallowing test apparatus | 2008-05-30 | Disclosed is a swallowing test apparatus including a laryngeal part displacement detection apparatus for detecting a distance of two positions located laterally at a laryngeal part of a subject, a swallowing sound detection apparatus for detecting a swallowing sound of the patient during swallowing, a display, a processing apparatus for displaying a waveform representing the distance of the two positions at the laryngeal part generated based on a signal obtained from the laryngeal part displacement detection apparatus, and a waveform representing the swallowing sound generated based on a signal obtained from the swallowing sound detection apparatus on the display. | 1. A swallowing test apparatus comprising: a laryngeal part displacement detection apparatus configured to detect a distance between two positions located laterally at a laryngeal part of a patient;a swallowing sound detection apparatus configured to detect a swallowing sound of the patient during swallowing;a processing apparatus configured to generate a waveform representing the distance of the two positions at the laryngeal part generated based on a signal obtained from the laryngeal part displacement detection apparatus, and a waveform representing the swallowing sound generated based on a signal obtained from the swallowing sound detection apparatus;a display configured to display the distance waveform and the swallowing sound waveform in response to the processing apparatus generating the distance waveform and the swallowing sound waveform; andan operation apparatus which is to be operated by the patient or a person who examines the patient when swallowing, the processing apparatus being configured to:detect a peak time of the swallowing sound when the swallowing sound is at maximum based on the signal obtained from the swallowing sound detection apparatus, andcalculate an amount of characteristics of the swallowing based on at least one of a time interval between a time when the operation apparatus is operated and the peak time of the swallowing sound, a time when the distance between the two positions at the laryngeal part is at minimum before the peak time of the swallowing sound, and a time when the distance between the two positions at the laryngeal part is at minimum after the peak time of the swallowing sound, whereinthe display is furthered configured to display the calculated amount of characteristics of the swallowing. 2. A swallowing test apparatus according to claim 1, wherein the swallowing sound detection apparatus being formed to include a microphone configured to be positionally disposed at the laryngeal part of the patient, the laryngeal part displacement detection apparatus being formed to include a oscillation coil and a detection coil configured to be positionally disposed at the laryngeal part of the patient,the processing apparatus is configured to detect the peak time of the swallowing sound based on an envelope of a signal from the microphone, and to calculate at least one ofthe time when the distance between the two positions at the laryngeal part is at the minimum before the peak time of the swallowing sound andthe time when the distance between the two positions at the laryngeal part is at the minimum after the peak time of the swallowing sound based on an envelope of a signal output from an amplifier circuit which amplifies voltage induced in the detection coil by the oscillation coil. 3. A swallowing test apparatus according to claim 2, further comprising: a full-wave rectifier and a low pass filter formed to full-wave rectify and filter the signal from the microphone so as to remove high frequency components,... | Hitachi, Ltd. | A61B007/00 A61B005/103 A61B005/117
US08246540 | B2 | 2012-08-21 | System for determining insulin dose using carbohydrate to insulin ratio and insulin sensitivity factor | 2008-07-01 | An apparatus and method are provided for determining a patient's carbohydrate to insulin ratio (CIR) and insulin sensitivity factor (ISF), and using these values, along with values for current blood glucose level and deviation from target blood glucose level, for determining insulin dose in view of carbohydrate intake during a particular time period. The apparatus and method employ algorithms that can be implemented in any of a personal computer, personal data assistant, hand held computing device, blood glucose monitor, infusion pump, medication delivery pen, meter, calculator, among other therapeutic, diagnostic or informational devices used for managing a patient's blood glucose levels. | 1. A computer-implemented method of determining factors used for insulin therapy comprising the steps of: storing at least one data set in a memory device, the data set comprising blood glucose readings taken, respectively, at substantially the beginning of and substantially the end of a selected period of time, a value corresponding to the patient's food intake for the selected period, and the amount of insulin administered to the patient during the selected period;calculating a delta blood glucose level corresponding to the difference between the blood glucose readings taken, respectively, at substantially the beginning of and substantially the end of the selected period; anddetermining, using the computer, a correct insulin amount using the delta blood glucose level divided by an insulin sensitivity factor, and the result of the division being added to the amount of insulin administered to the patient during the selected period. 2. A method as claimed in claim 1, wherein the selected period of time corresponds to a period of time occurring within one day. 3. A method as claimed in claim 1, wherein the selected period of time comprises one meal time selected from the group consisting of breakfast, lunch, dinner, or snack. 4. A method as claimed in claim 1, further comprising the step of generating an initial estimate of the insulin sensitivity factor using at least one number selected from the range of numbers equal to or between 1500 and 1800, and dividing the at least one number by the daily insulin dose. 5. A method as claimed in claim 1, further comprising the steps of: determining a plurality of daily carbohydrate to insulin ratios corresponding to respective ones of a plurality of days, the daily carbohydrate to insulin ratios each being based on the correct insulin amount and the value corresponding to the patient's food intake for the selected period from the data set of the corresponding one of the plurality of days; anddetermining a carbohydrate to insulin ratio for the selected period from an average of the daily carbohydrate to insulin ratios. 6. A method as claimed in claim 5, wherein the step of determining a carbohydrate to insulin ratio for the selected period comprises the step of calculating a weighted average of the plurality of daily carbohydrate to insulin ratios. 7. A method as claimed in claim 1, further comprising the step of determining a plurality of daily carbohydrate to insulin ratios corresponding to respective ones of a plurality of days, the daily carbohydrate to insulin ratios each being based on the correct insulin amount divided into the value corresponding to the patient's food intake for the selected period from the data set of the corresponding one of the plurality of days. 8. A method as claimed in claim 7, further comprising the step of determining a carbohydrate to insulin ratio for the selected period from an average of the daily carbohydrate to insulin ratios. 9. A method as claimed in claim 8, ... | Becton, Dickinson and Company | A61M031/00 A61B005/05
US08021303 | B2 | 2011-09-20 | System for extracting morphological information through a perfusion assessment process | 2007-06-25 | A perfusion assessment system is proposed. The system includes means for providing an echo-power signal indicative of a reperfusion of a contrast agent in a body-part of a living subject following destruction of a significant portion of the contrast agent in the body-part, means for associating the echo-power signal to a perfusion function with an S-shape based on a plurality of elementary perfusion functions with said S-shape each one for a corresponding value of at least one perfusion parameter, the elementary perfusion functions being weighted according to a probability density distribution of the at least one perfusion parameter, wherein the S-shape includes an initial portion with substantially zero first derivatives, a final portion with substantially zero first derivatives, and a central portion between the initial portion and the final portion changing monotonically from a value of the initial portion to a value of the final portion, means for determining at least one shape indicator of the probability density distribution, and means for comparing the at least one shape indicator with at least one predetermined further shape indicator to identify morphological characteristics of the body-part according to a result of the comparison. | 1. A perfusion assessment system including: a device configured to:provide an echo-power signal indicative of a reperfusion of a contrast agent in a body-part of a living subject following destruction of a significant portion of the contrast agent in the body-part; anda processor configured to:associate the echo-power signal to a perfusion function with an S-shape based on a plurality of elementary perfusion functions with said S-shape each one for a corresponding value of at least one perfusion parameter, the elementary perfusion functions being weighted according to a probability density distribution of the at least one perfusion parameter, wherein the S-shape comprises an initial portion with substantially zero first derivatives, a final portion with substantially zero first derivatives, and a central portion between the initial portion and the final portion changing monotonically from a value of the initial portion to a value of the final portion,determine at least one shape indicator of the probability density distribution,compare the at least one shape indicator with at least one predetermined further shape indicator to identify morphological characteristics of the body-part according to a result of the comparison. 2. The system according to claim 1, wherein the at least one shape indicator is indicative of a skewness of the probability density distribution. 3. The system according to claim 2, wherein each elementary perfusion function is a cumulative normal distribution function based on a first predetermined parameter indicative of an echo-power signal measure sensitivity and a second predetermined parameter indicative of a contrast agent destruction extent. 4. The system according to claim 2, wherein the processor is further configured to process the echo-power signal to be proportional to a concentration of the contrast agent in the body-part. 5. The system according to claim 2, wherein the associating comprises calculating a temporal sequence of samples representative of the echo-power signal in a region of interest, and fitting the samples by the perfusion function. 6. The system according to claim 1, wherein each elementary perfusion function is a cumulative normal distribution function based on a first predetermined parameter indicative of an echo-power signal measure sensitivity and a second predetermined parameter indicative of a contrast agent destruction extent. 7. The system according to claim 6, wherein the processor is further configured to process the echo-power signal to be proportional to a concentration of the contrast agent in the body-part. 8. The system according to claim 6, wherein the associating comprises for calculating a temporal sequence of samples representative of the echo-power signal in a region of interest, and fitting the samples by the perfusion function. 9. The system according to claim 1, wherein the processor is further configured to process the echo-power signal to be proportional to a concentr... | Bracco Research SA | A61B008/14 A61B008/06
US08126690 | B2 | 2012-02-28 | Algorithms to predict clinical response, adherence, and shunting with thiopurines | 2007-05-18 | A method of using a variable set from complete blood counts and blood chemistry panels to generate a machine learned algorithm for determining the effectiveness of thiopurine treatment on inflammatory bowel disease (IBD) patients using CART, boosted trees, random forest classification, RuleFit and/or logistic regression analysis. | 1. A method of determining the effectiveness of thiopurine treatment in a patient receiving the thiopurine treatment for inflammatory bowel disease (IBD) comprising: receiving on a computer device a set of data inputs related to characteristics of a patient receiving a thiopurine treatment for inflammatory bowel disease; andcalculating on the computer device a clinical response probability based on the set of data inputs using an algorithm that applies a weight to each data input of the set of data inputs, wherein each weight relates to a quantified importance of a particular data input associated with a fixed set of patient data from patients receiving a thiopurine treatment for inflammatory bowel disease in which the clinical response of the patients is known,wherein the set of data comprises at least three data inputs selected from the group consisting of hemoglobin (HGB), hematocrit (HCT), platelet count (PLT), red blood cell count (RBC), mean corpuscular hemoglobin (MCH), mean corpuscular hemoglobin concentration (MCHC), red cell distribution width (RDW), mean platelet volume (MPV), sodium level (SOD), potassium level (POT), chloride (CHLOR), carbon dioxide (CO2), urea nitrogen (UN), creatinine level (CREAT), glucose level (GLUC), calcium level (CAL), protein level (PROT), albumin level (ALB), aspartate transaminase (AST), alanine transaminase (ALT), alkaline phosphatase (ALK), bilirubin total (TBIL), neutrophil count (NEUT), monocyte count (MONO), basophil count (BASO), and age of patient (Age). 2. The method of claim 1, wherein the set of data inputs include at least four of red blood cell volume distribution width (RDW), eosinophil count (EOS), albumin level (ALB), alkaline phosphatase level (ALK), sodium level (SOD), mean corpuscular hemoglobin content (MCH), neutrophil count (NEUT), platelet count (PLT), creatinine level (CREAT), hematocrit (HCT), or age of patient (AGE). 3. The method of claim 1, wherein the probability is calculated using an algorithm generated by a modeling analysis selected from the group consisting of a logistic regression analysis, a classification and regression tree (CART) analysis, a random forest analysis, RuleFit, and a boosted tree regression. 4. The method of claim 1, wherein determining the clinical response probability is calculated based on a statistical regression model, and the set of data inputs for calculating the clinical response probability includes white blood cell count (WBC), red blood cell volume distribution width (RDW), eosinophil count (EOS), albumin level (ALB), alkaline phosphatase level (ALK), sodium level (SOD), and age of patient. 5. The method of claim 1, wherein determining the adherence probability is calculated based on a model derived from a logistic regression using white blood cell count (WBC), mean corpuscular hemoglobin content (MCH), neutrophil count (NEUT), red blood cell volume distribution width (RDW), and age of patient. 6. The method of claim 1, wherein determining ... | The Regents of the University of Michigan | G06G007/58 G01N033/48
US07938775 | B2 | 2011-05-10 | Device, system, and method for in-vivo analysis | 2005-06-28 | A device, system and method for in-vivo analysis. An autonomous in-vivo device may include a magnet to detain at least a portion of a sample collected from a body lumen; a sensor to sense a property of the detained sample portion; and a transmitter to transmit data of the sensed property. | 1. An autonomous in-vivo device comprising: a magnetized reagent to be mixed, within said autonomous in-vivo device, with a sample collected from a body lumen;a detector reagent to be mixed, within said autonomous in-vivo device, with said sample; whereby particles within a first portion of said sample, which contains a substance to be detected by said detector reagent, bind to said magnetized reagent and to said detector reagent, and particles within a second portion of said sample, which does not contain the substance to be detected by said detector reagent, do not bind to said magnetized reagent and to said detector reagent;a magnet to detain said first portion of said sample, said magnet not detaining said second portion of said sample;a sensor to sense a property of the detained sample portion; anda transmitter to transmit data of the sensed property. 2. The autonomous in-vivo device of claim 1, comprising an inlet to transfer said sample into said autonomous in-vivo device. 3. The autonomous in-vivo device of claim 1, comprising a duct to transfer said sample within said autonomous in-vivo device. 4. The autonomous in-vivo device of claim 1, comprising a micropump to pump in said sample into said autonomous in-vivo device. 5. The autonomous in-vivo device of claim 1, wherein the magnetized reagent includes a magnetized antibody. 6. The autonomous in-vivo device of claim 1, wherein the detector reagent includes a detector antibody. 7. The autonomous in-vivo device of claim 6, wherein said detector antibody comprises a color label. 8. The autonomous in-vivo device of claim 1, comprising an outlet to transfer out of the autonomous in-vivo device a portion of said sample not detained by said magnet. 9. The autonomous in-vivo device of claim 1, comprising an illumination source to illuminate said detained portion, and wherein said sensor comprises an imager to acquire in-vivo an image of said detained portion. 10. The autonomous in-vivo device of claim 1, comprising a first reservoir to store said magnetized reagent and a second reservoir to store said detector reagent. 11. The autonomous in-vivo device of claim 1, comprising an imager. 12. The autonomous in-vivo device of claim 1, wherein said autonomous in-vivo device comprises a swallowable capsule. 13. A system comprising: an autonomous in-vivo device including at least a magnetized reagent to be mixed, within said autonomous in-vivo device, with a sample collected from a body lumen,a detector reagent to be mixed, within said autonomous in-vivo device, with said sample; whereby particles within a first portion of said sample, which contains a substance to be detected by said detector reagent, bind to said magnetized reagent and to said detector reagent, and particles within a second portion of said sample, which does not contain the substance to be detected by said detector reagent, do not bind to said magnetized reagent and to said detector reagent;a magnet to detain said first ... | Given Imaging, Ltd. | A61B005/00
US07873398 | B2 | 2011-01-18 | Saccharide-measuring fluorescent monomer, saccharide-measuring fluorescent sensor substance, and implantable, saccharide-measuring sensor | 2008-08-12 | A saccharide-measuring fluorescence sensor substance including a copolymer of at least the following two compounds (I) and (II) is provided: (I) a fluorescent monomer compound represented by the following formula (1): formula (1); wherein: Q, Q′ and D3 may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups; andD1, D2 and D4 each represent a substituent, wherein at least one of D1, D2 and D4 is a substituent group comprising a vinyl group at an end thereof, and wherein the substituent group comprising a vinyl group at an end thereof enables the fluorescent monomer compound to be soluble in water, and(II) at least one hydrophilic, polymerizable monomer having a vinyl group. | 1. A saccharide-measuring fluorescence sensor substance comprising a copolymer of at least the following two compounds (I) and (II): (I) a fluorescent monomer compound represented by the following formula (1):wherein: Q, Q′ and D3 may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups;D1, D2 and D4 each represent a substituent, wherein at least one of D1, D2 and D4 is a substituent group comprising a vinyl group at an end thereof, and wherein the substituent group comprising a vinyl group at an end thereof enables the fluorescent monomer compound to be soluble in water, and (II) at least one hydrophilic, polymerizable monomer comprising a vinyl group, wherein said polymerizable monomer (II) is a compound represented by the following formula (6):wherein A is a hydrogen atom or a methyl group, and U and U′ may be the same or different and are each a substituted or unsubstituted alkyl group. 2. A saccharide-measuring fluorescence sensor substance comprising a copolymer of at least the following two compounds (I) and (II): (I) a fluorescent monomer compound represented by the following formula (1):wherein: Q and Q′ may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups;and in D1, D2, D3 and D4,(i) D1 and D2 may be the same or different and are each a substituted or unsubstituted alkyl group, D3 is a hydrogen atom, and D4 is a substituent group represented by the following formula (2), or(ii) D1 and D2 may be the same or different and are each a substituent group represented by the following formula (3), and D3 and D4 may be the same or different, may be combined together into a fused ring, and are each a substituent selected from the group consisting of a hydrogen atom, a halogen atom, a hydroxyl group, and substituted or unsubstituted alkyl, acyl, oxyalkyl, carboxyl, carboxylate ester, carboxamido, cyano, nitro, amino and aminoalkyl groups:wherein: X, in a case of the definition (i), is a substituent group selected from the group consisting of —COO—, —OCO—, —CH2NR—, —CH2S—, —CH2O—, —NR—, —NRCO—, —CONR—, —SO2NR—, —NRSO2—, —O—, —S—, —SS—, —NRCOO—, —OCONR— and —CO—, and in a case of the definition (ii), represents a C1 to C30 alkylene group comprising at least one substituent group selected from the group consisting of —COO—, —OCO—, —CH2NR—, —NR—, —NRCO—, —CONR—, —SO2NR—, —NRSO2—, —O—, —S—, —SS—, —NRCOO—, —OCONR— and —CO—, in which R represents a hydrogen atom or a substituted or unsubstituted alkyl group,Y is a... | Terumo Kabushiki Kaisha | A61B005/00
US08118752 | B2 | 2012-02-21 | Apparatus and methods for mapping retinal function | 2008-06-09 | The present invention provides an electrode array device for simultaneously detecting electrical potentials at five or more locations on the anterior surface of an eye. The device comprises a dielectric lens substrate having a concave inner surface conforming to the anterior surface of the eye, and at least five recording electrodes positioned in relation to the inner surface of the lens substrate so as to make electrical connection with the anterior surface of the eye when the lens substrate is placed on the anterior surface of eye. Each recording electrode is in electrically conductive communication with a corresponding conductive contact, there being one conductive contact for each recording electrode. Each conductive contact is adapted for operable connection to signal processor, and each conductive contact is electrically insulated from the anterior surface of the eye. A computational method for analyzing electrophysiological potentials recorded at five or more locations on the anterior surface of the eye, which reflect the spatial distribution of activity of the retina, is also provided. | 1. An electrode array device for simultaneously detecting electrical potentials at five or more locations on the anterior surface of an eye, the device comprising: a transparent or translucent dielectric lens substrate having a concave inner surface conforming to the cornea on the exterior surface of the eye;at least 5 recording electrodes positioned in relation to the inner surface of the lens substrate so as to make electrical connection with the cornea when the lens substrate is placed over the cornea on the exterior surface of the eye;wherein each recording electrode is in electrically conductive communication with a corresponding conductive contact, there being one conductive contact for each electrode, each conductive contact being adapted for operable connection to a signal processor, and each conductive contact being electrically insulated from the anterior surface of the eye. 2. The device of claim 1 wherein the lens substrate includes 5 to 500 recording electrodes. 3. The device of claim 1 wherein each recording electrode comprises a conductive material deposited on the concave surface of the lens substrate. 4. The device of claim 1 wherein each conductive contact is positioned in the peripheral region of the lens substrate, and is connected to its corresponding recording electrode via an electrically conductive trace that is electrically insulated from the surface of the eye. 5. The device of claim 4 wherein each recording electrode comprises a cylindrical through-hole formed in the lens substrate adjacent to an un-insulated portion of its corresponding conductive trace, such that an electrically conductive connection is made between the conductive trace and the anterior surface of the eye when the lens substrate is placed on the anterior surface of the eye in the presence of an amount of an electrically conductive liquid or hydrogel sufficient to fill each through-hole and connect with the un-insulated portion of the trace. 6. The device of claim 5 wherein each through-hole is lined with a conductive material in electrically conductive connection with the conductive trace. 7. The device of claim 6 wherein the conductive material lining each through-hole comprises a metal, conductive polymer, or semiconductor material. 8. The device of claim 4 wherein each conductive trace, active electrode, and contact pad comprises a metal, conductive polymer, or semiconductor material. 9. The device of claim 1 wherein the lens substrate is configured to provide corrective refractive power to correct for defects in the visual acuity of the eye. 10. The device of claim 1 wherein the lens substrate is plano-concave. 11. The device of claim 1 wherein each recording electrode comprises a conductive element passing through the lens substrate. 12. The device of claim 1 further comprising one or more scleral recording electrodes positioned to make electrical contact with the sclera of the eye when the lens substrate is placed on the anterior surf... | The Board of Trustees of the University of Illinois | A61B013/00 A61B005/04
US09138179 | B2 | 2015-09-22 | Flexible lancet | 2008-05-29 | A system for withdrawing body fluid is described which has a housing with at least one opening and at least one lancet which is elongate and substantially straight in its original and unused form. The lancet has a distal end with a lancet tip for carrying out a lancing operation. The system also includes a drive unit for driving the lancet in a drive direction and at least one guide element which makes contact with the lancet at least during the lancing operation and guides the lancet movement such that the lancet is bent relative to the drive direction. The lancet can have a high restoring force such that it returns to its original substantially straight shape after contacting the guide element. | 1. A system for withdrawing body fluid, comprising: a housing having an opening;a flexible lancet having a lancet body and a tip, the lancet being configured substantially straight before use;a drive unit for moving the lancet in a drive direction during a lancing operation; anda guide element which contacts and flexes the lancet relative to the drive direction during a movement of the lancet toward the skin for puncture, wherein the guide element is curved. 2. The system of claim 1, wherein the lancet comprises a restoring force which restores the lancet to the substantially straight configuration when the lancet is separated from the guide element. 3. The system of claim 1, wherein the lancet is flexible and substantially flat. 4. The system of claim 1, wherein the guide element comprises metal, ceramic, glass or plastic. 5. The system of claim 1, wherein the lancet is moved through the guide element during the lancing operation. 6. The system of claim 1, wherein the guide element is a rail. 7. The system of claim 6, wherein the rail comprises two rails. 8. The system of claim 1, wherein the guide element bends the lancet to a curved shape during the lancing operation. 9. The system of claim 1, wherein the guide element reorients the tip relative to the drive direction during the lancing operation. 10. The system of claim 1, wherein the guide element reorients the tip between 5° and 90° relative to the drive direction during the lancing operation. 11. The system of claim 1, wherein the guide element reorients the tip about 90° relative to the drive direction during the lancing operation. 12. The system of claim 1, wherein the guide element couples to the housing during the lancing operation. 13. The system of claim 1, wherein, prior to use, the lancet is disposed in a lancet magazine. 14. The system of claim 13, wherein the guide element is connected to the magazine. 15. The system of claim 1, wherein the lancet is arranged on a carrier tape. 16. The system of claim 15, further comprising a test element arranged on the carrier tape. 17. The system of claim 16, wherein the lancet and test element are arranged side by side on the carrier tape. 18. The system of claim 16, wherein, after the bending of the lancet, the lancet contacts the test element to transfer a body fluid sample. 19. The system of claim 15, wherein the lancet is aligned substantially parallel to a longitudinal direction of the carrier tape. 20. The system of claim 15, wherein the lancet is aligned substantially transverse to a longitudinal direction of the carrier tape. 21. The system of claim 15, further comprising a stop element that is integrated into the housing or guide element, wherein, when the tip reaches a maximum puncturing depth, the stop element engages a thickened portion of the lancet body. 22. The system of claim 21, wherein the position of the stop element can be varied to adjust the puncturing depth of the tip. 23. The system of claim 15, f... | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00 A61B017/14 A61B017/32 A61B005/15 A61B005/145 A61B005/151
US08190226 | B2 | 2012-05-29 | Sensor with signature of data relating to sensor | 2008-05-06 | A sensor has codes useful for a monitor which can be authenticated as accurate. The sensor produces a signal corresponding to a measured physiological characteristic and provides codes which can be assured of being accurate and authentic when used by a monitor. A memory associated with the sensor stores both data relating to the sensor and a digital signature. The digital signature authenticates the quality of the code by ensuring it was generated by an entity having predetermined quality controls, and ensure the code is accurate. | 1. A pulse oximeter sensor comprising: an emitter capable of emitting light into tissue;a detector configured to detect light from the emitter after passing through at least a portion of the tissue and capable of converting the detected light into an electrical signal for use by a monitor; anda memory storing a digital signature, wherein the digital signature comprises functional data residing completely within the digital signature, the functional data capable of facilitating proper operation of the monitor with the pulse oximeter sensor. 2. The pulse oximeter sensor of claim 1, wherein the functional data comprises a sensor-dependent calibration curve coefficient for use by the monitor to calculate oxygen saturation. 3. The pulse oximeter sensor of claim 1, wherein the emitter comprises one or more light emitting diodes. 4. The pulse oximeter sensor of claim 1, wherein the functional data comprises a threshold used to indicate when the sensor is not in proper contact with the tissue. 5. The pulse oximeter sensor of claim 1, comprising a thermistor, wherein the functional data comprises thermistor calibration coefficients. 6. The pulse oximeter sensor of claim 1, comprising a cable configured to communicatively couple the sensor with the monitor. 7. The pulse oximeter sensor of claim 6, wherein the memory is disposed within the cable. 8. The pulse oximeter sensor of claim 1, comprising a body housing the emitter and the detector. 9. A pulse oximeter comprising: a processing circuit;a sensor input configured to receive a signal from a pulse oximeter sensor indicative of a measured physiological characteristic;a first oximeter memory storing a digital signature received from a sensor memory of the pulse oximeter sensor;a second oximeter memory storing a signature translation program, the signature translation program configured to cooperate with the processing circuit to verify authenticity of the digital signature and to acquire functional data from the digital signature to facilitate operation of the pulse oximeter with the pulse oximeter sensor; andan output configured to present data indicative of a value of the measured physiological characteristic based at least in part upon the signal and/or the functional data. 10. The pulse oximeter of claim 9, wherein the first and second oximeter memories comprise different portions of the same physical memory. 11. The pulse oximeter of claim 9, wherein the processing circuit comprises a microprocessor. 12. The pulse oximeter of claim 9, wherein the signature translation program comprises a public key of a private key and public key pair. 13. The pulse oximeter of claim 9, wherein the functional data comprises a sensor-dependent calibration curve coefficient for use by the monitor to calculate oxygen saturation. 14. A method comprising: receiving a digital signature from a pulse oximeter sensor;verifying that the digital signature is authentic;retrieving functional data from within the digi... | Tyco Healthcare Group LP | A61B005/1455
US08273048 | B2 | 2012-09-25 | System and method for diverting flow to facilitate measurement of system parameters | 2008-04-11 | A method and various devices are disclosed that facilitates the measurement of hemodynamic parameters by injection of an indicator in an indicator dilution technique using an extracorporeal circuit connected to a patient. Specifically, the invention deals with problems caused by the spike in pressure in an extracorporeal line that result from the injection of a bolus. The method and various devices provide for diversion of blood during an indictor injection process and then return of the diverted blood back into the system after the injection is completed. The variations of the invention use diversion lines, accommodating cases and other devices that are designed to accumulate blood displaced during injection and then returning the blood to the extracorporeal circuit after injection. | 1. A system for facilitating the determination of hemodynamic parameters in a patient comprising: a. an extracorporeal circuit with a first end that connects to an arterial catheter of a patient and a second end connecting to a venous catheter of the patient;b. a fluid propelling apparatus to draw blood into the extracorporeal circuit from the first end;c. at least one injection port on the extracorporeal circuit for fluid access to the extracorporeal circuit;d. a dilution indicator bolus injector fluidly connected to the at least one injection port;e. a volume of dilution indicator in the dilution indicator bolus injector; andf. a flow volume accommodating case for diverting blood in the extracorporeal circuit while (i) the volume of the dilution indicator is being injected into the extracorporeal circuit at the at least one injection port (ii) blood is drawn into the extracorporeal circuit and (iii) an open flow path is maintained from the at least one injection port to the venous catheter, whereby the flow accommodating case includes a resilient member that will automatically expand and contract in response to pressure from the open flow path enabling pressure to be relieved from the open flow path. 2. A system for facilitating the determination of hemodynamic parameters comprising: a. an extracorporeal circuit with a first end that connects to an arterial catheter of a patient and a second end that connects to a venous catheter of the patient;b. a fluid propelling apparatus to move fluid into the extracorporeal circuit through the first end;c. at least one injection port on the extracorporeal circuit;d. a dilution indicator bolus injector fluidly connected to the at least one injection port;e. a volume of dilution indicator in the dilution indicator bolus injector; andf. a flow volume accommodating case including a resilient member that will automatically expand and contract enabling pressure relief and is in fluid communication with the extracorporeal fluid circuit at a position upstream on the extracorporeal circuit from the at least one injection port, the flow accommodating case (i) receiving accumulating fluid resulting from a constriction or interruption of flow downstream by the dilution indicator passing from the dilution indicator bolus injector into the extracorporeal circuit and (ii) releasing the accumulated fluid to flow into the extracorporeal circuit when the constriction or interruption of flow ceases, the flow accommodating case receiving accumulating fluid occurring during unrestricted flow from the at least one injection port to the second end. 3. The system of claim 1, wherein the flow volume accommodating case is fluidly intermediate the first end and the at least one injection port. 4. The system of claim 1, wherein the flow volume accommodating is upstream of the at least one injection port. 5. The system of claim 1, further comprising a one way check valve fluidly intermediate the at least one injection port and th... | Transonic Systems, Inc. | A61M037/00
US08251928 | B2 | 2012-08-28 | Method for carrying out a functional analysis of an artificial extremity | 2006-10-10 | A method for carrying out a functional analysis on a person equipped with an artificial extremity having adjustable settings. The method includes the steps of providing a sensor assembly configured as a replacement for a part of the artificial extremity and installing the sensor assembly in place of the part. Forces, accelerations and/or torques are then measured with the sensor assembly during use of the artificial extremity by the person and the settings of the artificial extremity are optimized based on these measurements. The sensor assembly is removed and the replaced part is installed back into the artificial extremity, while retaining the optimized settings. In one embodiment, the artificial extremity is a leg prosthesis having an artificial knee joint with a rotational adaptor mounted above it. The sensor assembly then replaces the rotational adaptor. | 1. A method for carrying out a functional analysis on a person equipped with an artificial extremity having adjustable settings and a removable part, the method comprising the steps of: providing a sensor assembly configured as a replacement for the removable part of the artificial extremity;installing the sensor assembly in place of the removable part;measuring at least one of forces, accelerations and torques with the sensor assembly during use of the artificial extremity by the person;optimizing the adjustable settings of the artificial extremity based on the measurements of the sensor assembly; andremoving the sensor assembly and installing the removable part back into the artificial extremity, while retaining the optimized settings. 2. The method of claim 1, wherein a first function of the artificial extremity is able to be adjusted based on the adjustable settings. 3. The method of claim 2, wherein the artificial extremity is formed in a modular fashion, including at least one removable functional part that provides a function in addition to the first function. 4. The method of claim 3, wherein the step of providing a sensor assembly further comprises configuring the sensor assembly with dimensions and connections consistent with the removable functional part and wherein the removable part of the artificial extremity comprises the removable functional part. 5. The method of claim 1, wherein the artificial extremity comprises a leg prosthesis. 6. The method of claim 5, wherein the leg prosthesis includes a rotational adaptor arranged above an artificial knee joint and wherein the sensor assembly is configured as a replacement for the rotational adaptor in the leg prosthesis. 7. A method for carrying out a functional analysis on a person equipped with an artificial extremity, a first function of the artificial extremity being able to be adjusted, with the artificial extremity being formed in a modular fashion, the artificial extremity including at least one removable functional part that provides a function in addition to the first function, the method comprising the steps of: providing a sensor assembly configured as a replacement for the removable functional part;installing the sensor assembly in place of the removable functional part;measuring at least one of forces, accelerations and torques with the sensor assembly during use of the artificial extremity by the person;optimizing adjustment settings of the first function of the artificial extremity based on the measurements of the sensor assembly; andreinstalling the removable functional part into the artificial extremity in place of the sensor assembly while retaining the optimized settings. 8. The method of claim 7, wherein the artificial extremity comprises a leg prosthesis. 9. The method of claim 8, wherein the leg prosthesis includes a rotational adaptor arranged above an artificial knee joint, and wherein the sensor assembly is configured as a replacement for the rotational a... | Otto Bock Healthcare GmbH | A61B005/103 A61B005/117 A61F002/80 G01S007/40 G12B013/00
US08047998 | B2 | 2011-11-01 | Non-invasive blood pressure determination method | 2007-04-17 | A method for estimating systolic and diastolic pressure is disclosed herein. The method includes obtaining a predetermined type of blood pressure data from a patient, and providing previously acquired blood pressure data obtained from a plurality of different subjects. The method also includes implementing the previously acquired blood pressure data to select systolic and diastolic amplitude ratios that most closely correlate with the predetermined type of blood pressure data obtained from the patient. The selected systolic and diastolic amplitude ratios are adapted to compensate for the effects of arterial compliance. The method also includes implementing the selected systolic and diastolic amplitude ratios to generate a systolic and diastolic blood pressure estimates. | 1. A method for estimating systolic blood pressure and diastolic blood pressure comprising: obtaining a pulse transit time slope from a patient;providing previously acquired blood pressure data obtained from a plurality of different subjects to a processor, said previously acquired blood pressure data adapted to convey the manner in which a systolic amplitude ratio and a diastolic amplitude ratio vary with respect to the pulse transit time slope obtained from the patient;implementing the previously acquired blood pressure data in the processor to select a systolic amplitude ratio and a diastolic amplitude ratio that most closely correlate with the pulse transit time slope obtained from the patient, wherein the selected systolic amplitude ratio and diastolic amplitude ratio are adapted to compensate for the effects of arterial compliance;implementing the selected systolic amplitude ratio in the processor to estimate a systolic blood pressure level; andimplementing the selected diastolic amplitude ratio in the processor to estimate a diastolic blood pressure level. 2. The method of claim 1, wherein said obtaining a predetermined type of blood pressure data includes implementing a non-invasive blood pressure monitor and a pulse oximeter device. 3. The method of claim 1, wherein said obtaining the pulse transit time slope includes obtaining at least two pulse transit times from the patient. 4. The method of claim 1, wherein said providing previously acquired blood pressure data obtained from the plurality of different subjects includes providing the previously acquired blood pressure data in the form of an equation. 5. The method of claim 1, wherein said providing previously acquired blood pressure data obtained from the plurality of different subjects includes providing the previously acquired blood pressure data in the form of a table. 6. A method for estimating systolic blood pressure and diastolic blood pressure comprising: providing a non-invasive blood pressure monitor comprising a cuff configured to apply a selectable pressure level to a patient;estimating a first pulse transit time in a processor at a first cuff pressure level, and estimating a second pulse transit time in the processor at a second cuff pressure level;calculating a pulse transit time ratio in the processor defined as the first pulse transit time divided by the second pulse transit time;providing blood pressure data in the processor adapted to correlate a plurality of pulse transit time ratios with a plurality of systolic amplitude ratios and a plurality of diastolic amplitude ratios;selecting one of the plurality of systolic amplitude ratios and one of the plurality of diastolic amplitude ratios that most closely correlate with the calculated pulse transit time ratio, wherein the selected systolic amplitude ratio and the selected diastolic amplitude ratio are adapted to compensate for the effects of arterial compliance;implementing the selected systolic amplitude ratio i... | General Electric Company | A61B005/022
US08475373 | B2 | 2013-07-02 | Transcutaneous analyte sensor | 2008-07-17 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A transcutaneous system for measuring an analyte concentration in a host, the system comprising: a sensor adapted for transcutaneous placement through a skin of a host;a housing adapted for placement adjacent to the skin of the host, wherein the housing comprises at least one electrical contact therein for providing an electrical connection between the sensor and sensor electronics; andan applicator configured for inserting the sensor through the housing and into the skin of the host, wherein the applicator comprises at least one needle adapted for inserting the sensor into the host, wherein insertion of the needle or retraction of the needle disposes the sensor adjacent to the electrical contact or within the electrical contact. 2. The system of claim 1, wherein the electrical contact comprises an elastomeric contact and wherein the sensor is configured to extend through the elastomeric contact such that, in use, the elastomeric contact relaxes around the sensor. 3. The system of claim 1, comprising two electrical contacts. 4. A transcutaneous system for measuring an analyte concentration in a host, the system comprising: a sensor adapted for transcutaneous placement through a skin of a host;a housing adapted for placement adjacent to the skin of the host, wherein the housing comprises a recess;an applicator comprising a needle comprising a lumen, wherein the applicator is configured for inserting the sensor through the housing and into the skin of the host, wherein the applicator comprises a base configured to fit within the recess of the housing, and wherein the applicator comprises at least one introducer adapted for inserting the sensor into the host; andan electronics unit releasably attachable to the recess in the housing, wherein the electronics unit comprises a base configured to fit within the recess of the housing, wherein the base of the applicator and the base of the electronics unit have a similar footprint. 5. The system of claim 4, wherein the footprint of the base of the applicator and the footprint of the base of the electronics unit are sized and arranged to mate with the recess. 6. The system of claim 4, wherein the applicator is configured to insert the sensor in response to an activation of an automated mechanism. 7. The system of claim 4, wherein the applicator is configured to insert the sensor in response to a manual force applied to the applicator by a user. 8. An applicator system for inserting a transcutaneous sensor into a host, the applicator comprising: a sensor configured for transcutaneous placement through a skin of a host;a needle comprising a lumen and configured for inserting the sensor into the host, wherein the sensor is disposed within the lumen of the needle prior to insertion into the host;a subassembly comprising a housing adapted for placement adjacent to the skin of the host and an electronics unit releasably attachable to the housing; andan applicator configured for inserting the needle and ... | DexCom, Inc. | A61B005/00
US08486070 | B2 | 2013-07-16 | Telemetric orthopaedic implant | 2006-08-23 | An instrumented orthopaedic implant, such as an intramedullary (IM) nail, is disclosed. The implant has the capacity to provide an accurate measurement of the applied mechanical load across the implant. The implant includes sensors and associated electronic components located in recesses on the outer surface of the implant. The implant houses the sensing apparatus, the interface circuitry, the data transmitter, and the power receiver. The hermetically sealed housing is adapted for implantation in the body of a patient. The implant is used with a controller which communicates with it by telemetry, and there is an acting unit connected to the electronic components which is adapted to carry out a function based upon a condition detected by the sensors. | 1. A telemetric orthopaedic implant system, the system comprising: a. an orthopaedic implant, the orthopaedic implant comprising: i. at least one sensor;ii. a first recess adapted to receive said at least one sensor;iii. an electronic component electrically connected to said at least one sensor, the electronic component including at least a power supply, a first transmitter, a first receiver, and a first microprocessor;iv. a second recess adapted to receive the electronic component;v. potting material to seal said first recess and said second recess;vi. a power source electrically connected to said electronic component; andvii. an acting unit electrically connected to said electronic component, said acting unit adapted to carry out a function based upon a condition; andb. a control unit, the control unit comprising: i. a second microprocessor;ii. a second transmitter electrically connected to said second microprocessor, the second transmitter adapted to send a signal to said first receiver of said electronic component; andiii. a second receiver electrically connected to said second microprocessor, the second receiver adapted to receive data from said first transmitter of said electronic component. 2. The telemetric orthopaedic implant system of claim 1, wherein said orthopaedic implant is an intramedullary nail. 3. The telemetric orthopaedic implant system of claim 1, wherein said at least one sensor is selected from the group consisting of a foil strain gauge, a semi-conductor strain gauge, a vibrating beam sensor, a force sensor, a piezoelectric element, a fibre Bragg grating, and a giant magneto-impedance (GMI) sensor. 4. The telemetric orthopaedic implant system of claim 1, wherein said control unit further comprises a graphical user interface. 5. The telemetric orthopaedic implant system of claim 1, wherein said orthopaedic implant further comprises a cap insert. 6. The telemetric orthopaedic implant system of claim 1, wherein said first transmitter and said first receiver are combined into a single first transceiver, and second transmitter and second receiver are combined into a single second transceiver. 7. The telemetric orthopaedic implant system of claim 1, further comprising a handheld device. 8. The telemetric orthopaedic implant system of claim 1, further comprising a reader. 9. The telemetric orthopaedic implant system of claim 1, further comprising a computing device. 10. The telemetric orthopaedic implant system of claim 1, further comprising a gait analysis tool. 11. The telemetric orthopaedic implant system of claim 1, wherein said function is delivery of a stored deliverable. 12. The telemetric orthopaedic implant system of claim 1, wherein said condition is based upon data acquired by said at least one sensor. 13. The telemetric orthopaedic implant system of claim 1, wherein said condition is based upon a command provided by said control unit. 14. The telemetric orthopaedic implant system of claim 1, wherein said ... | Smith & Nephew, Inc. | A61B017/56
US08184878 | B2 | 2012-05-22 | Image diagnosis support system and image diagnosis support method | 2008-04-22 | Specific information for specifying a specific study or series for which an object, which is obtained by adding at least one information of a scan condition, a scan range, and a key image position as a basis of a diagnosis in an image, is to be generated is specified. It is determined whether or not an image for references is present in the specific study or series on the basis of additional information of at least one image related to the specific study or series. When it is determined that the image for references is present, an object about the specific study or series is generated afterwards, for example, in the unit of series by using the image for references and at least one information of a scan condition, a scan range, and a key image position of the additional information. | 1. An image diagnosis support system comprising: a specifying unit that specifies specific information for specifying a specific study or series for object data to be generated, the object data being obtained by adding at least one information of a scan condition, a scan range, and a key image position as a basis of a diagnosis in a reference image;a first determination unit that determines whether or not a specific image for generating the reference image is present in the specific study or series on the basis of additional information of at least one image related to the specific study or series; anda first generation unit that generates the object data by using the specific image and at least one information of a scan condition, a scan range, and a key image position of the additional information when it is determined that the specific image is present. 2. The image diagnosis support system according to claim 1, wherein the additional information further includes at least one of an image position, an image direction, an image type, a study UID, a series UID, a study instance UID, and a series instance UID in addition to the scan condition, the scan range, and the key image position. 3. The image diagnosis support system according to claim 1, further comprising: a second generation unit that generates the specific image by using at least one information of the scan condition, the scan range, and the key image position of the additional information and at least one image related to the specific study or series when it is determined that the specific image is not present. 4. The image diagnosis support system according to claim 1, further comprising: a first acquisition unit that acquires information included in object data about a study or series similar to the specific study or series on the basis of the specific information,wherein the first generation unit generates object data about the specific study or series by using the acquired information. 5. The image diagnosis support system according to claim 4, wherein when information on other object data to be referred is included in the acquired information, the first acquisition unit acquires the information included in the other object data. 6. The image diagnosis support system according to claim 1, further comprising: a second acquisition unit that acquires a report related to the specific study or series; anda second determination unit that determines whether or not link information, which indicates a storage place of an image related to other studies or series different from the specific study or series, is included in the acquired report,wherein the first generation unit causes information, which specifies an image related to the other studies or series, extracted from the link information to be further included when it is determined that the link information is present. 7. The image diagnosis support system according to claim 1, wherein the reference image is an image for indicatin... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | G06K009/00
US08118757 | B2 | 2012-02-21 | Methods and devices for ostium measurement | 2007-04-30 | Devices and methods for accurately determining the size of an ostium of a patient and in particular sinus ostium. Methods for measuring a target ostium comprise inserting the distal end portion of an ostium measuring device into a patient, locating the target ostium with the measuring device, positioning the distal end portion of the measuring device appropriately adjacent to or into the target ostium and determining the diameter(s) of the target ostium by comparing the length, diameter and/or the circumference of the portion of the distal end portion of the measuring device adjacent to or within the target ostium. | 1. A method for measuring the size of and treating a paranasal sinus ostium of a patient, said method comprising the steps of: advancing a guidewire through a targeted paranasal sinus ostium in a patient;inserting a paranasal sinus ostium measuring device into the patient over the guidewire, wherein the measuring device comprises: an elongate, at least partially rigid shaft having a proximal end and a distal end;an enlarged distal end portion at the distal end of the shaft; andat least one preformed bend in the shaft;locating the targeted paranasal sinus ostium with the distal end portion of the paranasal sinus ostium measuring device;positioning the distal end portion of the paranasal sinus ostium measuring device at least partially within or through the targeted paranasal sinus ostium;determining a diameter of the targeted paranasal sinus ostium using the paranasal sinus ostium measuring device;withdrawing the paranasal sinus ostium measuring device from within or through the targeted paranasal sinus ostium;subsequent to withdrawing the paranasal sinus ostium measuring device, selecting a balloon catheter based upon the diameter of the targeted paranasal sinus ostium such that a balloon of the balloon catheter, when inflated within the ostium, fractures, moves or remodels bony structures that surround or are adjacent to the targeted paranasal sinus ostium; andintra-operatively treating the targeted paranasal sinus ostium with the selected balloon catheter. 2. The method of claim 1, wherein said distal end portion comprises a spherical portion, and wherein determining the diameter includes feeling tactile feedback of the spherical portion passing into the ostium to determine whether an outer diameter of the spherical portion is approximately equal to the diameter of the paranasal sinus ostium. 3. The method of claim 1, where said locating is performed by visual navigation techniques. 4. The method of claim 3 where said visual navigation technique is fluoroscopy. 5. The method of claim 1, wherein said positioning comprises placing the distal end portion of the paranasal sinus ostium measuring device against tissue surrounding the sinus ostium. 6. The method of claim 1, wherein said determining the sinus ostium size comprises visualizing the target sinus ostium and sinus ostium measuring device by fluorometry techniques. 7. The method of claim 1, wherein said determining the ostium size comprises visualizing the target sinus ostium and sinus ostium measuring device by image guided surgery (IGS) systems. 8. The method of claim 1, wherein said determining the sinus ostium size comprises visualizing the target sinus ostium and sinus ostium measuring device by endoscopic visualization. 9. The method of claim 1, wherein said determining the sinus ostium size comprises visualizing an indicator on the sinus ostium measuring device. 10. The method of claim 1, wherein said locating is performed by physical familiarity of a sinus ostium with the si... | Acclarent, Inc. | A61B005/103
US08655425 | B2 | 2014-02-18 | Optical sensor biosignal measurement apparatus and method of controlling optical sensor of the apparatus | 2008-01-25 | An optical sensor biosignal measurement apparatus including an optical sensor having a luminous element to emit light towards the skin of a user and a photo detector to detect light from the skin of the user; a light quantity adjustment member to adjust an optical transmission area of the photo detector; and a light quantity controller to detect the quantity of light detected by the photo detector, and control the light quantity adjustment member to adjust the optical transmission area of the photo detector according to the quantity of light is provided. | 1. An optical sensor biosignal measurement apparatus comprising: an optical sensor having a luminous element emitting light towards skin of a user and a photo detector detecting light from the skin of the user;a light quantity adjustment member adjusting an optical transmission area of the photo detector; anda light quantity controller detecting a quantity of the light detected by the photo detector, and controlling the light quantity adjustment member to adjust the optical transmission area of the photo detector,wherein the quantity of the light detected exceeding a threshold value is excluded when the quantity of the light detected exceeds the threshold value for less than a first predetermined period of time, and the optical transmission area is adjusted when the quantity of the light detected continuously exceeds the threshold value for a second predetermined period of time. 2. The optical sensor biosignal measurement apparatus of claim 1, wherein the first predetermined period of time and the second predetermined period of time are the same. 3. The optical sensor biosignal measurement apparatus of claim 1, wherein the light quantity adjustment member is a diaphragm, and the light quantity controller controls an open area and a closed area of the diaphragm to adjust the optical transmission area of the photo detector. 4. The optical sensor biosignal measurement apparatus of claim 1, wherein the light quantity adjustment member is at least one polymer-dispersed liquid crystal (PDLC) unit, and the light quantity controller controls a current supply with respect to the at least one PDLC unit to adjust the optical transmission area of the photo detector. 5. The optical sensor biosignal measurement apparatus of claim 4, wherein the at least one PDLC unit includes a plurality of PDLC units, and the current supply is selectively supplied to at least one of the plurality of PDLC units and is not supplied to at least one of the plurality of PDLC units. 6. The optical sensor biosignal measurement apparatus of claim 5, wherein when a first PDLC unit is not supplied current, the optical transmission area is a second area or a third area, excluding a first area. 7. The optical sensor biosignal measurement apparatus of claim 1, wherein the light quantity adjustment member is a liquid crystal display (LCD) unit, and the light quantity control unit controls an optical transmittance of the LCD unit to adjust the optical transmission area of the photo detector. 8. The optical sensor biosignal measurement apparatus of claim 1, wherein the light quantity controller detects a direct current (DC) level of the quantity of light detected by the photo detector, compares the DC level of the quantity of light with the threshold value for the second predetermined period of time, and controls the light quantity adjustment member to reduce the optical transmission area of the photo detector when the DC level of the quantity of light is greater than the threshold va... | Samsung Electronics Co., Ltd. | A61B005/00
US08121687 | B2 | 2012-02-21 | Cardiac motion characterization by strain measurement | 2005-10-31 | Methods for evaluating motion of a cardiac tissue location, e.g., heart wall, are provided. In the subject methods, timing of a signal obtain from a strain gauge stably associated with the tissue location of interest is employed to evaluate movement of the cardiac tissue location. Also provided are systems, devices and related compositions for practicing the subject methods. The subject methods and devices find use in a variety of different applications, including cardiac resynchronization therapy. | 1. A method for evaluating movement of a cardiac tissue location in a subject, the method comprising: (a) determining timing of a signal obtained from a resistive strain gauge of an implantable device stably associated with the cardiac tissue location, wherein the implantable device comprises:a lead of a length with a proximal end and a distal end, the lead comprising:a guide wire core formed within the lead along the length of the lead, wherein the resistive strain gauge is physically attached to the guide wire core; andat least one insulated conductor spirally wrapping the guide wire core from the proximal end to the distal end, wherein the resistive strain gauge is electrically coupled to an insulated conductor of the at least one insulated conductor; and(b) evaluating movement of the cardiac tissue location via a signal processing element communicatively coupled to the implantable device based on the timing of the signal obtained from the resistive strain gauge. 2. The method according to claim 1, wherein said strain gauge comprises multiple strain tranducers. 3. The method according to claim 1, wherein the resistive strain gauge comprises a substrate. 4. The method according to claim 3, wherein the substrate is compliant. 5. The method according to claim 1, wherein said strain gauge does not include a substrate. 6. The method according to claim 1, wherein said strain gauge comprises a transducer that is hermetically sealed. 7. The method according to claim 1, wherein the resistive strain gauge is not hermetically sealed. 8. The method according to claim 1, wherein the resistive strain gauge comprises a transducer in which strain induces an electric signal. 9. The method according to claim 3, wherein the resistive strain gauge is a fabricated using a resistive material. 10. The method according to claim 9, wherein the resistive material is platinum. 11. The method according to claim 3, wherein said strain transducer is fabricated from a piezoelectric material. 12. The method according to claim 1, wherein said strain gauge comprises a resonant marker band structure. 13. The method according to claim 1, wherein the lead further comprises a tissue securing element configured to stably associate the resistive strain gauge with the cardiac tissue location. 14. The method according to claim 1, wherein the resistive strain gauge is stably associated with the cardiac tissue location by a compressive force. 15. The method according to claim 1, wherein the implantable device is an acute device. 16. The method according to claim 1, wherein the implantable device is present on a permanently implantable device. 17. The method according to claim 1, wherein the cardiac tissue location is a septal wall location. 18. The method according to claim 1, wherein the cardiac tissue location is on a heart wall. 19. The method according to claim 18, wherein the heart wall is a chamber wall. 20. The method according to claim 19, wherein the chamber ... | Proteus Biomedical, Inc. | A61N001/18
US07922664 | B2 | 2011-04-12 | Method and system for improving physiologic status and health via assessment of the dynamic respiratory arterial pressure wave using plethysmographic technique | 2007-04-05 | The present invention specifies a method and system for assessing the dynamic respiratory arterial pressure wave using plethysmographic sensing techniques. The dynamic respiratory arterial pressure wave is measured and plotted for purposes of diagnosis and or remedial biofeedback. | 1. A method for assessing a dynamic respiratory arterial pressure wave, comprising: detecting the dynamic respiratory arterial pressure wave as a function of changing blood volume using a plethysmographic device;assessing at least one characteristic of the dynamic respiratory arterial pressure wave; andproviding at least one output for diagnostic purposes, wherein the at least one output represents the assessed at least one characteristic of the dynamic respiratory arterial pressure wave, whereby providing the at least one output further comprises defining at least one of a visual target and a visual threshold for the dynamic respiratory arterial pressure wave and providing the at least one of the visual target and the visual threshold to at least one of a clinician and a user. 2. The method of claim 1 wherein assessing the at least one characteristic of the dynamic respiratory arterial pressure wave further comprises measuring at least one parameter associated with the dynamic respiratory arterial pressure wave. 3. The method of claim 2 wherein measuring the at least one parameter associated with the dynamic respiratory arterial pressure wave further comprises measuring at least one of amplitude, frequency, phase, coherence, and centerline of the dynamic respiratory arterial pressure wave, wherein measuring the coherence further comprises measuring consistency of the amplitude, the frequency, and the phase of the dynamic respiratory arterial pressure wave. 4. The method of claim 3 wherein providing the at least one output further comprises displaying the at least one of the amplitude, the frequency, the phase, the coherence, and the centerline of the dynamic respiratory arterial pressure wave. 5. The method of claim 3 wherein measuring the amplitude further comprises measuring at least one of a peak amplitude and a valley amplitude of the dynamic respiratory arterial pressure wave. 6. The method of claim 1 wherein assessing the at least one characteristic of the dynamic respiratory arterial pressure wave further comprises calculating a centerline of the dynamic respiratory arterial pressure wave and filtering the centerline, and wherein providing the at least one output further comprises plotting the centerline of the dynamic respiratory arterial pressure wave. 7. The method of claim 6 wherein calculating the centerline of the dynamic respiratory arterial pressure wave further comprises calculating a root mean square (RMS) value of a pulse wave and adding a baseline blood volume value to the RMS value. 8. The method of claim 1 wherein assessing the at least one characteristic of the dynamic respiratory arterial pressure wave further comprises calculating a difference between a peak and a valley of a centerline of the dynamic respiratory arterial pressure wave to determine an amplitude of the dynamic respiratory arterial pressure wave. 9. The method of claim 1 further comprising detecting a peak blood volume and wherein defining the at le... | Coherence LLC | A61B005/02 A61B005/08
US08265745 | B2 | 2012-09-11 | Contact sensor and sheath exit sensor | 2006-12-29 | A system and method is provided that allows for determining the local impedance of one or more electrodes of an electrode catheter. Such local impedance may be utilized to identify the relative position of an electrode catheter to a sheath of a guiding introducer. In another arrangement, local impedance of a catheter electrode can be utilized to calibrate a catheter electrode to provide improved contact sensing. | 1. A method for performing a medical procedure, comprising the steps of: guiding a catheter to an internal patient location, wherein the catheter includes at least first and second electrodes;providing first and second separate electrical signals having equal frequency and phase to the first and second electrodes, respectively, wherein an amplitude of one of the first and second electrical signals is adjustable relative to an amplitude of the other of the first and second electrical signals;monitoring the first and second electrical signals as applied to the first and second electrodes, respectively, to identify first and second input responses;adjusting the amplitude of at least one of the first and second electrical signals as applied to a corresponding one of the electrodes to generate an initial relative response of the first and second electrodes;identifying a level of the initial relative response;identifying a change in the initial relative response;scaling the change based on the level of the initial relative response to generate a scaled change of response; andgenerating an output based on the scaled change of response. 2. The method of claim 1, wherein a return path for the electrical signals is via an external electrode. 3. The method of claim 1, wherein the step of generating an initial relative response further comprises: identifying a first impedance for the first electrode and identifying a second impedance for the second electrode. 4. The method of claim 1, further comprising: providing a plurality of separate electrical signals to a plurality of electrodes, respectively;identifying relative responses between one of the plurality of electrodes and at least two other electrodes, wherein the output is generated based on a change of at least one of the relative responses. 5. The method of claim 1, wherein the output indicates a contact condition with tissue. 6. The method of claim 1, further comprising: adjusting the amplitude of the one electrical signal to substantially match the first and second input responses of the first and second electrodes to the first and second electrical signals, respectively. 7. The method of claim 6, wherein the step of adjusting is performed while the first and second electrodes are in a fluid pool. 8. The method of claim 7, further comprising: moving the catheter to contact at least one of the electrodes with a wall of the internal patient locationwherein generating an output comprises generating a contact response of the electrode contacting the wall. 9. The method of claim 8, further comprising: in conjunction with generating the contact response, registering information with a location system. 10. The method of claim 8, further comprising: assessing contact between the catheter and the wall based on the contact response. 11. The method of claim 1, wherein the step of identifying the relative responses comprises: generating a differential of the first and second electrical signals. 12. Th... | St. Jude Medical, Atrial Fibillation Division, Inc. | A61B005/05
US08277387 | B2 | 2012-10-02 | System and method for determining the blood pressure of a patient | 2006-10-13 | In order to provide a technique for determining high-quality blood pressure values of a patient, especially in cases of unsupervised blood pressure measurements in a home environment, it is suggested to use a system (1) for determining the blood pressure of a patient (2) comprising a blood pressure measuring device (3, 4) for measuring a blood pressure value, an auxiliary device (6, 7, 8, 9) for measuring the motor activity of the patient (2) during a defined period of time prior to the blood pressure measurement, and a processing device (13, 14), said processing device (13, 14) being adapted to obtain the motor activity information and the blood pressure value, said processing device (13, 14) being further adapted to automatically assess the blood pressure value using the motor activity information, and said processing device (13, 14) being further adapted to provide a measuring result depending on the result of the assessment. | 1. A system for determining blood pressure of a patient, comprising: a blood pressure measuring device for measuring a blood pressure value;an auxiliary device for measuring motor activity of the patient during a defined period of time prior to the blood pressure measurement; anda processing device, said processing device being configured to obtain the motor activity information and the blood pressure value, said processing device being further configured to automatically make a determination whether the blood pressure value is reliable based on whether the motor activity information is at or below a predefined level stored in a look-up table, and said processing device being further configured to provide a measuring result based on the determination. 2. The system as claimed in claim 1, wherein the auxiliary device is configured to measure the motor activity of the patient during the blood pressure measurement. 3. The system as claimed in claim 1, wherein the processing device is configured to provide a measuring result in form of a corrected blood pressure value. 4. The system as claimed in claim 1, wherein the processing device is configured to provide a measuring result in form of a blood pressure value with a confidence interval. 5. The system as claimed in claim 1, wherein the processing device is configured to provide a measuring result in form of a signal to the patient requesting the patient to repeat the blood pressure measurement. 6. The system as claimed in claim 1, further comprising a number of additional auxiliary devices, said additional auxiliary devices being configured to provide additional measuring-related information which has been obtained before or during the blood pressure measurement, and the processing device being configured to obtain at least one additional piece of measuring-related information for making the determination of the reliability of the blood pressure value. 7. The system as claimed in claim 6, wherein said additional auxiliary device comprises a device for measuring ECG data of the patient. 8. The system as claimed in claim 6, wherein said additional auxiliary device comprises a device for measuring the galvanic skin response of the patient. 9. The system as claimed in claim 6, wherein said additional auxiliary device comprises a device for measuring the patient's silence. 10. The system as claimed in claim 6, wherein said additional auxiliary device comprises a device for measuring the patient's posture. 11. The system as claimed in claim 6, wherein said additional auxiliary device comprises a device for obtaining the sex or age of the patient. 12. The system as claimed in claim 6, wherein said additional auxiliary device comprises a device for obtaining the history of blood pressure measurements of the patient. 13. The system as claimed in claim 6, wherein said additional auxiliary device comprises a device for measuring the ambient temperature during the blood pressure measurement. 14. Th... | Koninklijke Philips Electronics N.V. | A61B005/02
US08144948 | B2 | 2012-03-27 | Method and apparatus for transporting images | 2007-03-30 | A portable image transport device and corresponding techniques for image transport are disclosed. The portable image transport device accommodates an automatic initiation of image transport when the device is in a designated location by automatically recognizing the location condition and then engaging in communication with a medical device having an image capture capability. A variety of image transport may be performed, such as sending the received image to a printer for printing a hard copy. The portable image transport device also manages restrictions on image transport. Various restrictions may apply including but not limited to the type of devices to which the transport may be made, or certain registered devices to ensure transport only within a defined set of recipient classes or particularly identified recipients. | 1. A method for transporting medical images through a portable medical image device that receives and transports medical images, the method comprising: receiving an indication that the portable medical image device is present at a designated location;communicating with a medical device having an image capture capability upon receiving the indication that the portable medical image device is present at the designated location;verifying that the medical device is registered for transporting images through the portable medical image device;receiving an image from the medical device; andtransporting the received image to an other device,wherein the medical device at a given time is associated with a patient, and wherein image transport parameters incorporate identification of the patient in association with recipient classification information to restrict distribution of the images accordingly. 2. The method of claim 1, wherein transporting the received image is to a printer through a wireless connection, and a confirmation of successful printing is returned to the portable medical image device using the wireless connection while it is at the designated location. 3. The method of claim 1, further comprising: correlating the medical device to information comprising image transport parameters that define restrictions in transporting images using the registered medical device. 4. The method of claim 1, wherein transporting the received image to the other device is performed using a wireless connection. 5. The method of claim 4, wherein the wireless connection is a wireless LAN connection. 6. The method of claim 1, wherein an audio annotation is associated to the received image prior to transporting the received image. 7. An apparatus for transporting medical images through a portable medical image device that receives and transports medical images, the apparatus comprising: means for receiving an indication that the portable medical image device is present at a designated location;means for communicating with a medical device having an image capture capability upon receiving the indication that the portable medical image device is present at the designated location;means for verifying that the medical device is registered for transporting images through the portable medical image device;means for receiving an image from the medical device; andmeans for transporting the received image to an other device, wherein the medical device at a given time is associated with a patient, and wherein image transport parameters incorporate identification of the patient in association with recipient classification information to restrict distribution of the images accordingly. 8. The apparatus of claim 7, wherein transporting the received image is to a printer through a wireless connection, and a confirmation of successful printing is returned to the portable medical image device using the wireless connection while it is at the designated location. 9. The appar... | Sony Corporation Sony Electronics Inc. | G06K009/00
US08065017 | B2 | 2011-11-22 | Method and apparatus for obtaining and registering an Electrical Cochlear Response (“ECR”) | 2008-02-26 | This invention relates to a method and apparatus for obtaining and registering the Electrical Cochlear Response (“ECR”). Particularly, the ECR refers to the electrical activity of the auditory nerve and the intracochlear residual tissue located in the neighborhood of an intracochlear stimulation electrode, when an external sound is presented to a Cochlear Implant user, human or animal. Applicants recognize that an ECR can be thought of as a far-field electrical potential registered on the scalp of an implanted patient. Applicants, through their method and apparatus, can extract that potential from the patient's spontaneous electroencephalographic activity (EEG) and measure it by using an averaging algorithm. The Electrical Cochlear Response can be applied for adaptation, calibration, performance evaluation and failure detection of the Cochlear Implant of the implanted patient. Also Applicants' method can be applied to estimate the audiometric thresholds of the Cochlear Implant even without the implanted patient's knowledge. | 1. An apparatus for obtaining an electrical cochlear response (“ECR”) of a patient using a cochlear stimulation system, the apparatus comprising: a sound generating system for generating an external sound stimulus to the cochlear stimulation system;an electro encephalographic (EEG) acquisition system for registering EEG electrical activity generated when the external sound is processed by the cochlear stimulation system; anda computer-controlled system having hardware and software configured to: receive the EEG activity as EEG epochs corresponding to a repeated sound stimulation, the EEG epochs having a time sufficient for at least one waveform; andaveraging software configured to average a plurality of EEG epochs to generate the ECR where the ECR is characterized by an ECR waveform having parameters that measure the patient's objective perception of the external sound. 2. The apparatus of claim 1 where the sound generating system includes: a signal generator for generating stimulation signals;an amplifier for amplifying the stimulation signals; andan electroacoustic transducer coupled to the amplifier for providing the patient with the external sound stimulus. 3. The apparatus of claim 1 where the EEG acquisition system includes a register electrode system placed on the scalp of the implanted patient for picking up the implanted patient's EEG activity. 4. The apparatus of claim 1 further comprising an audiometric enclosure in which the implanted patient is housed. 5. The apparatus of claim 1 further comprising a reduced dimension test sound chamber in which the electroacoustic transducer is mechanically coupled to a microphone of the cochlear stimulation system. 6. A method for obtaining an electrical cochlear response (“ECR”) of an implanted patient with a cochlear stimulation system comprising: generating an external sound stimulus at a predetermined dynamic range;registering EEG activity arising from the external sound stimuli, the EEG activity detected from electrical signals registered on the scalp of the patient; andprocessing the EEG activity as EEG epochs corresponding to a repeated sound stimulation; andaveraging a plurality of EEG epochs to generate the ECR where the ECR is characterized by an ECR waveform having parameters that measure the patient's objective perception of the external sound. 7. The method of claim 6 further comprising: before the step of registering EEG activity, mounting a register electrode system on the scalp of the patient. 8. The method of claim 6 further comprising fitting the cochlear stimulation system by a method comprising: controlling the cochlear stimulation system to operate according to an initial set of operational parameter values;after the step of processing the EEG activity to obtain the ECR, adjusting the operational parameters of the cochlear stimulation system according the ECR parameters obtained from the register electrode system;setting the dynamic range of the electrical stimulation cu... | Universidad Autonoma Metropolitana Unidad Iztapalapa | A61B005/0484
US08327841 | B2 | 2012-12-11 | Method and device for determining the location of snoring | 2008-04-04 | The treatment of snoring related sleep disorders require the knowledge of the location of origin of snoring in a patient. A method and device are provided for the determination of a location of origin of a primary vibration signal generated by snoring in an upper airway of a patient. At least two sensors are used to respectively detect the primary vibration signal and to generate respective intermediate signals. The sensors are spaced apart in a longitudinal direction of the patient's neck. The respective intermediate signals are processed to generate an output signal. The output signal is indicative of the location of origin. | 1. A method for determining a location of origin of a primary vibration signal generated by snoring in an upper airway of a human patient, the method comprising: placing at least two vibratory sensors spaced apart along a longitudinal direction of the human patient's neck to detect the primary vibration signal;generating respective intermediate signals from the at least two vibratory sensors, the respective intermediate signals being within a pre-determined frequency range, the pre-determined frequency range being 20 to 1000 Hz;processing the respective intermediate signals by comparing times at which the primary vibration signal is detected by the at least two vibratory sensors thereby generating an output signal; anddetermining the location of origin of said primary vibration signal generated by snoring in said upper airway of said human patient based on the output signal. 2. The method of claim 1, wherein said processing the respective intermediate signals comprises: selecting at least one characteristic of the respective intermediate signals from a plurality of characteristics of the respective intermediate signals that vary with the location of origin. 3. The method of claim 1, wherein the output signal is a time differential signal based on the times at which the primary vibration signal is detected by the at least two vibratory sensors. 4. The method of claim 1, wherein said determining the location of origin of said primary vibration signal generated by snoring in said upper airway of said human patient comprises determining an angle of incidence of the primary vibration signals. 5. The method of claim 1, wherein said processing the respective intermediate signals comprises conditioning the respective intermediate signals to reduce electrical noise. 6. The method of claim 1, wherein said processing the respective intermediate signals comprises: determining a peak occurrence in each of the respective intermediate signals; andcomparing the peak occurrence in each of the respective intermediate signals. 7. The method of claim 1 further comprising: correlating the primary vibration signal to an independent secondary vibration signal indicative of snoring by the human patient. 8. The method of claim 7, wherein the independent secondary vibration signal is a filtered pressure wave generated by a nasal pressure cannula or a vibration signal generated by a snoring microphone. 9. The method of claim 1, wherein the at least two vibratory sensors are placed on the back of the patient's neck. 10. The method of claim 1, wherein the at least two vibratory sensors are placed on a side of the patient's neck. | BRAEBON Medical Corporation | A62B007/00 A62B009/00 A62B018/00 A61F005/56 A61B005/08
US08075601 | B2 | 2011-12-13 | Deformity correction using neural integrity monitoring | 2007-04-30 | A method is provided for detecting neural interference between a bone anchor and a neural element during or subsequent to the application of force on the bone anchor to adjust the orientation of one or more vertebral bodies relative to the spinal column. The method includes engaging at least one bone anchor to the vertebral body and applying force to the bone anchor to adjust the positioning or orientation of the vertebral body. A nerve monitoring system provides an electrical signal and is operable to detect a neural element and its proximity to the bone anchor as a function of a characteristic of the electrical signal. | 1. A method comprising: providing a bone anchor;engaging the bone anchor to a location on a vertebral body when the vertebral body is in a first orientation;coupling the bone anchor with a correction tool;manipulating the correction tool to move the vertebral body toward a second orientation with the bone anchor, wherein manipulating of the correction tool comprises applying a rotational force to the correction tool which rotates the correction tool and correspondingly rotates the vertebral body toward the second orientation;electrically monitoring the bone anchor with a nerve monitoring system to detect interference between a neural element and the bone anchor while the vertebral body is moved toward the second orientation;wherein the nerve monitoring system is operable to provide an electrical signal to the bone anchor and detect interference between the neural element and the bone anchor by measuring neural element threshold as a function of the electrical signal; andwherein the correction tool is engaged to the bone anchor and the electrical signal travels through the correction tool to the bone anchor. 2. The method of claim 1, further comprising returning the vertebral body from the second orientation toward the first orientation upon detection of interference between the neural element and the bone anchor. 3. The method of claim 2, further comprising engaging the bone anchor in the vertebral body with the bone anchor in at least one of a differing location, differing orientation, and differing trajectory in the vertebral body after returning the vertebral body toward the first orientation. 4. The method of claim 3, wherein after engaging the bone anchor in at least one a differing location, differing orientation, and differing trajectory, further comprising manipulating the correction tool to move the vertebral body toward the second orientation through the bone anchor while monitoring the bone anchor with the nerve monitoring system to detect interference between the neural element and the bone anchor. 5. The method of claim 1, further comprising preparing a hole in bony tissue of the vertebral body to receive the bone anchor. 6. The method of claim 5, wherein the bony tissue is located within a spinal pedicle of the vertebral body. 7. The method of claim 6, wherein the bone anchor includes a longitudinally threaded stem structured to threadingly engage bony tissue around the hole. 8. The method of claim 1, further comprising retaining the vertebral body in the second orientation with a spinal fixation system and disengaging the correction tool from the bone anchor. 9. The method of claim 8, wherein fixing the vertebral body in the second orientation includes interconnecting the bone anchor with one or more implants engaged to one or more vertebral bodies to form a rigid construct between one or more vertebral levels along a spinal column. 10. The method of claim 1, wherein the bone anchor is electrically coupled to the nerve mon... | Warsaw Orthopedic, Inc. | A61B017/88 A61B017/70 A61B005/05
US07993398 | B2 | 2011-08-09 | Angle indicator for capsular bag size measurement | 2007-04-24 | Implantation of an accommodating intraocular lens in an eye may require the accurate measurement of the size, circumference, or diameter of the capsular bag of the eye. After the natural crystalline lens has been surgically removed, a flexible ring may be temporarily or permanently inserted into the capsular bag for measuring the circumference of the capsular bag. The ring is generally compressible to fit through a surgical incision, then expands to fill the capsular bag along an equatorial region. The ring has a central component that changes shape as the ring is compressed. The shape change is generally correlated to the size of the capsular bag and may be measured visually or with a camera through the cornea, the measurement being generally independent of corneal magnification. | 1. A method for measuring the size of a capsular bag of an eye, comprising: inserting an angle indicator into the capsular bag, the angle indicator comprising:an incomplete annulus and a pair of joined arms disposed within the incomplete annulus, the joined arms operably coupled to first and second ends of the angle indicator;reading an angle from the angle indicator;removing the angle indicator from the capsular bag; andconverting the angle to a capsular bag size. 2. The method of claim 1, further comprising positioning the angle indicator such that the joined arms are at least partially visible through the pupil. 3. The method of claim 1, wherein reading comprises: forming a physical angle between two segments of the angle indicator;superimposing a protractor onto the physical angle;comparing the physical angle to at least one calibration feature on the protractor; andidentifying the angle from the at least one calibration feature. 4. The method of claim 1, wherein converting comprises: identifying a predetermined relationship between the angle and the capsular bag size described by at least one of a formula, a table, and a graph; andusing the predetermined relationship to estimate the capsular bag size. 5. The method of claim 1, further comprising reading the angle from a central portion of the angle indicator. 6. The method of claim 1, further comprising expanding the angle indicator to coincide with an equator of the capsular bag. 7. An angle indicator for measuring the size of a capsular bag of an eye, comprising: a compressible, incomplete annulus having a first end, a second end, an interior and an outer surface that remains essentially circular when the incomplete annulus is compressed, the outer surface having a diameter that varies in response to compression of the incomplete annulus;a first segment attached to the first end of the incomplete annulus and extending into the interior of the incomplete annulus; anda second segment attached to both the second end of the incomplete annulus and a distal end of the first segment;wherein the first and second segments form an angle, the angle being indicative of the diameter of the outer surface of the incomplete annulus. 8. The angle indicator of claim 7, wherein the first and second segments are configured to form an angle visible through the pupil of an eye when the angle indicator is disposed in the capsular bag of the eye. 9. The angle indicator of claim 7, wherein the first segment is hingedly attached to the first end of the incomplete annulus; andwherein the second segment is hingedly attached to both the second end of the incomplete annulus and the first segment. 10. The angle indicator of claim 7, wherein the first segment comprises a first straight portion, and wherein the second segment comprises a second straight portion. 11. The angle indicator of claim 7, wherein the incomplete annulus has a varying radial thickness along its length, the radial thickness being greater a... | Abbott Medical Optics Inc. | A61F002/14 A61B005/103 A61B005/117
US08082029 | B2 | 2011-12-20 | Adjustable sensors for use with exercise apparatus | 2007-05-17 | Adjustable sensors for use with exercise apparatus are described. An example apparatus includes a support frame to assist a user in an exercise and a sensor integral with the support frame. The sensor is to detect a physiological condition of the user through physical contact with the user, and a position of the sensor is adjustable by the user. | 1. An exercise apparatus comprising: a support frame including a rotatable adjustment mechanism; anda sensor assembly comprising an aperture configured to receive a segment of the support frame, a sensor configured to detect a physiological condition of a user through direct physical contact with the user, and an element to engage the adjustment mechanism such that a position of the sensor relative to the segment of the support frame received by the sensor assembly is rotatably adjustable by the user. 2. The apparatus as defined in claim 1, further comprising a processor system. 3. The apparatus as defined in claim 2, wherein the processor system is to cause the display of information associated with the physiological condition. 4. The apparatus as defined in claim 1, wherein the position of the sensor is adjustable to a plurality of fixed positions. 5. The apparatus as defined in claim 4, further comprising a first indicator on the sensor assembly to be visible when the sensor is in one of the fixed positions and a second indicator to block visibility of the first indicator when the sensor is not in one of the fixed positions. 6. The apparatus as defined in claim 5, wherein the first indicator is a first color and the second indicator is a second color. 7. The apparatus as defined in claim 1, the wherein the element comprises a protrusion to engage the adjustment mechanism of the support frame to fix the position of the sensor in one of a plurality of predetermined positions. 8. The apparatus as defined in claim 1, wherein the position of the sensor is adjustable by rotating the sensor assembly about the segment of the support frame received by the aperture of the sensor assembly. 9. The apparatus as defined in claim 1, wherein the sensor comprises an electrode and the physiological condition comprises a heart rate. 10. The apparatus as defined in claim 1, wherein the adjustment mechanism is positioned at a junction of a first portion of the support frame having a first outer diameter and a second portion of the support frame having a second outer diameter, the first outer diameter being greater than the second outer diameter. 11. An exercise apparatus comprising: a frame including a first portion to be grasped by a user and a second portion having an adjustment mechanism to receive a sensor assembly, the sensor assembly comprising an aperture configured to receive the second portion of the frame, the sensor assembly further comprising a sensor configured to detect a physiological condition of the user, wherein a position of the sensor relative to the second portion of the frame is rotatable around the frame via the adjustment mechanism of the frame. 12. The exercise apparatus as defined in claim 11, further comprising a processor system. 13. The exercise apparatus as defined in claim 12, wherein the processor system is to cause the display of information associated with the physiological condition. 14. The exercise apparatus as def... | Brunswick Corporation | A61B005/0402
US08260401 | B2 | 2012-09-04 | Non-invasive in-vivo imaging of mechanoreceptors in skin using confocal microscopy | 2007-07-25 | Meissner's Corpuscles (MCs) are touch-pressure sensation receptors in glabrous skin. They are imaged by reflectance confocal microscopy to provide a non-invasive, in vivo quantification of their density or size to allow screening for, diagnosis or monitoring of sensory neuropathy and other peripheral nervous system disorders related to diabetes, HIV, or other conditions. | 1. A method for non-invasive, in vivo quantification of Meissner's Corpuscles in a region of interest on a patient's skin, the method comprising: (a) using confocal microscopy to image the region of interest to produce image data;(b) analyzing the image data to detect the Meissner's Corpuscles; and(c) quantifying the Meissner's Corpuscles in accordance with a result of step (b). 2. The method of claim 1, wherein step (c) comprises quantifying a density of the Meissner's corpuscles. 3. The method of claim 1, wherein step (c) comprises quantifying a size of the Meissner's corpuscles. 4. The method of claim 1, wherein the Meissner's corpuscles are monitored over time. 5. A system for non-invasive, in vivo quantification of Meissner's Corpuscles in a region of interest on a patient's skin, the system comprising: a confocal microscope for imaging the region of interest to produce image data; anda computer, configured to receive the image data, analyze the image data to detect the Meissner's Corpuscles and quantify the Meissner's Corpuscles. 6. The system of claim 5, wherein the computer quantifies a density of the Meissner's corpuscles. 7. The system of claim 5, wherein the computer quantifies a size of the Meissner's corpuscles. 8. The system of claim 5, wherein the computer monitors the Meissner's corpuscles over time. 9. A method for non-invasive, in vivo quantification of Meissner's Corpuscles in a region of interest on a patient's skin in order to diagnose or monitor a condition in the patient, the method comprising: (a) using confocal microscopy to image the region of interest to produce image data;(b) analyzing the image data to detect the Meissner's Corpuscles;(c) quantifying the Meissner's Corpuscles in accordance with a result of step (b); and(d) diagnosing or monitoring the condition in accordance with a result of step (c). 10. The method of claim 9, wherein step (c) comprises quantifying a density of the Meissner's corpuscles. 11. The method of claim 9, wherein step (c) comprises quantifying a size of the Meissner's corpuscles. 12. The method of claim 9, wherein the Meissner's corpuscles are monitored over time. 13. The method of claim 9, wherein step (d) is performed for a case in which the condition is a peripheral nervous system disorder. 14. The method of claim 13, wherein the peripheral nervous system disorder is sensory neuropathy. 15. The method of claim 14, wherein the sensory neuropathy is caused by diabetes. 16. The method of claim 14, wherein the sensory neuropathy is caused by HIV. 17. A system for non-invasive, in vivo quantification of Meissner's Corpuscles in a region of interest on a patient's skin in order to diagnose or monitor a condition in the patient, the system comprising: a confocal microscope for imaging the region of interest to produce image data; anda computer, configured to receive the image data, analyze the image data to detect the Meissner's Corpuscles, quantify the Meissner's Corpuscles and ... | University of Rochester Lucid, Inc. | A61B006/00
US08313432 | B2 | 2012-11-20 | Surgical data monitoring and display system | 2008-06-20 | A surgical data monitoring and display system is described. In some embodiments, the system includes a data storage module that stores retrospective data and real-time surgical data concerning a patient, a first processing module that receives and processes the retrospective data into processed retrospective data, and a second processing module that receives and processes the real-time data into processed real-time data. Each of the first processing module and the second processing module transmits their processed data to first and second display modules, respectively, before or during performance by a healthcare provider of a medical or surgical procedure on the patient. | 1. A surgical data monitoring and display system for use in an operating room or another patient care room, comprising: a data storage module that stores retrospective data and real-time data concerning a patient;a first processing module that receives the retrospective data, processes the retrospective data into processed retrospective data, and transmits the processed retrospective data to a first display module before or during performance of a medical or surgical procedure on the patient by a healthcare provider;the first display module, which displays the processed retrospective data;a first gateway that receives the retrospective data from a server and transmits the retrospective data to the data storage module;a second processing module that receives the real-time data, processes the real-time data into processed real-time data, and transmits the processed real-time data to a second display module before or during performance of the medical or surgical procedure on the patient by the healthcare provider; andthe second display module, which displays the processed real-time data;wherein the real-time data concerning the patient comprises at least two of the following: electrocardiographic data, electroencephalographic data, pulse oximetry data, videoscopic data, and vital sign data;wherein the retrospective data comprises at least one of imaging data, patient identification information, past medical history information, physical examination information, data concerning a past procedure performed on the patient, or data concerning another patient or a teaching case;wherein the real-time data is acquired during performance of the medical or surgical procedure on the patient;wherein the first display module and the second display module are positionable in the operating room or another patient care room such that they are viewable by the healthcare provider during the performance by the healthcare provider of the medical or surgical procedure on the patient;a third processing module configured to receive postoperative data, the retrospective data, and the real-time data, and further configured to process at least one of the postoperative data, the retrospective data, or the real-time data into report data; anda third display module, which displays the report data; andan alert module, coupled to the second processing module and configured to provide, on a single display, a visible alert when a predetermined threshold of the real-time data is exceeded; wherein the visible alert comprises at least one of: moving displayed real-time data from a first location, less prominent on a display of the second display module to the healthcare provider, to a second location, more prominent on the display of the second display module to the healthcare provider while maintaining display of the processed retrospective data on the first display module to be visible to the health care provider; andenlarging a display of a first set of the processed real-time data... | SurgMatix, Inc. | A61B005/00
US08449473 | B2 | 2013-05-28 | Gas sensor | 2008-04-10 | A method of determining the respiratory function of a subject is disclosed, the method comprising measuring the concentration of water vapor in the gas stream exhaled by the subject, and from the measured water vapor concentration determining the respiratory function of the subject. Mathematical techniques for analyzing the data obtained from the measurement of concentrations of gaseous components in an exhaled gas stream are also disclosed and their use in determining the lung function of a subject. | 1. A method of determining the respiratory function of a subject, the method comprising: measuring the concentration of water vapour in the gas stream exhaled by the subject using an electrochemical sensor, wherein said measuring comprises causing the gas stream to impinge on a sensing element of said sensor, the sensing element comprising a working electrode, a counter electrode and a layer of ion exchange material extending between the working electrode and the counter electrode, the ion-exchange layer being configured to form an electrical contact between the working electrode and the counter electrode when exposed to water vapour in said gas stream, and applying a pulsed electric potential across the working electrode and counter electrode whereby the sensor generates an output representative of the change of concentration of water vapour over time in said gas stream;employing a processor to use the output of the sensor to generate data providing a profile representative of said change of water vapour concentration over time in the exhaled gas stream; andfrom said profile directly determining the respiratory function of the subject. 2. The method according to claim 1, wherein the concentration of water vapour is measured for the duration of at least one complete exhalation. 3. The method according to claim 2, wherein the concentration of water vapour is measured during tidal breathing by the subject. 4. The method according to claim 1, wherein the exhaled gas stream is exhaled from the mouth of the subject. 5. The method of claim 1, wherein said pulsed potential alternates between a rest potential and a potential above the reaction threshold potential. 6. The method of claim 1, wherein the voltage is pulsed at a frequency of from 1 to 500 Hz. 7. The method of claim 1, wherein the slope of the ascending phase of the profile is measured and said slope is used to determine the respiratory function of the subject. 8. The method according to claim 7, further comprising measuring the slope of the plateau phase of the profile. 9. The method according to claim 8, wherein the ratio of the slope of the ascending phase and the plateau phase is determined. 10. The method according to claim 1 which further comprises analyzing the portion of the profile in the region of the transition from the ascending phase to the plateau phase and using the results of the analysis to make the determination of the respiratory function of the subject. 11. The method of claim 1 wherein said processor generates a display of said profile. 12. The method according to claim 1 wherein storage means are provided for storing a plurality of pre-existing profiles representative of change of concentration of water vapour over time in an exhaled gas stream and means are provided for comparing the profile thus obtained with one or more pre-existing profiles retrieved from the storage means. 13. The method according to claim 12 wherein said profile is compared with one or m... | Anaxsys Technology Limited | A61B005/08
US08744564 | B2 | 2014-06-03 | Oedema detection | 2005-06-17 | A method of detecting tissue oedema in a subject. The method includes determining a measured impedance for first and second body segments. An index indicative of a ratio of the extra-cellular to intra-cellular fluid is then calculated for each body segment, with these being used to determine an index ratio based on the index for the first and second body segments. The index ratio can in turn be used to determine the presence, absence or degree of tissue oedema, for example by comparing the index ratio to a reference or previously determined index ratios. | 1. A method of detecting tissue oedema in a subject, the method comprising: a) determining a measured impedance for a first body segment;b) determining a measured impedance for a second body segment, the second body segment being different from the first body segment;c) for each body segment, and using the respective measured impedance, determining, using a processor of a processing system, an index indicative of a ratio of the extra-cellular to intra-cellular fluid;d) determining, using the processor of the processing system, an index ratio based on the index for the first and second body segments; ande) determining the presence, absence or degree of tissue oedema based on the index ratio. 2. A method according to claim 1, wherein the method includes, in the processing system: a) comparing the index ratio to at least one reference; andb) determining the presence, absence or degree of tissue oedema based on the results of the comparison. 3. A method according to claim 2, wherein the reference includes at least one of: a) a predetermined threshold;b) a tolerance determined from a normal population; andc) a predetermined range. 4. A method according to claim 2, wherein the reference includes an index ratio previously determined for the subject. 5. A method according to claim 4, wherein the previously determined index ratio is determined prior to the subject undergoing at least one of: a) surgery; andb) treatment. 6. A method according to claim 1, wherein the first and second body segments are different types of body segments. 7. A method according to claim 1, wherein the first and second body segments are limbs. 8. A method according to claim 1, wherein the first body segment is a leg and the second body segment is an arm. 9. A method according to claim 1, wherein the method includes, in the processing system: a) determining a plurality of measured impedances for each body segment, each measured impedance being measured at a corresponding measurement frequency; and,b) determining the index ratio based on the plurality of measured impedances. 10. A method according to claim 9, wherein the method includes, in the processing system: a) determining values for parameters R0 and R∞ from the measured impedance values; and,b) calculating the index (I) using the equation:I=R∞R0-R∞where: R0 is the resistance at zero frequency;and, R∞ is the resistance at infinite frequency. 11. A method according to claim 10, wherein the method includes, in the processing system, determining the parameter values using the equation:Z=R∞+R0-R∞1+(j⁢⁢ω⁢⁢τ)(1-α)where: Z is the measured impedance at angular frequency ω,τ is a time constant,and α has a value between 0 and 1. 12. A method according to claim 11, wherein the method includes, in the processing system: a) determining the impedance of each body segment at four discrete frequencies; and,b) determining values for the parameters by solving the equation using four simultaneous equations. 13. A method according to... | Impedimed Limited | A61B005/00
US07935499 | B2 | 2011-05-03 | Method and device for predicting physiological values | 2008-06-30 | The invention relates generally to methods, systems, and devices for measuring the concentration of target analytes present in a biological system using a series of measurements obtained from a monitoring system and a Mixtures of Experts (MOE) algorithm. In one embodiment, the present invention describes a method for measuring blood glucose in a subject. | 1. A method to measure an amount of concentration or analyte present in a biological system, comprising: determining a calibration ratio (CalRatio) value, whereinCalRatio=BGcp(activecp+offset)where BGcp is a blood glucose concentration at the calibration point, activecp is an active signal that corresponds to an electrochemical sensor signal at the calibration point, and offset is a value that takes into account a non-zero y-intercept value;providing two or more ranges of CalRatio values;identifying the range in which said determined CalRatio value falls; employing an algorithm for prediction of further measurement values selected from one of a first algorithm and a second: algorithm, the first algorithm comprising: BG=w1BG1+w2BG2+w3BG3¦whereBG1=p1(time)+q1(active)+r1(signal)+s1(BG¦cp)+t1 BG2=p2(time)+q2(active)+r2(signal)+s2(BG¦cp)+t2 BG3=p3(time)+q3(active)+r3(signal)+s3(BG¦cp)+t3w1=ⅇd1ⅇd1+⁢ⅇd2+ⅇd3w2=ⅇd2ⅇd1+⁢ⅇd2+ⅇd3w3=ⅇd3ⅇd1+⁢ⅇd2+ⅇd3d1=τ1(time)+β1(active)+γ1(signal)+δ1(BG¦cp)+∈1 d2=τ2(time)+β2(active)+γ2(signal)+δ2(BG¦cp)+∈2 d3=τ3(time)+β3(active)+γ3(signal)+δ3(BG¦cp)+∈3 in which BGi is the analyte predicted, BG/cp is the blood glucose value at a calibration point, time is the elapsed time, active is the active signal, signal being the calibrated signal, pi, qi, ri are coefficients, ti is a constant, e indicates an exponential function, di is a parameter set usable to determine weightings wi, with τi, βi, γi, δi, and ∈i are constants, and the second equation comprising: BG=w1BG1+w2BG2+w3BG3 whereBG1=⁢p1⁡(timec)+q1⁡(active)+r1⁡(signal)+s1⁡(BG❘cp)+t1BG2=⁢p2⁡(timec)+q2⁡(active)+r2⁡(signal)+s2⁡(BG❘cp)+t2BG3=⁢p3⁡(timec)+q3⁡(active)+r3⁡(signal)+s3⁡(BG❘cp)+t3w1=ⅇd1ⅇd1+⁢ⅇd2+ⅇd3w2=ⅇd2ⅇd1+⁢ⅇd2+ⅇd3w3=ⅇd3ⅇd1+⁢ⅇd2+ⅇd3d1=τ1(timec)+β1(active)+γ1(signal)+δ1(BG¦cp)+∈1 d2=τ2(timec)+β2(active)+γ2(signal)+δ2(BG¦cp)+∈2 d3=τ3(timec)+β3(active)+γ3(signal)+δ3(BG¦cp)+∈3 in which BGi is the analyte predicted, timec is the elapsed time since calibration, active is the active signal, signal is the calibrated signal, BG/cp is the blood glucose value at a calibration point, pi, qi, ri are coefficients, ti is a constant, e indicates an exponential function, di is a parameter set usable to determine weightings wi, with τi, βi, γi, δi, and ∈i are constants, wherein said algorithm is optimized for performance in the identified range; and generating further measurement values indicative of amount or concentration of analyte present in the biological system, said generating comprising obtaining a raw signal specifically related to analyte amount or concentration in the biological system and using said algorithm to correlate the raw signal with a measurement value. | Animas Corporation | G01N033/66
US07942783 | B2 | 2011-05-17 | Exercise aid device | 2005-07-21 | An exercise aid device has a hip supporting member movable relative to a base, footrests movable relative to the base, drive means for driving the hip supporting member, body constitution estimating unit for estimating at least one of fat mass and muscle mass of a user, and a controller for the drive means. The controller controls the drive means such that a load acting on a femoral region by own weight of the user supported on the hip supporting member changes according to a relative positional displacement between the user's toe and trochanter major, the positional displacement is allowed in a direction of flexion and extension of knee joint of the user, and an angle of the knee joint is maintained substantially constant. In addition, since the controller controls the drive means by use of an output of the body constitution estimating unit, it is possible to provide an exercise with less burden to the knee joint and a suitable strength for the user. | 1. An exercise aid device comprising: a base;a supporting member, movable relative to said base in a forward right direction and a forward left direction, which is configured to support a hip of a user;a footrest movable relative to said base at least in the up and down directions;a drive means configured to drive said supporting member and at least one of said footrest anda controller configured to control said drive means such that a load acting on a femoral region by own weight of the user supported on said supporting member changes according to a relative positional displacement between the user's toe and trochanter major, said positional displacement is allowed in a direction of flexion and extension of knee joint of the user, and an angle of the knee joint is maintained substantially constant,wherein the supporting member comprises a saddle for supporting the user's hip,wherein said drive means drives said supporting member so that:the saddle is oscillated between an upright posture relative to the base and an inclined posture in the forward right direction or the forward left direction. 2. The exercise aid device as set forth in claim 1, further comprising a body constitution estimating unit configured to estimate at least one of fat mass and muscle mass of the user, and wherein said controller controls said drive means by use of an output of said body estimation unit. 3. The exercise aid device as set forth in claim 2, further comprising a grip which is held by the user, a pair of first electrodes disposed on said footrest, a pair of second electrodes provided on said grip, and an impedance measuring unit configured to measure a bioelectrical impedance of the user by detecting a potential difference between one of said first electrodes and one of said second electrodes, while applying a high frequency current between the other first electrode and the other second electrode, under the condition that the user's foot is placed on said footrest and said grip is held by the user, and wherein said body constitution estimating unit estimates at least one of fat mass and muscle mass of the user by use of an output of said impedance measuring unit. 4. The exercise aid device as set forth in claim 3, further comprising a body weight input unit configured to input the user's body weight, and said body constitution estimating unit estimates at least one of fat mass and muscle mass of the user by use of the output of said impedance measuring unit and the user's body weight input by said body weight input unit. 5. The exercise aid device as set forth in claim 3, further comprising a body information input unit configured to input body weight and body height of the user, and an energy expenditure operation unit configured to calculate one of an energy expenditure per unit time of the user during exercise and a target energy expenditure per unit time of the user by use of an output of said body constitution estimating unit and the user's body weight ... | Panasonic Electric Works, Ltd. | A63B071/00
US08948471 | B2 | 2015-02-03 | Image registration of multiple medical imaging modalities using a multiple degree-of-freedom-encoded fiducial device | 2007-11-30 | Disclosed is a system and method for registering images from two medical imaging modalities in a six-degree-of-freedom coordinate space, for the purpose of providing real time registered imagery for optimization of medical procedures. The system and method uses fiducial that is visible to the first imager and in a fixed position and orientation relative to the second imager. The first imager may be an X-ray device such as a C-arm fluoroscope, and the second imager may be a Transrectal Ultrasound (TRUS) device. | 1. A system for registering images from two medical imaging modalities comprising: a first imager having a first coordinate frame;a second imager having a second coordinate frame; anda data system,wherein the data system includes a non-transitory computer readable medium encoded with computer instructions for performing an image registration process including: acquiring a first image from the first imager;acquiring a second image from the second imager;determining a first coordinate transformation between the first coordinate frame and a fiducial coordinate frame corresponding to a fiducial; andregistering the first image to the second coordinate frame,wherein the determining the first coordinate transformation includes: providing a current estimated coordinate transformation;generating a current estimated image corresponding to the current estimated coordinate transformation;determining an error function between the first image and the current estimated image;determining a multidimensional derivative corresponding to the current estimated coordinate transformation; anddetermining a new estimated coordinate transformation corresponding to the error function and the multidimensional derivative. 2. The system of claim 1, wherein the first imager is a C-arm fluoroscope. 3. The system of claim 1, wherein the second imager is an ultrasound system. 4. The system of claim 1, wherein the fiducial includes a Tantalum wire. 5. The system of claim 1, wherein the image registration process further includes: acquiring a third image from the first imager; determining a second coordinate transformation between the first coordinate frame and the fiducial coordinate frame; andregistering the third image to the second coordinate frame. 6. The system of claim 5, wherein the image registration process further includes: identifying an object in the first and third images; andlocating the object in a three-dimensional space corresponding to the second coordinate frame. 7. The system of claim 6, wherein the image registration process further includes: displaying the object in the second image. 8. The system of claim 1, wherein the data system including the non-transitory computer readable medium encoded with computer instructions for performing the image registration process further includes: a fiducial having features visible to the first imager, the fiducial having the fiducial coordinate frame that has a known position and orientation relative to the second coordinate frame. 9. The system of claim 8, wherein the features include: a point;a line; andan ellipse. 10. The system of claim 9, wherein the features are arranged to provide a substantially distinct image corresponding to a viewing angle of the first imager. 11. A method for registering images from a first imaging modality to a second imaging modality, the method comprising: acquiring a first image from a first imager having a first coordinate frame;acquiring a second image from a second imager havin... | The John Hopkins University Acoustic MedSystems, Inc. | A61B005/00 A61B006/00 A61B008/12
US08262554 | B2 | 2012-09-11 | Apparatus and method for determining an optimized path traversal for radiation treatment delivery system | 2007-03-30 | An apparatus and method for determining optimized path traversal in a radiation delivery system is described. In one embodiment, determining an optimized path traversal in a radiation delivery system involves providing a plurality of spatial nodes used in a treatment plan, where each of the plurality of spatial nodes represents a position of a radiation source made available to the treatment plan for delivering radiation to a target, identifying a number of unused spatial nodes, from among the plurality of spatial nodes, at which radiation is not delivered according to the treatment plan, and skipping travel to one or more of the unused nodes by the radiation source when administering the treatment plan. Other embodiments are also described. | 1. A method, comprising: providing a plurality of spatial nodes used in a treatment plan, wherein the plurality of spatial nodes is a partial set selected from a superset of all possible spatial nodes, each of the plurality of spatial nodes representing a position of a radiation source made available to the treatment plan for delivering radiation to a target;identifying a number of unused spatial nodes, from among the plurality of spatial nodes, at which radiation is not delivered according to the treatment plan; andskipping travel to one or more of the unused nodes by the radiation source when administering the treatment plan. 2. The method of claim 1 further comprising determining a collision-free path of travel for the radiation source that does not contain an obstruction between any two consecutive spatial nodes according to the treatment plan. 3. The method of claim 1 further comprising determining a collision-free path of travel for the radiation source that contains no obstruction, by skipping the one or more of the unused spatial nodes. 4. The method of claim 3 further comprising moving the radiation source along the collision-free path of travel that does not contain an obstruction between any two consecutive spatial nodes in the administering of the treatment plan after the skipping travel to one or more of the unused nodes by the radiation source. 5. The method of claim 3 further comprising moving the radiation source according to the collision-free path of travel to deliver radiation treatments at a number of selected spatial nodes. 6. The method of claim 3 wherein the determining of a collision-free path of travel for the radiation source is based upon at least one of a look-up table or a reference database that is generated from an off-line simulation. 7. The method of claim 3, further comprising moving the radiation source according to treatment plan where the radiation source skips one or more unused spatial nodes and travels in a shorter path defined by a path without the one or more unused spatial nodes relative to a path with the one or more unused spatial nodes. 8. The method of claim 1 wherein the plurality of spatial nodes is ordered in sequence from a lowest sequenced spatial node to a highest sequenced spatial node. 9. The method of claim 8 wherein the treatment plan begins from the home position to the lowest sequenced spatial node and terminates at the home position after the highest sequenced spatial node. 10. The method of claim 1 wherein the one or more unused nodes is skipped if a path of travel between a spatial node before and a spatial node after the one or more unused nodes has no obstruction. 11. The method of 1 wherein the radiation source is coupled to a robotic arm having 5 or more degrees of freedom. 12. An apparatus, comprising: means for providing a plurality of spatial nodes used in a treatment plan, wherein the plurality of spatial nodes is a partial set selected from a superset of all possibl... | Accuray Incorporated | A61N005/10
US08046048 | B2 | 2011-10-25 | Resonator with adjustable capacitance for medical device | 2008-07-29 | Systems and methods for a resonator with an adjustable capacitance for a medical device. In one embodiment, a resonator system includes a resonator device with an LC resonator circuit that has an adjustable capacitance, an inductor coil in series with the adjustable capacitance, and an adjustable capacitance control that can control the adjustable capacitance to obtain different particular capacitance values. This embodiment also includes a medical device, positioned with the resonator device, so that at least a portion of the inductor coil surrounds a space that is surrounded by at least a portion of the medical device. | 1. A resonator system, comprising: a resonator device with an LC resonator circuit, including: an adjustable capacitance;an inductor coil in series with the adjustable capacitance; anda processor coupled to the adjustable capacitance, wherein: the processor can determine a resonant capacitance based at least in part on one or more signals transmitted by at least one sensor, wherein the at least one sensor includes a magnetic flux sensor; andthe processor can adjust the adjustable capacitance to the resonant capacitance subsequent to determining the resonant capacitance; anda device deliverable into a body, wherein at least a portion of the device forms at least a portion of the inductor coil. 2. The resonator system of claim 1, wherein: the one or more signals represent one or more sensed fluxes of one or more magnetic fields. 3. The resonator system of claim 1, wherein: the at least one sensor further includes a current sensor; andthe one or more signals represent one or more sensed electrical currents through a path of the LC resonator circuit. 4. The resonator system of claim 1, wherein: the at least one sensor further includes a voltage sensor; andthe one or more signals represent one or more electrical voltage differentials across the adjustable capacitance. 5. The resonator system of claim 1, wherein the resonator device includes a memory coupled to the processor, wherein the memory can store one or more values that represent the one or more signals transmitted to the processor. 6. The resonator system of claim 1, wherein the resonator device includes a selector coupled to the processor, wherein the selector can be set to one or more settings that represent one or more user inputs. 7. A resonator system, comprising: a resonator circuit that includes an inductor coil in series with an adjustable capacitance;a processor coupled to the resonator circuit;a magnetic flux sensor coupled to the processor, wherein: the magnetic flux sensor can sense a flux from a magnetic field;the magnetic flux sensor can transmit a signal based on the sensed flux to the processor;the processor can determine a resonant capacitance based at least in part on the signal; andthe processor can adjust the adjustable capacitance to the resonant capacitance subsequent to determining the resonant capacitance; andan implantable medical device, wherein at least a portion of the inductor coil surrounds a space that is surrounded by at least a portion of the implantable medical device. 8. The resonator system of claim 7, wherein the adjustable capacitance includes a plurality of individually switched parallel capacitors. 9. The resonator system of claim 8, wherein the plurality of individually switched parallel capacitors are of increasing size. 10. The resonator system of claim 7, wherein the adjustable capacitance includes a varactor. 11. The resonator system of claim 10, wherein the varactor includes a range of capacitances of 1 picoFarad. 12. The resonator syste... | Boston Scientific Scimed, Inc. | A61B005/05
US08301254 | B2 | 2012-10-30 | Cross-band communications in an implantable device | 2008-07-14 | An ambulatory monitoring device includes a sensor to monitor a physiological signal and a battery power source. The device also includes a wireless receiver adapted to monitor a first frequency band having frequencies below 1 MHz and configured to detect and receive, using less than 10 micro-amps of current from the battery power source when operating, wireless communications within the first frequency band from a remote device at least one meter away. The device further includes a wireless transmitter adapted to transmit—after receipt from the remote device of a first wireless communication within the first frequency band that includes an invitation for further communication—a second wireless communication in a second frequency band having frequencies above 10 MHz, the second wireless communication comprising data indicative of the physiological signal as sensed by the sensor. | 1. An ambulatory monitoring device, comprising: a) a battery power source;b) a sensor powered by the battery and configured for monitoring a physiological signal of a subject;c) a wireless receiver that uses less than 10 micro-amps of current from the battery power source when operating and is configured to receive wireless communications from a remote device within a first frequency band of a range of frequencies below 1 MHz when the monitoring device is at least 1 meter from the remote device; andd) a wireless transmitter powered by the battery and that is separate from the receiver,e) wherein after receipt by the receiver of a first wireless communication from the remote device at the first frequency band, the wireless transmitter is configured to transmit a second wireless communication in a second frequency band of a range of frequencies above 10 MHz, and wherein the second wireless communication is of the physiological signal sensed by the sensor. 2. The ambulatory monitoring device of claim 1, wherein the wireless receiver is configured to duty-cycle as it monitors for the first frequency band at some predetermined fraction of continuous time to conserve battery power. 3. The ambulatory monitoring device of claim 1, wherein the receiver is configured to monitor the first frequency band for 50 ms each second. 4. The ambulatory monitoring device of claim 1, wherein the wireless receiver is configured to continuously monitor the first frequency band. 5. The ambulatory monitoring device of claim 1, wherein the first frequency band is from about 125 kHz to about 134 kHz. 6. The ambulatory monitoring device of claim 1, wherein the second frequency band is from about 402 MHz to about 405 MHz. 7. The ambulatory monitoring device of claim 1, wherein the second frequency band is from about 902 MHz to about 928 MHz. 8. The ambulatory monitoring device of claim 1, wherein the first wireless communication within the first frequency band includes a selection of a channel within the second frequency band, and wherein the wireless transmitter is configured to transmit the second wireless communication at a frequency corresponding to the channel. 9. The ambulatory monitoring device of claim 1, wherein following the wireless transmitter's transmission of the second wireless communication within the second frequency band, the wireless receiver is configured to receive a third wireless communication from a remote device, wherein receipt of the third wireless communication within the first frequency band indicates to the ambulatory monitoring device that the remote device received the second wireless communication. 10. The ambulatory monitoring device of claim 1, wherein the receiver is programmed to periodically receive the first wireless communication within the first frequency band. 11. The ambulatory monitoring device of claim 1 wherein the receiver is programmed to receive the first wireless communication within the first frequency band at least... | Greatbatch Ltd. | A61N001/00
US08233961 | B2 | 2012-07-31 | Magnetic resonance device and method for perfusion determination | 2008-06-20 | In a method and a device for automatic determination of perfusion by using a magnetic resonance system, multiple first MR data sets are thereby acquired from a volume element over time with a perfusion-sensitive imaging sequence, and multiple second MR data sets of the same volume element are acquired over time with a control imaging sequence, in particular a perfusion-insensitive imaging sequence. These first MR data sets and the second MR data sets are subjected to a statistical time series analysis in order to determine the perfusion in the volume element. | 1. A method for automatic determination of perfusion in a magnetic resonance (MR) system, comprising the steps of: (a) generating multiple first MR data sets from a volume element of a subject over time using a perfusion-sensitive imaging sequence;(b) generating multiple second MR data sets of said volume element of said subject over time using a control MR imaging sequence;(c) obtaining a statistical analysis result by statistically analyzing said first MR data sets and said second MR data sets with at least one model function embodying at least one regressor, and selecting said at least one regressor from the group consisting of regressors representing breathing of the subject, regressors representing a heartbeat of the subject, regressors representing a physiological state of the subject, and regressors representing functional activity of the subject, and, using said statistical analysis result, determining perfusion in said volume element, as determined perfusion;(d) tracking quality of the determined perfusion in said volume element by also statistically analyzing said first MR data sets and said second MR data sets to obtain at least one further result, other than said statistical analysis result from which said perfusion is determined, selected from the group consisting of a reliability indicator for said determination of said perfusion of said volume element, information describing a contrast-to-noise ratio for an image point of said volume element, and information describing at least one characteristic of said subject that may produce an artifact in an image reconstructed from said first MR data sets and said second MR data sets and using said at least one further result to exclude at least a portion of at least one of said first and second MR data sets for which said at least one further result indicates a poor quality, to produce quality-tracked first and second MR data sets; and(e) continuously reconstructing MR images of said volume from said quality-tracked first MR data sets and said second MR data sets, and continuously updating said MR images with newly acquired multiple first MR data sets and multiple second MR data sets to which steps (c) and (d) are also applied. 2. A method as claimed in claim 1 comprising employing a perfusion-insensitive MR imaging sequence as said control MR imaging sequence. 3. A method as claimed in claim 2 comprising generating said perfusion-sensitive MR imaging sequence using a perfusion-sensitive magnetic field gradient adjustment, and generating said perfusion-sensitive MR imaging sequence using a perfusion-sensitive magnetic field gradient adjustment. 4. A method as claimed in claim 1 comprising reconstructing a sequence of said MR images from said first MR data sets and said second MR data sets and identifying a signal curve over time of at least one individual image point in said sequence of said MR images, and using said signal curve to produce said signal analysis result. 5. A method as cla... | Siemens Aktiengesellschaft | A61B005/055 G01V003/00
US07894885 | B2 | 2011-02-22 | Coherent signal rejection in ECG | 2007-05-02 | A method for monitoring an electrocardiogram (ECG) signal of a subject, includes digitally sampling an average signal from at least a first ECG electrode, determining an average interference frequency, and digitally sampling and buffering a raw ECG signal from at least a second ECG electrode. The method further includes: filtering the raw ECG signal to generate a residual signal; calculating, based on the residual signal, a first amplitude and a first phase shift of a primary interference signal at the average interference frequency and a second amplitude and a second phase shift of one or more harmonic interference signals at respective multiples of the average interference frequency; and digitally subtracting the primary interference signal and the one or more harmonic interference signals from the raw ECG signal so as to generate and output a clean ECG signal. | 1. A method for monitoring an electrocardiogram (ECG) signal of a subject, comprising: digitally sampling an average signal from ECG electrodes attached to the subject, the average signal comprising a Wilson Central Terminal (WCT) signal;processing the WCT signal with a band-pass filter to establish a band-pass filter signal, the band-pass filter signal having a line interference component;processing the band-pass filter signal with a frequency meter and determining a frequency of the line interference component;determining an average interference frequency of the WCT signal based on the frequency of the line interference component;digitally sampling and buffering a raw ECG signal from another ECG electrode attached to the subject;calculating a moving average of the raw ECG signal;calculating a residual signal using the moving average of the raw ECG signal wherein the residual signal comprises frequency components at and above the average interference frequency;calculating, based on the residual signal, a first amplitude and a first phase shift of a primary interference signal at the average interference frequency and a second amplitude and a second phase shift of one or more harmonic interference signals at respective multiples of the average interference frequency; anddigitally subtracting the primary interference signal and the one or more harmonic interference signals from the raw ECG signal so as to generate and output a clean ECG signal. 2. The method according to claim 1, and comprising using an artifact processor for imposing a limit on a rate of change in the first amplitude over time prior to subtracting the primary interference signal from the raw ECG signal. 3. The method according to claim 2, wherein the limit on the rate of change in the first amplitude over time is a first limit and further comprising imposing a second limit on a rate of change in the second amplitude over time prior to subtracting the one or more harmonic interference signals from the raw ECG signal. 4. The method according to claim 1, wherein filtering the raw ECG signal comprises applying to the raw ECG signal a moving average filter, and wherein generating the residual signal comprises subtracting an output of the moving average filter from the raw ECG signal. 5. The method according to claim 1, wherein calculating the first amplitude and the first phase shift of the primary interference signal comprises finding a correlation between the residual signal and a periodic function having a frequency equal to the average interference frequency. 6. The method according to claim 1, wherein calculating the second amplitude and the second phase shift of the one or more harmonic interference signals comprises finding a correlation between the residual signal and a periodic function having a frequency equal to a multiple of the average interference frequency. | Biosense Webster, Inc. | A61B005/0402