## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US20100286569 | A1 | 2010-11-11 | MASSAGE APPARATUS | 2008-12-23 | A massage apparatus includes: a supporting body 1 that supports a body of a massage recipient on his front side and a massage mechanism 2 that causes a massage head 28 which is disposed behind the supporting body 1 to massage a back of the body of the massage recipient. The massage apparatus further includes a contour measurement unit 3 that measures contour of the body of the massage recipient supported by the supporting body 1, and the massage mechanism 2 includes a controller C that controls the movement of the massage head 28 based on information of the body contour obtained by the contour measurement unit 3. The massage apparatus can give a suitable massage to the back of the massage recipient in accordance with a different size of the back different per each massage recipient and a position of a part to be massaged depending on the size. | 1. A massage apparatus, comprising: a support capable of supporting a front face of a human body;a massage mechanism that massages a back of the human body supported by the support, the massage mechanism including a massage head disposed with a space kept from the support and being movable toward the support for coming into contact with the back of the human body; a driving source to control a movement of the massage head; anda contour measurement unit that measures a contour of the human body supported by the support,wherein the controller drives the driving source to control the movement of the massage head based on measurement values measured by the contour measurement unit. 2. The massage apparatus as set forth in claim 1, wherein the contour measurement unit comprises a distance calculation unit that calculates a distance between the massage head and the back of the human body supported by the support,wherein the controller controls the movement of the massage head based on calculation results calculated by the distance calculation unit. 3. The massage apparatus as set forth in claim 1, further comprising a memory that stores information of muscle fiber directions of the back of the human body, wherein the controller extracts a plurality of characteristic parts of the human body from the measurement values measured by the contour measurement unit, andthe distance calculation unit calculates lengths between the characteristic parts, andthe controller divides the back of the human body into regions based on the calculated lengths between the characteristic parts and the stored information of muscle fiber directions, each of the regions including muscle fibers having substantially a same direction, and controls a moving direction of the massage head for each of the regions. 4. The massage apparatus as set forth in claim 1, further comprising a force sensor measuring a force applied to the massage head, wherein the controller feeds back measurement valves measured by the force sensor to the driving source to control a driving output of the driving source. 5. The massage apparatus as set forth in claim 4, wherein the massage head is configured to move in three axial directions, and the force sensor measures each axial component of the force applied to the massage head. 6. The massage apparatus as set forth in claim 5, wherein the memory stores a predetermined threshold to be compared with measurement values measured by the force sensor, and the controller compares the each axial component measured by the force sensor with the threshold and suppresses a movement of the massage head in an axis direction where the force exceeds the threshold. 7. The massage apparatus as set forth in claim 1, further comprising a force sensor measuring a force applied to the massage head; andan operation unit determining a muscle fiber direction of the back of the human body from a force measured by the force sensor,wherein the massage head is configured to mov... | Panasonic Electric Works Co., Ltd. | A61B005/103 A61H007/00 A61H015/00
US20100324402 | A1 | 2010-12-23 | Analyte Monitoring Device and Methods of Use | 2010-08-12 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for processing sensor data from a continuous analyte sensor, comprising: receiving sensor data from a continuous analyte sensor for a first time period, wherein the sensor data comprises at least one sensor data point;processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period; andproviding output, wherein providing output comprises displaying the target analyte value or range of target analyte values on a user interface. 3. The method of claim 2, wherein, the target glucose value is based on an individual physiological pattern. 4. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining whether the target analyte value or range of target analyte values is based on a clinical acceptability. 5. The method of claim 2, wherein the continuous analyte sensor is a continuous glucose sensor. 6. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on at least one parameter that is customizable by the user. 7. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on an amount and/or timing of a medicament delivery and/or caloric consumption. 8. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on a host's physiological pattern. 9. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on the sensor data. 10. The method of claim 2, wherein processing the at least one sensor data point to obtain a target analyte value or range of target analyte values for a future time period comprises determining the target analyte value or range of target analyte values based at least in part on a variation analysis of the sensor data using a clinical, physiological, and/or statistical variation analysis. 11. The method of claim 2, wherein providing output comprises displaying the target analyte value or range of target analyte values by an object representative of the target analyte value or range of target analyte values. 12. The method of claim 2, wherein provid... | Abbott Diabetes Care Inc. | A61B005/145
US20100305427 | A1 | 2010-12-02 | LONG-RANGE PLANAR SENSOR ARRAY FOR USE IN A SURGICAL NAVIGATION SYSTEM | 2009-06-01 | A planar sensor array for use in a surgical navigation system comprising at least one substrate and at two layers of a plurality of planar sensor coils formed on or within the at least one insulating substrate. The planar sensor array may be an electromagnetic planar transmitter coil array or an electromagnetic planar receiver coil array that includes at two layers of a plurality of planar electromagnetic transmitter or receiver spiral-shaped coils formed on or within the at least one substrate. The surgical navigation system includes the use of at least one magnetoresistance reference sensor attached to a fixed object; at least one magnetoresistance sensor attached to an object being tracked; and a planar sensor coil array communicating with the at least one magnetoresistance reference sensor and the at least one magnetoresistance sensor to determine the position and orientation of the object being tracked. | 1. A planar sensor array comprising: at least one insulating substrate; andat least two layers of a plurality of planar sensor coils formed on or within the at least one insulating substrate. 2. The planar sensor array of claim 1, further comprising a first layer of a plurality of planar sensor coils arranged on or within a first layer of the at least one insulating substrate. 3. The planar sensor array of claim 2, further comprising a second layer of a plurality of planar sensor coils arranged on or within a second layer of the at least one insulating substrate. 4. The planar sensor array of claim 3, wherein the second layer of the plurality of planar sensor coils is positioned above the first layer of the plurality of planar sensor coils, and overlaps a center portion of the first layer of the plurality of planar sensor coils. 5. The planar sensor array of claim 1, wherein the plurality of planar sensor coils are made of a conductive material forming a plurality of conductor traces with spaces in-between. 6. The planar sensor array of claim 5, wherein the conductor traces are made of a conductive, low density, low resistivity, radiolucent material. 7. The planar sensor array of claim 6, wherein the conductive, low density, low resistivity, radiolucent material is selected from the group consisting of aluminum, magnesium, carbon nanotubes, graphene, titanium, and their various alloys. 8. The planar sensor array of claim 1, wherein the plurality of planar sensor coils are rectangular-shaped spiral coils. 9. The planar sensor array of claim 1, wherein the plurality of planar sensor coils are circular or elliptical shaped spiral coils. 10. The planar sensor array of claim 1, wherein the plurality of planar sensor coils are a plurality of planar electromagnetic transmitter spiral-shaped coils formed on or within the at least one insulating substrate. 11. The planar sensor array of claim 1, wherein the plurality of planar sensor coils are a plurality of planar electromagnetic receiver spiral-shaped coils formed on or within the at least one insulating substrate. 12. The planar sensor array of claim 1, further comprising a conductive radiolucent material filling open areas within non-coil regions of the planar sensor array. 13. The planar sensor array of claim 12, wherein the conductive radiolucent material is selected from the group consisting of aluminum, magnesium, carbon nanotubes, graphene, titanium, and their various alloys. 14. The planar sensor array of claim 6, further comprising an epoxy or other similar material having an x-ray density approximately matched to the conductor trace material filling open areas in-between the conductor traces of the planar sensor array. 15. The planar sensor array of claim 1, wherein the at least one insulating substrate is rigid. 16. The planar sensor array of claim 1, wherein the at least one insulating substrate is flexible. 17. The planar sensor array of claim 1, wherein the planar sensor ar... | General Electric Company | A61B005/05
US20100312145 | A1 | 2010-12-09 | DEVICE FOR THE FUNCTIONAL DIAGNOSIS OF VESTIBULAR REFLEX ARCS USING MYOGENIC POTENTIALS | 2009-02-23 | The invention relates to a device for the detection of at least one vestibular evoked myogenic potential and the use of this device for the diagnosis of the otolithic organs, particularly for the measurement and/or evaluation of vertigo phenomena in a patient. The use particularly relates to a device for the functional diagnosis of acoustically, mechanically, or electrically evoked vestibular reflexes. | 1. A device for detecting at least one vestibular evoked myogenic potential (VEMP) in a patient, whereinthe device comprises at least one active electrode, a reference electrode, a grounding electrode, an acoustic, mechanical and/or electrical signal generator and a feedback system. 2. The device according to claim 1, whereina muscle tone and/or a visualized pressure is displayed in form of a perceptible, signal on a patient display. 3. The device according to claim 16, whereinthe muscle tone is measured using a pressure sensor, said pressure sensor being an essentially U-shaped, tubular element that can be filled with gas and/or liquid. 4. The device according to claim 3, whereinthe pressure arising in the tubular element is measured with the pressure sensor and can be displayed in a patient display. 5. A method for generating, measuring and/or evaluating at least one vestibular evoked myogenic potential comprising providing the device of claim 1 and generating, measuring and/or evaluating said at least one vestibular evoked myogenic potential via said device. 6. A method for obtaining intermediate results as a basis of otolith organ diagnostics comprising providing the device of claim 1 and obtaining said intermediate results as a basis of otolith organ diagnostics via said device. 7. A method for obtaining intermediate results for establishing a diagnosis of vertigo, especially rotary, postural, lift, movement and/or non-systematic vertigo comprising providing the device of claim 1 and obtaining said intermediate results for establishing said diagnosis of vertigo via said device. 8. A method for detecting intermediate results in finding a diagnosis regarding a function of vestibular reflexes comprising providing the device of claim 1 and detecting said intermediate results for finding a diagnosis regarding the function of vestibular reflexes via said device. 9. The method of claim 7, wherein the muscle tone is measured using a pressure sensor, said pressure sensor being an essentially U-shaped, tubular element that can be filled with gas and/or liquid and the muscle tone is recorded continuously between the VEMP measurements by measuring the potential and presented to the patient instead of and/or in addition to the pressure measurement in the tube in the form of a signal that can be perceived by the patient. 10. The method of claim 10, wherein acoustic, mechanical or electrical stimuli are presented to the patient for the measurement of VEMPs. 11. The method of claim 10, wherein the smallest and largest voltage values of the VEMP are determined within a time window following the signal. 12. The method of claim 10, wherein the points in time of the smallest and largest voltage values of the VEMP and their difference in magnitude are compared and matched with age- and/or sex-related standard values, taking into account the intensity of the muscle potential. 13. The method of claim 7, wherein the influence of age on the calculated qu... | - | A61B005/103
US07842031 | B2 | 2010-11-30 | Bioimpedance measurement system and method | 2005-11-18 | Methods and systems for detecting fluid egress, assessing lesion quality, determining tissue composition or structure, determining ice coverage of catheter tip and providing tissue contact assessment, by providing a catheter having a shaft with a proximal end portion and a distal end portion, the proximal end portion and the distal end portion define at least one fluid pathway therebetween with the shaft having a plurality of electrodes, positioning the catheter at a tissue treatment site, applying an electrical current between at least two of the plurality of electrodes, measuring impedance voltage between the at least two of the plurality of electrodes and, processing the measured impedance voltage caused by the applied electrical current to determine if fluid egress is present, to assess lesion quality, to determine tissue composition, ice cover of catheter tip, and to provide contact assessment. The system may have a control unit, a microprocessor, an impedance measuring device or the like to perform processing of impedance data. | 1. A method for determining fluid egress comprising: positioning a catheter at a tissue treatment site, the catheter comprising: a shaft having a proximal portion, a distal tip, and at least one fluid pathway therebetween;an expandable element defining a cooling chamber therein in fluid communication with the at least one fluid pathway;a first electrode located on an exterior surface of the shaft between the proximal portion and the expandable element; anda second electrode located on the exterior surface of the shaft between the expandable element and the distal tip;applying an electrical current between the first and second electrodes to form an electrical field about the expandable element;measuring impedance between the first and second electrodes; andcorrelating the measured impedance to a condition of fluid egress from the cooling chamber to a point exterior to the catheter, and generating a signal indicating the condition of fluid egress from the cooling chamber to the point exterior to the catheter. 2. The method of claim 1, wherein correlating the measured impedance includes: establishing a normal impedance range;comparing the measured impedance to the normal impedance range; andgenerating a signal if the measured impedance is outside of the normal impedance range. 3. The method of claim 1, wherein the electrical current is a low frequency signal. 4. The method of claim 1, wherein correlating the measured impedance is performed by a microprocessor. 5. The method of claim 1, wherein measuring impedance is performed by an impedance measuring circuit. | Medtronic Cryocath LP | A61B018/18 A61B005/05
US07844322 | B2 | 2010-11-30 | System and method for correlation of patient health information and implant device data | 2006-04-24 | A system and method for correlating health related data for display. The system includes a medical device recording data and a display producing device which correlates the data and simultaneously displays different types of data or displays two sets of the same type of data along with the circumstances at which the two sets of data were recorded. Such displays aid a physician in prescribing and ascertaining the efficacy of cardiac therapies. | 1. A method, comprising: implanting a cardiac rhythm management device, the cardiac rhythm management device including a housing and electrodes;storing, in a memory circuit disposed in the housing, a first data set of cardiac rhythm management data associated with an application of therapy by the cardiac rhythm management device;programming the memory circuit with further therapy parameters to control application of a further therapy by the cardiac rhythm management device;storing a second data set of cardiac rhythm management data associated with an application of the further therapy by the cardiac rhythm management device according to the further therapy parameters;associating the second data set with the further therapy parameters;receiving, with an output unit, the first data set of cardiac rhythm management data associated with implanted operation of the cardiac rhythm management device according to the further therapy parameters;receiving, with the output unit, the second data set of cardiac rhythm management data associated with implanted operation of the cardiac rhythm management device;correlating, with the output unit, the first data set to the second data set;displaying the first and second data sets in correlation simultaneously on the output unit; anddisplaying the further therapy parameters associated with the second data set on the output unit. 2. The method of claim 1, wherein correlating includes transmitting data over a communication network to a correlation data unit. 3. The method of claim 2, wherein transmitting data over a communication network includes transmitting over a global computer network. 4. The method of claim 1, wherein receiving the second data set includes sensing atrial events and ventricular events and transmitting the sensed events to the correlation data unit. 5. The method of claim 4, wherein transmitting the sensed events includes sending the second data over a communication network includes transmitting over a global computer network. 6. The method of claim 4, wherein receiving the second data set includes calculating an average atrial rate from the sensed atrial events. 7. The method of claim 4, wherein receiving the second data set includes determining an average ventricular rate from sensed ventricular events. 8. The method of claim 1, wherein correlating includes relating the first data set to the second data set based on a non-time event. 9. A method, comprising: implanting a cardiac rhythm management device including a housing and electrodes;storing, in a memory circuit disposed in the housing, a first data set relating to patient activity monitored via implanted operation of the cardiac rhythm management device;programming therapy parameters to control operation of the cardiac rhythm management device in the memory circuit;storing a second data set of cardiac rhythm management data associated with implanted operation of the cardiac rhythm management device according to the therapy parameter... | Cardiac Pacemakers, Inc. | A61B005/04
US07860558 | B2 | 2010-12-28 | ECG lead misplacement detection and correction | 2007-03-29 | A physiological parameter analysis system (10) detects ECG electrode wire misplacement. The system (10) includes a transformation component (32) that transforms original ECG information and combinations thereof from a first ECG lead system to a second ECG lead system and an inverse transformation component (34) that derives ECG information in the first ECG lead system from the transformed ECG information and the combinations thereof. The system (10) further includes an analysis component (38) that determines a correct ECG lead configuration in the first ECG lead system from among the original ECG information and the combinations thereof based on the derived ECG information. | 1. A physiological parameter analysis system (10) that detects ECG electrode wire misplacement, comprising: a receiver for receiving original ECG information corresponding to a plurality of electrode leads in a first lead system;a combinationalyzer/decombinationalyzer (28) that is operatively connected to the receiver and that generates one or more combinations from the original ECG information, each combination representing a permutation of the received electrode leads;a processor that is operatively connected to the combinationalyzer/decombinationalyzer, the processor having; a transformation component (32) that transforms the original ECG information and each combination from a first lead system to a second lead system different from the first lead system; andan inverse transformation component (34) that reverse transforms the transformed original ECG information and each combination from the second lead system to the first lead system;a comparator (36), operatively connected to the combinationalyzer/decombinationalyzer and the inverse transformation component, which computes a difference between the original ECG information and the reverse transformed original ECG information, and which computes a difference between each combination and each reverse transformed combination;an analyzer (38), operatively connected to the comparator, that determines a misplaced electrode wire in the original ECG information based on the computed differences; anda notification component (56), operatively connected to the analyzer, that transmits a notification of the analyzer determination. 2. The system (10) as set forth in claim 1, wherein the first lead system is one of a three, five, twelve, fifteen, and sixteen lead ECG system. 3. The system (10) as set forth in claim 1, wherein the second lead system is an EASI ECG lead system. 4. The system (10) as set forth in claim 1, wherein the notification comprises an alarm based on the analyzer determination of the misplaced electrode wire in the original ECG information. 5. The system (10) as set forth in claim 1, wherein the notification includes a message sent to at least one of an email address, a cell phone, a conventional telephone, a pager, a personal data assistant (PDA), a central monitoring station, a bedside monitor, a satellite, a viewing screen, an alarm light, and a network. 6. The system (10) as set forth in claim 5, wherein the notification includes information about the original ECG information, a subject patient, an operator who obtained the data, and a location of the patient. 7. The system (10) as set forth in claim 1, further including: a correction component (52, 54), operatively connected to the analyzer, that automatically corrects the original ECG information based on the analyzer determination of the misplaced electrode wire in the original ECG information. 8. The system (10) as set forth in claim 7, wherein the notification component (56) transmits a message indicating that the orig... | Koninklijke Philips Electronics N.V. | A61B005/0402
US20100317985 | A1 | 2010-12-16 | SYSTEM FOR QUANTITATIVE ASSESSMENT OF CARDIAC ELECTRICAL EVENTS | 2009-06-12 | Systems and methods for characterizing aspects of an electrocardiogram signal are presented, wherein primary and secondary analysis schemas are utilized to determine the timing of the end of a signal wave, such as a descending Twave, with precision. In one embodiment, the primary analysis schema involves comparing voltage amplitudes within a given sampling window and the secondary analysis schema involves comparing the results of primary analysis for successive sampling windows. The system may comprise a processor or microcontroller embedded into a system such as an electrocardiogram hardware system, personal computer, electrophysiology system, or the like. | 1. A system for determining a signal wave transition point, comprising: a. a memory device configured to store data pertinent to one or more signal waves sampled from electrodes operably coupled to one or more tissue structures; andb. a processor operably coupled to the memory device and configured to access the data and determine a transition point associated with the one or more signal waves, the processor configured to 1) sample a first plurality of points of a signal wave in a first time window, the first plurality comprising at least a first-in-time point and a last-in-time point within the first time window;2) sample a second plurality of points of the signal wave in a second time window different in time from the first time window, the second plurality comprising at least a first-in-time point and a last-in-time point within the second time window;3) compare the values of the first plurality relative to each other to determine whether an intra-window patterning rule has been broken; and4) conduct a secondary analysis subsequent to determining that an intra-window patterning rule has been broken, the secondary analysis comprising comparing the values of the second plurality relative to each other to determine whether the intra-window patterning rule has been broken. 2. The system of claim 1, further comprising a display operatively coupled to the processor, wherein the processor is further configured to cause a graphical image of at least one of the one or more signal waves to be depicted on the display, including a graphical indication of a respective pertinent transition point on each of the one or more signal waves. 3. The system of claim 1, wherein the processor and memory device are operatively coupled to one or more printed circuit boards. 4. The system of claim 3, wherein the one or more printed circuit boards comprise a card having a common housing and an electronic interface bus. 5. The system of claim 3, wherein the card is configured to interface with a personal computer. 6. The system of claim 1, wherein the processor and memory device comprise an application specific integrated circuit configured to be embedded into a parent electronic device. 7. The system of claim 1, wherein the processor and memory device comprise a field programmable gate array configured to be embedded into a parent electronic device. 8. The system of claim 1, wherein the processor and memory device are operatively coupled to an operating room electronic device selected from the group consisting of an electrophysiology mapping system, an echocardiography system, and a fluoroscopic imaging system. 9. The system of claim 8, wherein the processor and memory device are operably coupled to the operating room electronic device by an Ethernet connection. 10. The system of claim 9, wherein at least two of the processor, memory device, and operating room electronic device are configured to communicate with each other via a protocol selected from the group ... | NEWCARDIO, INC. | A61B005/044 A61B005/0402
US20100298683 | A1 | 2010-11-25 | PHYSIOLOGICAL MONITORING DEVICES AND METHODS | 2008-09-17 | Devices and methods are described for wirelessly monitoring an emergency responder. In some embodiments, a sensor acquires values of carboxyhemoglobin in blood. The values are recorded and are used to provide feedback to a user. The feedback includes at least one of visible, tactile, and audible information. | 1. A method, of monitoring emergency personnel, comprising: acquiring values of levels of carboxyhemoglobin in a person's blood during a sampling period, the person being in a region near or at a source of carbon monoxide;recording the values on machine-readable media;determining a regression function for the values with respect to time;determining at least one of an attribute of the function and a future value of a carboxyhemoglobin level in the person's blood, the future value being determined by extrapolation of the function; andproviding a user with safety information based on the at least one of the attribute of the function and the future value; wherein the safety information comprises at least one of visible, tactile, and audible information; andwherein the safety information comprises at least one of a time of remaining safety in the region, a warning to exit the region, an indication of continuing safety in the region, and an indication of impending danger. 2. The method of claim 1, wherein the attribute of the function comprises at least one of a derivative and an integral. 3. A machine-readable medium having machine-executable instructions for performing the method of claim 1. 4. A method, of monitoring emergency personnel, comprising: during a sampling period and of a person in a region near or at a source of carbon monoxide, acquiring (a) values of levels of carboxyhemoglobin in the person's blood, and (b) a value of a parameter selected from the group consisting of the person's heart rate, an indicator of the person's peripheral perfusion, and ambient carbon monoxide;recording the values of (a) and (b) on machine-readable media;determining a regression function for at least one of: (i) each of (a) and (b) with respect to time; and (ii) a combination of (a) and (b) with respect to time;determining at least one of: (i) an attribute of the function; and (ii) a future value of at least one of (a) and (b), the future value being determined by extrapolation of the function; andproviding a user with safety information based on the at least one of the attribute of the function and the future value; wherein the safety information comprises at least one of visible, tactile, and audible information; andwherein the safety information comprises at least one of a time of remaining safety in the region, a warning to exit the region, an indication of continuing safety in the region, and an indication of impending danger. 5. The method of claim 4, wherein the attribute of the function comprises at least one of a derivative and an integral. 6. The method of claim 4, further comprising: determining an expected level of carboxyhemoglobin present in the person based on the value of ambient carbon monoxide. 7. A machine-readable medium having machine-executable instructions for performing the method of claim 4. 8. A device for monitoring emergency personnel comprising: means for acquiring values of levels of carboxyhemoglobin in a person's blood d... | Deltin Corporation, A California Corporation | A61B005/145 G08B023/00
US20100324404 | A1 | 2010-12-23 | ICG/ECG MONITORING APPARATUS | 2009-06-22 | An ICG/ECG electrode includes a first electrode and a second electrode, wherein at least one of the first or second electrode senses both an ICG signal and an ECG voltage signal. A physiologic parameter monitoring apparatus includes a set of electrodes, including an electrode for sensing both an ICG voltage signal and an ECG voltage signal corresponding to a patient. The apparatus further includes an ICG monitor for processing the ICG voltage signal sensed by the electrode and an ECG monitor for processing the ECG voltage signal sensed by the same electrode. | 1. An ICG/ECG electrode, comprising: a first electrode contact; anda second electrode contact;wherein one of the first or the second electrode contacts senses both an ICG voltage signal and an ECG voltage signal. 2. The ICG/ECG electrode of claim 1, further comprising: a substrate, wherein the first and second electrode contacts are affixed to the substrate; anda barrier disposed on the substrate between the first and second electrode contacts. 3. The ICG/ECG electrode of claim 2, wherein the barrier includes an adhesive material for affixing the electrode contacts to a patient. 4. The ICG/ECG electrode of claim 2, wherein the barrier mitigates cross-coupling of an ICG electrical current signal applied to one of the first or second electrode contacts and the other of the first or second electrode contacts sensing the ECG voltage signal. 5. The ICG/ECG electrode of claim 2, wherein the first and second electrode contacts are affixed to wells of the substrate. 6. The ICG/ECG electrode (128) of claim 1, further comprising: first and second barriers, wherein the first and second electrode contacts are surrounded by the first barrier and second barriers (212, 214). 7. The ICG/ECG electrode of claim 6, wherein the first barrier is a non-adhesive barrier and the second barrier is an adhesive barrier. 8. The ICG/ECG electrode of claim 7, wherein the first and second barriers mitigate cross-coupling of an electrical current signal applied to one of the first or second electrode contacts and the other of the first or second electrode contacts sensing the ECG voltage signal. 9. The ICG/ECG electrode of claim 1, wherein no more than three ICG/ECG electrodes are used for ICG/ECG monitoring. 10. The ICG/ECG electrode of claim 1, wherein the ICG/ECG electrode is part of a carrier that includes no more than three ICG/ECG shared electrodes. 11. A physiologic parameter monitoring apparatus, comprising: a set of electrodes, including an electrode for sensing both an ICG voltage signal and an ECG voltage signal corresponding to a patient;an ICG monitor for processing the ICG voltage signal sensed by the electrode; andan ECG monitor for processing the ECG voltage signal sensed by the same electrode. 12. The physiologic parameter monitoring apparatus of claim 11, further comprising: a synchronizer that time-synchronizes the sensed ICG and ECG voltage signals. 13. The physiologic parameter monitoring apparatus of claim 12, wherein the ICG and ECG voltage signals are acquired based on a common acquisition clock. 14. The physiologic parameter monitoring apparatus of claim 12, wherein the ICG and ECG voltage signals are acquired based on different acquisition clocks, and the synchronizer time-synchronizes the clocks, thereby time-synchronizing the sensed ICG and ECG voltage signals. 15. The physiologic parameter monitoring apparatus of claim 11, wherein the ICG and ECG voltage signals are concurrently sensed. 16. The physiologic parameter monitoring apparat... | Analogic Corporation | A61B005/0408 A61B005/02
US20100286547 | A1 | 2010-11-11 | BREATHING DETECTION DEVICE AND METHOD FOR DETECTING ASYNCHRONOUS OR ARRESTED BREATHING | 2010-05-05 | The present invention in one aspect is a breathing detection device with a signal source and receiver. A card between the source and receiver alternates between a first state of signal reception and a second state of signal interference, on card movement. A cord moves the card in one direction on pulling, and in reverse on retraction. The card can be a disc, belt, or any configuration conforming to a principle of card movement to affect a difference in signal receiving state. Another aspect the invention is an asynchronous breathing detection method providing a first and a second detection device as described, both in communication. In various configurations, one or more processors is provided to receive and process data to determine magnitude and direction of cord and card movement. | 1. A breathing detection device comprising: a signal source, to emit a signal;a signal receiver, to receive the signal;a card disposed between the source and the receiver to alternate between a first state of receiver signal reception and a second state of receiver signal interference, on movement of the card;a movement means to move the card; anda cord defining two ends, a first end engaging the movement means to move the card in one direction on pulling of the cord, and in a reverse direction on retraction of the cord. 2. The device in claim 1 further comprising a housing, the device being encompassed within the housing. 3. The device in claim 1 wherein the movement means is an axle and cogwheel configuration. 4. The device in claim 1 further comprising the cord guide attached to the movement means. 5. The device in claim 1 wherein the card is disc-shaped. 6. The device in claim 1 wherein the movement means comprises a conveyor adapted to receive a conveyor belt. 7. The device in claim 1 wherein the card is comprised of a conveyor belt. 8. The device in claim 1 further comprising an outside processor to process receiver data. 9. The device in claim 1 further comprising the power source to power the signal source. 10. The device in claim 1 further comprising a speaker to emit sound related to data received from the receiver. 11. The device in claim 1 further comprising a light-emitting diode to indicate on and off state. 12. The device in claim 1 further comprising an alarm to provide alerts related to receiver data. 13. The device in claim 1 further comprising an input and output to receive and transmit data, the input being in communication with the processor. 14. An asynchronous breathing detection method comprising the following steps: providing a first breathing detection device comprising a signal source, to emit a signal; a signal receiver, to receive the signal; a card disposed between the source and receiver to alternate between a first state of receiver signal reception and a second state of receiver signal interference, on movement of the card; a movement means to move the card; and a cord defining two ends, the first end engaging the movement means to move the card in one direction on pulling of the cord, and in a reverse direction on retraction of the cord;providing a second breathing detection device same as the first; the first and second devices being in communication; andproviding a processor in communication with the first and second device, to receive and process data to determine magnitude and direction of cord and card movement. 15. The method in claim 14 wherein the first and second detection devices are attached to clothing. 16. The method in claim 14 wherein the first and second detection devices are embedded in clothing. 17. The method in claim 14 wherein the first and second devices are connected about a chest and an abdomen respectively. 18. The device in claim 1 wherein the device is embedded in cloth... | - | A61B005/08
US20100310469 | A1 | 2010-12-09 | VACCINE IMMUNOTHERAPY FOR IMMUNE SUPPRESSED PATIENTS | 2010-05-26 | A method of immunotherapy to treat cancer or a synergistic anti-cancer treatment by administering an effective amount of natural cytokine mixture (NCM), an effective amount of cyclophosphamide (CY), or an effective amount of indomethacin (INDO), wherein the NCM, CY, or INDO are administered singly or in communications thereof. An anti-metastatic treatment method by promoting differentiation and maturation of immature dendritic cells in a lymph node; allowing presentation thereof; and preventing development of metastasis. A method of using NCM as a diagnostic skin test for predicting treatment outcome. A method of pre-treating dendritic cells (DC) and a method of treating monocyte defects characterized by sinus histiocytosis or a negative NCM skin test. Compositions and method for eliciting an immune response to endogenous or exogenous tumor antigens. | 1. A method of immunotherapy to treat cancer by administering an effective amount of a natural cytokine mixture (NCM) including cytokines selected from the group consisting essentially of IL-1, IL-2, IL-6, IL-8, IL-12, IFN-δ, TNF-α, GM-CSF, G-CSF, recombinants thereof, and combinations thereof. 2. The method according to claim 1, wherein said administering step is defined as administering 75 to 500 units IL-2 equivalence. 3. The method according to claim 1, wherein said administering step is defined as bilaterally administering the NCM into lymphatics that drain into lymph nodes. 4. The method according to claim 1, wherein said administering step is defined as unilaterally administering the NCM. 5. The method according to claim 1, wherein said administering step is defined as administering the NCM for at least 1 to 10 days. 6. The method according to claim 5, wherein said administering step is further defined as administering the NCM up to about 20 days. 7. The method according to claim 6, wherein said administering step is further defined as administering the NCM bilaterally and for about 10 days. 8. The method according to claim 1, wherein said administering step is defined as administering the NCM prior to surgery or radiotherapy. 9. The method according to claim 1, wherein said administering step is defined as administering the NCM during recurrence of tumors. 10. The method according to claim 1, further including the step of administering an effective amount of cyclophosphamide (CY). 11. The method according to claim 1, further including the step of administering an effective amount of a nonsteroidal anti-inflammatory drug (NSAID) selected from the group consisting of indomethacin (INDO), Ibuprofen, celecoxib (Celebrex®), rofecoxib (Vioxx®), CoxII inhibitors, and combinations thereof. 12. A method of immunotherapy to treat cancer by administering an effective amount of CY and an effective amount of INDO. 13. A synergistic anti-cancer treatment method by administering an effective amount of CY and an effective amount of NSAID selected from the group consisting essentially of indomethacin (INDO), Ibuprofen, celecoxib (Celebrex®), rofecoxib (Vioxx®), CoxII inhibitors, and combinations thereof. 14. A method of immunotherapy to treat cancer by administering an effective amount of CY in combination with an effective amount of INDO and an effective amount of IFN-δ, IL-2, IL-1, and TNF-α. 15. A method of immunotherapy to treat cancer by administering an effective amount of CY in combination with an effective amount of INDO and an effective amount of recombinant IL-2, recombinant IFN-δ, recombinant TFN-α, and recombinant IL-1. 16. A synergistic anti-cancer treatment comprising the steps of administering an effective amount of CY and INDO in combination with an NCM including cytokines selected from the group consisting essentially of IL-1, IL-2, IL-6, IL-8, IL-12, IFN-δ, TNF-α, GM-CSF, G-CSF, recombinants thereof, and combinations there... | - | A61B005/00 A61K038/20 A61K038/21 A61K039/00 C12N005/02 A61K038/00 A61K031/66 A61P029/00 A61P035/02
US20100324533 | A1 | 2010-12-23 | CATHETER SYSTEM AND METHOD FOR DELIVERY OF THERAPEUTIC COMPOUNDS TO CARDIAC TISSUE | 2010-08-26 | This is a method and an apparatus for the treatment or introduction of contrast fluids into tissue, particularly cardiac tissue. The apparatus includes a catheter having an elongated flexible body and a tissue infusion apparatus including a hollow infusion needle configured to secure the needle into the tissue when the needle is at least partially inserted into the tissue to help prevent inadvertent removal of the needle from the tissue. This permits the selected treatment or contrast fluid to be confined to a specific site. The catheter may also include a visualization assembly including a transducer at the distal end of the body. | 1. A method for delivering a therapeutic compound to an endocardial delivery site, said method comprising: advancing an elongated catheter body comprising at least one electrode and a hollow penetrating structure into a cardiac chamber;sensing and selecting an area for treatment by measuring electric signals within the cardiac chamber;correlating the electric signals with an electrophysiological map;and delivering the therapeutic compound into the cardiac tissue through the hollow penetrating structure. 2. The method as in claim 1, wherein the at least one electrode comprises a plurality of electrodes. 3. The method as in claim 2, wherein the plurality of electrodes comprise a plurality of ring electrodes on the elongated catheter body. 4. The method as in claim 2, wherein the plurality of electrodes comprise a tip electrode on the catheter body. 5. The method as in claim 1, further comprising steering the catheter to engage the location of the endocardial delivery site. 6. The method as in claim 5, wherein steering comprises pulling cables to steer a distal portion of the catheter body. 7. The method as in claim 1, further comprising securing the end of the hollow penetrating structure into tissue at the endocardial delivery site while the therapeutic compound is being delivered. | The Regents of the University of California | A61B005/05 A61M025/00
US07837628 | B2 | 2010-11-23 | Electro-mechanical system for electrophysiology and hemodynamic studies | 2006-09-26 | An electrical-mechanical system may be used in electrophysiology and hemodynamic studies. The system may include a mounting arrangement couplable to a support surface for supporting a patient. At least one catheter input module may be electrically and mechanically coupled to the mounting arrangement, the catheter input module being connectable to catheters for use in the patient. A central processing device may be attached to the support surface and placed in communication with the mounting arrangement. | 1. An electro-mechanical system used in combination with at least one catheter in performing a medical diagnostic procedure on a patient situated on a support structure, comprising: a mounting arrangement configured to be coupled to the support structure for sustaining the patient, wherein the mounting arrangement extends either above or below the patient support structure;at least one catheter input module mechanically coupled to the mounting arrangement, the catheter input module being connected in communication with the at least one catheter; anda central processing device mechanically coupled at the support surface structure and connected in communication with the at least one catheter input module,wherein the mounting arrangement has an attachment structure that is removably coupled by a snap fit connector frictionally retained and slidably supported on the support structure without the need for manually tightening and loosening the attachment structure to the support structure. 2. The system of claim 1, wherein the mounting arrangement and the catheter input module have a cooperating alignment and a cooperating locking structure. 3. The system of claim 1, wherein a plurality of catheter input modules are electrically and mechanically coupled to the mounting arrangement. 4. The system of claim 3, wherein the stacked catheter input modules and the mounting arrangement are interconnected to one another by a latching arrangement. 5. The system of claim 3, wherein the stacked catheter input modules and the mounting arrangement are electrically connected together by a harness arrangement. 6. The system of claim 1, wherein the central processing device is coupled to the mounting arrangement by a single cable. 7. The system of claim 1, wherein the mounting arrangement and the at least one catheter input module are interconnected to one another by a latching arrangement. 8. The system of claim 7, wherein the latching arrangement comprises a ramped tongue on the catheter input module and a slotted dome on the mounting arrangement. 9. The system of claim 1, wherein the mounting arrangement and the at least one catheter input module are electrically connected together by a harness having a ribbon cable. 10. The system of claim 1, wherein the at least one catheter input module has a removable cover formed with openings and secured to connectors for receiving input jacks of catheters. 11. The system of claim 10, wherein the removable cover is provided with an LED arrangement for verifying the electrical coupling of the mounting arrangement to the at least one catheter input module. 12. The system of claim 1, wherein the mounting arrangement includes a mounting platter rotatably attached to a top end of a vertical spindle, and an offset extension with an attachment bracket fixed to a lower end of the spindle. 13. The system of claim 12, wherein a top surface of the mounting platter comprises elements that are configured to align with and engag... | The General Electric Company | A61B005/02
US07840251 | B2 | 2010-11-23 | Device and method for registering a position sensor in an anatomical body | 2004-08-30 | An apparatus having an insertable portion for holding a position sensor is disclosed. The position sensor can transmit a signal indicative of its position with respect to a field generator. The insertable portion of the catheter has fiducial markings that are detectable by an imaging modality when the insertable portion is inserted into the anatomical body. The fiducial markings are asymmetrical about at least a first axis of the insertable portion. After the insertable portion has been inserted into the anatomical body, the fiducial markings can be detected to facilitate registration of the position sensor held in the insertable portion to the anatomical body. The apparatus also has a fixing mechanism for releasably fixing the insertable portion to the anatomical body. The non-symmetrical fiducial markings facilitate unambiguous registration of the insertable portion to the anatomical body. When the insertable portion is inserted into the anatomical body to a location of interest, the fixing mechanism substantially rigidly fixes the insertable portion of the catheter to a part of the anatomical body near the location of interest. | 1. An apparatus insertable into an anatomical body, said apparatus comprising: an insertable portion holding a position sensor that can transmit a signal indicative of its position in a frame of reference;fiducial markings on the insertable portion, said fiducial markings being detectable by an imaging modality when the insertable portion is inserted in the anatomical body and permitting the position and orientation of the insertable portion to be determined, said insertable portion substantially rigidly holding the position sensor at a known spatial position with respect to the fiducial markings;wherein the fiducial markings are non-symmetrical about at least a first axis of the insertable portion;wherein, after insertion in the anatomical body, the fiducial markings can be detected by the imaging modality to facilitate unambiguous registration to the anatomical body of the position sensor rigidly held in the insertable portion by correlating the fiducial markings detected by the imaging modality to the determined position of the position sensor in the frame of reference. 2. The apparatus as defined in claim 1 further comprising: a fixing mechanism for fixing the insertable portion to the anatomical body;wherein, when the insertable portion is inserted into the anatomical body to a location of interest, the fixing mechanism rigidly fixes the insertable portion to a part of the anatomical body near the location of interest; andwherein the fiducial markings can be detected while the insertable portion is fixed to the part of the anatomical body near the location of interest. 3. The apparatus as defined in claim 2 further comprising a catheter for inserting the insertable portion into the anatomical body; and wherein the fixing mechanism comprises a plurality of longitudinally extending fingers affixed to the insertable portion, each finger having a gripping element for gripping a surface of a passageway near the location of interest, said fixing mechanism having a collapsed configuration where the gripping element is near the catheter, and a deployed configuration where at least one gripping element grips the surface of the passageway; andwherein when the fingers are in the deployed configuration, fluid can flow around the catheter and through the passageway. 4. The apparatus as defined in claim 2 wherein the fixing mechanism is selected from the group consisting of a screw for securing the insertable portion to the anatomical body at the location of interest, and at least one barb for rigidly fixing the insertable portion of the catheter to the anatomical body at the location of interest. 5. The apparatus as defined in claim 2 further comprising a catheter for inserting the insertable portion into the anatomical body; wherein the fixing mechanism comprises an inflatable member which inflates when the insertable portion is at the location of interest to releasably fix the insertable portion of the catheter to the part of the anatomical body ne... | Philips Electronics Ltd | A61B005/05
US07826895 | B2 | 2010-11-02 | Organ rejection monitoring | 2006-07-28 | The invention is directed to techniques for monitoring organ rejection. An implanted device monitors the impedance of the transplanted organ. When the impedance measurements indicate that the organ is being rejected, the device provides early warning of rejection. | 1. An implantable medical device comprising: an impedance measuring circuit contained within the implantable medical device, wherein the impedance measuring circuit is configured to make measurements of the impedance across at least a portion of a transplanted organ in a patient;a processor contained within the implantable medical device, wherein the processor is configured to determine whether the patient is experiencing rejection of the transplanted organ as a function of the impedance measurements;a first electrode for placement in contact with the transplanted organ;a second electrode for placement in contact with the transplanted organ;a first electrode for placement in contact with a non-transplanted organ; anda second electrode for placement in contact with the non-transplanted organ, wherein the impedance measuring circuit is further configured to make measurements of the impedance of at least a portion of the transplanted organ with the first and second electrodes for placement in contact with the transplanted organ, and make reference measurements of the impedance of at least a portion of the non-transplanted organ with the first and second electrodes for placement in contact with the non-transplanted organ. 2. The device of claim 1, further comprising a communication module that generates a notification signal when the processor determines that the patient is experiencing rejection of the transplanted organ. 3. The device of claim 2, wherein the communication module comprises a telemetry module for communication with an external communication device. 4. The device of claim 2, wherein the communication module comprises an alert module that generates an audible alert. 5. The device of claim 1, further comprising: a current driver that delivers an excitation current to the transplanted organ via the first and second electrodes for placement in contact with the transplanted organ. 6. The device of claim 5, wherein the impedance measuring circuit measures the impedance by comparing the voltage difference between the first and second electrodes for placement in contact with the transplanted organ with the excitation current. 7. The device of claim 5, further comprising: a third electrode for placement in contact with the transplanted organ; anda fourth electrode for placement in contact with the transplanted organ, wherein the impedance measuring circuit is further configured to measure the impedance by comparing the voltage difference between the third and fourth electrodes during delivery of the excitation current. 8. The device of claim 1, wherein the processor is further configured to monitor variations in the sensed impedance on a periodic basis, and determine that the patient is experiencing rejection of the transplanted organ when the variations decrease over time. 9. An implantable medical device comprising: an impedance measuring circuit contained within the implantable medical device, wherein the impedance measuring circuit... | Medtronic, Inc. | A61B005/05
US20100324441 | A1 | 2010-12-23 | Brain-Related Chronic Pain Disorder Treatment Method and Apparatus | 2009-01-30 | A method for treating brain-related chronic pain disorders in human subjects includes assessing the brain function of a subject suffering from chronic pain, diagnosing a chronic pain-related abnormal brain condition, and mitigating the abnormal brain activity by applying an electrical stimulation signal to tissues corresponding to at least one area of abnormal brain activity. | 1. A method for treating brain-related chronic pain disorders in human subjects, the method including the steps of: assessing the brain function of a subject suffering from chronic pain;diagnosing a chronic pain-related abnormal brain condition;locating at least one area of abnormal brain activity associated with the abnormal brain condition; andmitigating the abnormal brain activity by applying an electrical stimulation signal to tissues corresponding to the at least one area of abnormal brain activity. 2. The method of claim 1 in which the electrical stimulation signal comprises waveforms configured to minimize tissue impedance. 3. The method of claim 2 in which the electrical stimulation signal comprises an AMPWM signal. 4. The method of claim 1 in which: at least one of the assessing and locating steps includes obtaining an electroencephalogram (EEG) of the subject's electrical brain activity;the diagnosing step includes obtaining a quantitative assessment (qEEG) of the subject's EEG and making a statistical comparison between the subject's qEEG and a database of qEEGs of either healthy normal individuals' brain functions or the brain functions of individuals suffering from the chronic pain-related abnormal brain function condition; andthe mitigating step includes generating and applying to the subject an electrical stimulation signal having at least one parameter configured to modulate at least one abnormal aspect of the subject's EEG, which corresponds to at least one statistically significant difference found in the statistical comparison of qEEGs. 5-7. (canceled) 8. The method of claim 1 in which: the diagnosing step includes diagnosing fibromyalgia by making a statistical comparison between the subject's qEEG and a database of qEEGs of either normal, healthy individuals or individuals suffering from fibromyalgia, andthe mitigating step includes generating and applying to the subject a neuromodulation signal. 9. The method of claim 8 in which at least one neuromodulation signal parameter to be configured for modulation is selected from the group of parameters consisting of carrier signal frequency, neuromodulation signal frequency, amplitude, waveform, duty cycle, application times, and phase. 10. The method of claim 8 in which the neuromodulation signal is applied using an apparatus comprising: a signal generation and interface module (121) including a microcontroller (106) configured to generate signal waveforms and coupled to a signal generator circuit (107) configured to transform the signal waveforms into desired AMPWM neuromodulation signals;a cap (3) configured to be worn on a subject's head and carrying first and second electrodes (1, 2) in respective cap locations that, when the cap is placed on a subject's head, position the electrodes adjacent a predetermined area of brain tissues (11) to be stimulated; andelectrical conductors (5) that provide signal paths between the signal generation and interface module (121) and the... | - | A61B005/04
US20100278391 | A1 | 2010-11-04 | Apparatus for behavior analysis and method thereof | 2006-10-12 | In the present invention, an apparatus for behavior analysis and method thereof is provided. In this apparatus, each behavior is analyzed and has its corresponding posture sequence through a triangulation-based method of triangulating the different triangle meshes. The two important posture features, the skeleton feature and the centroid context, are extracted and complementary to each other. The outstanding ability of posture classification can generate a set of key postures for coding a behavior sequence to a set of symbols. Then, based on the string representation, a novel string matching scheme is proposed to analyze different human behaviors even though they have different scaling changes. The proposed method of the present invention has been proved robust, accurate, and powerful especially in human behavior analysis. | 1. An apparatus for posture recognition comprising: a triangulation unit for dividing a posture of a body into a plurality of triangular meshes; anda recognition unit for forming a spanning tree correspond to the meshes to recognize the posture. 2. The apparatus for posture recognition according to claim 1 further comprises a background subtraction unit to extract and define a boundary of the body. 3. The apparatus for posture recognition according to claim 2, wherein the background subtraction unit is a video. 4. The apparatus for posture recognition according to claim 1, wherein the meshes are triangles. 5. The apparatus for posture recognition according to claim 1, wherein the recognition unit is achieved via a skeleton analysis or a centroid context analysis. 6. The apparatus for posture recognition according to claim 5, wherein the skeleton analysis is defined via a graph search scheme. 7. An apparatus for behavior analysis comprising: a clustering unit for key postures selection via merging a plurality of postures iteratively;a coding unit for translating all the input postures into a plurality of correspondent symbols according to the selected key postures; anda matching unit, which takes advantages of the coding unit, for unscrambling the correspondent symbols to distinguish a behavior. 8. The apparatus for behavior analysis according to claim 7, wherein the clustering unit is programmable. 9. The apparatus for behavior analysis according to claim 7, wherein the clustering unit is user-defined. 10. The apparatus for behavior analysis according to claim 7, wherein the postures are obtained from an apparatus for posture recognition comprising: a triangulation unit for dividing a body posture into a plurality of triangular meshes; anda recognition unit for forming a spanning tree from the meshes to recognize the posture. 11. The apparatus for behavior analysis according to claim 10, wherein the apparatus for posture recognition further comprises a background subtraction unit to extract and define boundaries of the body. 12. The apparatus for behavior analysis according to claim 10, wherein the meshes can be defined via a triangle-mesh algorithm. 13. The apparatus for behavior analysis according to claim 10, wherein the recognition unit is achieved via a skeleton analysis or a centroid context analysis. 14. The apparatus for behavior analysis according to claim 13, wherein the skeleton analysis is defined via a graph search scheme. 15. The apparatus for behavior analysis according to claim 13, wherein the centroid context is formed by labeling each mesh with a number. 16. The apparatus for behavior analysis according to claim 7, wherein the correspondent symbols are unscrambling via a string matching method. 17. A method for posture recognition, comprising the steps of: triangulating a posture of a body into a plurality of triangular meshes;forming a skeleton analysis and a centroid context analysis correspond to the triangula... | - | G06T007/00 A61B005/00
US08708907 | B2 | 2014-04-29 | Method and apparatus for determining one or more blood parameters from analog electrical signals | 2010-05-05 | Embodiments of the present invention relate to a system and method for in vivo measurement of blood parameters by processing analog electrical signals from a plurality of photodetectors. In some embodiments, it is possible to determine one or more blood parameters according to (i) a first electrical signal from a first detector and (ii) a second electrical signal from a second photodetector. A difference analog electrical signal is generated, indicative of a difference between the light response signal at the first location and light response signal at the second location, is generated. One or more blood parameters may be detected according to the difference analog electrical signal. | 1. A method for obtaining physiological information from a subject in accordance with a light field whose intensity varies according to location, the method comprising: a) applying partially or entirely coherent light to a target region of the subject to induce a light response signal from the illuminated target region which contributes to a location-dependent light field LF(x, y, z);b) detecting the light field LF(x1, y1, z1) prevailing at a first location (x1, y1, z1) to generate a first analog electrical signal A1(t), the first analog electrical signal A1(t) including a dynamic light scattering (DLS) component DLS(A1(t)) that is indicative of a first DLS measurement;c) detecting the light field signal LF(x2, y2, z2) prevailing at a second location (x2, y2, z2) to generate a second analog electrical signal A2(t), the second analog electrical signal A2(t) including a DLS component DLS(A2(t)) that is indicative of a second DLS measurement and that is uncorrelated with the DLS component of the first analog signal A1(t);d) processing, by analog electrical circuitry, the first and second analog electrical signals to generate a processed analog electrical signal processed(t) from a difference between the first and second analog electrical signals, wherein: (i) a first DLS contribution ratioDLS⁡(A1⁡(t))A1⁡(t)is defined as the ratio between a magnitude ¦DLS(A1(t))¦ of the DLS component of the first analog electrical signal and a magnitude ¦A1(t)¦ of the first electrical analog signal; (ii) a second DLS contribution ratioDLS⁡(A2⁡(t))A2⁡(t)is defined as the ratio between a magnitude ¦DLS(A2(t))¦ of the DLS component of the second analog electrical signal and a magnitude ¦A2(t)¦ of the second analog electrical signal; (iii) processed analog electrical signal processed(t) includes a DLS component that is indicative of a combination of the first and second dynamic scattering measurements;(iv) a processed signal DLS contribution ratioDLS⁡(processed⁡(t))processed⁡(t)is defined as a ratio between a (i) magnitude ¦DLS(processed(t))¦ of the DLS component of the processed analog electrical signal and (ii) a magnitude ¦processed(t)¦ of the processed analog signal; and (v) the detecting and the processing are performed so that a ratio between: i) the processed signal DLS contribution ratioDLS⁡(processed⁡(t))processed⁡(t);ii) at least one of the first DLS contribution ratioDLS⁡(A1⁡(t))A1⁡(t)and the second DLS contribution ratioDLS⁡(A2⁡(t))A2⁡(t)is at least 10. 2. The method of claim 1 further comprising: e) digitizing the processed analog signal to generate a digital signal indicative of a combination of the first and second dynamic scattering measurements;f) processing the indicative digital signal to compute at least one physiological parameter. 3. The method of claim 2 wherein the computing includes computing at least one of a blood velocity, a blood rheological parameter, a blood pressure, a heartbeat, a blood oxygen concentration, and ... | Elfi-Tech | A61B005/02 A61B005/14551 A61B005/145 A61B008/06 A61B008/12 A61B005/026
US08540633 | B2 | 2013-09-24 | Identifier circuits for generating unique identifiable indicators and techniques for producing same | 2009-08-13 | The present invention provides for safe and reliable electronic circuitry that can be employed in ingestible compositions. The ingestible circuitry of the invention includes a solid support; a conductive element; and an electronic component. Each of the support, conductive element and electronic component are fabricated from an ingestible material. The ingestible circuitry finds use in a variety of different applications, including as components of ingestible identifiers, such as may be found in ingestible event markers, e.g., pharma-informatics enabled pharmaceutical compositions. | 1. A method of producing an identifier, the method comprising: laminating a non-conducting element between first and second dissimilar materials to create a laminate component, wherein the first and second dissimilar materials generate a voltage potential when in contact with a conducting fluid;removing a portion of the first and second dissimilar materials from the perimeter of the laminate component to expose the non-conducting element and to create a skirt; andsecuring a circuit within an aperture formed in the laminate component, wherein the circuit is electrically coupled to each of the first and second dissimilar materials, and wherein the circuit is operable by the voltage potential generated by the first and second dissimilar materials. 2. The method of claim 1, further comprising selecting the first and second dissimilar materials to create a voltaic cell through oxidation and reduction reactions when the first and second dissimilar materials are in contact with the conducting fluid. 3. The method of claim 2, further comprising configuring the circuit to generate a detectable signal that identifies unique information associated with the identifier. 4. The method of claim 2, wherein the conducting fluid is a physiological fluid. 5. The method of claim 1, further comprising securing the identifier to a pharmaceutical product to produce a tagged product such that activation of the circuit is an indication that the pharmaceutical product of the tagged product is in contact with physiological fluid. 6. The method of claim 5, further comprising sealing the tagged product with a coating to isolate the tagged product from the physiological fluid for a predefined period of time until the coating is dissolved by the physiological fluid to allow the tagged product to reach a target site at which time the tagged product is exposed to the physiological fluid. 7. The method of claim 1, further comprising laminating a current collector element between the non-conducting element and at least one of the first and second dissimilar materials. 8. The method of claim 1, further comprising forming the aperture through the laminate component to receive the circuit within the aperture. 9. The method of claim 1, further comprising fixing the circuit within the aperture with a conductive adhesive. 10. The method of claim 1, further comprising removing a portion of the laminate component encompassing the circuit to produce the identifier. | Proteus Digital Health, Inc. | A61B005/07
US08668652 | B2 | 2014-03-11 | Airway bypass site selection and treatment planning | 2010-08-13 | This invention relates to systems and methods for site selection and placement of extra-anatomic passages altering gaseous flow in a diseased lung. | 1. A method of assessing a plurality of treatment sites in airways of a lung of a patient, the method comprising: identifying the plurality of sites within the airways of the lung;determining at least one anatomic characteristic of each site;assessing a diseased region to determine a degree of disease of the diseased region;automatically calculating on a computer a score for each site based on: the degree of disease of the diseased region in the lung, at least one anatomic characteristic of each site and a proximity of the site to the diseased region; anddisplaying each score in association with each site in graphical form. 2. The method of claim 1, where identifying the plurality of sites comprises selecting at least one site per each segmental airway of the bronchi. 3. The method of claim 1, further comprises displaying information regarding at least one anatomic characteristic of the site when displaying the score of the site. 4. The method of claim 1, where determining at least one anatomic characteristic of each site comprises identifying a location of the blood vessels within a distance from the site. 5. The method of claim 1, where calculating the score for each site comprises adjusting the score proportionally to the degree of the diseased region. 6. The method of claim 5, where calculating the score for each site further comprises increasing the score as the proximity of the site to one of the diseased regions decreases. 7. The method of claim 1, where determining at least one anatomic characteristic of each site comprises determining a size of airway diameter of each site. | Broncus Medical Inc. | A61B005/08
US08886280 | B2 | 2014-11-11 | Nerve monitoring device | 2010-08-05 | A nerve monitoring device, including a cannula, a sensor for monitoring a nerve, and an optional support element, can be inserted into an anatomic space. The cannula, sensor and/or support element can automatically conform to and match the geometry of the anatomic space, which can enhance desired contact between the sensor and anatomic features, such as muscles, nerves or tissue in the space in an atraumatic manner. The sensor, cannula and/or support element can include one or more strips or other elements that convert from a retracted mode to an expanded mode in which the strips or elements expand, enlarge or otherwise move to match the geometry of the space and place the sensors in precise monitoring proximity relative to the target muscles, nerves or tissues. An exemplary application of the device is in the field of laryngeal monitoring, however, the device is well suited for other applications. | 1. A nerve monitoring device comprising: a support element joined with a plurality of sensors and arranged to join a tube, the support element including a plurality of strips oriented generally parallel to the longitudinal axis of a tube, each strip having a strip sensor, the plurality of strips joined with a first connection element that at least partially surrounds a tube, the strips moveable between a retracted mode, in which the strips are positioned adjacent a tube, and an extended mode, in which at least a portion of the strips extend outwardly from a tube so that the strips conform to the anatomical geometry of at least a portion of a space of a subject in an atraumatic manner;wherein two or more of the strips are placed immediately adjacent to each other so that the strips provide sufficient circumferential coverage such that rotational orientation does not affect the monitoring of at least one nerve; andwherein at least one of the sensors is adapted to measure the activity of at least one nerve when placed in monitoring proximity so that a health care provider can be provided with information concerning the location of a nerve. 2. The device of claim 1 comprising: an endotracheal tube configured for insertion into a laryngeal space of a subject, the endotracheal tube including a longitudinal axis;the sensors joined with the endotracheal tube at a predetermined location corresponding to at least one of a target laryngeal muscle and a laryngeal nerve when the endotracheal tube is positioned through a glottic opening of the subject; andan output element in communication with at least one of the sensor, the output element providing output indicative of measured activity of at least one of the target laryngeal muscle and the laryngeal nerve when a probe is positioned adjacent or on the laryngeal nerve;wherein the support element includes a second connection element joined with the plurality of strips distal from the first connection element, wherein each of the plurality of strips are located between and joined with both the first connection element and the second connection element, wherein the second connection element is moveable toward the first connection element. 3. The device of claim 2 wherein each of the plurality of strips include a first end joined with the first connection element, a second end joined with the second connection element, and a central portion located between the first end and the second end. 4. The device of claim 3 wherein the central portion moves away from the endotracheal tube a greater distance then the first end moves away from the endotracheal tube when the strips are in the extended mode. 5. The device of claim 1 wherein the plurality of strips are flexible, and wherein the plurality of strips include a portion that bends outwardly to a generally convex configuration when the strips are moved from the retracted mode to the extended mode. 6. The device of claim 5 wherein the strips are formed in an accu... | The Magstim Company Limited | A61B005/04 A61B005/00 A61N001/36 A61N001/05 A61B005/0488 A61B018/18 A61B019/00 A61M016/04 A61B001/267
US08467859 | B2 | 2013-06-18 | Method and device for deriving and evaluating cardiovascular information from curves of the cardiac current, in particular for applications in telemedicine | 2007-09-04 | The invention relates to a method and a device for deriving and evaluating cardiovascular information from curves of the cardiac current, in particular for applications in telemedicine. According to the invention, for data reduction purposes, a mean value and a standard deviation across the whole recording time are initially determined from a recorded ECG from the high-resolution signal, after which the RR intervals are assessed by means of an identification algorithm and only the results obtained are transmitted and/or recorded, for the purpose of data reduction. | 1. Method for deriving and evaluating cardiovascular information from electrocardiograms to determine sympatho-vagal balance, especially for telemedical applications, comprising: using a signal processor for the reduction of data, a mean value and a standard deviation are initially determined as signal-averaged ECG from the high-resolution signal of a recorded ECG over the total recording time and the R-R intervals are subsequently evaluated by means of a detection algorithm, and only the received results are transmitted and/or recorded for the purpose of the data reduction for which purpose from an FFT spectrum of the determined heart period durations area integrals F are determined for the defined LF and HF frequency ranges of 0.04 to 0.15 Hz and 0.15 to 0.4 Hz and are related to the mean area-normal-values, so that spectral indexes for the evaluation are obtained as follows:MLF=F0,04⁢…0,15⁢⁢Hzmean⁢-⁢area⁢-⁢normal⁢-⁢value⁢100⁡[%]MHF=F0,15⁢…0,4⁢⁢Hzmean⁢-⁢area⁢-⁢normal⁢-⁢value⁢100⁡[%]wherein furthera weighting of the square sTH2 of the standard deviation, the absolute heart rate variability, with the square of the mean heart rate fH is performed[STH]2·[fH_]2,so that the cardiovascular state can be evaluated with respect to the influence of the autonomic nervous system with its components sympathicus and parasympathicus,anddetermining by weightings of the indexes MLF and MHF with the square of the mean heart rate fH how the sympatho-vagal balance B varies, so that the sympathetic activity components occurring in the LF frequency range become quantifiable as compared to the parasympathetic activity components. 2. Method according to claim 1, characterized in thata quantification is performed ;with the weighted indexes as follows; normalized dimension for sympathetic activity, stress arid strain:MLF·[fH_70⁢⁢min-1]2=MLF,normalized dimension for parasympathetic activity, relaxation and recovery:MHf,sympatho-vagal balance:sympatho-vagal⁢⁢⁢balance=MLF*MHF=MLFMHF·[fH_70⁢⁢min-1]2. 3. Method according to claim 1, characterized in thatbeside the R-R. determination it is determined by a frequency analysis of the derived ECG total time function in the range of 0 to about 2 Hz whether two dominant frequency lines f1,2 in the from of a frequency doubling occur at a same amplitude Af1;2, so as to define a (stability) criterion Kμ for the actually non-linear cardiovascular system, derived from the ECG time function, as follows:Kfrequency=f1f2&lt;2Kamplitude=Af⁢⁢2Af⁢⁢1&lt;1. | Telozo GmbH | A61B005/0452
US09042964 | B2 | 2015-05-26 | System and method for fiducial deployment via slotted needle | 2010-04-21 | Embodiments include a fiducial deployment system and method for use thereof. A fiducial may include one or more protuberances configured to engage one or more slots in a needle of the system. The needle may be configured to deliver a plurality of fiducials to a target location in serial fashion, one at a time. In certain embodiments, echogenic placement of fiducials may present certain advantages. | 1. A fiducial deployment system comprising: a needle including a tubular cannula body defining a needle lumen disposed through at least a lengthwise portion of the cannula body and a distal needle end region, the distal end region comprising a distal needle end opening at a distal end of the needle lumen; andat least one generally longitudinal needle slot extending radially through at least a thickness portion of the cannula body and open to the needle lumen, where at least one detent structure extends across and within the needle slot near the distal needle end opening;a plurality of fiducials disposed in the lumen, each comprisinga generally columnar body including a central fiducial portion slidably disposed in the needle lumen; andat least one side protuberance projecting into the needle slot so as to engage a proximal surface of the at least one detent; anda stylet extending through a portion of the needle lumen and configured to advance each of the plurality of fiducials past the at least one detent and out of the distal needle end opening in a controlled, one at a time, serial manner. 2. The fiducial deployment system of claim 1, where the at least one protuberance of at least one of the plurality of fiducials comprises a plurality of protuberances. 3. The fiducial deployment system of claim 2, where more than one of the plurality of protuberances are generally longitudinally aligned with each other. 4. The fiducial deployment system of claim 1, where each fiducial includes a proximal end and a distal fiducial end, where the proximal end of at least one of the fiducials is immediately adjacent the distal end of another of the fiducials, and where a longitudinal length of each fiducial is greater than a longitudinal length of its protuberance. 5. The fiducial deployment system of claim 1, where the at least one generally longitudinal needle slot comprises a plurality of needle slots including a first needle slot and a second needle slot. 6. The fiducial deployment system of claim 5, where the at least one protuberance of at least one of the plurality of fiducials comprises a plurality of protuberances including a first protuberance and a second protuberance, and where the first protuberance at least partially occupies the first slot and the second protuberance at least partially occupies the second slot. 7. The fiducial deployment system of claim 1, where a lengthwise gap between the protuberances of the fiducials is at least as long as the at least one detent. 8. The fiducial deployment system of claim 1, where the central fiducial portion comprises a generally non-cylindrical portion. 9. The fiducial deployment system of claim 1, where at least one of the plurality of fiducials comprises at least one echogenic surface. 10. The fiducial deployment system of claim 1, further comprising being at least partially disposed through a working channel of an endoscope. 11. A method of placing a fiducial into a patient body, the method com... | Cook Medical Technologies LLC | A61B005/05 A61B017/34 A61B019/00 A61M037/00 A61B001/018 A61B017/00 A61B017/06 A61B017/04 A61N005/10
US08469965 | B2 | 2013-06-25 | Tool for detecting planes of a bone and assigned data processing method | 2010-05-06 | The present application relates to a tool for detecting planes which are defined by anatomical landmarks, comprising a first and a second level abutment area for placing onto a body structure, wherein the first and second abutment areas have a predetermined position relative to each other, and comprising a tool marker device which has a stationary position relative to the first and second abutment areas. | 1. A tool for detecting planes which are defined by anatomical landmarks, comprising: a first and a second level abutment area for placing onto a body structure, wherein the first and second abutment areas have a predetermined position relative to each other;a tool marker device which has a stationary position relative to the first and second abutment areas; anda pressure sensor which detects whether a pressure is being exerted in predetermined partial regions of the first and second abutment areas which lie on the mutually facing sides of the first and second abutment areas. 2. A data processing method for use with a tool for detecting planes which are defined by anatomical landmarks, the tool including a first and a second level abutment area for placing onto a body structure, wherein the first and second abutment areas have a predetermined position relative to each other, and a tool marker device which has a stationary position relative to the first and second abutment areas, the data processing method determining the variable position of the first and second abutment areas of the tool relative to the body structure, wherein the first abutment area abuts a first side of the body structure, and the second abutment area abuts a second side of the body structure, wherein the following data is provided:tool marker position data which describes the position of the tool marker device;body structure marker position data which describes the position of a body structure marker device which is attached to the body structure; andshape data which describes the position of a shape representative of the body structure,and wherein the following data processing steps are performed:determining whether the position of the shape representative fulfills a position criterion which relates to the position of the shape representative relative to the first and/or second abutment area;determining the tool marker position data and the body structure marker position data which obtain when the position criterion is determined to be fulfilled. 3. The data processing method according to claim 2, wherein the relative position between the shape representative and the first and/or second abutment area is determined on the basis of the determined tool marker position data and the determined body structure marker position data and on the basis of the shape data and tool data which describes the relative position of the first and/or second abutment area relative to the tool marker device, and/or the relative position between at least a part of the body structure and the first and/or second abutment area is determined on the basis of the determined tool marker position data, the determined body structure marker position data, the tool data and the body structure data which describes the relative position of the body structure marker device relative to at least the part of the body structure. 4. The data processing method according to claim 3, wherein contact data is also provi... | Brainlab AG | A61B017/58 A61B017/60 A61F002/00
US08463351 | B2 | 2013-06-11 | Electrochemical analyte sensor | 2010-08-06 | An electrochemical analyte sensor formed using conductive traces on a substrate can be used for determining and/or monitoring a level of analyte in in vitro or in vivo analyte-containing fluids. For example, an implantable sensor may be used for the continuous or automatic monitoring of a level of an analyte, such as glucose, lactate, or oxygen, in a patient. The electrochemical analyte sensor includes a substrate and conductive material disposed on the substrate, the conductive material forming a working electrode. In some sensors, the conductive material is disposed in recessed channels formed in a surface of the sensor. An electron transfer agent and/or catalyst may be provided to facilitate the electrolysis of the analyte or of a second compound whose level depends on the level of the analyte. A potential is formed between the working electrode and a reference electrode or counter/reference electrode and the resulting current is a function of the concentration of the analyte in the body fluid. | 1. An in vivo glucose sensor, comprising: a substrate having a narrow distal region for partial implantation through a skin surface and a wider proximal region, the substrate including at least a first surface and a second surface;at least one working electrode disposed on the narrow distal region of the substrate and in fluid contact with interstitial fluid under the skin surface, the at least one working electrode formed on the first surface of the substrate;at least one reference electrode disposed on the narrow distal region of the substrate, the at least one reference electrode formed on the first surface of the substrate;at least one counter electrode disposed on the narrow distal region of the substrate, the at least one counter electrode formed on the second surface of the substrate;a plurality of first conductive contacts disposed on the wider proximal region of the substrate and on the first surface of the substrate, wherein a first one of the plurality of first conductive contacts is in electrical communication with the working electrode, a second one of the plurality of first conductive contacts is in electrical communication with the reference electrode, and a third one of the plurality of first conductive contacts is in electrical communication with the counter electrode;at least one via formed through the substrate that electrically connects the counter electrode to the third one of the plurality of first conductive contacts;a sensing layer proximate to the at least one working electrode; anda mass transport limiting layer proximate to the at least one working electrode;wherein the flexibility of the sensor is varied along a continuum of the substrate. 2. The glucose sensor of claim 1, wherein the varied flexibility of the sensor is based at least in part on the composition of the substrate. 3. The glucose sensor of claim 1, wherein the varied flexibility of the sensor is based at least in part on the thickness of the substrate. 4. The glucose sensor of claim 1, wherein the varied flexibility of the sensor is based at least in part on the composition and the thickness of the substrate. 5. The glucose sensor of claim 1, wherein the sensor is a planar sensor. 6. The glucose sensor of claim 1, wherein the sensor is a wire sensor. 7. The glucose sensor of claim 1, wherein the substrate comprises a non-conducting plastic or polymeric material. 8. The glucose sensor of claim 7, wherein the non-conducting plastic or polymeric material is a thermoplastic material. 9. An in vivo glucose sensor, comprising: a substrate having a narrow distal region for partial implantation through a skin surface and a wider proximal region, the substrate including at least a first surface and a second surface;at least one working electrode disposed on the narrow distal region of the substrate and in fluid contact with interstitial fluid under the skin surface, the at least one working electrode formed on the first surface of the substrate;at least on... | Abbott Diabetes Care Inc. | A61B005/04 A61B005/00
US07927277 | B2 | 2011-04-19 | Methods for non-invasively monitoring health | 2010-08-23 | The time between heartbeats is measured over a series of such heartbeats. The time interval between two successive events is calculated and stored as a first array. The time difference between adjacent heartbeat intervals is also calculated from the first array and recorded as a differential array. The differential array is subjected to frequency analysis. First the differential array data is linearly interpolated to increase the number of data samples. The interpolated data is then subjected to a fast fourier transform (FFT) yielding a power spectrum. Characteristic frequency ranges are then integrated and the resulting frequency domain spectrum(s) are analyzed for dominant frequency characteristics. | 1. A method of monitoring health by physiological monitoring of body motion using a non-invasive sensor comprising: creating a first array of data using a processor based on discretely recorded body motion events detected by the sensor in which each element of the first array is representative of a time when a body motion event took place;creating a second array of data using a processor in which each element of the second array is an interval representative of the difference between successive elements of the first array;creating a third array of data using a processor in which each element of the third array is a delta interval representative of the difference between successive elements of the second array;performing a fast Fourier transform (FFT) on the third array using a processor to obtain power spectrum data representative of the third array;integrating the power spectrum data over frequency ranges of interest using a processor to obtain discrete power values for said frequency ranges of interest; anddisplaying the power spectrum data and discrete power values using a display device to monitor health. 2. A method of monitoring health by physiological monitoring of body motion using a non-invasive sensor to assist in cardiac evaluation comprising: creating a first array of heart vibrations using a processor based on discretely recorded heartbeats detected by the sensor in which each element of the first array is representative of a time when a heartbeat took place;creating a heart period array using a processor in which each element is a heart period interval representative of the difference between successive elements of the first array of heart vibrations;creating a delta heart period interval array using a processor in which each element is a delta heart period interval representative of the difference between successive elements of the heart period interval array;performing a fast Fourier transform (FFT) on the delta heart period interval array using a processor to obtain power spectrum data representative of the delta heart period interval array;integrating the power spectrum data over one or more frequency ranges of interest using a processor to obtain discrete power values for said one or more frequency ranges of interest; anddisplaying the power spectrum data and discrete power values using a display device to evaluate cardiac health. 3. The method of claim 2 wherein the frequency ranges of interest include a low frequency (LF) range, a high frequency (HF) range, and a very high frequency (VHF) range. 4. The method of claim 3 further comprising: calculating and displaying a total power, TP, value that is the sum of the LF, HF, and VHF power values. 5. The method of claim 3 further comprising: calculating and displaying a power ratio value that is equal to LF/HF. 6. The method of claim 4 further comprising: calculating and displaying a normalized LF power value that is equal to LF/TP. 7. The method of claim 4 further comprisin... | The Johns Hopkins University | A61B008/00
US08311601 | B2 | 2012-11-13 | Reflectance and/or transmissive pulse oximeter | 2009-06-30 | According to various embodiments, a medical sensor assembly may be configured to switch between transmission and reflectance mode. Such sensors may include multiple optical sensing components that may be activated or silent, depending on the mode in use. A practitioner may switch between modes based on the particular situation of the patient or based on the signal quality. | 1. A sensor comprising: a sensor body adapted to be applied to a patient's tissue;a first emitter and a second emitter disposed on the sensor body;a detector disposed on the sensor body, wherein the first emitter and the detector are capable of operating in transmission mode and the second emitter and the detector are capable of operating in reflectance mode;a switch disposed on the sensor body and coupled to the first emitter and the second emitter; anda processor-based controller disposed on the sensor body and configured to receive an input from the detector, wherein the controller is configured to control the switch to alternate between activation of the first emitter and the second emitter based on the input. 2. The sensor, as set forth in claim 1, wherein the sensor comprises at least one of a pulse oximetry sensor or a sensor for measuring a water fraction. 3. The sensor, as set forth in claim 1, wherein first emitter or the second emitter comprises at least one light emitting diode and wherein the detector comprises at least one photodetector. 4. The sensor, as set forth in claim 1, wherein the first emitter and the detector are spaced apart about 20 mm to about 25 mm or wherein the second emitter and the detector are spaced apart about 8 mm to about 14 mm. 5. The sensor, as set forth in claim 1, wherein the controller is configured to receive a signal from the detector and detect a signal artifact in the signal. 6. The sensor, as set forth in claim 5, wherein the controller is configured to switch from transmission mode or reflectance mode if the controller detects the signal artifact. 7. The sensor, as set forth in claim 5, wherein the controller is configured to arbitrate between transmission mode and reflectance mode to mitigate the signal artifact in the signal. 8. The sensor, as set forth in claim 5, wherein the controller is configured to switch from reflectance mode to transmission mode in the presence of a signal artifact comprising a η-artifact. 9. The sensor, as set forth in claim 5, wherein the controller is configured to switch between reflectance mode and transmission mode during the time an α-artifact or a ΔPtiss-artifact is detected. 10. The sensor, as set forth in claim 1, comprising an encoder, wherein the encoder comprises stored data related to the sensor. 11. A system comprising: a sensor comprising: a sensor body adapted to be applied to a patient's tissue;a first emitter and a second emitter disposed on the sensor body;a detector disposed on the sensor body, wherein the first emitter and the detector are capable of operating in transmission mode and the second emitter and the detector are capable of operating in reflectance mode, and wherein the sensor is capable of alternating between transmission mode and reflectance mode; anda monitor comprising a processor configured to: detect a signal artifact in a signal from the sensor;determine a characteristic of the signal artifact; andcontrol the activation of ... | Nellcor Puritan Bennett LLC | A61B005/1455
US08249684 | B2 | 2012-08-21 | Calibration techniques for a continuous analyte sensor | 2010-09-01 | Disclosed herein are systems and methods for calibrating a continuous analyte sensor, such as a continuous glucose sensor. One such system utilizes one or more electrodes to measure an additional analyte. Such measurements may provide a baseline or sensitivity measurement for use in calibrating the sensor. Furthermore, baseline and/or sensitivity measurements may be used to trigger events such as digital filtering of data or suspending display of data. | 1. A method for continuous glucose sensing, the method comprising: generating a glucose signal associated with a glucose concentration in a host using a continuous glucose sensor;measuring a sensitivity change of the continuous glucose sensor over a time period; anddetermining using electronic circuitry, a stability of the continuous glucose sensor based on the measured sensitivity change, wherein the continuous glucose sensor comprises an electrode, and wherein measuring the sensitivity change comprises switching, cycling or pulsing a voltage applied to the electrode and evaluating a response to the applied voltage. 2. The method of claim 1, wherein the continuous glucose sensor comprises a membrane system, and wherein the measured sensitivity change is representative of a change in glucose transport through the membrane system. 3. The method of claim 1, wherein measuring the sensitivity change comprises measuring an analyte other than glucose. 4. The method of claim 1, wherein measuring the sensitivity change comprises measuring a change of oxygen transport through a membrane system of the continuous glucose sensor. 5. The method of claim 1, wherein measuring the sensitivity change comprises monitoring the glucose signal over a time period. 6. The method of claim 1, wherein measuring the sensitivity change is performed in vivo. 7. The method of claim 1, wherein determining the stability further comprises detecting oxygen deprivation. 8. The method of claim 7, wherein detecting oxygen deprivation comprises examining the glucose signal for oxygen deprivation-induced artifacts. 9. The method of claim 1, further comprising calibrating the glucose signal using the measured sensitivity change. 10. The method of claim 1, further comprising controlling a display of the glucose signal responsive to the stability determination. 11. The method of claim 1, further comprising suspending display of the glucose signal responsive to the stability determination. 12. A system for continuous glucose sensing, the system comprising: a continuous glucose sensor configured to generate a glucose signal associated with a glucose concentration in a host; andelectronic circuitr configured to determine a stability of the continuous glucose sensor by measuring a sensitivity change over a time period, wherein the continuous glucose sensor comprises an electrode, and wherein the electronic circuitry is configured measure the sensitivity change based on a response to a switching, cycling or pulsing of a voltage applied to the electrode. 13. The system of claim 12, wherein the continuous glucose sensor comprises a membrane system, and wherein the measured sensitivity change is representative of a change in glucose transport through the membrane system. 14. The system of claim 12, wherein the electronic circuitry is configured to calibrate the glucose signal using the measured sensitivity change. 15. The system of claim 12, wherein the electronic circuitry is conf... | DexCom, Inc. | A61B005/05
US07990691 | B2 | 2011-08-02 | Mobile computer workstation | 2010-08-06 | A mobile workstation for use with a computer network. The mobile workstation may include a medical monitoring device, a radio transceiver in communication with the medical monitoring device operable for receiving and sending data to the computer network, a display screen, and a wheeled chassis for mounting the medical monitoring device, the radio transceiver and the display screen. | 1. A mobile computer workstation, comprising: a wheeled chassis including a beam having a substantially vertical first member and a second member slidably connected to the first member;a substantially unencumbered work surface supported by the second member at a vertical height;a display screen located above and off the work surface;a computing device located off of the work surface;an input device tray located proximate to and off of the work surface;a power unit supported by the wheeled chassis; andwherein the vertical height of the work surface changes when the second member moves relative to the first member. 2. The mobile computer workstation of claim 1, wherein the input device tray positions an input device in front of and below the work surface. 3. The mobile computer workstation of claim 1, wherein a power cable is routed within said vertical beam. 4. The mobile computer workstation of claim 1, wherein the power unit is disposed below the beam. 5. The mobile computer workstation of claim 1, further comprising: a video camera, wherein the video camera is mounted on the display screen; anda video-conferencing system. 6. The mobile computer workstation of claim 1, further comprising: a second display screen located off of the work surface. 7. The mobile computer workstation of claim 1, further comprising a patient vital signs capture device comprising: a sensor device; anda control device, wherein the control device controls information from the sensor device. 8. The mobile computer workstation of claim 7, wherein the sensor device comprises one of a blood pressure cuff, a thermometry sensor, and a pulse oximetry sensor. 9. The mobile computer workstation according to claim 1, further comprising: a radio transceiver; anda medical monitoring device; andwherein the radio transceiver is in communication with the medical monitoring device. 10. The mobile computer workstation of claim 1, further comprising means for controlling movement of the second member of the beam relative to the first member. 11. The mobile computer workstation of claim 10, wherein the means for controlling movement of the second member of the beam relative to the first member comprises one of a rack and pinion, a cable and pulley, a ratchet, a ball screw, a removable pin and holes arrangement, and a gas-spring. 12. The mobile computer workstation of claim 1, wherein the power unit is located at least partially below an upper surface of the wheels of the chassis. 13. The mobile computer workstation of claim 1, wherein the computing device is located within a housing of the display screen. 14. The mobile computer workstation of claim 1, wherein the work surface is defined by an upper surface of a substantially horizontal tray connected to the second member. 15. The mobile computer workstation of claim 14, wherein the computing device is disposed below the upper surface and at least partially within the tray. 16. A mobile computer workstation, comprising: a ch... | InterMetro Industries Corporation | G06F001/16
US09314180 | B2 | 2016-04-19 | Heart electrophysiological signal analysis system | 2009-12-01 | A system automatically detects and measures ST deviation of a heart wave ECG signal in the presence of noise and accommodates baseline variation of the signal and other artefacts. A system identifies a particular point in an electrophysiological signal representing heart electrical activity using an interface for receiving an electrical signal waveform comprising an R-wave and including an ST segment portion associated with heart electrical activity of a patient over a heart beat cycle. A signal processor processes data representing the electrical signal waveform by identifying an S point and T point in the electrical signal waveform and determining a first candidate J point in the electrical signal waveform having substantially a maximum distance from a line between the identified S and T points, the distance being measured perpendicularly to the line. | 1. A system for identifying a particular point in an electrophysiological signal representing heart electrical activity, comprising: an interface for receiving an electrical signal waveform comprising an R-wave and including an ST segment portion associated with heart electrical activity of a patient over a heart beat cycle; anda signal processor for processing data representing said electrical signal waveform by (a) identifying an S point in said electrical signal waveform, wherein the S point is an approximate fiducial point at a local minimum position following an R peak in the waveform and wherein the R peak is an approximate maximum position in the R-wave;(b) identifying a T point in said electrical signal waveform, wherein the T point is an approximate fiducial point at a local maximum position following an S point;(c) defining a line between the identified S and T points, and determining a maximum distance to the line between the identified S and T points, said distance being measured substantially perpendicular to said line;(d) determining a first candidate J point in said electrical signal waveform as having the maximum perpendicular distance from the line, wherein the J point is subsequently used in deviation detection of the electrophysiological signal; and(e) computing an ST segment variation of the ST segment value from the baseline of said electrical signal, using said determined J point to identify the start of the ST segment; and(f) generating an alert message to a worker in response to the ST segment variation exceeding a predetermined threshold value. 2. The system according to claim 1, wherein said signal processor processes data representing said electrical signal waveform by (a) identifying an R point in said electrical signal waveform, wherein the R point is an approximate fiducial point at the R peak and(b) defining a second line between the identified R and T points and determining a second maximum distance to the second line, said distance being measured substantially perpendicularly to said second line and(c) determining a second candidate J point in said electrical signal waveform as having the maximum perpendicular distance from the second line. 3. The system according to claim 2, wherein said signal processor selects the closest of the first and second candidate J points to a point substantially 40 milliseconds after the R point, the selected J point being used for further analysis. 4. The system according to claim 1, wherein said signal processor determines J points for individual heart cycles of a plurality of sequential cycles andsaid signal processor uses the identified J points in computing corresponding ST segment values for individual heart cycles of the plurality of sequential cycles. 5. The system according to claim 1, wherein said signal processor determines said electrical signal waveform baseline by (a) identifying a T peak point of a first heart cycle in said electrical signal waveform,(b) identifying... | Siemens Medical Solutions USA, Inc. | A61B005/0456 A61B005/0452 A61B005/00
US09439735 | B2 | 2016-09-13 | MRI-guided interventional systems that can track and generate dynamic visualizations of flexible intrabody devices in near real time | 2010-06-08 | MRI guided cardiac interventional systems are configured to generate dynamic (interactive) visualizations of patient anatomy and medical devices during an MRI-guided procedure and may also include at least one user selectable 3-D volumetric (tissue characterization) map of target anatomy, e.g., a defined portion of the heart. | 1. A circuit for an MRI-guided cardiac interventional system, comprising: at least one processor configured to: (a) generate near real time (RT) MRI images of at least a portion of a heart of a patient using relevant anatomical scan planes associated with a 3-D MRI image space having a coordinate system; (b) identify coordinates in the 3-D MRI image space coordinate system associated with a location of at least a distal end portion of at least one flexible intrabody catheter in the 3-D MRI image space using tracking coil data from a plurality of longitudinally spaced apart tracking coils on the distal end portion of the at least one flexible intrabody catheter; and (c) render interactive near RT visualizations with MR image data and at least the distal end portion of the at least one flexible intrabody catheter shown with a three-dimensional physical representation; andat least one display in communication with the at least one processor configured to display the rendered interactive near RT visualizations,wherein the at least one processor is configured to direct an MRI scanner to generate a tracking sequence and an imaging sequence, the imaging sequence using the anatomical scan planes used to generate the near RT MRI images, wherein the tracking sequence is carried out independent of the imaging sequence, wherein the tracking sequence generates X, Y, Z coordinates of the tracking coils of the distal end portion of the flexible catheter from signals from different MR scanner channels associated with coaxial cables from the tracking coils of the at least one flexible intrabody catheter which, during operation, are connected to the different channels of the MR scanner, wherein the at least one processor is configured to generate the three dimensional physical representation of at least the distal end of the flexible catheter showing orientation and shape using the X, Y and Z coordinates of the tracking coils and pre-defined information of the at least one flexible intrabody catheter including a distance to a tip of the at least one flexible intrabody catheter from one or more of the tracking coils and a spatial relationship of the tracking coils to select relevant anatomical scan planes for the near RT MRI images that can include scan volumes that exclude the catheter, wherein the at least one processor is configured to calculate a location in the 3-D MRI image space coordinate system associated with the tip and/or distal end portion of the at least one flexible intrabody catheter using the tracking coil data to identify at least one of the relevant anatomical scan planes,wherein the at least one processor is configured to calculate a catheter-tissue interface location proximate the tip of the at least one flexible intrabody catheter in the 3-D MRI image space using the X, Y, Z coordinates of locations of the tracking coils, wherein the calculation is carried out to place the catheter tissue interface location a distance of between about 0-5 mm b... | MRI Interventions, Inc. | A61B005/05 A61B019/00 A61B005/00 A61B005/055 A61B018/14 A61B017/00 A61B018/00
US20100284930 | A1 | 2010-11-11 | FLUORESCENT IMAGING OF SKIN CANCERS USING PHOSPHOLIPID ETHER COMPOUNDS | 2009-05-11 | The invention generally relates to novel fluorescent phospholipid compounds, compositions comprising these compounds, and their use in a variety of diagnostic applications, including fluorescence imaging of tumors. A preferred compound of the present invention has the following structural formula:; | 1. A method of diagnosing a skin malignancy comprising (a) administering to a patient a fluorescent phospholipid ether compound having one of the following formulas:or formula II:wherein n is an integer between 7 and 23; Y is H, Me or Et;Z is a fluorophore; andX is selected from the group consisting of(b) visualizing the anatomy of the skin; and (c) distinguishing a benign tissue from a malignant tissue, wherein the malignant tissue displays significantly more fluorescence caused by the fluorescent phospholipid compound. 2. The method according to claim 1, wherein in the fluorescent phospholipid ether compound, Z is selected from the group consisting of:wherein R is selected from the group consisting of H, Me, Et, Br and I; Q is selected from the group consisting of N, CH, C-Me, C-Et, C—CF3, C-Ph;R1 is selected from the group consisting ofL is selected from the group consisting of O, S, NH and NMe;R2 is selected from the group consisting of H, Me, and Et;R3 is selected from the group consisting of H, OMe, OEt, and NMe2;R4 is selected from the group consisting of H, Me, Et, Ph or p-methoxy-phenyl; andm is an integer from 1 to 5. 3. The method according to claims 1 or 2 wherein the fluorescent phospholipid analog is selected from the group consisting of: 4. The method according to claim 1, wherein the fluorescent phospholipid compound has formula: 5. The method of claim 1, wherein step (a) takes place from about 1 hour to about 4 hours prior to step (b). | - | A61B005/00
US20100331640 | A1 | 2010-12-30 | USE OF PHOTODETECTOR ARRAY TO IMPROVE EFFICIENCY AND ACCURACY OF AN OPTICAL MEDICAL SENSOR | 2009-06-26 | A system and method for determining physiological parameters of a patient based on light transmitted through the patient. The light may be transmitted via an emitter and received by a detector array that includes a plurality of detector elements. The emitter and the detector may both be located on a flexible substrate. | 1. A physiological sensor comprising: an emitter adapted to transmit light;a detector array comprising a plurality of detector elements each configured to receive the transmitted light via a respective light path; anda flexible substrate comprising both the emitter and the detector array. 2. The physiological sensor, as set forth in claim 1, comprising a second emitter configured to transmit light at a wavelength different from a wavelength of the emitter. 3. The physiological sensor, as set forth in claim 2, wherein the plurality of detector elements are arranged around the emitter and the second emitter and wherein the plurality of detector elements are configured to receive the transmitted light from each of the emitter and the second emitter. 4. The physiological sensor, as set forth in claim 1, wherein the flexible substrate comprises a material capable of maintaining its shape once adjusted. 5. The physiological sensor, as set forth in claim 4, wherein the material comprises a thermoplastic polymer. 6. The physiological sensor, as set forth in claim 1, wherein the detector elements are organized into a line or into a two dimensional grid. 7. A pulse oximetiy system comprising: a pulse oximetry monitor; anda sensor assembly configured to be coupled to the monitor, the sensor assembly comprising: an emitter configured to transmit light;an array of detector elements each configured to receive the transmitted light from the emitter; anda flexible substrate comprising both the emitter and the array of detector elements. 8. The pulse oximetry system, as set forth in claim 7, wherein each of the detector elements is configured to transmit an electrical signal to the pulse oximetry sensor based on the light received from the light emitter. 9. The pulse oximetry system, as set forth in claim 8, comprising a processor configured to scan each of the detector elements for the electrical signal corresponding to the strongest light transmission received from the emitter. 10. The pulse oximetry system, as set forth in claim 9, wherein the processor is configured to select in real time the electrical signal corresponding to the strongest light transmission received from the emitter for calculation of physiological parameters. 11. The pulse oximetry system, as set forth in claim 7, comprising a second emitter configured to transmit light at a wavelength different from a wavelength of the emitter. 12. The pulse oximetry system, as set forth in claim 11, comprising a processor configured to determine a first physiologic parameter based on the light transmitted from the emitter and a second physiologic parameter based on the light transmitted from the second emitter. 13. The pulse oximetry system, as set forth in claim 7, comprising a second emitter configured to transmit light at an identical wavelength to the light transmitted from the emitter, wherein the plurality of detector elements are configured to receive the light from the emitter and the... | Nellcor Puritan Bennett LLC | A61B005/1455
US20100298665 | A1 | 2010-11-25 | LIVING BODY INFORMATION DETECTION APPARATUS AND BLOOD-PRESSURE METER | 2010-07-30 | A living body information detection apparatus including a pair of arms opposed to each other. The apparatus including a spindle for connecting between the arms of the pair at each end of the arms. The apparatus further including a distance variable mechanism, provided in the spindle, for adjusting an interval between the other ends of the pair of arms. The apparatus also including a detection part, for detecting living body information, attached to the other end of at least one arm of the pair of arms on a side opposed to another arm. | 1. A living body information detection apparatus, comprising: a pair of arms opposed to each other;a spindle for connecting between the arms of the pair at each end of the arms;a distance variable mechanism, provided in the spindle, for adjusting an interval between the other ends of the pair of arms; anda detection part, for detecting living body information, attached to the other end of at least one arm of the pair of arms on a side opposed to another arm. 2. The living body information detection apparatus as claimed in claim 1, further comprising: a rotation mechanism for rotating at least one arm of the pair of arms around the spindle as a center axis. 3. The living body information detection apparatus as claimed in claim 1, wherein the detection part is mounted on the top of an adjustment screw attached to a screw hole passing through the other end of the arm. 4. The living body information detection apparatus as claimed in claim 1, wherein at least one arm of the pair of arms is variable in length from the spindle to the other end. 5. The living body information detection apparatus as claimed in claim 1, wherein the pair of arms is configured so as to pinch a part of an ear part of a human body with an arm placed on one side of the part of the ear part and an arm placed on another side of the part of the ear part. 6. The living body information detection apparatus as claimed in claim 5, wherein the part of the ear part is a part of an auricle and a periphery of the auricle. 7. The living body information detection apparatus as claimed in claim 6, wherein the part of the auricle is a part of a tragus and a periphery of the tragus. 8. The living body information detection apparatus as claimed in claim 1, wherein one arm of the pair of arms has a shape, on an opposite side of a side opposed to another arm of the pair of arms, that follows the concha auriculae or the cavity of the concha of the human body 9. The living body information detection apparatus as claimed in claim 1, wherein one arm of the pair of arms is provided with a cushion on an opposite side of a side opposed to another arm of the pair of arms. 10. The living body information detection apparatus as claimed in claim 1, further comprising an ear suspension for suspending the living body information detection apparatus from a base of the auricle of the human body. 11. The living body information detection apparatus as claimed in claim 10, wherein the ear suspension is detachable from a main body of the living body information detection apparatus. 12. The living body information detection apparatus as claimed in claim 10, wherein the ear suspension and the cushion pull against each other using a force of magnets. 13. The living body information detection apparatus as claimed in claim 1, comprising a light shielding cover for shielding at least the detection part from the outside. 14. The living body information detection apparatus as claimed in claim 1, comprising a l... | NIPPON TELEGRAPH AND TELEPHONE CORP. | A61B005/00
US20100331654 | A1 | 2010-12-30 | SYSTEMS FOR DIABETES MANAGEMENT AND METHODS | 2010-06-30 | Described herein are systems and methods to utilize factual information based on stored blood glucose data to allow greater insight into the management of diabetes of a user. | 1. A diabetes management system comprising: a blood glucose test strip having an inlet port to receive a blood sample; anda blood glucose measurement device having a test strip port configured to receive the test strip and a microcontroller coupled to a memory, display screen, and user interface buttons so that upon completion of a current blood glucose measurement with the test strip, the display screen presents a value of the current blood glucose measurement and at least one of two selectable message headers on the display screen comprising a review of blood glucose values stored in the memory over a predetermined time period, or a trend pattern of the stored blood glucose values. 2. The system of claim 1, in which the message indicative of a distribution of stored blood glucose values comprises an identification of a percentage of stored blood glucose values within a first distribution range, a percentage of stored blood glucose values, if any, lower than a minimum in the first distribution range and a percentage, if any, higher than a maximum in the first distribution range. 3. The system of claim 1, in which the message indicative of a distribution of stored blood glucose values comprises a first distribution range, a second distribution range with a value smaller than a minimum in the first distribution range, and a third distribution range with a value greater than a maximum in the first distribution range, actual number of measurements having values within respective first, second and third distribution ranges, and a percentage of the measurements within respective first, second and third distribution ranges. 4. The system of claim 1, in which the message indicative of a distribution of stored blood glucose values above a first predetermined threshold comprises measured values and the dates and times corresponding to the respective measured values. 5. The system of claim 1, in which the message indicative of a distribution of stored blood glucose values below a second predetermined threshold comprises measured values and the dates and times corresponding to the respective measured values. 6. The system of claim 1, in which the message indicative of a trend of the stored values over a second predetermined time period comprises a text indicating one of an increasing trend or decreasing trend in the measured values of the blood glucose. 7. The system of claim 1, in which the message indicative of a trend of the stored values comprises a table having the date, time, and measured blood glucose value at such date and time. 8. The system of claim 1, in which the message comprises a textual message. 9. An analyte test measurement device comprising: a housing;a measurement sensor disposed in the housing and configured to receive an analyte test strip; anda microcontroller coupled to a memory, display screen, user interface buttons and the measurement sensor which are mounted to the housing so that upon completion of a current analyte meas... | LifeScan Scotland Ltd. | A61B005/145
US20100312094 | A1 | 2010-12-09 | MRI-GUIDED SURGICAL SYSTEMS WITH PRESET SCAN PLANES | 2010-06-08 | The systems include a display with a User Interface in communication with a circuit, the display is configured to display a patient model/map or model of target anatomical structure. The User Interface is configured to allow a user to select at least one target site (e.g., ablation site) on the model/map whereby scan planes for a corresponding location in 3-D space for the selected target ablation site are pre-set for future use. The circuit can be configured to automatically direct the MRI Scanner to use the pre-set scan planes to obtain MR image data when an intrabody device such as an ablation catheter is at the location that corresponds with the selected target site. | 1. An MRI guided system, comprising: a circuit; anda display in communication with the circuit, the display having a User Interface (UI), the display configured to display a volumetric planning model of a target anatomical structure of a patient, wherein the UI is configured to allow a user to select at least one target site on the planning model, and wherein the circuit associates at least one pre-set scan plane for the at least one selected target site for subsequent use. 2. The system of claim 1, wherein the circuit is in communication with an MR Scanner having a 3-D imaging space with a coordinate system, and wherein the circuit is configured to electronically register the planning model to the 3-D MRI imaging space to define the at least one pre-set scan plane used by the MR Scanner based on the selected at least one selected site on the registered planning model. 3. The system of claim 1, wherein the circuit is configured to allow a user to select one of a plurality of different pre-defined surgical procedures, each procedure having a defined set of proposed target treatment sites that can be electronically automatically placed on the planning model as the target sites. 4. The system of claim 1, wherein the circuit is configured to provide an adjustable template or grid of suggested target treatment sites that is electronically applied to the planning model. 5. The system of claim 2, wherein the circuit is configured to automatically electronically direct the MRI Scanner to use the at least one pre-set scan plane associated with a corresponding selected target site to obtain an image slice of MR image data when a distal end portion of an intrabody device is determined to be proximate a location that corresponds to the selected target site. 6. The system of claim 1, wherein the planning model is an interactive model, and wherein the UI is configured to allow a user to rotate the planning model and electronically mark target sites to select the target sites at different locations. 7. The system of claim 1, wherein the circuit is configured to present suggested scan planes that will generate at least one image slice that will include relevant local tissue associated with the at least one selected target site based on a contour of the tissue. 8. The system of claim 1, and wherein, in response to the user selection of the at least one target site on the planning model, the circuit also electronically associates a corresponding 3-D perspective viewing window of the model that is subsequently automatically displayed during an MRI guided procedure when an intrabody device is at the location in the body that corresponds to the selected target site. 9. The system of claim 1, wherein the display comprises a viewing window with near real-time MR images of local tissue being actively treated by an intrabody device, wherein the circuit generates the near real-time MR images using at least one pre-set scan plane associated with a selected target s... | - | A61B005/055
US20100298681 | A1 | 2010-11-25 | Analyte Monitoring Device and Methods of Use | 2010-07-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for calibration of a continuous glucose sensor, the method comprising: receiving sensor data from a continuous analyte sensor, the sensor data comprising one or more sensor data points;receiving and processing calibration information;evaluating a predictive accuracy of calibration information; anddetermining when to request reference data based at least in part on the predictive accuracy of calibration information. 3. The method of claim 2, wherein the step of determining when to request reference data comprises determining a time period. 4. The method of claim 3, wherein the time period is between about 0 minutes and 7 days. 5. The method of claim 2, wherein the step of receiving and processing calibration information comprises receiving one or more matched data pairs, wherein the step of evaluating a predictive accuracy comprises evaluating a correlation of at least one matched data pair with at least some of the calibration information. 6. The method of claim 5, wherein the step of determining is based at least in part on the correlation of the at least one matched data pair and the calibration information. 7. The method of claim 2, wherein the step of receiving and processing calibration information comprises receiving reference data from a reference analyte monitor, forming at least one matched data pair, forming a calibration set including said at least one matched data pair, and forming a calibration line from said calibration set, wherein the step of evaluating a predictive accuracy comprises evaluating a correlation of the matched data pairs in the calibration set with the calibration line. 8. The method of claim 7, wherein the step of determining is based at least in part on the correlation of the matched pairs in the calibration set and the calibration line. 9. The method of claim 2, wherein the step of receiving and processing calibration information comprises receiving a matched data pair and forming a calibration set including said matched data pair, wherein the step of evaluating a predictive accuracy comprises evaluating a discordance of the matched data pair and/or the matched data pairs in the calibration set. 10. The method of claim 9, wherein the step of determining is based at least in part on the discordance of the matched data pair and/or the matched data pairs in the calibration set. 11. The method of claim 2, wherein the step of receiving and processing calibration information comprises forming one or more matched data pairs by matching time corresponding sensor and reference data and forming a calibration set including one or more matched data pairs wherein the step of evaluating a predictive accuracy comprises iteratively evaluating a plurality of combinations of matched data pairs in the calibration set to obtain a plurality of calibration lines. 12. The method of claim 11, further comprising removing matched data pairs from the calibration set in response to the iterative evaluatio... | Abbott Diabetes Care Inc. | A61B005/05
US20100292555 | A1 | 2010-11-18 | Analyte Monitoring Device and Methods of Use | 2010-07-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for calibrating an analyte sensor, the method comprising: receiving sensor data from an analyte sensor, the sensor data comprising at least one sensor data point;receiving reference data from a reference analyte monitor, the reference data comprising at least one reference analyte value;matching at least one sensor data point with at least one reference analyte value to form a matched data pair;converting sensor data into at least one estimated analyte value, utilizing the matched data pair; anddisplaying the estimated analyte value within about 10 minutes of receiving the reference analyte value. 3. The method of claim 2, wherein the step of displaying the estimated analyte value is performed within about 5 minutes of receiving the reference analyte value. 4. The method of claim 2, wherein the step of displaying the estimated analyte value is performed within about 1 minute of receiving the reference analyte value. 5. A method for calibrating glucose sensor data from a continuous glucose sensor, the method comprising: obtaining a reference glucose value; andmatching the reference glucose value with a sensor glucose value without compensating for a time lag between the reference glucose value and the sensor glucose value such that a time stamp for the reference glucose value is as close as possible to a time stamp of the sensor glucose value. 6. The method of claim 5, wherein the time stamp of the reference glucose value is within about 5 minutes of the time stamp of the sensor glucose value. 7. The method of claim 5, wherein the time lag between the reference glucose value and the sensor glucose value is determined, at least in part, by a filter applied to raw glucose sensor data measured by the continuous glucose sensor. 8. A method for calibrating glucose sensor data from a continuous glucose sensor, the method comprising: calibrating a continuous glucose sensor by matching a reference glucose value with a first sensor glucose value; andsubsequently calibrating the continuous glucose sensor by matching the reference glucose value with a second sensor glucose value. 9. A method for calibrating glucose sensor data from a continuous glucose sensor, the method comprising: immediately calibrating a continuous glucose sensor by matching a reference glucose value with a first sensor glucose value; andsubsequently calibrating the continuous glucose sensor by matching the reference glucose value with a second sensor glucose value. 10. The method of claim 9, wherein the first sensor glucose value and the second sensor glucose value are different. 11. The method of claim 9, wherein the step of immediately calibrating comprises matching the reference glucose value with the first sensor glucose value without compensating for a time lag. 12. The method of claim 9, wherein the step of subsequently calibrating comprises matching the reference glucose value with the second sensor glucose value. 13. The method of claim 12 w... | Abbott Diabetes Care Inc. | A61B005/05
US20100331645 | A1 | 2010-12-30 | METHODS AND SYSTEMS FOR WIRELESS COMMUNICATION BETWEEN A BLOOD GLUCOSE METER AND A PORTABLE COMMUNICATION DEVICE | 2009-06-25 | A blood glucose measuring system which comprises a blood glucose (bG) meter and a portable communication device (PCD) and methods thereof are disclosed. The blood glucose meter comprises a measurement module and a wireless module, wherein the measurement module is operable to measure the blood glucose level of a blood sample, the wireless module is an embeddable module and communicates to the measurement module via a serial interface, and the wireless module is operable to wirelessly communicate to the portable communication device. The portable communication device is operable to wirelessly receive information from the blood glucose meter related to the blood glucose measurement. | 1. A blood glucose measuring system comprising a blood glucose (bG) meter and a portable communication device (PCD), wherein: the blood glucose meter comprises a measurement module and a wireless module, wherein: the measurement module measures the blood glucose level of a blood sample;the wireless module is an embeddable module and communicates to the measurement module via a serial interface and wirelessly to the portable communication device; andthe portable communication device receives information from the blood glucose meter related to the blood glucose measurement. 2. The system of claim 1 wherein the wireless communication is Bluetooth, Zigbee, or infrared light. 3. The system of claim 1 wherein the PCD is a cellular phone, a smart phone, or a personal digital assistant. 4. The system of claim 1 wherein the measurement module comprises a measurement controller, user switches, a measurement application-specific integrated circuit (ASIC), a display, a serial interface to the communication module, and a battery. 5. The system of claim 4 wherein the measurement ASIC receives a measurement strip and determines the blood glucose level of a blood sample placed on the measurement strip. 6. The system of claim 1 wherein the communication module comprises a communication controller, a communication transceiver, an antenna, a voltage regulator, and an interface to the measurement module. 7. A method for establishing wireless communication between a blood glucose (bG) meter and a portable communication device (PCD), wherein the bG meter comprises a measurement module and a wireless module, wherein the measurement module measures the blood glucose level of a blood sample, and the wireless module is an embeddable module and communicates to the measurement module via a serial interface, the method comprising: switching on power to the bG meter and the PCD;sending a pairing beacon by the bG meter to the PCD, wherein the pairing beacon comprises information related to a type of the bG meter and a serial number of the bG meter;receiving the pairing beacon by the PCD;determining whether the PCD will establish wireless communication to the bG meter based on the type and the serial number of the bG meter;sending a response to the pairing beacon by the PCD to the bG meter indicating whether the PCD will establish wireless communication to the bG meter; andindicating on the bG meter whether the wireless communication was established. 8. The method of claim 7 wherein the sending of the pairing beacon and the sending of the response to the pairing beacon are performed by way of Bluetooth, Zigbee, or infrared light. 9. The method of claim 7 wherein the PCD is one of a cellular phone, a smart phone, and a personal digital assistant. 10. The method of claim 7 wherein the measurement module comprises a measurement controller, user switches, a measurement application-specific integrated circuit (ASIC), a display, a serial interface to the communication module... | Roche Diagnostics Operations, Inc. | A61B005/1468
US07846110 | B2 | 2010-12-07 | Self-contained test unit for testing body fluids | 2007-08-01 | A self contained disposable test unit for testing body fluid comprises a body member and a support member. The support member is moveable with respect to the body member between a first position and a second position. The support member includes a body part receiving surface for receiving a patient's body part. A lancet is carried by the body member and includes a lancet tip capable of piercing the skin of a patient to produce fluid flow. A test member is capable of interacting with body fluid to aid in the determination of information about body fluid components. A capillary member is capable of directing fluid flow to the test member. A pressure cup is capable of exerting pressure on a body part to foster fluid flow out of a lanced site and into the capillary, and a calibration member is provided for containing information for facilitating calibration of the test unit. | 1. A self contained disposable testing unit for testing body fluid comprising: a body member,a support member moveable with respect to the body member between a first position and a second position, the support member including a body part receiving surface for receiving a patient's body part,a lancet movable with respect to the support member, the lancet including a first end portion pivotably coupled to the body member for pivotal movement with respect to the body member, and a second end portion having a second end and a base end, the base end being disposed closer to the first end portion than the second end and the second end including a lancet tip capable of piercing a patient's skin to produce fluid flow, wherein during pivotal movement of the lancet, the base end and lancet tip of the second end portion move in an arcuate path and extend along a tangent of an arc scribed by the pivotal movement, wherein the base end follows the lancet tip;a test member capable of interacting with body fluid to aid in determining information about body fluid components;a capillary member capable of directing fluid flow to the test member,a pressure cup for exerting pressure on a body part to foster fluid flow out of a lanced site and into the capillary; anda calibration member capable of containing information for facilitating calibration of the testing unit. 2. The testing unit of claim 1 wherein the lancet is pivotably moveable between (1) a storage position wherein the lancet tip is disposed below the body part receiving surface of the support member,(2) a piercing position wherein the lancet tip is disposed above the body part receiving surface of the support member, and(3) a retracted position wherein the lancet tip is disposed below the body part receiving surface of the support member,wherein the lancet is moved into the retracted position through an engagement of the support member with the lancet as the support member moves between the first and the second position. 3. The testing unit of claim 2 wherein the lancet moves generally perpendicularly to a patient's skin when the lancet pierces a patient's skin as the lancet moves between the storage position and the piercing position, and wherein the lancet assumes the piercing position when the support member is positioned between the first position and the second position. 4. The testing unit of claim 2 wherein the body member includes a lancet movement resisting surface for providing resistance to movement of the lancet into the retracted position. 5. The testing unit of claim 2 wherein the lancet includes a first end pivotably coupled to the body member to permit the lancet tip to move in an arcuate path between the piercing position and the retracted position. 6. The testing unit of claim 2 wherein the body member and the support member share a common axis and wherein the support and body members are rotatable relative to each other about the common axis for permitting the user to vary a dis... | Advanced Medical Products GmbH | A61B005/00 A61B017/14 A61B017/32 A61B017/34 B65D081/00
US07848939 | B1 | 2010-12-07 | Methods for validating participants in a pension insurance program | 2003-04-15 | A method for ensuring valid participation of a plan participant in a defined contribution pension plan provides data query for identifying the plan participant. A pension debt coverage policy protects an outstanding balance of a defined contribution plan loan secured by participant's plan assets, eliminating adverse tax consequences to the plan beneficiary in the event of the death or disability of the participant and ensuring that the full plan account balance passes intact to beneficiary in such event. | 1. A computer-implemented method for ensuring valid participation of a plan participant in a defined contribution pension plan, the method comprising the steps of: establishing a pension debt mitigation coverage in response to a requirement for life insurance covering a balance of a loan from a retirement account of the plan participant, wherein the debt mitigation coverage comprises an insurance policy or debt cancellation coverage;wherein the debt mitigation coverage protects assets in the retirement account of the plan participant in the defined contribution pension plan from unnecessary dissipation from default due to a known exposure inherent within a risk event;wherein the debt mitigation coverage is imposed as a pre-condition to making the loan;wherein the assets in the retirement account serve as collateral for the loan to the participant; andwherein the debt mitigation coverage is held for the benefit of the plan and the participant in a plan trust or by a trustee for an initial amount of the loan;creating data structure of a computer for identifying the status of the plan participant;receiving, by a computer, an election from the plan participant regarding acceptance of the debt mitigation coverage;entering into the data structure an indicator indicating whether the plan participant has a debt mitigation coverage covering the plan or whether the plan participant has not accepted the offered pension debt mitigation coverage but has separately arranged for debt mitigation coverage from an offerer not associated with the plan;configuring, by a computer, a premium of the loan based upon the indicator in the data structure; andin the occurrence of the risk event, paying to the policy holder from the plan trust the balance of the loan from the debt mitigation coverage. 2. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises an insurance policy. 3. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises a credit insurance policy. 4. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises an individual credit insurance policy. 5. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises an individual credit life insurance policy. 6. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises an individual credit disability insurance policy. 7. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises an individual credit unemployment insurance policy. 8. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises an individual credit life and disability insurance policy. 9. The computer-implemented method according to claim 1, wherein the debt coverage policy comprises an individual credit life and unemployment insurance policy. 10. The computer-implemented metho... | Pentech, LLC | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07860288 | B2 | 2010-12-28 | X-ray device having a dual energy mode and method to analyze projection images detected in the dual energy mode | 2007-05-31 | A sequence of groups of projection images shows an object under examination comprising a vascular system and its environment. A computer determines a 2-dimensional evaluation image having a plurality of pixels based on combination images determined from the projection images of a group. The combination images have a plurality of pixels with pixel values. The sequence of the combination images shows the time characteristic of the distribution of a contrast medium in the object. The pixels of the evaluation image correspond to those of the projection images. The computer assigns each pixel, at least in a part area of the evaluation image, a type that is characteristic of whether the respective pixel corresponds to a vessel of the vascular system, a perfusion area or a background. It performs the assignment of the type on the basis of the time characteristic of the pixel values of the combination images. | 1. A method for evaluating a group of projection images of an object under examination, comprising: defining combination images having a plurality of pixels with pixel values based on the projection images in the group, a sequence of the combination images indicating a time characteristic of the pixel values showing a distribution of a contrast medium over time in the object;determining an evaluation image having a plurality of pixels with pixel values based on the combination images, the pixels of the evaluation image corresponding to the pixels of the combination images;assigning a type to each pixel in an area of the evaluation image corresponding to a type selected from the group consisting of: a vessel of the object, a perfusion area of the object, and a background of the object; and performing the type assignment based on the time characteristic of the pixel values of the combination images for identifying a vascular system of the object,wherein the object is an iteratively moving object,wherein the combination images are selected from a series of projection images of the iteratively moving object each having an assigned phase information that deviates by a maximum of one phase boundary from a reference phase,wherein the reference phase or the phase boundary is specified by a user, andwherein a number of the selected combination images is determined and outputted to the user. 2. The method as claimed in claim 1, wherein the pixel values of the pixels are defined in at least essentially identical areas of the object in the projection images. 3. The method as claimed in claim 1 , wherein one of the projection images is selected by the user and outputted together with the assigned phase information or the deviation of the assigned phase information to the user. 4. The method as claimed in claim 1, wherein a type evaluation core is defined by pixels of an evaluation image, andwherein the type assignment is performed to pixels of the combination images that are in the type evaluation core based on a time characteristic of pixel values of the pixels. 5. The method as claimed in claim 4, wherein the type assignment is performed by: type lots that are divided in an area of the evaluation image, ora time characteristic of a weighted or unweighted mean value or a maximum of pixel values of the pixels occurring in the type evaluation core. 6. The method as claimed in claim 5, wherein the type evaluation core corresponds to the type lots. 7. The method as claimed in claim 4, wherein one of the projection or combination images is displayed in the evaluation image. 8. The method as claimed in claim 4, wherein a degree of perfusion for pixels of the perfusion area in the evaluation image is determined and assigned to the pixels,wherein an extent evaluation core is defined by the pixel of the evaluation image, andwherein the assignment of the degree of perfusion is performed to pixels of the combination images that are in the extent evaluation core ... | Siemens Aktiengesellschaft | G06K009/00 H05G001/70 H05G001/10 A61B005/05 A61B005/02
US07835927 | B2 | 2010-11-16 | Medication management system | 2002-12-27 | A system and method for confirming that a medication administration device has been programmed with the correct medication administration parameters. A medical database carrier is used compare medication delivery parameters entered into a medication administration device to institutionally established guidelines or more widely accepted protocols to ensure that the medication is delivered in accordance to those guidelines. The medical database carrier may also be configured to communicate information regarding medication delivery and other patient information between a control system in communication with the care-giving facility's other information systems and a patient specific asset such as an infusion pump. The medical database carrier may be a smartcard, a PDA such as a Palm™ Pilot, laptop computer, pager, mobile phone, or other device capable of storing, processing and communicating information. The system may use either wired or wireless connections to communicate information between the components of the system. | 1. A system for use with an infusion pump for reducing the possibility of medication errors, the infusion pump being capable of being programmed with pumping parameters, the system comprising: a first program configured to accept user inputs related to pump operating parameter limits and to generate a data base from those user inputs; anda second program located in an infusion pump configured to: receive the pump parameter data base generated by the first program;compare pumping parameters programmed into the infusion pump by an operator with the data base; andmonitor the comparison and if the comparison exceeds the limits established in the data base, provide an indication to a pump operator that the programmed pumping parameter is outside the limits in the data base. 2. The system of claim 1 wherein the first program is configured to be executed on a first computer separate from the infusion pump. 3. The system of claim 1 wherein the first program includes names of drugs and associates multiple pumping parameters with at least one of said drugs. 4. The system of claim 1 wherein the second program causes the pump to display a visual indication to the operator that the programmed pumping parameter is outside the limits selected by the facility in which the pump is located. 5. The system of claim 4 wherein the indication presented to the operator further queries the operator as to whether the operator desires to override the indication. 6. The system of claim 1 wherein the second program creates a log of all comparisons that exceed the limits established in the data base. 7. The system of claim 1 wherein: the first program comprises an editor subprogram that enables an operator of a first computer to generate a data base of pumping parameters associated with multiple drugs; andthe first program is also configured to have a second subprogram that, when commanded, will install the data base in a selected infusion pump. 8. The system of claim 7 wherein the infusion pump comprises a processor that automatically compares the programmed pumping parameters to the data base that has been installed in the pump and immediately presents an indication of a program parameter that is outside the data base as it is detected. 9. The system of claim 7 wherein the infusion pump comprises a processor that automatically compares the programmed pumping parameters to the data base that has been installed in the pump one-by-one as they are entered into the pump and presents an indication of a program parameter that is outside the data base if such is detected. 10. The system according to claim 1 wherein the second program is further configured to: prevent an operator from initiating pumping action of the pump once an indication of a parameter being outside the data base is presented until; andpermitting an operator to initiate pumping action of the pump once the operator has either changed the outside parameter to be within the limits in the data base or indica... | Carefusion 303, Inc. | A61B005/00 G06F019/00
US07844089 | B2 | 2010-11-30 | Image interpretation report creating apparatus | 2007-05-31 | An image interpretation report creating apparatus for displaying a case in combination of an image interpretation report and a blood test result even if they are stored in different computers. The apparatus includes: a display unit for displaying an image interpretation report creation screen along with a blood test result; an input unit for specifying finding and a blood test result; a search key generating unit for generating search keys for searching an image interpretation report DB and a blood test result DB; a search processing unit for causing an image interpretation report server and an examination result server to obtain cases that match the search keys; a similarity degree determining unit for displaying the cases in an order of the degree of similarity; and an image interpretation report preservation unit for causing the image interpretation report server to store image interpretation report data. | 1. An image interpretation report creating apparatus to be used for creating an image interpretation report when connected directly or via a network to at least one terminal for displaying a medical image, an image server for storing image data of medical images, an image interpretation report server for storing image interpretation report database accumulating image interpretation report data of image interpretation reports created based on medical images, and at least one examination result server for storing examination result database accumulating examination result data of examination results obtained by examination using no medical image, said apparatus comprising: a display unit for displaying an image interpretation report creation screen for creating an image interpretation report of a medical image obtained by examination using a medical image along with an examination result obtained by examination using no medical image;input means to be used for inputting finding of the medical image and specifying first information and second information in the finding and an examination result displayed on said display unit, respectively;search key generating means for generating at least one first search key for searching said image interpretation report database and at least one second search key for searching said examination result database, based on the first and second information, respectively;search processing means for causing said image interpretation report server to obtain at least one case that matches said at least one first search key and causing said examination result server to obtain at least one case that matches said at least one second search key;similarity degree determining means for determining a degree of similarity between (i) the first and second information specified in the finding and the examination result displayed on said display unit and (ii) the at least one case obtained by said image interpretation report server and said examination result server, and displaying the obtained at least one case in a list on said display unit in an order of the degree of similarity;case display means for displaying contents of a case, which is selected by a user from among the at least one case displayed in the list, on said display unit; andimage interpretation report preservation means for causing said image interpretation report server to store image interpretation report data of the image interpretation report created based on the image interpretation report creation screen displayed on said display unit. 2. The image interpretation report creating apparatus according to claim 1, wherein said search processing means includes: image interpretation report search processing means for causing said image interpretation report server to search said image interpretation report database for image interpretation report data representing a case that matches said at least one first search key to obtain the image interpretation report data... | FUJIFILM Corporation | G06K009/00 A61B005/00
US20100280403 | A1 | 2010-11-04 | RAPID SERIAL PRESENTATION COMMUNICATION SYSTEMS AND METHODS | 2009-01-12 | Embodiments of the disclosed technology provide reliable and fast communication of a human through a direct brain interface which detects the intent of the user. An embodiment of the disclosed technology comprises a system and method in which least one sequence of a plurality of stimuli is presented to an individual (using appropriate sensory modalities), and the time course of at least one measurable response to the sequence(s) is used to select at least one stimulus from the sequence(s). In an embodiment, the sequence(s) may be dynamically altered based on previously selected stimuli and/or on estimated probability distributions over the stimuli. In an embodiment, such dynamic alteration may be based on predictive models of appropriate sequence generation mechanisms, such as an adaptive or static sequence model. | 1. A method, comprising: presenting at least one sequence of a plurality of stimuli to a user; andmeasuring a time course of at least one measurable response to the plurality of stimuli by monitoring at least one input signal of the user. 2. The method of claim 1, wherein monitoring at least one input signal comprises monitoring at least one of an electroencephalogram, an electromyogram, and an evoked response of the user, wherein the response of the user comprises at least one of a neural, physiological, and behavioral response. 3. The method of claim 1, wherein monitoring at least one input signal comprises detecting a positive recognition input signal distinguished from one or more negative input signals. 4. The method of claim 1, wherein monitoring at least one input signal comprises monitoring at least one input signal to correlate the at least one input signal to a nonbinary degree of similarity of the at least one sequence of a plurality of stimuli, wherein the nonbinary degree of similarity comprises at least one of a degree of recognition and a degree of acceptability. 5. The method of claim 1, wherein presenting at least one sequence of a plurality of stimuli comprises presenting a first stimulus followed by a second stimulus, wherein the second stimulus is selected for presentation according to a probability determination that the second stimulus is a next desired element in a sequence desired by the user. 6. The method of claim 5, wherein the probability determination is provided according to a natural language model. 7. The method of claim 6, wherein the natural language model is an adaptive natural language model using input from the user to modify the probability determination. 8. The method of claim 5, wherein the probability determination is provided according to a sequence model. 9. The method of claim 1, further comprising monitoring an attention level of the user. 10. The method of claim 1, wherein presenting at least one sequence of a plurality of stimuli comprises presenting a sequence of linguistic components. 11. The method of claim 1, wherein presenting at least one sequence of a plurality of stimuli comprises presenting a sequence of letters or words. 12. The method of claim 1, wherein presenting at least one sequence of a plurality of stimuli comprises presenting a sequence of nonlinguistic components. 13. The method of claim 1, wherein presenting at least one sequence of a plurality of stimuli comprises presenting a sequence visually. 14. The method of claim 1, wherein presenting at least one sequence of a plurality of stimuli comprises presenting a sequence audibly. 15. A communication system, comprising: rapid serial presentation of stimuli to a user;a user intent detection mechanism that employs one or more user signals to indicate intent of the user as to recognition of at least one of the presented stimuli, wherein the one or more user signals comprise at least one of an electrophysiological signal, ... | Oregon Health & Science University | A61B005/0482
US07844321 | B2 | 2010-11-30 | Rotating catheter probe using a light-drive apparatus | 2006-12-06 | The invention is a rotating tip catheter-imaging probe where electromagnetic energy is delivered to the distal end of a catheter and converted to mechanical energy using a light drive apparatus. The mechanical energy is then used to rotate a mirror that redirects light in fixed pattern on a sample. The rotating element of the light drive apparatus contains vanes, which rotate about an axis and positioned with bearings to minimize friction. A chamber encompasses the rotating element and is set to a vacuum pressure. The rotational speed of the catheter tip can be controlled by varying the optical power delivered to the vanes, the vacuum pressure in the chamber, or by a braking mechanism applied to the rotating element. The vanes may be shaped in a particular geometry to increase forces on the vanes from thermally driven gas flow. | 1. A device for directing optical energy, comprising: a) a body member having a proximal and a distal end;b) an optical conduit coaxial to, and within the body member;c) a reflector element in optical communication with the distal end of the optical conduit;d) a rotational element operably coupled to the reflector element; ande) at least one vane member associated with the rotational element and in optical communication with the optical conduit, wherein the at least one vane member includes a thermal gradient between a first surface and a second surface when in optical communication with the optical conduit, wherein a vacuum chamber surrounds at least a portion of the rotational element and the vane member. 2. The device of claim 1, further comprising an anterior bearing operably associated with the reflector element and a posterior bearing operably associated with the rotational element. 3. The device of claim 2, further comprising a lens in optical communication with the reflector element. 4. The imaging device of claim 3, wherein the lens comprises a gradient index lens. 5. The device of claim 2, wherein the reflector element further comprises a dichroic mirror. 6. The device of claim 2, wherein the at least one vane member further comprise radial projections from the rotational element. 7. The device of claim 2, wherein the rotational element further comprises a capillary tube having an inner lumen in optical communication with the optical conduit, wherein the inner lumen includes a light scattering material directing incident light toward the thermal gradient. 8. The device of claim 2, further comprising means for controlling rotation speed of the rotational element. 9. The device of claim 2, wherein the thermal gradient comprises a first energy absorbing surface and a second energy reflecting surface. 10. A method for directing optical energy for therapeutic or diagnostic purposes, comprising a) coupling a rotatably mounted vane member to a rotatably mounted reflecting element;b) directing optical energy to the rotatably mounted vane member within a vacuum sealed chamber, whereby the vane member and the reflecting element rotate by operation of a thermal gradient between the a first surface and a second surface of the vane member,c) directing optical energy to the reflecting element, whereby the reflecting element further directs optical energy to a region of interest. 11. The method of claim 10, further comprising focusing optical energy to the reflecting element. 12. The method of claim 11, further comprising varying the rotational speed of the vane member. | Board of Regents, the University of Texas System | A61B005/00
US08172809 | B2 | 2012-05-08 | Safety shield apparatus and mounting structure for use with medical needle devices | 2010-07-15 | A safety shield apparatus including a safety shield and mounting structure for mounting the safety shield to a medical needle. The safety shield includes a distal segment, a proximal segment and a retention member. The distal segment is pivotally secured to the proximal segment and the proximal segment is pivotally secured to the retention member. The retention member includes an opening for receiving a nose of a medical needle device, e.g., a blood collection device. In one embodiment, a retaining collar is provided which is dimensioned to be press-fit about the nose of the medical needle device to secure the safety shield to the medical needle device. Alternatively, mounting structure is formed directly on the retention member to secure the safety shield to the medical needle device. | 1. A manually advanceable safety shield apparatus comprising: a safety shield including a distal segment having a distal end and a proximal end, a proximal segment having a distal end and a proximal end, and a retention member, the proximal end of the distal segment being pivotally connected to the distal end of the proximal segment and the retention member being pivotally secured to the proximal end of the proximal segment;wherein the retention member includes an opening dimensioned to be slidably received about a nose of a medical needle device such that the distal segment and the proximal segment of the safety shield are manually movable from a retracted position to an advanced position to shield a needle supported on the medical needle device;wherein, when in the retracted position, the distal end of the distal segment extends around the nose of the medical needle device, and a longitudinal axis of the distal segment is non-parallel to a longitudinal axis of the needle, andwherein, when in the advanced position, the longitudinal axis of the distal segment is parallel to the longitudinal axis of the needle; anda retention collar dimensioned to be frictionally engaged about the nose of the medical needle device to secure the safety shield to the medical needle device;wherein the distal segment includes an angled extension having a distal end extending outwardly from a wall, the angled extension being configured to engage and slide along at least one of the needle and the nose of the medical needle device. 2. The safety shield according to claim 1, wherein the distal segment of the safety shield includes a body portion having a top wall, a bottom wall and a bearing member extending outwardly from the bottom wall, the retention member including a wall extension, wherein the bearing member is positioned to rest on the wall extension when the safety shield is in the retracted position. 3. The safety shield according to claim 1, wherein the proximal segment and the distal segment are pivotally connected by a hinge. 4. The safety shield according to claim 3, wherein the hinge is formed by cooperating elements on the proximal and distal segments. 5. The safety shield according to claim 3, wherein the retention member is integrally formed with the proximal segment. 6. The safety shield according to claim 1, wherein the proximal segment and the distal segment are manufactured as a single piece having a thinned transition region which pivotally interconnects the proximal segment and distal segment to one another, and wherein the retention member is also integrally formed with the proximal and distal segments. 7. The safety shield according to claim 1, wherein the safety shield includes a series of ribs disposed on either side of an inner surface thereof. 8. The safety shield according to claim 7, wherein the series of ribs define an inner channel configured to receive the needle therein. 9. The safety shield according to claim 1, wherein the dist... | Tyco Healthcare Group LP | A61M005/32
US08082127 | B2 | 2011-12-20 | Method for determining the spatial distribution of magnetic resonance signals through multi-dimensional RF excitation pulses | 2008-09-13 | A method for determining the spatial distribution of magnetic resonance signals from an imaging area, wherein nuclear spins are excited in a spatially encoded fashion through multi-dimensional RF pulses, is characterized in that in a definition step, a resolution grid with resolution grid cells is predetermined, and in accordance with a predetermined phase encoding scheme, an excitation pattern is defined for each phase encoding step, in which the amplitudes within the imaging area are set in accordance with a predetermined distribution identically for each phase encoding step. In a preparatory step, the amplitude and phase behavior of the RF pulses to be irradiated is calculated in accordance with a predetermined k-space trajectory for each defined complex excitation pattern. | 1. A method for determining the spatial distribution of magnetic resonance signals from a predetermined imaging area, which consists of at least one extended region within a volume of investigation of a nuclear magnetic resonance apparatus, wherein a transverse magnetization of nuclear spins is excited through multi-dimensional RF (radio frequency) pulses using magnetic field gradients and RF transmitting antenna means, said antenna means comprising at least one transmitting element, wherein in a definition step, a k-space trajectory which is to be generated through magnetic field gradient switching and which must be followed during excitation, as well as a phase encoding scheme with phase encoding steps are determined for spatial encoding using a computer or controller,in a preparatory step, the amplitude and phase behavior with time of the RF pulses to be irradiated for exciting the transverse magnetization of the nuclear spins is calculated for each transmitting element of the transmitting antenna means for the selected k-space trajectory using the computer or controller,in an execution step, transverse magnetization of nuclear spins within the volume of investigation is excited during each phase encoding step by a multi-dimensional RF pulse, phase encoding takes place according to the phase encoding scheme, and magnetic resonance signals are acquired using receiver antenna means,in a reconstruction step, a two- or three-dimensional spatial distribution and/or spatial allocation of the magnetic resonance signals and/or values derived therefrom are calculated from the magnetic resonance signals acquired in all phase encoding steps using the computer or controller, andin a display/storage step, the results of reconstruction and/or the one or more values derived therefrom are stored and/or displayed, wherein furtherin the definition step, a two- or three-dimensional resolution grid with resolution grid cells that covers the object under investigation is predetermined,an at least two-dimensional grid of directly neighboring encoding cells is determined for each region of the imaging area, wherein each encoding cell consists of one or more resolution grid cells of the imaging area, for which the same spatial encoding is to be performed,and for each phase encoding step, a complex excitation pattern is defined which is given by amplitudes and phases of the transverse magnetization to be excited in all resolution grid cells, in which the amplitudes assigned to the resolution grid cells are set within the imaging area in correspondence with a predetermined distribution identically for each phase encoding step, the amplitudes assigned to the remaining resolution grid cells are set to zero, and the phases assigned to the resolution grid cells within the imaging area are set in accordance with the defined phase encoding scheme,in the preparatory step, the amplitude and phase behavior of the RF pulses to be irradiated is calculated for each defined complex... | Bruker BioSpin MRI GmbH | G01R033/54 G06F019/00
US07869853 | B1 | 2011-01-11 | Analyte monitoring device and methods of use | 2010-08-06 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. An in vivo glucose monitoring system, comprising: an in vivo glucose sensor; anda receiver unit external to a user in signal communication with the in vivo glucose sensor, the receiver unit comprising: one or more processors; anda computer-readable storage medium including instructions which, when executed by the one or more processors, causes the one or more processors to: receive a data stream of sensor data points from the in vivo glucose sensor, the sensor data points indicative of a glucose level of the user;process the data stream of sensor data points indicative of the glucose level; anddetermine the acceleration of the rate of change of the glucose level. 2. The system of claim 1, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to determine whether the acceleration of the rate of change of the glucose level exceeds or meets one or more threshold values. 3. The system of claim 2, wherein the receiver unit further comprises an alarm, and wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to activate the alarm when the acceleration of the rate of change of the glucose level exceeds or meets the one or more threshold values. 4. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 5. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a single datapoint indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 6. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a predetermined number of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 7. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when an average of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 8. The system of claim 3, wherein the computer-readable storage further includes instructions which, when executed by the one or more processors, causes the alarm t... | Abbott Diabetes Care Inc. | A61B005/00
US08231555 | B2 | 2012-07-31 | Therapy system including multiple posture sensors | 2009-04-30 | Posture-responsive therapy is delivered by the medical system based on posture state input from only one of multiple posture sensors at any given time. An example implantable medical system includes a first posture sensor and a second sensor. A processor controls therapy delivery to the patient based on at least one of a patient posture state or a patient activity level determined based on input from only one of the first or second posture sensors. In some examples, one of multiple posture sensors of an implantable posture-responsive medical system is used to automatically reorient another posture sensor (of the system), which has become disoriented. The disoriented posture sensor may be automatically reoriented for one or more posture states at a time. | 1. A method comprising: detecting, by a processor of a therapy system, a change in orientation of a first posture sensor of the therapy system, wherein the therapy system provides posture responsive therapy to a patient; andautomatically reorienting, by the processor, the first posture sensor based on posture data from a second posture sensor of the therapy system. 2. The method of claim 1, wherein reorienting the first posture sensor comprises: receiving posture data from the first posture sensor;associating the posture data from the first posture sensor with a posture state determined based on the posture data from the second posture sensor; anddefining a reoriented posture region corresponding to the posture state with posture reference data based at least in part on the posture data from the first posture sensor. 3. The method of claim 2, wherein the posture region comprises a posture cone, and wherein the posture reference data comprises one of: a posture coordinate vector and a tolerance angle that define the posture cone, ora posture coordinate vector and a range of cosine values corresponding to an angle between the posture coordinate vector and a boundary vector of the posture cone. 4. The method of claim 2, wherein the posture region comprises at least one of a posture cone or a posture toroid. 5. The method of claim 1, further comprising: determining therapy information associated with a first posture state definition that is based on first posture data from the first posture sensor before reorienting the first posture sensor; andassociating the therapy information with a second posture state definition that is based on second posture data from the first posture sensor after reorienting the first posture sensor. 6. The method of claim 1, wherein reorienting the first posture sensor comprises reorienting the first posture sensor for a plurality of posture states based on the posture data from the second posture sensor. 7. The method of claim 1, wherein reorienting the first posture sensor comprises: determining a first sense posture vector based on posture data received from the second posture sensor;comparing the first sense posture vector to a posture region associated with the second posture sensor to determine a posture state of the patient, wherein the posture region is defined by posture reference data comprising a first reference coordinate vector;subtracting the first reference coordinate vector from the first sense posture vector to determine a difference vector;determining a second sense posture vector based on posture data received from the first posture sensor while the patient occupies the posture state;adding the difference vector to the second sense posture vector to determine a second reference coordinate vector for the first posture sensor; anddefining a reoriented posture region for the first posture sensor corresponding to the posture state of the patient with posture reference data based at least in part on the... | Medtronic, Inc. | A61B005/103 A61B005/117 A61B001/00 G06T015/00 A61M031/00
US08463346 | B2 | 2013-06-11 | Absolute calibrated tissue oxygen saturation and total hemoglobin volume fraction | 2010-06-10 | A medical device for monitoring a patient condition includes a sensor capable of being advanced transvascularly to be positioned along a volume of tissue, the sensor including a first combination of a light source and a light detector to emit light into a volume of tissue and to detect light scattered by the volume of tissue and to generate a first output signal corresponding to an intensity of the detected light. A control module is coupled to the light source to control the light source to emit light at least four spaced-apart light wavelengths, and a monitoring module is coupled to the light detector to receive the output signal and compute a measure of tissue oxygenation using the light detector output signal. | 1. A method for monitoring a patient condition in a medical device having an optical sensor, the method comprising: advancing the sensor intravascularly to position the sensor along a volume of tissue;positioning the sensor along a renal vein, the volume of tissue comprising kidney tissue;controlling the optical sensor to emit light comprising at least four spaced-apart light wavelengths;detecting light scattered by the volume of tissue, wherein detecting light comprises measuring an optical sensor output signal corresponding to an intensity of the scattered light;computing a measure of tissue oxygenation in response to the detected light; anddetecting hypoxia in the kidney tissue in response to the measurement of tissue oxygenation. 2. The method of claim 1, further comprising computing a measure of total hemoglobin volume fraction using the measure of tissue oxygen saturation. 3. The method of claim 1, further comprising adjusting a therapy in response to the measure of tissue oxygenation. 4. The method of claim 1, further comprising positioning the sensor along an inner vessel wall to orient the sensor to emit light through the vessel wall into the volume of tissue. 5. The method of claim 4, wherein the volume of tissue comprises a volume of arterial blood. 6. The method of claim 1, further comprising positioning the sensor through a vessel wall into tissue adjacent to the vessel wall to orient the sensor to emit light into the volume of tissue. 7. The method of claim 1, further comprising positioning the sensor through a vessel wall to orient the sensor to emit light into a volume of arterial blood. 8. The method of claim 1, further comprising positioning the sensor along a peripheral vein. 9. The method of claim 1, further comprising adjusting a heart failure therapy in response to detecting the tissue hypoxia. 10. The method of claim 1 further comprising positioning the sensor on a patient's skin. 11. The method of claim 10, further comprising delivering a therapy to the patient and adjusting the therapy in response to the tissue oxygenation measurement. 12. The method of claim 11, wherein the therapy comprises cardiopulmonary resuscitation. 13. The method of claim 1, wherein computing the measure of tissue oxygenation comprises: computing an attenuation for each of the at least four wavelengths of detected light;computing a second derivative with respect to wavelength for at least two intermediate wavelengths of the at least four wavelengths; andcomputing a scaled second derivative of one of the intermediate wavelengths as a ratio of the at least two intermediate wavelengths. 14. A medical device system, comprising: a sensor capable of being advanced transvascularly to be positioned along a volume of tissue, the sensor comprising a first combination of a light source and a light detector to emit light into a volume of tissue, to detect light scattered by the volume of tissue, and to generate a first output signal corresponding... | Medtronic, Inc. | A61B005/00
US08423117 | B2 | 2013-04-16 | System and method to process an acquired image of a subject anatomy to differentiate a portion of subject anatomy to protect relative to a portion to receive treatment | 2009-06-22 | A system and method to process an acquired image to differentiate a portion of subject anatomy to protect relative to a portion of anatomy to receive treatment is provided. The system includes a controller with program instructions to perform the steps of receiving an instruction to indicate a zone interest of the imaged anatomy to protect associated with treatment of the subject anatomy; calculating a first set of image elements associated with a first circulatory network in communication to supply a first zone of interest to treat and not in communication to supply a second zone of interest to protect; calculating a second set of image elements of the acquired image associated with a second circulatory network in communication to supply the second zone of interest to protect; and creating a display that differentiates the second set of image elements from a remainder of the acquired image. | 1. A method of processing an acquired image of a body portion intended to receive a treatment associated with a tool, the body portion including a first zone of interest to receive the treatment, the method comprising: using a processor to perform the steps of: receiving an instruction to indicate a second zone of interest of the imaged body portion to avoid impact associated with the treatment;identifying a first set of image elements of the acquired image associated with a first part of a circulatory network that circulates flow to the first zone of interest;identifying a second set of image elements of the acquired image associated with a second part of the circulatory network that circulates flow to the second zone of interest; and creating a display that differentiates the second set of image elements associated with the second part of the circulatory network from a remainder of the acquired image. 2. The method of claim 1, further comprising the step of: highlighting the first set of image elements associated with the first part of the circulatory network to differentiate from the remainder of the acquired image. 3. The method of claim 2 further comprising the step of: highlighting the second image elements associated with the second part of the circulatory network that circulates flow to the second zone of interest to differentiate the second image elements from the highlighted first image elements as well as the remainder of the acquired image. 4. The method of claim 1, further comprising the step of: highlighting acquired image data associated with an intersection of the first part and second part of the circulatory network. 5. The method of claim 1, further comprising the steps of: identifying and creating an illustration of at least one point in the acquired image that lies within the first set of image elements associated with the first zone of interest and outside the second set of image elements associated with the second zone of interest; andcalculating and displaying a path through the illustration of at least one point within the first set of image elements associated with the first zone of interest, wherein the path is displayed to provide a guide for the tool to reach the point to receive medical intervention. 6. The method of claim 1, further comprising the steps of: acquiring instructions of limits of a contour that defines a third zone of interest in the acquired image, the third zone of interest corresponding to an already treated portion of the body portion;identifying at least a part of the circulatory network that circulates flow to the third zone of interest; andhighlighting the acquired image data of at least part of the circulatory network associated with the third zone of interest to differentiate from the remainder of the acquired image as to be protected. 7. A system comprising: an imaging system to acquire an image of a body portion intended to receive a treatment, the body portion including a first zone of ... | General Electric Company | A61B005/05
US08452373 | B2 | 2013-05-28 | Diffusion tensor imaging-based alzheimer's diagnosis method | 2010-04-16 | A system and method is provided in which a diffusion tensor image of the gray matter of a subject's brain is recorded. The diffusion tensor image is used to determine the Fractional Anisotropy of that area of the subject's brain. It has been determined that a decrease in Fractional Anisotropy is related to a decrease in dendritic quality and/or quantity in the subject's brain, which is typically caused by the presence of Alzheimer's Disease. Consequently, the Fractional Anisotropy determination is used to diagnosis the presence of Alzheimer's Disease. | 1. A method for determining an increased likelihood that a subject has Alzheimer's Disease (AD) relative to a normal subject, said method including: imaging an area of the hippocampus of the brain of a subject to acquire a diffusion tensor age of said area;determining the Fractional Anisotropy (FA) from said diffusion tensor image;comparing said Fractional Anisotropy to normative data; andidentifying an increased likelihood of the presence of AD relative to a normal subject when said FA is less than said normative data by a statistically significant amount, wherein said statistically significant amount has a p-value of less than 0.05. 2. The method of claim 1 wherein said diffusion tensor image is acquired using magnetic resonance imaging. 3. The method of claim 1 further including calculating the mean diffusitivity of said area. 4. The method of claim 3 wherein said mean diffusivity is also compared to normative data and an increased likelihood of the presence of AD relative to a normal subject is determined when said mean diffusivity is greater than said normative data. 5. The method of claim 1 further including calculating the eigenvectors for said area. 6. A method for identifying when a subject has reduced dendritic quantity in the hippocampus of the subject's brain, said method including: imaging an area of the hippocampus of the brain of a subject to acquire a diffusion tensor image of said area;determining the Fractional Anisotropy (FA) from said diffusion tensor image;comparing said Fractional Anisotropy to normative data; andidentifying reduced dendritic quantity when said FA is less than said normative data by a statistically significant amount, wherein said statistically significant amount has a p-value of less than 0.05. 7. The method of claim 6 wherein said diffusion tensor image is acquired using magnetic resonance imaging. 8. The method of claim 6 further including calculating the mean diffusitivity of said area. 9. The method of claim 8 wherein said mean diffusivity is also compared to normative data and reduced dendritic quantity is determined when said mean diffusivity is greater than said normative data. 10. The method of claim 6 further including calculating the eigenvectors for said area. | - | A61B005/055
US08571620 | B2 | 2013-10-29 | Tissue oxygenation monitoring in heart failure | 2010-06-10 | A medical device system and associated method are used for monitoring a heart failure patient. A medical device for monitoring delivery of a therapy includes a sensor sensing an optical sensor signal corresponding to light attenuation by a volume of body tissue of a patient, a therapy delivery module to deliver a therapy, and a processor configured to compute a first tissue oxygenation measurement from the optical sensor signal prior to initiating delivery of the therapy, compute a second tissue oxygenation measurement from the optical sensor signal subsequent to initiating delivery of the therapy, compare the first and the second tissue oxygenation measurements, and determine whether the delivered therapy was successful in response to the first tissue oxygenation measurement and the second tissue oxygenation measurement. | 1. A medical device for monitoring delivery of a therapy, comprising: a sensor sensing an optical sensor signal corresponding to light attenuation by a volume of body tissue of a patient;a therapy delivery module to deliver the therapy; anda processor configured to compute a first tissue oxygenation measurement from the optical sensor signal prior to initiating delivery of the therapy, compute a second tissue oxygenation measurement from the optical sensor signal subsequent to initiating delivery of the therapy, compare the first and the second tissue oxygenation measurements, and determine whether the delivered therapy was successful in response to the first tissue oxygenation measurement and the second tissue oxygenation measurement, wherein the processor determines whether the delivered therapy was successful in response to the second tissue oxygenation measurement being greater than the first tissue oxygenation measurement. 2. The device of claim 1, wherein the processor is further configured to adjust a parameter controlling delivery of the therapy in response to the second tissue oxygenation measurement being not greater than the first tissue oxygenation measurement, compute a tissue oxygenation measurement from the sensor signal for the adjusted parameter, and continue to adjust the parameter and compute the tissue oxygenation measurement until a parameter setting resulting in a tissue oxygenation measurement that is greater than the first tissue oxygenation measurement is identified and determine whether the delivered therapy was successful in response to the identified parameter setting. 3. The device of claim 1, wherein the processor is further configured to adjust a parameter controlling delivery of the therapy to a plurality of settings, compute a tissue oxygenation measurement for each of the plurality of settings, determine a setting for the parameter corresponding to a maximum tissue oxygenation, and deliver the therapy using the determined parameter setting. 4. The device of claim 3, further comprising an activity sensor to sense an activity sensor signal corresponding to activity of the patient, wherein the processor is further configured to sense detect a change in activity level of the patient from the activity sensor signal, compute a tissue oxygenation measurement in response to detecting the change in activity level, detect a decrease in tissue oxygenation in response to the tissue oxygenation measurement computed in response to detecting the change in activity level, and adjust a parameter controlling delivery of the therapy in response to detecting the decreased tissue oxygenation. 5. The device of claim 1, wherein computing the first tissue oxygenation measurement and the second tissue oxygenation measurement comprises computing a second derivative of light attenuation with respect to wavelength and computing a tissue oxygen saturation using the second derivative. 6. The device of claim 5, wherein computing the first... | Medtronic, Inc. | A61B005/00
US08454522 | B2 | 2013-06-04 | Method and apparatus for estimating blood pressure | 2010-01-13 | A method for estimating blood pressure includes sensing a sphygmus wave at a body part of a user to which vibration is applied to generate a sensed sphygmus wave, filtering the sensed sphygmus wave to generated a filtered sphygmus wave, and estimating blood pressure of the user based on time differences between peaks of the sensed sphygmus wave and peaks of the filtered sphygmus wave. | 1. A method for estimating blood pressure, the method comprising: sensing a sphygmus wave at a body part of a user to which vibration is applied for a predetermined time to generate a sensed sphygmus wave;filtering the sensed sphygmus wave to generate a filtered sphygmus wave; andestimating blood pressure of the user based on time differences between peaks of the sensed sphygmus wave and peaks of the filtered sphygmus wave. 2. The method of claim 1, wherein the filtering the sensed sphygmus wave comprises filtering out a high-frequency signal having a frequency of the vibration from the sensed sphygmus wave to generate a filtered-out high-frequency signal, andthe estimating the blood pressure comprises estimating the blood pressure of the body part based on time differences between points of time corresponding to the peaks of the sensed sphygmus wave and points of time corresponding to peaks of the filtered-out high-frequency signal. 3. The method of claim 1, further comprising controlling application of the vibration to the body part and pressurization of the body part with different constant pressures, wherein the sensing the sphygmus wave comprises sensing sphygmus waves from the body part when the body part subjected to the vibration and pressurized with the different constant pressures. 4. The method of claim 3, wherein the sensing the sphygmus wave further comprises sensing the sphygmus wave at the body part when application of the vibration to the body part is stopped after the predetermined time, andthe estimating the blood pressure based on the time differences between the peaks of the sensed sphygmus wave and the peaks of the filtered sphygmus wave comprises estimating the blood pressures of the body part based on the time differences and a result of the sensing the sphygmus wave at the body part when application of the vibration to the body part is stopped after the predetermined time. 5. The method of claim 4, wherein the estimating of the blood pressure based on the time differences and the result of the sensing the sphygmus wave at the body part when application of the vibration to the body part is stopped after the predetermined time comprises: determining voltages corresponding to the sphygmus wave sensed when the application of the vibration to the body part is stopped after the predetermined time to generate determined voltages; andpressure of the body part by calculating blood pressures based on the determined voltages. 6. The method of claim 2, wherein the filtering the sensed sphygmus wave further comprises filtering out a low-frequency signal having a frequency lower than the frequency of the vibration from the sensed sphygmus wave,the estimating the blood pressure further comprises:determining voltages corresponding to a filtered-out low-frequency signal based on the time differences; andcalculating the blood pressure based on the determined voltages, andthe blood pressure of the body part is estimated based on the cal... | Samsung Electronics Co., Ltd. | A61B005/02
US08795206 | B2 | 2014-08-05 | Systems, methods and apparatuses for recording anatomic orientation and position | 2008-10-21 | Systems, methods and apparatuses for recording anatomic orientation and position. | 1. A system for determining the orientation of a desired head position of an animal, the system comprising: a processor;a database for storing a plurality of data, wherein the plurality of data comprises three-dimensional data for determining the orientation of a desired head position of the animal;a head orientation determination module that comprises instructions for obtaining three-dimensional data for determining the orientation of the desired head position;a head position recorder in communication with the head orientation determination module, that comprises: an attachment device that includes a bite-jig;a measuring device comprising a gyroscope and an inclinometer that measures the orientation of the desired head position; wherein the inclinometer is configured to detect and measure yaw and the gyroscope is configured to detect and measure pitch and roll, and wherein the attachment device operably couples the measuring device to at least one body part of the animal; andinstructions for obtaining a plurality of three-dimensional head positioning data from the measuring device to determine the desired head position of the animal;a first medical imaging device in communication with the head orientation determination module and with the head position recorder, wherein the imaging device comprises instructions to obtain and generate one or more images of the head of the animal; anda device to display the one or more obtained images. 2. The system of claim 1, further comprising a facial skeleton image generator that comprises instructions to generate an image of a facial skeleton of the animal. 3. The system of claim 2, further comprising instructions to couple the image of the head position recorder to the image of the facial skeleton of the animal. 4. The system of claim 3, further comprising a facial skeleton reorientation module that comprises instructions to reorient the one or more obtained images of the facial skeleton of the animal to the desired head position. 5. The system of claim 1, wherein the animal is a human patient. 6. The system of claim 1, further comprising at least three, non-collinear fiducial markers coupled to the measuring device. 7. The system of claim 6, wherein the at least three, non-collinear fiducial markers are comprised of a CT- or MRI-compatible material. 8. The system of claim 6, wherein the at least three non-collinear fiducial markers create a predetermined angle and a predetermined distance with respect to the measuring device, wherein both the predetermined distance and the predetermined angle are constant. 9. The system of claim 1, wherein the attachment device comprises an attachment band. 10. The system of claim 1, wherein the plurality of head positioning data comprises the pitch, the roll and the yaw of a measuring device. 11. A method for recording a desired head position in three-dimensions, comprising: obtaining a head position recorder as in claim 1 wherein the attachment device operably ... | The Methodist Hospital System | A61B005/00
US08929968 | B2 | 2015-01-06 | Dual electrode system for a continuous analyte sensor | 2010-07-19 | Disclosed herein are systems and methods for a continuous analyte sensor, such as a continuous glucose sensor. One such system utilizes first and second working electrodes to measure analyte or non-analyte related signal, both of which electrode include an interference domain. | 1. A sensor for continuously measuring a concentration of an analyte in a biological fluid, the sensor comprising: a substrate;a first conductive layer disposed over at least a portion of the substrate and comprising a first electrode, the first electrode comprising a first electroactive surface;a first insulating layer disposed over at least a portion of the first conductive layer;a second conductive layer disposed over at least a portion of the first insulating layer and comprising a second electrode, the second electrode comprising a second electroactive surface;a second insulating layer disposed over at least a portion of the second conductive layer, wherein the first electroactive surface and the second electroactive surface are separated by a distance such that diffusion of measured species between the first electroactive surface and the second electroactive surface is substantially reduced or eliminated; anda third conductive layer disposed over at least a portion of the second insulating layer and comprising a third electrode;wherein the first insulating layer is configured to provide insulation between the first conductive layer and the second conductive layer, and wherein the second insulating layer is configured to provide insulation between the second conductive layer and the third conductive layer. 2. The sensor of claim 1, wherein the first electroactive surface and the second electroactive surface are arranged in a staggered configuration along a longitudinal axis of the sensor. 3. The sensor of claim 1, wherein the distance is from about 0.02 inches to about 0.12 inches. 4. The sensor of claim 1, wherein the first electrode comprises a first electrical connector, wherein the second electrode comprises a second electrical connector, and wherein the first electrical connector and the second electrical connector are longitudinally spaced apart along a longitudinal axis of the sensor. 5. The sensor of claim 1, wherein at least one of the first electrode, the second electrode, or the third electrode is a reference electrode. 6. The sensor of claim 5, wherein the first electrode comprises a first electrical connector, wherein the second electrode comprises a second electrical connector, wherein the reference electrode comprises a third electrical connector, and wherein the first electrical connector, the second electrical connector, and the third electrical connector are longitudinally spaced apart along a longitudinal axis of the sensor. 7. The sensor of claim 1, wherein the first electrode comprises a first wire and the second electrode comprises a second wire, wherein the first wire and the second wire both have a substantially cylindrical shape. 8. The sensor of claim 1, further comprising a membrane. 9. The sensor of claim 8, wherein the membrane comprises a layer configured to substantially restrict or block a flow of interfering species therethrough. 10. The sensor of claim 8, wherein the membrane comprises a layer confi... | DexCom, Inc. | A61B005/05 A61B005/145 A61B005/1495 A61B005/1486 A61B005/00
US09656092 | B2 | 2017-05-23 | Methods and systems for managing, controlling and monitoring medical devices via one or more software applications functioning in a secure environment | 2010-05-12 | Systems and methods that include configurations of a medical device, user device and service platform are described. Embodiments may include a secure network to run medical applications that control and/or monitor the medical device. An online store may be provided for storing and distributing medical applications to the user device and medical device. A secure environment may be provided within the user device and medical device that protects the integrity of medical applications running on those devices. A service platform may provide a service that enables a medical authority to certify and monitor the medical applications. In some implementations, various third parties and the user of the user device may be allowed to manage and monitor the medical device. | 1. A computer-implemented method of interacting with a medical device in wireless communication with a user device, comprising: receiving, at the user device, a certified medical application wherein the user device is physically separated from the medical device and includes: a secure environment processor included within a secure environment of the user device wherein the secure environment is provided with guaranteed processor and memory resources in the event of at least one of a denial of service attack and an attempt to overload processing or memory usage in the secure environment,a nonsecure environment processor wherein the secure environment processor is physically isolated from the nonsecure environment processor,a secure environment memory coupled to the secure processor and accessible only to the secure environment processor, the secure environment memory being included within the secure environment and including a security monitor executed by the secure environment processor wherein the security monitor is configured to: identify secure traffic on the user device wherein the secure traffic is associated with a function of the certified medical application requiring security,identify a security requirement associated with the function,manage execution of the certified medical application within the secure environment in accordance with the security requirement,identify other traffic on the user device wherein the other traffic is associated with a nonsecure function of a noncertified application,determine that the nonsecure function does not require security,allow the noncertified application to run in the nonsecure environment,a nonsecure environment memory coupled to the nonsecure environment processor;at least one of a Bluetooth radio, a Bluetooth Low Energy radio and a WiFi radio for communicating with the medical device, wherein the at least one of the Bluetooth radio, the Bluetooth Low Energy radio and the WiFi radio is exclusively controlled by the secure environment processor when communicating with the medical device;a wide area network radio wherein the wide area network radio receives the certified medical application from a service platform;storing code for at least a portion of the certified medical application and code for a first operating system in the secure environment memory wherein the secure environment memory segment is isolated from the nonsecure environment memory and wherein the nonsecure environment memory is configured to store code for a second operating system and the noncertified applications, wherein: the first operating system supports execution of the certified medical application by the secure environment processor,the second operating system supports execution of the one or more noncertified applications,the first operating system operates independently of the second operating system wherein the first operating system remains operational when the second operating system is turned off or becomes non-op... | Chronicmobile, Inc. | A61N001/372 G06F019/00 G06Q010/06 G06Q050/22 A61B005/00
US09596999 | B2 | 2017-03-21 | Body-worn pulse oximeter | 2009-09-14 | The invention provides a body-worn system that continuously measures pulse oximetry and blood pressure, along with motion, posture, and activity level, from an ambulatory patient. The system features an oximetry probe that comfortably clips to the base of the patient's thumb, thereby freeing up their fingers for conventional activities in a hospital, such as reading and eating. The probe secures to the thumb and measures time-dependent signals corresponding to LEDs operating near 660 and 905 nm. Analog versions of these signals pass through a low-profile cable to a wrist-worn transceiver that encloses a processing unit. Also within the wrist-worn transceiver is an accelerometer, a wireless system that sends information through a network to a remote receiver, e.g. a computer located in a central nursing station. | 1. A system for monitoring a patient's blood oxygen saturation, comprising: a pulse oximetry sensor comprising: a first radiation source configured to emit a first optical radiation at a first wavelength; a second radiation source configured to emit a second optical radiation at a second wavelength; a photodetector configured to detect the first optical radiation emitted by the first radiation source after the first optical radiation passes through a portion of the patient to generate a first signal, and the second optical radiation emitted by the second radiation source after the second optical radiation passes through a portion of the patient to generate a second signal; and a finger-ring housing comprising a mounting portion that comprises the first and second radiation sources and the photodetector, the mounting portion configured to wrap around a base of the patient's thumb while leaving a tip of the thumb uncovered such that the first optical radiation emitted by the first radiation source and the second optical radiation emitted by the second radiation source pass through the thumb tissue and are received by the princeps pollicis artery before reaching the photodetector; anda processing unit configured to be worn on the patient's wrist and connected to the first sensor to receive the first and second signals, the processing unit comprising a motion sensor configured to detect a motion signal, and a processor configured to process: i) the first and second signals to determine AC signals; ii) at least one of the AC signals and the motion signal to determine selected AC signals; and iii) the selected AC signals, or signals derived therefrom, to determine the patient's blood oxygen saturation. 2. The system of claim 1, wherein the mounting portion comprises a curved, ring-shaped portion. 3. The system of claim 2, wherein the curved, ring-shaped portion is configured to partially surround a portion of the base of the patient's thumb. 4. The system of claim 3, wherein the curved, ring-shaped portion connects to a flexible strap. 5. The system of claim 1, wherein the processing unit further comprises an input port configured to receive an electrical signal representing the patient's heartbeat. 6. The system of claim 5, wherein the electrical signal is an ECG signal. 7. The system of claim 5, wherein the input port is a serial port configured to receive a digitized version of the electrical signal. 8. The system of claim 7, further comprising a cable comprising an electrical circuit that measures a digital ECG signal and connects to the serial port. 9. The system of claim 6, wherein the processor is further configured to process the ECG signal to determine a time-dependent marker that precedes both the first and second signals by less than pre-determined time period. 10. The system of claim 9, wherein the time-dependent marker is a QRS complex. 11. The system of claim 9, wherein the predetermined time period is less than 500 ms. 12. The... | SOTERA WIRELESS, INC. | A61B005/1455 A61B005/0245 A61B005/00 A61B005/021 A61B005/022 A61B005/0452 A61B005/11
US20100324437 | A1 | 2010-12-23 | DEVICE AND METHOD FOR ASSESSING PHYSIOLOGICAL PARAMETERS | 2008-09-12 | This invention is directed to a device and a method for assessing a subject. The device has at least one impedance measuring element functionally connected to a programmable element, programmed to analyze an impedance measurement, and to provide an assessment of at least one respiratory parameter of the subject. In another embodiment, the device has at least one small-scale motion measuring element. In another embodiment, the device preferably has a high resolution lens, a camera, and a programmable element, which is programmed to analyze at least one physiological parameter of the subject. The physiological parameter is preferably obtained by measuring differential displacements in the camera's field of view. The method according to the present invention involves recording a physiological parameter, analyzing the parameter to make a prediction, and providing an indication of the prediction to a user. | 1. A device for assessing a subject, the device comprising: at least one impedance measuring element functionally connected to a programmable element, programmed to analyze an impedance measurement, and to provide an assessment of at least one respiratory parameter of the subject. 2. A device according to claim 1, wherein the at least one impedance measuring element is one or more remote probes, wherein the one or more remote probes measure body wall movements, wherein the one or more remote probes are arranged as a net, vest, or array, and wherein one or more probes are placed on the thorax or abdomen of the subject. 3.-5. (canceled) 6. The device according to claim 1, wherein the at least one respiratory parameter is selected from the group consisting of the subject's respiratory rate, the subject's respiratory pressure, the subject's respiratory flow, the subject's end tidal CO2, the subject's sublingual CO2, intensity of respiration, variability of intensity of respiration, depth of respiration, variability of depth of respiration, localization of respiration, variation in localization of respiration, shape of a respiratory curve, change in shape of a respiratory curve, a respiratory curve based on inhaled volume, a respiratory curve based on exhaled volume, a respiratory curve based on inhaled pressure, a respiratory curve based on exhaled pressure, a respiratory curve based on inhaled flow, a respiratory curve based on exhaled flow, a respiratory curve based on motion of the subject's chest as measured by imaging, a respiratory curve based on motion of the subject's chest as measured by, contact sensors placed on the chest, and combinations thereof, and wherein the impedance measurement is based on a plurality of remote probe data sets, and wherein the programmable element is further programmed to enhance at least one of the plurality of remote probe data sets; or to stabilize at least one of the plurality of remote probe data sets; or to analyze each of the plurality of remote probe data sets for dynamic range and signal to noise ratio (SNR) values; or to evaluate and remove global motion; or to build a motion vector map to determine target displacements; or to calculate a differential motion map. 7.-8. (canceled) 9. The device according to claim 1, wherein the at least one respiratory parameter is selected from the group consisting of intensity of an acoustic signal, variability of an acoustic signal and is recorded for a duration of 30 seconds; continuously; intermittently; up to at least 10 of the subject's breaths; up to at least 100 of the subject'breaths; or up to at least 1000 of the subject's breaths. 10.-13. (canceled) 14. A device for assessing a subject, the device comprising: at least one small-scale motion measuring element functionally connected to a programmable element, programmed to analyze a small-scale motion measurement and to provide an assessment of the subject based on the small-scale motion measurement, where... | - | A61B005/08
US20100324428 | A1 | 2010-12-23 | METHOD AND DEVICE FOR THE NON-INVASIVE MEASUREMENT OF DYNAMIC CARDIOPULMONARY INTERACTION PARAMETERS | 2009-02-13 | A method for a non-invasive determination of cardiopulmonary interaction parameters in a patient includes fitting a pressure cuff on the patient, setting a volume of the pressure cuff in a pulsatile range of the patient, measuring pulsatile signals over time, and evaluating the measured pulsatile signals so as to ascertain the cardiopulmonary interaction parameters. | 1-10. (canceled) 11. A method for a non-invasive determination of cardiopulmonary interaction parameters in a patient comprising: fitting a pressure cuff on the patient;setting a volume of the pressure cuff in a pulsatile range of the patient;measuring pulsatile signals over time; andevaluating the measured pulsatile signals so as to ascertain the cardiopulmonary interaction parameters. 12. The method as recited in claim 11, wherein the cardiopulmonary interaction parameters include at least one of a stroke volume variation, a pulse pressure variation and a pre-ejection phase variation. 13. The method as recited in claim 11, wherein the measuring is performed over at least one breathing cycle of the patient, 14. The method as recited in claim 11, wherein the measuring is performed over at least three breathing cycles of the patient. 15. The method as recited in claim 11, further comprising ascertaining a respirator variation range of the cardiopulmonary interaction parameters. 16. The method as recited in claim 11, wherein the measuring is performed while keeping the volume of the pressure cuff constant. 17. The method as recited in claim 11, wherein the setting of the volume includes setting the volume between a volume for ascertaining a systolic blood pressure of the patient and a volume for ascertaining the diastolic blood pressure of the patient. 18. The method as recited in claim 17, further comprising deriving a value for performing a pulse contour method from the pulsatile signals. 19. A device for a non-invasive determination of cardiopulmonary interaction parameters in a patient comprising: a pressure cuff configured to measure a cuff pressure in a pulsative range of the patient over at least one breathing cycle of the patient; anda control device configured to detect a measured cuff pressure value and to evaluate a measured pulsatile signal so as to ascertain the cardiopulmonary interaction parameters. 20. The device as recited in claim 19, further comprising an output device configured to output the cardiopulmonary interaction parameters. 21. The device as recited in claim 20, further comprising a volume-regulating device configured to regulate a volume in the pressure cuff. | UP-MED GMBH | A61B005/022
US20100281982 | A1 | 2010-11-11 | Hearing Test and Screening System and Its Method | 2009-05-07 | A hearing test and screening system for testing and screening the hearing acuity of a person which includes a computer having built therein a graphic-control program for generating a pure tone signal required for a hearing test, an audio apparatus having a connection port coupled to the computer for receiving the pure tone signal from the computer and outputting a set of test signals, each having 3 frequencies minimum, at a constant sound volume upon receipt of the pure tone signal and a response button for operation by the test person to provide a feedback signal to the computer, and earphones coupled to the audio apparatus for enabling the test person to hear the set of test signals outputted by the audio apparatus. | 1. A hearing test and screening system for testing and screening the hearing acuity of a person, comprising: a computer, said computer comprising a program with graphic user interface for generating a pure tone-signal required for a hearing test;an audio apparatus coupled to said computer and adapted for receiving said pure-tone signal from said computer and outputting a set of test signals at a constant sound volume upon receipt of said pure tone signal, each said test signal comprising at least 3 frequencies, said audio apparatus comprising a connection port connected to said computer and a response button operable by a test person to provide a feedback signal to said computer; anda set of earphone means coupled to said audio apparatus and adapted for enabling a test person to hear said a set of test signals outputted by said audio apparatus. 2. The hearing test and screening system as claimed in claim 1, wherein said computer is selected from the group of desk computer and notebook computer; said earphone means is a calibrated test earphone. 3. The hearing test and screening system as claimed in claim 1, wherein said graphic-control program comprises a first application program and a second application program developed subject to the graphical programming of Lab VIEW from National Instrument Corp. 4. The hearing test and screening system as claimed in claim 3, wherein said first application program is adapted for measuring the case-value of the noise floor of the test environment and, during application of said first application program, the normal or better ear is firstly used, assisting achievement of tolerable test environment noise floor; said second application program is adapted for measuring hearing screening standard test values, including the parameters of audio frequency and sound intensity. 5. The hearing test and screening system as claimed in claim 1, wherein each said audio signal has an audio frequency selected from a group of 500 Hz, 1000 Hz, 2000 Hz and 4000 Hz or 6000 Hz, a sound intensity selected from a group of 0 dBHL, 2.5 dBHL, 5 dBHL, 7.5 dBHL, 10 dBHL, 12.5 dBHL, 15 dBHL, 17.5 dBHL, 20 dBHL, 22.5 dBHL, 25 dBHL, 27.5 dBHL, 30 dBHL, 32.5 dBHL, 35 dBHL, 37.5 dBHL, 40 dBHL, 42.5 dBHL and 45 dBHL; said second application program select at least three audio signals from a group of 500 Hz, 1000 Hz, 2000 Hz, 4000 Hz or 6000 Hz for test subject to the identity of the test person to be a young child, juvenile or adult. 6. The hearing test and screening system as claimed in claim 3, wherein said second application program generates a minimum measurement threshold value quantization table after the end of the test, said minimum measurement threshold value quantization table dividing the hearing screening results into 10 measurement threshold values (S1˜S10) subject to standards for hearing screening set by American Speech-Language-Hearing Association. 7. The hearing test and screening system as claimed in claim 6, wherein sa... | - | A61B005/12
US20100280363 | A1 | 2010-11-04 | Electromagnetic Navigation of Medical Instruments for Cardiothoracic Surgery | 2010-04-26 | The present invention provides devices, instruments, systems, and methods to navigate medical instruments within the thoracic cavity. More specifically, the present invention provides a navigation system comprising medical instruments having electromagnetic tracking functionality and the integration of previously acquired imaging into a user interface of the navigation system. | 1. A method of monitoring the position of a medical instrument in three-dimensional space and in real-time during a medical procedure being performed on a patient, the method comprising: acquiring imaging of a predetermined portion of the anatomy of a patient;providing a medical instrument, the medical instrument comprising one or more electromagnetically detectable receiver coils, the position of which can be determined in three-dimensional space and in real-time;placing the at least a portion of the medical instrument within the predetermined portion of the anatomy of the patient;identifying the position, in three-dimensional space and in real-time, of the at least a portion of the medical instrument placed within the predetermined portion of the anatomy of the patient;displaying the previously acquired imaging of the predetermined portion of the anatomy of the patient on a display screen; andindicating the position of the at least a portion of the medical instrument on the display screen. 2. The method of claim 1, comprising acquiring imaging of the predetermined portion of the anatomy of the patient using one or more of x-ray fluoroscopy, computed tomography, magnetic resonance, and ultrasound imaging. 3. The method of claim 1, comprising acquiring imaging of the predetermined portion of the anatomy of the patient prior to the medical procedure being performed on the patient. 4. The method of claim 1, comprising displaying an image of at least a portion of the predetermined portion of the anatomy of the patient acquired from an endoscopic instrument together with the previously acquired imaging and the indication of the position of the at least a portion of the medical instrument on the display screen. 5. The method of claim 4, wherein the image acquired from the endoscopic instrument comprises a real-time image. 6. The method of claim 1, comprising using information comprising the position, in 3-dimensional space, of at least one fiducial marking device positioned within the predetermined portion of the anatomy of the patient to register the previously acquired imaging of the predetermined portion of the anatomy of the patient with the actual anatomy of the patient. 7. The method of claim 1, comprising transmitting an electromagnetic signal to the one or more electromagnetically detectable receiver coil. 8. The method of claim 1, wherein the medical procedure comprises cardiothoracic surgery. 9. A medical instrument configured so the position of at least a portion of the medical instrument can be monitored in three-dimensional space and in real-time during a medical procedure being performed on a patient, the medical instrument comprising: a body portion; andan electromagnetically detectable receiver coil integrated with the body portion and configured so the position of at least a portion of the one or more electromagnetically detectable receiver coils can be determined in real-time and in three-dimensional space. 10. The medical i... | Medtronic, Inc. | A61B005/05
US20100324455 | A1 | 2010-12-23 | DEVICES FOR MANAGEMENT OF FOOT INJURIES AND METHODS OF USE AND MANUFACTURE THEREOF | 2010-05-22 | The present invention provides orthotic devices for use in managing the treatment and prevention of lower extremity injuries, including foot ulcers. In various aspects, the present invention provides foot-worn orthotics which provide for improved compliance monitoring, and methods of their manufacture and use. | 1. A footwear orthotic for monitoring delivery of therapy to an injured foot, comprising: (i) at least one proximity sensor configured to generate an electronic signal indicative of placement of the orthotic on a foot of a wearer;(ii) a computer processor operably connected to said at least one proximity sensor and said at least one pressure sensor, wherein said computer processor receives data indicative of periods of placement of the orthotic on said foot and stores said data for future retrieval;(iii) a power supply operably connected to said at least one proximity sensor, said at least one pressure sensor, and said computer processor; and(iv) a communications circuit configured to provide communication of data received by said computer processor, or a processed form thereof, to a display or to a second computer processor external to said orthotic. 2. A footwear orthotic according to claim 1, further comprising: (v) at least one pressure sensor configured to generate an electronic signal indicative of weight-bearing use of the orthotic by said wearer, wherein said computer processor receives data indicative of periods of weight bearing use of the orthotic by said wearer and stores said data for future retrieval. 3. The orthotic of claim 2, wherein the at least one pressure sensor generates electronic signals indicative of pressures detected at a plurality of locations within the orthotic. 4. The orthotic of claim 3, wherein the electronic signals indicative of pressures detected at a plurality of locations are used to determine a pressure profile, and said pressure profile is used to identify the wearer for initiation of a therapy protocol. 5. The orthotic of claim 1, comprising at least two proximity sensors, wherein said computer processor is configured to store data indicative of placement of the orthotic on the foot of said wearer when said computer processor receives indicative electronic signals from each of said at least two proximity sensors simultaneously. 6. The orthotic claim 1, further comprising a therapeutic for delivery to the foot of said wearer, wherein delivery of said therapeutic is controlled by said computer processor such that delivery of said therapeutic is limited to periods when said at least one proximity sensor indicates placement of the orthotic on said foot. 7. The orthotic of claim 1, further comprising one or more one or more emitters of low intensity electromagnetic radiation having a peak emission wavelength of between 400 nm and 1200 nm for delivery of electromagnetic radiation to the foot of said wearer, wherein delivery of said electromagnetic radiation is controlled by said computer processor such that delivery of said electromagnetic radiation is limited to periods when said at least one proximity sensor indicates placement of the orthotic on said foot. 8. The orthotic of claim 1, further comprising one or more one or more electrodes for delivery of electrical current to the foot of said wearer, whe... | LASERCURE SCIENCES, INC. | A61B005/103 A61F005/00 A43B013/38
US07830257 | B2 | 2010-11-09 | Long-range cattle identification system | 2006-04-24 | A long-range cattle identification system used to monitor a plurality of animals. The system uses a plurality of electronic information tags that are attachable to animals. The electronic information tags transmit and receive electronic information with at least one other electronic information tag. The electronic information tags are spatially disposed so that one electronic information tag can transmit and receive at least one other electronic information tag's individual electronic information to form a mesh network. The system includes a communication device that can communicate with at least one electronic information tag and receive information from that tag. The communication device can also receive the information contained in all the other electronic information tags in the mesh network through the at least one electronic information tag. | 1. A method for making a tag system for attachment to a plurality of animals, comprising the steps of: making a plurality of electronic information tags, with each tag attachable to a different animal and each tag made by the process of: installing a circuit board in a housing configured to be attached to an animal; mounting a wireless sensor networking chip to the circuit board, the wireless sensor networking chip being configured to work in a mesh network using tag-to-tag relaying of electronic information between tags attached to different animals; and installing an electronic memory to communicate with the wireless sensor networking chip to store algorithms for supporting software and to store the electronic information. 2. The method of claim 1, wherein communication devices are mounted to fixed locations, within a preselected area, so as to accumulate electronic information from the electronic information tags. 3. A method for making a tag system for attachment to a plurality of animals, comprising the steps of: making a plurality of electronic information tags, with each tag attachable to a different animal and each tag made by the process of: installing a circuit board in a housing configured to be attached to an animal;mounting a wireless sensor networking chip to the circuit board, the wireless sensor networking chip being configured to work in a mesh network using tag-to-tag relaying of electronic information between tags attached to different animals;installing a power source in the housing to supply the wireless sensor networking chip with power; andinstalling an electronic memory to communicate with the wireless sensor networking chip to store algorithms for supporting software and to store the electronic information. 4. A method for recording data onto a tag system for a long-range animal identification system, the method comprising the steps of: attaching tags to a plurality of animals such that each animal has a tag attached thereto, wherein the tags comprise a housing encapsulating a circuit board, a wireless sensor networking chip configured to work in a mesh network using tag-to-tag relaying of electronic information, a power source, and an electronic memory used to store the electronic information, the mesh network encompassing multiple animals; andusing a communication device to record the electronic information onto the tag. 5. The method of claim 4, wherein the electronic information comprises premise identification data. 6. The method of claim 4, wherein the electronic information comprises ownership data. 7. The method of claim 4, wherein the electronic information comprises identification data to identify a particular animal upon which the tag is connected. | The Board of Regents of the University of Oklahoma | G08B013/14 G08B023/00 G08B001/08 A61B005/00
US07833164 | B2 | 2010-11-16 | System and method for monitoring autonomic balance and physical activity | 2009-08-10 | An implantable device monitors the balance between sympathetic tone and parasympathetic tone as a function of an activity level. Cardio-neurological healthy users exhibit a generally sympathetic tone in conjunction with heavy activity level and a generally parasympathetic tone in conjunction with periods of low activity level. Deviations from expected results are associated with a health problem. Measured conditions are stored and available for subsequent reporting to a remote programmer. Therapy delivered by an implantable device is determined as a function of the relationship between autonomic balance and activity level. | 1. An implantable device comprising: an autonomic tone sensor to provide a tone signal as a function of a sympathetic nervous system and as a function of a parasympathetic nervous system for a predetermined period of time;an activity sensor to provide an activity signal as a function of physical activity corresponding to the predetermined period; andan information processing unit connected to the autonomic tone sensor and connected to the activity sensor and adapted to generate an output signal based upon a relationship between the tone signal and the activity signal, compare the output signal to a baseline that correlates different tone signals with different activity signals over a predetermined period of time, and classify a patient's health status based upon the comparison. 2. The device of claim 1 wherein the information processing unit is adapted to use a baseline established with population norms. 3. The device of claim 1 wherein the information processing unit is adapted to establish multiple patient-specific baselines for different exertion levels that include a resting baseline. 4. The device of claim 1 wherein the information processing unit is adapted to establish multiple patient-specific baselines for different exertion levels that include a maximum exertion baseline. 5. The device of claim 1 wherein the autonomic tone sensor includes at least one sensor selected from any combination of a group including a heart rate sensor, a nerve electrode, a biochemical sensor, a ventilation sensor and a muscle activity sensor. 6. The device of claim 1 wherein the activity sensor includes at least one sensor selected from any combination of a group including a minute ventilation sensor, an accelerometer, a respiratory sensor, a QT interval sensor, an impedance sensor, a contractility sensor and a depolarization sensor. 7. The device of claim 1 further including a telemetry circuit connected to the information processing unit to communicate with a remote device. 8. The device of claim 1 further including a therapy circuit connected to the information processing unit. 9. The device of claim 1 further including a transducer for generating an audible or tactile warning message when a dangerous condition is detected. 10. The device of claim 1 wherein the autonomic tone sensor and information processing unit are adapted to determine heart rate variability. 11. A method comprising: generating an autonomic tone signal as a function of autonomic tone detected during an epoch using an implantable sensor;obtaining an activity signal as a function of physical activity during the epoch; andgenerating an output signal based upon a relationship between the tone signal and the activity signal, compare the output signal to a baseline that correlates different tone signals with different activity signals over a predetermined period of time, and classify a patient's health status based upon the comparison. 12. The method of claim 11 wherein the baseline... | Cardiac Pacemakers, Inc. | A61B005/00
US07835785 | B2 | 2010-11-16 | DC magnetic-based position and orientation monitoring system for tracking medical instruments | 2005-10-04 | Miniaturized, five and six degrees-of-freedom magnetic sensors, responsive to pulsed DC magnetic fields waveforms generated by multiple transmitter options, provide an improved and cost-effective means of guiding medical instruments to targets inside the human body. The end result is achieved by integrating DC tracking, 3D reconstructions of pre-acquired patient scans and imaging software into a system enabling a physician to internally guide an instrument with real-time 3D vision for diagnostic and interventional purposes. The integration allows physicians to navigate within the human body by following 3D sensor tip locations superimposed on anatomical images reconstructed into 3D volumetric computer models. Sensor data can also be integrated with real-time imaging modalities, such as endoscopes, for intrabody navigation of instruments with instantaneous feedback through critical anatomy to locate and remove tissue. To meet stringent medical requirements, the system generates and senses pulsed DC magnetic fields embodied in an assemblage of miniaturized, disposable and reposable sensors functional with both dipole and co-planar transmitters. | 1. Apparatus for determining the position and orientation of one or more points on a medical instrument relative to a reference frame in at least five degrees of freedom, comprising: a) a pulsed DC magnetic field transmitter, including at least one DC magnetic field generating element;b) a signal generation module, connected to said at least one generating element, said signal generation module comprising a digital to analog converter and a signal processor connected thereto, said digital to analog converter converting a digital signal to an analog signal and providing said analog signal to said at least one generating element of said pulsed DC magnetic field transmitter;c) a medical instrument including at least one pulsed DC magnetic field sensor, the sensor configured to be attached to said medical instrument by mounting the sensor on or in said medical instrument, said medical instrument configured to be located within a patient and said sensor not visible outside said patient;d) a 3D guidance electronics unit connected to said sensor;e) wherein said signal generation module is configured to produce one or more pulsed DC transmit signals whose pulse time is adjusted for optimal performance and is configured to provide said one or more transmit signals to said at least one generating element of said pulsed DC magnetic field transmitter, said 3D guidance electronics unit is configured to receive one or more detected signals from said sensor, wherein at least a portion of said one or more detected signals: (1) corresponds to said one or more pulsed DC transmitted signals, (2) is used to determine position and orientation of said sensor in at least five degrees of freedom from strength of said one or more pulsed DC transmitted signals and at least a portion of said one or more detected signals, and (3) indicates the position and orientation of one or more points on said medical instrument;f) said DC magnetic field transmitter being integrated with an optical tracking system to produce a hybrid system of continuously tracking and registering location of the DC magnetic field transmitter when it is moved in space;g) an imaging modality, connected to an imaging computer, said imaging modality capturing a plurality of 2D or 3D images of an anatomical region of interest of the patient; and said at least one pulsed DC magnetic field sensor connected to said imaging computer;h) wherein said imaging computer in communication with said 3D guidance electronics unit is configured to perform an operation to register the plurality of 2D or 3D images of the patient's anatomical region of interest with the said instrument within a patient; said imaging computer is configured to perform further operation to superimpose a cursor or icon representing said instrument onto the plurality of 2D or 3D images of the patient's anatomical region of interest; said cursor or icon is configured to move on the plurality of 2D or 3D images of the patient's anatomical region of... | Ascension Technology Corporation | A61B005/05
US07850615 | B2 | 2010-12-14 | System and method of determining cardiac pressure | 2005-07-19 | A pressure sensor is deployed in the right atrium and is in contact with the tissue of the fossa ovalis. The fossa ovalis acts as a membrane and the pressure sensor determines the relative and/or absolute pressure within the left atrium while remaining within the right atrium. A variety of embodiment are provided to deploy and anchor the sensor into the proper position. | 1. An implantable medical device (IMD) comprising: a housing;a pressure sensor disposed within the housing and having a pressure sensing interface disposed along a first plane;a pressure sensor mount disposed within the housing and supporting the pressure sensor, wherein the mount permits the pressure sensor to move in a direction generally perpendicular to the first plane;an anchor prong tract disposed within the housing;an anchor prong disposed within the tract and moveable from a retracted position to an extended position;an anchor prong piercing head coupled to a distal end of the anchor prong and configured to pierce a substrate when moved in a first direction corresponding to moving the anchor prong from the retracted position to the extended position; andan anchor retention mechanism coupled to the anchor prong to preclude movement of the housing in a second direction that is opposite to the first direction after piercing the substrate, wherein the pressure sensor mount includes a threaded track and the pressure sensor includes a threaded portion that engages the threaded track such that rotation of the pressure sensor causes the pressure sensor to move in the direction generally perpendicular to the first plane. 2. The IMD of claim 1, wherein the anchor retention mechanism includes a pivotable member that is adapted to be generally parallel to the substrate after piercing the substrate. 3. The IMD of claim 1, further comprising an anchor prong spring disposed within the housing and positioned to bias the anchor prong in the second direction. 4. The IMD of claim 1, further comprising: a vacuum channel disposed within the housing;a vacuum grid disposed on the housing and in a plane generally parallel to the first plane. 5. The IMD of claim 4, further comprising a check valve disposed within the housing and in the vacuum channel such that when the check valve is in a closed position a generated vacuum is maintained between the check valve and the vacuum grid when the vacuum grid is occluded. 6. The IMD of claim 5, further comprising: a lead body coupled with the housing; anda lumen within the lead body coupled with the vacuum channel such that evacuation of the lumen at a proximal end of the lead body generates negative pressure within the vacuum channel. 7. The IMD of claim 1, further comprising: an abrasive ring coupled with the housing and disposed in a plane generally parallel to the first plane. 8. The IMD of claim 7, wherein the abrasive ring includes an uneven surface structure that facilitates tissue growth and encapsulation. 9. An implantable medical device (IMD) comprising: an implantable cardiac device housing that includes components operable to sense cardiac signals and deliver cardiac therapy;a lead body having a proximal end coupleable to the implantable cardiac device and a distal end;a pressure sensor assembly housing coupleable to the distal end of the lead body;a pressure sensor disposed within the pressure sensor ... | Medtronic, Inc. | A61B005/04 A61B005/02 A61N001/00
US07854228 | B2 | 2010-12-21 | Bronchial flow control devices and methods of use | 2005-03-28 | Methods and systems for lung volume reduction of a patient are described. The methods include implanting a flow control device in a bronchial passageway of the lung. The flow control device regulates fluid flow through the bronchial passageway and includes a valve protector that at least partially surrounds a valve member. The valve protector has sufficient rigidity to maintain the shape of the valve member against compression. | 1. A flow control device for a bronchial passageway, comprising: a valve member that regulates fluid flow through the flow control device;a seal member that forms a seal with the interior wall of the bronchial passageway when the flow control device is implanted in the bronchial passageway;an anchor member coupled to the seal member, wherein the anchor member exerts a radial force against the interior wall of the bronchial passageway when the flow control device is implanted in the bronchial passageway and retains the flow control device in the bronchial passageway; anda valve protector that at least partially surrounds the valve member, the valve protector having sufficient rigidity to maintain the shape of the valve member against compression, wherein the valve protector is manufactured of the same material as that of the anchor member and is integrally formed with the anchor member. 2. The flow control device of claim 1, wherein the opening extends radially though the valve protector. 3. The flow control device of claim 1, wherein the valve protector is manufactured of a different material than a material of the seal member. 4. The flow control device of claim 1, wherein the valve protector is entirely surrounded by the seal member. 5. The flow control device of claim 1, wherein the valve protector is structurally detached from the seal member. 6. The flow control device of claim 1, wherein the valve protector comprises a tube. 7. The flow control device of claim 1, wherein the valve protector is more rigid than the seal member. 8. A flow control device for a bronchial passageway, comprising: a valve member that regulates fluid flow through the flow control device;a seal member that forms a seal with the interior wall of the bronchial passageway when the flow control device is implanted in the bronchial passageway;an anchor member coupled to the seal member, wherein the anchor member exerts a radial force against the interior wall of the bronchial passageway when the flow control device is implanted in the bronchial passageway and retains the flow control device in the bronchial passageway; anda valve protector that at least partially surrounds the valve member, the valve protector being positioned within the seal member and having sufficient rigidity to maintain the shape of the valve member against compression, wherein the valve protector is more rigid than the seal member. 9. The flow control device of claim 8, wherein the valve protector includes at least one opening that extends through the valve protector and provides location where a removal device can grasp the flow control device. 10. The flow control device of claim 8, wherein the valve protector is manufactured of a different material than a material of the seal member. 11. The flow control device of claim 8, wherein the valve member is concentrically contained within the valve protector. 12. A flow control device for a bronchial passageway, comprising: a valve member tha... | Pulmonx Corporation | A61M015/00 A61M016/00 A62B018/08 A62B018/10 A61C005/14 A61F002/20 A61B005/08
US08129547 | B2 | 2012-03-06 | Anticoagulation of human blood ex vivo | 2007-01-26 | The present invention relates to a method for the anticoagulation of human blood ex vivo in which blood calcium concentration remains the same, no thrombin is formed and thrombocyte function is not affected. The invention further relates to a blood coagulation factor Xa inhibitor that can be advantageously used in this method, an agent that contains an inhibitor of the blood coagulation factor Xa, and a kit for diagnosing cellular blood components. The anticoagulated blood obtained according to the invention can be used for examining cellular blood components, for example for the diagnosis of thrombocyte function. | 1. An inhibitor of blood coagulation factor Xa comprising a compound of formula (I):or a salt with acids thereof,whereinn is 2, 3 or 4,R is —(CH2)m—NH—Y, wherein m is 1, 2, 3, 4 or 5, and Y is H orand X1, X2, X3, X4, X5 are, independently, H, F, Cl, Br, I, NO2, NH2, OH, alkyl or alkoxy. 2. The inhibitor according to claim 1 or a salt with acids thereof, wherein n is 3 or 4,m is 2, 3 or 4, Y is H orX1, X2, X3, X4, X5 are independently H, F, Cl, Br, I, NO2, NH2, OH, alkyl or alkoxy. 3. The inhibitor according to claim 1 or a salt with acids thereof, wherein n is 3, m is 3, Y is and X1, X2, X3, X4, X5 are each H, orn is 2, m is 3, Y is and X1, X2, X3, X4, X5 are each H, orn is 3, m is 4, Y is H and X1, X2, X3, X4, X5 are each H, orn is 3, m is 3, Y is X1, X2, X4, X5 are each H, and X3 is Cl. 4. The inhibitor according to claim 3 or a salt with acids thereof, wherein n is 3,m is 3, Y isX1, X2, X3, X5 are each H. 5. A method for anticoagulation of human blood ex vivo, comprising adding at least one inhibitor according to claim 1 or a salt with acids thereof to the blood. 6. The method according to claim 5, wherein at least one inhibitor or a salt with acids thereof is present in physiological NaCl solution; or a buffer solution with a physiological ionic strength, a physiological pH value, or both. 7. A tube for taking blood samples, comprising a tube and an inhibitor according to claim 1 or a salt with acids thereof. 8. An apheresis hose/bag system, comprising at least one collection bag for receiving blood, wherein the collection bag has an effective quantity of an inhibitor according to claim 1 or a salt with acids thereof. 9. An infusion comprising a) a thrombocyte concentrate,b) the inhibitor according to claim 1 or a salt with acids thereof, andc) a sterile collection bag. 10. A kit for testing cellular blood components, comprising an inhibitor of the blood coagulation factor Xa according to claim 1 and a means for testing cellular blood components. 11. The kit according to claim 10, wherein the means for testing cellular blood components is a means for thrombocyte function diagnosis. |  | C07D207/08 A61K031/40 G01N033/86
US08180440 | B2 | 2012-05-15 | Alarm system that processes both motion and vital signs using specific heuristic rules and thresholds | 2009-05-20 | The invention provides a body-worn monitor that measures a patient's vital signs (e.g. blood pressure, SpO2, heart rate, respiratory rate, and temperature) while simultaneously characterizing their activity state (e.g. resting, walking, convulsing, falling). The body-worn monitor processes this information to minimize corruption of the vital signs by motion-related artifacts. A software framework generates alarms/alerts based on threshold values that are either preset or determined in real time. The framework additionally includes a series of ‘heuristic’ rules that take the patient's activity state and motion into account, and process the vital signs accordingly. These rules, for example, indicate that a walking patient is likely breathing and has a regular heart rate, even if their motion-corrupted vital signs suggest otherwise. | 1. A system for processing at least one vital sign from a patient and generating an alarm, the system comprising: a first sensor comprising a first detector configured to detect a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;a second sensor comprising a second detector configured to detect a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;at least two motion-detecting sensors configured to be positioned on locations selected from a forearm, upper arm, and a body location other than the forearm or upper arm of the patient, each of the motion-detecting sensors generating at least one motion waveform indicative of motion of the location on the patient's body to which the motion-detecting sensor is affixed;a first processing component configured to receive the time-dependent physiological and motion waveforms and determine therefrom: (i) at least one vital sign calculated using at least one of the first and second time-dependent physiological waveforms; and (ii) at least one motion parameter selected from the patient's posture, activity state, and degree of motion calculated using at least one motion waveform generated by each of the motion-detecting sensors, wherein the motion parameter is calculated by processing a time-dependent component from at least one motion waveform, and a frequency-dependent component from at least one motion waveform; anda second processing component configured to receive the at least one vital sign and motion parameter and determine therefrom: (i) a first alarm condition, calculated by comparing the at least one vital sign to an alarm threshold; (ii) a second alarm condition, calculated from the at least one motion parameter; and (iii) an alarm rule, determined by collectively processing the first and second alarm conditions with an alarm algorithm, the alarm rule indicating whether or not the system generates an alarm. 2. A system for processing at least one vital sign from a patient and generating an alarm, the system comprising: a first sensor comprising a first detector configured to detect a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;a second sensor comprising a second detector configured to detect a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;at least two motion-detecting sensors configured to be positioned on locations selected from a forearm, upper arm, and a body location other than the forearm or upper arm of the patient, each of the motion-detecting sensors generating at least one motion waveform indicative of motion of the location on the patient's body to which the motion-detecting sensor is affixed;a first processing component configured to receive the time-dependent physiological and motion waveforms and determine therefrom: ... | Sotera Wireless, Inc. | A61B005/0205
US08346332 | B2 | 2013-01-01 | Absolute calibrated tissue oxygen saturation and total hemoglobin volume fraction | 2010-06-10 | A medical device for monitoring a patient condition includes a first combination of a light source and a light detector to emit light into a volume of tissue, detect light scattered by the volume of tissue, and provide a first output signal corresponding to an intensity of the detected light. A control module is coupled to the light source to control the light source to emit light at least four spaced-apart light wavelengths, and a monitoring module is coupled to the light detector to receive the output signal, compute a measure of tissue oxygenation in response to the light detector output signal, and detect tissue hypoxia using the measure of tissue oxygenation. | 1. A method for monitoring a patient condition in a medical device having an optical sensor, the method comprising: controlling the optical sensor to emit light comprising at least four spaced-apart light wavelengths;detecting light scattered by a volume of tissue wherein detecting light comprises measuring an optical sensor output signal corresponding to an intensity of the scattered light, the scattered light comprising light scattered by arterial and venous blood volumes present in the volume of tissue;computing a measure of tissue oxygenation correlated to the partial pressure of oxygen in the tissue from the detected light; anddetecting tissue hypoxia in response to the computed measure of tissue oxygenation. 2. The method of claim 1, wherein computing the measure of tissue oxygenation comprises: computing an attenuation for each of the at least four wavelengths of detected light;computing a second derivative with respect to wavelength for at least two intermediate wavelengths of the at least four wavelengths; andcomputing a scaled second derivative of one of the intermediate wavelengths. 3. The method of claim 1, further comprising computing a measure of total hemoglobin volume fraction in response to the measure of tissue oxygen saturation. 4. The method of claim 1, further comprising: emitting light into a volume of arterial blood;detecting light scattered by the volume of arterial blood;computing a measure of arterial oxygen saturation from the detected light scattered by the volume of arterial blood; anddetecting tissue hypoxia in response to the measure of arterial oxygen saturation. 5. The method of claim 1, further comprising adjusting a therapy in response to the measure of tissue oxygenation. 6. The method of claim 1, further comprising adjusting a heart failure therapy in response to detecting the tissue hypoxia. 7. The method of claim 1, further comprising detecting an episode of disordered breathing in response to the detected tissue hypoxia. 8. The method of claim 7, further comprising controlling a therapy delivery in response to the detecting episode of disordered breathing. 9. The method of claim 8, wherein controlling a therapy delivery comprises adjusting a continuous positive airway pressure. 10. The method of claim 1, further comprising delivering a therapy to the patient and adjusting the therapy in response to the tissue oxygenation measurement. 11. The method of claim 1, further comprising: computing an absolute tissue oxygen saturation in response to the detected light;computing a total tissue hemoglobin volume fraction in response to the detected light and the computed absolute tissue oxygen saturation;the tissue oxygenation measurement comprising both the absolute tissue oxygen saturation and the total tissue hemoglobin volume fraction. 12. The method of claim 11, further comprising: computing a tissue uniformity index in response to the tissue oxygen saturation and the total tissue hemoglobin volume fra... | Medtronic, Inc. | A61B005/00
US07884729 | B2 | 2011-02-08 | Method and system for providing data management in data monitoring system | 2010-08-02 | Method and system for providing a fault tolerant data receiver unit configured with a partitioned or separate processing units, each configured to perform a predetermined and/or specific processing associated with the one or more substantially non-overlapping functions of the data monitoring and management system is provided. | 1. A data receiver unit, comprising: a first processing unit configured to perform a first predetermined routine including generating a predetermined signal prior to a data reception period; anda second processing unit operatively coupled to the first processing unit, the second processing unit configured to receive the predetermined signal and to perform a second predetermined routine including: completing a first processing cycle and subsequently entering an inactive mode if the predetermined signal is received during the first processing cycle; orentering into an active mode and determining whether a second processing cycle is to be performed during the impending data reception period if the predetermined signal is received other than during the first processing cycle. 2. The data receiver unit of claim 1, wherein the second predetermined routine further includes, upon entering into the active mode: expediting the second processing cycle, generating a time flag associated with the expediting of the second processing cycle and entering an inactive mode if it is determined that the first processing cycle is to be performed during the impending data reception period; andperforming the second processing cycle and subsequently entering an inactive mode if it is determined that the first processing cycle is to be performed during the impending data reception period. 3. The data receiver unit of claim 2, further comprising a timer, wherein the generated time flag is provided to the timer and wherein the second processing unit is maintained in the inactive mode during the data reception period. 4. The data receiver unit of claim 1, further comprising a third processing unit operatively coupled to the second processing unit, the third processing unit configured to perform a third predetermined routine. 5. The data receiver unit of claim 4, further comprising a power supply unit operatively coupled to the second processing unit, wherein the power supply unit is configured to provide power to the first, the second and the third processing units. 6. The data receiver unit of claim 5, further comprising a memory unit operatively coupled to the second processing unit. 7. The data receiver unit of claim 6, wherein the memory unit includes a non-volatile memory. 8. The data receiver unit of claim 7, wherein the memory unit is configured to store one or more programming instructions for execution by one or more of the first processing unit, the second processing unit and the third processing unit. 9. The data receiver unit of claim 1, wherein the data receiver unit is configured to receive data related to one or more analyte levels during the data reception period. 10. The data receiver unit claim of 9, wherein the data receiver is configured to receive data from a transmitter unit coupled to an analyte sensor. 11. The data receiver unit of claim 1, wherein the second processing cycle includes a housekeeping routine. 12. An analyte monitoring and ma... | Abbott Diabetes Care Inc. | G08B023/00
US08114026 | B2 | 2012-02-14 | Methods and devices for remote ischemic preconditioning and near-continuous blood pressure monitoring | 2010-06-22 | Single- or dual-bladder devices for remote ischemic preconditioning and blood pressure monitoring are disclosed along with various oscillometry-based and other methods for detecting systolic and diastolic blood pressure while the ischemic preconditioning treatment is in progress. The devices and methods of the invention provide for delivery of ischemic preconditioning at the lowest effective cuff pressure while closely monitoring patient's hemodynamics. Advantageously, the device of the invention allows both ischemic preconditioning and blood pressure monitoring to be done on the same limb. Disposable battery-powered version of the device of the present invention is especially useful for emergency use with patients suffering from acute myocardial infarction, acute stroke, or acute trauma. Additional device configurations are described for use in a percutaneous intervention and vascular sealing settings. | 1. A device for remote ischemic preconditioning, the device comprising: a cuff sized to fit about a limb of a subject in a deflated state, the cuff configured to be inflated from the deflated state to an inflated state to at least substantially reduce a blood flow through the limb, said cuff is configured to assure a limb occlusion pressure equal or less than a systolic blood pressure of said subject,a controller connected to said cuff, said controller including a cuff inflation assembly configured to inflate and deflate said cuff according to a preconditioning treatment protocol, said treatment protocol includes a plurality of treatment cycles, each cycle comprising the intervals of: cuff inflation, during which said cuff inflation assembly inflates said cuff to at least said limb occlusion pressure,an ischemic duration, during which said controller maintains said cuff at or above said limb occlusion pressure but at or below the systolic blood pressure of the subject for at least a portion of said ischemic duration, the ischemic duration lasting for at least about a minute,cuff deflation, during which said controller deflates said cuff to said deflated state, anda reperfusion duration. 2. The device as in claim 1, wherein said controller is further configured to detect said systolic blood pressure of said subject using said cuff during said ischemic duration interval while simultaneously delivering remote ischemic preconditioning treatment. 3. The device as in claim 1, wherein said cuff is maintained at or above said limb occlusion pressure but below the systolic blood pressure of the subject throughout said ischemic duration interval. 4. The device as in claim 1, wherein said controller is further configured to periodically bring said cuff to or above said limb occlusion pressure but below the systolic blood pressure of the subject during said ischemic duration interval. 5. The device as in claim 4, wherein during said ischemic duration interval said controller is further configured to bring said cuff to or above said limb occlusion pressure but below the systolic blood pressure of the subject with intervals of about one minute or less. 6. The device as in claim 2, wherein said controller is further configured to detect said systolic blood pressure when said cuff is at or above said limb occlusion pressure but below the systolic blood pressure of the subject, whereby said systolic blood pressure is detected without reperfusion of said limb. 7. The device as in claim 6, wherein said controller is further configured to detect a diastolic blood pressure of said subject during said reperfusion duration. 8. The device as in claim 7, wherein said detection of said diastolic blood pressure is accomplished by bringing said cuff to a cuff pressure not exceeding a previously detected or estimated diastolic blood pressure, whereby said subject is under continuous hemodynamic surveillance throughout said preconditioning treatment protocol, said hemo... | Infarct Reduction Technologies Inc. | A61B005/02 A61B017/00
US08386001 | B2 | 2013-02-26 | Rocker-chute type finger-clipped oximeter | 2008-05-08 | The present invention discloses a finger-clipped oximeter comprising an upper case (2) and a lower case (1), wherein one of the upper case (2) and the lower case (1) is provided with a guide post (12) which enables the upper case (2) and the lower case (1) to depart from and approach to each other vertically, the other of the upper case (2) and the lower case (1) is provided with a guide sleeve (13) which enables the upper case (2) and the lower case (1) to depart from and approach to each other vertically, and a return spring (9) is provided between an end of the guide post (12) and the case provided with the guide sleeve (13). One of the upper case (2) and the lower case (1) is provided with rockers (7, 8) which are used to facilitate the upper case (2) and the lower case (1) to depart from and approach to each other vertically, the other of the upper case (2) and the lower case (1) is provided with chutes (14, 15) which are used to facilitate the upper case (2) and the lower case (1) to depart from and approach to each other vertically, and ends of the rockers (7, 8) fit with the chutes (14, 15). | 1. A finger-clipped oximeter comprising an upper case and a lower case, characterized in that one of the upper case and the lower case is provided with a guide post which enables the upper case and the lower case to depart from and approach to each other vertically, the other of the upper case and the lower case is provided with a guide sleeve which enables the upper case and the lower case to depart from and approach to each other vertically, and a return spring is provided between an end of the guide post and the case provided with the guide sleeve; andone of the upper case and the lower case is provided with rockers which are used to facilitate the upper case and the lower case to depart from and approach to each other vertically, the other of the upper case and the lower case is provided with chutes which are used to facilitate the upper case and the lower case to depart from and approach to each other vertically, and ends of the rockers fit with the chutes. 2. The finger-clipped oximeter according to claim 1, characterized in that the upper case is provided with the guide post which enables the upper case and the lower case to depart from and approach to each other vertically;the lower case is provided with the guide sleeve which enables the upper case and the lower case to depart from and approach to each other vertically; andthe return spring is a cantilevered plate spring. 3. The finger-clipped oximeter according to claim 1, characterized in that the lower case is provided with rockers which are used to facilitate the upper case and the lower case to depart from and approach to each other vertically; andthe upper case is provided with chutes which are used to facilitate the upper case and the lower case to depart from and approach to each other vertically. 4. The finger-clipped oximeter according to claim 3, characterized in that the rockers which are provided in the lower case are a U-shaped structure which extends through the lower case;a bottom of the U-shaped structure extends through the lower case along a horizontal direction which extends through the lower case; and two side arms of the U-shaped structure are kept parallel to longitudinal side surfaces of the lower case. 5. The finger-clipped oximeter according to claim 4, characterized in that two side arms of the U-shaped structure are provided with ends which are bent toward the center of the U-shaped structure respectively;the chutes which are provided in the upper case face both outer sides of the upper case; andthe ends of the two side arms of the U-shaped structure fit with the chutes of the upper case respectively. 6. The finger-clipped oximeter according to claim 1, characterized in that two sets of the same rockers are provided along a longitudinal direction of the lower case, and two sets of the same chutes are provided along a longitudinal direction of the upper case. 7. The finger-clipped oximeter according to claim 1, characterized in that the chutes of the uppe... | Beijing Choice Electronic Technology Co., Ltd. | A61B005/00
US08475389 | B2 | 2013-07-02 | Methods and devices for assessment of pneumostoma function | 2010-06-08 | A pneumostoma assessment and treatment system includes methods and devices for aftercare of a pneumostoma. In particular, methods and devices are provided for standardized monitoring and assessment of pneumostoma patency thereby facilitating the establishment of standards-based protocols for pneumostoma management. Methods of standardized pneumostoma patency assessment includes the use of devices for raising alveolar pressure to a controlled repeatable pressure above ambient while analyzing gas flow through the pneumostoma. In response to an assessment the functionality of the pneumostoma, the tissues of the pneumostoma may be treated with a treatment device utilizing one or more different modalities to preserve or enhance the health and function of the pneumostoma. | 1. A method for assessing function of a pneumostoma in a chest of a patient, the method including: (a) receiving an elevated-exhalation-pressure control device into the patient's mouth;(b) inserting a gas collection device into the pneumostoma;(c) assessing qualities of gas in the pneumostoma during a time period in which a patient is exhaling at an elevated pressure through the elevated-exhalation-pressure control device; and(d) removing the gas collection device from the pneumostoma. 2. The method of claim 1, wherein step (a) comprises: (a1) inserting a pneumostoma management device into the pneumostoma, (a2) receiving an elevated-exhalation-pressure control device into the patient's mouth; step (b) comprises: (b1) removing a pneumostoma management device from the pneumostoma; and then (b2) inserting a gas collection device into the pneumostoma. 3. The method of claim 1, wherein step (d) comprises: (d1) removing the gas collection device from the pneumostoma; and then (d2) inserting a pneumostoma management device into the pneumostoma. 4. The method of claim 1, wherein step (c) comprises: assessing qualities of gas in the pneumostoma during a time period in which a patient is exhaling at elevated pressure through the elevated-exhalation-pressure control device, said assessing including measuring the amount of gas flowing from the pneumostoma. 5. The method of claim 1, wherein step (c) comprises: assessing qualities of gas in the pneumostoma during a time period in which a patient is exhaling at elevated pressure through the elevated-exhalation-pressure control device, said assessing including measuring the pressure of gas in the pneumostoma. 6. The method of claim 1, further comprising: (e) comparing the qualities of gas assessed in step (d) with qualities of gas assessed in the patient on different days and determining a trend with respect to function of the pneumostoma. 7. The method of claim 1, further comprising: (e) comparing the qualities of gas assessed in step (d) with predetermined qualities of gas and taking a further treatment action in response. 8. The method of claim 1, wherein steps (a) through (d) are performed by the patient and wherein the method performed by the patient further comprises: (e) comparing the qualities of gas assessed in step (d) with a threshold and contacting a physician if the qualities of gas assessed in step (d) compare unfavorably with the threshold. 9. The method of claim 1, wherein steps (a) through (d) are performed by a medical provider and wherein the method performed by the medical provider further comprises: (e) comparing the qualities of gas assessed in step (d) with a threshold and making a treatment decision if the qualities of gas assessed in step (d) compare unfavorably with the threshold. 10. A system for assessing pneumostoma function of a pneumostoma of a patient, the system comprising: an airway interface adapted to make a seal with a natural airway of a patient;a pressure elevating... | Portaero, Inc. | A61B005/08
US08370070 | B2 | 2013-02-05 | Integrated biosensor and simulation system for diagnosis and therapy | 2008-10-24 | BioMEMS/NEMS appliance biologically monitors an individual, using biosensors to detect cellular components. Data is simulated or analyzed using systems-biology software, which provides diagnostic or therapeutic guidance. | 1. Integrated biosensor platform comprising: a multifunctional array comprising a microelectronic array of biosensor electrodes in which at least one microelectrode sensor cavity generates controllable electric current or voltage for drawing probes, samples, or reagents to locations on a sensor platform thereby allowing faster, controlled hybridization or analysis thereof, whereby the multifunctional array serves as programmable or logical interconnect for coupling or switching various sensor devices, and interacts with samples and detection system to interpret samples into data for automated analysis by a systems-biology platform that electronically configures the microelectronic array to couple or interconnect selectively according to simulation or modeling to access actual host condition via one or more biosensor signals, such signals being compared by the systems-biology platform against a simulation model or other software prediction. 2. Platform of claim 1 wherein one or more of the sensors sense food. 3. System biology platform comprising: a systems-biology platform; anda multifunctional array comprising a microelectronic array of biosensor electrodes in which at least one microelectrode sensor cavity generates controllable electric current or voltage for drawing probes, samples, or reagents to locations on a sensor platform thereby allowing faster, controlled hybridization or analysis thereof, whereby the multifunctional array serves as programmable or logical interconnect for coupling or switching various sensor devices, and interacts with samples and detection system to interpret samples into data for automated analysis by a systems-biology platform that electronically configures the microelectronic array to couple or interconnect selectively according to simulation or modeling to access actual host condition via one or more biosensor signals, such signals being compared by the systems-biology platform against a simulation model or other software prediction. 4. Platform of claim 3 further comprising a regulation overlay. 5. A systems biology platform for computational modeling and simulation of multi-cellular systems or unicellular systems comprising: internal or external databases, wherein the databases comprise genomics database, protein database, gene expression database, transcriptomics database, proteomics database, metabolomics database, pharmacogenomics database, pharmacokinetics database, disease database or drug database; anda multifunctional array comprising a microelectronic array of biosensor electrodes in which at least one microelectrode sensor cavity generates controllable electric current or voltage for drawing probes, samples, or reagents to locations on a sensor platform thereby allowing faster, controlled hybridization or analysis thereof, whereby the multifunctional array serves as programmable or logical interconnect for coupling or switching various sensor devices, and interacts with samples and detection syste... | - | G01N033/48 G06G007/48
US09218453 | B2 | 2015-12-22 | Blood glucose management and interface systems and methods | 2009-06-29 | A blood glucose management device is operable to seamlessly provide a sequence of at least two displays of data relating to one or more blood glucose measurements to a user, upon minimal user interaction, with the blood glucose management device. | 1. A portable blood glucose management device, comprising: a housing including a display and a user entry means for receiving user input;a processor in the housing operatively connected to the display and the user entry means; anda memory connected with the processor,wherein the processor is configured to:produce on the display a representation of a blood glucose measurement that may be stored in the memory,associate the blood glucose measurement with at least one other blood glucose measurement that may be stored in the memory to obtain associated blood glucose measurements, andseamlessly produce on the display a graphical representation of the associated blood glucose measurements for a period of time automatically and without user input to the user input means after producing the representation of the blood glucose measurement, andwherein the processor is further configured to:associate the blood glucose measurement and a plurality of blood glucose measurements in the memory with a context group in the memory, andseamlessly and sequentially produce on the display a plurality of graphical representations of blood glucose measurements in the context group for a period of time automatically and without user input to the user input means after producing the graphical representation of the associated blood glucose measurements. 2. The device of claim 1, wherein the processor is further configured to seamlessly produce on the display advice directed to a currently displayed one of the graphical representations upon user input into the user entry means. 3. The device of claim 2, wherein the advice includes an interpretation of the blood glucose measurements in the currently displayed graphical representation relative to pre-defined blood glucose limits. 4. The device of claim 3, wherein the advice includes a recommended dietary change. 5. The device of claim 2, wherein the user entry means includes at least two buttons, and the user input consists of one press of one button of the user entry means. 6. The device of claim 1, wherein the plurality of graphical representations includes xy-graphs of blood glucose measurements in the context group for at least one of a seven day time period and a thirty day time period. 7. The device of claim 1, wherein the entry means includes at least two buttons located adjacent to the display, and further wherein the processor is configured to assign a function to each of the at least two buttons and to produce the assigned functions on the display, and wherein the assigned functions depend on the graphical representation being currently produced on the display. 8. The device of claim 1, wherein the context groups include at least two or more of a fasting context group, a prandial context group, an exercise context group, a pre-sleep context group, an episodic testing context group, a pre-office visit context group, a glucose tolerance context group, an Average Daily Risk Ratio context group, and a predicted Hb... | Roche Diabetes Care, Inc. | A61B005/145 G06F019/00 A61B005/00 A61B005/042
US08444570 | B2 | 2013-05-21 | Signal processing techniques for aiding the interpretation of respiration signals | 2009-06-09 | According to embodiments, a respiration signal may be processed to normalize respiratory feature values in order to improve and/or simplify the interpretation and subsequent analysis of the signal. Data indicative of a signal may be received at a sensor and may be used to generate a respiration signal. Signal peaks in the respiration signal may be identified and signal peak thresholds may be determined. The identified signal peaks may be adjusted based on the signal peak threshold values to normalize the respiration signal. | 1. A method for processing a respiration signal comprising: using a processor for: obtaining the respiration signal, wherein the respiration signal comprises a plurality of signal peaks;identifying one or more signal peaks of the plurality of signal peaks;determining a signal peak threshold for the respiration signal based at least in part on the identified one or more signal peaks;adjusting one or more signal peaks of the plurality of signal peaks based at least in part on the signal peak threshold to generate an adjusted respiration signal, wherein at least one of the plurality of signal peaks remains unchanged in the adjusted respiration signal;determining a respiration parameter based at least in part on the adjusted respiration signal; andreporting the determined respiration parameter. 2. The method of claim 1, wherein the respiration signal is obtained based at least in part on data received from a sensor. 3. The method of claim 1, further comprising generating the respiration signal from a PPG photoplethysmograph (PPG) signal, a transthoracic impedance signal, a capnograph signal, a nasal thermistor signal, or an electrocardiogram (EKG) signal. 4. The method of claim 3, wherein generating the respiration signal from the PPG signal comprises: obtaining a PPG signal;generating a scalogram based at least in part on received PPG signal;analyzing one or more features within the scalogram; andgenerating the respiration signal based at least in part on the one or more analyzed features. 5. The method of claim 1, wherein determining a signal peak threshold comprises determining an upper signal peak threshold value. 6. The method of claim 5, wherein adjusting one or more signal peaks of the plurality of signal peaks comprises: comparing an amplitude value of a signal peak of the plurality of signal peaks with the upper signal peak threshold value; andreducing the amplitude value of the signal peak to a value less than the upper signal peak threshold value. 7. The method of claim 5, wherein determining a signal peak threshold further comprises determining a lower signal peak threshold value. 8. The method of claim 7, wherein adjusting one or more signal peaks of the plurality of signal peaks comprises: comparing an amplitude value of a signal peak of the plurality of signal peaks with the lower signal peak threshold value; andincreasing the amplitude value of the signal peak to a value greater than the lower signal peak threshold value. 9. The method of claim 1, wherein adjusting one or more signal peaks of the plurality of signal peaks comprises: selecting a portion of the respiration signal surrounding a signal peak of the plurality of signal peaks; andadjusting the selected portion of the respiration signal. 10. The method of claim 9, wherein adjusting the selected portion of the respiration signal comprises resealing the selected portion by a constant factor. 11. The method of claim 1, wherein the respiration parameter comprises a resp... | Nellcor Puritan Bennett Ireland | A61B005/08
US08417329 | B2 | 2013-04-09 | Analysis and regulation of food intake | 2008-05-11 | Apparatus (18, 18′) is provided, which includes one or more fundic sensors (74), configured to be applied to a fundus of a subject (10), and to generate a fundic signal. The apparatus further includes a control unit (90), which is configured to make a tentative determination of ingestion by the subject (10), determine fundic impedance responsively to the fundic signal, and calculate a duration of a period during which the fundic impedance rose during the tentatively-determined ingestion. Upon finding that the duration is less than a fundic rise duration threshold value, the control unit (90) determines that the tentative determination of ingestion is a false positive. Other embodiments are also described. | 1. Apparatus comprising: one or more antral sensors, configured to be applied to an antrum of a subject, and to generate an antral signal; anda control unit, configured to: detect ingestion of food by the subject,determine antral impedance responsively to the antral signal, andinterpret a change in antral impedance vs. a as line value of less than a threshold value as indicating that the ingested food is predominantly solid. 2. no apparatus according to claim 1, further comprising: gastrointestinal (GI) tract attachment element, configured to be coupled to a portion of a (GI)tract of a subject such that the element surrounds less than 360 degrees of the GI tract, and applies a constrictive force to the GI tract portion that causes at least a 5% reduction in a cross-sectional area of the GI tract in a vicinity of the portion; andat least one electrode, electrically coupled to the control unit and coupled to the attachment element such that the electrode contacts the GI tract portion when the attachment element is coupled to the GI tract portion. 3. The apparatus according to claim 1, further comprising a gastric band, wherein the control unit is configured to cause tightening of the gastric band in response to a determination that the ingested food is predominantly solid. 4. The apparatus according to claim 1, wherein the control unit is configured to detect the ingestion of the food by; making a tentative determination of ingestion of solid food by the subject, anddetermining that the tentative determination is not a false positive upon finding that the ingested food is predominantly solid. 5. The apparatus according to claim 4, further comprising one or more fundic sensors, configured to be applied to a fundus of a subject, and to generate a fundic signal, wherein the control unit is configured to determine fundic impedance responsive to the fundic signal, calculate a duration of a period during which the fundic impedance rose during the tentatively-determined ingestion, and determine that the tentative determination is a false positive upon finding that the duration is less than a fundic rise duration threshold value. 6. The apparatus according to claim 1, further comprising: an external device, which comprises a device wireless communication module; andan implantable wireless device communication module, which is coupled to control unit,wherein the control unit is configured to dike the implantable wireless device communication module to wirelessly transmit, to the device wireless communication module, an indication that the ingested food is predominantly solid. 7. A method comprising: detecting ingestion of food by a subject;determining antral impedance of antral impedance of the subject; andinterpreting a change in anti-al impedance vs. a baseline value of less than a threshold value as indicating that the ingested hand is predominantly solid. 8. Apparatus comprising: one or more fundic sensors, configured to be applied to a fundus of... | Metacure Ltd. | A61B005/05 A61F002/00 A61F013/00 A61N001/00
US20100324451 | A1 | 2010-12-23 | BLOOD TESTING APPARATUS | 2008-02-08 | Provided is a blood testing apparatus capable of measuring a blood sugar level while keeping the operability, even if the remaining battery life is short. This blood testing apparatus includes a casing (12) having a cylindrical opening cylinder (12b), a blood sensor (22) mounted removably in the opening cylinder (12b) of the casing (12) for analyzing the blood, a blood introducing portion (34) formed in the blood sensor (22) and having an opening for storing this opening with the blood having flown from the skin by a pierce, a laser emitting device (13) disposed in the casing (12) for causing a laser beam to pierce the skin through the inside of the opening cylinder (12b) of the casing and the opening of the blood introducing portion (34), and a needle piercing unit (18) disposed in the casing for piercing the skin with the needle through the opening of the blood introducing portion (34). | 1. A blood test apparatus comprising: a housing that comprises an open body of a cylindrical shape;a blood sensor that can be attached detachably to the open open body of a cylindrical shape of the housing and that analyzes blood;a blood guiding part that is formed in the blood sensor, that comprises an opening part and that stores blood which flows out from skin as a result of puncturing, in the opening part;a laser emitting apparatus that is provided in the housing and that punctures skin by means of laser light which passes an interior of the open body of a cylindrical shape of the housing and the opening part of the blood guiding part and punctures skin; anda needle-puncturing apparatus that is provided in the housing and that punctures skin by means of a needle which passes the opening part of the blood guiding part and punctures skin. 2. The blood test apparatus according to claim 1, wherein the needle-puncturing apparatus can be attached detachably to the open body of a cylindrical shape of the housing and comprises an open body of a same cylindrical shape as the housing. 3. The blood test apparatus according to claim 2, further comprising controlling means that, upon reception of a control signal for commanding execution of puncturing as input, executes puncturing by means of the laser emitting apparatus, if the blood sensor is attached to the open body of a cylindrical shape of the housing, andexecutes puncturing by means of the needle-puncturing apparatus, if the needle-puncturing apparatus is attached to the open body of a cylindrical shape of the housing. 4. The blood test apparatus according to claim 3, further comprising a voltage detecting section that detects a voltage of an electric battery that supplies power to the laser emitting apparatus, wherein, when the voltage detected in the voltage detecting section is less than a predetermined setting level, the controlling means stops the execution of puncturing by means of the laser emitting apparatus. 5. The blood test apparatus according to claim 2, further comprising controlling means that detects which one of the blood sensor and the needle-puncturing apparatus is attached to the open body of a cylindrical shape of the housing, and displays a detection result. 6. The blood test apparatus according to claim 5, further comprising a voltage detecting section that detects a voltage of an electric battery that supplies power to the laser emitting apparatus, wherein, when the voltage detected in the voltage detecting section is less than a predetermined setting level, the controlling means displays that puncturing is not possible by means of the laser emitting apparatus. 7. The blood test apparatus according to claim 1, further comprising: a voltage detecting section that detects a voltage of an electric battery that supplies power to the laser emitting apparatus; andcontrolling means that displays a selection mode for selecting one of the laser emitting apparatus and the needle-... | PANASONIC CORPORATION | A61B005/151
US07841991 | B2 | 2010-11-30 | Methods and apparatus for expressing body fluid from an incision | 2003-06-26 | A sample of a body fluid such as blood or interstitial fluid is obtained from a body by lancing a portion of a user's skin, preferably in an area other than a finger tip, to form an incision. After the needle has been removed from the incision, a force is applied to depress the skin in a manner forming a ring of depressed body tissue in surrounding relationship to the incision, causing the incision to bulge and the sides of the incision to open, whereby body fluid is forced out through the opening of the incision. A stimulator member is mounted to an end of a lancet-carrying housing for applying the force. The stimulator member can be movable relative to the housing, and can be either heated or vibrated to promote movement of the body fluid. | 1. A device for expressing bodily fluid from an incision in the skin, the incision being formed at an incision site in the skin surface, the device comprising: a skin stimulating member including at least three pins and at least three spaced-apart skin-engaging surfaces arranged around the incision site, said skin stimulating member having a first position with the skin-engaging surfaces contacting the skin and located at initial, outer positions, anda second position with the skin-engaging surfaces engaging the skin and located at constricting, inner positions, said inner position of each skin-engaging surface being in a radially-inward direction toward the incision site from the corresponding outer position,wherein each skin-engaging surface pivots around one of the pins when moving between the first and second positions; anda skin penetrating member connected to said skin stimulating member and positioned between the skin-engaging surfaces; andwherein said skin stimulating member includes at least three levers, each lever being pivotally mounted to said skin-stimulating member and each lever having a lower end with a roller, each said roller defining one of said spaced-apart skin-engaging surfaces and each said roller adapted to rotate about an axis extending approximately parallel to the skin surface at the incision site. 2. The device of claim 1, wherein said skin stimulating member includes at least three elongated elements spaced-apart from one another, wherein each elongated element includes a skin-engaging surface. 3. The device of claim 2, wherein each elongated element is rigid. 4. The device of claim 3, wherein each elongated element pivots to move between the inner and outer positions. 5. The device of claim 1, wherein said skin-engaging surfaces are adapted to move across the skin surface as said skin-engaging surfaces move from said outer positions to said inner positions while said skin stimulating member is held against the skin. 6. The device of claim 1, wherein said skin stimulating member is adapted to repeatedly move said skin-engaging surfaces from said outer positions to said inner positions while said skin stimulating member is held against the skin. 7. The device of claim 1, wherein the skin-engaging surfaces are adapted to draw the skin inward as the skin-engaging surfaces move from the outer positions to the inner positions, the skin being constricted by the skin-engaging surfaces in the inner positions. 8. The device of claim 1, wherein said skin penetrating member is oriented orthogonally to the skin surface when said skin stimulating member is pressed against the skin and in said first and second positions. 9. The device of claim 1, wherein the skin-engaging surfaces are adapted to bulge the skin as the skin-engaging surfaces move from the outer positions to the inner positions. 10. The device of claim 1, wherein the skin stimulating member forms a pressurized pinch of skin when the skin stimulating member c... | Roche Diagnostics Operations, Inc. | A61B005/00
US07831096 | B2 | 2010-11-09 | Medical navigation system with tool and/or implant integration into fluoroscopic image projections and method of use | 2006-11-17 | Certain embodiments of the present invention provide systems and methods for representing a tool or implant in an image. Certain embodiments provide a user interface system for displaying a representation of a tool or implant with respect to an image. The system includes a processor configured to establish a surface boundary for a region of interest depict in the image and configured to determine a position of the tool or implant with respect to the surface boundary, and a display configured to dynamically display the image and the representation to a user. The processor generates a representation of the tool or implant based on the position with respect to the surface boundary. A portion of the tool or implant inside the surface boundary is depicted in the representation with a degree of transparency compared to a portion of the tool or implant outside the surface boundary. | 1. A method for depicting a tool or implant with respect to an image for surgical guidance, the method comprising: establishing a surface boundary for a region of interest depicted in the image;determining a position of the tool or implant with respect to the surface boundary; anddisplaying a representation of the tool or implant on the image, wherein a first portion of the tool or implant inside the surface boundary is depicted in the representation with a first degree of transparency that is different than a second degree of transparency or opacity of a second portion of the tool or implant outside the surface boundary, and wherein the first portion of the tool or implant is depicted in the representation with a plurality of degrees of transparency based on a distance inside the surface boundary such that the degrees of transparency increase to become more transparent as the distance inside the surface boundary increases. 2. The method of claim 1, wherein the second portion of the tool or implant is depicted in the representation as opaque and the first portion of the tool or implant is depicted as semi-transparent. 3. The method of claim 1, wherein the transparency or opacity of the first and second portions of the tool or implant are dynamically adjusted as the tool or implant is moved relative to the surface boundary. 4. The method of claim 1, wherein the image comprises a fluoroscopic image. 5. The method of claim 1, wherein the tool or implant is depicted in the representation using various shades of grayscale. 6. The method of claim 1, wherein the representation depicts a three dimensional view of the tool or implant on a two dimensional image. 7. The method of claim 1, wherein the surface boundary is established using a tracking device. 8. The method of claim 1, wherein the step of establishing a surface boundary for a region of interest depicted in the image further comprises establishing a surface boundary for a region of interest depicted in the image based on identification of a location of one or more points on a surface of the region of interest. 9. A user interface system for displaying a representation of a tool or implant with respect to an image for surgical guidance, the system comprising: a processor configured to establish a surface boundary for a region of interest depicted in the image and configured to determine a position of the tool or implant with respect to the surface boundary, wherein the processor generates a representation of the tool or implant on the image, wherein a first portion of the tool or implant inside the surface boundary is depicted in the representation with a first degree of transparency that is different than a second degree of transparency or opacity of a second portion of the tool or implant outside the surface boundary, and wherein the first portion of the tool or implant is depicted in the representation with a plurality of degrees of transparency based on a distance inside the surface b... | General Electric Company | G06K009/00 G06K009/46 A61B005/05
US07840259 | B2 | 2010-11-23 | Method and system for electrocardiogram evaluation | 2007-01-11 | A method for evaluating an electrocardiogram is disclosed herein. The method includes measuring an electrical activity of a patient, processing the measured electrical activity to form a multi-lead signal, and extracting a segment of the multi-lead signal. The method for evaluating an electrocardiogram also includes transforming the segment of the multi-lead signal into a synthesized signal that is most representative of the patient's electrical activity, and evaluating the synthesized signal. A corresponding system for evaluating an electrocardiogram is also disclosed. | 1. A method for evaluating an electrocardiogram (ECG) comprising: measuring an electrical activity of a complete heartbeat of a patient using a cardiac monitoring system;processing the measured electrical activity in the cardiac monitoring system to form a twelve lead ECG signal for the complete heartbeat;obtaining a multi-dimensional ECG representation for the complete heartbeat having orthogonal X, Y and Z lead data from the twelve lead signal;extracting a segment of the complete heartbeat from the orthogonal X, Y and Z lead data;transforming the segment of the orthogonal X, Y and Z lead data into a synthesized signal that is most representative of the patient's electrical activity during the extracted segment; andevaluating the orientation of a projection vector for the extracted segment defined by the synthesized signal. 2. The method of claim 1, wherein said transforming the orthogonal X, Y and Z lead data includes performing a principal component analysis on the orthogonal X, Y and Z lead data. 3. The method of claim 1, wherein said extracting a segment includes extracting a T-wave segment, and wherein said evaluating the synthesized signal includes evaluating the shape of the T-wave segment defined by the synthesized signal. 4. The method of claim 1, wherein said extracting a segment includes extracting a T-wave segment. 5. The method of claim 4, wherein said evaluating the orientation of a projection vector includes evaluating the orientation of a projection vector with respect to a physical orientation of the patient. 6. The method of claim 1, wherein said extracting a segment includes extracting a T-wave segment and a QRS complex segment, and wherein said evaluating the synthesized signal includes evaluating the length of the QT interval as defined by the synthesized signal. 7. A method for evaluating an electrocardiogram (ECG) comprising: measuring an electrical activity of a complete heartbeat of a patient using a cardiac monitoring system;processing the measured electrical activity in the cardiac monitoring system to form a twelve lead ECG signal for the complete heartbeat;converting the twelve lead ECG signal of the complete heartbeat into a multi-dimensional representation of the complete heartbeat having orthogonal X, Y and Z lead data;extracting a segment of the complete heartbeat from the orthogonal X, Y and Z lead data;performing a principal component analysis on the extracted segment of the orthogonal X, Y and Z lead data to provide a synthesized signal that is most representative of the patient's electrical activity; andevaluating the orientation of a projection vector for the extracted segment defined by the synthesized signal. 8. The method of claim 7, wherein said extracting a segment includes extracting a T-wave segment, and wherein said evaluating the synthesized signal includes evaluating the shape of the T-wave segment defined by the synthesized signal. 9. The method of claim 7, wherein said extracting a segmen... | General Electric Company | A61B005/04
US20100331664 | A1 | 2010-12-30 | AUTOMATIC POSITIONING OF A SLICE PLANE IN MR ANGIOGRAPHY MEASUREMENTS | 2010-06-30 | In a method and magnetic resonance (MR) system to automatically determine a position of a slice plane in an examination region for an MR angiography measurement. MR image data are acquired from the examination region with a flow-sensitive overview imaging sequence. A three-dimensional image data set is automatically generated with the use of the acquired MR signals. Signal intensity profiles that run through the three-dimensional image data set are automatically determined. The position of a blood vessel is determined from the signal intensity profiles, and the position of the slice plane for the MR angiography measurement is automatically determined using the position of the blood vessel. | 1. A method to automatically determine a position of a slice plane in an examination region for acquiring magnetic resonance (MR) angiography data, comprising the steps of: operating a magnetic resonance data acquisition unit, with an examination subject therein, to acquire MR image data from an examination region, representing an overview image of the examination region, by implementing a flow-sensitive overview imaging sequence;in a processor supplied with said MR image data, automatically generating a three-dimensional data set;in said processor, automatically generating signal intensity profiles proceeding through said three-dimensional image data set;in said processor, automatically determining position of a blood vessel from the signal intensity profiles; andin said processor, automatically determining a position of a slice plane for implementing an MR angiography examination using said position of said blood vessel. 2. A method as claimed in claim 1 comprising, in said processor, automatically determining said position of said blood vessel in said three-dimensional data set by determining an intensity maximum in each of the signal intensities profiles, and associating the position of the blood vessel with the respective intensity maxima. 3. A method as claimed in claim 1 comprising, in said processor, automatically determining a thickness of a volume occupied by each slice plane, that completely contains at least a predetermined portion of said blood vessel. 4. A method as claimed in claim 1 comprising operating said magnetic resonance data acquisition unit to acquire said MR image data with the examination subject located on a table moving through successively different positions in said data acquisition unit, and determining the position of the slice plane at each different table position. 5. A method as claimed in claim 4 comprising determining the position of the blood vessel at a first of said table positions and at a second of said table positions, and when a discontinuity exists in a transition of the blood vessel from said first of said table positions to said second of said table positions, automatically determining the position of the blood vessel by interpolating, in said processor, the position of the blood vessel at said first of said table positions and the position of the blood vessel at said second of said table positions. 6. A method as claimed in claim 4 comprising, for each table position, determining an angle of the slice plane in said processor relative to a mean longitudinal access averaged over respective longitudinal axis of the slice planes at the different table positions and, in said processor, automatically checking whether angles of the respective slice planes lie within a predetermined angle range. 7. A method as claimed in claim 1 comprising, in said processor, determining said signal intensity profiles in an anterior-posterior direction proceeding through said examination region. 8. A method as claime... | - | A61B005/055
US20100292552 | A1 | 2010-11-18 | Analyte Monitoring Device and Methods of Use | 2010-07-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A system for processing data from a continuous analyte sensor, comprising: a continuous analyte sensor configured to continuously measure a concentration of analyte in a host and provide continuous analyte sensor data associated therewith; anda processor module configured to set a mode selected from a plurality of predetermined modes, wherein the processor module is configured to process the continuous analyte sensor data based at least in part on the mode. 3. The system of claim 2, wherein the processor module is configured to set the mode at least in part responsive to receipt of a user input. 4. The system of claim 3, wherein the system comprises one or more buttons, and wherein the processor module is configured to receive the user input by selection of one or more buttons. 5. The system of claim 3, wherein the system comprises a screen, wherein the processor module is configured to display one or more menus on the screen, and wherein the processor module is configured to receive the user input by selection of one or more items from the one or more menus. 6. The system of claim 3, wherein the system is configured to operably connect with an external system such that data can be transmitted from the external system to the system, and wherein the processor module is configured to receive the user input by a data transmission received from the external system. 7. The system of claim 6, wherein the operable connection is a wired connection. 8. The system of claim 6, wherein the operable connection is a wireless connection. 9. The system of claim 2, wherein the processor module is configured to automatically set the mode based at least in part on a comparison of data with one or more criteria. 10. The system of claim 9, further comprising a temperature sensor, wherein the data comprises data received from the temperature sensor. 11. The system of claim 2, wherein the continuous analyte sensor comprises a continuous glucose sensor, wherein the data comprises glucose sensor data, and wherein the one or more criteria comprise one or more thresholds associated with hyperglycemia and/or hypoglycemia. 12. The system of claim 2, wherein the processor module comprises programming configured to automatically set the mode at least in part responsive to an adaptive mode learning module, wherein the adaptive mode learning module is configured to process sensor data and time-corresponding mode over time and subsequently modify the automatic mode-setting programming at least in part responsive thereto. 13. The system of claim 2, wherein the system is further configured to provide an alarm responsive to the sensor data meeting one or more criteria. 14. The system of claim 13, wherein the one or more criteria are based at least in part on the mode. 15. The system of claim 13, wherein the system is configured to provide the alarm via an audible sound, visual display, vibration, alphanumeric message, and/or wireless transmission based... | Abbott Diabetes Care Inc. | A61B005/05
US20100331838 | A1 | 2010-12-30 | TRANSMURALITY CLAMP SYSTEMS AND METHODS | 2010-06-24 | Tissue clamp systems and methods assess the transmurality of a lesion created by ablation. Exemplary clamp systems for ablating tissue include a first jaw, an ablation element mounted on the first jaw, a second jaw, and a sensing element mounted on the second jaw. The sensing element can be configured to assess transmurality of a lesion created by the ablation element. Systems may also include a connection mechanism that connects the ablation element and the sensing element to a connector for connecting the ablation element and the sensing element to an ablation system. | 1. A clamp assembly for ablating tissue, comprising: a first jaw;an ablation element mounted on the first jaw;a second jaw;a sensing element mounted on the second jaw, wherein the sensing element is configured to assess transmurality of a lesion created by the ablation element; anda connection mechanism that connects the ablation element and the sensing element to a connector for connecting the ablation element and the sensing element to an ablation system. 2. A clamp assembly as in claim 1, wherein the connection mechanism comprises one or more wires. 3. A clamp assembly as in claim 1, wherein the sensing element comprises a temperature sensor. 4. A clamp assembly as in claim 1, wherein the sensing element comprises a pacing electrode. 5. A clamp assembly as in claim 1, wherein the sensing element comprises an optical sensor. 6. A clamp assembly as in claim 5, wherein the optical sensor detects tissue color change that occurs as tissues are heated to above 60° C. 7. A clamp assembly as in claim 5, wherein the optical sensor detects changes to near-field microwave that occur as tissues temperatures rise or fall. 8. A clamp assembly as in claim 1, wherein the sensing element penetrates a tissue surface of a tissue as the tissue is clamped between the first and second jaws of the clamp assembly. 9. A clamp assembly as in claim 8, wherein the sensing element comprises a temperature sensor. 10. A clamp assembly as in claim 8, wherein the sensing element comprises a pacing electrode. 11. A clamp assembly as in claim 8, wherein the sensing element comprises an electrode that senses tissue impedance. 12. A clamp assembly as in claim 1, wherein the ablation element heats tissue using radiofrequency energy. 13. A clamp assembly as in claim 1, wherein the ablation element heats tissue using microwave energy. 14. A clamp assembly as in claim 1, wherein the ablation element heats tissue using ultrasonic energy. 15. A clamp assembly as in claim 1, wherein the ablation element freezes tissue to achieve ablation. 16. A clamp assembly as in claim 8, further comprising one or more ablation elements coupled with the second jaw. 17. A system for ablating tissue, comprising: an ablation clamp comprising a first jaw, a second jaw, one or more ablation elements mounted on the first jaw, and a sensing element mounted on the second jaw that assesses transmurality of a lesion created by the ablation element;an ablation control system for controlling an ablation process provided by the ablation clamp; andmeans for connecting the ablating element and the sensing element with the ablation control system. 18. A system as in claim 17, wherein the ablation control system stops the ablation process when the sensing element carried by the second jaw indicates that a transmural lesion has been created. 19. A system as in claim 17, wherein the ablation control system stops the ablation process if a predetermined maximum time limit has been reached. 20. A clamp a... | ESTECH, Inc. (Endoscopic Technologies, Inc.) | A61B017/28 A61B018/18 A61B005/00
US20100286549 | A1 | 2010-11-11 | System and Method for Assessing Efficacy of Therapeutic Agents | 2008-12-12 | A method for assessing an effect of a therapeutic agent, comprises the steps of detecting brain electrical activity of a subject to generate a first set of brain wave data and extracting from the first set of brain wave data first data features sensitive to a neurological disorder in combination with the steps of comparing the first data features to control data to define a baseline profile of brain electropathophysiology and computing one of a first classifying score and a first discriminant score based on the baseline profile to estimate a probability that the baseline profile corresponds to a predetermined pathophysiological condition. | 1. A method for assessing an effect of a therapeutic agent, comprising the steps of: detecting brain electrical activity of a subject to generate a first set of brain wave data;extracting from the first set of brain wave data first data features sensitive to a neurological disorder; andcomparing the first data features to control data to define a baseline profile of brain electropathophysiology; andcomputing one of a first classifying score and a first discriminant score based on the baseline profile to estimate a probability that the baseline profile corresponds to a predetermined pathophysiological condition. 2. The method according to claim 1, further comprising: administering a dosage of a therapeutic agent to the subject; andafter a predetermined time interval has elapsed since the administration of the dosage, detecting brain electrical activity of the subject to generate a second set of brain wave data, the predetermined time interval being determined based on pharmocokinetic properties of the therapeutic agent. 3. The method according to claim 2, further comprising: subjecting the second set of brain wave data to spectral analysis to select second data features sensitive to a neurological disorder; andcomparing the second data features to control data to define a post-treatment profile of brain electropathophysiology; andcomputing one of a second classifying score and a second discriminant score based on the post-treatment profile to evaluate an effect of the therapeutic agent. 4. The method according to claim 1, wherein the control data includes population norm data, the method further comprising: compiling first population brain wave data from each of a first plurality of individuals, the first population brain wave data including data corresponding to the first data features;sorting the first population brain wave data based on criteria including at least one of age, medical history, gender and ethnicity of the first plurality of individuals; andselecting portions of the first population brain wave data for use as the population norm data based on a comparison of the criteria and corresponding characteristics of the subject. 5. The method according to claim 4, wherein the first plurality of individuals are selected from a group including one of individuals not suffering from the neurological disorder, individuals with a form of the neurological disorder more manageable than that of the subject and individuals substantially free of symptoms of the neurological disorder. 6. The method according to claim 4, further comprising: compiling second population brain wave data from each of a second plurality of individuals, the second population brain wave data including data corresponding to the first data features;sorting the second population brain wave data based on second criteria including at least one of a therapeutic agent used to treat the disorder of the corresponding individual and an observed therapeutic effect of the agent in t... | New York University | A61B005/0476 G06F019/00
US20100324405 | A1 | 2010-12-23 | ELECTRODE FOR ACQUIRING PHYSIOLOGICAL SIGNALS OF A RECIPIENT | 2008-07-21 | The present invention relates to an electrode for acquiring physiological signals of a recipient. Furthermore the present invention relates to a textile fabric for use in a garment to be worn by a recipient, and to a monitoring system for monitoring of physiological parameters of a recipient. In order to provide an electrode for acquiring physiological signals of a recipient, which on the one hand provides a soft, and comfortable skin contact, whilst on the other hand assures a high signal quality, an electrode (1) for acquiring physiological signals of a recipient is suggested, which comprises at least two conductive textile layers (2, 3) positioned on top of each other, wherein the first layer (2) is made of a woven material, and the second layer (3) having a working surface (4) to be brought into contact with the recipient's skin is made of a knitted material. | 1. An electrode (1) for acquiring physiological signals of a recipient, said electrode (1) comprising at least two conductive textile layers (2, 3) positioned on top of each other, wherein the first layer (2) is made of a woven material, and the second layer (3) is made of a knitted material having a working surface to be brought into contact with the recipient's skin. 2. The electrode (1) as claimed in claim 1, wherein at least one of the conductive textile layers (2, 3) are made by using a yarn, which comprises a metallic component and a synthetic component. 3. The electrode (1) as claimed in claim 2, wherein the metallic component is stainless steel or silver and the synthetic component is polyester. 4. The electrode (1) as claimed in claim 3, wherein the yarn is comprised of about 20 to about 30 weight % stainless steel and about 80 to about 70 weight % polyester. 5. The electrode (1) as claimed in claim 1, further comprising a support member (5), adapted to serve as a support for the conductive textile layers (2, 3). 6. The electrode (1) as claimed in claim 5, wherein the support member (5) has a cushion-like shape. 7. The electrode (1) as claimed in claim 5, wherein the conductive textile layers (2, 3) are stretched over the support member (5). 8. The electrode (1) as claimed in claim 5, wherein the support member (5) is flexible. 9. A textile fabric (7) for use in a garment to be worn by a recipient, said fabric (7) being adapted to serve as a carrier for an electrode (1) as claimed in claim 1. 10. The textile fabric (7) as claimed in claim 9, comprising an opening (8), said opening (8) being adapted to fit the size of the working surface (4) of the electrode (1). 11. The textile fabric (7) as claimed in claim 9, to which the electrode (1) is connected either by lockstitch sewing or by heat bonding. 12. The textile fabric (7) as claimed in claim 9, made of a non-stretchable material. 13. A monitoring system for monitoring of physiological parameters of a recipient, comprising an electrode (1) as claimed in claim 1. 14. A monitoring system for monitoring of physiological parameters of a recipient, comprising a garment comprising a textile fabric as claimed in claim 9. | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/04
US20100312089 | A1 | 2010-12-09 | METHOD AND SYSTEM FOR MARKER LOCALIZATION | 2010-05-28 | A method and system for locating a source of a signal. The source may be a wireless marker that is implanted in an object, such as a human body. The signal generated by the marker is a magnetic field. The location system uses an array of sensors to measure the magnetic field of the marker at various sensor locations. The location system compares the set of actual measurements to sets of reference measurements for various known locations within a bounding volume. Based on the comparisons, the location system identifies the set of reference measurements that most closely matches the set of actual measurements. The location system then uses sets of reference measurements for known locations near the closest known location to more accurately determine the marker location when it is not actually at one of the known locations. | 1-49. (canceled) 50. A method for locating a dipole source of a magnetic field, the dipole source being a marker implanted in or affixed to an organism, the method comprising: selecting a location associated with a set of reference measurements that most closely matches a set of actual measurements, an actual measurement being indicative of the magnetic field, a set of reference measurements being indicative of actual measurements expected when the dipole source is at the associated location; andidentifying a location associated with a set of interpolated reference measurements that most closely matches the set of actual measurements wherein the set of interpolated reference measurements is derived from sets of reference measurements. 51. The method of claim 50 wherein the identifying of the location includes searching for the location by iteratively determining how closely sets of interpolated reference measurements match the set of actual measurements. 52. The method of claim 50 wherein the marker is wireless. 53. The method of claim 50 wherein sets of reference measurements are calculated. 54. The method of claim 50 wherein the sets of reference measurements are experimentally derived. 55. The method of claim 50 wherein the locations associated with the sets of reference measurements are at grid points within a bounding box that contains the source. 56. The method of claim 50 wherein the closeness of a match between two sets of measurements is based on an evaluation of an objective function. 57. The method of claim 56 wherein the objective function is E2=sTH(HTH)−1HTs. 58. The method of claim 50 wherein the approximation is a trilinear interpolation. 59. The method of claim 50 wherein the approximation is a combination of a trilinear and a second order difference interpolation. 60. The method of claim 50 wherein the identifying of a location includes identifying a location near a candidate location wherein a derivative of an objective function is near zero. 61. The method of claim 50 wherein the identifying of a location includes calculating a spatial offset from a current location according to ρ=−C−1b. 62-71. (canceled) | - | A61B005/05
US20100277119 | A1 | 2010-11-04 | Medical Device Charging System | 2009-05-01 | A medical device charging system includes a charging cable to electrically couple with a power source and has a connector. A medical device includes a rechargeable battery electrically coupled to a medical device interface. The medical device interface and the connector on the charging cable are incompatible. An adapter includes a first interface to electrically couple with the connector on the charging cable, and a second interface to electrically couple with the medical device interface. The adapter conducts charging power from the charging cable to the medical device to charge the rechargeable battery in the medical device. | 1. A medical device charging system, comprising: a charging cable to electrically couple with a power source and includes a connector;a medical device having a rechargeable battery electrically coupled to a medical device interface, wherein the medical device interface and the connector on the charging cable are incompatible; andan adapter having a first interface to electrically couple with the connector on the charging cable, and a second interface to electrically couple with the medical device interface, wherein the adapter conducts charging power from the charging cable to the medical device to charge the rechargeable battery in the medical device. 2. The system of claim 1, wherein the charging cable includes a power plug to electrically couple to a power outlet. 3. The system of claim 1, wherein the charging cable includes a Universal Serial Bus (USB) plug to electrically couple to a USB port. 4. The system of claim 3, wherein the USB port is a communications link between the medical device and another device. 5. The system of claim 1, wherein the charging cable includes a vehicle connector to electrically couple to a vehicle power outlet. 6. The system of claim 1, wherein the connector on the charging cable is selected from the group consisting of a USB Standard-A plug, a USB Standard-B plug, a USB Mini-A plug, a USB Mini-B plug, a USB Mini-AB plug, a USB Micro-A plug, a USB Micro-B plug, and a USB Micro-AB plug. 7. The system of claim 1, wherein the power source is a battery. 8. The system of claim 1, wherein the medical device interface is adapted to couple with a data cable to transmit and receive data via the medical device interface. 9. The system of claim 8, wherein the data cable is the charging cable. 10. The system of claim 1, wherein the adapter further includes circuitry to convert charging power from a first power format received from the power source into a second power format to charge the rechargeable battery in the medical device. 11. The system of claim 1, wherein the medical device is selected from the group consisting of a glucose sensor, a blood glucose meter, a medical device controller, a medical device programmer, a medication delivery device, an insulin infusion device, and a sensor transmitter. 12. The system of claim 1, wherein the charging cable is a universal power charger. 13. The system of claim 1, wherein the adapter includes an indication device to communicate status of how the rechargeable battery is charging. 14. A medical device charging system, comprising: a charging cable to electrically couple with a power source and includes a Universal Serial Bus (USB) connector;a medical device having a rechargeable battery electrically coupled to a medical device interface, wherein the medical device interface and the USB connector on the charging cable are incompatible; andan adapter having a USB interface to electrically couple with the USB connector on the charging cable, and an adapter interface to ... | Medtronic MiniMed, Inc. | H02J007/00 A61B005/00
US20100292562 | A1 | 2010-11-18 | Method for generating MR (Magnetic resonance) images of a moving partial area of an object | 2010-07-27 | A method for generating MR (magnetic resonance) images of a moving partial area of an object with a repeating motion sequence over comparable motion states, wherein an MR data set, which is encoded for generating an individual MR image of the object, is provided for each motion state from a plurality of successive individual MR measurements with shorter time intervals than a repetition rate of the motion sequence, and wherein at least one navigator data point is generated for each individual MR measurement as an indicator for the comparability of several motion states, is characterized in that a position of the partial area is determined for each individual MR image, from which a function f(t) of the time shift of the position is determined, and the measuring data of the individual MR measurement is phase-corrected in correspondence with its respective motion state using the function f(t) to keep the position of the partial area in a spatially stationary state. This permits observation of a moving partial area irrespective of its motion state, such that rapid changes within the partial area can be observed. | 1. A method for generating MR (magnetic resonance) images of a moving partial area of an object, the partial area exercising a repeating motion sequence through a plurality of comparable motion states, the method comprising the steps of: a) recording a plurality of successive individual MR measurements in time intervals which are shorter than a repetition rate of the motion sequence, each individual MR measurement having one of a sequence of phase encoding steps;b) generating at least one navigator data point for each individual MR measurement as an indicator of one of the plurality of motion states;c) analyzing step b) to construct an individual MR image of each motion state using individual MR measurements corresponding to that motion state;d) determining a one, two, or three dimensional position of the partial area for each individual MR image constructed in step c)e) determining a function f(t) of a time shift in the positions determined in step d);f) phase correcting measuring data of the individual MR measurements in correspondence with respective motion states thereof using the function f(t) of step e) to keep positions of the partial area in a substantially stationary state; andg) generating a time course type of film of the substantially stationary partial area using the phase corrected measuring data of step f). 2. The method of claim 1, wherein the film has a time scale which is much larger than a time scale of the motion sequence. 3. The method of claim 1, wherein individual MR measurements of the film are acquired during and/or after administering an active substance or a contrast medium. 4. The method of claim 1, wherein the film depicts a non-periodic motion sequence or a peristalsis. 5. The method of claim 1, wherein the method is performed for several different motion sequences which overlap in time. 6. The method of claim 5, wherein the several different motion sequences comprise a heart beat and a breathing of a living entity. 7. The method of claim 1, wherein the method is a nuclear magnetic resonance (NMR) method. 8. The method of claim 1, wherein the indicator of step b) comprises at least one coherent area of successive data points within the individual MR measurement, this coherent area being identically repeated with respect to irradiated RF (radio frequency) pulses and switched gradients for all individual MR measurements within a respective MR measuring sequence. 9. The method of claim 8, wherein the method is a nuclear magnetic resonance method (NMR), the coherent area comprising a partial area, in which a rephasing gradient or a slice selection gradient is applied, thereby rephasing a nuclear spin system in this partial area. 10. The method of claim 7, wherein further individual MR measurements are performed with a second type of nucleus simultaneously with individual MR measurements of a first type of nucleus to determine the indicator. | Bruker BioSpin MRI GmbH | A61B005/055
US20100324408 | A1 | 2010-12-23 | IMPLANTABLE SYSTEM FOR DETERMINING THE ACCOMMODATION REQUIREMENT BY MEASURING THE EYEBALL ORIENTATION USING AN EXTERNAL MAGNETIC FIELD | 2008-02-18 | The present invention relates to an implantable system for determining the accommodation requirement in an artificial accommodation system by measuring the eyeball orientation using an external magnetic field, comprising a) at least one optical system (3),b) at least one data acquisition system (8) which does not contact the ciliary muscle and has means for measuring a spatial orientation of both eyeballs as a physical control signal for the accommodation requirement,c) at least one data processing system (9) for generating an actuating signal for the optical system (3) from the measured physical control signals,d) at least one energy supply system (10), ande) one fixing system, wherein the system in each case has means for measuring a magnetic field in both eyes and provision is made for transfer means for mutual information exchange between the means. | 1. An implantable system for determining the accommodation requirement in an artificial accommodation system by measuring the eyeball orientation using an external magnetic field, comprising a) at least one optical system (3),b) at least one data acquisition system (8) which does not contact the ciliary muscle and has means for measuring a spatial orientation of both eyeballs as a physical control signal for the accommodation requirement,c) at least one data processing system (9) for generating an actuating signal for the optical system (3) from the measured physical control signals,d) at least one energy supply system (10), ande) one fixing system,wherein the system in each case has means for measuring a magnetic field in both eyes and provision is made for transfer means for mutual information exchange between the means. 2. The system as claimed in claim 1, wherein the means for measuring a magnetic field are in each case fixedly connected to the respectively associated eyeball or are inserted therein. 3. The system as claimed in, claim 1, wherein the magnetic field is formed by a terrestrial magnetic field. 4. The system as claimed in claim 1, wherein the magnetic field is formed by a magnetic field that is fixed with respect to the head. 5. The system as claimed in claim 1, wherein the means for measuring the magnetic field respectively comprise a compass sensor or two magnetic field sensors which span a plane with their measurement directions. 6. The system as claimed in claim 5, wherein the compass sensors or the planes in both eyes are aligned parallel with respect to a straight line, which connects the pivotal points of both eyes, and are mutually parallel. 7. The system as claimed in, claim 5, wherein the means for measuring the magnetic field are formed by magnetoresistive or Hall sensors. 8. The system as claimed in claim 1, wherein the measurement means can be used as transfer means and hence as data receivers for alternating fields and signals. | - | A61B005/05
US07853456 | B2 | 2010-12-14 | Systems and methods for risk stratification of patient populations | 2005-03-04 | A statistical processing system includes a server operably configured with program instructions implementing a plurality of statistical models to at least one of (a) predict a health outcome based on questionnaire responses, (b) assist a patient's choice of therapeutic modality based on questionnaire responses, and (c) assess a health risk or status based on questionnaire responses. Also provided is a research agency communicating with the server and contracted to provide the statistical models using a visual interface communicated by the server. The server is configured to analyze requests received from users relating to a plurality of said statistical models to reduce redundancy in requests for patient data. | 1. A method of utilizing a statistical health model, the method comprising: storing in computer readable memory associated with a statistical health model translation system a first plurality of statistical health models from respective sources in a plurality of different sources, wherein at least one statistical health model in the first plurality of statistical health models from the respective sources is a medical prognostic risk stratification model and/or a medical prognostic outcomes prediction model in the form of at least one of a: linear model,a generalized linear model,a cumulative multinomial model,a generalized multinomial model,a proportional hazard model;providing via the statistical health model translation system one or more user interfaces including a plurality of fields that enable one or more users to specify for the at least one statistical health model: an outcome predicted by the at least one statistical health model;one or more outcome predictors;a mathematical relationship between: the outcome predicted by the at least one statistical health model, andthe one or more outcome predictors;providing via the statistical health model translation system one or more user interfaces that enable one or more users to select models from the first plurality of statistical health models for execution;receiving a user selection of one or more of the first plurality of statistical health models;automatically generating data-input interfaces for collecting patient-specific predictors utilized when executing the selected one or more of the first plurality of statistical health models based at least in part on the one or more outcome predictors;processing, via a computing device, the one or more outcome predictors and information regarding a patient received via the automatically generated data-input interfaces, where the information received via the automatically generated data-input interfaces includes one or more patient medical conditions, using the selected one or more of the first plurality of statistical health models; andproviding, via the computing device, an output from the selected one or more of the first plurality of statistical health models. 2. The method of claim 1, the method further comprising providing via the statistical health model translation system a user interface including a plurality of fields configured to: receive one or more predictor coefficients; andreceive one or more predictor coefficient covariances for calculating confidence intervals. 3. The method of claim 1, wherein the first plurality of statistical health models includes a statistical health model configured to determine a statistical outcome of a medical procedure, medical treatment or intervention, and/or medical condition with respect to the patient. 4. The method of claim 3, wherein at least one outcome includes risks associated with genetic factors or biochemical factors. 5. The method of claim 1, the method further comprising: receiving an u... | Health Outcomes Sciences, LLC | G06Q010/00 G06F007/60 G06F009/45 G06F019/00 G06G007/58 A61B005/00
US07860557 | B2 | 2010-12-28 | Radiolucent chest assembly | 2005-04-13 | A wireless monitoring system and, more particularly, a wireless monitoring system for detecting and transmitting physiological data. The present invention detects physiological data relating to a patients cardiac activity and respiration rate and transmits the data to a remote base station via telemetry. The base station processes the data so that the data can be displayed by an ECG monitor. | 1. A chest assembly for collecting physiological data, comprising: a one-piece flexible circuit comprising: a substantially radiolucent base layer having a first side and a second side, said first side containing a plurality of substantially radiolucent electrical traces, said electrical traces configured to be coupled to a plurality of electrodes, said base layer disposed adjacent at least one electrically insulating layer;at least one aperture formed in said base layer; andat least one electrode housing disposed adjacent said aperture in said base layer and comprising a resilient portion defining a female void for receiving a portion of an electrode therein. 2. The assembly of claim 1 wherein said base layer is disposed between a first electrically insulating layer and a second electrically insulating layer. 3. The assembly of claim 1 wherein said electrical traces connect to at least one electrode connector and an assembly connector. 4. The assembly of claim 3 wherein each of said plurality of electrical traces form and define a separate electrode. 5. The assembly of claim 1 wherein the electrical traces are printed ink or anisotropic adhesive. 6. The assembly of claim 1 wherein the electrical traces are silver epoxy. 7. The assembly of claim 1 further comprising a substantially radiolucent shielding layer. 8. The assembly of claim 7 wherein said substantially radiolucent shielding layer comprises at least one layer of dielectric material. 9. The assembly of claim 1 wherein one side of the assembly further comprises a fabric layer. 10. The assembly of claim 1 wherein said aperture forms at least one flap for engaging and securing the male portion of an electrode positioned through the aperture. 11. The assembly of claim 1 wherein said female void accommodates a plurality of different sized male portions of conventional electrodes. 12. The assembly of claim 1 wherein said elastomeric member securely retains the male portions of conventional electrodes. 13. The assembly of claim 1 wherein said elastomeric member removably retains the male portions of conventional electrodes. 14. The assembly of claim 1 wherein the assembly connects to a telemetry transceiver via an adapter assembly. 15. The assembly of claim 1 wherein the assembly is directly connected to an electrocardiogram monitor via an adapter assembly. 16. The assembly of claim 1 wherein the assembly is disposable. 17. The assembly of claim 1 wherein the assembly connector comprises a sense pin for turning electronics on. 18. A radiolucent precordial assembly for collecting physiological data comprising: a substantially radiolucent base layer having a first side and a second side, the first side attached to a plurality of radiolucent electrically conductive elements or traces, the electrically conductive elements or traces configured to be removably coupled to a plurality of precordial electrodes through a plurality of corresponding electrode connectors; said electrically ... | LifeSync Corporation | A61B005/04
US07840248 | B2 | 2010-11-23 | Online source reconstruction for eeg/meg and ecg/mcg | 2004-01-27 | In one embodiment, the present invention includes an EEG system comprising of sensors which are utilized to gather an electromagnetic signal from a patient; a signal processing system; and a computer system. The computer system is configured to support multiple threads of execution. The computer system initiates a first thread of execution, a measurement module, where the data is filtered and in some cases averaged. Typically, the data is filtered for a particular latency period. Once the data has been filtered and averaged, the result is given to a second thread of execution, the source reconstruction module, which then proceeds to generate a source reconstruction for the trial. The measurement module then acquires and processes new electromagnetic data from a new trial while the source reconstruction is being performed on the most recent trial. | 1. A method of evaluation comprising: acquiring an electromagnetic physiological signal with a sensor on a subject;performing a source reconstruction of the electromagnetic physiological signal via a signal processing circuit in communication with said sensor, with said circuit including a computer system with a processor and a memory, the computer system being configured to support multiple parallel threads of execution with one thread being a measurement module and a second thread being a source reconstruction module;feeding source reconstruction data through a feedback loop to the measurement module to modify the analysis of a new electromagnetic physiological signal, wherein the new electromagnetic physiological signal is being acquired while the source reconstruction is being performed; anddisplaying source reconstruction results overlayed onto anatomical data of said patient. 2. The method of claim 1, wherein the sensor acquires MEG data. 3. The method of claim 1, wherein the sensor acquires EEG data. 4. The method of claim 1, wherein the sensor acquires ECG data. 5. The method of claim 1, wherein the sensor acquires MCG data. | Compumedics Limited | A61B005/05
US07850619 | B2 | 2010-12-14 | Apparatus and method for breathing pattern determination using a non-contact microphone | 2005-07-21 | A method is provided for analyzing respiration of a subject. Using a non-contact microphone, a raw signal indicative of air-flow sounds of the respiration is generated. The raw signal is analyzed to determine a first set of one or more parameters of the respiration. An algorithm is applied to the first set of parameters to derive a second set of one or more estimated parameters of the respiration that are not generally directly measurable in the raw signal. Other embodiments are also described. | 1. Apparatus for analyzing respiration of a subject, comprising: a non-contact microphone that does not contact the subject, adapted to generate a raw signal representing airflow sounds that are generated at the microphone by airflow of the respiration; anda control unit, adapted to: analyze the airflow sounds, by deriving an expiratory airflow sound signal from the raw signal, and analyzing the expiratory airflow sound signal, to determine a first set of one or more parameters of the respiration, andapply an algorithm to the first set of parameters to derive a second set of one or more estimated parameters of the respiration, including an inspiration time of the subject. 2. The apparatus according to claim 1, wherein the control unit is adapted to set the inspiration time equal to a phenomenological function of the first set of one or more parameters. 3. The apparatus according to claim 1, wherein the first set of parameters includes a measure of breathing amplitude of the respiration, and wherein the control unit is adapted to analyze the airflow sounds by integrating the airflow sounds for a breath of the respiration to determine the measure of the breathing amplitude. 4. The apparatus according to claim 1, wherein the control unit is adapted to analyze a parameter selected from the group consisting of: the first and second sets of parameters, to derive at least one additional breathing-related parameter of the subject selected from the group consisting of: breathing amplitude, a geometrical property of airflow of the subject, a characteristic of the airflow indicative of pursed lips breathing, a characteristic of the breathing indicative of relaxed breathing, a characteristic of the breathing indicative of passive elastic recoil of lungs of the subject, a characteristic of breathing with effort, and a characteristic of breathing during which the lungs of the subject undergo a functional change. 5. The apparatus according to claim 1, comprising an output generator, wherein the control unit is adapted to drive the output generator to guide the subject to perform breathing in a plurality of respiration phases determined at least in part responsively to the inspiration time. 6. The apparatus according to claim 1, wherein the non-contact microphone comprises a non-contact microphone of a consumer electronics device capable of performing at least one function that does not facilitate analyzing respiration of the subject. 7. The apparatus according to claim 1, wherein the non-contact microphone is integrated into a medical device in fluid communication with respiration-related airflow of the subject. 8. The apparatus according to claim 7, wherein the medical device comprises a drug inhaler. 9. The apparatus according to claim 1, comprising an output generator, wherein the control unit is adapted to drive the output generator to generate a real-time indication for the subject that indicates whether expiration has been detected. 10. The appar... | Intercure Ltd. | A61B005/08
US08280749 | B2 | 2012-10-02 | 12-lead ECG measurement and report editing system | 2009-12-24 | The present invention provides a real-time interactive information system for measurement of 12-lead electrocardiogram waveforms and report editing, comprising a server providing cross-platform web page browsing for user to input commands or to review ECG reports and to edit or to measure information from patient further, a device for supporting internet interacting protocol which is a WEB SERVICE allowing for receiving commands issued by the server of cross-platform web page browsing and giving the server the feedback of processed ECG information by assistance from the database device, and a database device allowing for accessing patient's information to the internet interacting protocol. | 1. A real-time interactive information system for measurement of 12-lead electrocardiogram (ECG) waveforms and report editing comprising: (a) a server providing cross-platform web page browsing for a user to input commands or to review ECG reports and to edit or to measure information from patient further;(b) a device for supporting internet interacting protocol allowing for receiving commands issued by the server of cross-platform web page browsing and giving the server the feedback of processed ECG information by assistance from a database device;(c) the database device allowing for accessing patient's information to the internet interacting protocol, wherein a patient's information of the ECG reports is transformed into an XML file which comprises the ECG information and noise is removed from a ECG signal in the XML file; and wherein the server providing cross-platform web page browsing further comprises a function storing means of CLIENTBIN, providing function of computing a waveform height and a time period width or performing message accessing and editing through assistance of a dynamic vernier with customized scale. 2. The system of claim 1, wherein the server providing cross-platform web page browsing comprises: (a) a device of a hypertext markup language (HTML) suite to process web page creation and performance of information on the web browser; and(b) a device of a dynamic language suite used for improving the function and convenience of dynamic web page processing. 3. The system of claim 1, wherein the server providing cross-platform web page browsing is a server providing technical effects from SILVERLIGHT web. 4. The system of claim 1, wherein the device for supporting internet interacting protocol uses a communication method of software from WEB SERVICE. 5. The system of claim 1, wherein the patient's information in the database device is accessed from hospital information system (HIS). 6. The system of claim 1, wherein an application of the database device comprises outpatient appointment or emergency on-site registration, price calculation of physician order, accessing and printing medication bags and medical charts, bed management, control of length of stay, price calculation from nursing station, hospital billing or submission of outpatient visit and hospitalization. 7. The system of claim 2, wherein the dynamic language suite is ASP.NET suite. 8. The system of claim 1, wherein the device for supporting internet interacting protocol further comprises a device of wavelet transform used for removing noise from ECG signal. 9. The system of claim 1, wherein the user is health care professionals who provide service for patients. 10. The system of claim 9, wherein the health care professional is a physician, a nurse, a pharmacist, a medical technologist or a radiologic technologist. 11. The system of claim 5, wherein the HIS facilitates logging in, inquiring, browsing or editing patient's information by computers in or outsi... | Yuan Ze University | G06Q010/00
US08105208 | B2 | 2012-01-31 | Portable fitness monitoring systems with displays and applications thereof | 2009-05-18 | Portable fitness monitoring systems with displays, and applications thereof, are disclosed. In an embodiment, a method of providing training feedback to an individual using a heart rate sensor and a display module supported by the individual during a physical activity, the method including the steps of: (a) determining a maximum heart rate value for the individual; (b) defining a heart rate zone as a range of heart rate values that correspond to a range of percentages of the maximum heart rate value; (c) associating a color with the heart rate zone; (d) wirelessly transmitting heart rate data from the heart rate sensor to the display module during the physical activity; and (e) visually displaying the color associated with the heart rate zone to the individual on the display module during the physical activity in response to the heart rate data. | 1. A method of providing training feedback to an individual using a portable fitness device and a sensor in communication with the fitness device, the method comprising the steps of: (a) defining a range of performance parameter values;(b) wirelessly transmitting performance data from the sensor to the fitness device during a physical activity;(c) requesting feedback from the individual about the physical activity; and(d) selectively adjusting the range of performance parameter values based on feedback received from the individual. 2. The method of claim 1, wherein step (c) includes the step of querying the individual about the difficulty of the physical activity. 3. The method of claim 1, further comprising the step (e) of selectively adjusting the range of performance parameter values based on the performance data. 4. The method of claim 1, wherein step (c) occurs via a remote computer. 5. The method of claim 1, wherein the range of performance parameter values is a range of heart rate values. 6. The method of claim 5, wherein the performance data is heart rate data. 7. The method of claim 1, wherein the range of performance parameter values is a range of speed values. 8. The method of claim 1, wherein the range of performance parameter values is a range of pace values. 9. A method of providing training feedback to an individual using a fitness monitoring system including a portable fitness device, a sensor in communication with the portable fitness device, and a computer, the method comprising the steps of: (a) defining a plurality of performance zones, wherein each performance zone comprises a range of heart rate values;(b) wirelessly transmitting heart rate data from the sensor to the portable fitness device during a physical activity;(c) requesting feedback from the individual about the physical activity via the computer;(d) receiving feedback from the individual about the physical activity via the computer; and(e) adjusting the range of heart rate values of at least one of the performance zones based on the feedback from the individual. 10. The method of claim 9, wherein step (a) comprises the steps of: determining a maximum heart rate for the individual; and defining the plurality of performance zones based on ranges of percentages of the maximum heart rate. 11. The method of claim 9, further comprising the step of displaying a visual output indicative of at least one performance zone on the portable fitness device during the physical activity. 12. The method of claim 9, wherein the portable fitness device and the computer are separate devices. 13. The method of claim 9, wherein step (c) includes the step of querying the individual about the difficulty of the physical activity. | adidas AG | A63B071/00
US08727985 | B2 | 2014-05-20 | Medical system for providing external trigger signal | 2009-12-21 | There are disclosed embodiments for a medical system capable of providing an external trigger signal. The medical system comprises at least one of an ECG signal providing unit to provide an ECG signal indicative of an electrical change in a target object, a reference signal providing unit to provide at least one reference signal for triggering an operation of the medical system, and a medical image providing unit to provide a medical image of the target object. The medical system may further comprise a processor to form an external trigger signal using at least one of the ECG signal, reference signal and medical image. The medical system may further comprise a signal interface unit to transmit the external trigger signal to an external device coupled to the medical system. | 1. An ultrasound system, comprising: an ultrasound image provider configured to provide an ultrasound image of a target object;a processor configured to analyze the ultrasound image and form an external trigger signal including information on an operating condition and an operating timing of the ultrasound system when a predetermined pattern is detected based on a result of the analysis; anda signal interface unit configured to transmit the external trigger signal to the external device coupled to the ultrasound system,wherein the ultrasound system is synchronized with the external device based on the external trigger signal. 2. The ultrasound system of claim 1, further comprising: an ECG signal provider configured to provide an ECG signal indicative of an electrical change in the target object. 3. The ultrasound system of claim 2, wherein the processor is further configured to analyze the ECG signal and form the external trigger signal each time a predetermined ECG signal pattern is detected. 4. The ultrasound system of claim 1, further comprising: a reference signal provider configured to provide a reference signal for triggering an operation of the ultrasound system; andan ultrasound data former configured to form ultrasound data of the target object. 5. The ultrasound system of claim 4, wherein the reference signal includes: a first reference signal for triggering an operation of transmitting an ultrasound signal to the target object; anda second reference signal for triggering an operation of forming the medical image using the ultrasound data. 6. The ultrasound system of claim 4, wherein the processor is further configured to form the external trigger signal each time the reference signal is provided by the reference signal provider. | Medison Co., Ltd | A61B008/00
US08369930 | B2 | 2013-02-05 | MRI-guided devices and MRI-guided interventional systems that can track and generate dynamic visualizations of the devices in near real time | 2010-06-16 | An MRI-compatible catheter includes an elongated flexible shaft having opposite distal and proximal end portions. A handle is attached to the proximal end portion and includes an actuator in communication with the shaft distal end portion that is configured to articulate the shaft distal end portion. The distal end portion of the shaft may include an ablation tip and includes at least one RF tracking coil positioned adjacent the ablation tip that is electrically connected to an MRI scanner. The at least one RF tracking coil is electrically connected to a circuit that reduces coupling when the at least one RF tracking coil is exposed to an MRI environment. Each RF tracking coil is a 1-10 turn solenoid coil, and has a length along the longitudinal direction of the catheter of between about 0.25 mm and about 4 mm. | 1. An MRI-compatible catheter, comprising: an elongated flexible shaft having a distal end portion, and an opposite proximal end portion;a handle attached to the proximal end portion, wherein the handle includes an electrical connector interface configured to be in electrical communication with an MRI scanner;an ablation tip at the shaft distal end portion, wherein an RF conductor extends longitudinally within the shaft to the electrical connector interface at the handle and that connects the ablation tip to an RF generator; andat least one RF tracking coil positioned adjacent the distal end portion, wherein the at least one tracking coil includes a conductive lead extending between the at least one RF tracking coil and the electrical connector interface and configured to electrically connect the at least one tracking coil to an MRI scanner, wherein the conductive lead has a length sufficient to define an odd harmonic/multiple of a quarter wavelength of the operational frequency of the MRI Scanner. 2. The device of claim 1, further comprising a thermocouple at the shaft distal end portion. 3. The device of claim 1, wherein the at least one RF tracking coil is electrically connected to a circuit that reduces coupling when the at least one RF tracking coil is exposed to an MRI environment. 4. The device of claim 1, wherein the at least one RF tracking coil comprises first and second RF tracking coils in adjacent spaced-apart relationship. 5. The device of claim 1, wherein the at least one RF tracking coil is a 1-10 turn solenoid coil. 6. The device of claim 1, wherein the at least one RF tracking coil has a length along the longitudinal direction of the catheter of between about 0.25 mm and about 4 mm 7. The device of claim 1, wherein the conductive lead is a coaxial cable. 8. The device of claim 1, wherein the at least one tracking coil comprises at least three spaced apart tracking coils, with at least two of the tracking coils being closely spaced together in fixed relationship on a substantially rigid distal end portion of the shaft. 9. The device of claim 1, wherein the RF conductor includes a series of back and forth segments along its length. 10. The device of claim 1, wherein the conductive lead includes a series of back and forth segments along its length. 11. The device of claim 1, wherein the at least one RF tracking coil is recessed within the catheter shaft. 12. The device of claim 11, wherein a layer of MRI-compatible material overlies the at least one RF tracking coil and is substantially flush with an outer surface of the catheter shaft. 13. The device of claim 1, wherein the shaft distal end portion comprises a biasing member that is configured to urge the shaft distal end portion to a non-articulated position. 14. The device of claim 13, wherein the biasing member is non-metallic. 15. The device of claim 1, further comprising at least one fluid exit port at the shaft distal end portion, wherein the at least one fluid... | MRI Interventions, Inc. | A61B005/05
US08529599 | B2 | 2013-09-10 | Tissue tension detection system | 2010-08-17 | An apparatus for indicating the amount of tension applied to separate internal organ sections includes a housing, a shaft extending from the housing, and a fastener applier. The fastener applier is supported on the shaft. The fastener applier includes an anvil assembly and a cartridge assembly. The anvil assembly is positionable between first and second positions relative to the cartridge assembly. The anvil assembly and the cartridge assembly include a tissue detection device positioned to contact separate internal organ sections to generate a signal indicative of an amount of tension between the organ sections as the anvil assembly moves between the first and second positions. The tissue detection device is operably coupled to a gauge configured and dimensioned to generate an output based upon the signal. The output indicates the amount of tension between the organ sections. | 1. An apparatus for indicating the amount of tension applied to separate internal organ sections comprising: a housing;a shaft extending from the housing; anda fastener applier supported on the shaft, the fastener applier including an anvil assembly and a cartridge assembly, the anvil assembly being positionable between first and second positions relative to the cartridge assembly, the anvil assembly and the cartridge assembly including a tissue detection device positioned to contact separate internal organ sections to generate a signal indicative of an amount of tension between the organ sections as the anvil assembly moves between the first and second positions, the tissue detection device being operably coupled to a gauge configured and dimensioned to generate an output based upon the signal, the output indicating the amount of tension between the organ sections, the tissue detection device including a contact supported on one of the anvil assembly and the cartridge assembly, the contact being selectively movable relative to the respective anvil or cartridge assembly upon which the contact is supported. 2. The apparatus according to claim 1, wherein the tension detection device includes a first contact supported on the anvil assembly and a second contact supported on the cartridge assembly, the first contact being selectively movable relative to the anvil assembly and the second contact being selectively movable relative to the cartridge assembly. 3. The apparatus according to claim 1, wherein the contact is positioned to extend through at least one opening formed in one of the anvil assembly and the cartridge assembly. 4. The apparatus according to claim 1, wherein the shaft defines a longitudinal axis that extends distally through the shaft, wherein the contact is selectively movable between extended and retracted positions in a direction substantially transverse to the longitudinal axis. 5. The apparatus according to claim 4, wherein the contact extends or retracts through the at least one opening in response to movement of the anvil assembly between the first and second positions. 6. The apparatus according to claim 1, wherein the housing includes a switch that extends and retracts the contact upon the actuation of the switch. 7. The apparatus according to claim 1, wherein the contact is connected to the gauge by at least one wire. 8. The apparatus according to claim 1, wherein the anvil assembly is configured to move a first organ section towards a second organ section as the anvil assembly moves between the first and second positions such that an amount of tension is generated between the first and second organ sections. 9. The apparatus according to claim 1, wherein in the first position, the anvil assembly is spaced apart from the cartridge assembly, and in the second position, the anvil assembly is in close cooperative alignment with cartridge assembly. 10. The apparatus according to claim 1, wherein the gauge is supported wi... | Covidien LP | A61B017/068
US08260408 | B2 | 2012-09-04 | Post-exercise arrhythmia detection | 2009-06-25 | Post-exercise arrhythmias are detected by an implantable medical device. In some aspects, post-exercise arrhythmia may be prognostic of a worsening cardiovascular condition. Thus, the detection of post-exercise arrhythmia may be used as an indicator for adjusting the therapy prescribed for a patient. In some aspects post-exercise arrhythmia are detected if a patient is exercising at a level that equals or exceeds a threshold exercise level. In some aspects, therapy for a patient is modified if the detected post-exercise arrhythmia exceeds a threshold arrhythmia level. In some aspects therapy for a patient is modified if ischemia is detected in conjunction with post-exercise arrhythmia. | 1. A method of detection for an implantable cardiac device, comprising: monitoring exercise by a patient; comparing the monitored exercise with an exercise threshold; detecting an end of the exercise;detecting arrhythmia during a period of time after the end of the exercise if the monitored exercise equals or exceeds the exercise threshold;comparing the detected arrhythmia with an arrhythmia threshold;generating an indication of post-exercise arrhythmia if the detected arrhythmia equals or exceeds the arrhythmia threshold; andadapting therapy for the patient based on the indication of post-exercise arrhythmia. 2. The method of claim 1, wherein the exercise threshold corresponds to a level of exercise that results in post-exercise arrhythmia data being prognostic of cardiovascular mortality. 3. The method of claim 1, wherein the arrhythmia threshold corresponds to an arrhythmia that is prognostic of cardiovascular mortality. 4. The method of claim 1, further comprising determining an ischemic condition of the patient, wherein the arrhythmia threshold is adapted based on the ischemic condition. 5. The method of claim 1, further comprising determining an ischemic condition of the patient, wherein the detection of the arrhythmia is based on the ischemic condition. 6. The method of claim 1, further comprising: determining an ischemic condition of the patient; anddetermining, based on the ischemic condition, whether to report the detected arrhythmia. 7. The method of claim 1, further comprising: determining an ischemic condition of the patient; andtransmitting at least one indication of the detected arrhythmia and the determined ischemic condition to a monitor that is external to the patient. 8. The method of claim 1, wherein the detected arrhythmia comprises one or more ventricular contractions. 9. The method of claim 1, wherein the detected arrhythmia comprises sustained ventricular tachycardia. 10. The method of claim 1, wherein the detected arrhythmia comprises one or more atrial contractions. 11. An implantable cardiac device, comprising: an activity sensing circuit configured to monitor exercise by a patient;an exercise detector configured to compare the monitored exercise with an exercise threshold, and further configured to detect an end of the exercise;an arrhythmia detector configured to detect arrhythmia during a period of time after the end of the exercise if the monitored exercise equals or exceeds the exercise threshold;wherein the arrhythmia detector is further configured to compare the detected arrhythmia with an arrhythmia threshold, the apparatus further comprising;an indication generator configured to generate an indication of post-exercise arrhythmia if the detected arrhythmia equals or exceeds the arrhythmia threshold; anda therapy module configured to adapt therapy for the patient based on the indication of post-exercise arrhythmia. 12. The apparatus of claim 11, wherein the exercise threshold corresponds to a level of ... | Pacesetter, Inc. | A61B005/0452 A61B005/0464
US08114021 | B2 | 2012-02-14 | Body-associated receiver and method | 2009-12-15 | Receivers, which may be external or implantable, are provided. Aspects of receivers of the invention include the presence of one or more of: a high power-low power module; an intermediary module; a power supply module configured to activate and deactivate one or more power supplies to a high power processing block; a serial peripheral interface bus connecting master and slave blocks; and a multi-purpose connector. Receivers of the invention may be configured to receive a conductively transmitted signal. Also provided are systems that include the receivers, as well as methods of using the same. Additionally systems and methods are disclosed for using a receiver for coordinating with dosage delivery systems. | 1. A physiological communication receiver that can be removably secured to a subject's skin, the receiver comprising: a housing;a power source secured within the housing;a processing unit electrically coupled to the power source and secured within the housing; andat least two electrodes electrically coupled to the processing unit and secured to the perimeter of the housing such that the electrodes will contact the subject's skin, wherein the processing unit is configured to receive data encoded in an electrical alternating current in a frequency range of 3 kHz to 300 kHz conducted through the body tissue of the subject, wherein the at least two electrodes are configured to detect a differential voltage therebetween corresponding to the electrical alternating current conducted through the body tissue of the subject intervening the at least two electrodes, and wherein the electrical alternating current is associated with a device associated with the subject. 2. The receiver of claim 1, wherein the processing unit is configured to receive encoded physiological information as a physiological signal at a frequency greater than 0 Hz and less than 5 kHz conducted through the body tissue of the subject, wherein the at least two electrodes are configured to detect a differential voltage therebetween corresponding to the physiological signal conducted through the body tissue of the subject intervening the at least two electrodes, and wherein the physiological signal is associated with the subject's physiology. 3. The receiver of claim 2, further comprising a communication module electrically coupled to the processing unit, wherein the communication module allows for wireless communication between the receiver and a device external to the subject such that the receiver can provide the physiological information to the device external to the subject. 4. The receiver of claim 3, wherein the communication module is configured to communicate using a frequency hopping spread spectrum communication protocol. 5. The receiver of claim 2, further comprising a communication module electrically coupled to the processing unit, wherein the communication module allows for wireless communication between the receiver and a device external to the subject such that the receiver can provide control information to the device external to the body and wherein the control information is derived from the data encoded in the electrical alternating current provided by the device associated with the subject. 6. The receiver of claim 5, wherein the device external to the subject is a control unit associated with a pharmaceutical delivery system that includes a pharmaceutical product and can alter the delivery dosage of a pharmaceutical product. 7. The receiver of claim 6, wherein the pharmaceutical delivery system comprises a fluid containment unit having a chamber configured to contain a fluid and a plunger secured to the fluid containment unit, wherein the plunger is controlled... | Proteus Biomedical, Inc. | A61M005/20 A61B005/00
US08424716 | B2 | 2013-04-23 | Wristband applicator | 2008-09-18 | The present invention concerns apparatus for efficiently dispensing and applying a wristband identification bracelet that is meant for attachment around an extremity of an animal or person. The apparatus is provided with a supply of identity bands that are adapted to become a bracelet by having a first portion of the band affixed to a second portion to form a continuous loop around the extremity. The extremity is positioned within a receptacle in the apparatus, permitting the identification band to be wrapped closely around the extremity by the band fixing mechanism of the apparatus. The identification band is fixed to itself to form a bracelet, so as to prevent its removal from the extremity without damaging the integrity of the bracelet. The identification bracelet comprises an integral RFID inlay component, and particularly an RFID component whereby all of the RFID circuit is realised in low cost fully printed silicon ink TFT semiconductor form as opposed to a conventional integrated circuit (IC) chip used in prior art RFID inlays. | 1. An apparatus for dispensing and applying an identification bracelet that is attachable around an extremity of an animal, including a person, the apparatus being provided with a supply of identity bands that are adapted to become a bracelet by having a first portion of the band affixed to a second portion of the band to form a continuous loop around the extremity, characterised in that the extremity is positioned within a receptacle in the apparatus to permit the identification band to be wrapped closely around the extremity by band fixing means of the apparatus, and fixed to itself to form a bracelet, so as to prevent its removal from the extremity without damaging the integrity of said bracelet and wherein the band has identification information applied in the form of an RFID tag or inlay, and the identification includes RFID data. 2. The apparatus of claim 1, wherein the band is provided in the form of a continuous strip feedstock, which is cut to a suitable length by a cutting means of the apparatus. 3. The apparatus of claim 1, further comprising: tensioning means to tauten the band close to the surface of the extremity, wherein the tensioning means is auto-tensioning to cater for a full range of the circumference size of the extremity; or self-adaptive length dispensing means to cater for a full range of the circumference size of the extremity. 4. An apparatus for dispensing and applying an identification bracelet that is attachable around an extremity, of an animal. including a person, the apparatus being provided with a supply of identity bands that are adapted to become a bracelet by having a first portion of the band affixed to a second portion of the band to form a continuous loop around the extremity, characterised in that the extremity is positioned within a receptacle in the apparatus to permit the identification band to be wrapped closely around the extremity by band fixing means of the apparatus, and fixed to itself to form a bracelet, so as to prevent its removal from the extremity without damaging the integrity of said bracelet, and wherein the band has identification information including at least one of printed characters, barcode, OCR, or photo images applied to its outer surface by an information application device of the apparatus. 5. An apparatus for dispensing and applying an identification bracelet that is attachable around an extremity of an animal, including a person, the apparatus being provided with a supply of identity bands that are adapted to become a bracelet by having a first portion of the band affixed to a second portion of the band to form a continuous loop around the extremity, characterised in that the extremity is positioned within a receptacle in the apparatus to permit the identification band to be wrapped closely around the extremity by band fixing means of the apparatus, and fixed to itself to form a bracelet, so as to prevent its removal from the extremity without damaging the integrity of said ... | RFID n Print Pty Ltd. | B65D083/08 B65D083/00 A61B005/117 B65C003/00 G06K019/06 G06K019/067 G07F011/00
US08909330 | B2 | 2014-12-09 | Body-worn device and associated system for alarms/alerts based on vital signs and motion | 2009-05-20 | The invention provides a body-worn monitor that measures a patient's vital signs (e.g. blood pressure, SpO2, heart rate, respiratory rate, and temperature) while simultaneously characterizing their activity state (e.g. resting, walking, convulsing, falling). The body-worn monitor processes this information to minimize corruption of the vital signs by motion-related artifacts. A software framework generates alarms/alerts based on threshold values that are either preset or determined in real time. The framework additionally includes a series of ‘heuristic’ rules that take the patient's activity state and motion into account, and process the vital signs accordingly. These rules, for example, indicate that a walking patient is likely breathing and has a regular heart rate, even if their motion-corrupted vital signs suggest otherwise. | 1. A system for processing at least one vital sign from a patient and generating an alarm, the system comprising: a first sensor comprising a first detector worn on the patient's body and configured to detect a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;a second sensor comprising a second detector worn on the patient's body and configured to detect a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;a first motion-detecting sensor positioned on a first location on the patient's body, the motion-detecting sensor configured to generate a first set of time-dependent motion waveforms indicative of motion of the location on the patient's body to which it is affixed;a processing component configured to be worn on the patient's body, the processing component comprising: a second motion-detecting sensor positioned on a second location on the patient's body, the second motion-detecting sensor configured to generate a second set of time-dependent motion waveforms indicative of motion of the location on the patient's body to which the processing component is affixed; anda microprocessor configured to receive the time-dependent physiological waveforms and at least of portion of the first and second set of time-dependent motion waveforms and continuously monitor therefrom: (i) the patient's blood pressure and heart rate calculated using the first and second time-dependent physiological waveforms; (ii) the patient's posture, activity state, and degree of motion, calculated using the first set of motion waveforms and the second set of motion waveforms and (iii) an alarm condition, determined by collectively processing the patient's blood pressure, the patient's heart rate, and the patient's posture, activity state, and degree of motion with an alarm algorithm and determining therefrom if one or more alarm criteria have been met,wherein the alarm algorithm is configured to continuously adjust the one or more alarm criteria according to the patient's posture, activity state, and degree of motion by setting an alarm criterion for the patient's heart rate which reflects the patient's activity state, and setting an alarm criterion for the patient's blood pressure which reflects the patient's posture and degree of motion. 2. The system of claim 1, wherein each motion sensor comprises an accelerometer configured to generate a time-dependent motion waveform. 3. The system of claim 2, wherein each motion sensor further comprises an analog-to-digital converter configured to digitize the time-dependent waveform. 4. The system of claim 1, wherein the first sensor comprises at least two ECG electrodes and an ECG circuit. 5. The system of claim 4, wherein the first sensor is configured to attach to a location selected from the patient's arm and chest. 6. The system of claim 4, wherein the first time-dependent physiological waveform ... | Sotera Wireless, Inc. | A61B005/04 A61B005/11 A61B005/021 G08B021/04 A61B005/0205 A61B005/08 G06F019/00 A61B005/0456 A61B005/024 A61B005/0452 A61B005/00 A61B005/01 A61B005/145 A61B005/0245 A61B005/0404
US09687681 | B2 | 2017-06-27 | Therapy system with temperature control | 2009-01-09 | A therapy system comprises a therapy module to direct a therapeutic action to a target along successive trajectories in a target zone that includes the target. A thermometry module is provided to measure temperature in a measurement field and to compute a thermal dose. The measurement field at least partially covers the target zone. A control module controls the therapy module to switch the therapeutic action to a next successive trajectory on the basis of the measured temperature or thermal dose. The successive trajectories are located inside of or outside of one another within the target zone. The therapeutic action comprises application of a focused ultrasound beam to the target. Temperature measurement is done on the basis of magnetic resonance signals. | 1. A therapy system comprising: a therapy module directing a therapeutic action to a target along a first trajectory and along successive trajectories in a target zone that includes the target, wherein the first trajectory and the successive trajectories are concentric and comprise closed geometric shapes,a thermometry module measuring temperature in a measurement field and computing a thermal dose applied by the therapy module, the measurement field at least partially covering the target zone, anda control module controlling the therapy module to direct the therapeutic action from the first trajectory to a one of the successive trajectories on the basis of at least one of the measured temperature and the computed thermal dose, whereinthe first trajectory and the successive trajectories are located entirely inside of or entirely outside of one another within the target zone. 2. A therapy system as claimed in claim 1, wherein the control module directs the therapy module to apply the therapeutic action along one of the successive trajectories on the basis of the measured temperature and the computed thermal dose exceeding respective pre-determined threshold values,the respective pre-determined threshold values being dependent on respective associated trajectories. 3. A therapy system as claimed in claim 1, wherein the thermometry module computes an applied thermal dose to the target zone from the measured temperature. 4. A therapy system as claimed in claim 1, wherein the control module controls the duration of the therapeutic action to not exceed a preset maximum duration. 5. A therapy system as claimed in claim 1, in which the therapy module is a high-intensity focused ultrasound system. 6. A therapy system as claimed in claim 1, in which the thermometry module is a magnetic resonance examination system which is arranged to derive the temperature of the target zone from magnetic resonance signals. 7. A therapy system as claimed in claim 1, in which the first trajectory or the one of the successive trajectories is broken into sub-trajectories that are transitioned on the basis of at least one of the measured temperature and the computed thermal dose to compensate for non-uniform energy deposition within the respective first trajectory or the one of the successive trajectories. 8. A method for applying a therapeutic action to a target, comprising: directing a therapeutic action to a target along a first trajectory in a target zone that includes the target,measuring temperature in a measurement field, the measurement field at least partially covering the target zone, anddirecting application of the therapeutic action to a successive trajectory in the target zone that includes the target on the basis of the measured temperature,wherein the first trajectory and the successive trajectory are concentric and comprise closed geometric shapes,wherein the first trajectory and the successive trajectory are located entirely inside of or entirely outsi... | Koninklijke Philips N.V. | A61N007/02 A61B005/01 A61B005/055 A61B017/00 A61B090/00
US10064560 | B2 | 2018-09-04 | System for cardiac pathology detection and characterization | 2010-04-06 | A system for heart performance characterization and abnormality detection includes an interface for receiving digitized electrical signals representing blood pressure waveforms over one or more heart beat cycles. The digitized electrical signals comprise, a first digital data sequence representing normal blood pressure of a patient, a second digital data sequence representing random blood pressure of a normal patient and a third digital data sequence representing a potentially abnormal blood pressure of a patient. A complexity processor calculates first, second and third complexity indices for the corresponding first, second and third digital data sequences respectively. A correlation processor uses the calculated first, second and third complexity indices to calculate one or more measures indicating deviation of the potentially abnormal blood pressure of the patient from a normal value. | 1. A system for heart performance characterization and abnormality detection included in a medical device, the system comprising: an interface configured to receive, from a blood pressure sensor coupled to the interface, digitized electrical signals representing blood pressure waveforms over one or more heartbeat cycles comprising, a first digital data sequence representing a baseline blood pressure waveform of at least one of a first patient and a second patient having common characteristics as the first patient including at least one of age, weight, gender, height, and medical conditions,a second digital data sequence representing a random blood pressure waveform that is derived by a computer anda third digital data sequence representing an ongoing blood pressure waveform of the first patient;a complexity processor configured to calculate first, second and third complexity values for the corresponding first, second and third digital data sequences respectively;a correlation processor configured to determine one or more measures based on nonlinear changes of the ongoing blood pressure waveform with respect to the baseline blood pressure waveform and based on the first, second and third complexity values;a data processor configured to determine a first anatomical location of the first patient and a type of cardiac pathology and/or associated disease based on the one or more determined measures and a second anatomical location of the first patient at which the ongoing blood pressure waveform is acquired; anda transmitter configured to transmit a message to a user indicating the determined type of cardiac pathology and/or associated disease of the first patient. 2. The system according to claim 1, wherein said interface substantially synchronizes the first, second and third digital data sequences. 3. The system according to claim 2, wherein said interface substantially synchronizes the first, second and third digital data sequences in response to a peak or minimum value in the corresponding blood pressure waveforms. 4. The system according to claim 1, wherein the data processor is further configured to monitor said first, second and third complexity values and in response to a variation in the values exceeding a predetermined threshold, generate an alert message. 5. The system according to claim 4, wherein said data processor substantially continuously performs the monitoring for at least a 24 hour period. 6. The system according to claim 4, wherein said data processor uses predetermined mapping information, associating ranges of said values with corresponding medical conditions, in comparing said values with said ranges and generates an alert message indicating a potential medical condition. 7. The system according to claim 6, wherein said predetermined mapping information associates ranges of said values with particular demographic characteristics for a population of patients and with corresponding medical conditions and said data processor... | Siemens Healthcare GmbH | A61B005/02 A61B005/021 A61B005/0215 A61B005/00
US09687194 | B2 | 2017-06-27 | Closed-loop glucose and/or insulin control system | 2009-06-17 | Disclosed are a method and/or system for determining a suggested change in a recommended therapy for a patient based, at least in part, on sensor measurements, and generating an alert to an attendant in a hospital environment upon detection of the suggested change. In another embodiment, a method and/or system is directed to automatically determining a maximum interval to alert an attendant following receipt of a measurement at an operator interface. In yet another embodiment, a method and/or system is directed to blood-glucose sensor calibration. | 1. A method comprising: determining a recommended therapy for a patient derived from signals representative of blood-glucose sensor measurements; andgenerating a signal to initiate an alarm to an attendant in response to detection of a suggested change in said recommended therapy based, at least in part, on signals representative of subsequent blood-glucose sensor measurements and an indication of a predetermined predisposition for hypoglycemia in said patient obtained independently of said subsequent blood-glucose sensor measurements. 2. The method of claim 1, wherein said recommended therapy comprises infusion of insulin in said patient at a set infusion rate. 3. The method of claim 2, and further comprising: calculating an insulin infusion rate based, at least in part, on a proportional, plus integral, plus derivative (PID) command associated with a subsequent command cycle; andestablishing a new insulin infusion rate for said subsequent command cycle as said calculated infusion rate if a difference between an insulin infusion rate in a current command cycle and said calculated infusion rate exceed a predetermined threshold. 4. The method of claim 1, and further comprising: forecasting a blood-glucose level in said patient in a subsequent command cycle; anddetermining said suggested change commencing in said subsequent command cycle based, at least in part, on said forecasted blood-glucose level. 5. The method of claim 4, and further comprising: determining a proportional, plus integral, plus derivative (PID) command associated with said subsequent command cycle;determining a rate of insulin infusion for said suggested change in said recommended therapy based, at least in part, on said ND command if said forecasted blood glucose level exceeds a predetermined threshold blood glucose level. 6. The method of claim 1, wherein said recommended therapy comprises an infusion of a bolus of glucose. 7. The method of claim 6, and further comprising determining a size of said glucose bolus based, at least in part, on a magnitude of at least one proportional, plus integral, plus derivative (PID) command associated with a command cycle. 8. The method of claim 6, and further comprising: forecasting a blood-glucose level in said patient in a subsequent command cycle; andselectively providing a command for infusion of a bolus of glucose based, at least in part, on a proportional, plus integral, plus derivative (PID) command associated with said subsequent command cycle if said forecasted blood-glucose level does not exceed a threshold blood glucose level. 9. The method of claim 1, and further comprising: determining at least one current proportional, plus integral, plus derivative (PID) command based, at least in part, on blood-glucose sensor measurements processed in a current command cycle; anddetermining at least one subsequent PID command based, at least in part, on blood-glucose sensor measurements processed in a subsequent command cycle. 10. Th... | Medtronic Minimed, Inc. | A61B005/00 A61M005/142 A61M005/172 A61B005/145 A61M005/158
US20100331633 | A1 | 2010-12-30 | SYSTEM AND METHOD FOR SHORT RANGE TELEMETRY TO MEASURE A PARAMETER OF THE MUSCULAR-SKELETAL SYSTEM | 2010-06-29 | A sensing insert device (100) is disclosed for measuring a parameter of the muscular-skeletal system. The sensing insert device (100) can be temporary or permanent. Used intra-operatively, the sensing insert device (100) comprises an insert dock 202 and a sensing module 200. The sensing module (200) is a self-contained encapsulated measurement device having at least one contacting surface that couples to the muscular-skeletal system. The sensing module (200) comprises one or more sensors (303), electronic circuitry (307), and communication circuitry (320). The electronic circuitry (307) operatively couples to the one or more sensors (303) to measure the parameter. The communication circuitry (320) couples to the electronic circuitry (307) to wirelessly transmit measurement data. The sensing module (200) wirelessly sends one or more data packets where each data packet comprises sensor data and a priority level. The data packet can further include a cyclic redundancy check. | 1. A system to measure a parameter of the muscular-skeletal system comprising: a sensing module comprising: a substrate having a contacting surface for coupling to at least one bone of the muscular-skeletal system;one or more sensors coupled to the substrate; andelectronic circuitry operatively coupled to the sensors; anda transmitter coupled to the electronic circuitry where the sensor module wirelessly sends one or more data packets and where each data packet comprises measured parameter data and a priority level. 2. The system of claim 1 where the sensing module is operatively fitted within an orthopedic prosthesis such that the contacting surface couples to the at least one bone. 3. The system of claim 2 where the sensing module is fitted within a final insert having at least one bearing surface. 4. The system of claim 3 where the sensing module is placed in a trial insert and where the trial insert is substantial equal in dimensions to the final insert. 5. The system of claim 1 where the sensing module evaluates data bounds of the measured parameter and where a priority level is assigned corresponding to the data bounds. 6. The system of claim 5 where a communication mode is based on priority level and operating mode. 7. The system of claim 6 where an operating mode includes one or more power saving modes. 8. The system of claim 6 where the priority level corresponds to when the one or more data packets are transmitted. 9. The system of claim 1 where each data packet includes a cyclic redundancy check. 10. The system of claim 1 where the one or more sensors measures a transit time, frequency, or phase. 11. A method of transmitting measurement data from an orthopedic component comprising the steps of: measuring a parameter of the muscular-skeletal system with a sensing module coupled to the muscular-skeletal system;evaluating sensor data bounds; andassigning priority level based on the data bounds and measured parameter data. 12. The method of claim 11 further including the steps of: generating packet data where the packet data comprises measured parameter data and the priority level; andsending one or more data packets to a receiver where the transmission can be immediate or delayed based on the priority level. 13. The method of claim 12 further including the steps of: assessing an operating mode where the operating mode comprises one or more power saving modes; anddetermining a communication mode based on the priority level and the operating mode. 14. The method of claim 13 further including the step of applying a cyclic redundancy check. 15. The method of claim 14 further including a step of measuring a force, pressure, or load applied by the muscular-skeletal system. 16. A prosthetic component to measure a parameter of the muscular-skeletal system comprising: an insert having at least one bearing surface that allows articulation with at least one bone of the muscular-skeletal system where the insert includes:a sensing modul... | OrthoSensor | A61B005/07
US20100292610 | A1 | 2010-11-18 | CIRCUIT BOARD FOR BODY FLUID COLLECTION | 2008-03-19 | A circuit board for body fluid collection includes a circuit board portion including an insulating layer and a conductor pattern which is supported by the insulating layer and in which an electrode, a terminal to be connected to a device for measuring the component of a body fluid, and a wire which electrically connects the electrode and the terminal are integrally provided. The circuit board for body fluid collection comprises a puncture needle for extracting the body fluid by performing puncturing, which is formed integrally with the circuit board portion and which is protruded from the circuit board portion, and a guard portion for guarding the tip of the puncture needle, which is arranged in opposition to the puncture needle in the lower flow side in the puncture direction of the puncture needle. | 1. A circuit board for body fluid collection comprising: a circuit board portion including an insulating layer, and a conductive pattern supported on the insulating layer and integrally including an electrode, a terminal for connecting to a device for measuring a component of a body fluid, and a wire electrically connecting the electrode and the terminal;a puncture needle integrally formed with the circuit board portion, and protruding from the circuit board portion in order to collect the body fluid by puncture; anda guard portion disposed downstream of the puncture needle in a puncture direction in opposing relation thereto in order to guard a tip of the puncture needle. 2. The circuit board for body fluid collection according to claim 1, further comprising: a releasing portion for releasing the guarding of the tip of the puncture needle by the guard portion, by bringing the guard portion and the puncture needle away from each other. 3. The circuit board for body fluid collection according to claim 2, wherein the guard portion is supported by the circuit board portion, andthe releasing portion is bendably provided in a portion of the guard portion located upstream of the tip of the puncture needle in the puncture direction. 4. The circuit board for body fluid collection according to claim 2, wherein the releasing portion is bendably provided in a portion of the puncture needle or the circuit board portion located upstream of the tip of the puncture needle in the puncture direction. | - | A61B005/157
US20100292596 | A1 | 2010-11-18 | System, Method and Computer-Readable Storage Medium For Heart Signal Detection | 2010-05-12 | A system, method and computer-readable storage medium are configured for the detection of electrical signals originating from a human or animal heart. In particular, for monitoring devices, it is desired to obtain electrical signals from a human or animal heart with electrical contacts at the body of an implantable medical device, hence without the need to implant electrical leads to the hearts. Hence, a method, a system and a computer-readable storage medium for detecting electrical signals originating from a human or animal heart is proposed. The method includes the steps of receiving electrical signals in at least two sensing channels, combining the electrical signals for forming a combined channel, extracting a template from the signals of the combined channel, comparing incoming electrical signals with the template, and depending from the result of the comparison, performing at least one of controlling one or more devices and signaling the result. | 1. A method for detecting electrical signals originating from a heart, comprising the following steps: receiving electrical signals in at least two sensing channels;combining the electrical signals for forming a combined channel;extracting a template from the signals of the combined channel;comparing incoming electrical signals with the template; anddepending on the result of the comparison, performing at least one of controlling one or more devices and signaling the result. 2. The method according to claim 1, where the step of combining the electrical signals comprises summing weighted values of the at least two sensing channels. 3. The method according to claim 1, where the comparing step comprises detecting patterns of signals. 4. The method according to claim 3, where comparing step comprises performing a cross-correlation between the template and the incoming electrical signals. 5. The method according to claim 3, where detecting the patterns of signals is performed with respect to a threshold. 6. The method according to claim 5, where the threshold is adapted if during a predefined period of time no pattern has been detected. 7. The method according to claim 1, where detected patterns of signals are combined with the template. 8. A system for detecting electrical signals originating from a heart, the system comprising at least one data processing device arranged in such a manner to execute a method for detecting electrical signals originating from a heart according to claim 1. 9. The system according to claim 8, where the system comprises at least three electrodes. 10. The system according to claim 9, where the at least three electrodes are connected by a switch matrix to sensing channels. 11. The system according to claim 8, where the system comprises at least one of an implantable device and an external device. 12. The system according to claim 11, where the implantable device comprises an enclosure of electrically conductive or of electrically non-conductive material. 13. The system according to claim 12, where the electrically non-conductive enclosure material covers an electrically conductive body, wherein the electrically non-conductive enclosure has one or more holes that allow the conductive body to contact surrounding tissue. 14. The system according to claim 11, where the implantable device is a pacemaker, cardioverter-defibrillator, or a monitoring device 15. A computer-readable storage medium on which program code is stored, the program code configured such that a data processing device is able to execute the method of claim 1 for detecting electrical signals originating from a heart after the data processing device has the program code loaded into a memory of the data processing device. | BIOTRONIK CRM PATENT AG | A61B005/0402
US07860607 | B2 | 2010-12-28 | Method of estimating joint moment of bipedal walking body | 2004-06-29 | The joint rotational angles of joints 9, 11 and 13 of each leg 2 of a bipedal walking body 1 are detected to grasp the positions/postures of the corresponding rigid bodies 10, 12 and 14 of each leg 2 on a leg plane passing the joints 9, 11 and 13 of each leg 2. At the same time, the acceleration of a reference point (the origin of a body coordinate system BC) of the bipedal walking body 1, the floor reaction force acting on each leg 2 and the position of an acting point thereof are grasped in terms of three-dimensional amounts. Two-dimensional amounts obtained by projecting the acceleration, the floor reaction force and the position of the acting point thereof, and the positions/postures of the corresponding rigid bodies of each 2 onto the leg plane are used to estimate the moments acting on joints of each leg on the basis of an inverse dynamic model. The stability of the estimated values of joint moments can be improved while securing the accuracy of estimating the joint moments in the bending and stretching directions of each leg, considering three-dimensional motions of the bipedal walking body. | 1. A method of estimating a joint moment of a bipedal walking body, comprising: a first step for sequentially grasping the displacement amounts of a plurality of joints, including at least an ankle joint, a hip joint and a knee joint of each leg of a bipedal walking body;a second step for sequentially grasping the positions and/or postures of corresponding rigid bodies of the bipedal walking body that are associated with rigid elements of a rigid link model using at least the rigid link model, the rigid link model being established beforehand to express the bipedal walking body in the form of a link assembly composed of a plurality of the rigid elements and a plurality of joint elements and the grasped displacement amounts of the joints;a third step for grasping the acceleration of a preset reference point of the bipedal walking body by using at least an output of an acceleration sensor attached to a predetermined region of the bipedal walking body;a fourth step for sequentially grasping a floor reaction force acting on each leg and the position of an acting point of the floor reaction force, the grasped positions and/or the postures of the corresponding rigid bodies of the bipedal walking body, the floor reaction force and the position of the acting point of the floor reaction force being changeable every moment, the acceleration of the reference point, the floor reaction force, and the position of the acting point of the floor reaction force being used to estimate a joint moment acting on at least one joint of each leg,wherein at least the displacement amounts of the hip joint, the knee joint, and the ankle joint of each leg that are grasped in the first step include the amount of rotation about an axis substantially perpendicular to a leg plane as a plane passing through these three joints;fixedly setting a posture of a leg coordinate system to the leg plane;changing the posture of the leg coordinate system with respect to a body coordinate system that is fixedly set to the bipedal walking body; andmeasuring the hip joint displacement amount about three axes including an axis perpendicular to the leg plane, wherein the displacement amount of the hip joint is a three-dimensional amount, the positions and/or postures of the corresponding rigid bodies grasped in the second step include at least the positions and/or the postures of the corresponding rigid bodies of the leg on the leg plane,the acceleration of the reference point grasped in the third step and the floor reaction force and the position of the acting point of the floor reaction force grasped in the fourth step are three-dimensional amounts, anda component of a joint moment acting on at least one joint of the leg about the axis that is substantially perpendicular to the leg plane is estimated on the basis of an inverse dynamic model representing the relationship between the motions of the corresponding rigid bodies of the leg and the translational forces and the moments acting on the c... | Honda Motor Co., Ltd. | G06F019/00 A61B005/103
US20100292602 | A1 | 2010-11-18 | SEIZURE FORECASTING, MICROSEIZURE PRECURSOR EVENTS, AND RELATED THERAPEUTIC METHODS AND DEVICES | 2008-07-11 | A multiscale recording and stimulation system for identifying and responding to epileptiform activity. The system includes the use of microelectrode arrays (e.g., 10-100 micron electrodes with 100-500 micron spacing) to monitor iEEG activity of submillimeter regions or islands in the epileptogenic zone of brain tissue (e.g., independent microdomains in the range of 100-1000 micron diameter). | 1. A therapeutic method including monitoring microseizures within neural microdomains in the brain. 2. The therapeutic method of claim 1 wherein the microdomains are sub-millimeter islands of brain. 3. The therapeutic method of claim 2 wherein the microdomains are on the order of 1000 μm or less in size. 4. The therapeutic method of claim 1 including monitoring a plurality of the microdomains. 5. The therapeutic method of claim 1 wherein monitoring microseizures includes monitoring iEEG through microelectrodes. 6. The therapeutic method of claim 5 wherein the microelectrodes are on the order of 200 μm or less in size. 7. The therapeutic method of claim 1 wherein the microelectrodes are spaced from one another on the order of 1 mm or less. 8. The therapeutic method of claim 1 and further including performing interventions as a function of the monitored microseizures. 9. The therapeutic method of claim 8 wherein the interventions include one or more of drug delivery, stimulation, and warning. 10. The therapeutic method of claim 8 and further including monitoring macrodomain iEEG and performing interventions as a function of the macrodomain iEEG and the monitored microseizures. 11. The therapeutic method of claim 10 wherein the macrodomain iEEG includes one or both of high-frequency epileptoform oscillations and DC fluctuations. 12. A multiscale electrode array for use with the therapeutic method of claim 10. 13. A passive recording and stimulation-response multiscale electrophysiology method, including monitoring microseizures in microdomain structures, for identifying epileptic brain, periods of increased seizure probability, and therapeutic intelligent control stimulation to prevent and/or abort seizures. 14. The method of claim 13 wherein the monitoring is performed through venous access. 15. The method of claim 14 wherein the simulation is performed through venous access. 16. The method of claim 13 including monitoring the electrical stimulation-induced response and epileptiform activity from a plurality of the microdomains and macrodomain structures to identify epileptogenic brain and probability of seizure occurrence. 17. The method of claim 14 and further including performing interventions including electrical stimulation as a function of the monitored microseizures, DC-shifts or HFEO, and including intelligent stimulation that is tuned to effectively cancel the local field activity over selected frequency bands or the entire spectrum (DC-1000 Hz) | MAYO FOUNDATION FOR MEDICAL EDUCATION AND RESEARCH | A61N001/36 A61B005/0476
US20100298688 | A1 | 2010-11-25 | PHOTOACOUSTIC IMAGING USING A VERSATILE ACOUSTIC LENS | 2009-10-15 | To image various soft tissues in the body using pulsed laser optical excitation delivered through a multi-mode optical fiber to create photoacoustic impulses, and then image the generated photoacoustic impulses with an acoustic detector array, a probe includes either a mirror and an acoustic lens or a special acoustic lens of variable focal length and magnification that can operate in a liquid environment that is aberration-corrected to a sufficient degree that high resolution images can be obtained with lateral as well as depth resolution. | 1. A method for imaging an object, the method comprising: (a) stimulating the object with laser light to produce ultrasound waves through the photoacoustic effect;(b) focusing the waves through an acoustic system comprising a multi-element acoustic lens; and(c) imaging the focused waves in two dimensions. 2. The method of claim 1, wherein the multi-element acoustic lens comprises a movable element or group of elements which provides the multi-element acoustic lens with variable focal length and magnification. 3. The method of claim 2, wherein the focal length and magnification are varied in order to provide depth resolution. 4. The method of claim 1, wherein the multi-element acoustic lens is configured to correct aberrations so as to provide nearly diffraction-limited acoustic imaging. 5. The method of claim 1, wherein the multi-element acoustic lens comprises an element made of a hydrogel material. 6. The method of claim 1, wherein the object is a soft tissue. 7. The method of claim 6, wherein the soft tissue is in a prostate. 8. The method of claim 1, wherein the acoustic system further comprises an acoustic mirror. 9. The method of claim 8, wherein the acoustic mirror is curved. 10. A probe for imaging an object, the probe comprising: a housing;an acoustic and optical window in the housing;optics for applying laser light to the object to produce ultrasound waves through the photoacoustic effect;an acoustic system for focusing the waves, the acoustic system comprising a multi-element acoustic lens; anda detector array, disposed so that the acoustic system focuses the waves onto the detector array, for imaging the focused waves in two dimensions. 11. The probe of claim 10, wherein the multi-element acoustic lens comprises a movable element or group of elements which provides the multi-element acoustic lens with variable focal length and magnification. 12. The probe of claim 10, wherein the multi-element acoustic lens is configured to correct aberrations so as to provide nearly diffraction-limited acoustic imaging. 13. The probe of claim 10, wherein the multi-element acoustic lens comprises an element made of a hydrogel material. 14. The probe of claim 10, wherein the acoustic system further comprises an acoustic mirror. 15. The probe of claim 14, wherein the acoustic mirror is curved. 16. A multi-element acoustic lens comprising: a plurality of acoustic lens elements, the plurality of acoustic lens elements comprising:at least one acoustic lens element having a positive power; andat least one acoustic lens element having a negative power;the plurality of lens elements being arranged to be coaxial. 17. The multi-element acoustic lens of claim 16, wherein at least one of the plurality of acoustic lens elements is configured as a movable element or group of elements which provides the multi-element acoustic lens with variable focal length and magnification. 18. The multi-element acoustic lens of claim 16, wherein the multi-element... | - | A61B005/05
US20100317983 | A1 | 2010-12-16 | METHOD FOR QUANTITATIVE ASSESSMENT OF CARDIAC ELECTRICAL EVENTS | 2009-06-12 | Systems and methods for characterizing aspects of an electrocardiogram signal are presented, wherein primary and secondary analysis schemas are utilized to determine the timing of the end of a signal wave, such as a descending Twave, with precision. In one embodiment, the primary analysis schema involves comparing voltage amplitudes within a given sampling window and the secondary analysis schema involves comparing the results of primary analysis for successive sampling windows. The system may comprise a processor or microcontroller embedded into a system such as an electrocardiogram hardware system, personal computer, electrophysiology system, or the like. | 1. A method for determining a signal wave transition point, comprising: a. sampling a first plurality of points of a signal wave in a first time window, the first plurality comprising at least a first-in-time point and a last-in-time point within the first time window;b. sampling a second plurality of points of the signal wave in a second time window different in time from the first time window, the second plurality comprising at least a first-in-time point and a last-in-time point within the second time window;c. comparing the values of the first plurality relative to each other to determine whether an intra-window patterning rule has been broken within the first window; andd. conducting a secondary analysis subsequent to determining that an intra-window patterning rule has been broken, the secondary analysis comprising comparing the values of the second plurality relative to each other to determine whether the intra-window patterning rule has been broken within the second window. 2. The method of claim 1, wherein the secondary analysis further comprises determining whether an inter-window patterning rule has been broken. 3. The method of claim 2, further comprising associating a signal wave transition point with the location on the signal wave where an inter-window patterning rule has been broken. 4. The method of claim 1, further comprising characterizing a level of noise in the signal wave. 5. The method of claim 4, where characterizing a level of noise comprises fitting a curve through datapoints comprising the signal wave. 6. The method of claim 5, wherein fitting a curve comprises fitting a polynomial equation to best fit the datapoints. 7. The method of claim 5, wherein characterizing a level of noise further comprises determining the root mean square variance of the datapoints relative to the curve. 8. The method of claim 1, further comprising selecting the intra-window patterning rule based, at least in part, upon a level of noise in the signal wave. 9. The method of claim 2, further comprising selecting the inter-window patterning rule based, at least in part, upon a level of noise in the signal wave. 10. The method of claim 8, wherein the intra-window patterning rule is selected automatically based upon computer-based analysis of the signal wave. 11. The method of claim 9, wherein the inter-window patterning rule is selected automatically based upon computer-based analysis of the signal wave 12. The method of claim 1, wherein the signal wave comprises an analog-to-digital converted electrocardiogram signal associated with one of a plurality of electrodes operatively coupled to a patient. 13. The method of claim 1, wherein the signal wave comprises a vector magnitude Twave signal representation derived from electrocardiogram voltage amplitudes associated with a plurality of electrodes operatively coupled to a patient. 14. The method of claim 1, wherein the signal wave comprises voltage amplitudes plotted versus time, and w... | NEWCARDIO, INC. | A61B005/0452 A61B005/0476 A61B005/0488
US20100280329 | A1 | 2010-11-04 | ESTIMATING A NUTRITIONAL PARAMETER FOR ASSISTING INSULIN ADMINISTRATION | 2008-07-18 | A device for estimating a nutritional parameter of a meal consumed by an individual is disclosed. The apparatus comprises processing means adapted to obtain input values of at least a physiological parameter of the user measured prior to and after intake of a meal by the user, and of at least a dose of medication administered to the user. Based on the input values the apparitus is adapted to determine from at least the obtained input values, an estimate of a nutritional parameter of the meal and to generate an output to a user indicative of the determined estimate. | 1. An apparatus (100) for estimating a nutritional parameter, comprising processing means (102, 103) and input means (101), the apparatus being adapted to: obtain input values indicative of at least a first and a second measurement of a physiological parameter of a user, and of at least a dose of medication administered to the user,determine, from at least the obtained input values, an estimate of a nutritional parameter of a meal consumed by the user between the first and the second measurement, andgenerate an output indicative of the determined estimate. 2. An apparatus as in claim 1, wherein the medication includes a medicament for affecting at least the physiological parameter of the user. 3. An apparatus as in claim 1, wherein the medication includes a medicament for the treatment of diabetes, and wherein the physiological parameter is a glucose level. 4. An apparatus as in claim 1, further comprising an output device (104) for generating a user-detectable output indicative of the determined estimate. 5. An apparatus as in claim 1, wherein the processing means is further adapted to receive an input value indicative of a preliminary estimate of the nutritional parameter or a value derived from the preliminary estimate of the nutritional parameter. 6. An apparatus as in claim 5, wherein the processing means is further adapted to generate, from the received input values an output indicative of a deviation between the determined estimate and the received preliminary estimate of the nutritional parameter. 7. An apparatus as in claim 5, wherein the processing means is further adapted to calculate deviations between respective determined estimates and respective received preliminary estimates corresponding to a plurality of meal intakes; and to determine an overall deviation from the calculated deviations. 8. An apparatus as in claim 1, wherein the processing means is adapted to determine the estimate of the nutritional parameter of the meal from the obtained input values and from a physiological model. 9. An apparatus as in claim 8, wherein the physiological model is a user-specific model. 10. An apparatus as in claim 8, wherein the physiological model is a physiological prediction model for predicting a value of the physiological parameter after the meal intake from at least a value of the physiological parameter before the meal intake, from the dose of medication administered to the user, and from the estimated nutritional parameter of the meal. 11. An apparatus as in claim 10, wherein the processing means is adapted to determine the estimate of the nutritional parameter as a value of the nutritional parameter for which a difference between the predicted value of the physiological parameter and a measured value of the physiological parameter after the meal intake is smaller than a threshold. 12. An apparatus as in claim 1, wherein the nutritional parameter is a carbohydrate content. 13. An apparatus as in claim 1, wherein the process... | Novo Nordisk A/S | A61B005/00
US20100280351 | A1 | 2010-11-04 | APPARATUS AND METHOD FOR ASSESSING ROTATIONAL BONE KINEMATICS | 2008-09-19 | Apparatus and method are disclosed for assessing rotational bone kinematics. Instantaneous center of rotation (ICR) of a bone segment is derived from a digitalized image of a bone and it is compared with the normal distribution of an ICR to determine rotational bone kinematics. The apparatus and method may be used to assess kinematic status of a normal bone and pathological bone. They also may be used to design an improved artificial disc to provide an appropriate kinesis. | 1. An apparatus for assessing a rotational bone kinematics, the apparatus comprising: an image acquiring unit which acquires digitalized images of a bone including at least two bone segments in a flexion position and an extension position;a normal distribution database which stores normal distribution data of instantaneous centers of rotation (ICRs) of the bone segments;a rotational bone kinematics assessing unit which receives digital image data of the bone from the image acquiring unit, derives the ICRs of the bone segments, compares the derived ICRs with normal distributions of the ICRs read from the normal distribution database, and outputs an assessed result; anda display unit which displays the acquired images and the assessed result. 2. The apparatus according to claim 1, wherein the bone is a cervical vertebra, a thoracic vertebra and a lumber vertebra. 3. The apparatus according to claim 1, wherein the image acquiring unit is an X-ray imaging device, a sonography device, an X-ray computed tomography (CT) device or a magnetic resonance imaging (MRI) device. 4. The apparatus according to claim 1, wherein the normal distribution database includes data on normal distribution ranges of the ICRs of the bone segments of a specific bone. 5. The apparatus according to claim 1, wherein the rotational bone kinematics assessing unit obtains the ICRs on a one-dimensional coordinate system, a two-dimensional coordinate system or a three-dimensional coordinate system. 6. The apparatus according to claim 1, wherein the rotational bone kinematics assessing unit includes: a central control unit which derives the ICRs of the bone segments of the bone to be assessed and outputs the derived result to an image processing unit;a memory which stores information on the ICRs of the bone derived by the central control unit;the image processing unit which converts the image data received from the central control unit into an image signal to be displayed on the display unit;an input unit which is connected to the central control unit and receives a command of a user; andan operating program storage unit which stores an operation program of the central control unit and a whole operating program of the apparatus. 7. The apparatus according to claim 6, wherein the central control unit derives an ICR from an intersection of perpendicular bisectors from bisectors of lines, which are connected between two corresponding points in an image of the bone in an extension position and an image of the bone in a flexion position received from the image acquiring unit, to a lower bone segment, obtains the ICRs of the bone segments by repeatedly performing the deriving of the ICR with respect to the bone segments, and calculates ratios of the coordinates of the ICRs to total lengths of the bone segments. 8. The apparatus according to claim 6, wherein the central control unit determines whether the coordinates of the ICRs of the bone segments are in the respective normal distr... | - | A61B005/05 G06K009/00
US20100292598 | A1 | 2010-11-18 | Device for Monitoring Physical Fitness | 2007-01-12 | The invention relates to a device for monitoring the physical fitness using a pulse watch part (1) for detecting heart pulse rate data and an evaluation apparatus for providing fitness information which can be displayed by means of a display device and for whose derivation a processing unit (16) is present which is designed for reading in and including personal supplementary data. The accuracy of the fitness information made available is favoured by the fact that the personal supplementary data comprises physiological additional data, for whose detection additional equipment with an additional measuring apparatus (2) is provided, which can be connected, in terms of data transfer, to the pulse watch part (1) directly or indirectly for the automatic transmission of the physiological additional data and/or heart pulse rate data. | 1. A device for monitoring physical fitness having a heart rate monitor component (1) for detecting heart rate data and an evaluation device for providing fitness information that can be displayed by a display device and is derived by a processing unit (16) for reading in and including supplemental personal data, the device comprising: the supplementary personal data including additional physiological data detected by an additional apparatus having an additional measuring device (2) and can be brought directly or indirectly into a data transfer connection with the heart rate monitor component (1) for an automatic transmission of the additional physiological data and/or the heart rate data. 2. The device as recited in claim 1, wherein the heart rate monitor component (1) and the additional measuring device (2) have matching interfaces, the processing component is situated in the heart rate monitor component (1) while the heart rate monitor component (1) and the additional measuring device (2) are embodied for the automatic transmission of the additional physiological data from the additional measuring device (2) to the heart rate monitor component (1), or the processing component is situated in the additional measuring device (2) while the heart rate monitor component (1) and the additional measuring device (2) are for an automatic transmission of heart rate data from the heart rate monitor component (1) to the additional measuring device (2). 3. The device as recited in claim 2, wherein the heart rate monitor component (1) and the additional measuring device (2) have respective interface devices via which the heart rate data can be transmitted from the heart rate monitor component (1) to the additional measuring device (2) by a memory cache or the additional physiological data can be transmitted from the additional measuring device (2) to the heart rate monitor component (1). 4. The device as recited in claim 3, wherein the additional measuring device (2) is a scale and the additional physiological data are current weight data. 5. The device as recited in claim 4, wherein the scale has an impedance measuring unit for determining a body fat percentage and information about the body fat percentage is included in the additional physiological data. 6. The device as recited in claim 5, wherein the additional physiological data also include information regarding a muscle mass percentage, a body water percentage, and/or a bone mass percentage. 7. The device as recited in claim 6, wherein the additional apparatus has a computer (3) that can be brought into a data exchange with the heart rate monitor. 8. The device as recited in claim 7, wherein the fitness information can be displayed on a display unit situated on the heart rate monitor (1), the scale (2) and/or the computer (3). 9. The device as recited in claim 8, wherein the additional measuring device (2) is a diet scale. 10. The device as recited in claim 9 wherein the supplemental personal... | BEURER GMBH & CO. KG | A61B005/04
US20100280338 | A1 | 2010-11-04 | EAR-WORN BIOFEEDBACK DEVICE | 2008-12-29 | An ear-worn biofeedback device is provided including an ear-worn structure for mounting on at least an ear of the user, at least a housing attached to the ear-worn structure, at least a sensor/electrode, for acquiring biosignals, and a circuit system accommodated in the housing. The circuit system includes an analog/digital converter for receiving and digitizing biosignals acquired by the sensor/electrode, a processor, a battery, for providing power, and a sound module, disposed around the ear, wherein during the biofeedback process, the digitized biosignals are analyzed, and when a preset condition is matched, the processor notifies the user by driving the sound module to generate audio signals for being received by the ear. | 1. An ear-worn biofeedback device, comprising: an ear-worn structure, for mounting on at least an ear of the user;at least a housing, attached to the ear-worn structure;at least a sensor/electrode, for acquiring biosignals; anda circuit system, accommodated in the housing, comprising: an analog/digital converter, for receiving and digitizing biosignals acquired by the sensor/electrode;a processor;a battery, for providing power; anda sound module, disposed around the ear,whereinduring the biofeedback process, the digitized biosignals are analyzed, and when a preset condition is matched, the processor notifies the user by driving the sound module to generate audio signals for being received by the ear. 2. The device as claimed in claim 1, wherein the audio signals are generated continuously or intermittently. 3. The device as claimed in claim 1, wherein the audio signals are used to guide the user to change from a current physiological condition to another physiological condition, to drive the user to enter a physiological condition, or to notify the user a physiological condition is achieved. 4. The device as claimed in claim 1, wherein the analysis of the biosignals is performed by the processor. 5. The device as claimed in claim 1, further comprising a wireless module for performing wirelessly transmission. 6. The device as claimed in claim 5, wherein the wireless module wirelessly receives biosignals from a biosignal acquisition device for being analyzed by the processor. 7. The device as claimed in claim 6, wherein the analysis of the received biosignals is implemented together with the biosignals acquired by the sensor/electrode. 8. The device as claimed in claim 5, wherein the biofeedback device is wirelessly communicated with an external device through the wireless module. 9. The device as claimed in claim 8, wherein the external device wirelessly receives the biosignals for further analysis, and after analysis, the analysis result is transmitted back to the biofeedback device. 10. The device as claimed in claim 8, wherein the external device wirelessly receives the biosignals for displaying and/or storage. 11. The device as claimed in claim 8, wherein the external device is a handheld device or a wrist-worn device. 12. The device as claimed in claim 8, wherein the external device is a mobile phone, a PDA or a display device. 13. The device as claimed in claim 8, wherein the external device is a second biofeedback device used by another user. 14. The device as claimed in claim 13, wherein the biofeedback device and the second biofeedback device have a real-time interaction therebetween. 15. The device as claimed in claim 14, wherein the real-time interaction is used for achieving a game. 16. The device as claimed in claim 13, wherein the biofeedback device and the second biofeedback device achieve a wireless communication therebetween by having a compatible pair of wireless modules, by commonly connecting to a computer device... | - | A61B005/00
US20100280358 | A1 | 2010-11-04 | SINGLE BREATH-HOLD SYSTEM AND METHOD FOR DETECTION AND ASSESSMENT OF MULTI-ORGAN PHYSIOLOGIC, MORPHOLOGIC AND STRUCTURAL CHANGES | 2010-04-30 | In a single breath-hold the physiologic, morphologic and structural changes in multiple organs are detected and assessed. The main steps process include: polarizing the Xe-129 gas; inhalation or introduction of a certain pre-calculated amount of the gas and start of the breath-hold; acquisition of multiple spatially oriented spectrums, localized in the same plane (2D), in multiple different planes (3D) or in 3D plus at multiple time intervals (4D); stop the breath-hold; post-processing of the acquired data; evaluation of the multiple spectrums by comparison of the values in a region of interest with those in surrounding tissues or known as normal. | 1. A method for assessing structural changes in a volume of interest, the method comprising: (a) polarizing Xe-129 gas to form hyperpolarized Xe-129 gas;(b) introducing the hyperpolarized Xe-129 gas into the volume of interest;(c) acquiring multiple spatially oriented spectra from the volume of interest, using a magnetic resonance pulse sequence capable of producing said multiple spatially oriented spectra within a predetermined time;(d) post-processing the multiple spatially oriented spectra in a computing device to obtain magnetic resonance spectra and chemical shift maps; and(e) evaluating the magnetic resonance spectra, the chemical shift maps, or the magnetic resonance spectra and the chemical shift maps to assess the structural changes in the volume of interest. 2. The method of claim 1, wherein the volume of interest is in a lung of a patient. 3. The method of claim 2, wherein the predetermined time is less than a breath hold of the patient. 4. The method of claim 1, wherein the predetermined time is less than 15 seconds. 5. The method of claim 1, wherein the magnetic resonance pulse sequence has a repetition time which is minimized in accordance with the predetermined time. 6. The method of claim 1, wherein the magnetic resonance pulse sequence has crushing gradients whose times and amplitudes are selected in accordance with the predetermined time. 7. The method of claim 1, wherein the magnetic resonance pulse sequence has an echo time of less than 2.1 milliseconds. 8. The method of claim 1, wherein the magnetic resonance pulse sequence has an RF flip angle of less than 30 degrees per voxel. 9. A system for assessing structural changes in a volume of interest, the system comprising: a source of hyperpolarized Xe-129 gas;a processing device; anda plurality of magnetic resonance coils, under control of the processing device, for acquiring multiple spatially oriented spectra from the volume of interest, using a magnetic resonance pulse sequence capable of producing said multiple spatially oriented spectra within a predetermined time;the processing device being configured for post-processing the multiple spatially oriented spectra in a computing device to obtain magnetic resonance spectra and chemical shift maps and evaluating the magnetic resonance spectra, the chemical shift maps, or the magnetic resonance spectra and the chemical shift maps to assess the structural changes in the volume of interest. 10. The system of claim 9, wherein processing device is configured for use in situations in which the volume of interest is in a lung of a patient. 11. The system of claim, wherein the processing is configured such that the predetermined time is less than a breath hold of the patient. 12. The system of claim 9, wherein the processing device is configured such that the predetermined time is less than 15 seconds. 13. The system of claim 9, wherein the processing device is configured such that the magnetic resonance pulse sequence has ... | University of Virginia Patent Foundation | A61B005/055 G06T007/00
US07860550 | B2 | 2010-12-28 | Patient positioning assembly | 2004-06-30 | A robotic patient positioning assembly for therapeutic radiation treatment includes a robotic positioning device for moving and supporting the patient during treatment, a sensor system for detecting the position of the robotic positioning device, and a controller operatively connected with the sensor system for receiving position data of the robotic positioning device and operatively connected to the robotic positioning device for controlling the motions of the robotic positioning device. The controller is adapted for controlling the motion of the robotic positioning device in response to information representative of the position of the robotic positioning device received from the sensor system, so that the treatment target within a patient loaded on the robotic positioning device is properly aligned with a radiation source of a therapeutic radiation treatment system. | 1. A patient positioning assembly for adjusting patient position during a therapeutic radiation treatment using a therapeutic radiation treatment system, wherein said patient positioning assembly comprises:a robotic positioning device to support and to move the patient loaded on said robotic positioning device, wherein said robotic positioning device is configured to move said patient in at least five independent degrees of freedom, wherein said robotic positioning device comprises: a treatment table to support said patient thereon;a robotic arm assembly coupled to the treatment table, wherein the robotic arm assembly is configured to orient the treatment table by rotating the treatment table about roll-, pitch-, and yaw-axes for three rotational degrees of freedom, and wherein the robotic arm assembly is configured to position the treatment table by rotating about two axes to translate the treatment table for two translational degrees of freedom, wherein the robotic arm assembly comprises: an arm assembly extending between a first end and a second end, wherein the second end is rotatably connected to the treatment table to allow the treatment table to rotate about three axes, wherein rotations about the three axes rotate the treatment table in the three rotational degrees of freedom; anda plate member rotatably connected to the first end of the arm assembly to allow the arm assembly to rotate about a first axis, wherein rotations about the first axis translate the treatment table in a first of the two translational degrees of freedom; anda base member rotatably connected to the plate member to allow the plate member to rotate about a second axis, wherein rotations about the second axis translate the treatment table in a second of the two translational degrees of freedom. 2. A patient positioning assembly according to claim 1, further comprising the therapeutic radiation treatment system including: a therapeutic radiation source to generate at least one therapeutic radiation beam, andan image system to generate image data representative of at least one image of a treatment target, said image data containing information regarding the near real time position of said treatment target with respect to a treatment coordinate system, wherein said robotic patient positioning assembly further comprises a sensor system to detect the position of said robotic positioning device relative to said treatment coordinate system, and a controller operatively connected: (i) with said sensor system to receive position data of said robotic positioning device, (ii) said therapeutic radiation treatment system to receive near real time image data of said treatment target and information representative of the position of said therapeutic radiation source, and (iii) said robotic positioning device to control the movement of said robotic positioning device to align said treatment target with said therapeutic radiation source. 3. A patient positioning assembly according to... | Accuray, Inc. | A61B005/05
US07833172 | B2 | 2010-11-16 | Method and device for the extraction of a body fluid | 2007-10-07 | The present invention provides a method and a device for piercing a body part and determining whether a sufficient volume of blood has been withdrawn. A lancing element is rapidly inserted into a body part in a forward phase and retracted quickly to a lesser puncturing depth. Subsequently, the lancing element is retracted slower than during the first retraction movement and the distance retracted during the second retraction is shorter than the first retraction movement. During the second retraction movement, body fluid is collected in a collection phase by a capillary structure of the lancing element. Contact between the lancing element and the body fluid is detected after the forward phase at the beginning and the end of a waiting period. | 1. A method of collecting body fluid with a lancing element having a tip and a capillary structure, comprising: inserting the tip of the lancing element into the skin of a body part to a puncture depth;retracting the tip by a first distance during a retraction phase, the first distance being measured from the puncture depth to a lesser depth than the puncture depth;retracting the tip by a second distance during a collection phase, the second distance being measured from the lesser depth to the skin surface; andthe retraction phase being faster than the collection phase and the first distance being greater than the second distance. 2. The method of claim 1, further comprising collecting body fluid during the collection phase. 3. The method of claim 1, wherein the collection phase is about 10 to 10,000 times longer than the retraction phase. 4. The method of claim 1, wherein the retracting the tip during the collection phase is performed at substantially uniform speed. 5. The method of claim 1, further comprising decelerating the lancing element during the collection phase. 6. The method of claim 1, further comprising temporarily stopping the lancing element during the collection phase. 7. The method of claim 1, wherein the second distance is about 0.1 mm to about 2 mm. 8. The method of claim 1, wherein the tip is located in the stratum corneum during at least part of the collection phase. 9. The method of claim 1, further comprising controlling movement of the lancing element as a function of time. 10. The method of claim 1, further comprising controlling movement of the lancing element as a function of distance. 11. The method of claim 1, wherein the collection phase comprises collecting body fluid through the capillary structure. 12. The method of claim 11, further comprising creating a channel in the skin extending from the puncture depth to the lesser depth and through which the body fluid moves after the tip has reached the lesser depth. 13. The method of claim 12, wherein the channel is initially evacuated after the retraction from the puncture depth to the lesser depth. 14. The method of claim 1, further comprising: mechanically controlling movement of the lancing element during the inserting the tip to a puncture depth and during the retracting the tip by a first distance; andelectrically controlling movement of the lancing element during the collection phase. 15. The method of claim 1, further comprising maintaining the tip of the lancing element at the lesser depth while collecting body fluid with the capillary structure. 16. The method of claim 1, further comprising collecting body fluid with the capillary structure while the lancing element is being retracted from the lesser depth. 17. The method of claim 16, further comprising decelerating the lancing element while the lancing element is being retracted from the lesser depth. 18. A lancing device for collecting body fluid, comprising: a housing;a lancing element recip... | Roche Diagnostics Operations, Inc. | A61B005/00 A61B017/14 A61B017/32
US07828730 | B2 | 2010-11-09 | Device for hemorrhage detection | 2007-07-02 | A device for hemorrhage detection comprising a detector for detecting the presence of blood, the detector being connected to a fixing, apparatus. The fixing apparatus is fixed to the inner wall of a hollow organ. The device further has a transmitting unit by which data transmitted by the detector are adapted to be transmitted to a receiving unit disposed outside the body. | 1. A device for hemorrhage detection, comprising a fixing means mountable inside a hollow organ, and a detecting means connected to the fixing means for detecting a presence of blood, the device further comprising: a light source configured to emit light in an ultraviolet (UV) range and light in a red range of the visible spectrum;at least one photosensitive sensor configured to measure a first light intensity depending on absorption in the UV range and a second light intensity depending on absorption in the red range of the visible spectrum, and to output a first sensor signal depending on the first light intensity and a second sensor signal depending on the second light intensity; anda calculation unit configured to calculate a ratio of the first sensor signal and the second sensor signal, wherein the ratio comprises a parameter determining whether blood is present or not present. 2. The device for hemorrhage detection according to claim 1, further comprising a transmitting unit for sending data detected by the detecting means, and a receiving unit for receiving the data sent by the transmitting unit. 3. The device for hemorrhage detection according to claim 2, wherein the transmitting unit is formed to be encapsulated. 4. The device for hemorrhage detection according to claim 2, wherein the transmitting unit transmits data to the receiving unit wirelessly. 5. The device for hemorrhage detection according to claim 1, wherein the fixing means is made of a biodegradable material. 6. The device for hemorrhage detection according to claim 1, wherein the connecting member is made of a biodegradable material. 7. The device for hemorrhage detection according to claim 1, wherein the photosensitive sensor is a photodiode or a phototransistor. 8. The device for hemorrhage detection according to claim 1, wherein the light source is an LED. 9. The device for hemorrhage detection according to claim 1, wherein the detecting means comprises a measuring gap. 10. The device for hemorrhage detection according to claim 9, wherein the detecting means further comprises a light source wherein light emitted by the light source is able to pass through the interior of the hollow organ in the area of the measuring gap. 11. The device for hemorrhage detection according to claim 1, wherein the fixing means comprises a means for stanching blood of a bleeding source. 12. The device for hemorrhage detection according to claim 11, wherein the means for stanching blood of a bleeding source is an endoscopic hemostasis clip. 13. The device for hemorrhage detection according to claim 1, wherein the light source is configured to emit light successively in the UV range and in the red range of the visible spectrum. 14. The device for hemorrhage detection according to claim 1, wherein the device comprises two light sources, a first light source emitting light exclusively in the UV range and a second light source emitting light exclusively in the red range of the visible ... | Novineon Healthcare Technology Partners, GmbH | A61B005/00
US07845225 | B2 | 2010-12-07 | Analyzing grip pressure of a bowler | 2009-01-27 | Systems and methods for measuring and analyzing the grip of a bowler. In some embodiments, a bowling grip pressure device includes a glove configured to leave the second finger, the third finger, and the thumb substantially exposed. Pressure sensors are attached to the substantially exposed fingers. At least one pressure sensor is permanently affixed to the glove. Methods are taught for analyzing a bowler's performance. In some embodiments, the bowler performs a plurality of bowling motions. Pressures exerted on the second finger, the third finger, and the thumb are measured and recorded as a function of time. The bowler's performance is evaluated based upon the recorded pressures as a function of time. Methods are also taught for using recorded grip pressure data to fit a bowler with a proper bowling ball. | 1. A bowling grip pressure device comprising: a glove including a first finger sheath positioned to accept and cover a first finger,a second finger opening positioned to accept and leave substantially exposed a second finger,a third finger opening positioned to accept and leave substantially exposeda third finger,a fourth finger sheath positioned to accept and cover a fourth finger, anda thumb opening positioned to accept and leave substantially exposed a thumb; anda plurality of pressure sensors including a first pressure sensor configured to be temporarily attached to the substantially exposed second finger,a second pressure sensor configured to be temporarily attached to the substantially exposed third finger,a third pressure sensor configured to be temporarily attached to the substantially exposed thumb, andat least one pressure sensor permanently affixed to the glove. 2. The bowling grip pressure device according to claim 1, wherein the first pressure sensor is configured to be temporarily attached to the substantially exposed second finger between a second finger tip and a second distal interphalangeal joint when worn,the second pressure sensor is configured to be temporarily attached to the substantially exposed third finger between a third finger tip and a third distal interphalangeal joint when worn, andthe third pressure sensor is configured to be temporarily attached to the substantially exposed thumb between a thumb tip and a thumb proximal interphalangeal joint when worn. 3. The bowling grip pressure device according to claim 1, wherein the at least one pressure sensor permanently affixed to the glove includes: a fourth pressure sensor affixed to the first finger sheath and positioned so as to align between a first finger tip and a first distal interphalangeal joint when worn, anda fifth pressure sensor affixed to the fourth finger sheath and positioned so as to align between a fourth finger tip and a fourth distal interphalangeal joint when worn. 4. The bowling grip pressure device according to claim 1, wherein the at least one pressure sensor permanently affixed to the glove includes: a pressure sensor positioned so as to align with the metacarpophalangeal joint of at least one of the first finger, the second finger, the third finger, and the fourth finger when worn. 5. The bowling grip pressure device according to claim 1, wherein the at least one pressure sensor permanently affixed to the glove is positioned so as to align with a bowler's palm when worn. 6. The bowling grip pressure device according to claim 1, further comprising: a processor configured to record pressure data from each of the plurality of pressure sensors as a function of time; anda communication interface connecting the plurality of pressure sensors to the processor. 7. The bowling grip pressure device according to claim 6, wherein the communication interface includes a wireless communication device. 8. The bowling grip pressure device according to claim... | United States Bowling Congress, Inc. | A61B005/22
US07850624 | B2 | 2010-12-14 | Muscle thickness sensor | 2008-03-07 | A muscle thickness sensor for measuring muscle thickness over a bone structure having a resilient frame placed over the muscle of interest and compressible material within the frame to sense muscle contractions. The compressible material is strain sensitive and generates a detectable signal when compressed due to muscle flexure. The sensor is able to differentiate between muscle contraction and muscle relaxation and is generally immune to signal noise caused by movement of bone structure and adjacent muscle groups. | 1. A muscle thickness sensor for detecting muscle thickness over a bone structure, said muscle thickness sensor comprising: a rigid frame dimensioned for contact placement on a skin surface over a muscle of interest, said muscle of interest having a unique size and shape, wherein said rigid frame comprises a rigid top member having a top surface, a bottom surface, a size and a shape configurable to coincide with said unique size and shape of the muscle of interest and a plurality of supporting members for supporting said rigid top member off of said skin surface;an elastic compressive deformation sensor disposed within said rigid frame and comprising an elastically compressible central structure comprising a top surface in contact with said rigid top member bottom surface and a bottom surface having contact with the skin surface, wherein a contraction of a muscle of interest is sensed by an elastic compressive deformation of said central structure within the rigid frame when said muscle of interest undergoes said contraction;an electrical circuit disposed on said rigid top member top surface for converting said sensed elastic compressive deformation into a detectable electrical signal having a magnitude relative to the thickness of the muscle of interest; and,a computer for receiving said detectable electrical signal and processing the detectable electrical signal into a measurement of muscle thickness. 2. The muscle thickness sensor of claim 1 wherein the muscle of interest has a first relaxed state corresponding to a first muscle thickness and a first elastic compressive deformation and a second flexed state corresponding to a second muscle thickness and a second elastic compressive deformation. 3. The muscle thickness sensor of claim 2 wherein said first elastic compressive deformation is converted by said electrical circuit into a first electrical signal and said second elastic compressive deformation is converted by the electrical circuit into a second electrical signal. 4. The muscle thickness sensor of claim 3 wherein said first electrical signal and said second electrical signal are processed by said computer into a measurement of muscle thickness in said first relaxed state and said second flexed state respectively. 5. The muscle thickness sensor of claim 4 wherein a displacement of said bone structure causes no elastic compressive deformation of the elastic compressive deformation sensor resulting in the first and second electrical signals being relatively free of ambient noise from the movement of said bone structure. 6. The muscle thickness sensor of claim 5 wherein the flexure of the muscle of interest creates an adjacent active area of displacement of the covering skin so that the muscle thickness sensor provides between 50% and 75% coverage of said active area for optimal muscle flexure detection. 7. The muscle thickness sensor of claim 6 wherein the frame is fabricated from a suitable resilient material having a first Shore ... | - | A61B005/103
US07846104 | B2 | 2010-12-07 | Monitoring physiological condition and detecting abnormalities | 2007-05-02 | A method for monitoring an individual's physiological condition and detecting abnormalities therein, comprising concurrently receiving an electrocardiograph signal and a ballistocardiograph signal. The electrocardiograph and ballistocardiograph signals are conditioned to minimize background extraneous noise after which, each signal is concurrently processed and analyzed to detect repeating cyclical patterns and further characterized to identify individual components of the repeating cycles. At least one individual component in one signal is selected as a reference marker for a selected component in the other signal. The two signals are then synchronized, outputs produced therefrom and stored in a database. | 1. A method for monitoring an individual's physiological condition comprising: acquiring electrocardiograph and ballistocardiograph signals from said individual;receiving said electrocardiograph and ballistocardiograph signals at a computer; said computer including a microprocessor;converting said electrocardiograph and ballistocardiograph signals into electrocardiograph data and ballistocardiograph data;synchronizing said electrocardiograph data with said ballistocardiograph data; andgenerating an output indicative of said individual's physiological condition. 2. A method as claimed in claim 1, wherein said electrocardiograph and ballistocardiograph signals are acquired by detecting and transmitting said electrocardiograph and ballistocardiograph signals of said individual to said computer. 3. A method as claimed in claim 2, wherein said ballistocardiograph signal is detected by an accelerometer. 4. A method as claimed in claim 3, wherein said accelerometer is for communication with a sternum of said individual. 5. A method as claimed in claim 2, wherein said electrocardiograph signal is detected by an electrocardiograph. 6. A method as claimed in claim 1, wherein said electrocardiograph and ballistocardiograph signals are converted by an analog-digital converter. 7. A method as claimed in claim 1, wherein said computer includes a database in communication with said microprocessor, said database for storing said electrocardiograph data and said ballistocardiograph data. 8. A method as claimed in claim 7, wherein said output is a comparison of said electrocardiograph data and said ballistocardiograph data of said individual with reference electrocardiograph and ballistocardiograph wave pattern data, said reference electrocardiograph and ballistocardiograph wave pattern data being stored in said database. 9. A method as claimed in claim 7, wherein said output is a comparison of at least one resting-stage electrocardiograph-ballistocardiograph signal set and at least one post-exercise electrocardiograph-ballistocardiograph signal set. 10. A method as claimed in claim 9, wherein said at least one resting-stage electrocardiograph-ballistocardiograph signal set and said at least one post-exercise electrocardiograph-ballistocardiograph signal set are stored in said database. 11. A method as claimed in claim 7, wherein said computer includes a graphical user interface in communication with said microprocessor, said graphical user interface for allowing an operator to input data into said database. 12. A method as claimed in claim 1, wherein synchronizing said electrocardiograph data with said ballistocardiograph data includes aligning an R peak of an electrocardiograph waveform corresponding to said electrocardiograph data with an H peak of a ballistocardiograph waveform corresponding to said ballistocardiograph data. 13. A method as claimed in claim 1, wherein said output is provided on a device selected from the group consisting of: monito... | Heart Force Medical Inc. | A61B005/02
US07842281 | B2 | 2010-11-30 | Magnetic particle composition for therapeutic hyperthermia | 2005-05-10 | Magnetic nanoparticle compositions are provided which provide an inherent temperature regulator for use in magnetic heating, particularly for use in magnetic hyperthermia medical treatments. The composition includes magnetic nanoparticles having a Curie temperature of between 40 and 46° C., preferably about 42° C., and may further include a polymeric material and optionally a drug or radiosensitizing agent. Methods of hyperthermia treatment of a patient in need thereof are provided which include the steps of administering to the patient a composition comprising magnetic nanoparticles having a Curie temperature of between 40 and 46° C.; and exposing the magnetic nanoparticles in the patient to an alternating magnetic field effective to generate hysteresis heat in the nanoparticles. | 1. A composition comprising magnetic nanoparticles having a Curie temperature of 47° C., wherein the composition of the nanoparticles comprises an alloy of the formula Mn0.5Zn0.5Fe2O4 and the nanoparticles have an effective mean diameter of between 10 nm and 400 nm. 2. The composition of claim 1, further comprising a polymeric material. 3. The composition of claim 2, wherein the nanoparticles are coated by or dispersed in a biocompatible polymeric material. 4. The composition of claim 1, further comprising a drug. 5. The composition of claim 4, wherein the magnetic nanoparticles and the drug are contained in a biodegradable polymeric material. 6. The composition of claim 1, wherein the nanoparticles are made by a process comprising the steps of: forming a melt which comprises at least two different metals;solidifying the melt to form a bulk solid alloy of the metals;grinding the bulk solid alloy to form particles of the alloy; andsubjecting the particles to a ball milling process effective to form magnetic nanoparticles of the alloy. 7. The composition of claim 1, wherein the nanoparticles are synthesized by chemical coprecipitation process. 8. A pharmaceutical composition comprising: magnetic nanoparticles having a Curie temperature of 47° C., wherein the composition of the nanoparticles comprises an alloy of the formula Mn0.5Zn0.5Fe2O4 and the nanoparticles have an effective mean diameter of between 10 nm and 400 nm; anda pharmaceutically acceptable carrier. | The Florida State University Research Foundation | A61B005/055 A61M037/00 B32B005/16 H01F001/00 A61K033/24
US07853458 | B2 | 2010-12-14 | System for monitoring regulation of pharmaceuticals from data structure of medical and laboratory records | 2006-10-16 | A system is provided that integrates of records of clinical laboratory services into the assessment and optimization of patient health care and, in particular, regulation of the use of pharmaceuticals. Laboratory test result records are used in conjunction with other health care benefits records to monitor regulation of use of pharmaceuticals by patients. The incorporation of laboratory tests and results into such a utilization system allows improvement in the management of a patient's therapy based on a more precise picture of the patient's level of illness as revealed by the laboratory test results. The system of the present invention also allows optimization of the selection of laboratory tests to be performed, and also provides an outcome assessment of the risk of hospitalization due to pharmaceutical treatments resulting in physician intervention, leading to a change in physician prescribing behavior and, accordingly, a decrease in drug induced hospitalizations and improved quality of patient care and savings of health care costs. | 1. A computer system for reducing the risk of a drug therapy related illness in a patient, the system comprising: at least one medical records database comprising: i) drug therapy records, said drug therapy records including an indication of patient identification and an indication of prescribed drug therapy associated with said patient identification, and ii) laboratory test result records, said laboratory test result records including an indication of patient identification and an indication of laboratory test results associated with said patient identification;a medical information database, said medical information database including criteria pertaining to: i) interactions between a first drug and a second drug; ii) a relationship between the first drug and a first disease; and iii) administration of a laboratory test in connection with the administration of the first drug; anda search and determination engine, implemented by a computer, for reviewing said laboratory test result records, said drug therapy records and said medical information database criteria, and for determining an output for reducing the risk of the drug therapy related illness in the patient using said laboratory test result records, said drug therapy records and the medical information database criteria by determining a drug interaction indicating initiation of the first drug in a patient with current usage of the second drug, including the computer: (i) detecting the usage of the second drug concurrently with the first drug, (ii) monitoring for a laboratory report, (iii) generating an alert to the user of the system if such a report is not found, (iv) generating an alert to warn the user of the system about possible interaction effects between the first drug and the second drug; or (v) generating an alert to warn the user of the system of the need for particular tests from the laboratory to monitor the levels of at least one of the first and second drugs, and said search and determination engine determining the output for reducing the risk of the drug therapy related illness in the patient (PATIENT1) using said laboratory test result records, said drug therapy records and said medical information database criteria by performing one or more of the following functions:a monitor drug initiation function determining that PATIENT1 has been prescribed a first drug (DRUG A), detecting of initiation of DRUG A, monitoring for a first laboratory report (LABI) within a first predetermined number of days (X days) after initiation of DRUG A for duration period of a second predetermined number of days (Z days), and generating an alert is provided to the user if LABI indicates that no test has been run during the current period after initiation, and wherein the monitoring of lab testing after initiation of treatment with DRUG A provides information regarding level of DRUG A in patient and effectiveness of DRUG A as treatment;a monitor drug frequency function determining that PATIENT1 h... | Medco Health Solutions, Inc. | A61B005/00 G06Q010/00
US07848788 | B2 | 2010-12-07 | Magnetic resonance imaging probe | 2002-04-11 | Herein is disclosed a magnetic resonance imaging probe, having a probe shaft including a magnetic resonance antenna, and a spring tip attached to a distal end of the antenna. | 1. A magnetic resonance imaging probe, comprising: a probe shaft having a magnetic resonance antenna, the probe shaft comprising a core of non-magnetic material;a first insulator/dielectric covering the core;a shield layer covering the first insulator/dielectric;a proximal spring assembly coupled to a distal end of the shield layer and configured and arranged to act with the shield layer as a coaxial shield; anda magnetic resonance imaging coil conductively attached to the core and forming a part of the magnetic resonance antenna;a spring tip having increased flexibility relative to the probe shaft proximate the antenna; andan electrically insulating material affixing the spring tip to a distal end of the probe shaft and preventing electrical conduction between the spring tip and the imaging coil. 2. The probe of claim 1, wherein the core comprises a Nitinol wire and alternating layers of gold, silver, and gold, surrounding at least a portion of the Nitinol wire. 3. The probe of claim 1, wherein the probe shaft further comprises a second insulator/dielectric disposed over the shield layer. 4. The probe of claim 1, wherein the core of non-magnetic material is plated with alternating layers of gold and silver. 5. The probe of claim 1, wherein the core comprises a central core having at least one of carbon, glass fiber, and a polymer, and the core further comprises a radio frequency conductive covering surrounding the central core. 6. The probe of claim 5, wherein the covering comprises alternating layers of gold, silver, and gold, surrounding the central core. 7. The probe of claim 1, wherein the spring tip is attached to the imaging coil by an adhesive joint. 8. The probe of claim 1, wherein the imaging coil comprises a helical whip with a proximate end and a distal end, the helical whip having coils with a diameter and a spacing, and wherein the distal end of the helical whip is connected to the spring tip. 9. The probe of claim 8, wherein the helical whip is covered by a biocompatible material or covering. 10. The probe of claim 9, wherein an electrical length of the imaging coil is chosen so as to compensate for the biocompatible material or covering. 11. The probe of claim 1, wherein the probe is sized and shaped to be a guidewire. 12. The probe of claim 1, further comprising a connector coupled to a proximal end of the probe shaft. 13. The probe of claim 12, wherein the connector is removably coupled to the proximal end of the probe shaft. 14. The probe of claim 12, wherein the connector couples to an interface. 15. The probe of claim 14, wherein the interface comprises at least one of a tuning-matching circuit, a balun circuit, a decoupling circuit, or a variable capacitor. 16. The probe of claim 1, wherein the antenna comprises a loopless antenna. 17. The probe of claim 1, wherein the spring tip further comprises a ribbon extending from the spring tip to the core. 18. The probe of claim 1, wherein an outer surface of the pr... | The Johns Hopkins University SurgiVision, Inc. | A61B005/055
US07846157 | B2 | 2010-12-07 | Method and apparatus for control of ablation energy and electrogram acquisition through multiple common electrodes in an electrophysiology catheter | 2003-03-17 | Method and apparatus for control of ablation energy and electrogram acquisition through multiple common electrodes in an electrophysiology catheter. A device that routes ablation energy to and that routes mapping signals received from an electrophysiology catheter having a plurality of conductive filaments including circuitry that provides, for each conductive filament when ablation energy is being delivered, an electrical signal path that has a low impedance for ablation energy and a high impedance for mapping signals. The device also includes circuitry that provides, for each conductive filament, when mapping signals are being received, an electrical signal path that has a high impedance for ablation energy and low impedance for mapping signals. At least one switch is provided that selectively groups electrodes into sectors for delivery of ablation energy. | 1. A device for routing ablation energy to and for routing mapping signals received from an electrophysiology catheter having a plurality of conductive filaments comprising first and second filaments, the device comprising: circuitry that provides an electrical signal path between the first and second filaments that has a low impedance when the first and second filaments are used to transmit ablation energy and a higher impedance when the first and second filaments are used to transmit mapping signals. 2. The device of claim 1 further comprising switching circuitry for selectively coupling filaments in the electrophysiology catheter to a source of ablation energy. 3. The device of claim 2, wherein the switching circuitry selectively couples sectors of filaments to a source of ablation energy. 4. The device of claim 3, wherein there are four sectors. 5. The device of claim 2, wherein the switching circuitry selectively couples approximately a first half of the filaments in the electrophysiology catheter to a source of ablation energy. 6. The device of claim 5, wherein the switching circuitry selectively couples approximately a second half of the filaments in the electrophysiology catheter to a source of ablation energy. 7. The device of claim 6, wherein the first half of the filaments is interleaved with the second half of the filaments. 8. The device of claim 1, wherein the circuitry comprises a capacitor in the electrical signal path of each conductive filament. 9. The device of claim 1, further comprising a resistor respectively coupled to each conductive filament. 10. A catheter system comprising an electrophysiology catheter, an ablation energy generator, a recording device, and a device according to claim 1 that couples the electrophysiology catheter to the ablation energy generator and the recording device. | C.R. Bard, Inc. | A61B018/14 A61B005/04
US07828333 | B1 | 2010-11-09 | Label sheet with wristband | 2008-07-18 | A printable form comprising a printable face ply and a wristband releasably bonded to a surface of the printable face ply. In at least one embodiment the wristband may be detached from the surface of the printable face ply and secured to around a body part for use in identification. In at least one embodiment the printable face ply comprises one or more labels that may be detached therefrom and applied to the wristband. | 1. A printable form comprising: a substrate material, said substrate material comprising a leading edge, a trailing edge, and first and second side edges, said substrate material comprising a face ply and a liner ply, said face ply comprising a face ply surface and a second surface opposite said face ply surface, said liner ply removably adhered to said second surface of said face ply;a wristband, said wristband comprising a leading margin, a trailing margin, first and second side margins, a stub portion, a removeable portion, and a transverse line of weakness between said stub portion and said removeable portion, a top side, and an underside, said underside comprising an underside surface bounded by said leading margin, said trailing margin, and said first and second side margins; andat least one adhesive stripe adhered to said underside surface, at least one of said at least one adhesive stripes releasably bonded to said face ply surface such that said wristband is removably attached to said face ply surface,wherein at least one of said at least one adhesive stripes is adhered to said underside surface on each side of said transverse line of weakness and said removeable portion is separable from said stub portion at said transverse line of weakness, wherein said stub portion is permanently adhered to said face ply surface. | Chicago Tag & Label, Inc. | B42D015/00 A44C005/00 A61B005/103 A61B005/117 G09F003/10
US08192372 | B2 | 2012-06-05 | Test media cassette for bodily fluid testing device | 2010-07-21 | A bodily fluid sampling device includes a piercing device and a sensor enclosed in a housing. A cassette, which contains test media, is positioned proximal to the sensor so that the sensor is able to analyze a bodily fluid sample collected on the test media. The cassette includes a supply portion from which unused test media is supplied and a storage portion in which contaminated test media is stored after exposure to the bodily fluid. The cassette is adapted to collect a series of bodily fluid samples without requiring disposal of the test media. | 1. A cassette, comprising: a test tape having a plurality of test sections to test a plurality of body fluid samples;a cover protecting the test tape prior to use, wherein the cover is a layer over the test tape; andan exposure device configured to permit access to a test section of the test tape covered by the cover. 2. The cassette of claim 1, wherein the exposure device includes a cover reel configured to wind the cover therearound to peel the cover from the test tape. 3. The cassette of claim 2, further comprising: a supply reel comprising the test tape prior to use; anda take-up reel comprising the test tape after use. 4. The cassette of claim 1, wherein: the cover includes a plurality of blister packs covering the test sections; andthe exposure device includes a cutting device configured to pierce the blister packs to expose the test sections. 5. The cassette of claim 1, wherein: the cover includes a seal membrane; andthe exposure device includes a piercing device configured to pierce the seal membrane. 6. The cassette of claim 3, further comprising: an outer cover including an unexposed media supply portion, wherein the supply reel is located inside the unexposed media supply portion,an exposed media storage portion, wherein the take-up reel is located inside the exposed media storage portion, anda testing portion located between the unexposed media supply portion and the exposed media storage portion. 7. The cassette of claim 3, further comprising a peel tab configured to peel the cover from the test tape. 8. The cassette of claim 7, further comprising guides configured to guide the test tape between the supply reel and the take-up reel. 9. The cassette of claim 3, wherein: the supply reel has a supply gear;the cover reel has a cover gear; andthe supply gear and the cover gear are intermeshed to synchronize rotation of the supply reel and the cover reel. 10. The cassette of claim 1, further comprising: a supply reel comprising the test tape prior to use; anda take-up reel comprising the test tape after use. 11. The cassette of claim 10, further comprising: an outer cover including an unexposed media supply portion, wherein the supply reel is located inside the unexposed media supply portion,an exposed media storage portion, wherein the take-up reel is located inside the exposed media storage portion, anda testing portion located between the unexposed media supply portion and the exposed media storage portion. 12. The cassette of claim 4, wherein the test tape has a piercing device hole between adjacent test sections. 13. The cassette of claim 4, wherein the blister pack is a plastic film. 14. The cassette of claim 4, wherein the test sections include test media that is semitransparent or transparent to allow optical analysis of the body fluid samples. 15. The cassette of claim 5, wherein each of the test sections has a vacuum passageway and a test pad positioned between the seal membrane and the vacuum passageway. 16. A cas... | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00
US08257277 | B2 | 2012-09-04 | Test media cassette for bodily fluid testing device | 2010-08-02 | A bodily fluid sampling device includes a piercing device and a sensor enclosed in a housing. A cassette, which contains test media, is positioned proximal to the sensor so that the sensor is able to analyze a bodily fluid sample collected on the test media. The cassette includes a supply portion from which unused test media is supplied and a storage portion in which contaminated test media is stored after exposure to the bodily fluid. The cassette is adapted to collect a series of bodily fluid samples without requiring disposal of the test media. | 1. A cassette, comprising: a test tape having a plurality of test sections to test a plurality of body fluid samples;a cover tape configured to protect protecting the test tape prior to use; anda cover reel configured to peel the cover tape from the test tape. 2. The cassette of claim 1, further comprising: a supply reel comprising the test tape prior to use; anda take-up reel comprising the test tape after use. 3. The cassette of claim 2, further comprising: the supply reel having a supply gear; andthe cover reel having a cover gear intermeshed with the supply gear to synchronize rotation of the cover reel with the supply reel. 4. The cassette of claim 3, further comprising: a peel tab disposed along the cover tape to facilitate peeling of the cover tape from the test tape; andguides disposed between the supply reel and the take-up reel for guiding the test media tape between the supply reel and the take-up reel. 5. A device for analyzing a body fluid, comprising: a test media cassette includinga test media tape having chemistry adapted to analyze the body fluid,a cover tape configured to form an airtight seal over the test media tape to preserve the chemistry of the test media tape,a supply reel comprising an uncontaminated section of the test media tape,a take-up reel comprising a contaminated section of the test media tape, anda cover reel configured to peel the cover tape from the test media tape for exposing the test media tape to the body fluid. 6. The device of claim 5, further comprising: the supply reel having a supply gear; andthe cover reel having a cover gear intermeshed with the supply gear to synchronize rotation of the cover reel with the supply reel. 7. The device of claim 6, further comprising: a peel tab disposed along the cover tape to facilitate peeling of the cover tape from the test media tape; andguides disposed between the supply reel and the take-up reel for guiding the test media tape between the supply reel and the take-up reel. 8. The device of claim 7, further comprising: a sensor positioned between the supply reel and the take-up reel to sense at least one property of the body fluid on the test media tape. 9. The device of claim 5, further comprising: a peel tab disposed along the cover tape to facilitate peeling of the cover tape from the test media tape; andguides disposed between the supply reel and the take-up reel for guiding the test media tape between the supply reel and the take-up reel. 10. A method, comprising: providing a test media cassette that includes a test media tape having test chemistry, a cover tape forming an airtight seal over the test media tape to preserve the test chemistry of the test media tape, a supply reel around which an uncontaminated section of the test media tape is wound, a take-up reel around which a contaminated section of the test media tape is wound, and a cover reel around which the cover tape is wound;indexing the test media tape with the take-up reel;exposing the te... | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00
US08226680 | B2 | 2012-07-24 | System for embolizing a target site in a body by application of an external energy | 2010-08-31 | A vaso-occlusive device is introduced in a cavity of a patient's vasculature, such as an aneurysm, after which an external energy source, such as a MRI machine is activated to heat the vaso-occlusive device to assist in forming a thrombus or embolism within the treatment site, to release and/or activate a diagnostic or therapeutic agent carried by the vaso-occlusive device, and/or to fuse together portions of the vaso-occlusive device to help stabilize the device in a three-dimensional shape. | 1. A system for embolizing a target site in a vasculature of a body, comprising: a vaso-occlusive device detachably coupled to a distal end of a delivery wire insertable through a delivery catheter for deploying the vaso-occlusive device at a target site in a vasculature of a body, wherein the vaso-occlusive device is coated with a therapeutic bioactive agent, and additionally having a polymeric material coating substantially covering the bioactive agent coating; andan external energy emitting device configured to deliver energy from a location outside the body to thereby heat the vaso-occlusive device after the vaso-occlusive device has been deployed at the target site in the body, wherein, when heated by energy emitted from the external energy emitting device, the polymeric material coating at least partially melts or softens so that the bioactive agent is released or activated at the target site. 2. The system of claim 1, the target site comprising one of an aneurysm, a blood vessel lumen and a fistula. 3. The system of claim 1, the external energy emitting device comprising a magnetic resonance device. 4. The system of claim 3, wherein the vaso-occlusive device comprises a ferrous material in sufficient concentration to cause heating of the device in response to energy delivered by the magnetic resonance device. 5. The system of claim 1, wherein the external energy emitting device comprises an ultrasound device configured to be acoustically coupled to an exterior surface of the body. 6. The system of claim 1, wherein the external energy emitting device comprises a radio frequency energy generator. 7. A system for embolizing a target site in a body, comprising: a vaso-occlusive device detachably coupled to a distal end of a delivery wire insertable through a delivery catheter for deploying the vaso-occlusive device at a target site in a vasculature of a body, the vaso-occlusive device comprising a highly resistive element that is heated by application of a magnetic field, the vaso-occlusive device further comprising a polymeric material exterior coating and a therapeutic bioactive agent underlying the polymeric material coating; anda magnetic resonance imaging (“MRI”) device, wherein activating the MRI device to apply a variable magnetic field to the body after the vaso-occlusive device has been deployed at the target site heats the highly resistive element in the vaso-occlusive device, thereby at least partially melting or softening the polymeric material exterior coating on the vaso-occlusive device to release or activate the underlying therapeutic bioactive agent. 8. The system of claim 7, wherein the vaso-occlusive device may be sufficiently heated by application of magnetic field energy to cause coagulation of blood at the target site. 9. The system of claim 7, wherein the vaso-occlusive device is configured to assume a three-dimensional shape upon being deployed at the target site, and wherein the vaso-occlusive device may be suf... | Stryker Corporation Stryker NV Operations Limited | A61M029/00
US08906307 | B2 | 2014-12-09 | Apparatus for providing power management in data communication systems | 2010-08-18 | A blood glucose meter having a compact housing, a display unit disposed on the housing, the display unit including a display light source to illuminate the display unit, an input unit disposed on the housing, the input unit configured to provide input functions for the blood glucose meter, and a power source provided within the housing for providing power to the blood glucose meter, where the housing includes a port integrated on said housing configured to receive a blood glucose test strip, and corresponding methods of measuring blood glucose meter is provided. | 1. An analyte meter, comprising: a housing;a strip port disposed in the housing and sized to receive an analyte test strip;a processing unit for measuring an analyte level detected from the analyte test strip;an interface unit that operatively couples the analyte meter to an external computer to transmit data related to the analyte level detected from the analyte test strip, wherein the interface unit operatively couples the analyte meter to an external computer via a wireless connection;a liquid crystal display (LCD) screen;a non-transitory memory comprising stored instructions that cause the processing unit tooperate the analyte meter in a clock mode that displays current time and/or date information on the LCD screen, a test mode for performance of analyte measurements, a configuration mode that enables user selection of configuration preferences, and a measurement log review mode that enables reviewing of stored analyte measurements, andconfigure the analyte meter to automatically exit the clock mode and enter the test mode upon receiving the analyte test strip in the strip port. 2. The analyte meter of claim 1, wherein the analyte meter further comprises a mode switch to toggle between different modes of operation. 3. The analyte meter of claim 1, wherein the analyte meter exits all modes of operation after exceeding a predetermined period of non-activity. 4. The analyte meter of claim 1, wherein the analyte meter further comprises a speaker unit to output audio signals. 5. The analyte meter of claim 1, wherein the LCD displays a memory mode symbol to indicate that data has been stored in a memory unit of the analyte meter. 6. The analyte meter of claim 1, further comprising: a light source coupled to the housing to illuminate the strip port; anda power source to provide power to the light source. 7. The analyte meter of claim 6, wherein the light source is one or more LEDs. 8. The analyte meter of claim 6, wherein the analyte meter pulses the light source at a predetermined frequency and predetermined duty cycle, wherein the predetermined frequency is greater than 100 Hz, and wherein the predetermined duty cycle is approximately 10%. 9. The analyte meter of claim 6, wherein the light source illuminates the LCD. 10. The analyte meter of claim 6, wherein the light source provides backlighting to the LCD. 11. The analyte meter of claim 6, wherein the light source is coupled to the housing on the same end as the strip port. 12. The analyte meter of claim 1, wherein the analyte meter is a glucose meter. 13. The analyte meter of claim 1, measurement log review mode allows reviewing of glucose measurements stored in the meter. 14. The analyte meter of claim 1, wherein when the meter is in the measurement log review mode, the meter displays a glucose summary of stored results, average glucose level information, a number of measurements incorporated into the displayed average level, or a visual indication to the user that the displayed ... | Abbott Diabetes Care Inc. | G01N033/00 C12Q001/54 A61B005/145
US08491499 | B2 | 2013-07-23 | Biological fluid sampling and pretreating system and a method thereof | 2008-10-21 | Provided is a system for sampling and pretreating biological fluid. It comprises: a piercing unit having at a lower portion a capillary tip which is to be inserted into skin to a predetermined depth to take biological fluid therethrough; a dropper, connected to an upper portion of the piercing unit, having an injection tube at an upper portion thereof, the injection tube communicating with the capillary tip; and a reagent container, designed to accommodate the piercing unit therein in an airtight manner so as to seal an outer circumference of the piercing unit, functioning to contain a reagent for treating the biological fluid of the capillary tip of the piercing unit. The system allows even a novice to sample and pretreat biological fluid with high accuracy without the use of expensive precision devices. The system employs fewer expendable supplies, thus providing higher convenience for the user. | 1. A system for sampling and pretreating biological fluid, comprising: a piercing unit having at a lower portion a capillary tip which is to be inserted into skin to a predetermined depth to take biological fluid therethrough;a dropper, connected to an upper portion of the piercing unit, having an injection tube at an upper portion thereof, the injection tube communicating with the capillary tip; anda reagent container, designed to accommodate the piercing unit therein in an airtight manner so as to seal an outer circumference of the piercing unit, for containing a reagent for treating the biological fluid of the capillary tip of the piercing unit;wherein the dropper is provided with a cylindrical member the outer circumference of which is adapted to be in airtight contact with an inner circumferential wall of the reagent container, and the reagent container is provided at an upper portion thereof with a sealing plate for closing the reagent container, which is designed to be readily perforated by the capillary tip of the piercing unit and to function to fill a gap between the dropper and the reagent container. 2. The system according to claim 1, wherein the capillary tip is provided at an upper portion thereof with at least one stopper for preventing the capillary tip from being inserted into skin in excess of a predetermined depth. 3. The system according to claim 2, wherein the capillary tip is provided at the upper portion thereof with two or more circular stoppers. 4. The system according to claim 3, wherein the stopper is provided at an upper portion thereof with a ring for connecting the stopper to the dropper. 5. The system according to claim 1, wherein the injection tube is provided at a lower portion thereof with a circular plate in such a way that the circular plate is in airtight contact with an upper portion of the reagent container. 6. The system according to claim 1, wherein the dropper is provided at an upper portion thereof with a lid to which an injection tube is fixed. 7. The system according to claim 6, wherein the lid is provided at an inside thereof with a groove which is flexibly relaxed when the injection tube is fixed thereto and which closes the injection tube, and the injection tube has an outer diameter which becomes upwardly narrow. 8. A method for sampling and pretreating biological fluid using the device of claim 1, comprising: inserting the capillary tip into the skin to a predetermined depth to introduce a predetermined quantity of biological fluid into the capillary tip;fixedly inserting the piercing unit and a part of the dropper with the biological fluid impounded in the capillary tip into the reagent container containing a pretreatment reagent, to react the biological fluid with the reagent;applying the pretreated biological fluid to a diagnostic kit with both the piercing unit and the dropper fixed to the reagent container; andproviding a dropper with a lid to which the injection tube is inserted and i... | Boditechmed Inc | A61B005/00
US08501093 | B2 | 2013-08-06 | Portable handheld medical diagnostic devices with color-changing indicatior | 2009-06-11 | A portable handheld medical diagnostic device includes a housing forming a protective enclosure. A main circuit board is located in the protective enclosure. The main circuit board includes a controller facilitating a physiologic measurement. A display device is connected to the main circuit board that displays information related to the physiologic measurement. An electronic skin is on the housing. The electronic skin comprises a liquid crystal material and is configured to display a color. | 1. A portable handheld medical diagnostic device for a user, comprising: a housing forming a protective enclosure;a main circuit board located in the protective enclosure, the main circuit board including a controller which is configured to facilitate a physiological measurement;a single-piece frame located in the protective enclosure, the frame having a strip port formed therein;a display device mounted on the frame and connected to the main circuit board, wherein the display device is flat and rigid and displays information related to the physiological measurement; andan electronic skin on the housing separate from the display device, wherein the electronic skin is a color display and comprises a liquid crystal material, and is configured to display a color, wherein the controller is operatively connected to both the electronic skin and the display device, wherein the controller is configured to control the color of the electronic skin and configured to control the display of the information on the display device related to the physiological measurement, and wherein the controller is configured to permit selection by the user of the color of the electronic skin from a plurality of colors. 2. The portable handheld medical diagnostic device of claim 1, wherein the controller is configured to change the color of the electronic skin to a second, different color in response to the physiological measurement. 3. The portable handheld medical diagnostic device of claim 1, wherein the controller is configured to change the color of the electronic skin a second, different color based on a measurement value determined by the physiological measurement. 4. The portable handheld medical diagnostic device of claim 3, wherein the measurement value is a blood glucose concentration value. 5. The portable handheld medical diagnostic device of claim 4, wherein the controller is configured to change the electronic skin to the second, different color when the blood glucose concentration value is within a predetermined range of blood glucose concentration values. 6. The portable handheld medical diagnostic device of claim 1 further comprising a strip port assembly configured to receive a test strip for testing a blood sample. 7. The portable handheld medical diagnostic device of claim 6, wherein the strip port assembly comprises a strip port accessible from outside the protective enclosure for inserting a test strip, at least part of the strip port being formed of a material selected for distribution of light to illuminate the at least part of the strip port. 8. The portable handheld medical diagnostic device of claim 7, further comprising a light source used to illuminate the at least part of the strip port, and wherein the controller is configured to control the light source and is configured to permit selection by the user of a color illuminated by the light source. 9. The portable handheld medical diagnostic device of claim 7 further comprising: a light s... | Roche Diagnostics Operations, Inc. | G01N015/06
US09211074 | B2 | 2015-12-15 | System, method, apparatus, device and computer program product for automatically detecting positioning effect | 2009-11-17 | An automated system, method, apparatus, device and/or computer program product for detecting positioning effect is set forth, the apparatus according to an exemplary embodiment may include an output operable to couple to one or more stimulating electrodes to stimulate one or more peripheral nerves of the patient, an input operable to couple to one or more recording electrodes to record resultant electrical waveforms generated by a nervous system of a patient in response to the stimulating module, and one or more processors operable to identify the positioning effect based on the resultant electrical waveforms. | 1. An automated apparatus for detecting positioning effect, the apparatus comprising: an output operable to couple to at least one stimulating electrode to stimulate one or more peripheral nerves of a patient;an input operable to couple to at least one recording electrode to record resultant electrical waveforms generated by a nervous system in response to the one or more stimulating electrodes;at least one processor, coupled to said output and said input, operable to: detect a change in the resultant electrical waveforms; andanalyze the detected change in the resultant electrical waveforms to automatically identify the positioning effect, the positioning effect comprising an immediate or potential injury to the one or more peripheral nerves due to stress associated with a position in which the patient has been placed for medical care or treatment; anda display interface operable to forward data to a display, the data comprising information for displaying an anatomical diagram that depicts a location of the detected change. 2. The apparatus of claim 1, wherein the change in the resultant electrical waveforms comprises a change in a morphology of the resultant electrical waveforms. 3. The apparatus of claim 1, further comprising: an alert unit, coupled to said at least one processor, operable to alert a user to the positioning effect. 4. The apparatus of claim 1, further comprising: an interface, coupled to said at least one processor, operable to couple a table to said at least one processor. 5. The apparatus in claim 4, wherein said interface is operable to transfer information between the table and said at least one processor to reduce the positioning effect. 6. The apparatus of claim 4, wherein the apparatus further comprises the table. 7. The apparatus of claim 6, wherein the table comprises a surface on which a patient may be placed, selected from the group consisting of: a bed, a chair, an operating room table, a pre-op table, and a post-op table. 8. The apparatus in claim 1, wherein said apparatus further comprises at least one of: means for obtaining information from an anesthesia machine to determine when changes in the resultant electrical waveforms are due to anesthesia;an anesthesia machine;a blood pressure machine; ormeans for obtaining information from a blood pressure machine to determine when changes in the resultant electrical waveforms are due to blood pressure. 9. The apparatus of claim 1, wherein the resultant electrical waveforms are somatosensory evoked potential waveforms. 10. The apparatus of claim 1, wherein the at least one processor is operable to generate a recommendation for action to ameliorate the identified positioning effect. 11. The apparatus of claim 10, wherein the at least one processor is operable to generate the recommendation for action to ameliorate the identified positioning effect based on position information of a table. 12. The apparatus in claim 11, further comprising: an interface, coupled... | SAFEOP SURGICAL, INC. | A61B005/04 A61B005/05 A61G013/10 A61B005/00 A61B005/11 A61G013/00
US09848760 | B2 | 2017-12-26 | Devices for continual monitoring and introduction of gastrointestinal microbes | 2009-06-29 | Systems and methods described herein include those for the continual modification of intestinal microbes. Described herein are systems including sampling devices, analysis devices, computational devices and user interface devices as well as methods for the use of such devices in combination. | 1. A system for continual modification of intestinal microbes comprising: at least one ingestible sampling device including one or more electronic detectors, the at least one ingestible sampling device including one or more physical sample collectors and an internal storage region;at least one analysis device including a cell culture of the at least one sample;at least one computational device operably connected to the at least one analysis device, and including in memory a standard profile for gastrointestinal microbes;at least one user interface device operably connected to the at least one computational device, the user interface device including a monitor; andat least one material dispensing device connected to the at least one computational device. 2. The system of claim 1, wherein the at least one sampling device comprises: one or more chemical sensors. 3. The system of claim 1, wherein the at least one sampling device comprises: one or more containers. 4. The system of claim 1, wherein the at least one sampling device comprises: one or more absorbent pads. 5. The system of claim 1, wherein the at least one sampling device comprises: at least one additional ingestible sampling device. 6. The system of claim 1, wherein the at least one analysis device comprises: at least one nucleic acid analysis device. 7. The system of claim 1, wherein the at least one analysis device comprises: at least one protein analysis device. 8. The system of claim 1, wherein the at least one analysis device comprises: at least one chemical analysis device. 9. The system of claim 1, wherein the at least one analysis device comprises: at least one electronic detectors. 10. The system of claim 1, wherein the at least one computational device and the at least one user interface device are an integrated unit. 11. The system of claim 1, comprising: at least one additional sampling device. 12. The system of claim 1, comprising: at least one dietary sampling device. 13. The system of claim 12, wherein the at least one dietary sampling device is operably connected to the at least one computational device. 14. A system for monitoring and introducing gastrointestinal microbes into a mammalian gastrointestinal tract over time, comprising: at least one gastrointestinal microbe analysis device;at least one computational device connected to the at least one gastrointestinal microbe analysis device, the at least one computational device including logic with protocols to integrate information relating to an individual with information from the at least one gastrointestinal microbe analysis device;at least one user interface connected to the at least one computational device; andat least one material dispensing device connected to the at least one computational device, the at least one material dispensing device including control circuitry for the at least one material dispensing device and at least one data storage unit including dispensation protocols. 15. The system... | Gearbox, LLC | G01N033/48 G01N033/50 A61B001/04 G06F019/12 A61B005/00 G06F019/00 G06F019/18
US20100298671 | A1 | 2010-11-25 | METHOD FOR ANALYZING ANALYTE IN TISSUE FLUID, ANALYZER FOR ANALYZING ANALYTE IN TISSUE FLUID, CARTRIDGE FOR ANALYZING ANALYTE IN TISSUE FLUID, AND KIT FOR ANALYZING ANALYTE IN TISSUE FLUID | 2010-05-25 | A method for analyzing an analyte in tissue fluid is disclosed. The method comprises: placing a collecting member in contact with the skin of a subject, wherein the collecting member is configured to collect and retain tissue fluid containing the analyte from the subject; collecting tissue fluid from the subject into the collecting member; detaching the collecting member from the skin of the subject; causing the analyte in the tissue fluid collected by the collecting member to diffuse in a liquid by supplying the collecting member with the liquid; and detecting the analyte diffused in the liquid.;An analyzer for analyzing an analyte in tissue fluid, and cartridge and kit for use in the analyzing method are also disclosed. | 1. A method for analyzing an analyte in tissue fluid, comprising: placing a collecting member in contact with a skin of a subject, wherein the collecting member is configured to collect and retain tissue fluid containing an analyte from the subject;collecting tissue fluid from the subject into the collecting member;detaching the collecting member from the skin of the subject;causing the analyte in the tissue fluid collected by the collecting member to diffuse in a liquid by supplying the collecting member with the liquid; anddetecting the analyte diffused in the liquid. 2. An analyzing apparatus for analyzing an analyte in tissue fluid, comprising: an accommodating part configured to accommodate a collecting member containing tissue fluid collected from a subject;a reservoir configured to retain liquid, wherein the reservoir is arranged in communication with the accommodating part;a liquid transporter configured to transport liquid to the accommodating part and to transport liquid to the reservoir from the accommodating part; anda detector for detecting an analyte in the liquid transported into the reservoir by the liquid transporter. 3. The analyzing apparatus according to claim 2, further comprising: a cartridge comprising the accommodating part and the reservoir; anda cartridge holder configured to detachably hold the cartridge. 4. The analyzing apparatus according to claim 3, wherein the cartridge comprises a reagent. 5. The analyzing apparatus according to claim 4, wherein the reagent causes a change in optical characteristics by reacting with the analyte in the liquid retained in the reservoir; andthe detector is configured to detect the analyte based on the change in the optical characteristics of the reagent. 6. The analyzing apparatus according to claim 2, wherein the detector comprises an electrode configured to contact the liquid retained in the reservoir. 7. The analyzing apparatus according to claim 2, wherein the reservoir comprises a first reservoir and a second reservoir; andthe detector comprises a first detector for detecting a first analyte in the liquid retained in the first reservoir, and a second detector for detecting a second analyte in the liquid retained in the second reservoir. 8. The analyzing apparatus according to claim 7, wherein the accommodating part and the first reservoir are connected via a flow path;the second reservoir is disposed in the flow path; andthe first reservoir comprises a reagent. 9. The analyzing apparatus according to claim 8, wherein the first detector is disposed opposite a first surface of the cartridge set in the cartridge holder; andthe second detector is disposed opposite a second surface of the cartridge set in the cartridge holder. 10. The analyzing apparatus according to claim 3, wherein the cartridge has a channel that communicates with the accommodating part and the reservoir when the cartridge is set in the cartridge holder. 11. The analyzing apparatus according to claim 2, ... | SYSMEX CORPORATION | A61B005/00
US20100280343 | A1 | 2010-11-04 | MULTIPLE WAVELENGTH PHYSIOLOGICAL MEASURING APPARATUS, SENSOR AND INTERFACE UNIT FOR DETERMINATION OF BLOOD PARAMETERS | 2009-10-28 | A measuring apparatus, a physiological sensor, and an interface unit for determining blood parameters of a subject are disclosed. The sensor comprises an emitter unit comprising a first plurality of emitter elements configured to emit radiation at a second plurality of wavelengths and a detector unit configured to receive radiation generated by the emitter unit and transmitted through tissue of the subject. The sensor further comprises a sensor memory storing sensor-specific information about the sensor unit, wherein the sensor-specific information includes at least calibration data for a given measurement mode, and a memory access interface for enabling an entity external to the sensor to update at least part of the sensor-specific information in a sensor ability update process, thereby to update ability of the sensor unit to operate in the given measurement mode. | 1. A measuring apparatus for determining the amount of at least one substance in blood of a subject, the measuring apparatus comprising: a sensor unit comprising an emitter unit comprising a first plurality of emitter elements configured to emit radiation at a second plurality of wavelengths; anda detector unit configured to receive radiation generated by the emitter unit and transmitted through tissue of a subject, wherein the detector unit is further configured to produce measurement signals indicative of absorption caused by blood of the subject;a first memory storing sensor-specific information about the sensor unit, wherein the sensor-specific information includes at least calibration data for calibrating the apparatus for a selected measurement mode that corresponds to a given combination of wavelengths; anda sensor ability maintenance unit configured to perform a sensor ability update process in which at least part of the sensor-specific information is updated, thereby to update ability of the sensor unit to operate in the selected measurement mode. 2. The measuring apparatus according to claim 1, wherein the sensor-specific information includes service information indicating whether the sensor ability update process needs to be initiated, in which the sensor ability update process is intended to improve performance of the sensor unit in the selected measurement mode. 3. The measuring apparatus according to claim 1, wherein the sensor-specific information further includes activation status information indicating whether the sensor ability update process needs to be initiated, in which the sensor ability update process is intended to activate resources of the sensor unit for the selected measurement mode. 4. The measuring apparatus according to claim 1, wherein the first memory is in the sensor unit. 5. The measuring apparatus according to claim 1, wherein the apparatus further comprises a user information display unit configured to indicate when the sensor ability update process needs to be performed. 6. The measuring apparatus according to claim 2, wherein the sensor ability maintenance unit is configured to update the calibration data and the service information in the sensor ability update process. 7. The measuring apparatus according to claim 3, wherein the sensor ability maintenance unit is configured to update the activation status information. 8. The measuring apparatus according to claim 1, wherein the sensor-specific information further comprises emitter activation data indicating how to switch on at least some of the first plurality of emitter elements, and wherein the sensor ability maintenance unit is configured to update the emitter activation data. 9. The measuring apparatus according to claim 6, wherein the sensor ability maintenance unit is further configured to collect history data concerning sensor usage and sensor-specific characteristics and update the service information based on the history data. 10. The meas... | GENERAL ELECTRIC COMPANY | A61B005/1455
US20100298682 | A1 | 2010-11-25 | Analyte Monitoring Device and Methods of Use | 2010-07-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for processing continuous glucose sensor data, the method comprising: receiving sensor data from a continuous glucose sensor, wherein the sensor data comprises one or more sensor data points;obtaining an estimated sensor glucose value from the one or more sensor data points;calculating at least two rate of change values; andfiltering the at least two rate of change values to obtain a filtered rate of change value. 3. The method of claim 2, wherein the at least two rate of change values are point-to-point rate of change values. 4. The method of claim 2, further comprising determining a predicted value for a future time period based on the estimated sensor glucose value, the filtered rate of change value, and a time to the future time period. 5. The method of claim 4, further comprising comparing the predicted value against a threshold. 6. The method of claim 5, further comprising triggering an alarm when the predicted value passes the threshold. 7. The method of claim 5, wherein the threshold is user selectable. 8. The method of claim 2, further comprising determining a predicted time to a threshold, wherein the predicted time is based at least in part on the estimated sensor glucose value, the filtered rate of change value, and the threshold. 9. The method of claim 8, wherein the threshold is user selectable. 10. The method of claim 8, further comprising outputting information related to the predicted time to the threshold on a user interface. 11. The method of claim 8, further comprising displaying the predicted time to the threshold on a user interface. 12. The method of claim 11, wherein the step of displaying the predicted time to the threshold is performed when the predicted time is below a predetermined value 13. The method of claim 2, further comprising determining an insulin therapy based at least in part on the filtered rate of change value. 14. The method of claim 2, wherein the step of filtering to obtain a filtered rate of change value is performed continuously. 15. The method of claim 2, wherein the step of filtering to obtain a filtered rate of change value is not performed when a level of noise is above a threshold. 16. The method of claim 2, further comprising displaying a trend arrow representative of the filtered rate of change values. 17. A system for processing continuous glucose sensor data, the system comprising: a continuous glucose sensor configured to generate sensor data associated with glucose concentration in a host; anda computer system that obtains an estimated sensor glucose value, calculates at least two rate of change values, and filters the at least two rate of change values to obtain a filtered rate of change value. 18. The system of claim 17, wherein the at least two rate of change values are point-to-point rate of change values. 19. The system of claim 17, wherein the computer system determines a predicted value for a future time period based on the estimated sensor ... | Abbott Diabetes Care Inc. | A61B005/05
US20100324378 | A1 | 2010-12-23 | PHYSIOLOGIC SIGNAL MONITORING USING ULTRASOUND SIGNALS FROM IMPLANTED DEVICES | 2010-05-18 | Devices, systems, and methods for monitoring and analyzing physiologic parameters within the body using an intrabody ultrasound signal are disclosed. An illustrative method includes receiving an ultrasound signal transmitted from a remote device containing encoded sensor data, converting the ultrasound signal into an electrical signal, decoding the sensor data from the electrical signal and generating a first physiological waveform, generating a second physiological waveform by analyzing fluctuations of the electrical signal caused by physiologic modulation of the ultrasound signal during propagation through the body, and analyzing one or more characteristics of the first and second waveforms to determine one or more physiologic parameters within the body. | 1. A method for determining one or more time-varying physiologic parameters within the body of a patient using intrabody ultrasound signals, comprising: receiving an ultrasound signal transmitted from a remote device located within the body, the ultrasound signal including encoded sensor data measured by the remote device;transducing the ultrasound signal into an electrical signal;decoding the sensor data from the electrical signal and generating a first physiological waveform corresponding to the sensor data measured by the remote device; andgenerating a second physiological waveform by analyzing fluctuations of the electrical signal caused by physiologic modulation of the ultrasound signal during propagation through the body. 2. The method of claim 1, further including analyzing at least one characteristic of the first and second physiological waveforms to determine one or more physiological parameters within the body. 3. The method of claim 1, wherein the first physiological waveform is a pressure waveform. 4. The method of claim 3, wherein the second physiological waveform is a respiration waveform. 5. The method of claim 3, wherein the second physiological waveform is a cardiac waveform. 6. The method of claim 1, further comprising using the one or more physiologic parameters to calibrate another device within the body. 7. The method of claim 1, wherein the remote device is a pressure sensor implanted within a pulmonary artery, and wherein the encoded sensor data comprises blood pressure data measured by the remote device within the pulmonary artery. 8. The method of claim 1, wherein generating a second physiological waveform includes filtering the electrical signal with a low-pass or band-pass filter having a bandwidth corresponding to the frequency range of a physiologic signal of interest. 9. The method of claim 2, wherein analyzing at least one characteristic of the first and second physiological waveforms to determine one or more physiologic parameters within the body includes detecting one or more peaks in the electrical signal and correlating the amplitude and timing of the peaks in the electrical signal with the measured sensor data from the first physiological waveform. 10. The method of claim 2, wherein analyzing at least one characteristic of the first and second physiological waveforms includes determining the end expiration stage of the patient's respiration cycle. 11. The method of claim 2, wherein analyzing at least one characteristic of the first and second physiological waveforms includes determining a respiration rate of the patient's respiration cycle. 12. The method of claim 2, wherein analyzing at least one characteristic of the first and second physiological waveforms includes determining a tidal volume of the patient's respiration cycle. 13. The method of claim 2, wherein analyzing at least one characteristic of the first and second physiological waveforms includes determining a heart rate. 14. The method ... | - | A61B005/0215 A61B005/00 A61B005/02 A61B005/08
US20100312083 | A1 | 2010-12-09 | System for Monitoring Glucose and Measuring Wattage | 2010-04-19 | A computerized system for monitoring information related to the power being generated by a user and blood glucose levels of the user, the system comprising: a shoe worn by the user, an electronic force sensor associated with the shoe, and a computer electronically coupled to the force sensor. | 1. A computerized system for monitoring information related to the power being generated by a user and blood glucose levels of the user, the system comprising: a shoe worn by the user, an electronic force sensor associated with the shoe, and a computer electronically coupled to the force sensor. 2. The system of claim 1, wherein the force sensor is configured to measure force applied by a foot of a user within the shoe. 3. The system of claim 1, further comprising a glucose monitor configured for generating signals associated with blood glucose of the user, the glucose monitor being electronically coupled to the computer, wherein the computer processes the signals corresponding to the blood glucose and the force signals relative to time, and correlates the force and the blood glucose relative to one another. | - | A61B005/145 A61B005/103
US20100284928 | A1 | 2010-11-11 | METHODS FOR IN VIVO ATHEROSCLEROTIC PLAQUE CHARACTERIZATION USING MAGNETIC SUSCEPTIBILITY TO IDENTIFY SYMPTOM-PRODUCING PLAQUES | 2009-01-02 | Methods and uses of in vivo atherosclerotic plaque characterization in a subject in need thereof, that include detecting a level of one or more iron complexes in plaque in the subject, and comparing detected level to a known value determined from subjects without an atherosclerotic condition are disclosed. | 1. A method for in vivo atherosclerotic plaque characterization in a subject in need thereof, comprising: detecting a level of one or more iron complexes in plaque in the subject, andcomparing detected level to a known value determined from subjects without an atherosclerotic condition. 2. The method of claim 1, comprising measuring carotid plaque using a T2*-weighted imaging to distinguish symptom-producing plaque from clinically-silent plaque in the subject. 3. The method of claim 2, comprising distinguishing magnetic susceptibility to symptom-producing plaque from clinically-silent plaque in the subject. 4. The method of claim 2, wherein the iron complex comprises paramagnetic-Fe(III) complexes. 5. A method of claim 4, wherein the presence of decreased levels of paramagnetic-Fe(III) complexes is indicative of symptom-producing plaque. 6. The method of claim 2, wherein greater concentrations of one or more iron complexes are present atherosclerotic plaque compared with normal arterial tissue. 7. The method of claim 2, wherein change in iron concentration is qualitatively measured using magnetic resonance T2*-weighted imaging. 8. The method of claim 2, wherein change in iron concentration is quantified using a T2* relaxation parameter. 9. The method of claim 8, wherein T2* quantification allows for a substantially accurate estimation of tissue iron content in the plaque. 10. The method of claim 9, using MRI imaging of the plaque in vivo. 11. A method for detecting an atherosclerotic condition in a subject, comprising: a) acquiring MRI data from the subject with a magnetic resonance imaging system that measures T2* in plaque present in the subject; andb) comparing the MRI data of step a) to one or more known parameters;c) indicating a difference between the data of step a) with the known parameters. 12. The method of claim 11, comprising: obtaining at least a first set of image data created in response to an MRI scan of at least a portion of plaque in the subject, wherein the MRI scan includes T2* weighted imaging; anddetermining the presence of at least one iron complex present in the plaque;wherein the presence of the at least one iron complex at an altered level is indicative of symptom-producing plaque in the subject. 13. The method of claim 12, wherein the subject being scanned was injected with a contrast agent. 14. A system for plaque characterization, comprising: an imaging system generating at least a first set of image data of plaque in a subject in response to a first magnetic energy level; the first set of image data corresponding to a T2*-weighted imaging of the plaque, anda processing device in communication with the imaging system obtaining at least the first set of image data from the imaging system and displaying an indication of one or more iron complexes present in the plaque. 15. The system of claim 14, wherein the imaging system comprises a magnetic resonance imaging system. 16. The system of claim 15, wherein ... | THE OHIO STATE UNIVERSITY RESEARCH FOUNDATION | A61K049/06 A61K049/00 A61B005/055 G06K009/00
US20100292591 | A1 | 2010-11-18 | Venous measurement system | 2010-05-14 | A system and plethysmographic test procedure for evaluating venous blood volume/flow at a subject's legs. Occluding and sensing cuffs mounted to a subject's thighs and calves/ankles are inflated and deflated relative to controlled movement of a supporting chair. The chair provides hinged backrest, seat and footrest pieces and is controlled to manipulate the subject during several test phases as blood volume/flow is monitored. The chair is manipulated between a 1) baseline calibration position with the subject's back upright, seat horizontal to floor and feet lowered, 2) an outflow position with the subject's back tipped back flat and legs elevated above the heart, 3) a dependent filling position, rapid return of subject to back upright, seat horizontal to floor and feet lowered, and 4) an exercise position wherein the ankles are flexed with the back upright, seat horizontal to the floor and feet lowered. “Outflow”, “dependent filling”, “exercise” and “ejection fraction” flow values are determined relative to measured tracings exhibiting venous volume/flow changes in the legs. | 1. Apparatus for evaluating venous blood flow comprising: a) a chair including a framework supporting a backrest, a seat and a footrest and wherein said backrest and footrest are mounted to pivot relative to said seat;b) an air pump;c) deflation means;d) a pressure monitor adapted to be coupled to a leg of a test subject, wherein the monitor comprises 1) inflatable occlusion and sensor cuffs coupled to said air pump and said deflation means and wherein said sensor cuff includes a pressure sensor and 2) means for controlling i) said air pump to inflate said sensor cuff to establish at least two pressure levels to determine a baseline condition relative to air volume, ii) said air pump to inflate said occlusion cuff to occlude blood flow through the leg, and iii) said deflation means to deflate said occlusion and sensor cuffs;e) means for positioning said chair between a first condition wherein said backrest extends transverse to said footrest and seat and said seat and footrest are vertically displaced apart in parallel relation to each other and a second condition wherein said backrest and said seat align in coplanar horizontal alignment and said footrest transversely extends above said seat and backrest; andf) means coupled to monitor and visually display venous blood flow pressures continuously sensed by said sensor cuff at said first and second chair conditions relative to inflated and deflated conditions of said occlusion cuff and to leg movements and display time based blood flow volume tracings and % ratios determined from said tracings indicative of the vascular condition of the monitored leg. 2. Apparatus as set forth in claim 1 wherein occlusion and sensor cuffs of a second pressure monitor are coupled to a second leg of the subject. 3. Apparatus as set forth in claim 1 wherein said footrest is supported to said backrest to pivot to a third condition wherein said backrest, seat and foot rest are coplanar to each other. 4. Apparatus as set forth in claim 3 wherein said backrest includes a first support member, wherein said footrest includes a second support member and wherein said first and second support members are mounted to pivot relative to each other to establish said first, second and third chair conditions. 5. Apparatus as set forth in claim 4 wherein said footrest comprises a platform mounted to move to and fro relative to said second support. 6. Apparatus as set forth in claim 5 wherein said footrest comprises a leg member mounted to pivot between parallel and orthogonal orientations to said platform whereby the footrest platform is supported to a chair support surface at said third chair condition. 7. Apparatus as set forth in claim 4 wherein including means for biasing the rotation of said footrest to prevent freefall. 8. Apparatus for evaluating venous blood flow comprising: a) a chair including a framework supporting a backrest, a seat and a footrest and wherein said backrest and footrest are mounted to pivot relative... | - | A61B005/026
US07840421 | B2 | 2010-11-23 | Infectious disease surveillance system | 2002-07-31 | An infectious disease surveillance system is disclosed, which comprises a database system for storing data relating to at least one infectious disease; input means for providing data to the database system; a complex adaptive system associated with the database system; processing means for processing the data of the database system and converting the data into surveillance data; and output means for displaying the surveillance data. | 1. An infectious disease surveillance system, which comprises: (a) a database system for storing data relating to at least one infectious disease;(b) input means for providing biological data to the database system, said input means being remote from the database system and configured to be mobile;(c) computer processing means associated with the database system and for processing biological data of the database system and converting the processed data into surveillance data;(d) a computational complex adaptive system being operatively associated with the computer processing means and being associated with the database system, the complex adaptive system configured for adapting continuously according to the biological data provided by the input means and the surveillance data;(e) at least one model selected from the group comprising an analytical model and an empirical model associated with the complex adaptive system, the model having parameters which are capable of being continuously adapted;(f) output means for displaying the surveillance data generally instantly after said input means provides said biological data to said database system; and(g) decision-making means, responsive to said surveillance data, for making policy decisions regarding the surveillance data. 2. A system as claimed in claimed 1, in which the complex adaptive system is capable of adapting according to characteristics of the input means providing the data. 3. A system as claimed in claimed 1, in which the complex adaptive system comprises at least one of the elements selected from the group comprising a Bayesian network, a neural network, a genetic algorithm and an artificial life system. 4. A system as claimed in claimed 1, which is deployed in a distributed environment. 5. A system as claimed in claimed 1, in which the input means are connected real-time to the database system. 6. A system as claimed in claimed 1, in which the database comprises real-time data. 7. A system as claimed in claimed 1, which requires an identification code for transferring data from the input means to the database system. 8. A system as claimed in claimed 1, which comprises filtering means for filtering data received from the input means. 9. A system as claimed in claimed 1, in which the infectious disease is an epidemic disease. 10. A system as claimed in claimed 1 which comprises data selected from the group comprising incidence data, prevalence data, death data, diagnosed data, estimated data, extrapolated data and predicted data. 11. A system as claimed in claimed 10, in which the estimated data, extrapolated data and predicted data are obtained by means of the model. 12. A system as claimed in claim 1, which comprises pharmaceutical data relating to any one selected from the group comprising genetic differences of patients infected with the infectious disease, pharmaceutical interactions, pharmaceutical-disease interactions, effects of placebos, and pharmaceutical treatment c... | - | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US20100331630 | A1 | 2010-12-30 | MACHINE, MANUFACTURE, AND PROCESS FOR ANALYZING THE RELATIONSHIP BETWEEN DISPOSABLE DIAPER WEAR WITH SLEEP AND/OR DEVELOPMENTAL INDICATORS | 2009-06-25 | A machine or system, article of manufacture or computer program product, or process for facilitating the analysis of the effect of diaper wear on sleep and/or developmental indicators including verifying a test diaper was secured on at least one of a first plurality of subjects; monitoring the subject for sleep initiation data; evaluating the sleep of the subject based upon sleep evaluation data. | 1. An analysis machine configured to facilitate analysis of the effect of diaper wear on sleep and/or developmental indicators comprising: a disposable diaper verification module configured to accept verification input indicating that a disposable diaper was secured on at least one of a first plurality of subjects;a sleep initiation monitoring module configured to monitor the subject for sleep initiation data;a sleep evaluation module configured to evaluate the sleep of the subject based upon sleep evaluation data. 2. The machine according to claim 1, further comprising a developmental indicator input module configured to accept data of a developmental indicator of the subject of the first plurality, anda display module configured to display the developmental indicator data or the sleep evaluation data of the subject of the first plurality,wherein the developmental indicator data and the sleep evaluation data are collected during a coinciding time duration. 3. The machine according to claim 2, further comprising a cloth diaper verification module configured to accept verification input indicating that a cloth diaper was secured to at least one of a second plurality of subjects;a sleep initiation monitoring module configured to monitor the subject of the second plurality for sleep initiation data;a sleep evaluation module configured to evaluate the sleep of the subject of the second plurality based upon sleep evaluation data;a developmental indicator input module configured to accept data of a developmental indicator of the subject of the second plurality, anda display module configured to display the developmental indicator data or the sleep evaluation data of the subject of the second plurality,wherein the developmental indicator data and the sleep evaluation data of the subject of the second plurality are collected during a coinciding time duration. 4. The machine according to claim 3, further comprising a comparator module configured to compare the developmental indicator data of the subject of the first plurality and the developmental indicator data of the subject of the second plurality. 5. The machine according to claim 2, further comprising a comparator module configured to compare data of the sleep of the subject of the first plurality and the sleep of the subject of the second plurality. 6. The machine according to claim 1, wherein the sleep initiation monitoring module or the sleep evaluation module is any one or more of a group consisting of: a polysomnography device, an actigraphy device, or a video monitoring device, or a remote sensing device configured to utilize Doppler radiofrequency for remote sensing of physical activity and respiratory rate, or software configured to enable processing of a physical motion signal and/or a respiratory signal, or a computational algorithm configured to enable processing of a physical motion signal and/or a respiratory signal, or a remote sensing device configured to determine a sleep/wake st... | - | A61B005/00 G09B019/00
US20100312104 | A1 | 2010-12-09 | SYSTEM AND METHOD FOR CALIBRATING AND POSITIONING A RADIATION THERAPY TREATMENT TABLE | 2010-01-04 | A system and method for calibrating and positioning a radiation therapy treatment table with a high degree of precision and accuracy. The system and method comprising a treatment table for use with a radiation therapy system with radiographic indexing markers embedded or inlaid in the patient treatment table. The radiographic indexing markers having a density different from density of the table material, or a radio-frequency or magnetic contrasting matter, something that the imaging system and detector of the radiation therapy system can detect, so that when scanned, with or without a patient or phantom on the table, the precise location of the patient treatment table in the x, y and z planes is known. The radiographic indexing markers may also have the ability to be visually seen by the technician, doctor, physicist or other health care provider conducting the procedure. This is important so that the table may be reconstructed, alone or with a phantom on it, by the radiation therapy system software, providing a very good indication of where the table top is prior to treatment. | 1. A system for calibrating and positioning a radiation therapy treatment table of a radiation therapy system comprising at least one radiographic indexing marker included in the radiation therapy treatment table for precisely locating and calibrating the radiation therapy treatment table in three-dimensional space. 2. The system of claim 1, wherein the at least one radiographic indexing marker has a density different from the density of the radiation therapy treatment table material. 3. The system of claim 1, wherein the at least one radiographic indexing marker is detectable by the radiation therapy system. 4. (canceled) 5. (canceled) 6. The system of claim 1, wherein the at least one radiographic indexing marker is made of a high density material. 7. The system of claim 1, wherein the at least one radiographic indexing marker is made of a low density material. 8. The system of claim 1, wherein the at least one radiographic indexing marker is permanently fixed in the radiation therapy treatment table. 9. The system of claim 1, wherein the at least one radiographic indexing marker is removable from the radiation therapy treatment table. 10. (canceled) 11. The system of claim 1, wherein the at least one radiographic indexing marker is embedded or inlaid into the surfaces of the radiation therapy treatment table. 12. (canceled) 13. The system of claim 1, wherein the at least one radiographic indexing marker is an indentation or groove in the radiation therapy treatment table. 14. (canceled) 15. The system of claim 1, wherein the at least one radiographic indexing marker is magnetic and sensors may be used for positioning a patient and/or a fixation device. 16. The system of claim 1, wherein the at least one radiographic indexing marker is a laser. 17. The system of claim 1, wherein the at least one radiographic index marker is an RF emitter. 18. The system of claim 1, wherein the at least one radiographic index marker is an RF detector. 19. The system of claim 1, wherein the at least one radiographic index marker is a solid. 20. The system of claim 1, wherein the at least one radiographic index marker a disjoint solid. 21. The system of claim 1, wherein the at least one radiographic index marker is a liquid. 22. The system of claim 1, wherein the at least one radiographic index marker is interchangeable with different sets or subsets of markers utilized for different objectives. 23. The system of claim 1, wherein the at least one radiographic index marker is used to send and receive signals with other fixtures to detect changes in table height or deformations, such as sags, and automatically, semi-automatically, or manually adjust the table position before or during radiotherapy treatment procedures. 24. The system of claim 1, wherein the at least one radiographic index marker is used for calibration and verification of table speed, table position and virtual isocenter. 25. A system for precisely locating the position of a ... | - | A61B005/05
US07859671 | B2 | 2010-12-28 | Method for determining optical properties of turbid media | 2005-12-22 | There is provided a method for improving contrast and resolution of an optical image of an object obtained by time-resolved techniques such as Time Domain (TD) and Frequency Domain (FD). The method comprises obtaining a Temporal Point Spread Function (TPSF), and determining optical properties of volumes of interest (VOI), each volume being defined by an ensemble of equiprobable effective photon paths corresponding to a time point or time gate of the TPSF. | 1. A method for determining an optical absorption coefficient of a volume of interest (VOI) in a turbid medium using time resolved optical modalities, said method comprising: injecting light in said medium at an injection point;detecting light exiting said medium at a detection point;generating a Temporal Point Spread Function (TPSF) using said detected light;correlating a predetermined time point or time gate of said TPSF with said VOI wherein said VOI comprises equiprobable effective photon paths; anddetermining said optical absorption coefficient within said VOI using one or more characteristics of said TPSF. 2. The method as claimed in claim 1 wherein said absorption coefficient is determined as a function of a light transmission equation incorporating said one or more characteristics of said TPSF. 3. The method as claimed in claim 2 wherein said light transmission equation is the Beer-Lambert equation. 4. The method as claimed in claim 3 wherein said step of determining said absorption coefficient comprises the step of estimating an intensity value for said injected light to calculate said absorption coefficient using said Beer-Lambert equation. 5. The method as claimed in claim 4 further comprising the step of determining a concentration of a chromophore using said intensity value and said Beer-Lambert equation. 6. The method as claimed in claim 5 wherein the concentration of two or more chromophores is determined by obtaining two or more TPSF's at two or more wavelengths. 7. The method as claimed in claim 3 wherein said Beer-Lambert equation is used for determining concentration of a chromophore by using the ratio of intensity of said TPSF at said time point or time gate for said VOI with and without said chromophore. 8. The method as claimed in claim 4 wherein said intensity value is determined as a function of the rate of decay of the asymptotic portion of said TPSF and a time point or time gate within said asymptotic portion. 9. The method as claimed in claim 1 wherein said determination of said absorption coefficient is a relative determination of two VOI's. 10. The method as claimed in claim in claim 1 wherein said TPSF is obtained in transmission mode. 11. The method as claimed in claim 1 wherein said TPSF is obtained in the reflection mode. 12. The method as claimed in claim 1 wherein said VOI is an ellipsoid. 13. The method as claimed in claim 1 wherein said time point or time gate is equal to or greater than time point corresponding to the maximum of said TPSF. 14. The method as claimed in claim 1 wherein said optical absorption coefficient is mapped for more than one VOI thereby producing an image of optical absorption coefficient of said medium in a region of interest (ROI). 15. The method as claimed in claim 14 wherein a plurality of TPSF's are obtained from a plurality of injection/detection points configurations to produce a plurality of images of said VOI and wherein said images are combined to produce an image... | ART, Advanced Research Technologies Inc. | G01N021/00 G01N021/64 A61B005/00
US07826887 | B2 | 2010-11-02 | MR head coil | 2007-10-24 | An MR head coil has a helmet-shaped coil housing and an arrangement of a number of first antenna elements mounted on the coil housing. The helmet-shaped coil housing is shaped such that it completely frees the field of view of a patient, and second antenna elements or antenna element parts that can be temporarily arranged in the field of view are provided. | 1. A magnetic resonance head coil comprising: a helmet-shaped head coil housing comprising a frontal opening therein that, when said head coil housing is worn on the head of a patient, frees the patient's field of view;a plurality of magnetic resonance antenna elements mounted at said head coil housing, each of said magnetic resonance antenna elements being configured to at least receive magnetic resonance signals originating in the head of the patient;said plurality of antenna elements including two antenna elements each comprising a first antenna element part fixedly mounted to said head coil housing adjacent said frontal opening, and a second antenna element part mounted to said head coil housing for sliding movement relative to the first antenna element part, each of said first and second antenna element parts having an arcuate shape and each comprising a first free end and a second free end;for each of said two antenna elements, the second antenna element part thereof being slidably movable along an arcuate path from a standby position on or in said head coil housing to an operating position in said frontal opening that does not substantially obstruct the patient's field of view; andfor each of said two antenna elements, the first and second antenna element parts thereof comprising a slip contact, at the respective first ends of said first and second antenna element parts, that makes an electrical connection between the first and second antenna element parts while allowing sliding movement therebetween, and a plug connection, at the respective second ends of said first and second antenna element parts, that forms a closed loop between said first antenna element part and said second antenna element part in said operating position. 2. A magnetic resonance head coil as claimed in claim 1 wherein, for each of said two antenna elements, the first antenna element part thereof comprises a first segment of a circle, and the second antenna element part thereof comprises a second segment of said circle, and wherein said second antenna element part slidably moves along a circular path from said standby position to said operating position. 3. A magnetic resonance head coil as claimed in claim 2 wherein said two antenna elements are positioned at said head coil housing to cause the respective circles, formed by the respective first and second antenna element parts thereof, to surround the respective eyes of the patient when the head coil holder is worn at the head of the patient. | Siemens Aktiengesellschaft | A61B005/05 G01V003/00
US07841445 | B2 | 2010-11-30 | Dual-sensor stethoscope with electronic sensor | 2009-04-07 | A dual-sensor stethoscope, or retrofit device for a stethoscope, promotes anti-sepsis and stereoscopy through use of a substantially rigid, generally T-shaped tube for support of dual stethoscope sensors, wherein at least one sensor is electronic. Each sensor may be rotated away from a body independently for use of a single head or rotated into the same general plane for dual-head stereoscopy. A clinician can create a spatial, three-dimensional (3D) antiseptic barrier and avoid the need to carry multiple stethoscopes. During stereoscopy, the use of a common tube or earphones allows the transmission of sound with constructive interference of sound waves. The substantially rigid, generally T-shaped support tube allows a clinician to auscultate with one hand and monitor two locations without transference of pathogens. | 1. A dual-head anti-sepsis stethoscope device comprising: a connection tube dimensioned to engage a sound conduction tube of a stethoscope;a second tube connected to the connection tube, the second tube having a first end and a second end;a first sensor head-supporting tube coupled to the first end of the second tube;a second sensor head-supporting tube coupled to the second end of the second tube; wherein at least one of the first and second sensor head-supporting tubes is rotatably coupled relative to the second tube;an electronic sensor incorporating a microphone attached to the first sensor head-supporting tubes at an end opposite the second tube;a processor for processing signals from the microphone to produce processed audio signals; andan audio output means for receiving the processed audio signals, wherein said audio output means is positioned in one of the locations selected from the group consisting of the connection tube, the second tube, and the first sensor head-supporting tube. 2. The dual-head anti-sepsis stethoscope device of claim 1, wherein the first and second sensor head-supporting tubes are rotatably coupled relative to the second tube. 3. The dual-head anti-sepsis stethoscope device of claim 1, further comprising: a detent or locking means for restraining movement of the at least one sensor head-supporting tube that is rotatably coupled relative to the second tube. 4. The dual-head anti-sepsis stethoscope device of claim 1, further comprising: a second electronic sensor comprising a second microphone and attached to the second sensor head-supporting tubes at an end opposite the second tube;the processor further processing signals from the second microphone to produce the processed audio signals; andthe audio output means positioned in the connection tube. 5. The dual-head anti-sepsis stethoscope device of claim 1, wherein portions of the device other than the processor are constructed from disposable materials. 6. The dual-head anti-sepsis stethoscope device of claim 1, wherein the processor comprises manually operated controls. 7. The dual-head anti-sepsis stethoscope device of claim 1, wherein the second tube further comprises at least one telescoping element. 8. The dual-head anti-sepsis stethoscope device of claim 1, wherein the second tube is connected to the connection tube at a location between the first and second ends. 9. The dual-head anti-sepsis stethoscope device of claim 4, wherein the processor comprises separate manually operated controls for each electronic sensor. 10. The dual-head anti-sepsis stethoscope device of claim 1, wherein the connection tube is directly connected to one of the first and second sensor head-supporting tubes. 11. A dual-head anti-sepsis stethoscope, comprising: an eartip;an earpiece attached to the eartip;a common sound conduction tube attached to the earpiece;a connection tube dimensioned to engage the common sound conduction tube at an end opposite an end attached to the e... | - | A61B007/02 A61B007/00 A61B005/02
US20100331627 | A1 | 2010-12-30 | ADHERENCE INDICATION TOOL FOR CHRONIC DISEASE MANAGEMENT AND METHOD THEREOF | 2009-06-26 | An adherence indication tool for chronic disease self-management and method thereof for measuring adherence or compliance to following or achieving prescribed therapy steps to achieve stated target goals for improved chronic disease self-management are disclosed. | 1. A method for measuring adherence to following or achieving prescribed therapy steps to achieve stated target goals for improved chronic disease self-management comprising: defining a plurality of adherence units, each adherence unit containing a plurality of rules governing activities which need to be accomplish in order to complete the prescribed therapy steps;collecting data when the activities are accomplished;specifying a time window of interest in the collected data;determining total number of adherence units in the collected data which fall within the specified time window of interest;counting each of the adherence units in the specified time window of interest as an adhered unit when the collected data indicates the accomplished activities were in accordance to the rules;determining adherence as a percentage of the count for the adhered units to the total number of adherence units for the specified time window; andproviding at least one of the determined adherence percentage and adherence count for the specified time window. 2. The method of claim 1 further comprises selecting from a library contained in memory of a processing device the plurality of rules governing activities. 3. The method of claim 1 further comprises inputting into memory of a processing device the collected data. 4. The method of claim 1 further comprises setting a sequence and timing of the activities contained in each of the adherence units. 5. The method of claim 4 further comprising providing a user interface of a processing device by which to set the sequence and timing of the activities in each of the adherence units. 6. The method of claim 1 further comprises using a processing device to perform the collecting of the data. 7. The method of claim 1 further comprises programming a processing device to prompt for information regarding each of the activities when accomplished. 8. The method of claim 1 further comprises storing the collected data in memory of a processing device. 9. The method of claim 1 further comprises annotating the collected data with a timestamp of start and completion. 10. The method of claim 1 wherein the specified time window of interest is defined by a starting time and date, and an ending time and date. 11. An adherence indication tool measuring adherence to following or achieving prescribed therapy steps to achieve stated target goals for improved chronic disease self-management comprising: a memory containing data collected when the activities were accomplished;a user interface facilitating selection of a plurality of adherence units, each adherence unit containing a plurality of rules governing activities which need to be accomplish in order to complete the prescribed therapy steps, and inputting of a specified time window of interest for the collected data;a process determining total number of adherence units in the collected data which fall within the specified time window of interest;a process counting each of the adhere... | ROCHE DIAGNOSTICS OPERATIONS, INC. | A61B005/00
US07860545 | B2 | 2010-12-28 | Analyte measuring device | 2008-02-26 | An implantable analyte-measuring device including a membrane adapted to promote vascularization and/or interfere with barrier cell layer formation. The membrane includes any combination of materials, architecture, and bioactive agents that facilitate analyte transport to provide long-term in vivo performance of the implantable analyte-measuring device. | 1. A device for subcutaneous monitoring of glucose levels, comprising: a housing;a sensor; anda porous biointerface layer for promoting long term microcirculatory delivery of glucose and oxygen to the sensor for a time period of at least about one month, wherein the porous biointerface layer comprises an architecture that supports tissue ingrowth and is configured to maintain sufficient fluid flow to the sensor for long term glucose measurement, wherein the porous biointerface layer further comprises a first bioactive agent configured to be released over a first time period of from about one day to about one month, wherein release of the first bioactive agent is configured to increase a rate at which the device becomes functional by at least partially overcoming sleep period effects associated with a temporary decrease in sensor sensitivity associated with tissue ingrowth into the porous biointerface layer; wherein the porous biointerface layer further comprises a second bioactive agent configured to be released over a second time period, wherein release of the second bioactive a ent is configured to modulate a foreign body response to reduce a continual decline in sensor sensitivity over a life of a sensor. 2. The device of claim 1, wherein the biointerface layer comprises a material selected from the group consisting of hydrophilic polyvinylidene fluoride, mixed cellulose esters, polyvinyl chloride, polyvinyl alcohol, polyethylene, polytetrafluoroethylene, cellulose acetate, cellulose nitrate, polycarbonate, nylon, polyester, mixed esters of cellulose polyvinylidene difluoride, silicone, polyacrylonitrile, polypropylene, polysulfone, polymethacrylate, and mixtures thereof. 3. The device of claim 1, wherein the biointerface layer comprises expanded polytetrafluoroethylene. 4. The device of claim 1, wherein the biointerface layer comprises silicone. 5. The device of claim 1, wherein at least one of the first bioactive or the second bioactive agent is selected from the group consisting of anti-inflammatory agents, anti-infective agents, anesthetics, inflammatory agents, growth factors, immunosuppressive agents, antiplatelet agents, anticoagulants, ACE inhibitors, cytotoxic agents, anti-barrier cell compounds, vascularization-inducing compounds, anti-sense molecules, and mixtures thereof. 6. The device of claim 1, wherein at least one of the first bioactive agent or the second bioactive agent is an anti-inflammatory agent selected from the group consisting of nonsteroidal anti-inflammatory drugs (NTHEs), aspirin, celecoxib, choline magnesium trisalicylate, diclofenac potassium, diclofenac sodium, diflunisal, etodolac, fenoprofen, flurbiprofen, ibuprofen, indomethacin, ketoprofen, ketorolac, melenamic acid, nabumetone, naproxen, naproxen sodium, oxaprozin, piroxicam, rofecoxib, salsalate, sulindac, tolmetin, corticosteroids, cortisone, hydrocortisone, methylprednisolone, prednisone, prednisolone, betamethesone, beclomethasone dipropionate, bude... | DexCom, Inc. | A61B005/00
US07834890 | B2 | 2010-11-16 | Information processing method and image processing method | 2004-10-15 | An image processing method includes an object level acquisition step of acquiring an object level set for each computer graphics object, a display step of displaying a slide bar with a movable cursor to allow a user to set a user level by manipulating the cursor, and a user level acquisition step of acquiring a user level set by the user, based on a position of the cursor of the slide bar. In addition, a rendering step renders a virtual space including each computer graphics object by rendering each computer graphics object to have its visibility attribute depending on its object level acquired in the object level acquisition step and the user level. The rendering step renders a computer graphics object with an object level less than the user level as a visible object and renders a computer graphics object with an object level equal to or larger than the user level as an invisible object, and the rendering step is repeatedly performed in response to the position of the cursor being changed by the user. | 1. A method of processing a plurality of computer graphics objects in a 3-dimensional virtual space, comprising: an object level acquisition step of acquiring respective object levels set for each of a plurality of computer graphics objects;a user level determination step of determining a user level set by the user;a calculation step of calculating respective opacity attribute values of the plurality of computer graphics objects from a ratio of the acquired object levels of each of the plurality of computer graphics objects and the determined user level, wherein the opacity attribute values are proportional to the ratio and are calculated by multiplying the ratio of the determined user level to the acquired object levels by a predetermined maximum value of opacity;a comparing step of comparing the opacity attribute values calculated in the calculation step to the predetermined maximum opacity to determine an output opacity for each of the plurality of computer graphics objects, wherein (i) if the calculated opacity attribute value is equal to or larger than the predetermined maximum opacity, the opacity of the object is set to be the predetermined maximum opacity, (ii) if the calculated opacity attribute value is smaller than the predetermined minimum opacity, the opacity of the object is set to be the predetermined minimum opacity, and (iii) if the calculated opacity attribute value is smaller than predetermined maximum opacity and equal to or larger than the predetermined minimum opacity, the opacity of the object is set to be the opacity attribute value calculated for the plurality of computer graphics object in the calculation step; anda rendering step of rendering the 3-dimensional virtual space comprising the plurality of computer graphics objects by rendering, based on a predetermined angle of view and the opacity attribute values determined in the comparing step, each of the plurality of computer objects,wherein the rendering step is repeatedly performed in response to a determination of the user level in the user level determination step. 2. A non-transitory computer-readable medium encoded with a computer program for making a computer execute a method of processing a plurality of computer-graphics objects in 3-dimensional virtual space, comprising: an object level acquisition step of acquiring respective object levels set for each of a plurality of computer graphics objects;a user level determination step of determining a user level set by the user;a calculation step of calculating respective opacity attribute values of the plurality of computer graphics objects from a ratio of the acquired object levels of each of the plurality of computer graphics objects and the determined user level, wherein the opacity attribute values are proportional to the ratio and are calculated by multiplying the ratio of the determined user level to the acquired object levels by a predetermined maximum value of opacity;a comparing step of comparing the opac... | Canon Kabushiki Kaisha | G09G005/00 G06T015/00 G06K009/00 A61B005/05 A61B008/00
US07824333 | B2 | 2010-11-02 | Diabetes management methods and systems | 2006-03-31 | The present invention relates to methods and systems for monitoring the effectiveness of diabetes treatment. Methods and systems in accordance with the present invention provide information relating to variability of glucose levels and hypoglycemia and hyperglycemia. Such information is based on time-stamped blood glucose data obtained from a meter or the like and actual measurements of HbA1c levels are not required. | 1. A system for managing blood glucose data, the system comprising: a processor programmed to: receive a plurality of blood glucose data collected over a predetermined period of time;provide an estimation of an idiosyncratic HbA1c level from the plurality of blood glucose data;provide a calculation of an idiosyncratic variability index of glycemia from the plurality of blood glucose data; anddetermine whether at least one exclusion criteria has been met and report the idiosyncratic variability index of glycemia and the idiosyncratic HbA1c level when at least one exclusion criterion is not met and does not report the idiosyncratic variability index of glycemia and the idiosyncratic HbA1c level when the at least one exclusion criteria is met and further, wherein the at least one exclusion criterion includes a frequency of tests per day. 2. The system of claim 1, wherein the processor is programmed to report the idiosyncratic variability index of glycemia as a function of the idiosyncratic HbA1c level on a weekly basis. 3. The system of claim 1, wherein the processor is programmed to report the idiosyncratic variability index of glycemia as a function of the idiosyncratic HbA1c level on a monthly basis. 4. The system of claim 1, wherein the processor is programmed to report the idiosyncratic variability index of glycemia and the idiosyncratic HbA1c level in tabular form. 5. The system of claim 1, wherein the processor is programmed to estimate the idiosyncratic HbA1c level from a previously measured HbA1c level. 6. The system of claim 1, further comprising a blood glucose measurement device. 7. The system of claim 1, further comprising a memory device. | LifeScan, Inc. | A61B005/00
US07854071 | B2 | 2010-12-21 | Customized shoe and insole, method and apparatus for determining shape of a foot and for making a shoe or insole | 2010-03-24 | A method and kit of parts for measuring the comfortable shape of a foot for a person. The kit comprises front and rear platforms of adjustable height and a supporting surface for accommodating the plantar surface of a foot extending between said platforms. The aforementioned components can be adjusted to various settings with a view to determining a comfortable plantar foot shape. The foot shape can then be captured by scanning with an optical or mechanical probe or scanner. A method of making a customized shoe and shoe last on basis of the captured foot shape is also disclosed. | 1. A method of measuring the plantar surface which a foot will assume in a shoe, comprising the steps of: a) providing an apparatus having a rear platform for supporting the rear of foot, a front platform for supporting a fore foot and a supporting surface comprising a strip of flexible material extending between and mounted to the front and rear platforms;b) adjusting the height and/or angle of inclination of one or both of the front and rear platforms to a predetermined or comfortable setting;andc) placing the rear of a foot on the rear platform, the forefoot on the front platform and resting the middle of the foot on the supporting surface;d) capturing the shape of the plantar surface of the foot placed on the supporting surface. 2. The method of claim 1 wherein the foot is placed on the supporting surface before or after adjustment of the apparatus to a predetermined or comfortable setting and wherein the shape of the foot is captured in the predetermined or comfortable setting. 3. The method of claim 1 wherein the apparatus has a mid- foot supporting member for supporting a middle portion of the supporting surface and wherein the position of said middle foot supporting member is adjusted in order to adjust the curvature of the supporting surface. 4. The method of claim 3 wherein the height, angle or lateral position of the middle foot supporting member relative to the rear platform is adjusted. 5. The method of claim 1 wherein step d) is carried out by using a mechanical or optical probe or optical scanner. 6. The method of claim 1 wherein in step d) both the plantar and dorsal surfaces of the foot are captured. 7. A method of manufacturing customized or personalized shoes, insoles for shoes or a shoe last for making a shoe, comprising manufacturing a shoe, insole or shoe last on the basis of measurements of a foot and the shape of the plantar surface of a foot which has been captured by using the method of claim 1. 8. The method of claim 7 wherein the shape of the plantar surface has been captured by using a mechanical or optical probe or an optical scanner. 9. The method of claim 7 wherein the shape of the dorsal surface of the foot has been captured and the shoe is manufactured on the basis of said measurements and captured shapes of said dorsal and plantar surfaces of the foot. 10. The method of claim 7 comprising the step of making a shoe last on the basis of said captured plantar surface shape and measurements and manufacturing a shoe on the basis of said shoe last. 11. The method of claim 10 wherein the captured plantar foot shape and measurements are manually entered to a shoe last making machine or software. 12. The method of claim 10 wherein the captured plantar foot shape and measurements are automatically communicated to a shoe last making machine or software. 13. A customized or personalized shoe, shoe last or insole which has been manufactured according to the method of claim 7. 14. An apparatus for measuring or de... | The Hong Kong University of Science and Technology | A61B005/103
US08048321 | B2 | 2011-11-01 | Apparatus and method for separating and concentrating fluids containing multiple components | 2010-08-11 | An apparatus that allows for separating and collecting a fraction of a sample. The apparatus, when used with a centrifuge, allows for the creation of at least three fractions in the apparatus. It also provides for a new method of extracting the buffy coat phase from a whole blood sample. A buoy system that may include a first buoy portion and a second buoy member operably interconnected may be used to form at least three fractions from a sample during a substantially single centrifugation process. Therefore, the separation of various fractions may be substantially quick and efficient. | 1. A method of withdrawing a material directly from a patient and collecting a selected fraction of the material, comprising: forming an access port to the patient;forming a pressure differential between a collection container and the patient;connecting the collection container to the port;filling the collection container to a selected volume with the material; andcentrifuging the collection container containing the selected volume of the material to separate the material to form the selected fraction, wherein a separating member moves relative to a boundary of the selected fraction to substantially form a mechanical separation of the selected fraction. 2. The method of claim 1, wherein forming an access port to the patient includes at least one of positioning a member substantially intravenously in the patient, positioning a syringe substantially intravenously in the patient, or combinations thereof. 3. The method of claim 2, further comprising interconnecting the access port in the patient with the collection container such that a material in the intravenous portion is operable to move substantially directly to the collection container. 4. The method of claim 1, wherein forming a pressure differential in the collection container includes forming a pressure in the collection container substantially less than a pressure in the patient. 5. The method of claim 1, wherein forming a pressure differential includes removing a selected volume of a fluid from the collection container with at least one of a resilient bulb, a mechanical pump, a fluid pump, or combinations thereof. 6. The method of claim 1, wherein separating centrifuging the material to form the selected fraction includes: forming at least a first fraction and a second fraction; andapplying a force to the collection container to substantially urge the formation of the first fraction and the second fraction from the material. 7. The method of claim 1, wherein centrifuging the collection container includes: forming at least a first fraction and a second fraction of the material in the collection container;wherein the separating member moves relative to a boundary of the first fraction and the second fraction to substantially form the mechanical separation between the first fraction and the second fraction. 8. The method of claim 1, further comprising: forming a vacuum in the collection container relative to the patient to form the pressure differential in the collection container;urging a volume of a fluid to flow from the patient through the access port based on the pressure differential; andstopping the flow of the fluid being urged from the patient when the collection container is filled to the selected volume. 9. A method of withdrawing a material directly from a patient and collecting a selected fraction of the material, comprising: connecting a collection container directly to the patient;forming a pressure differential between the collection container and the patient;filling t... | Biomet Biologics, LLC | A61B005/153 B01D021/26 B01D017/038
US08688386 | B2 | 2014-04-01 | Analyte testing method and device for calculating basal insulin therapy | 2010-06-30 | Described herein are various methods to ensure safety and the compliance of therapeutic diabetes protocols. The method can be achieved by performing safeguards against hypoglycemia of the user prior to any change in basal insulin dosage based on the plurality of data. | 1. A method for management of diabetes for basal insulin dosing of a user with an insulin delivery device and a handheld diabetes data management unit having an analyte test sensor, a processor coupled to a memory and display screen, the method comprising: measuring a plurality of blood glucose concentration values of the user with the analyte test sensor over a plurality of time periods;collecting data representative of a plurality of fasting blood glucose concentration values with the handheld diabetes data management unit;ascertaining from the collected data whether the user has conducted a minimum number of fasting blood glucose concentration measurements within at least one of four prescribed time periods;determining whether the collected data indicate one of a first low blood glucose concentration pattern and a second low blood glucose concentration pattern lower than the first low blood glucose concentration pattern;evaluating whether basal insulin adjustment in the handheld diabetes data management unit is appropriate in the event that glucose concentration value results for two consecutive days are below a first low value for a first pattern;evaluating whether basal insulin adjustment in the handheld diabetes data management unit is appropriate in the event that glucose concentration value results for four consecutive days are each between about a low threshold and about a first low value for a second pattern;upon determination of at least one of the first and second low blood glucose concentration patterns of the user, displaying safety notifications on the display screen of the handheld diabetes data management unit; andprohibiting basal insulin adjustment in one of the insulin delivery device and handheld diabetes data management unit whenever either one of the first or second low blood glucose concentration pattern is determined. 2. The method of claim 1, in which consecutive days include a first day having a first glucose measurement and a second day having a second glucose measurement, where the two glucose measurements are apart by at least more than a first time interval. 3. The method of claim 2, in which the first time interval is greater than about two hours. 4. The method of claim 2, in which the determining comprises recommending insulin dosage that is decreased by a first decrement where a prior last glucose concentration value is less than a first low threshold. 5. The method of claim 2, in which the determining comprises recommending no changes to a current insulin dose stored in the memory of the handheld diabetes data management unit in the event that all three prior consecutive fasting blood glucose concentration values are not higher than a high threshold; and recommending, in the event that all three prior consecutive fasting blood glucose concentration values are greater than an upper threshold, an increase in insulin by a first increment, otherwise recommending an increase in insulin by a second increment. 6. ... | LifeScan, Inc. | G01N033/48 G01N031/00 G06G007/48 G06G007/58
US08731256 | B2 | 2014-05-20 | Automated image analysis for magnetic resonance imaging | 2009-02-02 | A magnetic resonance imaging (MRI) system, comprising: a magnetic resonance imaging scanner configured to generate a plurality of signals for forming at least one magnetic resonance image of a soft tissue region from a subject under observation, wherein the at least one magnetic resonance image provides at least one integrating feature to facilitate automatic segmentation; a signal processing system in communication with the magnetic resonance imaging scanner to receive the plurality of signals; and a data storage unit in communication with the signal processing system, wherein the data storage unit contains at least one template corresponding to the soft tissue region, wherein the signal processing system is adapted to process the plurality of signals received from the magnetic resonance imaging scanner to automatically perform segmentation for the soft tissue region of the subject under observation by utilizing the at least one template and the at least one integrating feature. | 1. A magnetic resonance imaging (MRI) system, comprising: a magnetic resonance imaging scanner configured to generate a plurality of signals for forming at least one magnetic resonance image of a soft tissue region from a subject under observation, wherein said at least one magnetic resonance image provides at least one integrating feature to facilitate automatic segmentation;a signal processing system in communication with said magnetic resonance imaging scanner to receive said plurality of signals;and a data storage unit in communication with said signal processing system, wherein said data storage unit contains at least one template corresponding to said soft tissue region,wherein said signal processing system is adapted to process said plurality of signals received from said magnetic resonance imaging scanner to automatically perform segmentation for said soft tissue region of said subject under observation by utilizing said at least one template and said at least one integrating feature,wherein performing said segmentation further comprises: transferring segmentation data from the at least one template to the at least one magnetic resonance image using a three-dimensional diffeomorphic transformation; and performing segmentation using the transferred segmentation data. 2. The magnetic resonance imaging system according to claim 1, wherein said at least one integrating feature provided by said at least one magnetic resonance image facilitates discerning said soft tissue region by at least one of: a correction of intensity homogeneity; a plurality of contrasts; a suppression of at least one tissue component; a quantification of at least one magnetic resonance (MR) contrast parameter, or variations thereof. 3. The magnetic resonance imaging system according to claim 1, wherein said signal processing system is at least partially incorporated within a structure housing said magnetic resonance scanner. 4. The magnetic resonance imaging system according to claim 1, wherein said signal processing system is at least partially incorporated in a workstation that is structurally separate and in communication with said magnetic resonance imaging scanner. 5. The magnetic. resonance imaging system according to claim 1, wherein said at least one template is obtained from magnetic resonance imaging. 6. The magnetic resonance imaging system according to claim 1, wherein said at least one template incorporates at least one magnetic resonance image from at least one subject that is different from said subject under observation. 7. The magnetic resonance imaging system according to claim 1, wherein said at least one template incorporates at least one magnetic resonance image from a previous scan of said subject under observation. 8. A workstation, comprising: a preprocessing engine to receive and preprocess magnetic resonance signals, from a magnetic resonance scanner, wherein said magnetic resonance signals form at least one magnetic resonance image show... | The Johns Hopkins University | G06K009/00 G06T007/00
US08521272 | B2 | 2013-08-27 | Method and device for monitoring breastfeeding | 2009-01-21 | A method of monitoring amount of milk consumed by an infant being breastfed by a breast is disclosed. The method comprises: determining variations in electric capacitance of the breast during breastfeeding, and correlating the electric capacitance variations to an amount of milk consumed by the infant. | 1. A method of monitoring amount of milk consumed by an infant being breastfed by a breast, comprising obtaining a multiplication between electric resistance of the breast and electric capacitance of the breast during breastfeeding, using a circuit for determining variations in said multiplication, and using a circuit for correlating said variations in said multiplication to an amount of milk consumed by the infant. 2. The method of claim 1, further comprising measuring said electric capacitance so as to determine said variations. 3. The method of claim 2, wherein said measurement of said electric capacitance is performed so as to estimate electric capacitance of an interior of the breast while reducing contribution of a skin of the breast to said electric capacitance. 4. The method of claim 1, wherein said determination of said variations comprises measuring a phase of voltage sampled from a skin of the breast in response to an electrical current applied to said skin. 5. The method of claim 4, wherein said phase is measured via at least four electrodes connected to a skin of the breast. 6. The method of claim 1, further comprising correcting said correlation using history data collected during previous breastfeeding sessions. 7. The method of claim 1, further comprising performing a calibration measurement prior to said breastfeeding so as to collect calibration data, wherein said multiplication is corrected based on said calibration data. 8. The method of claim 1, wherein said multiplication is determined for each of a plurality of frequencies, and wherein said amount of milk is correlated to a combination of at least two multiplications. 9. The method of claim 1, further comprising using said multiplication for searching over the breast for regions occupied clusters of alveoli. 10. The method of claim 2, wherein said multiplication is measured at a frequency of less than 100 MHz. 11. The method of claim 1, wherein said multiplication is determined using a plurality of electrodes, and wherein the method comprises employing at least one multiplexing cycle such that at different sub-cycles of said multiplexing cycle, a different set of electrodes is used for determining said multiplication. 12. The method of claim 11, further comprising using multiplication values measured during said at least one multiplexing cycle for determining measuring locations on the breast in future sessions. 13. The method of claim 11, further comprising analyzing multiplication values measured during said at least one multiplexing cycle so as to differentiate between measurement sensitivities in different depths in the breast. 14. A breastfeeding monitor system, comprising a measuring unit having circuit configured for measuring a multiplication between electric resistance of the breast and electric capacitance of a breast during breastfeeding, and a processing unit having circuit configured for determining variations in said multiplication and correlating... | Yeda Research and Development Co. Ltd. | A61B005/053
US08554297 | B2 | 2013-10-08 | Body-worn pulse oximeter | 2009-09-14 | The invention provides a body-worn system that continuously measures pulse oximetry and blood pressure, along with motion, posture, and activity level, from an ambulatory patient. The system features an oximetry probe that comfortably clips to the base of the patient's thumb, thereby freeing up their fingers for conventional activities in a hospital, such as reading and eating. The probe secures to the thumb and measures time-dependent signals corresponding to LEDs operating near 660 and 905 nm. Analog versions of these signals pass through a low-profile cable to a wrist-worn transceiver that encloses a processing unit. Also within the wrist-worn transceiver is an accelerometer, a wireless system that sends information through a network to a remote receiver, e.g. a computer located in a central nursing station. | 1. A method for monitoring a patient, comprising: (a) positioning a pulse oximetry sensor on a base of the patient's thumb, the pulse oximetry sensor comprising a first radiation source configured to emit a first optical radiation at a first wavelength; a second radiation source configured to emit a second optical radiation at a second wavelength; a photodetector configured to detect the first optical radiation emitted by the first radiation source after the first optical radiation passes through a portion of the patient to generate a first signal, and the second optical radiation emitted by the second radiation source after the second optical radiation passes through a portion of the patient to generate a second signal; and a finger-ring housing comprising the first and second radiation sources and the photodetector, wherein the pulse oximetry sensor is positioned such that the first optical radiation emitted by the first radiation source and the second optical radiation emitted by the second radiation source pass through the thumb tissue and are received by the princeps pollicis artery before reaching the photodetector;(b) causing the first radiation source to emit the first optical radiation and measuring a first time-dependent signal by detecting the first optical radiation with the photodetector;(c) causing the second radiation source to emit the second optical radiation and measuring a second time-dependent signal by detecting the second optical radiation with the photodetector;(d) measuring a time-dependent motion signal with at least one motion sensor;(e) transmitting the first time-dependent signal, the second time-dependent signal, and the time-dependent motion signal at a processing component;(f) processing the time-dependent motion signal with the processing component to determine the patient's posture and activity state;(g) processing both the first and second time-dependent signals with the processing component to determine an oxygen saturation value for the patient; and,(h) suppressing generation of an alarm based on the oxygen saturation value when the patient's posture and activity state meet a predetermined criteria. 2. The method of claim 1, wherein step (h) further comprises processing the patient's posture and activity state in combination with a parameter resulting from a comparison of the oxygen saturation value to a predetermined alarm criteria to suppress the alarm. 3. The method of claim 2, wherein step (h) further comprises suppressing the alarm if the patient's posture and activity state indicate that the patient is standing upright. 4. The method of claim 2, wherein step (h) further comprises suppressing the alarm if the patient's posture and activity state indicate a change from lying down to one of sitting and standing upright. 5. The method of claim 2, wherein step (h) further comprises suppressing the alarm if the patient's posture and activity state indicate a change from one of standing upright and sitting t... | Sotera Wireless, Inc. | A61B005/1455
US08206319 | B2 | 2012-06-26 | Tissue penetration device | 2010-08-26 | A moving coil tissue penetration device includes a magnetic source that produces a magnetic field in a magnetically active region. A cylindrical coil is secured to a translation substrate and disposed at least partially within the magnetically active region. A sharpened member is configured to penetrate tissue and mechanically coupled to the translation substrate. | 1. A moving coil tissue penetration device, comprising: a) a magnetic source that produces a magnetic field in a magnetically active region;b) a sharpened member configured to penetrate tissue and a lancet configured to penetrate tissue and mechanically coupled to a translation substrate;c) an iron core mounted to the lancet;d) a sharpened member guide that serves as a stop for the iron core; ande) an electromagnetic driver includes a driver coil pack that is divided into a plurality of coils along a path of the lancet with direct current configured to be alternated to the coils to advance and retract the lancet. 2. The tissue penetration device of claim 1, further comprising a feedback loop having a processor which is electrically coupled to an elongate electrical conductor of a cylindrical wire coil and configured to control the magnitude of the force on the translation substrate based on feedback from a position sensor. 3. The tissue penetration device of claim 2, wherein the processor is configured to control the magnetic force on the translation substrate so as to have the translation substrate and the lancet follow a predetermined tissue penetration profile. | Sanofi-Aventis Deutschland GmbH | A61B005/00
US08639309 | B2 | 2014-01-28 | Method and system for monitoring oxygenation levels of compartments and tissue | 2010-05-04 | A method and system for continually monitoring oxygenation levels in real-time in compartments of an animal limb, such as in a human leg or a human thigh or a forearm, can be used to assist in the diagnosis of a compartment syndrome. The method and system can include one or more near infrared compartment sensors in which each sensor can be provided with a compartment alignment mechanism and a central scan depth marker so that each sensor may be precisely positioned over a compartment of a living organism. The method and system may comprise hardware or software (or both) may adjust one or more algorithms based on whether tissue being monitored was traumatized or is healthy. The method and system can also monitor the relationship between blood pressure and oxygenation levels and activate alarms based on predetermined conditions relating to the oxygenation levels or blood pressure or both. | 1. A method for automatically monitoring oxygenation levels of a compartment of a human body for automatically detecting conditions of a compartment syndrome, comprising: displaying one or more visuals on a display device comprising anatomical locations in which to position a compartment sensor to monitor a compartment of injured tissue;receiving confirmation with a computing device that the compartment sensor is ready;automatically monitoring oxygenation levels of the compartment in a continuous manner with the computing device coupled to the compartment sensor;automatically monitoring oxygenation levels of healthy tissue with the computing device coupled to a sensor in a continuous manner, the sensor detecting systemic perfusion of the human body from the healthy tissue;automatically monitoring blood pressure of the human body in a continuous manner with the computing device, the blood pressure comprising diastolic and systolic blood pressure values;displaying oxygenation levels of the compartment on the display device;displaying diastolic and systolic blood pressure values simultaneously with the oxygenation levels of the compartment on the display device; andactivating an alarm with the computing device when both the blood pressure of the human body comprising the diastolic and systolic blood pressure values decreases and oxygenation levels of the compartment sensor start decreasing in value compared to the oxygenation levels of the sensor for the healthy tissue. 2. The method of claim 1, further comprising detecting lower serial hemoglobins (Hgb) or hematocrits (Hct) in blood of a patient with the computing device based on monitored oxygenation levels. 3. The method of claim 1, further comprising sensing at least one of respiration, pH levels, temperature, pulse, heart rate, and altitude with another sensor coupled to the computing device. 4. The method of claim 1, wherein monitoring oxygenation levels of the compartment with the compartment sensor further comprises measuring muscle oxygenation of the compartment with the computing device. 5. The method of claim 1, further comprising increasing a frequency at which the oxygenation levels and blood pressure are monitored if both the oxygenation levels and blood pressure fall within a predetermined range of values. 6. The method of claim 5, wherein the anatomical locations correspond to locations on one of a leg and arm. 7. The method of claim 1, further comprising calculating and displaying a mean arterial pressure (MAP) on the display device. 8. A system for automatically monitoring oxygenation levels of a compartment of a human body for automatically detecting conditions of a compartment syndrome, the system comprising: a compartment sensor configured for automatically detecting oxygenation levels of injured tissue of the human body in a continuous manner;a sensor configured for automatically detecting oxygenation levels of healthy tissue of the human body in a continuous manner;a bl... | J&M Shuler, Inc. | A61B005/0205
US09854995 | B2 | 2018-01-02 | Non-invasive, non contact system, electronic control unit, and associated methodology for minimizing muscle stress and improving circulation | 2009-06-05 | A method, system, and process are described to decrease muscle fatigue and improve circulation of an individual resting on a support device, such as an adjustable chair or bed. The body pose and heat distribution of the individual are determined and used to estimate muscle fatigue and circulation. The estimated muscle fatigue and circulation are then used to determine settings for the support device or the temperature of the environment. | 1. A method of minimizing muscle stress or increasing blood circulation of a person resting on a support device in an environment, the method comprising: sensing, by a body pose sensor, a body pose of the person using a first image sensor located above head level of the person and a second image sensor located below the head level of the person in order to capture a three-dimensional image of at least a torso of the person, the three-dimensional image being used to determine the body pose of the person;sensing, by a body temperature sensor, a body heat distribution of the person over a plurality of different body parts of the person;determining, by circuitry, at least one of a muscle stress estimate and a blood circulation estimate of the person based only on the sensed body pose and the sensed body heat distribution; andcontrolling a positional adjustment of the support device according to the at least one of the muscle stress estimate and the blood circulation estimate. 2. The method according to claim 1, wherein the support device is an adjustable seat. 3. The method according to claim 1, wherein the support device is a hospital bed. 4. An electronic control unit (ECU) implementing the method according to claim 1, wherein the support device is a vehicle seat. 5. The method according to claim 1, further comprising: determining a setting of a temperature of the environment based on the at least one of the muscle stress estimate and the blood circulation estimate. 6. A system configured to minimize muscle stress or improve blood circulation of a person on a support device in an environment, the system comprising: a body pose sensor configured to sense a body pose of the person, the body pose sensor including a first image sensor located above head level of the person and a second image sensor located below the head level of the person in order to capture a three-dimensional image of at least a torso of the person, the three-dimensional image being used to determine the body pose of the person;a body temperature sensor configured to sense a body heat distribution of the person over a plurality of different body parts of the person;circuitry configured to determine at least one of a muscle stress estimate and a blood circulation estimate of the person based only on the sensed body pose and the sensed body heat distribution; andcontrol a positional adjustment of the support device according to the at least one of the muscle stress estimate and the blood circulation estimate. 7. The system according to claim 6, wherein the body temperature sensor is a heat-detecting imaging device. 8. The system according to claim 6, wherein the support device is a vehicle seat. 9. The system according to claim 6, wherein the support device is a hospital bed. 10. The system according to claim 6, wherein the circuitry includes a controller that is part of an electronic control unit (ECU). 11. The system according to claim 6, wherein the circuitry is further ... | Toyota Motor Engineering & Manufacturing North America, Inc. | A61B005/103 A61B005/00 B60N002/00 B60N002/02 A61B005/01
US20100331646 | A1 | 2010-12-30 | Health Management Devices and Methods | 2009-06-30 | Methods and devices to detect analyte in body fluid are provided. Embodiments include analyte sensors designed so that at least a portion of the sensor is positionable beneath a skin surface during analyte monitoring. | 1. An in vivo glucose monitoring system, comprising: an in vivo glucose sensor comprising at least a portion that is positionable beneath a skin surface to contact biological fluid for a period of time of glucose monitoring, a working electrode and at least one additional electrode, at least one electrical contact coupled to the electrodes, and a glucose-responsive enzyme; andan on-body unit comprising a sensor contacting portion to contact the in vivo sensor when at least a portion of the sensor is positioned beneath skin, and comprising a selective depth controller to selectively adjust the depth of the sensor beneath the skin surface without adjusting the distance between the sensor and the on-body unit so that the distance between the sensor and on-body unit is fixed and the sensor extends the same distance from the on-body unit regardless of the selected depth. 2. The system of claim 1, wherein the on-body unit comprises a mounting unit, or a sensor control unit or a sensor control unit mountable in a mounting unit. 3. The system of claim 2, wherein the depth controller comprises an offset. 4. The system of claim 3, wherein the offset comprises one or more spacers. 5. The system of claim 4, wherein the system comprises at least two spacers, and the at least two spacers are the same size. 6. The system of claim 4, wherein the system comprises at least two spacers, and the at least two spacers are different sizes. 7. The system of claim 4, wherein a spacer has a variable dimension. 8. The system of claim 7, wherein the dimension is a thickness. 9. The system of claim 2, wherein the depth controller is integrated with the on-body unit. 10. The system of claim 9, wherein the depth controller is selectively moveable from a first position substantially within or on the on-body unit to a second position in which at least a portion of the depth controller extends from the on-body unit a selective distance. 11. The system of claim 10, wherein the depth controller is retractable to a position within or on the on-body unit. 12. The system of claim 2, wherein the depth controller is separate from the on-body unit. 13. The system of claim 2, wherein the depth controller selectively adjusts the angle that the sensor is positioned relative to skin. 14. The system of claim 13, wherein the angle ranges from about 1 to about 90 degrees. 15. The system of claim 2, wherein the depth controller is an adjustable bladder. 16. The system of claim 2, wherein the on-body unit comprises a user-actuatable depth selector coupled to the depth controller. 17. The system of claim 2, wherein the depth selector comprises an indexer that increments the depth ranges through a plurality of discrete positions. 18. The system of claim 2, wherein the on-body unit comprises an inflatable cuff 19. The system of claim 1, wherein the sensor is a planar sensor. 20. The system of claim 1, wherein the sensor is a wire-type sensor. 21. A method, comprising: selecting a... | Abbott Diabetes Care Inc. | A61B005/1468
US07848795 | B2 | 2010-12-07 | Determination of sleep depth | 2006-05-22 | The invention relates to a mechanism for determining the depth of sleep of a subject. In order to obtain information about the continuum of the depth of sleep in a user-friendly way without a need for high computational power, EEG signal data is obtained from a subject, REM sleep periods of the subject are detected, and a measure indicative of irregularity in the EEG signal data is derived. Based on the measure and the depicted REM sleep periods, a sleep depth index and state information are produced, where the index is indicative of the depth of sleep of the subject and the state information indicates whether or not the EEG signal data is obtained during a REM sleep period. The detection of the REM periods may be based on a bioimpedance measured from EEG electrodes simultaneously with the EEG measurement or on EOG signal data measured from the subject. | 1. A method for determining the depth of sleep of a subject, the method comprising the steps of: acquiring EEG signal data from a subject;detecting REM sleep periods of the subject from the EEG signal data using a control unit;deriving a sequence of a measure indicative of irregularity in the EEG signal data using the control unit;producing, based on the measure and the detecting step, a combined sleep depth curve depicting both a sequence of a sleep depth index indicative of the depth of sleep of the subject and state information indicative of whether the EEG signal data is obtained during a REM sleep period;discarding selected segments of the EEG signal data, the selected segments corresponding to the REM sleep periods detected in the detecting step; anddisplaying the combined sleep depth curve on a monitor, wherein the portions of the combined sleep depth curve obtained during the REM sleep periods are visually distinguished from other sleep periods. 2. A method according to claim 1, wherein the detecting step includes the sub-steps of: supplying an AC excitation signal through a signal path formed between two electrodes of an electrode set attached to the subject's skin surface in a measurement area of the subject's body;measuring an impedance signal through the electrode set, the impedance signal being indicative of the impedance of the signal path;determining at least one time period during which the impedance signal fulfills at least one predetermined criterion; anddefining the REM sleep periods based on the at least one time period. 3. A method according to claim 2, wherein the determining sub-step includes a sub-step of comparing the impedance signal with a predetermined threshold value. 4. A method according to claim 1, further comprising a step of replacing the selected segments with new signal values. 5. A method according to claim 1, wherein the deriving step includes deriving the measure indicative of irregularity in the EEG signal data, in which the measure is indicative of spectral entropy of the EEG signal data. 6. A method according to claim 1, wherein the detecting step includes the sub-steps of: acquiring EOG signal data from the subject; andutilizing the EOG signal data for defining the REM sleep periods. 7. A method according to claim 6, wherein the utilizing sub-step includes comparing the EOG signal data with a predetermined threshold. 8. A method according to claim 6, wherein the utilizing sub-step includes comparing the EOG signal data with the EEG signal data. 9. An apparatus for determining the depth of sleep of a subject, the apparatus comprising: means for acquiring EEG signal data from a subject;detection means for detecting REM sleep periods of the subject;first calculation means for deriving a sequence of a measure indicative of irregularity in the EEG signal data; andsecond calculation means, responsive to the first calculation means and the detection means, for producing a combined sleep depth curve depi... | General Electric Company | A61B005/04
US07840252 | B2 | 2010-11-23 | Method and system for determining a three dimensional representation of a tubular organ | 2005-09-23 | A method for determining a three dimensional representation of a tubular organ, the method comprising the procedures of registering a three dimensional medical positioning system (MPS) coordinate system with a two dimensional coordinate system, acquiring a plurality of MPS points within the tubular organ, the MPS points are associated with the MPS coordinated system, each of the MPS point is acquired with a respective organ timing signal reading, acquiring at least one two dimensional image of the tubular organ, the at least one two dimensional image is associated with the two dimensional coordinate system, the at least one two dimensional image is acquired from at least one perspective, each of the at least one two dimensional image is further acquired with a respective organ timing signal reading, associated each of the MPS points with the at least one two dimensional image according to the organ timing signals thereof, determining a temporal three dimensional representation of a portion of the tubular organ for each MPS point, the portion being identifiable in the respective two dimensional image. | 1. A method for determining a three dimensional representation of a tubular organ, said method comprising the procedures of: registering a three dimensional medical positioning system (MPS) coordinate system with a two dimensional coordinate system;acquiring a plurality of MPS points within said tubular organ, said MPS points are associated with said MPS coordinated system, each of said MPS points is acquired with a respective organ timing signal reading;acquiring at least one two dimensional image of said tubular organ, said at least one two dimensional image is associated with said two dimensional coordinate system, said at least one two dimensional image is acquired from at least one perspective angle, each of said at least one two dimensional image is further acquired with a respective organ timing signal reading;associating each of said MPS points with said at least one two dimensional image according to the organ timing signals thereof;determining a temporal three dimensional representation of a portion of said tubular organ for each MPS point, said procedure of determining a three dimensional representation comprising the sub-procedures of:selecting at least one two dimensional image, said at least one two dimensional image including said tubular organ;projecting the trajectory of said catheter on the selected two dimensional image, said trajectory being determined by said MPS points associated with said selected two dimensional image, thereby defining a centerline of said tubular organ;detecting the border points of tubular organ on said selected two dimensional image; andgenerating a three dimensional closed curve around each of said MPS points, associated with said selected image. 2. The method according to claim 1 further includes injecting a dye into said tubular organ prior to said acquiring at least one two dimensional image. 3. The method according to claim 1 further includes the procedure of superimposing said three dimensional representation on a two dimensional image. 4. The method according to claim 3, wherein said two dimensional image is a real time two dimensional image, said real time two dimensional image being associated with an organ timing signal readings. 5. The method according to claim 3, wherein said two dimensional image is one of said at least one two dimensional image. 6. The method according to claim 3, wherein the perspective of said two dimensional image is selected such that a foreshortening phenomenon is minimized, said foreshortening phenomenon caused due to the super position of said three dimensional representation on the plane of said two dimensional image. 7. The method according to claim 3, wherein a two dimensional projection of said superimposed three dimensional representation serves as a virtual dye instead of injecting a dye prior to acquiring said real time two dimensional image. 8. The method according to claim 1, wherein said plurality of MPS points are acquired by inserting a catheter, ... | MediGuide, Ltd. | A61B005/05
US07846106 | B2 | 2010-12-07 | Atrial fibrillation detection using SPO2 | 2006-04-26 | The method and system includes detecting atrial fibrillation in a patient by monitoring the blood oxygen saturation level over a period of time. The method and system produces a plethysmographic waveform from the monitored blood oxygen saturation level and analyzes the plethysmographic waveform and detected intervals and determines whether the patient is in atrial fibrillation. The method and system is preferably implemented in a software application and may be configured to report to the user on the current state of atrial fibrillation (AFIB) and a current trend. | 1. A method of detecting atrial fibrillation in a human or animal subject, the method comprising: acquiring a plethysmographic waveform of a patient from a monitoring device over a period of time;detecting a set of pulse intervals from the plethysmographic waveform;measuring a set of pulse features from the plethysmographic waveform; andanalyzing the set of pulse intervals and the set of pulse features from the plethysmographic waveform to determine an atrial fibrillation state of the patient. 2. The method as claimed in claim 1, wherein the monitoring device is a pulse oximeter sensor. 3. The method as claimed in claim 1, wherein one of the set of features is a set of pulse amplitudes. 4. The method as claimed in claim 1, further comprising analyzing the set of pulse intervals and the set of pulse features from the plethysmographic waveform to detect an atrial fibrillation trend of the patient. 5. The method as claimed in claim 4, further comprising reporting the atrial fibrillation trend. 6. The method as claimed in claim 1, further comprising reporting the atrial fibrillation state. 7. The method as claimed in claim 1, wherein analyzing the set of pulse intervals includes: conducting a hidden Markov model analysis on the set of pulse intervals; and conducting a contextual analysis on the set of pulse intervals. 8. The method as claimed in claim 7, wherein the hidden Markov model analysis determines a probability of the patient having an irregular heart rhythm. 9. The method as claimed in claim 7, wherein the contextual analysis rules out atrial fibrillation as a source of irregular heart rhythm. 10. The method as claimed in claim 1, wherein analyzing the set of pulse features includes: calculating amplitude variation between each consecutive one of the set of pulse amplitudes. 11. The method as claimed in claim 10, wherein the atrial fibrillation state determination is dependent on an analysis of pulse amplitude variation. 12. The method as claimed in claim 1, wherein analyzing the set of pulse amplitudes includes determining a set of filling time dynamics of the set of pulse amplitudes. 13. The method as claimed in claim 12, wherein the steps of analyzing the set of pulse amplitudes and determining the set of filling time dynamics provides an indication of the degree of atrial augmentation present. 14. A system for detecting atrial fibrillation in a human or animal subject, the system comprising: a remote sensing system configured to analyze a plethysmographic waveform from a monitoring device;a storage media for storing a computer application;a processing unit coupled to the remote sensing system and the storage media, and configured to execute the computer application, and further configured to receive the plethysmographic waveform from the remote sensing system,wherein when the computer application is executed, a set of pulse intervals from the plethysmographic waveform are detected and a set of pulse features are measured fro... | The General Electric Company | A61B005/02
US20100286571 | A1 | 2010-11-11 | System and Method for Providing Body Sway Feedback to a Body of a Subject | 2009-05-14 | The present invention relates to a system (10) and to a method for providing body sway feedback to a body of a subject. The system (10) comprises (a) at least one body sway sensor means (12) adapted to be attached the body and adapted for providing body sway signals indicative of a sway of the body, (b) a system processor means (14) adapted for deriving body sway information from the body sway signals and adapted for generating a body sway feedback signal from the body sway information, and (c) a subject body sway feedback means (40) having at least one transducer unit (42, 44) and adapted for providing a body sway feedback to the body via said at least one transducer unit (42, 44) based on said body sway feedback signal, (d) wherein said at least one transducer unit (42, 44) is or comprises at least one transducer, (e) wherein said at least one transducer is adapted to be contacted to a head of the body, adapted for vibrating, and adapted for transferring a respective vibration to the head, and (f) wherein said at least one transducer is adapted to evoke based on the vibration a multimodal body sway feedback as stimuli to the subject having a plurality of kinds of stimuli of the group comprising auditory stimuli, tactile stimuli, and vestibular stimuli. | 1-36. (canceled) 37. A system for providing multi-modal body sway feedback to a body of a subject, comprising: (a) at least one body sway sensor capable of being attached at a lower back of the body and providing body sway signals indicative of a sway of the body, the at least one body sway sensor comprising at least one of: an angular velocity sensor capable of providing body sway signals indicative of an angular velocity of the body;an acceleration sensor system capable of providing body sway signals indicative of an angular acceleration of the body; anda position sensor capable of providing body sway signals indicative of an angular displacement of the body;(b) a system processor capable of deriving body sway information from the body sway signals and capable of generating a body sway feedback signal from the body sway information;(c) a subject body sway feedback device having at least first and second transducer units and capable of providing a body sway feedback to the body through the at least first and second transducer units based upon the body sway feedback signal;(d) the first transducer unit: capable of being removably secured to a head of the body; andcapable of applying a relatively low-level vibration and transferring the low-level vibration to skin of the head; and(e) the second transducer: capable of applying a relatively high-level vibration; andcapable to evoke, with the first transducer, a multimodal body sway feedback as stimuli to the subject, the stimuli selected from at least one of the group consisting of: auditory stimuli;tactile stimuli; andvestibular stimuli,based upon the high-level vibration of a bone of a skull of the body, the high-level vibration being insufficient to cause eye movements through the vestibular ocular-reflex. 38. The system according to claim 37, wherein the at least one body sway sensor is capable of providing body sway signals indicative of the sway of the body in a pitch direction and is capable of providing body sway signals indicative of the sway of the body in a roll direction. 39. The system according to claim 37, wherein the system processor comprises: a derivation device capable of deriving body sway angle and body sway angular velocity information from the body sway signals; andfeedback generating device capable of generating the body sway feedback from the body sway angle and body sway angular velocity information. 40. The system according to claim 37, wherein the system processor comprises: means for deriving body sway angle and body sway angular velocity information from the body sway signals; andmeans for generating the body sway feedback from the body sway angle and body sway angular velocity information. 41. The system according to claim 37, wherein the subject body sway feedback device comprises a feedback device capable of providing the body sway feedback to the subject aurally through bone conduction to hearing receptors of an inner ear through at least one bone vibrating tra... | - | A61B005/11 A61H001/00
US20100298689 | A1 | 2010-11-25 | SYSTEMS AND METHODS FOR DETECTING FLOW AND ENHANCING SNR PERFORMANCE IN PHOTOACOUSTIC IMAGING APPLICATIONS | 2008-11-13 | The present disclosure provides systems and methods for combining photoacoustic/thermoacoustic imaging with power Doppler signal processing. More particularly, the disclosed systems and methods involve use of encoded Doppler signals in order to detect and image in vivo blood flow. The disclosed flow detection systems and methods may be used in photoacoustic imaging using PD to achieve, inter alia, enhanced signal-to-noise (SNR) and sensitivity performances. A method for detecting flow in a target region may involve (i) obtaining a encoded signal containing photoacoustic imaging data for the target region using a photoacoustic imaging system, (ii) decoding the encoded signal, (iii) passing the decoded signal through a demodulator and a low-pass filter, resulting in a base-band signal, (iv) passing the base-band signal through a wall filter, resulting in an uncluttered signal; and (iv) estimating the Ro value by integrating the power spectrum of the uncluttered signal. | 1. A method for detecting flow in a target region, the method including the steps of: a. obtaining an encoded signal containing photoacoustic imaging data for the target region using a photo acoustic imaging system;b. decoding the encoded signal;c. passing the decoded signal through a demodulator and a low-pass filter, resulting in a base-band signal;d. passing the base-band signal through a wall filter, resulting in an uncluttered base-band signal; ande. estimating an R0 value for an entire spectrum of the uncluttered base-band signal by integrating a power spectrum of the uncluttered base-band signal wherein the R0 value provides a power estimate that can be translated to or (ii) correlated with bulk flow properties/parameters in the target region not based on Doppler frequency shifts. 2. The method according to claim 1, wherein the decoded signal is a base-band signal. 3. The method according to claim 1, wherein the decoded signal is an analytic signal. 4. The method according to claim 1, wherein the decoded signal is a raw radio-frequency signal. 5. The method according to claim 1, wherein the wall filter removes tissue clutter from the decoded signal. 6. The method according to claim 1, wherein the R0 value is an estimate of bulk blood flow. 7. The method according to claim 1, wherein the R0 value is used to detect hyperemia conditions. 8. The method according to claim 1, wherein the R0 value is estimated using the following calculation: R0 =∫z(t)z*(t)dt=∫(x2(t)+y2(t))dt=∫P(ω)dω, wherein z is the uncluttered base-band signal and z(t)=x(t)+iy(t) and P is the power spectrum of the uncluttered base-band signal. 9. A photoacoustic imaging system comprising: a. one or more electromagnetic beam sources adapted to irradiate a target location, wherein the power spectrum of the one or more beams is encoded;b. one or more ultrasound detectors adapted to detect one of a photoacoustic signal or a signal that includes a an encoded base-band signal resulting from a target sample;c. means for synchronization of irradiation and detection functionalities, andd. means for processing the encoded base-band signal to derive flow information related to the target location, wherein processing the encoded base-band signal includes estimating an R0 value for an entire spectrum of a corresponding decoded base-band signal wherein the R0 value provides a power estimate that can be (i) translated to or (ii) correlated with bulk flow properties/parameters in the target region not based on Doppler frequency shifts. 10. The system according to claim 9, wherein the means for processing the encoded base-band signal includes a demodulator, a low-pass filter, a wall filter and means for R0 value estimation. 11. The system according to claim 10, wherein the demodulator and low-pass filter are adapted to decode and extract a base-band signal of a corresponding photoacoustic/thermoacoustic signal. 12. The system according to claim 10, wherein the wall filter is adapted... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/05 G06F019/00 A61B005/026
US07860544 | B2 | 2010-12-28 | Analyte monitoring device and methods of use | 2007-03-07 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. An analyte monitoring system, comprising: a first sensor, at least a portion of which is positionable in the body of a person, and including at least one electrode with an analyte-responsive enzyme associated therewith;a second sensor, at least a portion of which is positionable in the body of the person, and including at least one electrode with an analyte-responsive enzyme associated therewith;a data processing unit to receive data from the first and second sensors and programmed to verify the data validity; anda medication delivery unit in communication with the data processing unit; wherein when a predetermined physiological condition is detected, the data processing unit is configured to analyze data from the first or second sensor over a predetermined time period, and to determine one or more patterns associated with the detected predetermined physiological condition based on the stored data, and to predict an occurrence of a future episode associated with the physiological condition based on the determined one or more patterns; and furtherwherein the data processing unit is trainable to modify medication delivery based on the predicted occurrence of the future episode associated with the physiological condition. 2. The system of 1 wherein the analyte-response enzyme associated with the at least one electrode of the first and second sensors are the same. 3. The system of 2 wherein the analyte response enzyme is responsive to glucose. 4. The system of claim 1 wherein the analyte response enzyme associated with the at least one electrode of the first and second sensors are different. 5. The system of claim 4 wherein the analyte response enzyme is responsive to glucose. 6. The system of claim 4 wherein the analyte response enzyme is responsive to lactate. 7. The system of claim 1 wherein the data processing unit is configured to confirm validity of at least one data from the first or second sensors when the data from the first and the second sensor are within a predetermined parameter. 8. The system of claim 1 wherein the data from at least one of the first or the second sensors is calibrated using a blood sample. 9. The system of claim 1 wherein the data from at least one of the first or the second sensors is calibrated at the factory. 10. The system of claim 1 further including an output unit operatively coupled to the medication delivery unit to output one or more of the data from the first or the second sensor. 11. The system of claim 10 wherein the output unit presents the data visually, audibly, vibratorily, or any combination thereof. 12. The system of claim 1 wherein at least a portion of the at least one electrode of the first sensor or of the at least one electrode of the second sensor is in fluid contact with an analyte of the person. 13. The system of claim 12 wherein the at least the portion of the at least one electrode of the first sensor and the at least one electrode of the second sensor are in fluid contact sim... | Abbott Diabetes Care Inc. | A61B005/05
US07848797 | B2 | 2010-12-07 | Motor unit number estimation (MUNE) for the assessment of neuromuscular function | 2007-08-17 | A method for the assessment of neuromuscular function by motor unit number estimation, comprising: (i) determining and controlling stimulation and data acquisition process via pre-configured electrode array so as to acquire stable and less uncertainty MU responses; (ii) pre-processing acquired MUs responses so as to attenuate noise, determine MUs activity region, and improve processing speed and accuracy; (iii) minimizing alternation effects by globally searching and comparing SMUPs; (iv) eliminating alternation effects by identifying alternating MUs directly; and (v) computing and reporting MUNE results, as well as the statistical description of these MUN estimates to evaluate its robustness. | 1. A method for assessing neuromuscular function by estimating motor unit numbers, the method comprising: providing an electrode array for delivering one or more controlled stimuli to a subject using at least one stimulator electrode and for acquiring electric activity signals of one or more motor units of a subject using at least one detector electrode;positioning said electrode array on a subject, delivering a stimulus of varying characteristics to the subject and acquiring a set of electric activity signals of one or more motor units of the subject;identifying and eliminating any acquired signals due to alternation;automatically estimating motor unit number using an automation algorithm,wherein the automation algorithm consists of: (i) pre-processing acquired responses so as to consolidate and enhance the acquired signals; and(ii) identifying plausible new motor unit activation events and corresponding response changes; and(iii) determining representatives of single motor unit potential features, and computing approaches; and(iv) estimating motor unit number and distribution statistics associated with the motor unit number;wherein pre-processing the acquired signals comprises: (a) estimating background noise level; and(b) determining motor unit activity region in the acquired signal; and(c) ordering waveforms according to activity level, and removing noise-only waveforms; and(d) identifying and combining waveforms with insignificant morphological variations in relation to background noise;wherein a background noise level is determined by: (a) using waveform segments known to have no motor unit activity; and(b) estimating statistics of noise level, the estimated statistics comprising at least one of mean absolute deviation, standard deviation, and mean square error. 2. A method according to claim 1 wherein the electrode array consists of individual electrodes placed independently on a subject according to anatomical landmarks. 3. A method according to claim 1 wherein the electrode array consists of electrodes whose positions, relative to each other, are fixed for at least some of the electrodes. 4. A method according to claim 1 wherein the electrode array consists of one or more pairs of stimulator electrodes, wherein the relative positions of the stimulator electrodes are selected from the group consisting of: (i) fixed, and (ii) adjustable. 5. A method according to claim 1 wherein the varying characteristics of the stimuli are controlled and adjusted according to one selected from the group consisting of: (i) directly and manually by a user, and (ii) automatically by a computer algorithm. 6. A method according to claim 5 wherein the characteristics of the stimuli include at least one selected from the group consisting of: duration, intensity, polarity, frequency, and other features of electric current to be delivered via stimulus electrodes. 7. A method according to claim 5 wherein the control and adjustment of the stimuli is based on r... | NeuroMetrix, Inc. | A61B005/04 A61B005/05
US20100280366 | A1 | 2010-11-04 | CONTINUOUS FIELD TOMOGRAPHY SYSTEMS AND METHODS OF USING THE SAME | 2009-05-12 | Continuous field tomography systems are provided. Aspects of systems include a data aggregating module configured to receive both continuous field tomography data and non-continuous field physiological data and produce an aggregated data product from these disparate types of data. Also provided are methods of using systems of the invention in a variety of different applications, including diagnostic and therapeutic applications. The systems and methods of the invention find use in a variety of different applications, such as cardiac related applications. | 1. A system comprising: a data aggregating module configured to: receive continuous field tomography data;receive non-continuous field physiological data; andproduce an aggregated data product from the received continuous field tomography data and non-continuous field physiological data. 2. The system according to claim 1, wherein the continuous field tomography data is electric tomography data. 3. The system according to claim 1, wherein the aggregated data product comprises information that is configured to be employed by a user. 4. The system according to claim 3, wherein the aggregated data product is configured to be displayed to a user. 5. The system according to claim 4, wherein the aggregated data product is configured to be displayed to a user on an image display device. 6. The system according to claim 1, wherein the aggregated data product comprises information that is to be employed by a device. 7. The system according to claim 1, wherein the device is configured to modify an operating parameter in response to receiving the information. 8. The system according to claim 1, wherein the system further comprises a continuous field tomography data source 9. The system according to claim 1, wherein the system further comprises a non-continuous field physiological data source. 10. The system according to claim 9, wherein the non-continuous field physiological data source comprises an implantable device. 11. The system according to claim 10, wherein the implantable device is a cardiac device. 12. The system according to claim 9, wherein the non-continuous field physiological data source comprises an extra-corporeal body-associated device. 13. The system according to claim 12, wherein the extra-corporeal body-associated device comprises a conductively transmitted signal receiver. 14. The system according to claim 9, wherein the non-continuous field physiological data source comprises an ingestible event marker. 15. The system according to claim 9, wherein the non-continuous field physiological data source comprises a patient-related behavior parameter recordation device. 16. A method comprising: receiving continuous field tomography data and non-continuous field physiological data at a system comprising a data aggregating module configured to: receive continuous field tomography data;receive non-continuous field physiological data; andproduce an aggregated data product from the received continuous field tomography data and non-continuous field physiological data. 17. The method according to claim 16, wherein the method further comprises obtaining the continuous field tomography data. 18. The method according to claim 16, wherein the method further comprises obtaining the non-continuous field physiological data. 19. The method according to claim 16, wherein the method further comprises outputting the aggregated data product. 20. The method according to claim 19, wherein the aggregated data product is output to an implantable ... | - | A61B005/05
US20100292545 | A1 | 2010-11-18 | INTERACTIVE PSYCHOPHYSIOLOGICAL PROFILER METHOD AND SYSTEM | 2009-05-14 | An efficient, objective, flexible and easily deployable system for conducting evaluations of mental and physiological state and recommending individualized treatment to improve said state is described. The method and system are based on commensurate measurement of mental functions, levels of stress and anxiety, and/or biologically active molecules such as neurotransmitters, immune markers including cytokines and hormones. The method and system are designed to assess an individual's cognitive function and the underlying physiology in order to delineate various disease processes, injuries, drug states, training stages, fatigue levels, stress levels, aging processes, predict susceptibility to stress and/or sleep deprivation, identify aptitude for training and/or characterize effects of any experimental conditions. The system and method may be used in recommending individualized treatment protocols, as well as to guide the treatment process by assessing the efficacy of such therapies in the clinical trials process. | 1. A method of determining a subject's cognitive and emotional state comprising: obtaining a first physiological measure from the subject which provides baseline data;administering to the subject a test including a plurality of self-paced computerized tasks while obtaining a second physiological measure from the subject which provides task-related data;deriving one or more scores quantifying the subject's cognitive and emotional states utilizing the task-related data; andcreating a profile by determining the significance of the one or more scores with reference to normative data from a plurality of subjects. 2. The method of claim 1, wherein the obtaining the first and second physiological measures includes measuring biologically active molecules from the subject. 3. The method of claim 1, further comprising: deriving or more scores quantifying the subject's immune status utilizing the task-related data. 4. The method of claim 1, further comprising: providing individualized suggestions for treatment options based on the profile results. 5. The method of claim 1, wherein the first and second physiological measures include electroencephalogram (“EEG”) and electrocardiogram (“EKG”). 6. The method of claim 5, wherein the first and second physiological measures further include at least one physiological measure selected from the group including oxygen levels in the blood, brain, or peripheral tissues; respiration; muscle electrical activity; eye movement; galvanic skin reaction; body temperature; and skin temperature. 7. The method of claim 2, wherein the biologically active molecules are selected from the group including blood, saliva, sweat, urine, and hair. 8. The method of claim 1, wherein the test including a plurality of self-paced computerized tasks assesses a behavior selected from the group including learning speed and aptitude; verbal and visuospatial memory; sustained attention; selective attention; distractibility; impulsiveness; adaptability; flexibility in responding; speed in responding; and propensity for risk-taking. 9. The method of claim 5, wherein the subject's level of alertness, mental effort and engagement during each of the self-paced computerized tasks are characterized by: measuring on a second-by-second basis the power of the EEG at frequencies from 1 to 40 Hz in 1 Hz steps, andusing a four-class discriminant function to convert these power measures into probabilities that the subject's level of mental effort and engagement belongs to one of the four classes defined as High Engagement, Low Engagement, Distraction and Drowsiness. 10. The method of claim 5, wherein the subject's mental workload is characterized by measuring on a second-by-second basis EEG recorded at frequencies from 1 to 40 Hz in 1 Hz steps using a two-class discriminant function to convert these power measures into probabilities that the subject is at Low or High workload. 11. The method of claim 8, wherein the subject's working memory is character... | ADVANCED BRAIN MONITORING, INC. | A61B005/0205
US20100280355 | A1 | 2010-11-04 | SYSTEM AND METHOD TO CHARACTERIZE CARDIAC FUNCTION | 2008-12-11 | Systems and methods can quantify cardiac function. In one embodiment, a method (10) for quantifying cardiac function for a patient's heart includes determining (12) an end-systolic strain for each of a plurality of myocardial segments at end systole and determining (14) a peak strain in each of the plurality of myocardial segments. A difference between the peak strain and the end-systolic strain is computed (16) for each of the plurality of myocardial segments. A strain delay index is computed (18) from the computed differences. | 1. A method for quantifying cardiac function for a patient's heart, comprising: determining an end-systolic strain for each of a plurality of myocardial segments;determining a peak strain in each of the plurality of myocardial segments;computing a difference between the peak strain and the end-systolic strain for each of the plurality of myocardial segments; andcomputing a strain delay index from the computed differences. 2. The method of claim 1, generating strain curves for each of the plurality of myocardial segments, the end-systolic strain and the peak strain for each of the plurality of myocardial segments being ascertained from the respective strain curves. 3. The method of claim 2, further comprising determining a global strain curve by averaging the strain curves with respect to time, the global strain curve representing overall strain for ventricular function. 4. The method of claim 2, further comprising determining a timing of end systole as a time at which the global strain curve peaks. 5. The method of claim 1 further comprising: quantifying regional wall motion for a ventricle of the patient' heart; anddetermining longitudinal strain for the plurality of myocardial segments of the ventricle based on the quantified regional wall motion. 6. The method of claim 5, wherein the quantifying regional wall motion further comprises acquiring images of the patient's heart over time to provide corresponding image data, the corresponding image data including a representation of wall motion for the plurality of myocardial segments; and processing the corresponding image data to provide strain curves for the plurality of myocardial segments, the end-systolic strain and the peak strain for each of the plurality of myocardial segments being ascertained from the respective strain curves. 7. The method of claim 6, wherein the acquiring images further comprises employing an ultrasound imaging modality. 8. The method of claim 7, wherein the ultrasound imaging modality comprises two-dimensional speckle tracking echocardiography. 9. The method of claim 7, wherein the acquiring images further comprises employing one of a computed tomography imaging modality and a magnetic resonance imaging modality. 10. The method of claim 6, further comprising: determining a global strain curve by averaging the strain curves with respect to time, the global strain curve representing overall strain for the ventricle; anddetermining timing of end systole as the time at which the global strain curve peaks. 11. The method of claim 5, wherein the plurality of myocardial segments comprise at least twelve myocardial segments. 12. The method of claim 5, wherein the longitudinal strain for the plurality of myocardial segments comprises strain longitudinal strain for at least sixteen myocardial segments of the ventricle. 13. The method of claim 1, wherein the strain delay index is defined as follows:straindelayindex=∑i=1n(ɛpeaki-ɛESi)where: εpeak is the peak strai... | - | A61B005/055 A61B005/02 A61B008/00
US20100317980 | A1 | 2010-12-16 | METHOD AND DEVICE FOR USING A PHYSIOLOGICAL PARAMETER TO EXPRESS EVOLUTION | 2009-06-11 | A method for presenting a bioorganism's evolutionary passage through at least a portion of its life, by reference to a selected physiological parameter emanating from the bioorganism, as an alternative to use of conventional horology, which includes selecting a physiological parameter specific to the bioorganism itself, one which continuously and repeatedly occurs, for the most part involuntarily, throughout and only throughout the bioorganism's life, and then detecting the repeated occurrences generated by the parameter. | 1. A method for presenting a bioorganism's evolutionary passage through at least a portion of its life, by reference to a selected physiological parameter emanating from the bioorganism, as an alternative to use of conventional horology, comprising the steps of: selecting a physiological parameter specific to the bioorganism itself, one which continuously and repeatedly occurs, for the most part involuntarily, throughout and only throughout the bioorganism's life;detecting the repeated occurrences generated by the parameter;recording the occurrences successively and continuously and arranging them in basic units; anddisplaying the units so that the bioorganism's evolving life or segments thereof, over time, can be measured and interpreted by reference to the recorded units, as an alternative to using horology as a measure. 2. The method of claim 1 wherein the displaying step further includes presenting the units simultaneously with each occurrence. 3. The method of claim 2 wherein the recording step includes accumulating and retaining the units so that a cumulative measure of all occurrences generated by the parameter can be selectively displayed. 4. The method of claim 3 wherein the display step includes the capability of selectively displaying the units as they occur, along with a cumulative display of all the units generated over time, or selected period. 5. The method of claim 1 wherein the displaying step is presented digitally, and is selectively viewable by the bioorganism or another. 6. The method of claim 1 wherein the selected occurrences can be regulated at least partially over a selected duration by voluntary effort of the bioorganism. 7. The method of claim 1 wherein cardiac signals generated by the bioorganism itself are selected as the physiological parameter for detecting, recording and displaying. 8. The method of claim 7 wherein the cardiac signals are represented by the bioorganism's pulse, and wherein each basic unit is defined as being a single heartbeat. 9. A method for presenting a bioorganism's evolutionary passage through at least a portion of its life, by reference to a selected physiological parameter emanating from the bioorganism, as an alternative to use of conventional horology, comprising the steps of: selecting cardiac signals generated by the bioorganism itself as the physiological parameter;detecting the cardiac signals generated by the bioorganism;displaying the cardiac signals by arranging them in basic units; andaccumulating the basic units successively and continuously, so that the bioorganism's evolving life or portion thereof, over time, can be measured and interpreted by reference to the recorded basic units, or cardiac signals, as an alternative to using horology as a measure. 10. The method of claim 9 wherein the bioorganism's cardiac signals are represented by the bioorganism's pulse, and wherein each basic unit further defined as being a single heart beat. 11. The method of claim 10 wherein ... | - | A61B005/04
US20100331717 | A1 | 2010-12-30 | Seizure Sensing and Detection Using an Implantable Device | 2010-09-12 | A system and method for detecting and predicting neurological events with an implantable device uses a relatively low-power central processing unit in connection with signal processing circuitry to identify features (including half waves) and calculate window-based characteristics (including line lengths and areas under the curve of the waveform) in an electrographic signal received from a patient's brain. The features and window-based characteristics are combinable in various ways according to the invention to detect and predict neurological events in real time, enabling responsive action by the implantable device. | 1. A method for detecting a neurological event by analyzing an electrical signal from a patient's brain with an implantable device, the method comprising the steps of: receiving a first electrical signal and a second electrical signal from a plurality of electrodes;processing the electrical signals with a detection subsystem to obtain a first detection channel output and a second detection channel output;combining the first detection channel output and the second detection channel output to obtain an event detector output; andcausing the implantable device to perform an action in accordance with the event detector output. | NEUROPACE, INC. | A61B005/0476
US20100298680 | A1 | 2010-11-25 | METHOD AND DEVICE FOR DETERMINING A PROPERTY OF LIVING TISSUE | 2008-01-11 | The invention relates to a measurement of tissue properties, in particular glucose, by measuring the response of the tissue to an applied electric field. The tissue is modeled by a System of homogeneous layers. In one approach, a plurality of electrical fields are generated in the tissue at different frequencies. For each of the fields, a signal depending on the dielectric permittivity as seen by the electrode arrangement at the frequency is measured, thereby generating a measured dataset. In another approach the different electrode configurations can be used to achieve different penetration depths in the desired layers. A function is then fitted to the dataset by varying at least some parameters of the function. These parameters describe the dispersion of the dielectric permittivity of a plurality of layers in the tissue. At least part of the parameters obtained in this fitting procedure are then used for determining the desired tissue property. Furthermore a combination of these two approaches can be used to optimize the uniqueness of the Solution of the fitting procedure for changes at a specific depth. | 1. A method for measuring a property c of living tissue, which property c affects the complex dielectric permittivity ∈(ω) of said tissue, comprising the steps of applying an electrode arrangement to a skin region of said tissue,generating, by means of said electrode arrangement, a plurality of electrical fields in said tissue at different frequencies ωw with w=1 to W and measuring, for each of said frequencies, a signal sw with w=1 to W, depending on the dielectric permittivity ∈(ωw) as seen by said electrode arrangement at the frequency ωw, thereby generating a measured dataset {(s1, ω1), . . . (sW, ωW)},using dispersion parameters pmn with m=1 M with M>1 and n=1 . . . N, wherein said dispersion parameters pmn are parameters of a dispersion function H describing a dispersion of the dielectric permittivity ∈m of a virtual homogeneous tissue layer m in said skin region by ∈m(ω)=H(pm1, . . . , pmN,ω),with m=1 . . . M, fitting a function F1s(ω)=F1(p11, . . . , pMN,ω)to said measured dataset {(s1, ω1), . . . (sW, ωW)} by varying at least part of said dispersion parameters pmn, wherein said function F is given by F1(p11, . . . pWL,ω)=F0(∈1(ω), . . . ∈M(ω))with a function F0(∈1(ω), . . . ∈L(ω)) describing the signal s(ω) measured if said layers 1 . . . M have the dielectric permittivities ∈1(ω), . . . ∈M(ω), said method further comprising the step of using at least part of the varied dispersion parameters pmn for calculating said property c. 2. The method of claim 1, further comprising the steps of deriving said function F0 by obtaining a plurality of vectors vk=(∈1k, . . . ∈Mk, sk) with k=1 . . . K, wherein each vector vk comprises the signal sk that would be measured at said electrode arrangement if said layers had the dielectric permittivities ∈Mk,fitting a model function L s=L(r1, . . . rT, . . . ∈Mk)to said vectors vk by varying model parameters r1, . . . rT of said model function L and using the varied model parameters r1, . . . rT for calculating F0(∈1(ω), . . . ∈L(ω))=L(r1, . . . rT,∈1(ω),∈L(ω)). 3. The method of claim 2, wherein said model function L is linear in r1, . . . rT. 4. The method of claim 1, whereinɛm(ω)=H(pm1,…,pmN,ω)=ɛ∞,m+ɛ0,m-ɛ∞,m[1+(j·ω·τm)αm]βm,where0≤α,β≤1with ∈∞,m=pm1, αm=pm2, βm=pm3, τm=pm4, ∈0,m=pm5, and N=5. 5. The method of claim 1, whereinɛm(ω)=H(pm1,…,pmN,ω)=(ɛm(ω))α=∑q=1Qvm,q(ɛq)αwhere vm,q is a volume fraction of the q-th component of a mixture in layer m, ∈q the complex dielectric permittivity of the q-th component and Q the number of components in the mixture, ∈q is ∈q(ω) for at least some values of q and α a number between −1 and 1. 6. The method of claim 1, wherein M=2. 7. The method of any claim 1, wherein a thickness of a topmost layer of said skin area is between 10 and 300 μm. 8. A method for measuring a property c of living tissue, which property c affects the complex dielectric permittivity c of said tissue, comprising the steps of: applying an electrode arrangement to a skin region of... | - | A61B005/1477 A61B005/053
US07844317 | B2 | 2010-11-30 | Method and system for estimating three-dimensional respiratory motion | 2003-11-26 | One or more techniques are provided for measuring the motion of an organ in three dimensions. As provided by the technique, the motion of the organ along each dimension may be determined by a suitable methodology. Where sensor-based motion measurements are suitable, one or more sensors may be placed on a patient to measure internal motion of the organ of interest along one or more perpendicular axes. Where image-based techniques are suitable, the motion of the internal organ along a perpendicular axis may determined using pre-acquisition image data or acquisition image data when suitable. Concurrent motion vectors for all three dimensions may be obtained from the motion data acquired for the perpendicular axes by the disparate methodologies. The concurrent motion vectors may be combined to describe the three-dimensional motion of the organ over time. Validation of the motion data may be performed for each of the one-dimensional motion data sets using motion data acquired by image-based methods, or other image-based methods, for a respective axis. | 1. A method for determining the motion of an organ, comprising the steps of: acquiring a first set of one-dimensional motion data for an organ along a first axis by a first methodology;acquiring a second set of one-dimensional motion data for the organ along a second axis by a second methodology, wherein the first axis and the second axis are perpendicular;acquiring a third set of one-dimensional motion data for the organ along a third axis by a third methodology, wherein the third axis is perpendicular to the first axis and the second axis;deriving one or more concurrent motion vectors from each of the first, second, and third sets of one-dimensional motion data; andcombining the one or more concurrent motion vectors to generate a set of three-dimensional motion data for the organ. 2. The method as recited in claim 1, wherein: acquiring at least one set of one-dimensional motion data comprises measuring the motion along the one or more respective axes with a set of one or more sensors. 3. The method as recited in claim 1, wherein: acquiring at least one of the first, second, or third set of one-dimensional motion data comprises validating the one or more sets of one-dimensional motion data using one or more respective sets of validation motion data. 4. The method as recited in claim 1, wherein acquiring at least one set of one-dimensional motion data comprises determining the motion along the one or more respective axes from a respective set of motion data derived from an imager. 5. The method as recited in claim 4, wherein the respective set of motion data is derived from a set of acquisition image data or pre-acquisition image data. 6. The method as recited in claim 5, wherein the set of acquisition image data comprises a set of unreconstructed image data or a set of reconstructed image data. 7. The method as recited in claim 1, wherein two of the first methodology, the second methodology or the third methodology comprise the same methodology. 8. The method as recited in claim 1, wherein at least one of the first methodology, the second methodology or the third methodology comprises a sensor-based methodology, and wherein at least one of the first methodology, the second methodology or the third methodology comprises an image-based methodology. 9. The method as recited in claim 1, wherein the first methodology, the second methodology or the third methodology comprise one or more data-based methodologies, wherein the one or more data-based methodologies determine motion from one or more respective sets of acquisition image data. 10. The method as recited in claim 9, wherein the one or more respective sets of acquisition image data comprise one or more sets of unreconstructed image data. 11. The method as recited in claim 9, wherein the one or more respective sets of acquisition image data comprise one or more sets of reconstructed image data. 12. A—non-transitory—computer readable storage medium having executable code stored thereon,... | General Electric Company | A61B005/05
US20100282616 | A1 | 2010-11-11 | Method of Calibrating an Analyte-Measurement Device, and Associated Methods, Devices and Systems | 2010-07-26 | The invention relates to a method for calibrating an analyte-measurement device that is used to evaluate a concentration of analyte in bodily fluid at or from a measurement site in a body. The method involves measuring a concentration, or calibration concentration, of an analyte in blood from an “off-finger” calibration site, and calibrating the analyte-measurement device based on that calibration concentration. The invention also relates to a device, system, or kit for measuring a concentration of an analyte in a body, which employs a calibration device for adjusting analyte concentration measured in bodily fluid based on an analyte concentration measured in blood from an “off-finger” calibration site. | 1-22. (canceled) 23. A method for calibrating a signal from a subcutaneous sensor, comprising: obtaining a calibration measurement from a calibration sensor placed in contact with capillary blood from an off-finger calibration site within a body;evaluating whether the calibration measurement is within a predetermined range of analyte concentration; anddetermining that the calibration measurement is suitable for use in converting an analyte signal from the subcutaneous sensor into an analyte concentration if the calibration measurement is within the predetermined range,wherein the subcutaneous sensor comprises a working electrode, the working electrode comprising an analyte responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with a polymer disposed on the working electrode. 24. The method of claim 23, wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer by one or more covalent bonds to the polymer. 25. The method of claim 24, wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer by one or more carbon-carbon, carbon-nitrogen, or metal-carbon covalent bonds to the polymer. 26. The method of claim 25, wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer through one or more crosslinking agents. 27. The method of claim 26, wherein the crosslinking agent is one or more of substituted or unsubstituted epoxides, aldehydes, imidoesters, N-hydroxysuccinimide esters, carbodiimides, cyanuric chloride, tetrachlorobenzoquinone, benzoquinone and tetracyanoquinodimethane. 28. The method of claim 24, wherein the mediator is crosslinked with the polymer by at least one covalent bond to one or more ligands of the mediator. 29. The method of claim 28, wherein the one or more ligands comprises a substituted or unsubstituted heterocyclic nitrogen-containing moiety. 30. The method of claim 24, wherein one or more ligands of the mediator are crosslinkers of the polymer. 31. The method of claim 24, wherein one or more ligands of the mediator are substituents of the polymer backbone. 32. The method of claim 24, wherein the analyte-responsive enzyme is crosslinked with the polymer by at least one covalent bond to one or more side chains of the enzyme. 33. The method of claim 32, wherein the one or more side chains of the analyte-responsive enzyme is selected from the group consisting of substituted or unsubstituted amines, alcohols, thiols, phenols, imidazoles, indols, and carboxylic acids. 34. The method of claim 32, wherein the one or more side chains of the analyte-responsive enzyme are crosslinkers of the polymer. 35. The method of claim 32, wherein the one or more side chains of the analyte-responsive enzyme are substituents of the polymer backbone. 36. The method of claim 23, wherein the polymer forms a coordination complex with the mediator... | Abbott Diabetes Care Inc. | G01N027/26 A61B005/145
US20100305466 | A1 | 2010-12-02 | INCENTIVE SPIROMETRY AND NON-CONTACT PAIN REDUCTION SYSTEM | 2010-07-01 | A non-contact mechanism for encouraging and facilitating incentive spirometry, ensuring that it is performed adequately, in a timely manner, and for a sufficient duration is discussed. The embodiments also quantitatively and qualitatively keep a record of the incentive spirometry activity, including recording the performance in an electronic medical record. A non-contact monitoring system is used to generate a breathing waveform for a subject that may be compared to target waveforms. Visual and non-visual cues may then be provided to the subject to help guide the subject towards the desired breathing pattern. | 1. A non-contact monitoring system for monitoring incentive spirometry exercises, comprising: a respiratory waveform detection module, the respiratory waveform detection module performing non-contact monitoring of a subject performing incentive spirometry exercises, the respiratory waveform detection module generating a waveform based on detected respiratory motion of the subject;an analysis module programmatically analyzing the generated waveform based on a stored target waveform indicative of a target respiratory motion for an incentive spirometry exercise; anda biofeedback module providing biofeedback to the subject to assist the subject in obtaining or maintaining the target waveform, the biofeedback based on a result of the analyzing of the generated waveform. 2. The system of claim 1, further comprising: a display, the display used to provide the biofeedback in the form of a display of at least one of the generated waveform and the target waveform. 3. The system of claim 2 wherein the display of the generated waveform is accompanied by audible or visual instructions for the subject to increase or decrease a depth of breathing so as to attain the target waveform. 4. The system of claim 1, further comprising: a display, the display used to provide the biofeedback in the form of a display of an intermediate waveform that represents a waveform between the generated waveform and the target waveform. 5. The system of claim 1, further comprising: an auditory module that is used as an adjunct or to provide feedback in the form of audio transmissions detectable by the subject. 6. The system of claim 1 wherein the respiratory waveform detection module and the analysis module communicate over a network. 7. An integrated non-contact monitoring apparatus for monitoring incentive spirometry exercises, comprising: a respiratory waveform detection module, the respiratory waveform detection module performing non-contact monitoring of a subject performing incentive spirometry exercises, the respiratory waveform detection module generating a waveform based on detected respiratory motion of the subject;an analysis module programmatically analyzing the generated waveform based on a stored target waveform indicative of a target respiratory motion for an incentive spirometry exercise;a biofeedback module providing biofeedback to the subject to assist the subject in obtaining or maintaining the target waveform, the biofeedback based on a result of the analyzing of the generated waveform; anda display surface, the display surface used to provide the biofeedback in the form of at least one of a display of the generated waveform and the target waveform. 8. The apparatus of claim 7, further comprising: an auditory module that is used to provide feedback in the form of audio transmissions detectable by the subject. 9. The apparatus of claim 7 wherein the respiratory waveform detection module monitors the subject using radiated energy. 10. The apparatus of clai... | ENGINEERED VIGILANCE, LLC | A61B005/087
US07828795 | B2 | 2010-11-09 | Surgical ablation and pacing device | 2006-02-28 | A surgical device has an integral first tip having pair of electrodes configured to ablate tissue using electric energy. A second tip has a pair of electrodes configured to provide pacing signals to a heart and/or to sense electrical signals passing through heart tissue. The second tip is configured to snap onto the first tip, such that the same device may be used for ablation, pacing, and sensing. Alternatively, the second tip may be integral with the device and the first tip configured to snap onto the second tip. Alternatively, a single integral tip of the surgical device may be used for ablation, pacing, and sensing. Such a multipurpose tip may comprise a plurality of electrode pairs or an array of electrodes. A user interface on the device or elsewhere may be operable to provide selectable modes for selectively varying properties of signals communicated to the electrodes. | 1. A surgical device for ablating tissue and pacing a heart using electrical energy, the device comprising: a) a handle;b) a head connected to the handle, wherein the head has a transverse axis and comprises a first tip portion having a distal-facing surface convexly-curved relative to the transverse axis, wherein the first tip portion comprises a first pair of convexly-curved electrodes on the surface thereof adapted to contact tissue and operable to communicate an electrical signal;c) a pair of conductors, wherein each conductor is in communication with a respective one of the pair of electrodes of the first tip portion, wherein each of the conductors is configured to selectively deliver either pacing signals or ablation energy from a power source to tissue; andd) a second tip comprising a second pair of electrodes on the surface thereof adapted to contact tissue for the delivery of pacing signals, wherein the second tip is configured to be removably secured to the head, wherein the first pair of electrodes conducts electrical signals to the second pair of electrodes when the second tip is secured to the head. 2. The surgical device of claim 1, wherein the head is connected to the handle via a malleable shaft. 3. The surgical device of claim 2, wherein a portion of the pair of conductors is located within the malleable shaft. 4. The surgical device of claim 1, wherein the electrodes of one or both of the first tip portion or the second tip are configured to ablate tissue in accordance with an electrical signal communicated through the conductors. 5. The surgical device of claim 1, wherein the electrodes of one or both of the first tip portion or the second tip are configured to pace a heart in accordance with an electrical signal communicated through the conductors. 6. The surgical device of claim 1, wherein the electrodes of the first tip portion are configured to ablate tissue in accordance with an ablation signal communicated through the conductors, wherein the electrodes of the second tip are configured to pace a heart in accordance with a pacing signal communicated through the conductors. 7. The surgical device of claim 1, wherein each electrode of one or both of the first tip portion or the second tip is parallel with the other electrode of the corresponding pair. 8. The surgical device of claim 1, wherein each of the electrodes of the second tip are in communication with a respective leaf spring, wherein each of the leaf springs is configured to provide electrical continuity between the head and the electrodes of the second tip when the second tip is secured to the head. 9. The surgical device of claim 8, wherein each of the leaf springs is configured to engage a corresponding electrode of the first tip portion. 10. The surgical device of claim 1, wherein the second tip is configured to snap onto the head adjacent to the first tip portion. 11. The surgical device of claim 1, wherein one or both of the first tip portion or the s... | AtriCure, Inc. | A61B018/14 A61B005/04 A61N001/05
US07831302 | B2 | 2010-11-09 | Monitoring electrical muscular activity | 2004-03-16 | A method of monitoring electrical activity non-invasively (such as uterine activity) which includes applying a localized group of electrodes to a patient's skin and monitoring signals thereon. The electrodes are localized sufficiently such that their muscular signal contributions simulate a single source despite source non-stationarity. The signals are amplified, filtered and digitized. They are then digitally filtered and processed by independent component analysis (ICA) to separate a muscular activity source from other sources. The method may be used to monitor maternal uterine activity, fetal activity and maternal and fetal cardiac activity simultaneously with the aid of additional electrodes and associated circuitry. | 1. A method for non-invasive monitoring of electrical muscular activity which is at least partially due to a non-stationary muscular source, the method incorporating the steps of: a) providing a blind signal separation technique suitable for separating stationary signals,b) placing a plurality of low-noise signal electrodes externally upon a patient's skin for detection of electrical muscular activity, the signal electrodes being localised sufficiently such that: i) their muscular signal contributions simulate a single muscular source to the blind signal separation technique despite at least partial non-stationarity of the muscular source, andii) the number of sources detected by the blind signal separation technique is not more than the number of signal electrodes;c) using computer apparatus to apply the blind signal separation technique to digital signals derived from signals received from the signal electrodes to separate the muscular source, andd) using a display device to display the separated muscular source to a user. 2. A method according to claim 1 wherein the muscular activity is uterine activity. 3. A method according to claim 1 wherein the blind signal separation technique is based on an algorithm of a kind known as an instantaneous algorithm and suitable for addressing blind signal separation problems referred to as instantaneous mixing problems, the instantaneous algorithm incorporating an assumption that signals arrive synchronously at each sensor in a sensor array. 4. A method according to claim 3 wherein the instantaneous algorithm is independent component analysis (ICA). 5. A method according to claim 4 wherein the step of applying the blind signal separation technique applies ICA to processing data derived from signals from the signal electrodes, the data being arranged in successive overlapping blocks such that in pairs of adjacent blocks each subsequent block incorporates a proportion of the data in the respective preceding block, and a correlation scheme is applied to re-order independent sources derived in ICA processing of different blocks to correct for signal swapping. 6. A method according to claim 1 wherein the step of placing the signal electrodes comprises placing four or five signal electrodes at and above navel height with respect to an upright patient at positions close to the expected site of pacemaker activity. 7. A method according to claim 1 wherein the signal electrodes are a first set of signal electrodes and the step of placing the signal electrodes includes placing a second set of signal electrodes upon a patient's skin in positions not localised sufficiently for their muscular signal contributions to simulate a single source to the blind signal separation technique, and wherein the blind signal separation technique employs signals derived via the first set of signal electrodes for monitoring non-stationary muscular activity and signals derived via the first and second sets of signal electrodes for m... | Qinetiq Limited | A61B005/04 A61B005/0488 A61B005/103 A61B005/117
US07853459 | B2 | 2010-12-14 | Automated processing of electronic medical data for insurance and disability determinations | 2008-08-14 | A machine-implemented method of assisting medical evidence collection for adjudication of a medical disability request is described. In some embodiments, the method includes providing an electronic medical record of a patient, the electronic medical record includes evidence of at least one claimable medical disability condition and locating in the electronic medical record at least one keyword associated with the at least one claimable medical disability condition. In some embodiments, the method further includes processing an inference rule associated with the at least one claimable medical disability condition, such that the evidence of the at least one claimable medical disability condition is identified in the electronic medical record and outputting, to a computer-readable medium, information indicative of (1) the evidence of the at least one claimable medical disability condition, and (2) a location of the evidence in the electronic medical record. | 1. A machine-implemented method of assisting medical evidence collection for adjudication of a medical disability request, the method comprising: providing an electronic medical record of a patient, the electronic medical record comprising evidence confirming or disconfirming at least one claimable medical disability condition;locating in the electronic medical record at least one keyword associated with the at least one claimable medical disability condition;processing, by a computer, an inference rule associated with the at least one claimable medical disability condition, such that the evidence confirming or disconfirming the at least one claimable medical disability condition is identified in the electronic medical record; andoutputting, to a display, (1) information indicative of the identified evidence confirming or disconfirming the at least one claimable medical disability condition, and (2) a location reference of the identified evidence in the electronic medical record;wherein the processing an inference rule comprises locating an indicator of a value in the electronic medical record that tends to confirm an existence of the claimable medical disability condition in the patient; andwherein the outputting results in the display providing a visible association between the identified evidence confirming or disconfirming the at least one claimable medical disability condition, and the location reference of the identified evidence in the electronic medical record. 2. The method of claim 1, wherein the inference rule is processed if the at least one keyword is present in the electronic medical record. 3. The method of claim 1, further comprising indexing the electronic medical record prior to the locating in the electronic medical record the at least one keyword. 4. The method of claim 1, wherein the at least one keyword is selected from a plurality of keywords associated with the at least one claimable medical disability condition. 5. The method of claim 4, further comprising determining, for each of the plurality of keywords, whether that keyword is included in the electronic medical record. 6. The method of claim 1, wherein the inference rule comprises searching, in the electronic medical record, for information associated with the at least one keyword. 7. The method of claim 1, wherein the inference rule utilizes at least one of information on a category associated with the at least one keyword, a rating code, and a search query. 8. The method of claim 1, wherein the information indicative of the identified evidence confirming or disconfirming the at least one claimable medical disability condition includes at least one of a secondary condition associated with the at least one claimable medical disability condition, a category associated with the at least one claimable medical disability condition, and a rating code associated with the at least one claimable medical disability condition. 9. The method of claim 1, wherein the visib... | QTC Management, Inc. | A61B005/00 G06F019/00
US07957725 | B2 | 2011-06-07 | Methods for providing enhanced telecommunication services | 2010-06-08 | A first method for facilitating personal contacts among a plurality of subscribers to a location-based person contact service provided by a service provider; a second method for communicating information to a user of a wireless device by a server; and a third method for real-time monitoring of vital signs of a living being. In the first method, personal profiles are transmitted to subscribers. In the second method, a message specific to both an establishment and the user is transmitted to a wireless device or smart display proximal to the user in the establishment. In the third method, information relating to a violated vital sign parameter determined from the monitoring is communicated to a local communication node. | 1. A method for communicating information to a user of a wireless device by a customer relationship management (CRM) server, comprising: receiving data from the wireless device being carried by the user in an establishment, said data identifying the wireless device or relating to the user;cross referencing the received data against a transaction history associated with the device or the user to determine demographic information concerning the user;generating a message specific to both the establishment and the user, said message being a function of the determined demographic information concerning the user; andtransmitting the message to an instrument selected from the group consisting of the wireless device and smart display proximal to the user in the establishment, wherein said receiving, said cross referencing, said generating, and said transmitting are performed by the CRM server. 2. The method of claim 1, wherein the data identifies the wireless device. 3. The method of claim 1, wherein the data relates to the user. 4. The method of claim 1, wherein the transaction history is in a database comprised by the CRM server. 5. The method of claim 1, wherein the instrument is the wireless device. 6. The method of claim 1, wherein the instrument is the smart display. 7. The method of claim 6, wherein the smart display is a smart computer-driven display. 8. A system comprising a local server adapted to perform the method of claim 1, said local server being the CRM server. | - | H04M011/00
US08000763 | B2 | 2011-08-16 | Integrated devices having extruded electrode structures and methods of using same | 2009-06-30 | The present disclosure provides electrode structures and integrated electrode structures having one or more conductive materials coextruded with one or more dielectric materials. The disclosed electrode structures can be configured for use as analyte sensors. Also provided, are methods of making and using the electrode structures and integrated electrode structures described herein. | 1. An integrated medication delivery device and electrode structure, comprising: an electrode structure comprising a first conductive material, a second conductive material, and a dielectric material,wherein the first conductive material, the second conductive material, and the dielectric material are coextruded to provide an electrode structure having the first conductive material and the second conductive material electrically isolated by the dielectric material,wherein the first conductive material comprises a working electrode and the second conductive material comprises a reference/counter electrode, the working electrode comprising an analyte responsive enzyme disposed on, in or proximate thereto,wherein the electrode structure comprises a lumen, the lumen comprising a first opening, a second opening, and a lumen wall,wherein the dielectric material defines the lumen wall, andwherein the electrode structure includes a plurality of stand-offs along the lumen wall and extending inward into the lumen; anda medication delivery device coupled to the lumen and configured to deliver a medication through the lumen to the subject. 2. The integrated medication delivery device and electrode structure of claim 1, wherein the medication delivery device comprises a medication delivery tube comprising a third conductive material and a fourth conductive material, the third and fourth conductive materials being in electrical contact with the first and second conductive materials respectively. 3. The integrated medication delivery device and electrode structure of claim 1, wherein the medication is insulin. 4. The integrated medication delivery device and electrode structure of claim 1, wherein the first conductive material comprises carbon. 5. The integrated medication delivery device and electrode structure of claim 1, wherein the second conductive material comprises Ag/AgCl. 6. The integrated medication delivery device and electrode structure of claim 1, further comprising a membrane disposed over the working electrode. 7. The integrated medication delivery device and electrode structure of claim 1, wherein the analyte-responsive enzyme is a glucose-responsive enzyme. 8. The integrated medication delivery device and electrode structure of claim 1, wherein the electrode structure is configured for continuous glucose monitoring. 9. The integrated medication delivery device and electrode structure of claim 1, wherein the analyte-responsive enzyme is a ketone-responsive enzyme. 10. The integrated medication delivery device and electrode structure of claim 1, wherein the electrode structure is configured to engage a connector, the connector comprising first and second electrode contacts for engaging the first conductive material and the second conductive material respectively. 11. The integrated medication delivery device and electrode structure of claim 1, wherein the medication delivery device is a pump. 12. The integrated medication delivery devi... | Abbott Diabetes Care Inc. | A61B005/05
US08285408 | B2 | 2012-10-09 | System and method for designing an ear insert device | 2009-06-15 | A method for designing an ear insert device using anatomical information relating to the ear. A user record is that includes anatomical information relating to an ear of a user is received. The anatomical information comprises at least one sub-dermal feature of an ear canal. The user record is processed to obtain an ear insert device design record. The ear insert device design record comprises a three dimensional representation of a bounding surface shape having surface boundaries that substantially conform to surface boundaries of the ear canal and the at least one sub-dermal feature of the ear canal. An ear insert device may be produced using the ear insert design record. | 1. A method for designing an ear insert device comprising: receiving by a processor a user record, wherein the user record comprises anatomical information relating to an ear of a user and wherein the anatomical information comprises three dimensional measurement data of at least one sub-dermal feature of an ear canal; andprocessing by the processor the user record to obtain an ear insert device design record, wherein the ear insert device design record comprises a three dimensional representation of a bounding surface shape having surface boundaries that substantially conform to surface boundaries of the ear canal and the at least one sub-dermal feature of the ear canal. 2. The method of claim 1 further comprising using the ear insert device design record to produce an ear insert device that substantially conforms to the surface boundaries of the ear canal and the at least one sub-dermal feature of the ear canal. 3. The method of claim 2, wherein the ear insert device comprises a distal end and wherein using the ear insert device design record to produce an ear insert device comprises using the ear insert device design record to produce an ear insert device having the distal end terminated relative to the at least one sub-dermal anatomical feature of the ear canal. 4. The method of claim 2, wherein the ear insert device comprises a distal end and wherein using the ear insert device design record to produce an ear insert device comprises using the ear insert device design record to produce an ear insert device having the distal end shaped to accommodate the at least one sub-dermal feature of the ear canal. 5. The method of claim 1, wherein the at least one sub-dermal feature is selected from the group consisting of a most proximal surface of the tympanic part of the temporal bone, a soft tissue of the ear canal, and a Condyloid process of the mandible. 6. The method of claim 2, wherein using the ear insert device design record to produce an ear insert device that substantially conforms to the surface boundaries of the ear canal and the at least one sub-dermal feature of the ear canal comprises using the ear insert device design record to produce an ear insert device that when inserted into the ear canal clears the at least one sub-dermal feature by about 0 to about 3 millimeters. 7. The method of claim 2, wherein using the ear insert device design record to produce an ear insert device that substantially conforms to the surface boundaries of the ear canal and the at least one sub-dermal feature of the ear canal comprises using the ear insert device design record to produce an ear insert device that when inserted in the ear canal deforms the ear insert device by about 0 to about 3 millimeters. 8. The method of claim 2, wherein using the ear insert device design record to produce an ear insert device that substantially conforms to the surface boundaries of the ear canal and the at least one sub-dermal feature of the ear canal comprises using... | Adaptive Technologies Inc. | G06F019/00
US08792950 | B2 | 2014-07-29 | Oximetry with remote display | 2010-06-04 | A device includes a first sensor coupler that is configured to receive a first input signal from a first sensor. The first input signal corresponds to a first physiological parameter and is based on optical excitation of a tissue. The device includes a processor coupled to the first sensor coupler. The processor is configured to generate an output signal based on the first input signal. The first physiological parameter is encoded in the output signal. The output signal differs from the first input signal. The device includes an output coupler configured to communicate the output signal to a remote device. | 1. A device comprising: a first sensor coupler configured to receive a first input signal from a first sensor, the first input signal corresponding to a non-pulsatile direct current (DC) component of a first physiological parameter and based on optical excitation of a tissue;a processor coupled to the first sensor coupler, the processor configured to generate an output signal based on the first input signal, the first physiological parameter encoded in the output signal and the output signal differing from the first input signal, and wherein the output signal corresponds to a measure of tissue oxygenation; andan output coupler configured to communicate the output signal to a remote device. 2. The device of claim 1 wherein the first sensor coupler includes at least one of an electrical connector and an optical connector. 3. The device of claim 1 wherein at least one of the first sensor coupler and the output coupler includes at least one of a wireless receiver or a wireless transmitter. 4. The device of claim 1 further including a second sensor coupler configured to receive a second input signal from a second sensor, the second input signal corresponding to a second physiological parameter and wherein the processor is coupled to the second sensor coupler, the output signal based on the second input signal. 5. The device of claim 1 further including a user-operable switch coupled to the processor. 6. The device of claim 1 wherein the processor is configured to generate a timing signal. 7. The device of claim 1 wherein the processor includes at least one of an amplifier, an analog-to-digital converter, a digital-to-analog converter, a multiplexer, and a filter. 8. The device of claim 1 wherein the output coupler includes at least one of an electrical connector and a wireless transmitter. 9. The device of claim 1 wherein the output signal is encoded with a calculated value configured for rendering on a visual display. 10. A method comprising: exchanging a first timing signal with a first sensor, the first sensor configured to provide a detected signal corresponding to an optical property of a tissue correlated with the first timing signal;receiving the detected signal;processing the detected signal to generate an output signal, the output signal corresponding to a parameter of the tissue, the parameter based on a non-pulsatile direct current (DC) component of the detected signal, and the output signal different than the detected signal, the parameter corresponding to a measure of tissue oxygenation; andcommunicating the output signal to a remote device. 11. The method of claim 10 wherein exchanging the first timing signal includes at least one of sending or receiving relative to the first sensor. 12. The method of claim 10 further including exchanging a second timing signal relative to a second remote sensor based on the exchanging of the first timing signal. 13. The method of claim 10 wherein receiving the detected signal includes receiv... | Nonin Medical, Inc. | A61B005/00
US08644579 | B2 | 2014-02-04 | System and method for diagnosing onset of osteoarthritis | 2010-06-10 | The present invention relates to a method of diagnosing/predicting a degenerative disease in the hip of a subject. The method including the steps of applying a compressive force on at least one hip joint of the subject and calculating a compression index value for at least one hip joint, wherein a non-zero compression index value for the at least one hip joint is indicative of the subject developing/having a degenerative disease. The present invention also relates to a diagnostic test for determining incongruity in a hip joint of a subject, where the test includes the steps of applying a compressive force on at least one hip joint of the subject and calculating a compression index value, wherein a non-zero compression index value for the hip joint is indicative of at least mild hip joint incongruity. | 1. A method of predicting a degenerative disease in a hip of a canine, comprising: applying a compressive force on at least one hip joint of the canine;obtaining at least one radiographic image of the at least one hip joint;determining, via computer software, centers of the acetabulum and centers of the femoral heads from the at least one radiographic image;measuring, via computer software, the distance between the acetabulum and the femoral head centers from the at least one radiographic image;calculating, via computer software, a compression index value for the at least one hip joint of the canine by dividing the distance between the acetabulum and the femoral head centers by the radius of the femoral head;determining the at least one hip joint of the canine is developing a degenerative disease or is at risk for developing a degenerative disease based on a non-zero compression index value. 2. The method of claim 1, wherein the non-zero compression index value for the at least one hip joint is further indicative of hip joint incongruity. 3. The method of claim 1, wherein the degenerative disease is osteoarthritis. 4. The method of claim 3, wherein the non-zero compression index value for the at least one hip joint is indicative of osteoarthritis in the subject at or before osteoarthritis is radiographically apparent in the subject. 5. The method of claim 1, wherein the non-zero compression index value is greater than or equal to about 0.01. 6. The method of claim 1, wherein the non-zero compression index value is greater than or equal to about 0.03. 7. The method of claim 1, wherein the calculation of compression index value comprises use of circle gauges on the radiographic images to find the centers of the acetabulum and centers of the femoral heads. 8. The method of claim 7, wherein the calculation further comprises superimposing precision-machined circle gauges on the cortical margins (rims) of the acetabulum and femoral heads to find the respective geometric centers. 9. The method of claim 7, wherein the calculation further comprises a software program that uploads digital radiographic images to a high-resolution computer screen. 10. A method for determining incongruity in a hip joint of a canine, comprising: applying a compressive force on at least one hip joint of the canine;obtaining at least one radiographic image of the at least one hip joint;determining, via computer software, centers of the acetabulum and centers of the femoral heads from the at least one radiographic image;measuring, via computer software, the distance between the acetabulum and the femoral head centers from the at least one radiographic image;calculating, via computer software, a compression index value for the at least one hip joint of the canine by dividing the distance between the acetabulum and the femoral head centers by the radius of the femoral head;determining the at least one hip joint of the canine has at least mild hip joint incongruity based on... | The Trustees of the University of Pennsylvania | G06K009/00
US08487771 | B2 | 2013-07-16 | Personal health management device | 2009-05-21 | A personal health management device is described. The device is part of a system comprising at least one console, one or more devices in wireless communication with the console wherein the console acts as an intelligent gateway through which the one or more devices may communicate to external data and voice networks, and an alert system wherein an alert is triggered by one or more of a user's action and pre-set alarm criteria and wherein the triggering of an alert causes the console to take an action. The device may be a wearable device or it may be implemented in software and reside on third-party hardware. | 1. A personal health management device comprising: a radio providing two-way wireless communication with a console unit;a display showing alphanumeric and graphic information;a user interface to support functions of said personal health management device;a communication means coupled to external data and voice networks;a wireless communication connection means for receiving vital sign measurements from one or more stand-alone devices and for transmitting same to said console unit;a motion sensor detecting the presence or absence of motion in a user of said personal health management device; anda relay device maintaining wireless communication with said console unit, wherein said relay device provides connection between said personal health management device and said external data and voice networks when said console unit is out of range, and wherein said console unit provides connection between said personal health management device and said external data and voice networks when said console unit is within range. 2. The device according to claim 1 wherein said personal health management device is implemented in software and resides on third-party hardware wherein said hardware provides said two-way wireless communication, said showing alphanumeric and graphic information, said user interface, said connection to external data and voice networks, and said detecting motion. 3. The device according to claim 1 wherein said showing alphanumeric and graphic information further comprises showing current time. 4. The device according to claim 1 wherein said user interface to support functions of said personal health management device comprises push buttons on said personal health management device. 5. The device according to claim 1 wherein said personal health management device is a wearable device comprising a bracelet, pendant, wrist-watch, or a belt attachment and wherein said relay device is carried on said user's person. 6. The device according to claim 1 wherein said communication means coupled to said external data and voice networks comprises communicating with external data and voice networks by one or more of the following: analog telephony, Voice over Internet Protocol (VOIP) over Ethernet, WiFi, and cellular wireless networks. 7. The device according to claim 1 wherein said detecting the presence or absence of motion in a user comprises: providing a motion sampling interval, a polling interval, and alert criteria set at said console unit;activating said motion sensor at the beginning of said motion sampling interval;deactivating said motion sensor at the end of said motion sampling interval or upon detection of a motion in said user;sending motion information from said personal health management device to said console unit indicating motion or no motion for said polling interval; andsending an alert from said console unit to said personal health management device based on said motion information and said alert criteria. 8. A system com... | SilverPlus, Inc. | G08B023/00
US08209003 | B2 | 2012-06-26 | Morphology discrimination based on inflection point-related information | 2010-08-11 | A morphology discrimination scheme extracts shape characteristics from cardiac signals and identifies an associated cardiac condition based on the shape characteristics. For example, internal data structures may be updated to match the shape characteristics of a known condition (e.g., a patient's normal sinus rhythm). Similarly acquired shape characteristics obtained in conjunction with a later event (e.g., QRS complexes acquired during a tachycardia episode) may be compared with the previously stored shape characteristics to characterize the later event. In some aspects the shape characteristics relate to inflection points of cardiac signals. | 1. A method of processing cardiac signals, comprising: acquiring a first set of cardiac signals;using a processor of a cardiac device to: generate derivatives of the first set of cardiac signals to provide sets of optima associated with the signals, wherein each of the optima relates to a corresponding inflection point;rank the optima according to magnitude to provide rank vectors;generate derivatives of the rank vectors to provide vectors having elements indicative of magnitude of change in the ranking of each of the inflection points; andgenerate a baseline matrix based on the vectors having elements indicative of magnitude of change in the ranking of each of the inflection points;acquiring a second set of cardiac signals;generating a set of vectors based on inflection points of the second set of cardiac signals; andscoring the set of vectors based on inflection points of the second set of cardiac signals using the baseline matrix to characterize the second set of cardiac signals. 2. The method of claim 1, wherein: the first set of cardiac signals is associated with sinus rhythm; andthe second set of cardiac signals is associated with an arrhythmia. 3. The method of claim 1, wherein the second set of cardiac signals is characterized as being associated with either ventricular tachycardia or super-ventricular tachycardia. 4. The method of claim 1, further comprising commencing the acquisition of the second set of signals based on an indication of tachycardia. 5. The method of claim 1, wherein the scoring of the set of vectors using the baseline matrix provides an indication as to whether shapes associated with the second set of cardiac signals are similar to corresponding shapes associated with the first set of cardiac signals. 6. The method of claim 1, wherein the baseline matrix comprises a representation of frequency and relative position at which inflection point related characteristics of the first set of cardiac signals occur. 7. A method of processing cardiac signals, comprising: acquiring a first set of cardiac signals;using a processor of a cardiac device to generate a baseline matrix based on inflection points of the first set of cardiac signals;acquiring a second set of cardiac signals;using a vector generator component of an implantable cardiac device to generate a set of vectors based on inflection points of the second set of cardiac signals; andscoring the set of vectors using the baseline matrix to characterize the second set of cardiac signals, wherein:the baseline matrix is based on a first set of derivatives of the first set of cardiac signals;the set of vectors is based on a second set of derivatives of the second set of cardiac signals;the baseline matrix is further based on local optima of the first set of derivatives;the set of vectors is further based on local optima of the second set of derivatives;the baseline matrix is further based on a relative ranking of the local optima of the first set of derivatives; andthe ... | Pacesetter, Inc. | A61B005/0402
US08057400 | B2 | 2011-11-15 | System and method of measuring changes in arterial volume of a limb segment | 2009-06-12 | A medical diagnostic device performs diagnostics for assessing the ability of the arteries to respond to an increase in blood flow. The medical diagnostic device determines relative changes in arterial volume of the limb segment during a time period after a stimulus relative to the arterial volume of the limb segment during a baseline period using the amplitudes or other portions of the component pulse waves (such as early systolic components) of volume pulse waves during the baseline period and after the stimulus. | 1. A method for assessing change in arterial volume of a limb segment of a mammal, said method comprising: determining amplitudes of component pulse waves of detected volume pulse waves of a limb segment detected during a baseline period to determine a baseline arterial volume of the limb segment;applying a stimulus to the mammal to induce a period of change in the arterial volume of the limb segment;determining amplitudes of component pulse waves of detected volume pulse waves of the limb segment detected during a time period after the stimulus has been applied to the mammal to induce a period of change in the arterial volume of the limb segment; anddetermining relative change in arterial volume of the limb segment during the time period after the stimulus relative to the arterial volume of the limb segment during the baseline period from the amplitudes of the component pulse waves of the detected volume pulse waves during the baseline and after the stimulus;wherein determining relative change in arterial volume of the limb comprises: calculating an average component pulse wave amplitude during the baseline period;generating a curve fit of component pulse wave amplitudes during the after-stimulus period;calculating a maximum of the fitted curve;determining a time between beginning of the after-stimulus period and the time of the maximum of the fitted curve;determining a relative amplitude change from the baseline to the maximum of the fitted curve; anddetermining a relative change in arterial volume from the relative amplitude change from the baseline to the maximum of the fitted curve and the average component pulse wave amplitude during the baseline period. 2. The method of claim 1, wherein the component pulse wave is an early systolic component. 3. The method of claim 2 wherein the determining relative change in arterial volume includes comparing maximum amplitudes of the early systolic components of the volume pulse waves during the baseline period and maximum amplitudes of the early systolic components of the volume pulse waves after the stimulus. 4. The method of claim 2, wherein determining amplitude of early systolic component of a volume pulse wave comprises: determining an occurrence of an inflection point in the volume pulse wave;determining a time of the occurrence of the inflection point;determining a pressure value on the detected volume pulse wave at said time;determining a pressure value at the beginning of the volume pulse wave; andassessing the amplitude of the early systolic component of the volume pulse wave as the difference between said pressures values. 5. The method of claim 1, further comprising: monitoring the limb segment to record the detected volume pulse waves of the limb segment during the baseline period;monitoring the limb segment to record the detected volume pulse waves of the limb segment during a period of an after-stimulus period. 6. The method of claim 5 wherein the component pulse wave is an early sys... | Angiologix, Inc. | A61B005/02
US09952299 | B2 | 2018-04-24 | Method and apparatus to determine a magnetic resonance relaxation time in the heart muscle in a magnetic resonance examination | 2010-05-14 | In a method and magnetic resonance (MR) system to determine an MR relaxation time (for example a T1, T2 or T2* relaxation time) in the heart muscle in a magnetic resonance examination, a determination an annular slice image region of the heart muscle of the left heart chamber in MR image data with the use of an automatic image segmentation. Multiple sub-regions within the slice image region are automatically determined. Each sub-region respectively includes multiple pixels of the annular slice image region of the heart muscle of the left heart chamber. An MR relaxation time is determined automatically for each of the multiple sub-regions and associated with the corresponding sub-region. A characteristic MR relaxation time in the heart muscle is determined by a statistical analysis of the multiple MR relaxation times that are associated with the multiple sub-regions. | 1. A method to determine a magnetic resonance (MR) relaxation time in a heart muscle in a magnetic resonance examination, comprising the steps of: using a computer to operate an MR data acquisition apparatus in order to acquire MR data from an examination subject, while the examination subject is situated in the MR data acquisition apparatus, from heart muscle tissue of the left heart chamber of the examination subject;in said computer, reconstructing an image from said MR data, said image being comprised of pixels that each exhibit a signal intensity, said image depicting anatomy of the heart muscle tissue at the left heart chamber of the examination subject;in said computer, (a) implementing an automatic image segmentation of said anatomy in said image to determine an annular slice image region, comprised of pixels of said image of the heart muscle tissue of the left heart chamber so that said annular slice region has a width in a radial direction that is reduced by an inner periphery of the annular slice image region being increased by a predetermined value and an outer periphery of the annular slice image region being reduced by a predetermined value;in said computer, (b) automatically dividing said annular slice image region of said image of the heart muscle tissue of the left heart chamber into a plurality of sub-regions of said annular slice image region of said image, each sub-region comprising a sub-plurality of said plurality of said pixels of said annular slice image region of the heart muscle tissue of the left heart chamber, each of said pixels in each sub-region exhibiting said signal intensity;for each of said sub-regions, (c) automatically determining, in said computer, an MR relaxation time of the heart muscle tissue depicted in the respective sub-region from the respective signal intensities of the respective pixels thereof, thereby obtaining a plurality of MR relaxation times; andin said computer, (d) automatically quantifying a characteristic MR relaxation time, of all the heart muscle, by statistical analysis of said plurality of MR relaxation times of the heart muscle tissue depicted in the respective sub-regions, and making a representation of the quantified characteristic MR relaxation time of all the heart muscle tissue available in electronic form at an output of the computer. 2. A method as claimed in claim 1 comprising employing, as said MR relaxation times of the respective sub-regions and said MR relaxation time of the heart muscle, an MR relaxation time selected from the group consisting of a T2* relaxation time, a T1 relaxation time and a T2 relaxation time. 3. A method as claimed in claim 1 wherein the step of providing a computerized processor with an image comprises providing said computerized processor with comprising using said computer to operate said MR data acquisition apparatus in order to acquire multiple images respectively at different times, selected from the group consisting of different echo times ... | Siemens Aktiengesellschaft | A61B005/05 G01R033/50 A61B005/055 G01R033/24 G01R033/56 G01R033/561
US20100324439 | A1 | 2010-12-23 | HIGH FREQUENCY AIRWAY OSCILLATION FOR EXHALED AIR DIAGNOSTICS | 2010-06-09 | The present invention relates to noninvasive methods for obtaining exhaled breath condensate (EBC) samples from the airways and lungs of a subject for use in diagnosing various conditions and diseases associated with biomarkers present in EBC, including diagnosis of lung cancer. The methods of the subject invention include generating and/or applying an oscillating airflow to a subject during inhalation and/or exhalation to induce increased concentration of biomarkers in EBC; collecting EBC from exhaled breath; and analyzing the collected EBC for biomarkers associated with disorders and/or diseases. | 1. A method for increasing the concentration of exhaled breath condensates (EBC) in exhaled breath comprising: supplying an external means for generating and/or maintaining an oscillating airflow to a subject; collecting at least one exhaled breath sample from the subject following application of the oscillating airflow; and collecting an EBC sample from the exhaled breath sample. 2. The method of claim 1, further comprising the step of identifying and/or measuring the concentration of at least one biomarker present in the EBC sample. 3. The method of claim 2, further comprising the step of identifying any diseases or conditions associated with the biomarker(s) present in the EBC sample. 4. The method of claim 3, wherein the biomarker is one associated with lung cancer. 5. A system for increasing the concentration of EBC in exhaled breath comprising: an external oscillating means for application to a breathing circuit to a subject; and a collection means for collecting at least one exhaled breath and EBC sample from the subject. 6. The system of claim 5, further comprising a sensor for detecting any biomarker present in the EBC sample. 7. The system of claim 6, wherein the biomarker is one associated with lung cancer. | - | A61B005/08
US20100292564 | A1 | 2010-11-18 | System and Method For Magnetic-Nanoparticle, Hyperthermia Cancer Therapy | 2010-05-18 | A system and method is provided for performing and monitoring a magnetic nanoparticle hyperthermia process on a subject having received a dose of a magnetic nanoparticle configured to bind to tissue within a target area of the subject. A low-field MRI system is utilized having a static magnetic field and a radio frequency (RF) system configured to receive MRI data from the target area. A hyperthermia system is coupled to the low-field MRI system and configured to generate a hyperthermia excitation field configured to cause the magnetic nanoparticle to rotate at a lag with respect to a magnetic field experienced by the magnetic nanoparticle to cause a temperature increase in the target area. The MRI system is configured to acquire medical imaging data from the target area and reconstruct images therefrom indicating at least one of a spatial distribution of the nanoparticle in and a temperature of the target area. | 1. A system for performing and monitoring a magnetic nanoparticle hyperthermia process on a subject having received a dose of a magnetic nanoparticle configured to bind to tissue within a target area of the subject, the system comprising: a low-field magnetic resonance imaging (MRI) system having a static magnetic field configured to receive a subject arranged therein and a radio frequency (RF) system configured to receive MRI data from a target area in the subject;a hyperthermia system coupled to the low-field MRI system and configured to generate a hyperthermia excitation field configured to cause a magnetic nanoparticle administered to the subject to rotate at a lag with respect to a magnetic field experienced by the magnetic nanoparticle to cause a temperature increase in the target area; andwherein the MRI system is configured to acquire medical imaging data from the target area and reconstruct images therefrom indicating at least one of a spatial distribution of the nanoparticle in and a temperature of the target area. 2. The system of claim 1 wherein the hyperthermia system is configured to generate the hyperthermia excitation field to be substantially orthogonal to the static magnetic field of the MRI system. 3. The system of claim 1 wherein the hyperthermia system is configured to generate the hyperthermia excitation field that is at least one of a rotating and a sinusoidal magnetic field applied substantially perpendicular to the static magnetic field. 4. The system of claim 1 wherein the hyperthermia system is configured to generate the hyperthermia excitation field such that the hyperthermia excitation field and a magnetic moment of the magnetic nanoparticle are maintained substantially in a non-parallel alignment. 5. The system of claim 1 wherein the magnetic nanoparticle is configured to operate as a contrast agent during the acquisition of the medical imaging data. 6. The system of claim 1 wherein the hyperthermia system includes a tuner configured to tune at least one of a frequency and an amplitude of the hyperthermia excitation field to substantially match the a time constant of the magnetic nanoparticle. 7. The system of claim 1 wherein the hyperthermia system is further configured to generate another hyperthermia excitation field to create an alternating magnetic field substantially transverse to the static magnetic field of the MRI system. 8. The system of claim 7 wherein the hyperthermia excitation field and the another hyperthermia excitation field are displaced by a quarter cycle to create a rotating hyperthermia excitation field. 9. The system of claim 1 wherein the hyperthermia system is configured to utilize the RF system of the MRI system to generate the hyperthermia excitation field. 10. The system of claim 1 wherein the hyperthermia system includes an RF heating coil system configured to be driven in conjunction with a transmit and receive system of the MRI system by the RF system of the MRI system. 11. A m... | - | A61B005/055
US20100305420 | A1 | 2010-12-02 | FLEXIBLE CIRCUIT AND METHOD FOR FORMING THE SAME | 2010-05-28 | A flexible circuit is provided herein that includes conductive material on the top and bottom planar surfaces of a dielectric substrate. The flexible circuit can be used in various applications, including use as a sensor. A via is used to provide electrical communication between the top and bottom surface of the flexible circuit. A method of preparing a flexible circuit and a medical instrument including the flexible circuit are also provided. | 1. A sensor including a flexible circuit, comprising: a flexible dielectric substrate having opposing first and second planar surfaces defining longitudinal, transverse and normal directions;one or more conductive contacts adjacent the first planar surface of said flexible dielectric substrate;one or more conductive contacts adjacent the second planar surface of said flexible dielectric substrate;a first dielectric mask adjacent the first planar surface and substantially covering the first planar surface, said first dielectric mask having one or more mask openings corresponding to one of more of the conductive contacts adjacent the first planar surface;a second dielectric mask adjacent said second planar surface substantially covering said second planar surface;at least one conductive material provided within the mask openings of said first dielectric mask and in electrical communication with the one or more conductive contacts adjacent the first planar surface;a via extending through said dielectric substrate and providing electrical communication between a contact adjacent the first planar surface and a contact adjacent the second planar surface to provide electrical communication between the conductive material within one of the mask openings and the contact adjacent the second planar surface;wires in electrical communication with the contacts adjacent the first planar surface and the contacts adjacent the second planar surface; andone or more membrane layers applied in physical contact with at least a portion of the conductive material, said one or more membrane layers performing a chemical transduction that is communicated to the conductive material. 2. The sensor according to claim 1, wherein the second dielectric mask includes one or more mask openings corresponding to one or more conductive contacts adjacent the second planar surface and further comprising a conductive material applied to the second dielectric mask adjacent the mask openings in the second dielectric mask such that the conductive material is in electrical connection with the one or more conductive contacts adjacent the second planar surface. 3. The sensor according to claim 2, wherein the conductive material applied to the second dielectric mask adjacent the mask openings in the second dielectric mask is in electrical communication with two conductive contacts adjacent the second planar surface and forms a thermistor with the two conductive contacts. 4. The sensor according to claim 1, further comprising a third dielectric mask adjacent the first dielectric mask and substantially covering the first dielectric mask, said third dielectric mask having one or more mask openings corresponding to one of more of the conductive contacts adjacent the first planar surface, at least a portion of said one or more membrane layers provided within the mask openings in the third dielectric mask. 5. The sensor according to claim 1, wherein at least one contact and at least one membrane... | Edwards Lifesciences Corporation | A61B005/1473 H01L029/66 H01L021/02
US20100286546 | A1 | 2010-11-11 | GARMENT FOR DETECTING RESPIRATORY MOVEMENT | 2008-11-04 | A garment for detecting respiratory movement of a living being, wherein the garment can be pulled over the thorax of the living being, including an electric conductor integrable at the height of the thorax of the living being into the garment, which is attached to the garment in order to change an electrically measurable characteristic in dependence on the respiratory movement of the thorax, and a holder for fixing evaluation electronics integrable into the garment, which can be coupled to the electric conductor and which is implemented to detect the electrically measurable characteristic and which further includes an interface for outputting or storing data derived from the electrically measurable characteristic. | 1-13. (canceled) 14. Garment for detecting respiratory movement of a living being, wherein the garment can be pulled over the thorax of the living being, comprising: a straight electric conductor, which is integrable into the garment at the height of the thorax of the living being, which is attached to the garment in order to change an inductance of the electric conductor in dependence on the respiratory movement of the thorax, wherein the electric conductor is an elastic electrically conductive fiber, which can expand in dependence on the respiratory movement and thereby experiences a change of the inductance when the garment is pulled over the thorax of the living being; anda holder for fixing evaluation electronics, which is integrable into the garment; andthe integrable evaluation electronics, which can be coupled to the elastic electrically conductive fiber and is implemented to detect a change in inductance of the elastic electrically conductive fiber and further comprises an interface for outputting or storing the data derived from the inductance. 15. Garment according to claim 14, wherein the integrable evaluation electronics is implemented such that the inductance of the elastic electrically conductive fiber variable with the respiratory movement forms an inductance of an LC parallel resonant circuit, which is part of a Colpitts oscillator circuit, wherein the Colpitts oscillator circuit is within the integrable evaluation electronics. 16. Garment according to claim 15, wherein the integrable evaluation electronics comprises a frequency counter for detecting a frequency of the Colpitts oscillator circuit varying due to the respiratory movement. 17. Garment according to claim 14, wherein the electric conductor is integrated into the garment such that an area surrounded by the elastic electrically conductive fiber is dependent on the respiratory movement when the garment is pulled over the thorax of the living being. 18. Garment according to claim 14, wherein the evaluation electronics is fixed in the holder such that the evaluation electronics is part of the garment, wherein the evaluation electronics is implemented in a watertight manner. 19. Garment according to claim 14, wherein the integrable evaluation electronics comprises a wireless interface for radio communication. 20. Garment according to claim 14, wherein a fabric of the garment is elastic for allowing tight application of the electric conductor to the torso of the living being. 21. Garment according to claim 14, wherein the garment can additionally be pulled over the abdomen of the living being and comprises a further electric conductor at the height of the abdomen of the living being, which is attached to the garment in order to change an electrically measurable characteristic in dependence on the respiratory movement of the abdomen and wherein the further electric conductor can be coupled to the integrable evaluation electronics. 22. Usage of a garment for detecting... | Fraunhofer-Gesellschaft zur Foerderung der angewandten Forschung e.V. | A61B005/08
US20100298685 | A1 | 2010-11-25 | ADAPTIVE INSULIN DELIVERY SYSTEM | 2010-05-21 | A proactive system and method in which levels of glucose are monitored after a meal signal and compared to a safe range. If a monitored glucose level is outside the safe range, a post-prandial vertex of the glucose level is identified and an action is provided to more rapidly return the glucose level to a target level within the safe range than if no action was provided. In another aspect a control parameter in an IDM system is adjusted by determining a performance metric of the system as a function of the levels of glucose and a medication administration signal over a first window of time; and, if the performance metric is outside an expected range, adjusting the control parameter to adjust an amount of medication and to bring the performance metric inside the expected range. | 1. A system for proactively monitoring glucose levels, the system comprising: a sensor that measures an indication of glucose and provides levels of glucose over a monitoring time period;a memory having stored therein a safe range of glucose;a device that provides a meal signal indicating that a meal has been consumed, the device also providing a medication administration signal indicating an amount of medication and a time that it was administered;a processor configured to receive the meal signal, levels of glucose, and the safe range, wherein upon receiving the meal signal, the processor is further configured to: monitor the levels of glucose beginning after the meal signal;compare the levels of glucose to the safe range;if a monitored glucose level is outside the safe range, determine a post-prandial vertex of the glucose level, andonce the vertex is determined, provide an action to return the glucose level to a target glucose level within the safe range, wherein the action includes consideration of a user parameter. 2. The system for proactively monitoring glucose levels of claim 1, wherein the vertex is a peak and the action provided by the processor includes calculating a dose of medication sufficient to reduce the glucose level more rapidly to the target glucose level within the safe range and to decrease an amount of high glucose exposure than if the action was not taken. 3. The system for proactively monitoring glucose levels of claim 1, wherein the vertex is a valley and the action provided by the processor includes calculating an amount of carbohydrates sufficient to increase the glucose level more rapidly to the target glucose level within the safe range than if the action was not taken. 4. The system for proactively monitoring glucose levels of claim 1, wherein the vertex is a valley and the action by the processor includes calculating a reduction in a basal rate of a medication administration sufficient to raise the glucose level more rapidly to the target glucose level within the safe range than if the action was not taken. 5. The system for proactively monitoring glucose levels of claim 1, wherein the processor is configured to delay monitoring the levels of glucose for a selected time period after receiving the meal signal. 6. The system for proactively monitoring glucose levels of claim 1, wherein the the processor is configured to determine the vertex by comparing a first and a second glucose level trend, the vertex being at the point at which the first and second glucose level trends diverge. 7. The system for proactively monitoring glucose levels of claim 1, wherein the processor is configured to determine the vertex by comparing a first and a second glucose rate of change, the vertex being at the point at which the signs of the first and second glucose rate of change diverge. 8. The system for proactively monitoring glucose levels of claim 1, wherein the processor is further configured to: identify a first and second ... | ABBOTT DIABETES CARE INC. | A61B005/145 A61M005/00
US20100292589 | A1 | 2010-11-18 | HYPOTHENAR SENSOR | 2009-05-13 | A device having an arterial pulse sensor that is adhered to the hypothenar region of a palm using an adhesive patch. The patch has an adhesive surface that is covered by a removable film with an outer portion and a central portion. Also disclosed is a method of detecting an arterial pulse by providing an arterial pulse sensor, placing the sensor on the hypothenar region of the palm of a hand, and receiving an arterial pulse signal from the sensor. | 1. A device for detecting arterial pulse signals from an artery comprising: an arterial pulse sensor; andattaching means for holding the sensor on the hypothenar region of the palm of a hand. 2. The device of claim 1, further comprising a data processing unit for receiving arterial pulse signals from the pulse sensor wherein the data processing unit is in signal communication with the pulse sensor. 3. The device of claim 1, wherein the signal communication is wired or wireless. 4. The device of claim 1, wherein the arterial pulse sensor is a photoplethysmographic sensor. 5. The device of claim 1, further comprising a motion sensor. 6. The device of claim 5, wherein the motion sensor is selected from the group consisting of an accelerometer and a piezo sensor. 7. The device of claim 5, further comprising an adaptive filter. 8. The device of claim 7, wherein the motion sensor is a three-axis accelerometer and the adaptive filter is adapted to obtain reference signals from the three-axis accelerometer. 9. The sensor of claim 7, wherein the adaptive filter is selected from the group consisting of LMS, RLS and Kainan. 10. The device of claim 1, wherein the attaching means is selected from the group consisting of an adhesive, a glove, and a spring biased mechanism. 11. The device of claim 1, in combination with an ECG sensor whereby arterial pulse wave velocity can be aquired. 12. An adhesive patch for adhering a sensor to a hypothenar area of a palm, the adhesive patch comprising: a first adhesive region for adhering the sensor to the patch; anda second adhesive region for adhering the patch to the palm. 13. The patch of claim 12, further including: a first removable film for covering the first adhesive surface; anda second removable film for covering the second adhesive surface. 14. The adhesive patch of claim 12, wherein the first adhesive region defines an area at or near the centre of the adhesive patch. 15. The patch according to claim 12, wherein the first adhesive region is sized and shaped to receive the sensor. 16. The patch according to claim 12, further including an opening in the patch for receiving a cable. 17. The patch of claim 12, further comprising an adhesive pad fixed to the adhesive surface between the adhesive surface and the central portion of the film. 18. A method of detecting an arterial pulse comprising the steps of: providing an arterial pulse sensor;placing the sensor on the hypothenar region of the palm of a hand; andreceiving an arterial pulse signal from the sensor. 19. The method of claim 18, further comprising the step of providing a data processing unit for receiving and processing the arterial pulse signal. 20. The method of claim 18, further comprising the step of providing a display device and displaying the processed signal on the device. | - | A61B005/02
US07860551 | B2 | 2010-12-28 | Magnetic resonance diagnostic apparatus | 2005-08-29 | A magnetic resonance diagnostic apparatus includes an input unit which inputs information associated with the physique of a subject to be examined and information associated with an imaging region, a calculating unit which calculates a partial body SAR associated with a partial body including the imaging region on the basis of the input information associated with the physique and the input information associated with the imaging region, and a display unit which displays the calculated partial body SAR. | 1. A magnetic resonance diagnostic apparatus comprising: an input unit configured to input information associated with a physique of a subject to be examined and information associated with an imaging region;a calculating unit configured to calculate a partial body SAR associated with a partial body including the imaging region on the basis of the input information associated with the physique and the input information associated with the imaging region; anda display unit configured to display the calculated partial body SAR,wherein, when imaging is divisionally performed a plurality of times for a wide area, the calculating unit which calculates the partial body SAR selects a maximum value of a plurality of partial body SARs associated with a plurality of body portions. 2. An apparatus according to claim 1, wherein the information associated with the physique includes a weight of the subject. 3. An apparatus according to claim 2, wherein the information associated with the physique includes a height of the subject. 4. An apparatus according to claim 3, wherein the weight and height of the subject are input as numerical values. 5. An apparatus according to claim 3, wherein the weight and height of the subject are selected from a plurality of height ranges. 6. An apparatus according to claim 1, wherein the calculating unit includes an estimating unit which estimates a partial weight of a partial body including the imaging region on the basis of the input information associated with the physique and the input information associated with the imaging region. 7. An apparatus according to claim 6, wherein the estimating unit which estimates the partial weight includes a storage unit which stores a table in which a standard partial weight is made to correspond to the physique and imaging region. 8. An apparatus according to claim 6, wherein the calculating unit includes a determining unit which determines a partial body SAR upper limit value on the basis of the estimated partial weight and the input information associated with the physique. 9. An apparatus according to claim 8, wherein the calculating unit includes a comparing unit which compares the calculated partial body SAR with the determined partial body SAR upper limit value. 10. An apparatus according to claim 9, wherein the display unit displays a message corresponding to a result of comparison between the calculated partial body SAR and the determined partial body SAR upper limit value. 11. An apparatus according to claim 10, wherein when the calculated partial body SAR exceeds the upper limit value, the display unit displays a message indicating that the calculated partial body SAR exceeds the upper limit value, together with information indicating a recommended change in imaging condition. 12. An apparatus according to claim 11, wherein the information indicating the recommended change includes a decrease in the number of multi-slices based on the pulse sequence. 13. An apparatus... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | A61B005/055 G01V003/00
US07824343 | B2 | 2010-11-02 | Method and apparatus for blood sampling | 2004-10-01 | Methods and apparatus for collecting blood samples in vacuum sample tubes are disclosed. The samples, including the initial blood sample, are substantially free of excess air. | 1. A biological fluid sampling system comprising: a flexible plastic container comprised of two flat sheets melt sealed together substantially along their peripheries and defining an interior chamber;a sample access site external to and spaced from said container;an internal flow path providing the only path for biological fluid flow into and from said interior chamber, said internal flow path communicating with said access site;said internal flow path extending substantially into and substantially across the entire distance of the interior chamber to allow for substantially complete drainage of the container. 2. The sampling system of claim 1 further comprising an intermediate tube defining a flow path with one end of said tube directly communicating with said sample access site and the other end directly communicating with said internal flow path. 3. The sampling system of claim 1 wherein said interior chamber has a generally circular profile. 4. The sampling system of claim 1 wherein said interior chamber includes a fluid receiving well. 5. The sampling system of claim 4 wherein the distal end of said internal flow path extends substantially into said interior chamber, said internal flow path having a distal open end terminating within said well. 6. The sampling system of claim 1 wherein, at least one of said flat sheets is generally translucent and includes indicia thereon for indicating the volume of fluid within said chamber. 7. The sampling system of claim 1 wherein said internal flow path is defined by a plastic tube that extends substantially into said chamber interior. 8. The sampling system of claim 1 wherein said internal flow path is defined by an interior seal in said plastic bag. 9. The sampling system of claim 1 wherein said access site is adapted to receive a holder for a vacuum tube. 10. The sampling system of claim 9 further comprising a holder for receiving a vacuum tube. | Fenwal, Inc. | A61B005/00 A61B019/00 B65D081/00 A61M037/00
US20100324401 | A1 | 2010-12-23 | DIABETES MANAGEMENT METHODS AND SYSTEMS | 2010-08-10 | The present invention relates to methods and systems for monitoring the effectiveness of diabetes treatment. Methods and systems in accordance with the present invention provide information relating to variability of glucose levels and hypoglycemia and hyperglycemia. Such information is based on time-stamped blood glucose data obtained from a meter or the like and actual measurements of HbA1c levels are not required. | 1. A method of managing blood glucose data, the method comprising the steps of: providing a plurality of blood glucose data collected over a predetermined period of time;estimating an idiosyncratic HbA1c level from the plurality of blood glucose data using a first predetermined algorithm;calculating an idiosyncratic variability index of glycemia from the plurality of blood glucose data using a second predetermined algorithm; andcomparing the idiosyncratic variability index of glycemia and the idiosyncratic HbA1c level to at least one exclusion criteria. 2. The method of claim 1, wherein the at least one exclusion criterion comprises at least one of a history of testing, a frequency of tests per day, and a randomness of data on a daily basis. 3. The method of claim 1, comprising reporting the idiosyncratic variability index of glycemia and the idiosyncratic HbA1c level when the at least one exclusion criteria is not met. 4. The method of claim 3, wherein the reporting step comprises reporting the idiosyncratic variability index of glycemia as a function of the idiosyncratic HbA1c level. 5. The method of claim 4, wherein the reporting step comprises providing the idiosyncratic variability index of glycemia as a function of the idiosyncratic HbA1c level on a weekly basis. 6. The method of claim 4, wherein the reporting step comprises providing the idiosyncratic variability index of glycemia as a function of the idiosyncratic HbA1c level on a monthly basis. 7. The method of claim 3, wherein the reporting step comprises providing the idiosyncratic variability index of glycemia and the idiosyncratic HbA1c level in tabular form. 8. A method of managing blood glucose data, the method comprising the steps of: providing a plurality of blood glucose data collected over a predetermined period of time;estimating an idiosyncratic HbA1c level from the plurality of blood glucose data using a first predetermined algorithm;calculating an idiosyncratic variability index of glycemia from the plurality of blood glucose data using a second predetermined algorithm;determining the idiosyncratic variability index of glycemia as a function of the idiosyncratic HbA1c level when at least one exclusion criterion is not met; andproviding the reported idiosyncratic variability index of glycemia and idiosyncratic HbA1c level while indicating at least one aspect of glycemic control. 9. The method of claim 8, wherein the at least one aspect of glycemic control comprises target, cautionary, and danger regions of glycemic control. 10. The method of claim 8, wherein the at least one exclusion criterion comprises at least one of a history of testing, a frequency of tests per day, and a randomness of data on a daily basis. 11. The method of claim 8, wherein the determining step comprises determining the idiosyncratic variability index of glycemia as a function of the idiosyncratic HbA1c level on a weekly basis. 12. The method of claim 8, wherein the determining step compris... | LifeScan, Inc. | A61B005/145
US07826888 | B2 | 2010-11-02 | Dynamic radiographic image obtaining method and apparatus | 2005-03-28 | A method and apparatus for obtaining a dynamic radiographic image with enhanced resolution, in which an electrocardiogram of a subject is obtained, and five time phases are allocated to each cardiac beat period (1 second). The initial X-ray irradiation is performed at 1st time phase of 1st beat to obtain a radiographic image. Thereafter, X-ray irradiations are performed sequentially at 2nd time phase of 2nd beat, at 3rd time phase of 3rd beat, at 4th time phase of 4th beat, and at 5th time phase of 5th beat to obtain respective radiographic images. The five radiographic images obtained are combined at the time interval of 0.2 seconds to be displayed as a motion image. The time interval between the irradiations of the X-ray is 1.2 seconds, so that image signals corresponding to 1st to 5th time phases may be obtained with high resolution. | 1. A dynamic radiographic image obtaining apparatus for obtaining a dynamic radiographic image of a periodically moving subject, the apparatus: a dynamic signal detecting means for detecting a dynamic signal of the subject;an image obtaining means for obtaining a plurality of radiographic images of the subject by irradiating radiation on the subject from one direction and detecting the radiation passed through the subject;a control means for allocating a plurality of time phases to each of predetermined periods of the dynamic signal detected by the dynamic signal detecting means, and controlling the image obtaining means such that the radiographic images of the subject are obtained at respective time phases that differ from one predetermined period to another;a storing means for storing the radiographic images obtained with the time phase information correlated therewith; anda dynamic image generating means for generating the dynamic radiographic image of the subject by arranging the radiographic images obtained in the order of the time phases, each being stored and correlated with each corresponding radiographic image among the radiographic images, from a radiographic image obtained at a time phase closest to a starting point of the each of the predetermined periods of the dynamic signal further comprising an interpolated image generating means for generating an image through interpolation from the radiographic images already obtained, which corresponds to a time phase that differs from each of the time phases at which the radiographic images are obtained. 2. The dynamic radiographic image obtaining apparatus according to claim 1, wherein the apparatus further comprises a timing setting means for setting timing for obtaining the dynamic radiographic image on a periodical basis based on the dynamic signal, andthe control means is configured such that the dynamic radiographic image is obtained according to the timing set by the timing setting means. 3. The dynamic radiographic image obtaining apparatus according to claim 2, wherein the apparatus further comprises an unsteady state detecting means for detecting an unsteady state of the dynamic signal, andthe control means is configured such that the dynamic raiodgraphic image is obtained automatically when an unsteady state of the dynamic signal is detected by the unsteady state detecting means. 4. The dynamic radiographic image obtaining apparatus according to claim 3, wherein the subject is a region of a living organism which moves periodically by cardiac beats or respiration. 5. The dynamic radiographic image obtaining apparatus according to claim 2, wherein the subject is a region of a living organism which moves periodically by cardiac beats or respiration. 6. The dynamic radiographic image obtaining apparatus according to claim 1, wherein the apparatus further comprises an unsteady state detecting means for detecting an unsteady state of the dynamic signal, andthe control means is config... | FUJIFILM Corporation | A61B005/05 A61B006/00
US07848798 | B2 | 2010-12-07 | Live finger detection by four-point measurement of complex impedance | 2003-12-03 | Method and sensor assembly determining the condition of a structure, especially for confirming if a measured fingerprint is on a live finger, by measuring characteristics of close to the structure surface, the sensor comprising a first pair of current supply electrodes coupled to a current source, providing an electrical current to the skin, at least two pickup electrodes at chosen and different positions relative to the current supply electrodes, at least a first of said pickup electrodes being coupled to an instrument for measuring the voltage between said first pickup electrode and at least one of the pickup or current supply electrodes, storage means for a predetermined set of values characterizing a certain condition of the surface, and means for comparing the characteristics from each pickup electrode with the measurements of the other pickup electrodes and with the predetermined set of characteristics for determining the surface condition. | 1. Sensor assembly for determining the condition of a structure, especially for confirming if a measured fingerprint is on a live finger, by measuring characteristics of close to the structure surface, the sensor comprising: a current source;at least four electrodes at chosen positions relative to each other, said positions providing at least two relative distances between the electrodes, wherein a chosen first pair of said at least four electrodes constitutes current supply electrodes, and a chosen second pair of said at least four electrodes constitutes pickup electrodes, and of which at least one does not constitute a current supply electrode;a measuring instrument coupled to said at least four electrodes for measuring the impedance between said chosen pair of pickup electrodes for providing a value characterizing the structure;wherein said sensor is adapted to store a predetermined set of values characterizing a chosen condition for said structure;wherein said sensor is adapted to compare said characteristics from each of said at least one pair of pickup electrodes with said set of predetermined values for detecting if said structure is in a certain condition, and the sensor assembly is adapted to alternating coupling of at least one current supply and measuring instrument to different electrode pairs with different distances between them, for measuring characteristic values at different depths in said structure; andwherein the supplied current is oscillating within a chosen frequency range and said sensor is adapted to measure the impedance at each pickup electrode and to compare the imaginary and real parts of the impedance signals as functions of the applied frequency, by determining the slope of the resulting curve, and to compare this slope with a predetermined set of slopes indicating a live finger. 2. Sensor assembly for determining the condition of a structure, especially for confirming if a measured fingerprint is on a live finger, by measuring characteristics of close to the structure surface, the sensor comprising: a current source;at least four electrodes at chosen positions relative to each other, said positions providing at least two relative distances between the electrodes, wherein a chosen first pair of said at least four electrodes constitutes current supply electrodes, and a chosen second pair of said at least four electrodes constitutes pickup electrodes, and of which at least one does not constitute a current supply electrode;a measuring instrument coupled to said at least four electrodes for measuring the impedance between said chosen pair of pickup electrodes for providing a value characterizing the structure;wherein said sensor is adapted to store a predetermined set of values characterizing a chosen condition for said structure;wherein said sensor is adapted to compare said characteristics from each of said at least one pair of pickup electrodes with said set of predetermined values for detecting if said structure is... | Idex ASA | A61B005/05 G06T001/00 G06T007/00
US07833157 | B2 | 2010-11-16 | Multilumen catheter | 2008-05-15 | A multilumen catheter having tubings extending into lumens within the catheter. The lumens may be used for blood, drugs or other medicants. The lumens may also be used for sensors. The junction element, external to the patient, connects the tubings to the lumens. The tubings, also external to the patient, connect to infusion members, to which one or more infusion systems may be connected to deliver blood, drugs and other medicants to the patient. A sensor having a sensing element may extend through the sensor lumen and be positioned internal to the patient for physiological parameter sensing. An external portion of the sensor may be connected to associated electronics to provide automatic monitoring of the physiological parameters and automatic delivery and control of the infusants. Also, a central line catheter for delivering fluids directly into a main artery or vein near the heart, which contains a first lumen to deliver a fluid through the central line catheter and a second lumen containing a sensor capable indicating a characteristic level in blood. The first lumen delivers the fluid downstream of the sensor to prevent any interference between the fluid delivery and the sensor readings. In other versions, the central line catheter can have additional lumens for additional purposes. In addition, the central line catheter can further include a flush sleeve to remove debris around the sensor. | 1. A catheter for delivering fluids directly into a patient's body, the catheter comprising: a catheter body having a generally circular outer peripheral cross section;a first lumen in the catheter body; anda second lumen in the catheter body;wherein the first lumen is configured for communicating a fluid through the catheter, the first lumen having a site at which the fluid is able to enter or exit;the catheter further comprising at least a portion of a sensor lead extending in the second lumen, the sensor lead having a sensor capable of indicating a characteristic level in the patient's body, the second lumen having a size relative to the sensor lead that squeezes the sensor within the second lumen, the second lumen having an end that is adapted to be located in the patient's body when the first lumen is communicating the fluid to the patient's body, such that the sensor is able to extend past the site of the first lumen;wherein the second lumen has a generally oval shape; andwherein the first lumen has a cross sectional shape that has one edge facing the second lumen and having a curvature corresponding to that of the generally oval shaped second lumen, and the first lumen has a second edge facing the catheter body and having a curvature corresponding to that of the generally circular shaped catheter body. 2. The catheter according to claim 1, further comprising at least one additional lumen in the catheter body, the at least one additional lumen having a cross sectional shape that has one edge facing the second lumen and a curvature corresponding to that of the generally oval shaped second lumen, the at least one additional lumen further having a second edge facing the catheter body and a curvature corresponding to that of the generally circular shaped catheter body. 3. The catheter according to claim 1, further comprising at least one additional lumen. 4. The catheter according to claim 1, wherein the catheter comprises a central line catheter further comprising a flush sleeve configured to remove debris around the sensor. 5. The catheter according to claim 1, wherein the sensor of the sensor lead is able to extend up to several inches past the site of the first lumen and the end of the second lumen. 6. The catheter according to claim 1, wherein the sensor is an analyte sensor. 7. The catheter according to claim 1, wherein the sensor is a glucose sensor. 8. The catheter of claim 1, the sensor extends past the site of the first lumen. 9. The catheter of claim 1, wherein the shape of the second lumen is configured to provide a compression fit for the sensor lead in at least one dimension. 10. The catheter according to claim 8, further comprising: at least one additional lumen;wherein the at least one additional lumen has a cross sectional shape that has one edge facing the second lumen and a curvature corresponding to that of the generally oval shaped second lumen, and the at least one additional lumen has a second edge facing the cat... | Medtronic Minimed, Inc. | A61B005/05 A61B005/00
US07846107 | B2 | 2010-12-07 | Endoscopic apparatus with integrated multiple biopsy device | 2005-05-13 | An imaging endoscope comprising a shaft having a proximal end adapted to be secured to a handle, and a distal end having a biopsy forceps disposed therein. The biopsy forceps includes one or more end-effector elements that are actuated with a control cable that may be connected to the handle. The endoscope shaft includes a biopsy sample lumen that is configured to receive a biopsy sample obtained from the forceps assembly. A sample collection apparatus is attached to the handle to capture multiple biopsy samples. In some embodiments, the endoscope is a single-use endoscope. | 1. A multiple biopsy system comprising: an endoscope comprising: an elongated shaft;an imaging system at or adjacent a distal end of the elongated shaft;a biopsy forceps disposed within the elongated shaft and having a proximal end, a shaft, and a distal end, wherein the biopsy forceps includes a bite assembly that is movable to obtain a biopsy sample;a handle permanently fixed and stationary relative to the elongated shaft, the handle including one or more controls for operating the endoscope and including an actuator integrated therein for moving the bite assembly of the biopsy forceps;a biopsy lumen configured to receive the biopsy sample obtained with the biopsy forceps; anda sample vial cassette removably secured to the handle, the sample vial cassette having a plurality of sample chambers into which a sample vial is removably disposed, wherein the sample vial cassette is movable with respect to an opening in the handle so as to move a sample vial within the sample vial cassette to a position within the handle so as to be aligned with a proximal end of the biopsy lumen;wherein the biopsy lumen is coupleable to means for retrieving the biopsy sample through the biopsy lumen and into the sample vial of the sample vial cassette aligned with the proximal end of the biopsy lumen. 2. The system of claim 1, wherein multiple biopsy samples are individually received in a corresponding number of sample vials without removing the biopsy forceps from the elongated shaft. 3. The system of claim 1, wherein the means for retrieving the biopsy sample is a vacuum source applied to the proximal end of the biopsy lumen. 4. The system of claim 1, wherein the means for retrieving the biopsy sample is a liquid applied at a distal end of the biopsy lumen to flush the biopsy sample through the biopsy lumen. 5. The system of claim 1, wherein the means for retrieving the biopsy sample is a gas applied at a distal end of the biopsy lumen to blow the biopsy sample through the biopsy lumen. 6. The system of claim 1, wherein the shaft of the biopsy forceps includes a hollow lumen that forms the biopsy lumen. 7. The system of claim 1, wherein the bite assembly is actuated with a programmable firing mechanism. 8. The system of claim 1, wherein the bite assembly includes echogenic markings. 9. The system of claim 1, wherein the elongated shaft includes a tool lumen separate from the biopsy lumen, in which the biopsy forceps is placed. 10. The system of claim 1, wherein each sample chamber of the plurality of sample chambers includes a port for forming a connection with the biopsy lumen and a vacuum port configured to permit a vacuum to be selectively applied to the sample chamber. 11. The system of claim 1, wherein the system further comprises a position tracking system for recording a set of reference coordinates that indicate the position where the biopsy sample was obtained. 12. A kit for obtaining multiple biopsy samples, the kit comprising: i) a single-use e... | Boston Scientific Scimed, Inc. | A61B005/00
US07846095 | B2 | 2010-12-07 | Atraumatic sensor lead assemblies | 2004-08-17 | The invention is directed to lead configurations for sensors that allow for less invasive sensor replacement procedures. In one configuration, a sensor lead assembly includes an outer lead body and an inner lead including a sensor such as an electrochemical glucose sensor. The inner lead can be positioned in an inner conduit of the outer lead body. The outer lead body may be substantially permanently implanted in the patient, and the inner lead can be implanted through the inner conduit of the outer lead body. Once the sensor of the inner lead has worn out or otherwise exhausted its useful life, the inner lead can be removed, and a new inner lead can be implanted in place of the old inner lead. | 1. An implantable assembly, comprising: an implantable medical device housing including a channel within a connector module having an opening, and an electrical contact being housed within the channel proximal to the opening;an outer lead body including a proximal end, a distal end, and a conduit extending longitudinally from the proximal end to the distal end; andan inner lead including a first end, a second end, a plurality of sensors, and a plurality of electrical contacts; the plurality of sensors being positioned in proximity to the second end of the inner lead and electrically coupled to corresponding electrical contacts of the inner lead and the plurality of electrical contacts of the inner lead being positioned in proximity to the first end;wherein the opening of the channel is adapted for mechanical engagement of the proximal end of the outer lead body and for passage of the inner lead into the channel; anda lateral window opening formed in the outer lead body and positioned in proximity to the distal end of the lead body; wherein the plurality sensors of the inner lead includes a first sensor being longitudinally aligned with the window opening; andwherein the inner lead slideably engages the channel and the conduit of the outer lead body and at least a first one of the plurality of electrical contacts of the inner lead electrically couples to the electrical contact housed within the channel; andwherein the sensor is a glucose sensor. 2. The implantable assembly of claim 1, wherein the glucose sensor is an electrochemical sensor or an optical sensor. 3. An implantable assembly, comprising: an implantable medical device housing including a channel within a connector module having an opening, and an electrical contact being housed within the channel proximal to the opening;an outer lead body including a proximal end, a distal end, and a conduit extending longitudinally from the proximal end to the distal end; andan inner lead including a first end, a second end, a plurality of sensors, and a plurality of electrical contacts; the plurality of sensors being positioned in proximity to the second end of the inner lead and electrically coupled to corresponding electrical contacts of the inner lead and the plurality of electrical contacts of the inner lead being positioned in proximity to the first end;a lateral window opening formed in the outer lead body and positioned in proximity to the distal end of the outer lead body; anda rotatable structure within the medical device housing and coupled to the inner lead;wherein the opening of the channel is adapted for mechanical engagement of the proximal end of the outer lead body and for passage of the inner lead into the channel;wherein the plurality sensors of the inner lead includes a first sensor being longitudinally aligned with the window opening;wherein the inner lead slideably engages the channel and the conduit of the outer lead body and at least a first one of the plurality of electrical... | Medtronic, Inc. | A61B005/00 A61B005/05
US07841988 | B2 | 2010-11-30 | Nasal mask for carbon dioxide gas measurement | 2007-07-06 | A nasal mask is adapted to attached on a face of a subject so as to cover a nose of the subject. A shell has a contact edge adapted to be brought into close contact with the face of the subject, thereby forming a sealed space therein. An attachment portion is disposed in the shell and has an inner space. The attachment portion is configured to detachably engage with a sensor disposed outside the shell and operable to measure concentration of carbon dioxide gas expired from the nostrils. A guide member is disposed inside the shell and forming a passage adapted to introduce expired gas from the nostrils into the inner space. The shell is formed with an inlet adapted to introduce external gas into the sealed space. The expired gas is guarded by the guide member from the external gas introduced through the inlet. The attachment portion is formed with an outlet adapted to discharge the expired gas from the inner space to the sealed space. | 1. A nasal mask, adapted to be attached to a face of a subject so as to cover a nose of the subject, comprising: a shell, having a contact edge adapted to be brought into close contact with the face of the subject, thereby forming a sealed space therein;an attachment portion, disposed in the shell and having an inner space, the attachment portion being configured to detachably engage with a sensor disposed outside the shell and operable to measure concentration of carbon dioxide gas expired from the nostrils; anda guide member, disposed inside the shell and forming a passage adapted to introduce expired gas from nostrils into the inner space, wherein:the shell is formed with an inlet adapted to introduce external gas into the sealed space;the expired gas is guarded by the guide member from the external gas introduced through the inlet; andthe attachment portion is formed with an outlet adapted to discharge the expired gas from the inner space to the sealed space. 2. The nasal mask as set forth in claim 1, wherein: the guide member is a nasal tube member adapted to be placed within the nostrils so as to communicate the nostrils and the inner space. 3. The nasal mask as set forth in claim 1, wherein: the guide member is a cup member integrally provided with the attachment portion and adapted to cover the nose; andan inner space of the cup member is communicated with the inner space of the attachment portion. 4. The nasal mask as set forth in claim 1, wherein: the shell is formed with a recessed part adapted to receive the sensor, and the attachment potion is projected into the recessed part. 5. The nasal mask as set forth in claim 4, wherein: the attachment portion is monolithically formed with the shell. 6. The nasal mask as set forth in claim 4, wherein: the attachment portion is detachably fitted with an aperture formed in the recessed part. 7. The nasal mask as set forth in claim 1, wherein: the nasal mask is configured such that the expired gas travels first from the nostrils and then through the inner space and then to the sealed space. | Nihon Kohden Corporation | A61B005/08 A62B018/02 A62B007/00
US07828729 | B2 | 2010-11-09 | Method of measuring bleeding volume | 2008-03-10 | A method of measuring blood clotting ability of a patient. A known volume of blood is dispensed on a reference disc of blotter paper, such that at least one reference blotch is formed on the reference disc. The surface area of the reference blotch formed on the reference disc, such that a relationship between the volume of blood and the surface area of the reference blotch can be obtained. An incision is made at the patient. A piece of blotter paper to the incision is applied to the incision to collect blood flowing therefrom, such that at least one blotch is formed on the piece of blotter paper. The surface area of the blotch formed on the piece of blotter paper is measured. The volume of the blood absorbed blood by the blotter paper is calculated according to the relationship between the blood volume and the surface area of the blotch. | 1. A method of measuring blood clotting ability of a patient with a predetermined relationship between a reference surface area and a reference volume, the method comprising the steps of: a) breaking a skin of the patient to induce bleeding;b) applying blood of the patient obtained from step a) to a blotter paper, such that a blotch of the patient's blood is formed on the blotter paper;c) measuring the surface area of the blotch formed on the blotter paper in step b); andd) determining the blood volume in the blotter paper in accordance with the predetermined relationship between the reference surface area and the reference volume, the determined blood volume being indicative of the blood clotting ability of the patient. 2. The method of claim 1 wherein the breaking step a) comprises the step of making an incision on the patient to induce bleeding. 3. The method of claim 1 wherein the applying step b) comprises the step of applying the blotter paper to the incision to form the blotch of the patient's blood on the blotter paper. 4. The method of claim 1 wherein the applying step b) comprises the step of applying a plurality of blotter paper to the incision until the bleeding of step a) stops to form a plurality of blotches of the patient's blood on the plurality of blotter paper. 5. The method of claim 1 wherein the measuring step comprises the steps of electronically scanning the blotch formed on the blotter paper in step b) into a computer and electronically calculating the surface area of the blotch formed on the blotter paper in step b). 6. A method of aiding a medical personnel to measure a blood clotting ability of a patient, the method comprising the steps of: a) establishing a relationship between a reference volume and a reference surface area;b) providing the obtained relationship to the medical personnel; andc) advising the medical personnel to perform the following steps of: i) breaking a skin of the patient to induce bleeding;ii) applying blood of the patient obtained from step i) to a first blotter paper, such that a blotch of the patient's blood is formed on the first blotter paper;iii) measuring a surface area of the blotch formed on the first blotter paper; andiv) determining the blood volume in the first blotter paper in accordance with the provided relationship in step b), wherein the determined blood volume is indicative of the blood clotting ability of the patient. 7. The method of claim 6 further comprising the steps of: dispensing a known volume of reference blood on a second blotter paper, such that a reference blotch is formed on the second blotter paper; andmeasuring a surface area of the reference blotch formed on the second blotter paper. 8. The method of claim 7, wherein the dispensing step further comprises the step of using a micropipette to dispense a known volume of reference blood on the second blotter paper. 9. The method of claim 6, wherein the advising step further comprises the step of repeating steps ... | - | A61B005/00
US07840418 | B2 | 2010-11-23 | Tracking the availability of service providers across multiple platforms | 2008-05-29 | Amongst other things, the availability of service providers is tracked across one or more service provider networks. A status change of the service provider is identified on one or more of the service provider networks and the other service provider networks associated with the service provider is sent data indicating the changed status of the service provider. | 1. A computer-implemented method comprises: tracking by one or more computers a status of a medical service provider's present availability for a real-time consultation with a consumer of medical services;identifying by the one or more computers a status change in the medical service provider's present availability;determining by the one or more computers a list of health insurance networks associated with the medical service provider; andsending a message to one or more computer systems associated with the health insurance networks included in the list to update, based on the identified status change, the medical service provider's status in each of the health insurance networks. 2. The method of claim 1 further comprising: when the status of the medical service provider is available, logging the medical service provider into the health insurance networks associated with the medical service provider. 3. The method of claim 1 wherein tracking comprises: using one or more central computer-based availability tracking systems to determine the medical service provider's status. 4. The method of claim 1 further comprising: registering the medical service provider and the medical service provider's status with a central computer-based availability tracking system. 5. The method of claim 1 wherein identifying comprises: receiving data indicating that the medical service provider's status on a central computer-based availability tracking system has changed. 6. The method of claim 1 wherein the actions of tracking, identifying and sending are executed in response to: receiving a request from the consumer of medical services using a brokerage system to consult with a medical service provider having a medical service provider profile that satisfies at least some attributes associated with medical qualifications in a set of attributes that define a suitable medical service provider; andidentifying an available medical service provider associated with the brokerage system that has a medical service provider profile satisfying at least some of the attributes in the set of attributes. 7. The method of claim 1 further comprising: coordinating the medical service provider's status across one or more health insurance networks. 8. The method of claim 1 further comprising: pooling the status of one or more medical service providers in a central computer-based availability tracking system. 9. A computer program product residing on a computer readable medium for providing broker services to consumers and service providers, the computer program product comprising instructions for causing a computer to: track a status of a medical service provider's present availability for a real-time consultation with a consumer of medical services;identify a status change in the medical service provider's present availability;determine a list of health insurance networks associated with the medical service provider; andsend a message to one or more computer systems associated... | American Well Corporation | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07854704 | B2 | 2010-12-21 | Systems, methods and products for diagnostic hearing assessments distributed via the use of a computer network | 2008-10-22 | The systems, methods and associated devices performing diagnostic hearing tests which use a computer network to allow interaction between a test administration site and one or a plurality of remote patient sites. The test can be administered by an audiologist or clinician at a site remote from the patient, in a manner, which can allow interaction between the user and the clinician during at least a portion of the administration of the test. The diagnostic hearing tests can be performed such that they meet standardized guidelines such as ANSI requirements or certification standards and can include distortion product emission level measurements or middle ear compliance measurements. | 1. A method of controlling a diagnostic hearing test, the method comprising the steps of: connecting a portable hearing test device to a local computer to communicate with a remote web client associated with a clinician at a remote site using a computer network, wherein the portable hearing test device has a function generator and an attenuator held therein, and wherein the portable hearing test device is configured to receive commands from the remote web client during a diagnostic hearing test session, at least some of the commands controlling output of the function generator in the hearing test device;serving web pages from a web server in communication with the portable hearing test device to a remote web client which indicate a status of the diagnostic hearing test,receiving requests from the web client in the portable test device which provide parameters for performing the diagnostic hearing test using the portable test device; andcontrolling operation of the hearing test device based on the parameters of the received requests from the web client so as to allow the clinician to control the diagnostic hearing test from the remote site during the hearing test session. 2. A method according to claim 1, wherein the requests from the web client are common gateway interface (CGI) requests transmitted to the portable hearing device which specify parameters for performing the diagnostic hearing test using the portable hearing test device. 3. A method according to claim 1, wherein the serving, receiving and controlling steps are carried out to allow the clinician at the remote site to have substantially real-time interaction with the portable hearing test device and patient during the diagnostic hearing test, and wherein the diagnostic hearing test complies with predetermined standardized testing codes. 4. A method according to claim 1, wherein the portable hearing test device includes: (a) an Internet connection to allow the testing device to communicate with the web client over the Internet, and (b) an on-board processor that receives commands from the web client and directs the operation of the function generator and attenuator in the testing device. 5. A method according to claim 1, wherein the portable hearing test device is discrete from the local computer and is configured with the web-server on-board. 6. A method according to claim 1, wherein the controlling step is carried out in substantially real-time based on input from the clinician at the remote web client to control operation of the diagnostic hearing test, and wherein the portable hearing test device includes a user input device that allows a user to indicate a response to clinician-initiated test stimuli. 7. A method according to claim 1, further comprising providing visual or audiovisual communication between the patient and the clinician during the hearing test session. 8. A hearing evaluation device, comprising: a portable diagnostic hearing test device with an Internet com... | East Carolina University | A61B005/00
US08657757 | B2 | 2014-02-25 | System and methods for the measurement of lung volumes | 2010-07-06 | A method for determining a pulmonary volume change includes: receiving a respiration event from a subject in an airflow chamber; interrupting the respiration event by an occlusion of the airflow chamber initiated at a first time instant and terminated at a second time instant subsequent to the first time instant; taking a plurality of measurements of airflow rate through the airflow chamber between the second time instant and a third time instant subsequent to the second time instant; and determining a pulmonary volume change substantially equal to a reduction of a pulmonary air volume by a pulmonary response air volume and a normal air volume, wherein the pulmonary volume change is related to a change in density of air in the airflow chamber. | 1. A method for determining a pulmonary volume change, comprising: receiving a respiration cycle from a subject in an airflow chamber;interrupting the respiration cycle by an occlusion of the airflow chamber initiated at a first time instant and terminated at a second time instant subsequent to the first time instant;taking a plurality of measurements of airflow rate through the airflow chamber between the second time instant and a third time instant subsequent to the second time instant; anddetermining a pulmonary volume change substantially equal to a reduction of a pulmonary air volume by a pulmonary response air volume and a normal air volume, wherein the pulmonary volume change is related to a change in density of air in the airflow chamber. 2. The method of claim 1, wherein determining a pulmonary volume change substantially equal to a reduction of a pulmonary air volume by a pulmonary response air volume and a normal air volume comprises: determining a pulmonary air volume based on the plurality of measured airflow rates;determining a normal air volume; anddetermining a pulmonary response air volume between a fourth time instant and a fifth time instant. 3. The method of claim 2, wherein determining a normal air volume comprises: determining a normal airflow rate during the respiration cycle between the second time instant and the third time instant, wherein the normal airflow rate is substantially equal to an airflow rate that would have existed between second and third time instants in the absence of the interruption by the occlusion of the airflow chamber; andintegrating the normal airflow rate between the second and third time instants. 4. The method of claim 3, wherein determining a normal airflow rate during the respiration cycle between the second time instant and the third time instant comprises: measuring a normal airflow rate for a time period prior to the first time instant;measuring the normal airflow rate for a time period subsequent to the third time instant; anddetermining the normal airflow rate between the second and third time instants by interpolating from the measured normal airflow rates prior to the first time instant and subsequent to the third time instant. 5. The method of claim 4, wherein the interpolation comprises one of: a spline interpolation;an exponential interpolation; anda polynomial interpolation. 6. The method of claim 3, wherein determining a normal airflow rate during the respiration cycle between the second time instant and the third time instant comprises: measuring at least one of a normal airflow rate for a time period prior to the first time instant and the normal airflow rate for a time period subsequent to the third time instant; anddetermining the normal airflow rate between the second and third time instants by extrapolation from one of the measured normal airflow rate prior to the first time instant and the measured normal airflow rate subsequent to the third time instant. 7. The method... | Pulmone Advanced Medical Devices, Ltd. | A61B005/08
US08313339 | B2 | 2012-11-20 | Connection device and medical system for acquiring electric signals provided with such a connection device | 2008-12-09 | The invention relates to connection means including a female connector (2) and a male connector (1).;The female connector (2) includes: a case (23) including walls (9b, 10b, 11b, 12b) delimiting a cavity for receiving (4) a male connector (1); andcontact fingers (24) extending along at least one of said walls (10b, 11b) and intended to cooperate with the matching contacts (15) of the male connector (1);isolation partitions (27) made of a dielectric material protruding from the wall or walls (10b, 11b) towards the inside of the cavity (4) and positioned between the contact fingers (24).;The invention also relates to a male connector (1) intended to be introduced into the cavity of said female connector (2).;In addition, the invention also relates to a medical system of acquisition of electric signals provided with such a connection device. | 1. A connection device for the transmission of electric signals including: a female connector including: a case including walls delimiting a cavity for receiving a male connector; andfirst contact fingers extending parallel to a first wall and intended to cooperate with matching contacts of the male connector;a male connector including:a body made of a dielectric material intended to be introduced into the female connector, said body including:a housing provided with an opening for the passage of cables;outer faces intended to be positioned opposite the walls delimiting the cavity of the female connector;notches formed on at least a first outer face provided with a lumen opening into said housing; andcontact clamps inserted into said lumens including a contact track intended to cooperate with the contact fingers of the female connector and an insulation-displacement portion extending in the housing;said connection device being characterised in that the female connector includes isolation partitions made of a dielectric material protruding from the first wall towards the inside of the cavity and positioned between the first contact fingers, and in that each notch of the male connector has dimensions adapted for receiving a contact finger and one of the isolation partitions so as to build a partition between said notch and a neighbouring notch. 2. A connection device according to claim 1, characterised in that the female connector includes second contact fingers extending parallel to a second wall opposite the first wall and isolation partitions made of a dielectric material protruding from the second wall towards the inside of the cavity and positioned between the second contact fingers. 3. A connection device according to claim 2, characterised in that the first and second contact fingers are positioned in staggered rows with respect to each other. 4. A connection device according to claim 3, characterised in that the contact fingers are separated by a distance of more than 1 mm, preferably equal to or greater than 1.4 mm. 5. A connection device according to claim 1, characterised in that the isolation partitions are made of polyamide. 6. A connection device according to claim 3, characterised in that the contact fingers are made of brass. 7. A connection device according to claim 3, characterised in that the contact fingers have a cylindrical cross-section. 8. A connection device according to claim 1, characterised in that at least one of the walls of the case is provided with at least one polarising slot. 9. A connection device according to claim 1, characterised in that at least one of the walls is provided with abutting means making it possible to hold the male connector in the cavity. 10. A connection device according to claim 1, characterised in that each notch has, between the lumen and an edge of said notch, a first space for receiving an isolation partition. 11. A connection device according to claim 1, characterised in that ea... | Integral Process | H01R004/24
US08571619 | B2 | 2013-10-29 | Hemoglobin display and patient treatment | 2010-05-19 | The present disclosure describes embodiments of a patient monitoring system and methods that include the measure and display of hemoglobin statistics. In an embodiment, total hemoglobin trending is displayed over a period of time. Statistics can include frequency domain analysis, which may be unique for each patient monitored. The total hemoglobin trending and/or statistics can further be used to help control the treatment of a patient, such as being used to control IV administration. | 1. A patient monitoring system comprising: a noninvasive sensor for emitting energy into a patient measurement site and detecting the energy attenuated by the patient measurement site;a processing board comprising: an instrument manager including a memory buffer; anda FFT module in communication with the instrument manager;wherein the instrument manager is adapted to determine a first plurality of total hemoglobin readings over a first time period and to determine a second plurality of total hemoglobin readings over a second time period using the attenuated energy detected by the sensor and to store the first and second plurality of total hemoglobin readings in the memory buffer;wherein the FFT module is adapted to perform a Fourier transform on the first and second plurality of total hemoglobin readings stored in the memory buffer and to determine a variation between of the outputs of the Fourier transform of the first and second total hemoglobin readings;and wherein the instrument manager is adapted to trigger an alarm if the variation reaches a certain threshold. 2. The patient monitoring system of claim 1 wherein the patient monitoring system includes a display that simultaneously displays a graph of the output of the Fourier transform of the first and second plurality of total hemoglobin readings. 3. The patient monitoring system of claim 2 wherein the first and second time both incorporate at least a complete cycle of the patient's total hemoglobin readings. 4. The patient monitoring system of claim 1 wherein a display outputs a most current total hemoglobin reading. 5. The patient monitoring system of claim 1 wherein the FFT module comprises a special-purpose processor. 6. The patient monitoring system of claim 1 wherein the FFT module comprises an integrated circuit. 7. The patient monitoring system of claim 1 wherein the FFT module comprises a processor running software with a FFT algorithm. 8. The patient monitoring system of claim 1 further comprising: an administration unit adapted to administer treatment to a patient and in communication with the processing board, wherein the treatment is administered based at least in part on one of said outputs of said Fourier transforms of said first or second plurality of total hemoglobin readings. 9. The patient monitoring system of claim 8 wherein the treatment includes administration of at least one from the following: a drug;blood;plasma;nutrition; oran IV fluid. 10. A patient monitor device comprising: a processing device configured to accept signals indicative of optical energy attenuated by patient tissue detected from a noninvasive, optical sensor and further configured to interpret the signals and determine a first plurality of total hemoglobin readings and a second plurality of total hemoglobin readings, wherein the processing device is further configured to perform a Fourier transform on the first and second plurality of total hemoglobin readings and determine a variation betw... | Masimo Corporation | A61B005/00 A61B005/02
US08401605 | B2 | 2013-03-19 | Multiple wavelength physiological measuring apparatus, sensor and interface unit for determination of blood parameters | 2009-04-30 | A measuring apparatus, a physiological sensor, and an interface unit for determining blood parameters of a subject are disclosed. The sensor comprises an emitter unit comprising a first plurality of emitter elements configured to emit radiation at a second plurality of wavelengths and a detector unit configured to receive radiation generated by the emitter elements and transmitted through the tissue of the subject, wherein the detector unit is further configured to produce electric measurement signals indicative of absorption caused by the blood of the subject. The sensor or the interface unit is provided with a memory that stores emitter activation information for at least a third plurality of wavelengths, thereby to enable a monitoring unit operably connectable to the physiological sensor to employ a combination of wavelengths selected from the third plurality of wavelengths, wherein the third plurality is equal to or smaller than the second plurality. | 1. A measuring apparatus for determining the amount of at least one substance in blood of a subject, the measuring apparatus comprising: a sensor unit comprising an emitter unit comprising a first plurality of emitter elements configured to emit radiation at a second plurality of wavelengths; anda detector unit configured to receive radiation generated by the emitter elements and transmitted through tissue of a subject, wherein the detector unit is further configured to produce electric measurement signals indicative of absorption caused by blood of the subject;a first memory storing emitter activation information for at least a third plurality of wavelengths, wherein the third plurality of wavelengths belongs to the second plurality of wavelengths, and wherein the emitter activation information indicates, for each wavelength in the third plurality of wavelengths, input terminals of the sensor unit to which the corresponding emitter element is connected; anda monitoring unit configured to process the electric measurement signals, wherein the monitoring unit is further configured to produce, based on the emitter activation information, control information for employing a desired combination of wavelengths selected from the third plurality of wavelengths. 2. The measuring apparatus according to claim 1, wherein the first memory is in the sensor unit. 3. The measuring apparatus according to claim 1, wherein the emitter activation information indicates, for each wavelength in the third plurality of wavelengths, a polarity of current to be supplied to said input terminals. 4. The measuring apparatus according to claim 1, wherein the emitter activation information comprises at least one control code, and wherein each control code is linked to the control information required to employ a particular combination of wavelengths. 5. The measuring apparatus according to claim 2, wherein the emitter unit comprises at most n+1 input terminals and the first plurality of emitter elements comprises 2n emitter elements. 6. The measuring apparatus according to claim 1, further comprising an interface unit operably connected between the sensor unit and the monitoring unit. 7. The measuring apparatus according to claim 6, wherein the first memory is in the interface unit. 8. The measuring apparatus according to claim 6, wherein the monitoring unit comprises a single drive current source operably connectable to an emitter switching unit in the interface unit. 9. The measuring apparatus according to claim 1, wherein the monitoring unit is further configured to change the desired combination of wavelengths in time domain. 10. The measuring apparatus according to claim 1, wherein the monitoring unit is further configured to pre-check compatibility of the sensor unit with any combination of wavelengths intended to be employed in the measuring apparatus. 11. A physiological sensor for use in determining the amount of at least one substance in blood of a subject, ... | General Electric Company | A61B005/1455
US08708926 | B2 | 2014-04-29 | Double temperature sensor | 2010-03-31 | A double temperature sensor is provided for determining the body temperature. The sensor includes at least a first temperature sensor element for measuring a near-surface temperature of the ambient air and a second temperature sensor element for measuring a temperature of the skin surface. The first and the second temperature sensor elements are arranged in a sensor block. A holding element surrounds the sensor block at least partially. An adhesive element fixes the holding element on the skin surface. A spring element acts on the sensor block and provides a spring force directed in the direction of the skin surface. | 1. A double temperature sensor for determining body temperature, the double temperature sensor comprising: a first temperature sensor element for measuring a near-surface temperature of ambient air;a second temperature sensor element for measuring a temperature of a skin surface, said first temperature sensor element and said second temperature sensor element being arranged in a sensor block;a holding element at least partially surrounding said sensor block;an adhesive element connected on one side to said holding element, said adhesive element being adapted for fixing said holding element on the skin surface with another side of said adhesive element, said adhesive element being made with a layer of foamed material and a gel layer;a spring element, which acts on said sensor block and whose spring force directs said sensor block towards the skin surface when said adhesive element fixes said holding element to the skin surface,at least two connection cables for connecting the temperature sensor elements to an analyzing unit, wherein a first connection cable is arranged between said layer of foamed material and said gel layer and a second connection cable is arranged on said layer of foamed material. 2. The double temperature sensor in accordance with claim 1, wherein said spring element is arranged between said holding element and said sensor block. 3. The double temperature sensor in accordance with claim 1, wherein said spring element is integrated in said holding element. 4. The double temperature sensor in accordance with claim 1, wherein said spring element comprises at least one of an elastic material, foamed material and a mechanical spring. 5. The double temperature sensor in accordance with claim 1, wherein said sensor block, said holding element and said adhesive element are arranged coaxially with one another, wherein the temperature sensor elements are arranged along a central axis of said sensor block. 6. The double temperature sensor in accordance with claim 1, wherein at least one air inclusion is provided between said holding element and said sensor block. 7. The double temperature sensor in accordance with claim 1, further comprising a heat-insulating element comprising a foamed material coaxially surrounding said holding element. 8. The double temperature sensor in accordance with claim 7, wherein said heat-insulating element is of a cylindrical design. 9. The double temperature sensor in accordance with claim 1, wherein a pushbutton is provided in an area of the sensor block directed towards the skin surface. 10. A temperature sensor for an object with a surface, the temperature sensor comprising: a sensor block;a first temperature sensor element arranged in said sensor block and measuring a temperature at an ambient air surface of said sensor block;a second temperature sensor element arranged in said sensor block and adapted for measuring a temperature of the surface of the object;a holding element connected to said se... | Drägerwerk AG & Co. KGaA | A61B005/01 G01K013/00
US20100298696 | A1 | 2010-11-25 | SELF-CONTAINED, SELF-PIERCING, SIDE-EXPELLING MARKING APPARATUS | 2010-08-05 | A marking apparatus includes a rigid cannula having a peripheral wall forming a lumen that carries an imaging marker, a proximal end coupled to a handle, a lateral opening in the peripheral wall that is open to the lumen, and a closed-off distal portion having a ramp adjacent the lumen. The closed-off distal portion extends distally from the ramp to terminate at a tissue piercing pointed tip. The ramp of the rigid cannula is curved to transition from the peripheral wall of the rigid cannula to the lateral opening of the rigid cannula. The ramp of the rigid cannula is adapted to engage the distal end of the stylet as the stylet is advanced through the lumen of the rigid cannula to guide the distal end of the stylet to a position to substantially close off the lateral opening of the rigid cannula having the tissue piercing tip. | 1. A marking apparatus for the percutaneous placement of an imaging marker in a tissue mass, the marking apparatus comprising: a handle to be grasped by a user;a rigid cannula having: a peripheral wall forming a lumen that carries the imaging marker,a proximal end coupled to the handle,a lateral opening in the peripheral wall that is open to the lumen, anda closed-off distal portion having a ramp adjacent the lumen, the closed-off distal portion extending distally from the ramp to terminate at a tissue piercing pointed tip, the ramp of the rigid cannula being curved to transition from the peripheral wall of the rigid cannula to the lateral opening of the rigid cannula; anda stylet having a distal end, with at least the distal end of the stylet being flexible, the stylet being slidably received within the lumen of the rigid cannula for movement in the lumen, the ramp of the rigid cannula being adapted to engage the distal end of the stylet as the stylet is advanced through the lumen of the rigid cannula to guide the distal end of the stylet to a position to substantially close off the lateral opening of the rigid cannula having the tissue piercing tip. 2. The marking apparatus of claim 1, wherein only the distal end of the stylet closes off the lateral opening. 3. The marking apparatus of claim 1, wherein the stylet received in the lumen has a ready position, where the distal end of the stylet is spaced inwardly from the closed-off distal portion to form a marker recess that is in communication with the lateral opening and that is sized to receive the imaging marker, and an expelled position, where the distal end of the stylet is advanced a sufficient distance into the marker recess to expel the imaging marker contained in the lumen of the rigid cannula through the lateral opening, and the ramp being adapted to engage the distal end of the stylet and guide the distal end of the stylet into the lateral opening to substantially close off the lateral opening of the rigid cannula after the imaging marker is expelled. 4. The marking apparatus of claim 3, and further comprising multiple imaging markers contained within the marker recess. 5. The marking apparatus of claim 1, wherein the rigid cannula is 13 gage or less. 6. The marking apparatus of claim 1, wherein the handle, the rigid cannula, and the stylet are operably coupled to form a self-contained marking apparatus. 7. A marking apparatus for the percutaneous placement of an imaging marker in a tissue mass, comprising: a handle to be grasped by a user;a cannula having: a peripheral wall forming a lumen that carries the imaging marker,a proximal end coupled to the handle,a lateral opening in the peripheral wall that is open to the lumen, anda closed-off distal portion having a ramp adjacent the lumen, the closed-off distal portion extending distally from the ramp to terminate at a tissue piercing pointed tip; anda stylet having a distal end with an angled surface, the stylet being disposed in... | BARD PERIPHERAL VASCULAR, INC. BARD SHANNON LIMITED | A61B005/05
US20100331716 | A1 | 2010-12-30 | METHODS AND APPARATUS FOR MEASURING RESPIRATORY FUNCTION USING AN EFFORT SIGNAL | 2009-06-26 | Breathing effort of a patient, as determined (for example) from a photoplethysmograph (“PPG”) signal from the patient, can be used in conjunction with signals from a spirometer being used by the patient to provide additional information about various characteristics of the patient's breathing. For example, such additional information may include the relationship between the patient's breathing effort and breath volume. As another example, the additional information may include the relationship between the patient's breathing effort and rate of breath flow. Still other such additional information may be derived from various combinations of spirometer and PPG output signals. | 1. Apparatus for detecting a breathing characteristic of a patient comprising: a spirometer operatively coupled to the patient for producing a volume signal indicative of volume of the patient's breath;breathing effort monitoring apparatus operatively coupled to the patient for producing a breathing effort signal from the patient; andfirst circuitry for using the volume signal and the effort signal to determine a relationship between the volume of the patient's breath and the effort the patient exerted to produce that volume of breath. 2. The apparatus defined in claim 1 wherein the first circuitry produces a volume vs. effort signal indicative of the relationship determined by the first circuitry; and wherein the apparatus further comprises: output circuitry for outputting the volume vs. effort signal in human-readable form. 3. The apparatus defined in claim 1 wherein the spirometer additionally produces a flow signal indicative of a rate of flow of the patient's breath; and wherein the apparatus further comprises: second circuitry for using the flow signal and the effort signal to determine a relationship between the rate of flow of the patient's breath and the effort the patient exerted to produce that rate of flow. 4. The apparatus defined in claim 3 wherein the second circuitry produces a flow vs. effort signal indicative of the relationship determined by the second circuitry; and wherein the apparatus further comprises: output circuitry for outputting the flow vs. effort signal in human-readable form. 5. The apparatus defined in claim 1 wherein the breathing effort monitoring apparatus comprises: scalogram circuitry for producing scalogram signals; andbreathing effort extraction circuitry for extracting the effort signal from the scalogram signals. 6. The apparatus defined in claim 1 further comprising: compliance measure circuitry for determining a patient compliance measure signal from the relationship determined by the first circuitry. 7. The apparatus defined in claim 1 further comprising: maximum effort detection circuitry for determining a maximum effort signal from the effort signal. 8. The apparatus defined in claim 3 further comprising: respiratory resistance measure circuitry for determining a respiratory resistance measure signal from the relationship determined by the second circuitry. 9. The apparatus defined in claim 3 further comprising: FEF50 determination circuitry for determining an FEF50 signal from the relationship determined by the second circuitry. 10. The apparatus defined in claim 9 further comprising: effort ratio determination circuitry for determining a ratio between effort at FEF50 and maximum effort. 11. The apparatus defined in claim 1 wherein the breathing effort monitoring apparatus comprises: a photoplethysmograph (“PPG”) for producing a PPG signal from the patient. 12. Apparatus for detecting a breathing characteristic of a patient comprising: a spirometer operatively coupled to the patient for p... | Nellcor Puritan Bennett Ireland | A61B005/087
US20100280975 | A1 | 2010-11-04 | SYSTEMS FOR CLINICAL TRIALS | 2008-11-05 | The invention provides methods and systems for assessing the efficacy of a pharmaceutical which is putatively disease modifying of a cognitive disorder, for use in the treatment or prophylaxis of that cognitive disorder, the method comprising the steps of: (1) stratifying a subject group into at least 2 sub-groups according to a baseline indicator of likely disease progression, (2) treating members of each subject group with the pharmaceutical for a treatment time frame, (3) deriving psychometric and optionally physiological outcome measures for each treated patient group, (4) comparing the outcomes at (3) with a comparator arm of said sub-groups which is optionally a placebo or minimal efficacy comparator arm, (5) using the comparison in (4) to derive an efficacy measure for the pharmaceutical. The methods and systems of the invention address problems such as low rate of decline over the treatment time-frame of patients who have mild-disease severity at baseline and biased withdrawal, particularly in the placebo/comparator treatment arm. | 1. A method for assessing the efficacy of a pharmaceutical which is putatively disease modifying of a cognitive disorder, for use in the treatment or prophylaxis of that cognitive disorder, the method comprising the steps of: (1) stratifying a subject group into at least 2 sub-groups according to a baseline indicator of likely disease progression,(2) treating members of each subject group with the pharmaceutical for a treatment time frame,(3) deriving psychometric and optionally physiological outcome measures for each treated patient group,(4) comparing the outcomes at (3) with a comparator arm of said sub-groups which is optionally a placebo or minimal efficacy comparator arm,(5) using the comparison in (4) to derive an efficacy measure for the pharmaceutical. 2. A system for assessing the efficacy of a pharmaceutical which is putatively disease modifying of a cognitive disorder, for use in the treatment or prophylaxis of that cognitive disorder, the system comprising the steps of: (1) stratifying a subject group into at least 2 sub-groups according to a baseline indicator of likely disease progression,(2) selecting a treatment time frame over which members of each subject group are to be treated with the pharmaceutical,(3) selecting psychometric and optionally physiological outcome measures to be derived for each treated patient group and a comparator arm of said sub-groups which is optionally a placebo or minimal efficacy comparator arm, whereby the efficacy measure for the pharmaceutical may be derived from a comparison of the treated patient groups and the comparator arm. 3. A method as claimed in claim 1 wherein the cognitive disorder is a neurodegenerative disorder causing dementia. 4. A method as claimed in claim 3 wherein the pharmaceutical is a putative inhibitor of pathological protein aggregation, where the aggregation is associated with the neurodegeneration. 5. A method as claimed in claim 3 wherein the neurodegenerative disorder is a tauopathy. 6. A method as claimed in claim 5 wherein the neurodegenerative disorder is selected from: Alzheimer's disease, MCI, motor neurone disease, Fronto-temporal dementia, Lewy body disease, Pick's disease, Progressive Supranuclear Palsy. 7. A method as claimed in claim 6 wherein the neurodegenerative disorder is Alzheimer's disease or MCI and the psychometric measures are selected from the group consisting of: Alzheimer's Disease Assessment Scale-cognitive subscale (ADAS-cog), National Institute of Neurological and Communicative Disorders and Stroke-Alzheimer's Disease and Related Disorders Association (NINCDS-ADRDA), Diagnostic and Statistical Manual of Mental Disorders, 4th Edn (DSMIV). 8. A method as claimed in claim 3 wherein the baseline indicator into which the sub-groups are stratified is disease severity using the Clinical Dementia Rating (CDR) scale. 9. A method as claimed in claim 8 wherein the sub-groups are subjects having a CDR rating of 1 (mild sub-group) or 2 (moderate sub-... | WisTa Laboratories Ltd. | G06Q090/00 A61B005/00
US20100317949 | A1 | 2010-12-16 | OPTICAL COUPLER FOR NON-INVASIVE SPECTROPHOTOMETRIC PATIENT MONITORING | 2009-06-12 | Flexible, low-cost, physically robust optical coupling patches for use in spectrophotometric patient monitoring, and methods of fabrication thereof, are described. The optical coupling patch comprises a flexible base layer having a skin-contacting surface and a first aperture formed therethrough that establishes an optical interface with a skin surface when the base layer is placed against the skin surface. The optical coupling patch further comprises an elastomeric waveguiding member laterally disposed on a surface of the base layer opposite the skin-contacting surface. The optical coupling patch guides optical radiation between a laterally propagating state at a first location laterally distal from the first aperture and a generally vertically propagating state at the first aperture. | 1. An optical coupling patch for use in spectrophotometric patient monitoring, comprising: a flexible base layer having a skin-contacting surface and a first aperture formed therethrough, said flexible base layer comprising a first elastomeric material having a first refractive index, said first aperture establishing an optical interface with a skin surface when said flexible base layer is placed against said skin surface;an elastomeric waveguiding member laterally disposed on a surface of said flexible base layer opposite said skin-contacting surface, said elastomeric waveguiding member comprising a second elastomeric material having a second refractive index greater than said first refractive index and being configured to guide optical radiation between (i) a laterally propagating state at a first location laterally distal from said first aperture, and (ii) a generally vertically propagating state at said first aperture, wherein said elastomeric waveguiding member includes a substantially planar reflecting surface shaped integrally thereinto near said first aperture, said reflecting surface being oriented at an angle that causes reflective redirection of the optical radiation between said laterally propagating state and said vertically propagating state; anda flexible cladding material having a third refractive index less than said second refractive index and selectively covering said elastomeric waveguiding member such that a cavity occupied by one of air and a low-index material having a fourth refractive index less than said third refractive index is formed directly adjacent said reflecting surface for facilitating said reflective redirection of the optical radiation. 2. The optical coupling patch of claim 1, said elastomeric waveguiding member having a laterally facing end facet at said first location, said elastomeric waveguiding member thereby guiding the optical radiation between said end facet and said first aperture. 3. The optical coupling patch of claim 2, said laterally facing end facet being adapted for coupling to an optical radiation source external to said optical coupling patch, said elastomeric waveguiding member guiding the source optical radiation to said first aperture, the source radiation thereby propagating into tissue underlying the skin surface. 4. The optical coupling patch of claim 3, said laterally facing end facet being a first end facet and said elastomeric waveguiding member being a first elastomeric waveguiding member, said optical coupling patch further comprising a second elastomeric waveguiding member formed similarly to said first elastomeric waveguiding member and extending between a second laterally facing end facet and a second aperture formed through said skin-contacting surface, said second laterally facing end facet being adapted for coupling to an optical radiation detector external to said optical coupling patch, said second elastomeric waveguiding member guiding optical radiation received through... | 02 MEDTECH, INC. | A61B005/1455
US07835926 | B1 | 2010-11-16 | Method for conducting a home health session using an integrated television-based broadband home health system | 2002-08-29 | An integrated home health system includes a television-based patient station, a first provider station for providing telemedicine or other healthcare services to a patient located at the patient station, a second provider station for providing caregiver services to the patient, a third provider station for providing emergency response services to the patient and a system management station coupled together by a data network. In addition to various management operations performed on behalf of the integrated home health system, the system management station is further configured to provide various home health services to the patient located at the patient station, either alone, or in conjunction with one or more of the first, second and/or third provider stations. | 1. For a home health system having a home health provider station coupled to a patient station by a data network, a method for conducting a home health session between said home health provider station and said patient station, comprising: transmitting video data from said home health provider station to said patient station over said data network;receiving said video data via an interface device at said patient station;receiving, via the interface device at said patient station, an unrelated television broadcast signal originating from a television signal source;transmitting, via the interface device, a combination of said video data and said television broadcast signal to a television set which forms a first part of said patient station; anddisplaying, on a television screen of said television set at said patient station, video images constructed from said video data, said video images overlaying said television broadcastwherein said video data is transmitted to said patient station and said video images overlaying said television broadcast signal are displayed on said television screen absent any request from said patient station. 2. The method of claim 1, wherein said home health provider station is a healthcare provider station and said home health session is a telemedicine session. 3. The method of claim 1, wherein said home health provider station is a healthcare provider station and said home health session is a telehealth session. 4. The method of claim 1, wherein said home health provider station is a caregiver provider station and said home health session is a caregiver session. 5. The method of claim 1, wherein said interface device at said patient station is coupled to said data network by a broadband access device and wherein transmitting video data from said home health provider station to said patient station further comprises: generating video data at said home health provider station;placing said home health provider station video data onto said data network; andsaid broadband access device transferring said home health provider station video data from said data network to said interface device at said patient station. 6. The method of claim 5, and further comprising: generating audio data at said home health provider station;placing said home health provider station audio data onto said data network;said broadband access device transferring said home health provider station audio data from said data network to said interface device at said patient station; andgenerating, at said patient station, audible sound from said home health provider station audio data transmitted from said home health provider station. 7. The method of claim 6, wherein said audible sound is generated by said television set. 8. The method of claim 6, wherein said audible sound is generated by a telephone handset which forms a second part of said patient station. 9. The method of claim 6, and further comprising: generating video data at said patien... | Telehealth Broadband LLC | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00 G06Q040/00
US07853311 | B1 | 2010-12-14 | Surgical targeting system | 2000-04-21 | A targeting device for providing a series of coordinates/lines within a sheet of sterile, flexible material with an adherent surface which is applied to the skin (after suitable surgical preparation). The sheet is non-porous and may have a topical antiseptic on the side which is applied to the skin. A fluoroscope or roentgenographic image of the portion of the body to which the adherent film is applied will show the underlying skeletal and radiopaque elements as well as the overlying surgical grid. Once the targeting device is applied, the coordinates on the grid lines are clearly visible on the surface of the skin as well as on the fluoroscopic or radiographic image and by knowing the direction of the fluoroscopic or radiographic beam, the operator will be able to thereby correlate a specific locus on the skin with an underlying skeletal element or other underlying radiopaque structure. By applying the targeting device to the part of the body in a circumferential or nearly circumferential manner, and utilizing a radiolucent operating room table, and a radiograph or C-arm fluoroscope, targeting techniques using the targeting device can be utilized at surgery. Multiple targeting or percutaneous procedures can be performed at the same sitting with the application of a single device. | 1. A method for correlating the buttock with the femoral canal of the femur of a body, said method comprising the steps of: applying a radio-transparent drape having at least two indicia each comprising a radio-opaque longitudinal axis to the leg of the body such that a first portion of the drape extends in an anterior-posterior plane relative to the body, said applying step further providing for a second portion of the drape to extend laterally relative to the body, said applying step providing further for each of the indicia to be contained in respective first and second portions of the drape, said applying step providing further for each of the indicia to be longitudinally and centrally aligned relative to the leg;directing imaging radiation through said drape such that a radiographic image of said body and indicia is formed on a medium;comparing, by viewing the radiographic image, the relative positions of each of the indicia relative to the longitudinal axis of the femoral canal;translating the drape, as required, relative to the leg such that one of the indicias is contained in an anterior-posterior plane which coincides with the longitudinal axis of the femoral canal, and such that the other of the indicia is contained in a lateral plane which coincides with the longitudinal axis of the femoral canal; andlocating the intersection of the indicia on the buttock, the intersection of the indicia defining a start point for a reference axis which, when intersecting said start point and parallel to the indicia, coincides with the longitudinal axis of the femoral canal. 2. A method as set forth in claim 1 and further comprising the steps of: positioning a longitudinal nail relative to the leg such that a pointed end of the nail is adjacent to the start point on the buttock;orienting the nail relative to the leg such that the longitudinal axis of the nail coincides with the reference axis;inserting the nail through the tissue of the leg such that the longitudinal axis of the nail coincides with the reference axis, said inserting step being initiated by puncturing the outer surface of the leg at the start point with the pointed end of the nail;inserting the nail further through the tissue such that the longitudinal axis of the nail continues to coincide with the reference axis, and such that the pointed end of the nail punctures the proximal end of the femoral canal; andinserting the nail further into the femoral canal such that the longitudinal axis of the nail continues to coincide with the reference axis. 3. A sterile surgical drape comprising: a radio-lucent sheet, wherein the sheet comprises a central cutout and a slit extending outward from the central cutout;adhesive on a major surface of the radio-lucent sheet;a radio-opaque pattern on the sterile surgical drape, wherein the radio-opaque pattern comprises a perimeter, an interior within the perimeter, and a plurality of intersections distributed within the interior of the radio-opaque patt... | 3M Innovative Properties Company | A61B005/00
US20100292584 | A1 | 2010-11-18 | SPECTRUM ANALYTICAL METHOD FOR QUANTIFYING HEAT-LUNG INTERACTION | 2009-05-15 | The present invention is related to a spectrum analytical method for quantifying hear-lung interaction, which can estimate cardiac function by using a heart-associated monitoring signal. According to the method of the present invention, quantification of heart-lung interaction is conducted by choosing spectrum signals within a specified frequency band, such that the interference to the heart-associated monitoring signals by incidental events occurring at a low frequency, can be avoided. Therefore, the method of the present invention can be performed even in the subjects who are not in a state of general anesthesia or sedation, and hence is very useful in estimating the cardiac function of the test subjects. | 1. A spectrum analytical method for quantifying heart-lung interaction, comprising performing spectrum analysis of arterial blood pressure signals within a time domain by the following steps: (a) transforming the arterial blood pressure signals to pulse pressure signals;(b) subjecting the pulse pressure signals to spectrum transform to obtain power spectrum signals;(c) choosing a frequency band of 0.1˜1.5 Hz from the power spectrum signals to obtain a power spectral density distribution curve, integrating all energy values of the power spectral densities over the chosen frequency band to obtain an integrated energy value of the power spectral densities, which is used as a predictor of cardiac function. 2. The spectrum analytical method of claim 1, wherein there are at least 10 respiratory cycles in the time domain. 3. The spectrum analytical method of claim 1, wherein the pulse pressure signal in the step (a) is obtained according to the following equation: PPnorm=(PP−PPmean)/PPmean whereinPPnorm is a normalized pulse pressure signal,PP is an arterial blood pressure signal,PPmean is a mean value of the arterial blood pressure signals within the time domain. 4. The spectrum analytical method of claim 1, wherein the pulse pressure signals are subjected to spectrum transform by fast Fourier transform in the step (b). 5. The spectrum analytical method of claim 1, wherein the power spectrum signals in the step (b) are normalized. 6. The spectrum analytical method of claim 1, wherein the integrated energy value of the power spectral densities in the step (c) equal to or less than 4.62×10−4 (min−1) indicates that the cardiac blood volume is sufficient. 7. A spectrum analytical method for quantifying heart-lung interaction, comprising performing spectrum analysis of stroke volume signals within a time domain by the following steps: (a) subjecting the stroke volume signals to spectrum transform to obtain power spectrum signals;(b) choosing a frequency band of 0.1˜1.5 Hz from the power spectrum signals to obtain a power spectral density distribution curve, integrating all energy values of the power spectral densities over the chosen frequency band to obtain an integrated energy value of the power spectral densities, which is used as a predictor of cardiac function. 8. The spectrum analytical method of claim 7, wherein there are at least 10 respiratory cycles in the time domain. 9. The spectrum analytical method of claim 7, wherein the stroke volume signals are subjected to spectrum transform by fast Fourier transform in the step (b). 10. The spectrum analytical method of claim 7, wherein the stroke volume signal in the step (b) is normalized. 11. The spectrum analytical method of claim 7, wherein the integrated energy value of the power spectral densities in the step (c) equal to or less than 4.62×10−4 (min−1) indicates that the cardiac blood volume is sufficient. 12. A spectrum analytical method for quantifying heart-lung interaction, comprisin... | - | A61B005/02
US20100312128 | A1 | 2010-12-09 | SYSTEMS AND METHODS FOR MONITORING BLOOD PARTITIONING AND ORGAN FUNCTION | 2009-12-29 | Methods and systems for monitoring an organ of interest within a patient use one or more sensors to obtain one or more signals indicative of one or more of blood being provided to the organ of interest, blood being received from the organ of interest, and blood present in the organ of interest. Changes in an amount of blood being provided to the organ of interest, an amount of blood being received from the organ of interest, and/or an amount of blood present in the organ of interest are monitored based on changes in the obtained signal(s). Such methods and systems can be used to detect dysfunction of the organ of interest or tumor growth in the organ of interest, but are not limited thereto. | 1. A method for monitoring an organ of interest within a patient, the method comprising: (a) using one or more sensors to obtain one or more signals indicative of one or more of blood being provided to the organ of interest, blood being received from the organ of interest, and blood present in the organ of interest; and(b) monitoring changes in one or more of an amount of blood being provided to the organ of interest, an amount of blood being received from the organ of interest, and an amount of blood present in the organ of interest, based on changes in at least one of the one or more obtained signals. 2. The method of claim 1, wherein each of the one or more sensors comprises an impedance sensor, each impedance sensor including at least two electrodes, and each of the obtained signals comprises an impedance plethysmography signal. 3. The method of claim 1, wherein each of the one or more sensors comprises an optical sensor, each optical sensor including a light source and a light detector, and each of the obtained signals comprises a photoplethysmography signal. 4. The method of claim 1, wherein step (a) includes using said one or more sensors to obtain one or more plethysmography signals indicative of one or more of blood being provided to the organ of interest, blood being received from the organ of interest, and blood present in the organ of interest. 5. The method of claim 1, wherein step (b) includes: (b.1) determining, from time to time, one or more metrics based on at least one of the one or more obtained signals, wherein the one or more metrics is/are indicative of one or more of blood being provided to the organ of interest, blood being received from the organ of interest, and blood present in the organ of interest; and(b.2) monitoring changes one or more said determined metrics, over time, to thereby monitor changes in one or more of an amount of blood being provided to the organ of interest, an amount of blood being received from the organ of interest, and an amount of blood present in the organ of interest. 6. The method of claim 5, wherein step (b.2) includes monitoring whether one or more of the amount of blood being provided to the organ of interest, the amount of blood being received from the organ of interest, and the amount of blood present in the organ of interest is increasing, decreasing or staying relatively the same. 7. The method of claim 5, wherein step (b) includes (b.3) comparing a said metric to a corresponding baseline; and further comprising: (c) triggering an alert and/or therapy if the said metric falls below or rises above the corresponding baseline by at least a specified threshold. 8. The method of claim 5, wherein a said metric determined at step (b.1) is indicative of blood volume of blood vessels known to provide blood to the organ of interest. 9. The method of claim 5, wherein a said metric determine at step (b.1) is indicative of venous oxygen saturation or arterial oxygen saturation. 10. The met... | - | A61B005/0265
US20100312093 | A1 | 2010-12-09 | COMBINED APPARATUS FOR IMAGING THE INNER PART OF A BODY, AND METHOD THEREOF | 2010-06-07 | A combined apparatus for diagnostic imaging an inner part of a body, the apparatus including at least a device for imaging and displaying images according to at least two types of beams, waves or signals transmitted toward a body or a part of the body to be examined, the device including a Nuclear Magnetic Resonance imaging and displaying device and an ultrasound imaging and displaying device in an integrated configuration within the same apparatus. | 1. A magnetic resonance imaging apparatus comprising: a static field magnet defining a patient receiving space for receiving at least a part of a patient body; anda patient support adapted to hold a patient's body or at least a part of the patients body;wherein at least said patient support or at least said static field magnet or both said patient support and said static field magnet being displaceable so as to move said patient support and said static field magnet relative one to the other over a range of relative positions;at least one of the patient support and the static field magnet or both being displaceable along al least one of three orthogonal axis and/or rotatable or oscillatable relative one to the other around at least one axis of rotation so that the patient body and/or the part of the patient body is displaceable over a range of translational or angular positions within the static field magnet relative to the said static field magnet. 2. The magnetic resonance imaging apparatus according to claim 1, wherein mechanical links between the static field magnet and the patient support are provided, said links comprising one or more guides along which the magnet and/or the patient support are displaceable with a translational motion or rotary supports defining one or more axis of rotation of said magnet and/or of said patient support relative one to the other. 3. The magnetic resonance imaging apparatus according to claim 2, further comprising motor driven actuators provided for carrying out the translational and the rotational displacement of the static field magnet and/or of the patient support. 4. The magnetic resonance imaging apparatus according to claim 1, further comprising tracer means for detecting the relative position and orientation of the magnet and of the patient support table along the directions of translational and rotational displacement. 5. The magnetic resonance imaging apparatus according to claim 4, wherein the tracer means are proximity sensors. 6. The magnetic resonance imaging apparatus according to claim 4, wherein the tracer means are encoders for detecting the angles of rotation and the number of revolutions completed by the driving motors. 7. The magnetic resonance imaging apparatus according to claim 4, wherein the tracer means include means for reading the resistance between conductors associated to the guides and/or to the rotational links between the magnet and the patient support. 8. The magnetic resonance imaging apparatus according to claim 4, wherein mechanical links between the static field magnet and the patient support are provided, said links comprising one or more guides along which the magnet and/or the patient support are displaceable with a translational motion or rotary supports defining one or more axis of rotation of said magnet and/or of said patient support relative one to the other, and wherein the tracer means include code reading means associated with the guides and/or the rotatio... | ESAOTE S.p.A. | A61B005/055 A61B008/00 A61G013/00
US20100312076 | A1 | 2010-12-09 | POP BOX METHODS AND SYSTEMS FOR ASSESSING RISK OF PRESSURE RELATED SKIN INJURY AND DETERMINING A RECOMMENDED TREATMENT THEREOF | 2010-06-03 | A system or method for assessing a patient for pressure related skin injury risk and recommended treatment including measuring extrinsic and intrinsic skin pressure of the patient and assigning a rating responsive to the skin pressure measurement, measuring oxygenation of the patient and assigning a rating responsive to the oxygenation measurement, measuring perfusion of the patient and assigning a rating responsive to the perfusion measurement, summing the skin pressure, oxygenation and perfusion ratings to obtain a POP Box score, determining a risk of pressure related skin injury for the patient in response to the POP Box score, and determining or selecting a recommended treatment from among a plurality of recommended treatments for the patient in response to the determined risk of pressure related skin injury. | 1. A system for assessing a patient for pressure related skin injury risk and recommended treatment comprising: a memory;an input device;an output device; anda processor,wherein the memory and processor include computer executable instructions for implementing the method ofreceiving a measured skin pressure of the patient;assigning a rating responsive to the received skin pressure measurement;receiving a measured tissue oxygenation of the patient;assigning a rating responsive to the received oxygenation measurement;receiving a measured perfusion of the patient;assigning a rating responsive to the received perfusion measurement;summing the skin pressure, tissue oxygenation and perfusion ratings to obtain a POP Box score;determining a risk of pressure related skin injury for the patient in response to the POP Box score; anddetermining a recommended treatment from among a plurality of recommended treatments for the patient in response to the determined risk of pressure related skin injury. 2. The system of claim 1 wherein the computer executable instructions for receiving measured tissue oxygenation include receiving measured tissue oxygenation as determined by StO2 and wherein the assigned rating for oxygenation is responsive to the received StO2 measurement. 3. The system of claim 2 wherein assigning a rating for the measured tissue oxygenation based on StO2 includes: assigning a first StO2 rating where StO2 levels of WNL SO2 are less than 10 percent of normal;assigning a second StO2 rating where StO2 is between 10 and 20 percent of normal;assigning a third StO2 rating where StO2 is between 20 and 30 percent of normal; andassigning a fourth StO2 rating where StO2 is greater than 30 percent of normal. 4. The system of claim 1 wherein receiving the measured tissue oxygenation includes one or more of the following processes: receiving a measurement of a venous oxygen saturation (SvO2) or a central venous oxygen saturation (SvcO2) and assigning a rating responsive to the measured SvO2 or SvcO2;receiving a measurement of a ratio of arterial oxygen concentration to the fraction of inspired oxygen (P/F ratio) of the patient and assigning a rating responsive to the measured P/F ratio;receiving a measurement of a desaturation recovery time of oxygen saturation (SpO2) of the patient and assigning a rating responsive to the measured (SpO2);receiving a measurement of a blood glucose of a patient and assigning a rating responsive to the measured blood glucose;receiving a hemogolin/hematocrit (H/H) measurement of the patient and assigning a rating responsive to the measured H/H;receiving a measured vital sign of the patient including the blood pressure and heart rate and assigning a rating responsive to the measured vital sign;receiving a Cardiac Index (CI), Cardiac Output (CO), or Stroke Volume (SV) measurement of the patient and assigning a rating responsive to the measured CI, CO or SV;receiving a measured or identified presence of one or more pressor age... | - | A61B005/00 G06N005/02
US20100286488 | A1 | 2010-11-11 | METHOD AND SYSTEM FOR USING A MOBILE DEVICE AS A PORTABLE PERSONAL TERMINAL FOR MEDICAL INFORMATION | 2005-08-25 | A medical information and support system that allows user's to gain access to their medical information on a mobile device terminal that interfaces to an application and server system through a wireless network. The mobile device may also interface to medical measurement systems to obtain vital and/or physiological measurements of the user. They system can trigger alarms if the physiological measurements are of concern. In addition, the mobile device can be used to provide instructions to a user to perform certain medical actions such as taking medication, as well as identifying the dosages and types of medication. The user can mobile device or through a web interface to the application and server system. In addition, health care entities may also access and/or modify the medical information through a web interface to the application and server system | 1. A system for maintaining and providing medical data pertinent to a user, the system comprising: a mobile device;an application and storage server communicatively coupled to the mobile device and operative to maintain a personal file for the user, the personal file including the pertinent medical data; and the mobile device being operative to access the personal file and render the data to the user. 2. The system of claim 1, wherein a copy of the personal file is maintained in memory of both the application and storage server and the mobile device; and the application and storage server and the mobile device are operative to synchronize the copies of the personal file. 3. The system of claim 1, wherein the mobile device is further operative to receive additional data to be included in the personal file. 4. The system of claim 2, wherein the application and storage server includes a third party interface, wherein a entity can access the personal file through the third party interface and modify the data within the personal file. 5. The system of claim 2, wherein the mobile device includes a software module that is operative to prompt the user to take certain medical actions based on the contents of the personal file. 6. The system of claim 5, wherein the software module is further operative to instruct a user to take a particular dosage of medication. 7. The system of claim 5, wherein the software module is further operative to receive an indicator that the user has performed the medical action. 8. The system of claim 7, wherein the software module is further operative to trigger an alarm if the software module has not received the indicator within a threshold period of time. 9. The system of claim 7, wherein the software module is further operative to trigger an alarm if the software module receives an indicator that the user has not performed the medical action. 10. The system of claim 2, wherein the mobile device includes a medical measurement device interface over which the mobile device can receive signals from a medical measurement device relevant to the user. 11. The system of claim 10, wherein the medical measurement device can measure one or more vital signs of the user and provide the measurement to the medical measurement device interface of the mobile device. 12. The system of claim 10, wherein the software module is further operative to determine whether to trigger an alarm based on the measurement and the synchronized personal file. 13. The system of claim 2 wherein the mobile device is communicatively coupled to the application and server system through a wireless telecommunications network that includes data transmission. 14. The system of claim 2, wherein the mobile device is communicatively coupled to the application and server through a digital cellular telecommunications network 15. The system of claim 10, wherein the medical measurement device is an integral part of the mobile device. 16. The system of claim 2... | - | A61B005/00 G06Q050/00
US07835779 | B2 | 2010-11-16 | Magnetic tracking system | 2006-05-10 | The present invention provides an electromagnetic tracking system that includes a field generator and a field sensor arranged to generate and detect, respectively, an electromagnetic field. Both the transmitter and receiver coils are connected to signal conditioning and processing circuitry to provide outputs indicative of the coil signals. A processor operates on the signals to determine the coordinates of the sensing assembly relative to the generator assembly. The signal processor produces ratiometric outputs, and applies a mutual inductance model to solve for position/orientation coordinates. In some embodiments, a disturber in the form of a conductive ring or a sheath is disposed about an interfering piece of equipment to moderate and standardize disturbances due to eddy currents. | 1. An intra-operative imaging and tracking system for guiding a surgical tool during a surgical procedure performed on a patient, comprising: a fluoroscope having an x-ray source and an imaging assembly, said source and imaging assembly being movable about the patient to generate a plurality of two-dimensional x-ray images of the patient from different views;a magnetic tracker having a field generating unit driven by a drive signal to generate an electromagnetic field in an area of interest and a sensing unit that generates a sensing signal in response to said field, one of said units being secured against movement relative to said imaging assembly and the other unit being affixed to the surgical tool;a magnetic field distorter secured to said imaging assembly, a signal measurement circuit for measuring said drive and sensing signals to generate measured signal data; anda processor operative on said x-ray images and said measure signals, said processor modeling said field distorter and said generating and sensing units to derive modeled signal data and fitting said modeled signal data to said measured signal data to determine relative coordinates and orientations of said generating and sensing units, said processor further utilizing said x-ray images and said relative coordinates and orientations to determine position of the tool relative to the patient. 2. The system of claim 1, wherein said sensing unit is in a fixed position relative to said magnetic field distorter. 3. The system of claim 1, wherein said magnetic field distorter surrounds said imaging assembly. 4. An intra-operative imaging and tracking system for guiding a surgical tool during a surgical procedure performed on a patient, comprising: a fluoroscope having an x-ray source and a detector, said x-ray source and detector being movable relative to the patient so as to generate a plurality of two-dimensional x-ray images of the patient from different views;a magnetic tracker having a magnetic field generating unit driven by a drive signal to generate an electromagnetic field in an area of interest and a magnetic field sensing unit generating a sensing signal responsive to the electromagnetic field, one of said units being secured against motion relative to the detector and the other unit being affixed to the surgical tool;signal measurement circuitry electrically coupled to the tracker to measure said drive and sensing signals to form a matrix representing mutual inductance between said generating and sensing units, a processor operative with said mutual inductance matrix and said x-ray images to determine coordinates of the unit affixed to the surgical tool and position of the surgical tool relative to the patient; anda magnetic field distorter secured to said imaging assembly, a signal measurement circuit for measuring said drive and sensing signals to generate measured signal data. 5. An intra-operative imaging and tracking system for guiding a surgical tool during a surgical... | GE Medical Systems Global Technology Company LLC | A61B005/05
US07853313 | B2 | 2010-12-14 | Apparatus and method for radiosurgery | 2006-12-12 | A method and system is presented for treating moving target regions in a patient's anatomy by creating radiosurgical lesions. The method includes determining a pulsating motion of a patient separately from a determining of a respiratory motion, and directing a radiosurgical beam, from a radiosurgical beam source, to a target in the patient based on the determining of the pulsating motion. Directing the radiosurgical beam to the target may include creating a lesion in the heart to inhibit atrial fibrillation. The method may further include determining the respiratory motion of the patient, and compensating for movement of the target, due to the respiratory motion and the pulsating motion of the patient, in the directing of the radiosurgical beam based on the determining of the respiratory motion and the determining of the pulsating motion. | 1. A method, comprising: identifying a target in a patient;determining a respiratory motion of the target in the patient;determining a pulsating motion of the target in the patient separately from the determining of the respiratory motion, wherein determining the respiratory motion comprises: detecting a first motion of the target in the patient caused by a combined respiratory and pulsating motion of the patient while the patient breathes during a first period using one or more radiographic images of the target generated during the first period;detecting a second motion of the target in the patient caused substantially only by the pulsating motion while the patient holds a breath during a second period using one or more radiographic images of the target generated during the second period; andcalculating the respiratory motion of the target in the patient using the detected first and second motions; anddirecting a radiosurgical beam, from a radiosurgical beam source, to the target in the patient based on the determining of the pulsating motion and based on the determining of the respiratory motion. 2. The method of claim 1, wherein the target is a heart and the pulsation motion is due to a heartbeat of the heart, and wherein directing the radiosurgical beam to the target comprises creating a lesion in the heart to inhibit atrial fibrillation. 3. The method of claim 1, wherein said directing the radiosurgical beam comprises: determining a relative position of the radiosurgical beam source and the target based on the determining of the respiratory motion and the determining of the pulsating motion; andcompensating for movement of the target, due to the respiratory motion and the pulsating motion of the patient, in the directing of the radiosurgical beam based on the determining of the respiratory motion and the determining of the pulsating motion. 4. The method of claim 3, wherein determining the respiratory motion comprises detecting a respiratory cycle and determining the pulsating motion comprises detecting a pulsating cycle, and wherein the pulsating cycle is detected separately from the detecting of the respiratory cycle. 5. The method of claim 3, wherein detecting the first motion of the target in the patient caused by the combined respiratory and pulsating motion comprises: generating the one or more radiographic images of the target during the first period; andcombining information of the pulsating cycle and the respiratory cycle with the one or more radiographic images of the target during the first period to form the first motion. 6. The method of claim 5, wherein detecting the second motion of the target in the patient caused substantially only by the pulsating motion comprises: generating the one or more radiographic images of the target while the patient holds the breath during the second period; andcombining information of the pulsating cycle with the one or more radiographic images of the target during the second period to form ... | Accuray Incorporated | A61B005/05 G21K005/10
US07857771 | B2 | 2010-12-28 | Method and system for the derivation of human gait characteristics and detecting falls passively from floor vibrations | 2004-03-26 | The gait monitor system and method provides various basic gait parameters including step count, cadence, and step duration, in addition to its ability to distinguish between normal, limping and shuffling gait modes, as well as determine falls. Moreover, this gait monitor may be provided with additional sensors, e.g. beam break at the beginning and end of a corridor to estimate average walking velocity (with the distance between the beams known or determined); this enables the calculation of additional gait characteristics such as average step length and average stride length. These parameters can additionally be used to detect various gait anomalies and other diagnostic information. | 1. A gait monitoring system for monitoring gait characteristics of a subject, said system comprising: a sensor module configured to touch an upper surface of a floor, the sensor module configured to while touching the upper surface of the floor detect at least one of floor acceleration, floor vibration, and floor deflection and to provide at least one of an acceleration, vibration, and deflection signal, wherein the subject can walk on the upper surface of the floor in proximity to said sensor module; anda processor that is configured to analyze the acceleration, vibration, and deflection signal and to determine gait characteristics based on the signal,wherein, the processor is configured to determine the gait characteristics based on the at least one signal from a single sensor module; andthe processor is configured to detect a fall of a human being by distinguishing between steps of the human being and the fall of the human being and comparing the data to a threshold. 2. The system of claim 1, further comprising: an output module for receiving data indicative of the gait characteristics. 3. The system of claim 2, wherein said output module comprises at least one of display, alarm, memory storage, communication device, printer, buzzer, PDA, lap top computer, computer, audio or visual alarm, and light. 4. The system of claim 3, wherein said communication device comprises at least one of modem, pager, network interface, Ethernet card, serial communications port, parallel communications port, telephone, and PCMCIA slot and card. 5. The system of claim 1, wherein said sensor module and processor module are in wireless communication. 6. The system of claim 5, wherein said wireless communication comprises at least one of RF link, an infrared, cellular phone link, optical and electromagnetic. 7. The system of claim 1, wherein said sensor module and processor module are in a hard wired communication. 8. The system of claim 7, wherein said hard wired communication comprises at least one of electronic, integrated circuit, electromagnetic, wire, cable, fiber optics, a phone line, twisted pair, and coaxial. 9. The system of claim 1, further comprising: a rate-of-travel detector to determine the rate-of-travel of the subject. 10. The system of claim 9, wherein said rate-of-travel detector comprises at least one of a plurality of beam breaks, floor switches, and door switches. 11. The system of claim 9, wherein said rate-of-travel detector comprises at least one of ultrasonic communication, IR communication, laser communication, ground radar communication, wide band radar communication, and doppler communication. 12. The system of claim 1, wherein said gait characteristics of the subject includes at least two of step count, pace, normal gait condition, limp, shuffle, and falls. 13. The system of claim 1, further comprising an archival storage module. 14. The system of claim 13, wherein the processor module is configured to perform at least two of... | University of Virginia Patent Foundation | A61B005/103 A61B005/117 G01C022/00 G01B005/02 G01B005/14 G01B007/02 G01B007/14 G01B011/02 G01B011/14 G01B013/02 G01B021/02 G08B023/00
US07826882 | B2 | 2010-11-02 | Electrode lead set for measuring physiologic information | 2007-04-03 | An electrode lead set is provided for electrical connection to a body. The electrode lead set includes a flexible ribbonized cable core extending between a proximal end portion and a distal end portion. The distal end portion includes a plurality of branch end portions that are each configured to hold an electrode. The flexible ribbonized cable core is separable into a plurality of branches that are each joined to adjacent branches by a separable interface prior to separation. Each of the plurality of branches includes a corresponding one of the branch end portions. Each of the plurality of branches includes a generally planar ribbon conductor extending along the corresponding branch from the proximal end portion of the ribbonized cable core to the corresponding branch end portion. | 1. An electrode lead set for electrical connection to a body, said electrode lead set comprising: a flexible ribbonized cable core extending between a proximal end portion and a distal end portion, the distal end portion comprising a plurality of branch end portions that are each configured to hold an electrode, the flexible ribbonized cable core being separable into a plurality of branches that are each joined to adjacent branches by a separable interface prior to separation, each of the plurality of branches comprising a corresponding one of the branch end portions, each branch end portion comprising an electrode housing that is configured to hold the corresponding electrode, and each of the plurality of branches comprising a generally planar ribbon conductor extending along the corresponding branch from the proximal end portion of the ribbonized cable core to the corresponding branch end portion, each generally planar ribbon conductor being at least partially surrounded by an insulating material,the electrode housings each comprising a first opening for receiving a distal end portion of the corresponding generally planar ribbon conductor and a second opening for receiving an electrical contact of the corresponding electrode, the first and second openings being relatively arranged such that the distal end portion of the corresponding generally planar ribbon conductor is configured to electrically connect to the electrical contact of the corresponding electrode. 2. The electrode lead set according to claim 1, wherein each of the electrode housings comprises a cover configured to substantially cover the first opening. 3. The electrode lead set according to claim 1, wherein each of the electrode housings comprises a cover configured to substantially cover the corresponding first opening, and each of the covers comprises an extension configured to engage the distal end portion of the corresponding generally planar ribbon conductor to facilitate electrical connection between the corresponding generally planar ribbon conductor and the electrical contact of the corresponding electrode. 4. The electrode lead set according to claim 1, wherein the electrode housings each comprise a first opening for connecting with an electrical contact of the corresponding electrode with a snap-fit connection. 5. The electrode lead set according to claim 1, wherein each of the generally planar ribbon conductors is shielded. 6. The electrode lead set according to claim 1, wherein an electrically conductive material at least partially surrounds the insulating material, and an insulating jacket at least partially surrounds the electrically conductive material. 7. The electrode lead set according to claim 1, wherein prior to separation the plurality of branches are configured such that the branch end portions are maintained in an array at the distal end portion of the substrate core, and each branch end portion within the array is located at about the same distance fr... | Tyco Electronics Corporation | A61B005/0408
US07835892 | B2 | 2010-11-16 | Ultrasound simulation apparatus and method | 2004-09-28 | A method includes receiving data values associated with one of a position and orientation of a simulated scanner relative to an object. Image values are calculated, substantially in real-time, based on the data values. A simulated ultrasound image is rendered in a graphical display based on the image values. The simulated ultrasound image is representative of an interior or a simulated interior of the object on the ultrasound scan plane. | 1. A method comprising: placing a first physical object and a simulated scanner in contact with one another, the first physical object representing an object to be a subject of a simulated scan during an ultrasound simulation, the contact between the first physical object and simulated scanner established at least once during the simulation, wherein the simulated scanner is a second physical object;during the simulation, rendering a simulated ultrasound image of the first physical object based on a position of the simulated scanner during the simulation without the simulated scanner scanning the first physical object;while the first physical object and simulated scanner are in contact, detecting a change in position of the simulated scanner relative to a reference point using at least one sensing device and receiving a plurality of data values, wherein the change in position of the simulated scanner is caused by a movement of the first physical object occurring during the simulation;calculating, substantially in real-time, a plurality of image values based on the plurality of data values; andupdating based on the plurality of image values at least a portion of the simulated ultrasound image of the first physical object, the plurality of data values associated with a polygonal model of the first physical object. 2. The method of claim 1, wherein the updating includes rendering at least a portion of a simulated ultrasound image via a stencil buffer algorithm. 3. The method of claim 1, wherein one of the position and orientation of the simulated scanner relative to the first physical object defines a scan plane and a rendering direction. 4. The method of claim 2, wherein the rendering the simulated ultrasound image includes rendering a first set of data values from the plurality of data values to a stencil buffer of a graphics processor, and rendering a second set of data values from the plurality of data values to a frame buffer of the graphics processor, the first set of data values and the second set of data values being associated with a polygonal model of the first physical object. 5. A method comprising: placing a first physical object and a simulated scanner in contact with one another, the first physical object representing an object to be a subject of a simulated scan during an ultrasound simulation, the contact between the first physical object and simulated scanner established at least once during the simulation and wherein the simulated scanner is a second physical object;during the simulation, detecting a change in position of a simulated scanner using at least one sensing device without the simulated scanner scanning the first physical object;receiving a first set of data values and a second set of data values, wherein at least one of the first set and second set of data values indicates the change in position of the simulated scanner relative to a reference point, the change in position being caused by movement of the first physic... | Immersion Medical, Inc. | G06G007/48 G06G007/60 G06G007/62 G09B023/28 G09B023/32 A61B005/05 A61B008/00
US07857767 | B2 | 2010-12-28 | Lumen-traveling device | 2006-12-21 | Lumen-traveling devices and associated methods and systems are described. Lumen-traveling devices capable of traveling within a body lumen may include a propelling mechanism to produce movement of the lumen-traveling device within the lumen, as well as additional components such as a sensor, an active portion, and/or control circuitry. In some embodiments, a sensor may be used to detect a local condition, and an action may be performed within the body lumen. Actions that may be performed include, but are not limited to, transmitting information, releasing a material within the lumen, performing a surgical step, or collecting a sample, among others. | 1. A lumen-traveling device, comprising: a motion-arresting portion;a fluid-contacting portion configured to contact fluid within the body lumen and to at least intermittently permit flow of fluid through the body lumen;a propelling mechanism configured to produce movement of the lumen-traveling device through a body lumen in which the lumen-traveling device is deployed;motion control circuitry carried at least in part by said lumen-traveling device and configured to control the propelling mechanism to control movement of the lumen-traveling device through the body lumen;a sensor configured to detect a condition of interest in the body lumen and to generate a sense signal indicating detection of the condition of interest;response initiation circuitry operatively connected to the sensor and configured to generate a response initiation signal upon receipt of the sense signal indicating detection of the condition of interest in the body lumen; andan active portion operatively connected to the response initiation circuitry and configured to produce a response upon receipt of the response initiation signal. 2. The device of claim 1, wherein the motion control circuitry is operatively connected to the sensor and configured to control the propelling mechanism at least in part in response to receipt of the sense signal indicating detection of the condition of interest in the body lumen. 3. The device of claim 1, wherein the lumen-traveling device includes a structural element carrying at least one of the propelling mechanism, motion control circuitry, sensor, response initiation circuitry and active portion. 4. The device of claim 3, wherein the structural element includes at least one of an adjustable diameter, a substantially tubular structure, multiple lumens in fluid communication with the body lumen, a capsule-like structure, a self-expanding material, a resilient material, a mesh-like material, a slotted structure, a bioactive component, or a biocompatible material. 5. The device of claim 1, including a power source configured to provide power to at least one of the propelling mechanism, the motion control circuitry, the sensor, the response initiation circuitry, and the active portion. 6. The device of claim 5, wherein the power source includes at least one of a battery, a microbattery, a fuel cell, a biofuel cell, a nuclear battery, or an energy-scavenging device, an electrical power source located remote from the lumen-traveling device and connected to the lumen-traveling device by a wire, an optical power source located remote from the lumen-traveling device and connected to the lumen-traveling device by an optical fiber, or a power receiver configured to receive power from a remote source. 7. The device of claim 6, wherein the power source includes an energy-scavenging device including a pressure-rectifying mechanism, an acceleration-rectifying mechanism, or a flow-rectifying mechanism. 8. The device of claim 1, including: a steering me... | Invention Science Fund I, LLC | A61B005/02 A61B005/00 A61B005/07
US07824612 | B2 | 2010-11-02 | Bodily fluid analyzer, and system including same and method for programming same | 2007-04-24 | A bodily fluid analyzer, system and method for programming the same includes a sensor for sensing at least one analyte in a patient; a display; a data reader unit for reading information from a data storage unit that contains stored information concerning a particular drug being taken by or course of treatment for the patient; and a processor for setting the at least one threshold value for at least one analyte to be sensed by the sensing unit based on the information read by the data reader from the data storage unit, for processing the information concerning the analyte and for sending the processed information to the display, wherein the threshold value is associated with the particular drug taken by or course of treatment for the patient, the threshold value being one beyond which the display will display an alert. | 1. A method for programming a bodily fluid analyzer, comprising: selecting by a prescribing physician or a drug company at least one threshold value of at least one analyte to be sensed by the bodily fluid analyzer, the at least one threshold value of the at least one analyte being associated with a particular drug being or to be taken by the patient or course of treatment for the patientproviding a data storage unit separately from the bodily fluid analyzer and containing stored information including the at least one threshold value of at least one analyte to be sensed by the bodily fluid analyzer selected by the prescribing physician or the drug company;reading the stored information stored on the data storage unit into a data reader associated with a bodily fluid analyzer;setting the bodily fluid analyzer with the at least one threshold value for the at least one analyte to be sensed by the bodily fluid analyzer with the information read by the data reader from the data storage unit, wherein the at least one threshold value is a value such that a sensed analyte level beyond the value will cause the display to display an alert;sensing the analyte level; anddisplaying an alert if the sensed analyte level is beyond the threshold value. 2. The method according to claim 1, wherein the data storage unit is provided with a drug container. 3. The method according to claim 2, wherein the data reader is a bar code reader and the data storage unit is a bar code. 4. The method according to claim 2, wherein the data reader is a radio frequency receiver and the data storage unit is a radio frequency identification tag. 5. The method according to claim 4, wherein the drug is provided in a closed container and the method further comprises opening the container and activating the radio frequency identification tag upon opening of the container. 6. The method according to claim 2, wherein the data reader is a magnetic stripe reader and the data storage unit is a magnetic stripe. 7. The method according to claim 1, wherein the at least one threshold value of the at least one analyte to be sensed by the bodily fluid analyzer is selected by a prescribing physician for the patient or for a class of patients. 8. The method according to claim 1, wherein the at least one threshold value of the at least one analyte to be sensed by the bodily fluid analyzer is selected by a drug company for the patient or for a class of patients. 9. A system for monitoring a patient, comprising a bodily fluid analyzer and a data storage unit provided seperately from the bodily fluid analyzer and containing stored information set by a prescribing physician or a drug company and including at least one threshold value of at least one analyte to be sensed by the bodily fluid analyzer, the at least one threshold value of the at least one analyte being associated with a particular drug being or to be taken by the patient or course of treatment for the patient, wherein the bodily fluid... | - | G01N021/00 G01N035/02 A61B005/00 G01N015/06
US07826880 | B2 | 2010-11-02 | Electrochemical sensor for in-vivo or ex-vivio measurements of the carbon dioxide partial pressure of living tissue | 2005-07-18 | A physiological sensing device for the measurement of pCO2 includes a closed chamber bounded, at least partially, by a carbon dioxide permeate e membrane (12). There are two electrodes (10) within the chamber. The chamber contains a substantially electrolyte-free liquid in contact with electrodes (10) and the membrane (12). The liquid contains a non-ionic excipient in order to prevent egress of water due to an osmotic gradient across the membrane (12) in use. | 1. A physiological sensing device for the measurement of pCO2, the device comprising: a closed chamber bounded, at least partially, by a carbon dioxide permeable membrane; and at least two electrodes within the chamber, wherein the chamber contains a mixture in contact with the electrodes and the membrane and inclusive of at least a first part and a second part, the first part being a substantially electrolyte-free liquid and the second part being a non-ionic excipient, wherein the non-ionic excipient is present in the electrolyte-free liquid in at least isotonic concentration, such that the osmolality of the fluid in the chamber is at least equivalent to that of 0.9% w/v sodium chloride solution. 2. A physiological sensing device as claimed in claim 1, wherein the non-ionic excipient is propylene glycol. 3. A sensor as claimed in claim 1, wherein the electrolyte-free liquid is deionised or distilled water. 4. A sensor as claimed in claim 2, wherein the electrolyte-free liquid is deionised or distilled water. | Alertis Medical AS | A61B005/05
US07841986 | B2 | 2010-11-30 | Methods and apparatus of three dimensional cardiac electrophysiological imaging | 2007-05-10 | Generating a three-dimensional representation of cardiac electrical activity within a subject's heart includes collecting, from a catheter having a distal portion that includes sensors that sense electrical activity and which catheter distal portion is positioned within a cardiac chamber of a subject, data for electrical activity sensed by the catheter distal portion while positioned within the chamber of the heart. The method also includes executing an imaging algorithm on the collected electrical activity data to generate a three-dimensional representation of cardiac electrical activity within a volume of the subject including all or a portion of the subject's heart. The method further includes displaying the imaged three-dimensional cardiac electrical activity. | 1. A method of generating a three-dimensional representation of cardiac electrical activity for a subject, the method comprising: providing a three-dimensional cardiac model comprising a structure of a heart and finite elements on which cardiac electrophysiological properties are considered;recording, from a catheter having a distal portion comprising multiple electrodes, sensed intra-cardiac electrical potential data for electrical activity sensed by the multiple electrodes while the catheter distal portion is positioned within one or more chambers of the subject's heart;using the recorded intra-cardiac electrical potential data to estimate a value for an electrophysiological property for each of various finite elements throughout the three-dimensional cardiac model; andgenerating on a visual display device, and using the calculated values for the electrophysiological property for the various finite elements, a visual representation of estimated cardiac electrical activity for the subject. 2. The method of claim 1, wherein the electrophysiological property is a source model that predicts cardiac electrical activity, and estimates the calculated values for the electrophysiological property by minimizing a difference between the catheter recorded intra-cardiac electrical potential data and the source model predicted cardiac electrical activity. 3. The method of claim 2, wherein the source model is a heart biophysical model comprising three-dimensional equivalent source representations, including one or more of current density, transmembrane potential, local field potential, activation time, and repolarization time. 4. The method of claim 3, wherein the three-dimensional equivalent source representations include estimated equivalent current density representations, and wherein the method further comprises deriving a repolarization dispersion from the estimated equivalent current density representations to aid in diagnosis and management of cardiovascular diseases. 5. The method of claim 1, wherein the estimated cardiac electrical activity comprises cardiac activation. 6. The method of claim 1, wherein the estimated cardiac electrical activity comprises cardiac repolarization. 7. The method of claim 1, wherein the generated visual representation of the estimated cardiac electrical activity is used to guide ablation of cardiac arrhythmias. 8. The method of claim 1, wherein the generated visual representation of the estimated cardiac electrical activity is used to guide catheter ablation of cardiac arrhythmias. 9. The method of claim 1, wherein the generated visual representation of the estimated cardiac electrical activity is used to guide cardiac synchronized therapy. 10. The method of claim 1, wherein the visual representation of the estimated cardiac electrical activity includes representations of cardiac electrical activity both within the three-dimensional volume of the heart and over a surface of the heart. 11. The method of claim 1, ... | Regents of the University of Minnesota | A61B005/0476
US08498696 | B2 | 2013-07-30 | Health monitoring device and human electric signal processing method | 2010-06-08 | A health monitoring device is provided and includes a detecting unit for detecting a first electric signal from a human body and a processing unit for performing the following steps: receiving the first electric signal; performing wavelet transform on the first electric signal, to filter frequency bands thereof which include undesired noises, and then perform a reverse wavelet transform to obtain a second electric signal; combining frequency bands of the first electric signal which include primary characteristics and then perform a reverse wavelet transform on the combined frequency bands to obtain a fourth electric signal; and comparing the fourth electric signal with a plurality of electric signal patterns of a database to determine whether the fourth electric signal is an electric signal for healthiness. | 1. A health monitoring device, comprising: a detecting unit for detecting a first electric signal from a human body; anda processing unit for performing the following steps: receiving the first electric signal;performing wavelet transform on the first electric signal, to filter frequency bands thereof which include undesired noises, and then perform reverse wavelet transform to obtain a second electric signal;combining frequency bands of the first electric signal which include desired primary characteristics and then performing reverse wavelet transform on the combined frequency bands to obtain a third electric signal;obtaining period information of the second electric signal from the third electric signal, and then re-sampling the second electric signal according to the period information to obtain a fourth electric signal; andcomparing the fourth electric signal with a plurality of electric signal patterns in a database to determine whether the fourth electric signal is an electric signal for healthiness. 2. The health monitoring device as claimed in claim 1, wherein the processing unit further performs the following steps: performing spectral estimation on the fourth electric signal to obtain a plurality of characteristic coefficients;comparing the characteristic coefficients with a plurality of characteristic coefficient patterns to determine whether the fourth electric signal is an electric signal for healthiness. 3. The health monitoring device as claimed in claim 1, wherein the processing unit further performs the following steps: performing normalization on one of the first, second and fourth electric signals so that the magnitude range of the electric signal is consistent with the magnitude range of the electric signal patterns in the database. 4. The health monitoring device as claimed in claim 1, wherein the frequency bands which include undesired noises in the first electric signal are low frequency drift bands and high frequency noise bands. 5. The health monitoring device as claimed in claim 1, wherein the frequency bands of the second bands which include primary characteristics comprise one of a QRS band, P band, T band, or U band. 6. The health monitoring device as claimed in claim 1, wherein the first electric signal is one of an electrocardiogram signal, an electroencephalogram signal, an electromyogram signal or a blood sugar concentration signal. 7. The health monitoring device as claimed in claim 2, wherein spectral estimation is performed by using Autoregressive Model (AR Model). 8. The health monitoring device as claimed in claim 2, wherein step of comparing the fourth electric signal with a plurality of electric signal patterns in a database and step of comparing the characteristic coefficients with a plurality of characteristic coefficient patterns are performed by a Support Vector Machine (SVM). 9. The health monitoring device as claimed in claim 8, wherein the SVM is performed by using a mapping function to map ... | National Taiwan University | A61B005/04
US08670810 | B2 | 2014-03-11 | Regional oximetry analog front end | 2010-06-04 | An apparatus includes a first amplifier having a first input coupled to a first optical detector. The first amplifier includes a first output corresponding to a logarithm of the first input. The apparatus includes a second amplifier having a second input coupled to a second optical detector and having a second output corresponding to a logarithm of the second input. The apparatus includes a differential amplifier configured to amplify a difference between the first output and the second output. | 1. An apparatus comprising: a first amplifier having a first input coupled to a first optical detector and having a first output corresponding to a logarithm of the first input wherein the first input corresponds to light intensity along a first optical path through a tissue, the first optical detector on the first optical path;a second amplifier having a second input coupled to a second optical detector and having a second output corresponding to a logarithm of the second input wherein the second input corresponds to light intensity along a second optical path through the tissue, the second optical detector on the second optical path and wherein the first optical path differs from the second optical path, wherein the light intensity corresponds to concentration of an absorbing species in the tissue; anda differential amplifier having a differential output, the differential amplifier configured to amplify a difference between the first output and the second output and generate an output signal, the output signal provided at the differential output and the output signal corresponding to oxygenation of a region of the tissue. 2. The apparatus of claim 1 wherein the first optical detector is coupled to an inverting input of the first amplifier. 3. The apparatus of claim 1 wherein the second optical detector is coupled to an inverting input of the second amplifier. 4. The apparatus of claim 1 wherein at least one of the first amplifier and the second amplifier includes an operational amplifier. 5. The apparatus of claim 1 wherein at least one of the first amplifier and the second amplifier includes a feedback path having a transistor junction. 6. The apparatus of claim 1 wherein the first optical detector and the second optical detector are responsive to light intensity. 7. The apparatus of claim 1 wherein the first optical detector and the second optical detector are configured to provide a signal corresponding to light absorbance. 8. The apparatus of claim 1 wherein the first amplifier matches the second amplifier. 9. The apparatus of claim 1 wherein the first amplifier and the second amplifier are in a common package. 10. The apparatus of claim 1 wherein the differential output is coupled to a module, the module configured to determine a measure of a tissue characteristic. 11. The apparatus of claim 10 wherein the module is configured to determine a characteristic of blood in the tissue. 12. The apparatus of claim 10 wherein the module is configured to determine a measure of a chromophore. 13. The apparatus of claim 10 wherein the module is configured to determine a measure of hemoglobin. 14. The apparatus of claim 10 wherein the module is configured to determine a hemoglobin index corresponding to a ratio of tissue and blood. 15. The apparatus of claim 10 wherein the module includes a display. 16. A method comprising: receiving a first signal from a first optical detector, the first signal corresponding to a first optical parameter... | Nonin Medical, Inc. | A61B005/05
US09005233 | B2 | 2015-04-14 | Contact activated lancet device | 2010-06-14 | A lancet device including a housing and a lancet structure having a puncturing element. The lancet structure disposed within the housing and adapted for movement between a retaining or pre-actuated position wherein the puncturing element is retained within the housing, and a puncturing position wherein the puncturing element extends through a forward end of the housing. The lancet device includes a drive spring disposed within the housing for biasing the lancet structure toward the puncturing position, and a retaining hub retaining the lancet structure in the retracted position against the bias of the drive spring. The retaining hub includes a pivotal lever in interference engagement with the lancet structure. An actuator within the housing pivots the lever, thereby moving the lancet structure toward the rearward end of the housing to at least partially compress the drive spring, and releasing the lever from interference engagement with the lancet structure. | 1. A lancet device comprising: a housing;a lancet structure comprising a puncturing element, the lancet structure disposed within the housing and adapted for movement between a pre-actuated position wherein the puncturing element is retained within the housing and a puncturing position wherein the puncturing element extends through a forward end of the housing;a drive spring biasing the lancet structure toward the puncturing position;a retaining hub retaining the lancet structure in the pre-actuated position against the bias of the drive spring, the retaining hub comprising a lever pivotal about a fulcrum, the lever comprising a class 1 lever, the lever providing interference engagement with the lancet structure, the lever comprising a first lever element and a second lever element spaced from the first lever element, the first lever element and the second lever element disposed on opposing sides of the lancet structure; andan actuator adapted to pivot the lever about the fulcrum, thereby releasing the lever from interference engagement with the lancet structure, permitting the drive spring to drive the lancet structure toward the puncturing position. 2. The lancet device of claim 1, wherein the actuator comprises an actuation element extending through and into the housing, wherein movement of the actuation element with respect to the housing causes the actuation element to pivot the lever about the fulcrum. 3. The lancet device of claim 2, wherein the actuation element extends through a rearward end of the housing and is movable with respect to the housing to cause the actuation element to pivot the lever about the fulcrum. 4. The lancet device of claim 1, wherein the actuator comprises an internal contact integrally formed within the housing, wherein movement of the housing toward the retaining hub causes the internal contact within the housing to pivot the lever about the fulcrum. 5. The lancet device of claim 4, further comprising a shield extending through the forward end of the housing and movable with respect to the housing, with the retaining hub adjacent a rearward end of the shield. 6. The lancet device of claim 5, wherein the lancet structure is retained by the retaining hub at the rearward end of the shield, and wherein movement of the rearward end of the shield toward the rearward end of the housing causes the actuator to pivot the lever of the retaining hub. 7. The lancet device of claim 5, wherein the lancet structure and the shield include corresponding guiding surfaces for guiding the lancet structure through the shield. 8. The lancet device of claim 5, wherein the shield and the housing include corresponding guiding surfaces for guiding the shield through the housing. 9. The lancet device of claim 5, further comprising a retraction spring for retracting the lancet structure within the shield after the drive spring drives the lancet structure through the shield toward the puncturing position. 10. The lancet device of cla... | Becton, Dickinson and Company | A61B017/14 A61B005/151 A61B005/15 A61B017/00
US08821414 | B2 | 2014-09-02 | Device and method for predicting a loss of control over a muscle | 2008-06-16 | A device for predicting a loss of control over a muscle of a human being has a detector for detecting a sound of the muscle, an acquirer for acquiring an acceleration of the human being and an evaluator for evaluating the sound and the acceleration to determine an imminent loss of control over the muscle from the sound and the acceleration. | 1. A device for predicting a loss of control over a muscle of a human being, comprising: a detector for detecting a muscle sound;an acquirer for acquiring an acceleration of the human being; andan evaluator for evaluating the muscle sound and the acceleration and determine an imminent loss of control over the muscle from the muscle sound and the acceleration. 2. The device in accordance with claim 1, wherein the detector for detecting is implemented to detect the muscle sound of a supportive muscle or a leg muscle, and wherein the evaluator for evaluating is implemented to determine an imminent loss of control of the supportive muscle or the leg muscle. 3. The device in accordance with claim 1, wherein the evaluator for evaluating is implemented to output, in the case of determining an imminent loss of control over the muscle, a warning signal detectable by the human being. 4. The device in accordance with claim 3, wherein the warning signal comprises a photo signal and/or an acoustic signal and/or a tangible stimulus. 5. The device in accordance with claim 3, wherein the evaluator for evaluating is implemented to transfer the warning signal to a predetermined notification institution. 6. The device in accordance with claim 1, wherein the evaluator for evaluating is implemented to acquire the muscle sound and the acceleration at a sequence of time instances and combine same to form a data set. 7. The device in accordance with claim 6, wherein the evaluator for evaluating is implemented to compare the data set to a model data set and to determine a percentage matching value to draw conclusions about a current movement of the human being, the model data set corresponding to a data set which may be generated by a certain movement of the human being. 8. The device in accordance with claim 7, wherein the model data set comprises a parameter, and wherein the evaluator for evaluating is implemented to vary the parameter in order to maximize a percentage matching value. 9. The device in accordance with claim 1, wherein the evaluator for evaluating is implemented to store data of the detector and/or data of the acquirer. 10. The device in accordance with claim 1, wherein the evaluator for evaluating comprises a transferrer and the transferrer is implemented to acquire data from the detector and/or data from the acquirer and transfer same to a central unit. 11. The device in accordance with claim 10, wherein the transferrer acquires data in a wireless manner or transfers data in a wireless manner. 12. The device in accordance with claim 1, wherein the detector for detecting and the acquirer for acquiring and the evaluator for evaluating are housed together in a combined sensor. 13. The device in accordance with claim 12, wherein the combined sensor comprises an attacher for attaching to a thigh or leg of the human being. 14. The device in accordance with claim 1, wherein the detector for detecting comprises an attacher for attaching to a thigh... | Fraunhofer-Gesellschaft zur Foerderung der Angewandten Forschung E.V. Friedrich-Alexander-Universitaet Erlangen-Nuernberg | A61B007/00 A61B005/103 A61B005/117
US09072775 | B2 | 2015-07-07 | Fluorescent emulsions for optical imaging | 2008-02-14 | The invention relates to a fluorescent emulsion, to a diagnostic reagent containing the same, and to the use thereof in the preparation of a diagnostic reagent for in vivo fluorescence imaging. | 1. A metastable fluorescent emulsion of oil-in-water type that is stable at pH 7.3, comprising at least one aqueous continuous phase andat least one oily dispersed phase;wherein the dispersed phase comprises droplets of: at least one liquid biocompatible oil composed of at least one plant or animal oil andat least one crystallizable oil having the following composition: C8: 0.1 to 0.9%,C10: 0.1 to 0.9%,C12: 25 to 50%,C14: 10 to 24.9%,C16: 10 to 24.9%,C18: 10 to 24.9%,wherein said droplets have an average diameter greater than or equal to 10 nm and less than or equal to 200 nm and contain at least one lipophilic fluorophore which absorbs and emits at a wavelength between 640 and 900 nm, andwherein said droplets are stabilized by a surfactant layer located at the periphery of said droplets, wherein said layer comprises, as a mixture, at least one amphiphilic surfactant and at least one stealth cosurfactant. 2. The emulsion as claimed in claim 1, wherein the average diameter of the droplets is between 10 and 80 nm, limits included. 3. The emulsion as claimed in claim 1, wherein the plant oil is soybean oil. 4. The emulsion as claimed in claim 1, wherein the ratio by weight of liquid biocompatible oil/crystallizable oil ranges between 10/90 and 90/10, limits included. 5. The emulsion as claimed in claim 1, wherein the amphiphilic surfactants are selected from the group consisting of compounds of which the lipophilic part comprises a linear or branched, saturated or unsaturated chain containing from 8 to 30 carbon atoms. 6. The emulsion as claimed in claim 1, wherein the amphiphilic surfactants are selected from the group consisting of phospholipids, cholesterols, lysolipids, sphingomyelins, tocopherols, glucolipids, stearylamines, cardiolipins of natural or synthetic origin, molecules composed of a fatty acid coupled to a hydrophilic group by an ether or ester function, polymerized lipids, lipids conjugated to short chains of polyethylene oxide (PEG), sugar esters, and mixtures thereof. 7. The emulsion as claimed in claim 6, wherein the amphiphilic surfactants are selected from the group consisting of soybean lecithin, phospholipids and cholesterol. 8. The emulsion as claimed in claim 1, wherein the lipophilic fluorophores are selected from the group consisting of fatty acid analogs, sphingolipids, steroids, lipopolysaccharides and phospholipids functionalized with a group which absorbs and emits between 640 and 900 nm, and amphiphilic derivatives thereof. 9. The emulsion as claimed in claim 8, wherein the fluorophores are amphiphilic derivatives of dialkylcarbocyanines. 10. The emulsion as claimed in claim 1, wherein the stealth cosurfactant(s) is (are) selected from the group consisting of amphiphilic molecule(s) of which the hydrophilic part is completely or partly composed of a polyethylene oxide (PEO) chain in which the number of PEO units ranges from 2 to 500, and polysaccharide(s). 11. The emulsion as claimed in claim 10, wherein the... | - | A61K049/00 B01F017/00 A61K009/107
US08838216 | B2 | 2014-09-16 | Method of and apparatus for generating a model of a cardiac surface having a plurality of images representing electrogram voltages | 2008-05-02 | A method of generating a model of a cardiac surface having a plurality of images representing electrogram voltages for a plurality of measured points within a heart comprises measuring an electrogram voltage at a plurality of points within a heart, generating a first model of a cardiac surface of the heart, generating an image representing each electrogram voltage, each image having a characteristic representative of the electrogram voltage, and generating a further model of a cardiac surface. The images representing the electrogram voltages protrude from the further model of the cardiac surface at points on the further model corresponding to the points at which the electrogram voltages were measured. There is also disclosed an apparatus for generating a model of a cardiac surface. | 1. A method of generating a model of a cardiac surface having a plurality of images representing electrogram voltages for a plurality of measured points within a heart comprising: internally measuring an electrogram voltage at a plurality of points within a heart; generating a first model of a cardiac surface of said heart; generating an image representing each electrogram voltage, each image having a characteristic representative of the electrogram voltage; and generating a further model of a cardiac surface, said images representing said electrogram voltages protruding from said further model of said cardiac surface at points on said further model corresponding to the points at which said electrogram voltages were measured. 2. A method according to claim 1, wherein the step of generating the first model of said cardiac surface comprises: (a) obtaining three-dimensional coordinates of the locations of said plurality of points at which said electrogram voltages were measured;(b) fitting a surface of a three-dimensional representation of a solid to said three-dimensional coordinates;(c) locating on the surface of said solid a plurality of points, one or more of said points corresponding to a respective one of said plurality of points at which said electrogram voltages were measured to generate associated surface points on said representation of said solid;(d) dividing said surface of said representation of said solid into a plurality of triangles, each triangle having an associated apex at which one of said plurality of said surface points is located;(e) transforming the surface of said three-dimensional representation of a solid into a two-dimensional plane around a first selected triangle of said plurality of triangles using a predetermined transformation process;(f) transforming the three-dimensional coordinates of the locations of the plurality of points at which said electrogram voltages were measured using said transformation process applied in step (e) to said surface;(g) connecting a number of said transformed points at which said electrogram voltages were measured to other of said transformed points at which said electrogram voltages were measured to form a further surface using a bicubic spline interpolation method;(h) sampling an area within said first selected triangle that was interpolated using the bicubic spline interpolation method, the samples being adjusted with a weighted average to obtain a shape representative of the surface of said triangle;(i) repeating steps (e), (f), (g) and (h) for each triangle of said plurality of triangles to obtain a plurality of three-dimensional surfaces representative of each of said plurality of triangles; and(j) transforming the three-dimensional surfaces into a three- dimensional model to form the first model of said cardiac surface. 3. A method according to claim 2, wherein said step of measuring an electrogram voltage at a plurality of points within a heart comprises measuring a maximum volt... | Imperial Innovations Limited | A61B005/04 G06F019/00 G06T017/20 A61B006/00 A61B005/042
US08430821 | B2 | 2013-04-30 | Blood pressure measuring apparatus | 2010-07-27 | A blood pressure measuring apparatus includes a cuff member (2) which can be attached to and detached from a blood pressure measurement portion, and a cuff main body (1) including a compression air bladder (8) for pressing the blood pressure measurement portion, a sub air bladder (7), and a pulse wave detection air bladder (5). After pressurization is performed by compression air bladder pressurizing/depressurizing means (22, 23) and sub air bladder pressurizing/depressurizing means (27, 26), depressurization is controlled at the same depressurization rate, thereby reducing the cuff-edge effect and obtaining a pulse wave signal. | 1. A blood pressure measuring apparatus comprising: a cuff member configured to be attached to and detached from a blood pressure measurement portion;a cuff main body including a compression air bladder which is formed on a side of said cuff member, which is in contact with the blood pressure measurement portion, and presses the blood pressure measurement portion,a sub air bladder which is formed on a side of said compression air bladder, which is in contact with the blood pressure measurement portion, and presses a heart side of the blood pressure measurement portion, anda pulse wave detection air bladder which is formed on the side of said compression air bladder, which is in contact with the blood pressure measurement portion, and detects a pulse wave downstream of a central portion of the blood pressure measurement portion;a compression air bladder pressurizing/depressurizing unit configured to pressurize/depressurize said compression air bladder and said pulse wave detection air bladder;a sub air bladder pressurizing/depressurizing unit configured to pressurize/depressurize said sub air bladder;a cuff pressure detecting unit configured to obtain a cuff pressure signal from said pulse wave detection air bladder;a sub air bladder pressure detecting unit configured to obtain a sub air bladder pressure signal from said sub air bladder;a pulse wave detecting unit configured to detect a pulse wave superposed on the cuff pressure signal, and obtaining a pulse wave signal;a blood pressure determining unit configured to determine a blood pressure value based on the cuff pressure signal and the pulse wave signal;a blood pressure display unit configured to display the blood pressure value;a first tube connected between said pulse wave detection air bladder and said cuff pressure detecting unit;a second tube connected between said compression air bladder and said compression air bladder pressurizing/depressurizing unit, and connected to said cuff pressure detecting unit via a fluid resistance;a third tube connected between said sub air bladder and said sub air bladder pressurizing/depressurizing unit, and connected to said sub air bladder and said sub air bladder pressurizing/depressurizing unit; anda control unit configured to control, after said compression air bladder and said sub air bladder are pressurized by said compression air bladder pressurizing/depressurizing unit and said sub air bladder pressurizing/depressurizing unit, said compression air bladder pressurizing/depressurizing unit and said sub air bladder pressurizing/depressurizing unit so as to depressurize said compression air bladder and said sub air bladder at the same depressurization rate, and obtain the pulse wave signal. 2. The blood pressure measuring apparatus according to claim 1, wherein said control unit performs control of: pressurizing said sub air bladder to a prescribed pressure by said sub air bladder pressurizing/depressurizing unit, and then stopping the pressurization... | Terumo Kabushiki Kaisha | A61B005/02
US08657295 | B2 | 2014-02-25 | Device and method for measuring reaction time | 2010-06-15 | A manual neurological testing device for measuring a subject's reaction time comprising: a releasable holding member operable to be grasped by the subject; a stimulus device coupled with the releasable holding member, the stimulus device outputting a perceivable stimulus; and a measurement device coupled with the releasable holding member, the measurement device measuring the subject's reaction time for grasping the releasable holding member relative to the perceivable stimulus. | 1. A manual neurological testing device for measuring a human subject's reaction time comprising: a) a releasable holding member operable to be released by a human tester and free fall in response to the force of gravity and be subsequently grasped by the human subject;b) a stimulus device fixedly coupled with said releasable holding member so as to free fall therewith in response to the force of gravity, said stimulus device operable to output a perceivable stimulus in response to free fall; andc) a measurement device fixedly coupled with said releasable holding member so as to free fall therewith in response to the force of gravity, said measurement device operable to measure the human subject's reaction time for grasping said releasable holding member as said releasable holding member, said stimulus device, and said measurement device collectively free fall relative to said perceivable stimulus. 2. The manual neurological testing device according to claim 1 wherein said releasable holding member is a rigid member. 3. The manual neurological testing device according to claim 1 wherein said releasable holding member is a flexible member. 4. The manual neurological testing device according to claim 1 wherein said releasable holding member is a telescoping member. 5. The manual neurological testing device according to claim 1 wherein said measurement device operable to determine a recognition reaction time response accuracy. 6. The manual neurological testing device according to claim 1 wherein said measurement device comprises an accelerometer operable to detect acceleration of at least one of said releasable holding member and said measurement device. 7. The manual neurological testing device according to claim 1 wherein said stimulus device comprises a visual cue device. 8. The manual neurological testing device according to claim 1 wherein said stimulus device comprises an audible cue device. 9. The manual neurological testing device according to claim 1 wherein said stimulus device operable to output a plurality of distinct cues for eliciting a corresponding unique response by the subject. 10. The manual neurological testing device according to claim 1 wherein said stimulus device operable to output a plurality of distinct cues for eliciting a corresponding unique response by the subject, a first of said plurality of cues indicating a catch command to the subject, a second of said plurality of cues indicating a no-catch command to the subject. 11. The manual neurological testing device according to claim 1, further comprising: a spacer system for positioning the subject's appendage in a pre-test position. 12. The manual neurological testing device according to claim 1, further comprising: a spacer system for positioning the subject's appendage in a pre-test position, said spacer system having a proximity sensor operable to detect positioning of the subject's appendage. 13. The manual neurological testing device according to claim ... | The Regents of The University of Michigan | A63B067/00 A63B071/00 A63F009/00 A61B019/00 A61B005/117 A61B005/103 G09B019/00
US09302069 | B2 | 2016-04-05 | Device and method for visual stimulation | 2010-09-03 | A device and method for providing stimulation signals that reset the phase of the neuronal activity of neurons in a patient's brain. The device includes a control unit; and a stimulation unit that has a plurality of stimulation elements, and each stimulation element generates visual stimulation signals that reset the phase of the neuronal activity of the neurons when the signals are taken up via an eye of a patient and transmitted to neurons that are exhibiting a pathologically synchronous and oscillatory neuronal activity. The control unit is further capable of actuating the stimulation unit such that the stimulation elements generate the visual stimulation signals with a time offset in respect to one another and/or with differing phase and/or with differing polarity. | 1. A device for stimulating subpopulations of a patient's neuron population, the device comprising: a stimulation unit including N stimulation elements, each configured to generate respective visual stimulation pulses at different spatial locations corresponding to respective prescribed locations in a visual field of the patient such that the visual stimulation pulses are transmitted to respective subpopulations of the neuron population via an optic nerve of the patient when the visual stimulation pulses are visually received by the patient; anda control unit configured to sequentially actuate each of the N stimulation elements of the stimulation unit to generate the visual stimulation pulses at a repeating periodic frequency fstim that is between 1 and 30 Hz, where N is an integer and where Tstim is equal to 1/fstim,wherein the control until actuates each of the N stimulation elements to generate the visual stimulation pulses for a time period between Tstim/2−Tstim/10 and Tstim/2+Tstim/10, andwherein the N stimulation elements are sequentially actuated to generate the visual stimulation pulses with a time offset that is approximately 1/(fstim×N) seconds between respective visual stimulation pulses in the sequence of the visual stimulation pulses generated by the N stimulation elements. 2. The device as claimed in claim 1, wherein the stimulation unit is a pair of transmission glasses having a pair of lenses. 3. The device as claimed in claim 2, wherein the lenses include a plurality of segments and the control unit is further configured to control the transmission properties of the segments. 4. The device as claimed in claim 1, wherein the stimulation unit is a pair of glasses including a plurality of light sources. 5. The device as claimed in claim 4, wherein the glasses have a transparent region. 6. The device as claimed in claim 4, wherein the glasses are opaque. 7. The device as claimed in claim 1, further comprising a measurement unit with at least one sensor configured to record measurement signals relating to physiological activity of the patient. 8. The device as claimed in claim 7, wherein the control unit is further configured to actuate the stimulation elements based on the measurement signals such that the stimulation elements convert the measurement signals into visual stimulation pulses. 9. The device as claimed in claim 7, wherein the control unit is further configured to process the measurement signals linearly or nonlinearly and to the processed measurement signals into control inputs of the stimulation elements. 10. The device as claimed in claim 7, wherein the control unit is further configured to determine, depending on the measurement signals, whether the stimulation elements generate the respective visual stimulation pulses. 11. The device as claimed in one of claim 7, wherein the at least one sensor is at least one of an EEG electrode, an MEG sensor, an accelerometer, an EMG electrode, a sensor configured to dete... | Forschungszentrum Juelich GmbH Universitaet zu Koeln | A61N001/18 A61M021/00 A61N005/06 A61N005/073
US20100312072 | A1 | 2010-12-09 | METHOD AND SYSTEM FOR DETECTING AND GRADING PROSTATE CANCER | 2009-01-01 | A method of estimating a grade of a prostate cancer from zinc data associated with the prostate, the zinc data being arranged gridwise in a plurality of picture-elements representing a zinc map of the prostate. The method comprises, clustering the zinc map according to zinc levels associated with the picture-elements, and estimating a cancer grade of at least one tissue region, based, at least in part, on zinc levels associated with a cluster of picture-elements representing the tissue region. | 1. A method of estimating a grade of a prostate cancer from zinc data associated with the prostate, the zinc data being arranged gridwise in a plurality of picture-elements representing a zinc map of the prostate, the method comprising: clustering the zinc map according to zinc levels associated with said picture-elements; andestimating a cancer grade of at least one tissue region, based, at least in part, on zinc levels associated with a cluster of picture-elements representing said tissue region. 2. A method of estimating a grade of a prostate cancer, comprising: recording zinc data from the prostate so as to generate a zinc map represented by a plurality of gridwise arranged picture-elements;clustering the zinc map according to zinc levels associated with said picture-elements; andestimating a cancer grade of at least one tissue region, based, at least in part, on zinc levels associated with a cluster of picture-elements representing said tissue region. 3. The method according to claim 1, further comprising segmenting the zinc data into a plurality of segments, each corresponding to a predetermined range of zinc levels, wherein said clustering is according to said segments. 4. The method according to claim 1, further comprising displaying at least one of said clusters. 5. The method according to claim 1, further comprising determining a location of a tumor in the prostate based on said at least one cluster. 6. The method according to claim 1, further comprising estimating a cancer stage of said tissue region. 7. A method of guiding an invasive medical device in a prostate, comprising: determining a location of a tumor in the prostate using the method of claim 5;imaging the prostate to provide an image and marking said location on said image; andusing said image for guiding the medical device to said location. 8. A system for estimating a grade of a prostate cancer, comprising: an input module, configured for inputting zinc data associated with the prostate, the zinc data being arranged gridwise in a plurality of picture-elements representing a zinc map of the prostate;a clustering module, configured for clustering the zinc map according to zinc levels associated with said picture-elements; anda grade estimating module, configured for estimating a cancer grade of at least one tissue region, based, at least in part, on zinc levels associated with a cluster of picture-elements representing said tissue region. 9. The system of claim 8, further comprising a segmentation module configured for segmenting said zinc data into a plurality of segments, each corresponding to a predetermined range of zinc levels, wherein said clustering module is configured for clustering said zinc map according to said segments. 10. The system according to claim 8, further comprising a staging module, for estimating a cancer stage of said tissue region. 11. The system according to claim 8, further comprising a mapping module for generating said zinc map using sa... | Yeda Research And Development Co. Ltd. | A61B005/00 C40B030/02 C40B060/12
US20100331738 | A1 | 2010-12-30 | INTEGRATED SENSOR AND INTERCONNECT FOR MEASURING A PARAMETER OF THE MUSCULAR-SKELETAL SYSTEM | 2010-06-29 | A sensing insert device (100) is disclosed for measuring a parameter of the muscular-skeletal system. The sensing insert device (100) can be temporary or permanent. Used intra-operatively, the sensing insert device (100) comprises an insert dock (202) and a sensing module (200). The sensing module (200) is a self-contained encapsulated measurement device having at least one contacting surface that couples to the muscular-skeletal system. The sensing module (200) comprises one or more sensing assemblages (1802), electronic circuitry (307), and communication circuitry (320). An interconnect stack within the sensing module (1700) couples at least one circuit board (1612 and 1616) to one or more sensing assemblages (1802). The interconnect stack includes at least one flexible interconnect (1506) that couples to the circuit board (1612 and 1616). The flexible interconnect (1506) can be in a path for conducting an energy wave through the sensing assemblage (1802). | 1. An interconnect stack to reduce a form factor of a sensing module for measuring a parameter of the muscular-skeletal system comprising: a first printed circuit board;a second printed circuit board;a flexible interconnect coupled between the first and second printed circuit boards; anda sensing platform including at least one sensing assemblage where the flexible interconnect couples to the sensing assemblage and where the flexible interconnect includes at least one bend placing the flexible interconnect on at least two levels of the interconnect stack. 2. The interconnect stack of claim 1 further including a housing for the sensing module where the second printed circuit is suspended a predetermined height above a bottom surface of the housing and where electronic components mounted on the second printed circuit board are directed towards the bottom surface of the housing. 3. The interconnect stack of claim 2 where electronic components mounted on the first printed circuit board are directed towards an opening of the housing. 4. The interconnect stack of claim 3 further including a first support structure coupled to the housing overlying the first and second printed circuit board. 5. The interconnect stack of claim 4 further including a second support structure overlying the first support structure where the sensing assemblage is between the first and second support structure. 6. The interconnect stack of claim 5 where a major surface of the second support structure is above a surface of the sidewall of the housing. 7. The interconnect stack of claim 6 further including a cap overlying the second support structure. 8. The interconnect stack of claim 7 where the cap includes a lip overlying a portion of the sidewall of the housing and where an adhesive between the lip of the cap and the sidewall of the housing seals and attaches the cap to the housing. 9. The interconnect stack of claim 1 where a sensing assemblage comprises: a first transducer;an energy propagation medium; anda second transducer where the flexible interconnect includes a first bend to couple the flexible interconnect between the first transducer and a first location of the energy propagation medium and where the flexible interconnect includes a second bend to couple the flexible interconnect between the second transducer and a second location of the energy propagation medium. 10. The interconnect stack of claim 1 where the sensing assemblage comprises: a transducer;an energy propagation medium; anda reflective surface where the at least one bend places couples the flexible interconnect between the transducer and a first location of the energy propagation medium and where the reflective surface is located at a second location of the energy propagation medium. 11. A sensing module having an interconnect stack for measuring a parameter of the muscular-skeletal system comprising: at least one sensing assemblage for measuring a parameter of the muscular-skeletal system;el... | OrthoSensor | A61B005/103 H05K001/14
US20100298722 | A1 | 2010-11-25 | LIVING BODY INFORMATION DETECTION APPARATUS AND BLOOD-PRESSURE METER | 2010-07-30 | A living body information collecting apparatus with a part of the living body information collecting apparatus includes a shape composed of a cylinder, a cone, a prism, a pyramid, a truncated cone or a truncated pyramid. The living body information collecting apparatus further including a sensing part for collecting living body information. | 1. A living body information collecting apparatus, wherein a part of the living body information collecting apparatus includes a shape composed of a cylinder, a cone, a prism, a pyramid, a truncated cone or a truncated pyramid, the living body information collecting apparatus comprising: a sensing part for collecting living body information. 2. The living body information collecting apparatus as claimed in claim 1, the living body information collecting apparatus including a hollow part, that is a route of sound, in the axial orientation of the cylinder, the cone, the prism, the pyramid, the truncated cone or the truncated pyramid. 3. The living body information collecting apparatus as claimed in claim 1, wherein the sensing part includes a blood pressure sensor; and the blood pressure sensor includes at least a pair of an light-emitting element and a light-receiving element, and a pressure applying part. 4. A living body information collecting apparatus having a shape suitable for being worn in an external ear or an external auditory meatus of a human, comprising: a sensing part for collecting living body information from the external ear or the external auditory meatus. 5. The living body information collecting apparatus as claimed in claim 1, the living body information collecting apparatus further comprising a hollow part that is a route of sound when the living body information collecting apparatus is worn in the external ear or the external auditory meatus. 6. The living body information collecting apparatus as claimed in claim 4, wherein the shape is formed based on a shape of the external ear or the external auditory meatus of the human. 7. The living body information collecting apparatus as claimed in claim 1, further comprising: a drive control part for drive-controlling the sensing part and processing a signal from the sensing part. 8. The living body information collecting apparatus as claimed in claim 7, further comprising: a transmission part for transmitting information processed by the drive control part. 9. The living body information collecting apparatus as claimed in claim 8, further comprising a power source part for supplying power to at least one of the sensing part, the drive control part and the transmission part. 10. The living body information collecting apparatus as claimed in claim 1, wherein the sensing part includes at least one of a blood pressure sensor, a body temperature sensor, a pulse sensor, a posture sensor, an acceleration sensor, a blood oxygen levels sensor, and an electroencephalogram sensor. 11. The living body information collecting apparatus as claimed in claim 1, further comprising a suspension part for suspending the living body information collecting apparatus from the external ear. 12. The living body information collecting apparatus as claimed in claim 1, wherein the sensing part includes a blood pressure sensor; and the blood pressure sensor includes a detection part for detecting a pu... | NIPPON TELEGRAPH AND TELEPHONE CORP. | A61B005/02
US20100292585 | A1 | 2010-11-18 | LIVING BODY INFORMATION DETECTION APPARATUS AND BLOOD-PRESSURE METER | 2010-07-30 | A living body information detection apparatus for detecting living body information at an auricle of a human body. The living body information detection apparatus having a shape that follows a cartilage of the auricle in a periphery of a concha auriculae. | 1. A living body information detection apparatus for detecting living body information at an auricle of a human body, wherein the living body information detection apparatus has a shape that follows a cartilage of the auricle in a periphery of a concha auriculae. 2. A living body information detection apparatus for detecting living body information at an auricle of a human body, wherein the living body information detection apparatus has a shape that follows the auricle in a periphery of a concha auriculae. 3. The living body information detection apparatus as claimed in claim 1, wherein the shape that follows the cartilage of the auricle is a shape that follows an antihelix. 4. The living body information detection apparatus as claimed in claim 2, wherein the shape that follows the auricle is a shape that follows the concha auriculae and an antihelix. 5. The living body information detection apparatus as claimed in claim 1, wherein the shape that follows the cartilage of the auricle is a shape that follows an antihelix and a lamina of tragus. 6. The living body information detection apparatus as claimed in claim 2, wherein the shape that follows the auricle is a shape that follows the concha auriculae, a tragus, an antitragus and an antihelix. 7. The living body information detection apparatus as claimed in claim 2, wherein the shape that follows the auricle is a shape that follows the concha auriculae, a tragus, a crus helicis, a crus anthelicis, an antihelix, an antitragus and a cavity of the concha. 8. The living body information detection apparatus as claimed in claim 1, wherein the shape of the living body information detection apparatus also follows an outer surface of a lamina of tragus, and covers the tragus with an inside part and an outside part, wherein the inside part touches the tragus from the inside of the lamina of tragus and wherein the outside part touches the tragus from the outer side of the lamina of tragus. 9. The living body information detection apparatus as claimed in claim 2, wherein the shape of the living body information detection apparatus also follows an outer surface of the tragus, and covers the tragus with an inside part and an outside part, wherein the inside part touches the tragus from the inside of the tragus and wherein the outside part touches the tragus from the outer side of the tragus. 10. The living body information detection apparatus as claimed in claim 1, comprising: a light-emitting element for causing output light to enter a living body tissue of the auricle; anda light-receiving element for receiving scattered light from the living body tissue. 11. The living body information detection apparatus as claimed in claim 1, comprising: a pressure applying part for applying a pressure on a part of an auricle; anda detection part for detecting a pulse wave at the part of the auricle. 12. The living body information detection apparatus as claimed in claim 11, wherein the pressure applying part ... | NIPPON TELEGRAPH AND TELEPHONE CORP. | A61B005/0215 A61B005/00 A61B005/02
US20100280336 | A1 | 2010-11-04 | ANXIETY DISORDER MONITORING | 2010-03-31 | An anxiety episode may be identified as being an anxiety event that is attributable to an anxiety disorder of a patient based on the patient activity associated with the anxiety episode. The patient activity may include, for example, patient motion, patient posture or voice activity. Detection of the activity component during an anxiety episode can help distinguish between a general anxiety state and an anxiety event that differs from the general anxiety state. Examples of anxiety events include, for example, an occurrence of a compulsion or a panic attack. The detected anxiety events can be used to evaluate an anxiety disorder of a patient, evaluate therapy programs implemented by a medical device to treat the anxiety disorder, or control therapy delivery. In some examples, a mood state transition is detected based on patient activity information and therapy delivery is controlled based on the detection of the mood state transition. | 1. A method comprising: detecting, with a processor, an anxiety episode of a patient based on a physiological parameter of the patient; anddetermining, with the processor, whether the anxiety episode is an anxiety event attributable to an anxiety disorder of the patient based on a signal indicative of patient activity. 2. The method of claim 1, wherein the signal is indicative of at least one of motion, posture state or voice activity of the patient. 3. The method of claim 1, wherein the anxiety disorder is characterized by the occurrence of a motor activity and determining whether the anxiety episode is the anxiety event comprises: detecting, with the processor, the motor activity based on the signal; anddetermining, with the processor, the anxiety episode is the anxiety event based on the detection of the motor activity. 4. The method of claim 3, wherein the motor activity comprises at least one of a compulsive motor activity, a motor tic, or at least one of a patient motion or posture state resulting from a compulsion. 5. The method of claim 1, wherein the anxiety disorder is characterized by the occurrence of a voice activity and determining whether the anxiety episode is the anxiety event comprises: detecting, with the processor, the voice activity based on the signal; anddetermining, with the processor, the anxiety episode is the anxiety event based on the detection of the voice activity. 6. The method of claim 5, wherein the voice activity comprises at least one of a compulsive act or a vocal tic. 7. The method of claim 1, wherein the signal comprises a first signal, and detecting the anxiety episode comprises: receiving a second signal indicative of the physiological parameter; anddetermining, with the processor, whether the second signal is indicative of the anxiety episode. 8. The method of claim 7, wherein determining whether the second signal is indicative of the anxiety episode comprises at least one of comparing a peak amplitude of the second signal to a threshold amplitude, comparing an average amplitude of the second signal to the threshold amplitude, comparing a median amplitude of the second signal to the threshold amplitude, comparing a trend in a waveform of the second signal over time to a template, comparing a power level within one or more frequency bands of the second signal to a threshold power level or comparing a ratio of power levels in two or more frequency bands of the second signal to a threshold power level ratio, and detecting the anxiety episode based on the comparison. 9. The method of claim 7, wherein the second signal is indicative of at least one of brain activity, a heart rate, respiratory rate, electrodermal activity, a facial expression or facial flushing of the patient. 10. The method of claim 7, wherein determining whether the second signal is indicative of the anxiety episode comprises determining whether the second signal is indicative of the anxiety episode with a support vector machine algor... | Medtronic, Inc. | A61B005/00
US20100298728 | A1 | 2010-11-25 | Signal Processing Techniques For Determining Signal Quality Using A Wavelet Transform Ratio Surface | 2009-05-20 | According to embodiments, a wavelet transform ratio surface measure signal may be generated from two PPG signals. Values of the wavelet transform ratio surface measure signal at a given moment of time (i.e., instantaneous values) may be indicative of localized signal discrepancies within and/or between the PPG signals such as noise and signal artifacts. Spikes in the instantaneous values of the wavelet transform ratio surface measure signal may be located and used to determine a signal quality measure for the PPG signals. Characteristics of the spikes such as number, location, grouping, distribution, amplitude, and polarity may be used in the signal quality determination. | 1. A method for determining signal quality using a wavelet transformation ratio surface measure, the method comprising: obtaining a first photoplethysmograph (PPG) signal and a second PPG signal from a sensor;performing a continuous wavelet transform of the first PPG signal and the second PPG signal to produce a first wavelet transformed signal and a second wavelet transformed signal;generating a wavelet transform ratio surface scalogram based at least in part on a ratio of the first wavelet transformed signal to the second wavelet transformed signal;generating a wavelet transform ratio surface measure signal from a region in the wavelet transform ratio surface scalogram; andlocating one or more relatively high energy areas in the wavelet transform ratio surface measure signal;determining a signal quality measure for the first PPG signal and the second PPG signal based at least in part on the located relatively high energy areas in the wavelet transform ratio surface measure signal. 2. The method of claim 1 wherein the first PPG signal comprises a red PPG signal and the second PPG signal comprises an infrared PPG signal. 3. The method of claim 1, further comprising normalizing the first PPG signal and the second PPG signal obtained from the sensor prior to performing the continuous wavelet transforms. 4. The method of claim 1, wherein generating the wavelet transform ratio surface measure signal comprises identifying a pulse band within the wavelet transform ratio surface scalogram. 5. The method of claim 1, wherein the one or more relatively high energy areas located in the wavelet transform ratio surface measure signal are generally associated with noise in one of the first PPG signal and the second PPG signal at temporal locations corresponding to locations of the relatively high energy areas. 6. The method of claim 1, further comprising determining the signal quality measure based at least in part on a number of relatively high energy areas located in the wavelet transform ratio surface measure signal. 7. The method of claim 1, further comprising determining the signal quality measure based at least in part on an amplitude of the one or more relatively high energy areas located in the wavelet transform ratio surface measure signal. 8. The method of claim 1, further comprising determining the signal quality measure based at least in part on a polarity of the one or more relatively high energy areas located in the wavelet transform ratio surface measure signal. 9. The method of claim 8, wherein the polarity is based at least in part on a relative difference in a noise level associated with the first PPG signal and the second PPG signal. 10. The method of claim 1, further comprising determining that a plurality of the identified relatively high energy areas are part of a group of relatively high energy areas based at least in part on a proximity of the plurality of relatively high energy areas. 11. The method of claim 10, further compr... | Nellcor Puritan Bennett Ireland | A61B005/02
US20100298672 | A1 | 2010-11-25 | METHODS FOR EVALUATING A SUBJECT USING DATA ASSOCIATED WITH A FLUORESCENT ANALYTE | 2010-06-25 | Methods, systems, devices and computer program product include: (i) administering a fluorescent analyte to a subject; (ii) repetitively emitting excitation light from a sensor over a desired monitoring period; (iii) detecting fluorescence intensity in response to the excitation light using the sensor that outputs the excitation light; and (iv) using data associated with the detected fluorescence intensity to perform at least one of: (a) calculate the concentration or dose of the analyte received proximate to the sensor site; (b) evaluate the pharmacodynamic or pharmacokinetic activity of the fluorescent analyte; (c) confirm Ab attachment to a tumor site; (d) monitor a non-target site to confirm it is not unduly affected by a therapy; (e) monitor for changes in cellular properties; (f) use the calculated dose or concentration data to adjust or customize a therapeutic amount of the analyte administered to the subject; (g) confirm micelle concentration at a target site and then stimulate toxin release based on the confirmation; and (h) monitor for the expression of a protein produced from a gene therapy modification.;In particular embodiments, the intensity of the excitation signals emitted to the localized tissue can be varied in a predetermined manner to generate optical profiling data of the response of the tissue proximate the sensor. | 1. A method of evaluating a subject, comprising: administering a fluorescent analyte to a subject, the fluorescent analyte including at least one of a fluor-labeled analyte, a naturally fluorescent analyte and an analyte that exhibits fluorescence when internally administered to the subject;repetitively emitting excitation light from an implanted sensor over a desired monitoring period;electronically detecting fluorescence intensity over time to generate a time-dependent measurement profile of fluorescence in localized tissue in response to the excitation light using the implanted sensor that outputs the excitation light;electronically using data associated with the detected fluorescence intensity to perform at least one of: (a) calculate the concentration or dose of the analyte received proximate to the implanted sensor; (b) evaluate the pharmacodynamic or pharmacokinetic activity of the fluorescent analyte; (c) confirm Ab attachment to a tumor site; (d) monitor a non-target site to confirm it is not unduly affected by a therapy; (e) monitor for changes in cellular properties; (f) use a calculated dose or concentration data to adjust or customize a therapeutic amount of an fluorescent analyte administered to the subject; (g) confirm micelle concentration at a target site and then stimulate toxin release based on the confirmation; and (h) monitor for the expression of a protein produced from a gene therapy modification. 2. A method according to claim 1, further comprising varying the intensity of the excitation signals emitted to the localized tissue in a predetermined manner to generate optical profiling data of the response of the tissue proximate the sensor. 3. A method according to claim 1, wherein the sensor is implanted in or proximate to a tumor. 4. A method according to claim 1, wherein the sensor is adapted to be implanted at depths in the body up to about 5-25 cm below the skin of a patient. 5. A method according to claim 1, wherein the sensor is adapted to be implanted at depths in the body between about 5-20 cm below the skin of a patient, the method further comprising pulsing a laser diode disposed in the implanted sensor to generate the excitation light. 6. A method according to claim 1, wherein the sensor is adapted to be implanted at depths in the body from about 1 to about 25 cm below the skin of a patient, the method further comprising pulsing a laser diode disposed in the implanted sensor to generate the excitation light. 7. A method according to claim 5, wherein the laser diode is operated with between about a 1-10% duty cycle to generate the excitation light. 8. A method according to claim 7, wherein the repeated emissions of the excitation light and associated detecting steps are carried out at spaced apart intervals over at least 1 hour. 9. A method according to claim 8, wherein the emitting and detecting steps are repeated at desired intervals over between about at least one 24-48 hour monitoring period. 10. A met... | - | A61B005/1455
US20100305482 | A1 | 2010-12-02 | SENSOR APPLICATION COMPRISING A CARRIER STRUCTURE AND A SENSOR STRIP, CARRIER STRUCTURE, SENSOR STRIP AND METHOD FOR OPERATING THE SENSOR APPLICATION | 2008-11-04 | A sensor application (11) has modules consisting of a carrier strip (13) and a sensor strip (12). The sensor strip (12) is fixed to the carrier strip (13) which, in turn, is fixed, for example, to the back of a human proband. In this way, the spinal curvature can be determined, the carrier strip (13) and the sensor strip (12) being elastically embodied for this purpose. The carrier strip (13) has a data memory (22), for example an RFID label on which relevant information is stored. The data can be evaluated by an evaluation device (14), and the sensor strip (12) is only released when the carrier strip is suitable for the sensor strip used. In this way, defective functions and measuring errors are advantageously avoided. | 1. A sensor application for a moveable body, comprising an elastic carrier structure with a connecting surface, which can be fastened to the surface of the body by the connecting surface anda moveable sensor strip, which is held by the carrier structure following the contour of the surface,whereinthe sensor strip and the carrier structure are connected to one another by means of a detachable connection, with the carrier structure having a data storage device, in which a code is stored and which has a transmit interface for the code, andthe sensor strip having a switching unit, with which the sensor function can be disabled and enabled and which has a receive interface for an enabling and disabling signal obtained by evaluating the code. 2. A sensor application for a moveable body, comprising an elastic carrier structure with a connecting surface, which can be fastened to the surface of the body by way of the connecting surface, anda flexible sensor strip, which is held by the carrier structure following the contour of the surface,whereinthe sensor strip and the carrier structure are connected to one another by means of a detachable connection, with the carrier structure having a data storage device, in which a code is stored and which has a transmit interface for the code, andthe sensor strip being connected to a control device, which has a receive interface for the code and an evaluation unit for the code for generating an enabling or disabling signal, with which the sensor function can be disabled and enabled. 3. An elastic carrier structure with a connecting surface, which can be fastened to the surface of a moveable body by way of the connecting surface, comprising a receptacle for a flexible sensor strip, which can be held by the carrier structure following the contour of the surface, whereinthe carrier structure has a detachable connecting structure for fastening the sensor strip, anda data storage device, in which a code is stored and which is equipped with a transmit interface for the code. 4. The carrier structure according to claim 3, wherein the data storage device and the transmit interface are realized by means of an RFID transponder. 5. The carrier structure according to claim 3, wherein the code contains information relating to the condition of the sensor strips to be used. 6. The carrier structure according to claim 3, wherein the code contains a serial number. 7. The carrier structure according to claim 3, wherein the code contains a use-by date. 8. A flexible sensor strip, which can be held by an elastic carrier structure, which can be fastened for its part to a connecting surface on the surface of a moveable body, whereinthe sensor strip comprises has a detachable connecting structure relative to the fastening on the carrier strip following the contour of the surface, andhas a switching unit (28), with which the sensor function can be disabled and enabled and which is equipped with a receive interface for an enabling or ... | - | A61B005/11 H04Q005/22
US20100331643 | A1 | 2010-12-30 | Extruded Analyte Sensors and Methods of Using Same | 2009-06-30 | The present disclosure provides electrode structures and integrated electrode structures having one or more conductive materials coextruded with one or more dielectric materials. The disclosed electrode structures can be configured for use as analyte sensors. Also provided, are methods of making and using the electrode structures and integrated electrode structures described herein. | 1. An electrode structure, comprising: a first conductive material, a second conductive material, and a dielectric material,wherein the first conductive material, the second conductive material, and the dielectric material are coextruded to provide an electrode structure having the first conductive material and the second conductive material electrically isolated by the dielectric material, andwherein the electrode structure comprises a lumen, the lumen comprising a first opening, a second opening, and a lumen wall. 2. The electrode structure of claim 1, wherein the electrode structure is configured for in vitro use. 3. The electrode structure of claim 1, wherein the electrode structure is configured for in vivo use. 4. The electrode structure of claim 3, wherein at least a portion of the electrode structure is adapted to be implanted in a subject. 5. The electrode structure of claim 4, wherein at least a portion of the electrode structure is adapted to be subcutaneously positioned in the subject. 6. The electrode structure of claim 1, wherein the first conductive material and the second conductive material comprise conductive polymers. 7. The electrode structure of claim 1, wherein the first conductive material comprises a first color, the second conductive material comprises a second color and the first color and the second color are different. 8. The electrode structure of claim 1, wherein the dielectric material is a biocompatible material. 9. The electrode structure of claim 1, wherein the first conductive material comprises a working electrode and the second conductive material comprises a reference/counter electrode, the working electrode comprising an analyte responsive enzyme disposed on, in or proximate thereto. 10. The electrode structure of claim 9, wherein the first conductive material comprises carbon. 11. The electrode structure of claim 9, wherein the second conductive material comprises Ag/AgCl. 12. The electrode structure of claim 9, further comprising a membrane disposed over the working electrode. 13. The electrode structure of claim 9, wherein the analyte-responsive enzyme is a glucose-responsive enzyme. 14. The electrode structure of claim 9, wherein the analyte-responsive enzyme is a ketone-responsive enzyme. 15. The electrode structure of claim 9, wherein the working electrode comprises a redox mediator disposed on, in or proximate thereto. 16. The electrode structure of claim 15, wherein the redox mediator comprises a ruthenium-containing complex or an osmium-containing complex. 17. The electrode structure of claim 1, wherein the electrode structure further comprises an extruded outer layer. 18. The electrode structure of claim 17, wherein the extruded outer layer comprises a biocompatible material. 19. The electrode structure of claim 17, wherein the extruded outer layer is coextruded with the first conductive material, the second conductive material, and the dielectric material. 20. The electrode stru... | Abbott Diabetes Care Inc. | A61B005/1473 G01N027/26 G01N027/30 A61B005/15
US20100331727 | A1 | 2010-12-30 | Linear Device for Microdialysis, Method of Manufacturing the Device and Method for Studying a Tissue with the Device | 2008-07-09 | There is provided a linear device (100) for microdialysis, comprising a length of tubing (102) having at least one lumen (118) wherein at least one tubular membrane sleeve (104) is coaxially arranged outside the tubing periphery and extending at least partially along the periphery of the tubing (102). The membrane sleeve (104) is extending intermediate the ends of the tubing (102), and the tubing (102) extends at least through the tubular membrane sleeve (104). There is also provided a method of manufacturing the linear device (100) and a method wherein the linear device (100) is used. | 1. A linear device for microdialysis, comprising at least one dialysis chamber, a length of tubing having at least one lumen, at least one first element and at least one second element connecting one of said at least one lumen with one of said at least one dialysis chamber, wherein at least one tubular, sleeve-formed membrane is coaxially arranged outside the tubing periphery and extending at least partially along the periphery of the tubing, intermediate the ends of the tubing, and wherein the tubing extends at least through the membrane. 2. A linear device for microdialysis according to claim 1, wherein the tubing imparts sufficient rigidity to the linear device in order to eliminate the need for using supporting wires to support or protect the membrane when locating the linear device in target tissue. 3. A linear device for microdialysis according to claim 1, wherein the membrane extends sufficiently long along the periphery of the tubing to establish a suitable interaction between a perfusion fluid transported in the lumen and a surrounding body fluid, so that sufficient material can be collected from a target tissue and transported with the perfusion fluid for subsequent analysis. 4. A linear device for microdialysis according to claim 1, wherein the membrane is arranged on the tubing to provide said dialysis chamber between the outer surface of the length of tubing and the inner surface of the membrane. 5. A linear device for microdialysis according to claim 1 wherein said at least one first element comprises at least one inlet channel and said at least one second element comprises at least one outlet channel connecting one of said at least one lumen with one of said at least one dialysis chamber. 6. A linear device for microdialysis according to claim 5, wherein there is provided at least one blocking in the tubing between the at least one first element or inlet channel and the at least one second element or outlet channel, said at least one blocking the lumen for passage of liquid. 7. A method of manufacturing a linear device for microdialysis, the method comprising the following steps: a. providing a length of tubing having at least one lumen,b. providing at least one first element or inlet channel extending from the outside surface of the length of tubing to the at least one lumen,c. providing at least one second element or outlet channel at a distance from the at least one first element or inlet channel, the at least one second element or outlet channel extending from the outside surface of the length of tubing to the at least one lumen,d. providing at least one blocking between said at least one first element or inlet channel and the at least one second element or outlet channel, said blocking stopping the passage of liquid in the lumen,e. arranging a tubular, sleeve-formed membrane coaxially outside the tubing periphery and extending at least partially along the periphery of the tubing, said membrane covering the at least one f... | - | A61B005/00 A61M031/00 B21D051/16
US07841987 | B2 | 2010-11-30 | System and method for visualizing sleep-related information | 2005-03-31 | A method for improved visualization of information related to the physiology of a sleeping patient is disclosed. Physiological information from the patient is obtained by a device, converted to digital format, and transformed into physiological data of two or more types. Physiological data of two or more types are combined into a compact graphical display representing data from all physiological data types. | 1. A method for transforming sound information into a visual representation of the breathing of a patient comprising: identifying a first set of physiological data values derived from respiratory sound information collected from the patient during a time interval associated with a sleep period of the patient;using a computer, obtaining from the first set of physiological data values an envelope signal, the envelope signal being indicative of a loudness, level, or intensity of the respiratory sound;outputting within a spatial region having a first axis and a second axis an envelope trace derived from the envelope signal, with the time interval associated with the sleep period represented on the first axis and the magnitude of the envelope signal represented on the second axis; andoutputting within the spatial region provided by the first axis and the second axis a first plurality of graphical components, each of the first plurality of graphical components being associated with a different time period provided on the first axis, at least one of the first plurality of graphical components representing a degree of snoring during the time period associated with the graphical component by the use of a color or texture pattern of the graphical component, wherein at least three different degrees of snoring, one of which degrees of snoring may be the absence of snoring, are represented among the first plurality of graphical components; and each of the first plurality of graphical components having a top border conforming to the envelope trace. 2. The method of claim 1 wherein the duration of the time period associated with each graphical component in the first plurality of graphical components is less than or equal to the duration of an inhalation by the patient. 3. The method of claim 1 wherein the duration of the time period associated with a graphical component in the first plurality of graphical components is less than or equal to one second. 4. A method for transforming sound information into a visual representation of the breathing of a patient comprising: identifying a first set of physiological data values derived from respiratory sound information collected from the patient during a time interval associated with a sleep period of the patient;using a computer, obtaining from the first set of physiological data values an envelope signal, the envelope signal being indicative of a loudness, level, or intensity of the respiratory sound;outputting within a spatial region having a first axis and a second axis an envelope trace derived from the envelope signal, with the time interval associated with the sleep period represented on the first axis and the magnitude of the envelope signal represented on the second axis;outputting within the spatial region provided by the first axis and the second axis a first plurality of graphical components, each of the first plurality of graphical components being associated with a different time period provided on th... | Apneos Corporation | A61B005/08 A61B007/00
US07826879 | B2 | 2010-11-02 | Analyte sensors and methods of use | 2006-02-28 | An analyte sensor system including a substrate, a first electrode disposed on a first surface of the substrate, a second electrode disposed on a second surface of the substrate, a third electrode provided in electrical contact with at least one of the first or second electrodes, where at least a portion of the first electrode and the second electrode are subcutaneously positioned in a patient, and where the third electrode is substantially entirely positioned external to the patient, and corresponding methods are provided. | 1. An analyte sensor system, comprising: a substrate;a first electrode disposed on a first surface of the substrate;a second electrode disposed on a second surface of the substrate; anda transmitter unit configured for electrical communication with one or more of the first and second electrodes, and including a housing having a plurality of electrical contacts disposed thereon and configured for physical contact with a skin surface; wherein the plurality of electrical contacts are electrically coupled such that when the transmitter is positioned on the skin surface, one or more of the plurality of electrical contacts provide continuous electrical contact between the transmitter and the skin surface during a time period when at least a portion of the first electrode and the second electrodes are subcutaneously positioned under the skin surface. 2. The system of claim 1 wherein the first and second surfaces of the substrate are the same. 3. The system of claim 1 wherein the first surface of the substrate on the opposite side of the substrate as the second surface of the substrate. 4. The system of claim 1 wherein the analyte sensor includes a glucose sensor. 5. The system of claim 1 wherein the plurality of electrical contacts are coupled to a counter electrode contact of the transmitter unit. 6. The system of claim 1 wherein the plurality of electrical contacts are electrically coupled to each other such that when the transmitter unit is positioned on the skin surface, one or more of the plurality of electrical contacts provide continuous electrical contact between the skin surface and the transmitter unit. 7. The system of claim 1 wherein the time period includes one of approximately, three days, five days or seven days. 8. The system of claim 1 wherein the first electrode is a working electrode, and the second electrode is a reference electrode. 9. The system of claim 8 further including a second working electrode at least a portion of which is subcutaneously positionable in the patient. 10. The system of claim 1 further including an alarm unit, the alarm unit configured to alert the patient only when all of the plurality of the electrical contacts are no longer in physical contact with the skin surface at the same time. 11. The system of claim 10 wherein the alarm unit is integrated in the transmitter unit. 12. The system of claim 1 wherein the plurality of electrical contacts are disposed on the surface of the transmitter unit housing in a predetermined position. 13. The system of claim 12 wherein the predetermined position includes one or more of a concentric pattern or a circular pattern. 14. A method, comprising: subcutaneously positioning at least a portion of an analyte sensor including a first electrode disposed on a first surface of a substrate and a second electrode disposed on a second surface of the substrate such that the portion of the analyte sensor is in fluid contact with an analyte under a skin surface; andposition... | Abbott Diabetes Care Inc. | A61B005/05 A61B005/1473
US07837635 | B2 | 2010-11-23 | Device and method for determining at least one characteristic point of joint to be orthopedically measured | 2005-11-07 | A device determines at least one characteristic point of a joint which joins together two limbs and is to be measured orthopedically comprising a first coupling element which may be positioned on a first limb adjoining the joint, a second coupling element which may be moved relative to the first coupling element and may be positioned on a second limb adjoining the joint, and a measuring device for determining the position of the first coupling element relative to the second coupling element. The measuring device has at least one reference point rigidly connected to one of the first and second coupling elements, and a reference point detection device rigidly connected to the respective other of the first and second coupling elements. During a movement of the first and second limbs with respect to one another the reference point detection device determines the at least one characteristic point. | 1. Device for determining at least one characteristic point of a joint which joins together two limbs and is to be measured orthopedically, said device compromising: a first coupling element adapted to be attached to a first limb adjoining the joint, a second coupling element adapted to be pivoted and shifted relative to the first coupling element and adapted to be positioned on a second limb adjoining the joint, anda measuring device for determining the position of the first coupling element relative to the second coupling element,wherein the measuring device has at least one reference point which is rigidly connected to one of the first and second coupling elements, and also has a reference point detection device which is rigidly connected to the respective other coupling element of the first and second coupling elements, the reference point detection device having a detector integrated therein for detecting a plurality of positions of the reference point relative to the reference point detection device by an optical, electronic, and/or electromagnetic measurement during a pivoting and shifting movement of the first and second limbs with respect to one another and determining therefrom the at least one characteristic point, and wherein the measuring device includes an indicating device positioned on one of the first and second coupling elements that indicates the at least one determined characteristic point of the joint directly on the joint. 2. Device according to claim 1, wherein the reference point detection device detects a movement path of the respective reference point from the plurality of positions of the reference point. 3. Device according to claim 1, including a processing system which is coupled or adapted to be coupled to the reference point detection device for transmitting data. 4. Device according to claim 3, wherein the processing system is configured to be integral with the device. 5. Device according to claim 3, wherein the processing system determines the at least one characteristic point as at least one virtual pivotal point for the joint to be measured from the data received from the reference point detection device, in particular from the movement path of the respective reference point. 6. Device according to claim 3, wherein the processing system determines a plurality of characteristic points, from the data received from the reference point detection device, in particular from the movement path of the respective reference point. 7. Device according to claim 1, including at least two mutually spaced apart reference points which are rigidly connected to one of the first and second coupling elements, the reference point detection device detecting a plurality of positions of the at least two reference points relative to the reference point detection device during a movement of the first and second limbs with respect to one another. 8. Device according to claim 1, wherein the indicating device comprises an optical out... | Fraunhofer-Gesellschaft zur Förderung der Angewandten Forschung E.V. | A61B005/103
US07826881 | B1 | 2010-11-02 | Endocardial chamber mapping system | 2000-06-07 | An endocardial mapping system is described to map electric field activity in a heart chamber. The system includes a catheter assembly having an expandable array of electrodes and a plug with each electrode connected to a single connection in the plug. The system also contains an interface apparatus with a plug adapter to receive the catheter plug, a voltage acquisition apparatus, and a signal generator to generate signals in response to voltages received on the electrodes. Finally, the system contains a computer capable of computing a three-dimensional volumetric electric field distribution based on signals received from the signal generator. | 1. An endocardial chamber mapping system comprising: a) a catheter assembly having; i) an array of signal acquisition electrodes expandable from a substantially cylindrical shape to an expanded shape, andii) a catheter plug with multiple connections, each of the connections being electrically coupled to one of said array of acquisition electrodes;b) an interface apparatus having; i) an interface plug connected to the catheter plug to establish electrical connection to each of the electrodes,ii) a voltage acquisition apparatus in communication with the interface plug and coupled to said array of acquisition electrodes, said voltage acquisition apparatus having an analog to digital converter for digitizing voltages on said array of acquisition electrodes, andiii) a signal generator in communication with one or more of said array of voltage acquisition electrodes for generating low current pulses; andc) a computer having; i) electrical communication with the signal generator of the interface apparatus to control its function,ii) electrical communication with the voltage acquisition apparatus to receive the voltages acquired by the signal acquisition electrodes,iii) processing unit to compute a three-dimensional volumetric electric field distribution based on the voltages received from the signal acquisition electrodes, andiv) a display showing the computed field distribution. 2. The endocardial chamber mapping system of claim 1, wherein the computer further comprises: v) means for obtaining data relating to volume and shape of the endocardial chamber through the generation of low current pulses by the signal generator and for creating an image of said volume and shape, and wherein the display shows the computed field distribution displayed on the image of the volume and shape of the endocardial chamber. 3. The endocardial chamber mapping system of claim 2, wherein the display shows the computed field distribution in a continuously filled color-scale map shown over the volume and shape of the endocardial chamber. 4. An endocardial chamber mapping system comprising: a) a catheter assembly having i) an array of signal acquisition electrodes expandable from a substantially cylindrical shape to a substantially spherical shape, andii) an electrical connector plug with multiple connectors electrically coupled to each of the electrodes present in said array;b) voltage acquisition apparatus in communication with the electrodes of said array, having an analog to digital converter for converting the voltages on said array to corresponding signals; andc) a computer having i) an electrical communication with the voltage acquisition apparatus to receive the voltage acquired by the signal acquisition electrodes,ii) a processing unit capable of computing a three-dimensional volumetric electric field distribution based on the signals acquired at the signal acquisition electrodes, andiii) a display for displaying the three-dimensional volumetric electric field dis... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/042
US07842060 | B2 | 2010-11-30 | Lancet device for generating a puncture wound, and lancet drive assembly | 2005-12-09 | A lancet device for generating a puncture wound in a skin surface, comprising a housing, in which a lancet can be moved on a puncturing path, a lancet drive with a drive rotor that can be driven by a drive spring, a lancet coupling mechanism that transforms a rotational motion of the drive rotor into a puncturing and returning motion of the lancet during a working cycle of the lancet drive, a reference element that is mobile relative to the lancet and relative to the housing and is fixed in a defined position relative to the lancet drive during the puncturing and rests against the surface of the skin with a contact surface thereof such that a puncture wound with a reproducible puncturing depth is generated by the puncturing motion, and a reference element coupling mechanism that is coupled to the lancet drive in order to move the reference element. The invention further relates to a suitable lancet drive assembly for a lancet device of this type. | 1. A lancing device, comprising: a housing;a lancet drive disposed in the housing and including a drive rotor driven by a drive spring;a first coupling mechanism disposed in the housing and including a lancet connecting arm configured to connect to a lancet, the first coupling mechanism transforming motion of the drive rotor into reciprocating motion of the lancet connecting arm;a reference element that is moveable relative to the housing and relative to the lancet connecting arm, the reference element having a contact surface adapted to contact skin to be punctured during the reciprocating motion of the lancet connecting arm; anda second coupling mechanism which couples the reference element to the lancet drive, wherein movement of the lancet drive is transformed into movement of the reference element, the second coupling mechanism transforming movement of the drive rotor into a return movement of the reference element. 2. The lancing device of claim 1, wherein the lancet drive comprises a tensioning rotor for tensioning the drive spring. 3. The lancing device of claim 2, wherein, in a working cycle, the drive rotor runs through a total range of rotation that comprises a preparatory range of rotation during which the drive rotor rotates jointly with the tensioning rotor. 4. The lancing device of claim 3, wherein the reference element is moved by the second coupling mechanism while the drive rotor rotates through the preparatory range of rotation. 5. The lancing device of claim 1, wherein the first coupling mechanism comprises a steering curve along which a steering curve traveler travels in order to transform rotational motion of the drive rotor into the reciprocating motion of the lancet connecting arm. 6. The lancing device of claim 5, wherein the second coupling mechanism comprises a second steering curve along which a second steering curve traveler travels. 7. The lancing device of claim 2, wherein the tensioning rotor comprises a triggering cam that actuates a trigger mechanism when the tensioning rotor is rotated beyond a predetermined angle, whereupon the trigger mechanism releases the drive rotor to start the reciprocating motion of the lancet connecting arm. 8. The lancing device of claim 1, wherein the reference element comprises a pressure sensor that generates a signal for starting a puncturing motion once a pressure in excess of a predetermined minimal pressure is applied to the contact surface. 9. The lancing device of claim 1, wherein the second coupling mechanism transforms a rotational motion of the drive rotor into motion of the reference element. 10. The lancing device of claim 2, wherein the second coupling mechanism transforms rotational motion of the tensioning rotor into motion of the reference element. 11. The lancing device of claim 1, wherein the reference element is arranged in the housing behind a plane that extends through an inner rim that defines an opening in the housing, the inner rim defining a surface... | Roche Diagnostics Operations, Inc. | A61B010/00 A61B005/00 A61B017/14 A61B017/32
US07957792 | B2 | 2011-06-07 | Spatial resolution determination for cardiac mapping | 2010-06-01 | A non-contact cardiac mapping method is disclosed that includes: (i) inserting a catheter into a heart cavity having an endocardium surface, the catheter including multiple, spatially distributed electrodes; (ii) measuring signals at the catheter electrodes in response to electrical activity in the heart cavity with the catheter spaced from the endocardium surface; and (iii) determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the electrodes with respect to the endocardium surface. Related systems and computer programs are also disclosed. | 1. A method comprising: inserting a catheter into a heart cavity having an endocardium surface, the catheter comprising multiple, spatially distributed electrodes;measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;determining physiological information at multiple locations of the endocardium surface based on the measured signals and positions of the catheter electrodes with respect to the endocardium surface;computing values indicative of a degree of spatial resolution of the determined physiological information for at least some of the locations of the endocardium surface; anddisplaying information related to at least some of the determined physiological information and information related to at least some of the computed resolution values. 2. The method of claim 1, wherein the computed values are derived, at least in part, from a transformation function relating the physiological information at the multiple locations of the endocardium surface to the signals measured by the catheter electrodes. 3. The method of claim 1, further comprising: moving the catheter to each of multiple, different positions in the heart cavity; andfor each of the different catheter positions, determining the positions of the catheter electrodes with respect to the endocardium surface and measuring signals at the catheter electrodes in response to electrical activity in the heart cavity;wherein the determination of the physiological information at the multiple locations of the endocardium surface is based further on the positions of the catheter electrodes and the measured signals at the different catheter positions. 4. The method of claim 3, wherein determining the position of the catheter electrodes comprises measuring information about at least one of a position and orientation of the catheter within the heart cavity and determining the position of the catheter electrodes is based further on information about the distribution of the electrodes on the catheter. 5. The method of claim 3, wherein determining the position of the catheter electrodes comprises directly measuring the position of each catheter electrode within the heart cavity. 6. The method of claim 1, wherein displaying information related to at least some of the determined physiological information further comprises overlaying on a display device at least some of the physiological information determined at the multiple surface locations. 7. The method of claim 1, wherein displaying information related to at least some of the determined physiological information further comprises displaying at least a portion of the endocardium surface to include a selected subset of the physiological information determined at the multiple surface locations, wherein the subset is selected based on at least some of the computed resolution values. 8. The method of claim 1, wherein the determined physiological information comprises isopotential lines or bands correspondi... | Rhythmia Medical, Inc. | A61B005/02
US08238631 | B2 | 2012-08-07 | System and method for automatic registration between an image and a subject | 2009-05-13 | A patient defines a patient space in which an instrument can be tracked and navigated. An image space is defined by image data that can be registered to the patient space. A tracking device can be connected to a member in a known manner that includes imageable portions that generate image points in the image data. The tracking device can be tracked to register patient space to image space. | 1. A method for registering image data to a subject, comprising: connecting a base to a subject;connecting an imageable body having a plurality of physical imageable portions to the base for acquiring image data of the plurality of imageable portions relative to the subject to allow determining a location of a plurality of image fiducial points that relate to the plurality of physical imageable portions in an image space of the image data; andconnecting a tracking device to the base for determining a location of a plurality of virtual fiducial points defined relative to a tracked location of the tracking device in a subject space for registering the image space of the image data and the subject space by relating the plurality of virtual fiducial points to the plurality of image fiducial points. 2. The method of claim 1, wherein registering the image space of the image data and the subject space includes executing a program with a processor to determine a translation map based on automatically determining the determined location of the plurality of physical fiducial imageable portions defined by the plurality of image fiducial points in the image space of the image data and automatically determining the determined location of the plurality of virtual fiducial points defined relative to the tracking device in the subject space. 3. The method of claim 2, further comprising: tracking an instrument in the subject space;determining a location of the instrument in the image space; anddisplaying an icon representing the determined location of the instrument superimposed on a display of the registered image data. 4. The method of claim 1, wherein the determining the location of the plurality of virtual fiducial points defined relative to the tracking device in the subject space, includes: calibrating the tracking device to determine the plurality of virtual fiducial points relative to a physical portion of the tracking device; andtracking the tracking device to identify the location of the plurality of virtual fiducial points based on the calibration. 5. The method of claim 4, wherein calibrating the tracking device to determine the plurality of virtual fiducial points relative to a physical portion of the tracking device includes determining a location of the plurality of virtual fiducial points outside of a physical boundary of the physical portion of the tracking device. 6. The method of claim 5, wherein connecting the tracking device to the base includes fixing the tracking device in the base in a known configuration; and further including:determining the location of the calibrated plurality of virtual fiducial points relative to the base based on the known configuration of the fixation. 7. The method of claim 6, wherein the known configuration of the fixation orients the imageable portion with the determined plurality of virtual fiducial points. 8. The method of claim 7, wherein the fixing the tracking device in the base in the known configura... | Medtronic Navigation, Inc. | G06K009/00
US08167845 | B2 | 2012-05-01 | Catheter having distal sealing member | 2009-06-02 | A catheter includes a proximal end, a distal end, a tubular member, and an inner member extending through the tubular member. The inner member is movable axially relative to the tubular member. The catheter also includes a seal member including a first portion and a second portion. The seal member extends between the tubular member and the inner member. The first portion of the seal member remains stationary relative to the tubular member, and the second portion of the seal member remains stationary relative to the inner member during relative axial movement between the tubular member and the inner member. | 1. A catheter comprising: a proximal end;a distal end;a tubular member;an inner member extending through said tubular member, said inner member movable axially relative to said tubular member; anda seal member comprising a first portion and a second portion, said seal member extending between said tubular member and said inner member, said first portion of said seal member remaining stationary relative to said tubular member and said second portion of said seal member remaining stationary relative to said inner member during relative axial movement between said tubular member and said inner member. 2. A catheter in accordance with claim 1 wherein said inner member comprises at least one of a rod and a second tubular member. 3. A catheter in accordance with claim 1 wherein said seal member first portion is attached to an inner surface of said tubular member. 4. A catheter in accordance with claim 1 wherein said seal member second portion is attached to an outer surface of said inner member. 5. A catheter in accordance with claim 1 wherein said seal member comprises a first surface and a second surface, said seal member first portion attached to said tubular member at said first surface, said seal member second portion attached to said inner member at said second surface. 6. A catheter in accordance with claim 1 further comprising an outer tubular member, said inner member comprises an inner tubular member, said catheter further comprising a basket electrode assembly connected to a distal end of said inner tubular member and configured to be extended from, and retracted into, said outer tubular member. 7. A catheter in accordance with claim 6 wherein said seal member is configured to fold upon itself when said basket electrode is withdrawn into said outer tubular member. 8. A basket electrode mapping catheter comprising: a proximal end;a distal end;a tubular member;an inner member extending through said tubular member;a basket electrode assembly connected to a distal end of said inner member; anda seal member comprising a first portion and a second portion, said seal member first portion contacting said tubular member, said seal member second portion contacting said inner member, said second portion of said seal member remaining stationary relative to said inner member during relative axial movement between said tubular member and said inner member, said seal member configured to prevent fluid flow between said tubular member and said inner member. 9. A catheter in accordance with claim 8 wherein said inner member comprises at least one of a rod and a second tubular member. 10. A catheter in accordance with claim 8 wherein said seal member comprises a first surface and a second surface, said first surface of said seal member first portion contacting said tubular member, said second surface of said seal member second portion contacting said inner member. 11. A catheter in accordance with claim 8 further comprising an outer tubular member, ... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61M005/178
US09005140 | B2 | 2015-04-14 | Weight monitoring apparatus, weight monitoring system, and related methods thereof | 2008-10-31 | The present invention relates to embodiments which enable repetitive measurement and monitoring of the load applied to an apparatus, in a system, and/or by a method which may exceed a predetermined threshold of weight. Specifically, the present invention relates to weight monitoring apparatus, locomotion monitoring apparatus, a gait-response apparatus, a weight monitoring system, and a therapeutic load monitoring method. | 1. A weight monitoring apparatus for wearing on a foot of a user, comprising: a. a rigid first member having a length and width sized to be larger than the length and width of said foot, and having a top surface, said rigid first member designed to be a stable and stiff base that is unyielding when a load is applied thereon;b. a second member having a length and width sized to be larger than the length and width of said foot, and having a top surface and a bottom surface;c. a first sensor disposed between and configured to associate with the top surface of the first member and the bottom surface of the second member and disposed at a location proximal to a toe region of said foot;d. a second sensor disposed between and configured to associate with the top surface of the first member and the bottom surface of the second member and disposed at a location proximal to a heel region of said foot; ande. a support bar disposed between and in contact with the top surface of the first member and the bottom surface of the second member, and disposed at a location between the first and second sensors,wherein the support bar functions as a fulcrum for the second member when the load is applied to the top surface of the second member allowing for angular load transfer to the first or the second sensor. 2. The apparatus of claim 1, wherein one of said sensors is a mechanical sensor. 3. The apparatus of claim 2, wherein one of said sensors is a dome-shaped sensor. 4. The apparatus of claim 1, wherein said apparatus is capable of being inserted into footwear. 5. The apparatus of claim 1, wherein said apparatus comprises a cam boot. 6. The apparatus of claim 1, wherein at least one of said sensors provides a detectable feedback to a user if a predetermined mechanical force threshold is exceeded on at least one of said sensors. 7. The apparatus of claim 2, wherein said dome-shaped sensor is a snap-dome sensor. 8. The apparatus of claim 6, wherein said detectable feedback is one of audible, visual and tactile. 9. The apparatus of claim 8, wherein said detectable feedback is audible. 10. The apparatus of claim 8, wherein said detectable feedback is tactile. 11. The apparatus of claim 1, wherein said first and second members are flat. 12. The apparatus of claim 1, wherein said first member substantially conforms to an inner bottom portion of footwear. 13. The apparatus of claim 12, wherein said footwear is one of a shoe, a boot, a cast, a walking cast, a brace, or a cam boot. 14. The apparatus of claim 13, wherein said footwear is a cam boot. 15. The apparatus of claim 12, wherein said apparatus is a replacement for the foam insert of a cam boot. 16. The apparatus of claim 1, wherein at least one of said sensors generates an electrical signal. 17. The apparatus of claim 1, wherein at least one of said sensors generates a radio-frequency signal. 18. The apparatus of claim 6, wherein said predetermined mechanical force threshold is between ten and fift... | The Research Foundation for the State University of New York | G01G019/44 A61B005/103 A61B005/00 A61B005/117
US08634890 | B2 | 2014-01-21 | Device and method for monitoring of absolute oxygen saturation and tissue hemoglobin concentration | 2010-06-10 | A method and medical device for detecting signals that detects emitted light scattered by a volume of tissue delivered along a first pathway at a plurality of wavelengths to generate corresponding first detected light intensity output signals, detects emitted light scattered by the volume of tissue delivered along a second pathway different from the first pathway at a plurality of wavelengths to generate corresponding second detected light intensity output signals, determines whether a difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is greater than a predetermined threshold, and alters sensing by the device in response to the determining whether a difference is greater than the predetermined threshold. | 1. A method of detecting signals in a medical device, comprising: detecting emitted light scattered by a volume of tissue delivered along a first pathway to generate corresponding first detected light intensity output signals;detecting emitted light scattered by the volume of tissue delivered along a second pathway different from the first pathway to generate corresponding second detected light intensity output signals;determining whether a difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is greater than a predetermined threshold; andaltering sensing by the device in response to the determining whether a difference is greater than the predetermined threshold. 2. The method of claim 1, further comprising: determining a first oxygen saturation measurement corresponding to detected light emitted along the first pathway in response to the first detected light intensity output signals; anddetermining a second oxygen saturation measurement corresponding to detected light emitted along the second pathway in response to the second detected light intensity output signals, wherein determining whether a difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is greater than a predetermined threshold comprises comparing the first and second oxygen saturation measurement. 3. The method of claim 2, wherein altering sensing by the device comprises selecting between the first pathway and the second pathway. 4. The method of claim 2, wherein determining whether a difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is greater than a predetermined threshold comprises determining levels of uniformity of the volume of tissue. 5. The method of claim 4, wherein determining whether a difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is greater than a predetermined threshold comprises: determining whether a difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is less than a first level threshold;identifying the volume of tissue as having a first tissue uniformity in response the difference being less than the first level threshold;determining whether a difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is greater than a second level threshold different from the first level threshold; andidentifying the volume of tissue as having a second tissue uniformity in response the difference being both less than the second level threshold and not less than the first level threshold. 6. The method of claim 5, wherein altering sensing by the device comprises selecting between the first pathway and the second pathway in response to the identi... | Medtronic, Inc. | A61B005/00
US08784342 | B2 | 2014-07-22 | Shape sensing clothes to inform the wearer of a condition | 2010-02-26 | A garment may be configured with one or more sensors and circuitry to communicate information related to a physiological condition to the user. | 1. An apparatus comprising: a first garment configured to be worn by a user and including: a first fabric;a first set of one or more sensors, the one or more sensors in the first set of one or more sensors being arranged to detect a first shape-related parameter of the first fabric and responsive to the detected first shape-related parameter to produce a first signal, wherein at least one of the one or more sensors in the first set of one or more sensors is arranged to detect a change in electrical impedance in response to the detected first shape-related parameter; andone or more control elements responsive to a second signal and arranged to provide a force on the fabric corresponding to the second signal; andcircuitry responsive to the first signal to determine a first shape of the first fabric and configured to compare the first shape of the first fabric to a reference shape of the first fabric, and to produce the second signal corresponding to a difference between the first shape of the first fabric and the reference shape of the first fabric, the second signal being indicative of a physiological condition of the user. 2. The apparatus of claim 1 wherein the circuitry is further responsive to a third signal, different from the first signal, and wherein the circuitry is further configured to produce the second signal in response to the third signal. 3. The apparatus of claim 2 further comprising a physiological sensor arranged to sense a physiological parameter of the user, wherein the physiological sensor is configured to produce the third signal. 4. The apparatus of claim 2 wherein the first shape-related parameter of the first fabric corresponds to a first region of the first fabric, and wherein a second set of one or more sensors is further arranged to detect a second shape-related parameter of the first fabric corresponding to a second region of the first fabric, different from the first region of first fabric, and wherein the one or more sensors is further responsive to the detected second shape-related parameter to produce the third signal. 5. The apparatus of claim 2 wherein the first shape-related parameter of the first fabric corresponds to a first time interval, and wherein the first set of one or more sensors is further arranged to detect a second shape-related parameter of the first fabric corresponding to a second time interval, different from the first time interval, and wherein the first set of one or more sensors is further responsive to the detected second shape-related parameter to produce the third signal. 6. The apparatus of claim 1 wherein the circuitry is further responsive to stored information to determine the first shape of the first fabric. 7. The apparatus of claim 1 wherein the one or more sensors in the first set of one or more sensors are configured to provide one or more corresponding measurements, the one or more corresponding measurements being associated with one or more corresponding regions of the fir... | The Invention Science Fund I LLC | A61B005/103 A61B005/117
US08765099 | B2 | 2014-07-01 | Magnetic resonance imaging hyperpolarization of liquids or solids by light with orbital angular momentum | 2008-12-19 | In magnetic resonance imaging (MRI), selected magnetic dipoles in a subject are aligned with a main magnetic field for later manipulation, and signals received after such manipulations are used to create image representations of the subject. One drawback is that even powerful magnetic fields can only align a very small percentage of dipoles in the region of the field. Electromagnetic radiation endowed with orbital angular momentum (OAM) aligns dipoles along the direction of travel of the radiation, but at a much higher percentage; as high as 100% of the dipoles in the region can be aligned. Resultantly, resonance signals emanating from the region are several orders of magnitude stronger than signals emanated using traditional MRI techniques. All electromagnetic radiation, including visible light can be endowed with OAM and used to hyperpolarize a region of interest. | 1. A light-based hyperpolarization device comprising: a light source configured to emit light;a phase hologram configured to impart orbital angular momentum to the light;a spatial filter configured to filter out a portion of the light and to allow a portion of the light with a pre-determined amount of orbital angular momentum to pass; andat least one optical element configured to direct the light that passes the spatial filter to at least one of blood and/or tissue to be hyperpolarized. 2. A magnetic resonance system comprising: the light-based hyperpolarization device as claimed in claim 1 configured to polarize selected dipoles via transferred orbital angular momentum; andan RF system configured to induce resonance in the polarized dipoles and configured to receive resonance signals. 3. The magnetic resonance system according to claim 2, further comprising: a catheter configured to be inserted into a patient, the catheter, comprising a light channel configured to channel light from the light-based hyperpolarization device to a working end of the catheter. 4. The magnetic resonance system according to claim 2, further, comprising: a gradient magnet system configured to subject the resonating dipoles in an examination region of the magnetic resonance system to a gradient magnetic field. 5. The magnetic resonance system according to claim 2, further comprising: a gradient magnet system configured to subject the resonating dipoles in an examination region of the magnetic resonance system to a gradient magnetic field, the gradient magnet system being mounted to the catheter adjacent the working end. 6. The magnetic resonance system as set forth in claim 2, wherein the phase hologram is embodied in a spatial light modulator. 7. The magnetic resonance system as set forth in claim 6, wherein the spatial light modulator includes a liquid crystal on silicon panel. 8. The magnetic resonance system as set forth in claim 2, further comprising: a wave polarizer configured to circularly polarize to the light. 9. The magnetic resonance system according to claim 2, wherein the spatial filter blocks a central, undiffracted bright spot of a diffraction pattern, and allows higher order diffraction components carrying orbital angular momentum of a single polarity to pass. 10. A surface probe comprising: a light output unit configured to direct light to penetrate tissue of a patient; andthe light-based hyperpolarization system as claimed in claim 1 configured to impart orbital angular momentum to generated light, the light to which orbital angular momentum has been imparted being discharged through the light output unit to polarize selected dipoles in the patient. 11. A catheter comprising: an elongated portion terminating in a working end configured to be inserted in a patient; andthe light-based hyperpolarization system as claimed in claim 1 configured to impart orbital angular momentum to generated light, the light to which orbital angular momentum has ... | Koninklijke Philips N.V. | A61B005/055
US09339206 | B2 | 2016-05-17 | Adaptor for endovascular electrocardiography | 2010-06-14 | An adaptor for endovascular electrocardiography includes a member having two ends and an inner lumen such that when one of the ends is connected to a catheter and the other end is connected to a syringe, fluid from the syringe is flushed into the catheter through the inner lumen. The adaptor also includes a metal insert, disposed along the inner lumen of the member, which connects through a sealed electrical connection to the outer side of the member. | 1. An adaptor for endovascular electrocardiography, comprising: a member having a first end, a second end, and an inner lumen, the first end designed to be connected to a luer of a catheter and the second end designed to be connected to a syringe, such that fluid from the syringe is flushed into the catheter through the inner lumen; anda metal ring, disposed along the inner lumen of the member and located fully inside the member, the metal ring connecting through a sealed electrical connection to an outer side of the member, the metal ring separated from the catheter by a section of the member when the first end of the member is connected to the luer of the catheter. 2. The adaptor according to claim 1, wherein the metal ring comprises a metal ring disposed along an inner surface of the inner lumen. 3. The adaptor according to claim 1, wherein the metal ring terminates prior to the two ends of the member. 4. The adaptor according to claim 1, further comprising a wire having a first end connected to the metal ring. 5. The adaptor according to claim 4, further comprising a connector coupled to an opposing end of the wire. 6. The adaptor according to claim 5, wherein the connector comprises external insulation. 7. The adaptor according to claim 5, wherein the connector is configured to plug into an electronic module. 8. The adaptor according to claim 5, wherein the connector comprises a nipple shape. 9. The adaptor according to claim 1, wherein an exterior most surface of the one of the ends of the adaptor is tapered. 10. The adaptor according to claim 9, wherein the other end comprises a luer connector. 11. The adaptor according to claim 1, wherein a body portion provides a connection between the metal ring along an inner surface of the inner lumen of the member and a wire external to the member through a perforation in a wall of the member. 12. The adaptor according to claim 11, wherein the connection through the wall is waterproof. 13. The adaptor according to claim 1, wherein the two ends of the member comprise different external and internal diameters. 14. An assembly for endovascular electrocardiography, comprising: the adaptor according to claim 1; anda connector separate from a body of the adaptor coupled to the metal ring. 15. The assembly according to claim 14, wherein the connector is coupled to the adaptor by a wire between the metal ring and the connector, the wire configured to freely orient the connector with respect to the adaptor. 16. The assembly according to claim 14, wherein the connector is configured to plug into an electronic module. 17. The assembly according to claim 14, wherein the connector comprises a nipple shape. | Bard Access Systems, Inc. | A61B005/042 A61B005/06 A61M025/00 A61B005/04 A61B005/0452 A61M039/10 A61B005/00 A61B019/00
US09820684 | B2 | 2017-11-21 | Method and apparatus for a fluid sampling device | 2005-06-03 | Methods and apparatus are provided for manufacturing an analyte detecting device. In one embodiment, the apparatus comprises a housing; a penetrating member driver; a cartridge containing a plurality of penetrating members; a display on the cartridge; a linear slider on the housing, the slider coupled to a rod; and the rod moving with the slider, the rod having at least one roller. The device uses the linear motion of the slider to rotate the cartridge, punch open a new cavity and load a new penetrating member. | 1. A fluid sampling device comprising: a housing;a cartridge positioned in the housinga plurality of body fluid sensing strips integrated into a multi-strip roll;at least one penetrating member positioned in the cartridge;a gripper engaged to a penetrating member;a drive assembly coupled to the cartridge;a member for moving the cartridge to bring an unused penetrating member into position for launch, the member clearing the gripper and drive assembly to lift the gripper and drive assembly clear so that the cartridge can rotate; anda sterile blood barrier with an offset pad to maintain an air gap between a patients' finger and the sterile blood barrier, wherein the offset pad is a multi-layer laminate, the sterile blood barrier including a bend relief and one or more location features to position the sterile blood barrier. 2. The device of claim 1 wherein a cam surface is aligned parallel to a slider. 3. The device of claim 1 wherein a linear motion of a cam rotates the cartridge and moves a plunger to break the sterility barrier on the cartridge. 4. The device of claim 1 wherein a cam surface comprises a linear strip of maternal with at least two raised portions and two depressed portions. 5. A fluid sampling device comprising; a housing;a cartridge defining a plurality of cavities, said cartridge sized to fit within said housing; a plurality of body fluid sensing strips integrated into a multi-strip roll;a plurality of penetrating members at least partially contained in said cavities of the cartridge wherein each of a penetrating member is slidably movable, to extend outward from said cartridge to penetrate tissue; said cavities each having a lateral opening providing access to an elongate portion of the penetrating member;a gripper configured to be engaged with each of a penetrating member;a drive assembly coupled to the cartridge;a sterile blood barrier with an offset pad to maintain an air gap between a patients' finger and the sterile blood barrier, wherein the offset pad is a multi-layer laminate, the sterile blood barrier including a bend relief and one or more location features to position the sterile blood barrier; anda member for moving the cartridge to bring an unused penetrating member into position for launch, the member clearing the gripper and drive assembly to lift the gripper and drive assembly clear so that the cartridge can rotate. 6. A device comprising: a housing;a penetrating member driver;a cartridge containing a plurality of penetrating members;a plurality of body fluid sensing strips integrated into a multi-strip roll;a display on said cartridge;a linear slider on the housing, said slider coupled to a rod;said rod moving with said slider, said rod having at least one roller using a linear motion of the slider to rotate the cartridge, punch open a new cavity and load an unused penetrating member;a gripper configured to be engaged with each of a penetrating member;a driver coupled to the cartridge;a member for moving t... | SANOFI-AVENTIS DEUTSCHLAND GMBH | A61B005/00 A61B005/157 A61B005/151 A61B005/15
US09746419 | B2 | 2017-08-29 | Optical MRI device | 2010-05-05 | A dual modality imaging apparatus, comprising a magnetic resonance imaging (=MRI) system and a fluorescence molecular tomography (=FMT) system, for investigating a sample (42) located at a sample position (10), wherein the MRI system comprises a magnet (9) with a room temperature bore (8), with the sample position (10) located within the bore (8), and wherein the FMT system comprises means for directing a light beam (12, 17; 30) towards the sample position (10), and a position-sensitive detector (37) for collecting fluorescence light from the sample (42), is characterized in that the position-sensitive detector (37) is located within the bore (8), wherein at least part of the sample (42) is imaged onto the position-sensitive detector (37), and that the means for directing the light beam (12, 17; 30) comprise a focusing device for focusing the light beam (12, 17; 30) into a focal spot (41) on the sample (42), and a scanning device for scanning the focal spot (41) on the sample (42). The apparatus is simple in design and versatile in application, and in particular imposes less limitations on the maximum number of source/detector pairs during FMT imaging, thus allowing a better FMT image resolution. | 1. A dual modality imaging apparatus for investigating a sample located at a sample position, the apparatus comprising: a magnetic resonance imaging (=MRI) system, said magnetic resonance imaging system having a magnet with a room temperature bore within which the sample position is located, said room temperature bore having a longitudinal axis, said magnetic resonance imaging system further comprising a radio frequency coil having at least one window, said at least one window opening in a transverse direction, substantially perpendicular to said longitudinal axis of said room temperature bore, wherein said radio frequency coil is disposed within said room temperature bore and the sample is at least partially within said coil; anda fluorescence molecular tomography (=FMT) system having a mechanism for directing a light beam towards the sample position and a position-sensitive detector, said position-sensitive detector having electrical connections, said position-sensitive detector and said electrical connections being disposed within said bore and outside said radio frequency coil to detect fluorescence light from the sample, wherein at least part of the sample is imaged onto said position-sensitive detector, said mechanism for directing said light beam having a focusing device for focusing said light beam into a focal spot on the sample, a scanning device structured for continuously scanning said focal spot on the sample and a deflection mirror, said scanning device being disposed outside of said room temperature bore and said deflection mirror being disposed inside said room temperature bore, wherein said mechanism for directing said light beam, said radio frequency coil and said position-sensitive detector are disposed, structured and dimensioned such that said light beam enters said room temperature bore substantially parallel to said longitudinal axis and is deflected by said deflection mirror in a substantially transverse direction to pass through said at least one window of said radio frequency coil and illuminate the sample, wherein fluorescence light emanating from the sample passes through said at least one window in said radio frequency coil and is incident on said position sensitive detector, the apparatus thereby being adapted for simultaneous measurement of MRI data and FMT data. 2. The apparatus of claim 1, wherein said FMT system comprises a rigid linkage connecting said scanning device with a sample platform for supporting the sample at the sample position. 3. The apparatus of claim 2, wherein said FMT system comprises a guiding and/or a stop located within said bore for inserting and positioning said rigid linkage. 4. The apparatus of claim 1, wherein said scanning device is adapted to move an optical element into different positions corresponding to different positions of said focal spot of said light beam on the sample, wherein said optical element directs said light beam. 5. The apparatus of claim 4, wherein said optical ... | University of Zurich | A61B005/05 G01N021/64 A61B005/00 G01N021/47
US20100331736 | A1 | 2010-12-30 | WIRELESS SENSING MODULE FOR SENSING A PARAMETER OF THE MUSCULAR-SKELETAL SYSTEM | 2010-06-29 | A sensing insert device (100) is disclosed for measuring a parameter of the muscular-skeletal system. The sensing insert device (100) can be temporary or permanent. Used intra-operatively, the sensing insert device (100) comprises an insert dock (202) and a sensing module (200). The sensing module (200) is a self-contained encapsulated measurement device having a contacting surface that couples to the muscular-skeletal system. The sensing module (200) comprises one or more sensors (303), electronic circuitry (307), and communication circuitry (320). The electronic circuitry (307) operatively couples to the one or more sensors (303) to measure the parameter. The communication circuitry (320) couples to the electronic circuitry (307) to wirelessly transmit measurement data. The communication circuitry (320) comprises a data packetizer (422), a cyclic redundancy check circuit (413), a transmitter (416), a matching network (414), and an antenna (412). The sensing insert device (100) when inserted allows movement of the muscular-skeletal system. | 1. A sensing system for measurement of a parameter of the muscular-skeletal system comprising: a sensing module comprising: one or more sensors;electronic circuitry operatively coupled to the sensors; andcommunication circuitry coupled to the electronic circuitry to transmit data from the one or more sensors comprising: a data packetizer coupled to the electronic circuitry;a cyclic redundancy check circuit coupled to the data packetizer;a transmitter coupled to the cyclic redundancy check circuit. 2. The sensing system of claim 1 further including a receiving circuit external to the sensing module to receive data from the one or more sensors comprising: an antenna;a matching network coupled to the antenna;a receiver coupled to the matching network;a cyclic redundancy check circuit coupled to the receiver; anda data packetizer coupled to the cyclic redundancy check circuit 3. The sensing system of claim 1 where the sensing module further includes an encapsulated enclosure having a power source therein. 4. The sensing system of claim 3 where the communication circuitry further comprises: a matching network coupled to the transmitter; andan antenna coupled to the matching network where the antenna is within the encapsulated enclosure. 5. The sensing system of claim 3 where the enclosure includes a surface for receiving the applied parameter, where the one or more sensors underlie the surface, and where the electronic circuitry underlies the one or more sensors. 6. The sensing system of claim 5 where the sensing module is in a trial insert. 7. The sensing system of claim 5 where the sensing module is in a final insert having at least one bearing surface to promote movement of the muscular-skeletal system. 8. The sensing system of claim 1 where the electronic circuitry and the communication circuitry are on a single application integrated circuit to reduce the form factor and power consumption the sensor module. 9. The sensing system of claim 1 where the transmission range of the transmitter is 5 meters or less. 10. The sensing system of claim 1 where the sensing module is transmit only to prevent device tampering. 11. The sensing system of claim 1 where a cyclic redundancy check is performed on a transmission. 12. A prosthetic component for measuring a parameter of the muscular-skeletal system comprising: a trial insert comprising: a contacting surface;one or more sensors coupled to the contacting surface;electronic circuitry operatively coupled to the sensors; andcommunication circuitry coupled to the contacting surface to transmit data from the one or more sensors comprising: a data packetizer coupled to the electronic circuitry;a cyclic redundancy check circuit coupled to the data packetizer;a transmitter coupled to the cyclic redundancy check circuit. 13. The prosthetic component of claim 12 where the trial insert further comprises: a dock having a cavity; anda sensing module having the contacting surface where the one or more sensors... | OrthoSensor | A61B005/103
US20100305473 | A1 | 2010-12-02 | TISSUE CONFORMING MICRONEEDLE DEVICE FOR DRUG DELIVERY OR BIOLOGICAL FLUID COLLECTION | 2010-08-10 | Microneedle arrays are provided for use on a contoured or flexible tissue surface. In one embodiment, the microneedle array includes a plurality of microneedles, each having a base portion, a tip end portion distal to the base portion, and body portion therebetween; and a flexible substrate which comprises a plurality of apertures, each of which are defined by (i) a plurality of substrate elements which are integral with the base portions of the microneedles, and (ii) at least one spring element connecting at least two of the substrate elements. The spring element may include a curved element, such as a C-shaped, U-shaped, or S-shaped element. Apertures may be defined, for example, by two substrate elements, which connected to three or four spring elements. A skin patch is provided for therapeutic or diagnostic applications, which includes the microneedle array and an adhesive material. | 1. A device for administration of a drug to a contoured or flexible tissue surface, the device comprising: a microneedle array which comprises a plurality of microneedles, each microneedle having a base portion, a tip end portion distal to the base portion, and body portion between the base portion and the tip end portion, andan elastically stretchable substrate which comprises a plurality of apertures, each of which are defined by both (i) a plurality of substrate elements which are integral with the base portions of the microneedles, and (ii) at least one spring element integral with and connecting at least two of the substrate elements; andat least one drug storage element, which contains a drug formulation, disposed adjacent to the microneedle array. 2. The device of claim 1, wherein the drug storage element is attached to a first surface of the substrate, the first surface being opposed to a second surface of the substrate of the microneedle array, wherein the microneedles project from said second surface. 3. The device of claim 1, wherein the drug storage element is coated onto an outer surface of the microneedles. 4. The device of claim 1, wherein the drug storage element can flex and deform with substrate of the microneedle array. 5. The device of claim 1, further comprising a release mechanism for releasing the drug formulation from the drug storage element to permit the drug formulation to pass through the apertures of the substrate of the microneedle array. 6. The device of claim 1, wherein the drug storage element comprises a porous material, wherein the drug formulation is stored in pores of the porous material. 7. The device of claim 1, further comprising an adhesive surface suitable for securing the device to the skin of a patient during administration of the drug formulation to the patient. 8. The device of claim 1, further comprising a removable release liner, which covers the microneedles and apertures prior to use of the device. 9. The device of claim 1, wherein at least one of the spring elements of the substrate comprises a curved element. 10. The device of claim 9, wherein the curved element is C-shaped, U-shaped, or S-shaped. 11. The device of claim 1, wherein at least one of the apertures of the substrate is defined by two substrate elements. 12. The device of claim 1, wherein at least one of the substrate elements is connected to three or four spring elements. 13. The device of claim 1, wherein at least one of the plurality of microneedles has a channel extending substantially from the base portion through at least a part of the body portion, the channel being open along at least part of the body portion and in fluid communication with at least one of the apertures in the substrate. 14. The device of claim 13, wherein the base portion of the at least one of the microneedles is untapered and has a substantially rectangular cross-sectional shape in a plane parallel to the substrate. 15. The device of claim 14,... | - | A61M005/00 A61B005/00 B23P017/04
US20100298727 | A1 | 2010-11-25 | LIVING BODY INFORMATION DETECTION APPARATUS AND BLOOD-PRESSURE METER | 2010-07-30 | A cuff including a base composed of a non-elastic member. The cuff further including an elastic member provided on one surface of the base. The cuff also including an air supplying pipe, where a pressing surface of the elastic member swells only on the one surface by supplying air from the air supplying pipe. | 1. A cuff comprising: a base composed of a non-elastic member;an elastic member provided on one surface of the base; andan air supplying pipe,wherein a pressing surface of the elastic member swells only on the one surface by supplying air from the air supplying pipe. 2. The cuff as claimed in claim 1, wherein the base is plain-like or curved surface-like. 3. A cuff comprising: a case in which one face is open;an elastic member that covers the open face to form a pressing surface on the face; andan air supplying pipe provided to the case,wherein the pressing surface swells by supplying air from the air supplying pipe into the inside of the cuff surrounded by the case and the elastic member. 4. The cuff as claimed in claim 3, wherein an outer shape of the pressing surface is round, elliptic or rounded quadrangular. 5. The cuff as claimed in claim 3, wherein the shape of the pressing surface is concave in relation to the outside of the cuff. 6. The cuff as claimed in claim 3, wherein the shape of the pressing surface is convex to the outside of the cuff. 7. The cuff as claimed in claim 3, wherein the shape of the pressing surface is flat. 8. The cuff as claimed in claim 3, wherein the cuff includes slack between the pressing surface and the case in the elastic member for expanding or contracting a side part for supporting the pressing surface in a movement direction of the pressing surface. 9. The cuff as claimed in claim 8, wherein an expansion ratio of the pressing surface in the swelling direction of the pressing surface is less than an expansion ratio of the slack in the swelling direction. 10. The cuff as claimed in claim 8, wherein thickness of a part that forms the pressing surface in the elastic member is greater than thickness of a part of the slack. 11. The cuff as claimed in claim 8, wherein an elasticity of material of a part that forms the pressing surface in the elastic member is smaller than an elasticity of material of a part forming the slack. 12. The cuff as claimed in claim 3, wherein the side part of the elastic member is fixed to an outer wall of the case by an elastic body. 13. The cuff as claimed in claim 3, wherein the side part of the elastic member is fixed to an outer wall of the case by elasticity of the side part. 14. The cuff as claimed in claim 3, wherein the side part of the elastic member is fixed to an outer wall of the case by thermocompression bonding. 15. The cuff as claimed in claim 3, wherein the material of the elastic member is a material having elasticity such as silicone resin, natural rubber and butyl rubber, or general plastic material such as polyethylene, polypropylene, polyvinyl chloride, polyvinyl acetate, or copolymer of these, or airtight cloth or paper obtained by coating nonwoven fabric with polymer, wherein the material of the elastic member passes light. 16. The cuff as claimed in claim 1, the cuff further comprising a light-emitting element, in the inside of the case, for emittin... | NIPPON TELEGRAPH AND TELEPHONE CORP. | A61B005/022
US20100298666 | A1 | 2010-11-25 | LIVING BODY INFORMATION DETECTION APPARATUS AND BLOOD-PRESSURE METER | 2010-07-30 | A living body information collecting system includes a portable terminal and a living body information collecting apparatus. The living body information collecting apparatus includes a sensing part for collecting living body information from an external auditory meatus. The living body information collecting apparatus further includes a drive control part for drive-controlling the sensing part and processing a signal from the sensing part. The living body information collecting apparatus also includes a transmission part for transmitting information processed by the drive control part. The portable terminal includes a terminal receive part for receive-processing information from the transmission part. The portable terminal also includes a display part for displaying information from the terminal receive part. | 1. A living body information collecting system comprising a portable terminal and a living body information collecting apparatus, the living body information collecting apparatus comprising:a sensing part for collecting living body information from an external auditory meatus;a drive control part for drive-controlling the sensing part and processing a signal from the sensing part; anda transmission part for transmitting information processed by the drive control part,the portable terminal comprising:a terminal receive part for receive-processing information from the transmission part; anda display part for displaying information from the, terminal receive part. 2. The living body information collecting system as claimed in claim 1, wherein a part of the living body information collecting apparatus includes a shape composed of a cylinder, a cone, a prism, a pyramid, a truncated cone or a truncated pyramid. 3. The living body information collecting system as claimed in claim 1, wherein the living body information collecting apparatus has a shape suitable for being worn in an external ear or an external auditory meatus of a human. 4. The living body information collecting system as claimed in claim 1, wherein the living body information collecting apparatus includes a shape composed of a frame part including a first arm and a second arm that are opposed to each other. 5. A living body information collecting system comprising a portable terminal and a living body information collecting apparatus, the living body information collecting apparatus comprising:a sensing part for collecting living body information from an external auditory meatus;a drive control part for drive-controlling the sensing part and processing a signal from the sensing part; anda transmission part for transmitting information processed by the drive control part,the portable terminal comprising:a terminal receive part for receive-processing information from the transmission part; anda communication part for transmitting a signal from the terminal receive part to an information processing apparatus via a communication network. 6. The living body information collecting system as claimed in claim 1, wherein the transmission part is a wireless transmission part for transmitting a signal by wireless, and the terminal receive part is a wireless terminal receiving part for receiving a signal by wireless. 7. The living body information collecting system as claimed in claim 1, wherein information processed by the drive controlling part includes at least one of blood pressure, pulse, body temperature, posture, acceleration, blood oxygen level and electroencephalogram. 8. The living body information collecting system as claimed in claim 7, wherein a part of the living body information collecting apparatus includes a shape composed of a cylinder, a cone, a prism, a pyramid, a truncated cone or a truncated pyramid. 9. The living body information collecting system as claimed in claim 7,... | NIPPON TELEGRAPH AND TELEPHONE CORP. | A61B005/00
US20100324396 | A1 | 2010-12-23 | Analyte Monitoring Device and Methods of Use | 2010-08-26 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A sensor system for measuring an analyte in a host, the system comprising: a sensor configured to measure an analyte concentration in a host; andelectronics operably connected to the sensor and configured to initialize a sensor session after the sensor is in communication with a biological sample of the host, wherein the electronics are configured to confirm removal of the sensor from contact with the biological sample of the host prior to initialization of another sensor session. 3. The sensor system of claim 2, wherein the electronics are configured to render the sensor system inoperative after a predetermined time period. 4. The sensor system of claim 2, wherein the electronics are configured to identify a sensor insertion. 5. The sensor system of claim 2, wherein the electronics are configured to identify a removal of the sensor by evaluation of sensor data. 6. The sensor system of claim 2, wherein the electronics are configured to disallow display of sensor data. 7. The sensor system of claim 2, wherein the electronics are configured to disallow display of sensor data after removal of the sensor. 8. A sensor system for measuring an analyte in a host, the system comprising: a sensor configured to measure an analyte concentration in a host; andelectronics operably connected to the sensor, wherein the electronics configured to confirm a sensor insertion prior to initialization of a sensor session. 9. The sensor system of claim 8, wherein the electronics are configured to provide information indicative of a predetermined time period associated with the sensor. 10. The sensor system of claim 9, wherein the electronics are configured to receive the information indicative of a predetermined time period associated with the sensor and are configured to display sensor data for the predetermined time period. 11. The sensor system of claim 9, wherein the electronics are configured to receive the information indicative of a predetermined time period associated with the sensor and are configured to display sensor data only for the predetermined time period. 12. The sensor system of claim 9, wherein the electronics are configured to receive the information indicative of a predetermined time period associated with the sensor and wherein the sensor system is configured to become inoperative by discontinuing display of sensor data after the predetermined time period. 13. The sensor system of claim 8, wherein the electronics are configured to identify a sensor insertion by examining sensor data. 14. A sensor system for measuring an analyte in a host, the system comprising: a sensor configured to measure an analyte concentration in a host; andelectronics operably connected to the sensor and configured identify at least one of an insertion of the sensor and a removal of the sensor from contact with a biological sample of the host; anda user interface operably connected to the electronics, wherein the electronics are configured to co... | Abbott Diabetes Care Inc. | A61B005/05
US20100312147 | A1 | 2010-12-09 | Obesity Treatment Systems | 2010-08-16 | In some embodiments, methods to sense parameters associated with restriction procedures and/or devices are disclosed. In some embodiments, the parameters are sensed directly from the region of the procedure and/or device and in other embodiments, associated neural pathways are sensed. Methods for further optimization of the obesity treatment regimens are also disclosed. | 1. A method of quantifying the effectiveness of an implanted gastric restriction device in a patient undergoing a treatment for obesity comprising: a. Sensing a parameter associated with the stress on an implant caused by food passing through the restriction created by the implant wherein the implant resides between the external portion of the stomach and the abdominal wall and does not encircle the stomach. | - | A61B005/103
US20100312073 | A1 | 2010-12-09 | METHOD OF PREDICTING PAIN MEDICATION EFFICACY | 2009-01-29 | A predictive method that aims at matching a specific pain medication, particularly of the SNRI type, to a specific patient is disclosed. In accordance with the predictive method, DNIC pain modulation is first performed on the patient who is about to receive a pain medication treatment. The results of the test are then taken as is or optionally transformed to construct a value. By analyzing this value, a specific pain medication can be matched to the aforementioned patient. | 1. A method of predicting the efficacy of pain medications selected from the group consisting of SNRI and tricyclics, said method comprising the steps of: identifying a subject;subjecting said subject to at least a single DNIC test;obtaining scores of said DNIC test, andanalyzing said scores;wherein lower scores at said DNIC test are associated with a higher efficacy of said treatment, andwhereby said efficacy is predicted. 2. The method as in claim 1, further comprising the step of prescribing said subject with a pain medication. 3. The method as in claim 1, wherein said subject is expected to or undergoes a condition associated with pain. 4. The method as in claim 3, wherein said condition is selected from the group consisting of: idiopathic pain, neuropathic pain, diabetic neuropathy, uremic neuropathy, post herpetic neuralgia, post operative neruopathic pain, low back pains, traumatic nerve lesions, neuropathy due to chemotherapy, pain neuropathy, hereditary neuropathy, pesticide induced neuropathy; temporomandibular disorders, fibromyalgia, vulvar vestibulitis, irritable bowel, tension type headache, musculoskeletal pain, pain in the chronic fatigue syndrome, post-operative pain, chronic post-operative pain and any variation thereof. 5. A method as in claim 1, wherein a stimulus for said DNIC test is selected from the group consisting of: thermal, electrical, mechanical and chemical and any combination thereof. 6. A method as in claim 1, wherein a stimulus for said DNIC test is evoked by a means selected from the group consisting of: water bath, heating surface, a soft surface on which is contacted with said subject and it either heats or cools, a plurality of small surfaces in contact with said subject each heating or cooling, an array of electrodes that give electrical stimuli, a blood pressure cuff to be inflated for a short time, a thermal glove, combination and variation thereof. 7. A method as in claim 1, further comprising a step selected from the group consisting of: subjecting said subject to repetitive DNIC tests;subjecting said subject to additional non-DNIC testing;calculating a primary predictive value;calculating further predictive values, andcalculating a predictive efficacy indicator. 8. A method as in claim 7, wherein said additional non-DNIC testing is selected from the group consisting of: tests based on pain perception—TS test, QST tests, pain thresholds, supra-threshold stimuli, post-operative (PO) pain; tests that are not based on pain perception—personality, age, and gender. | - | A61B005/00
US07835556 | B2 | 2010-11-16 | System and method for diagnosing breast cancer | 2005-04-29 | The invention relates to a system and a method for the investigation of a body volume (3), particularly for the diagnosis of breast cancer According to the method, a sequence of X-ray projections (P, P′) from different directions is produced by a rotatable X-ray source (1) and a stationary digital X-ray detector (5). From these projections (P, P′), a set of sectional images (a, b, c, d) is calculated by tomosynthesis. A physician may indicate a suspicious structure on a reference image (R) that is derived from one of the projections (P, P′) or sectional images (a, b, c, d) and displayed on a monitor (6). The computer (7) may then locate the structure on all sectional images (a, b, c, d) and calculate the similarity of a corresponding image feature. The sectional 10 image at which the similarity is strongest then indicates the depth (zo) at which the structure (4) is positioned in the body volume (3). Based on this information, a biopsy device (9) with a needle (10) can be advanced into the body volume (3) until it reaches the suspicious structure (4). | 1. System for the investigation of a body volume for diagnosing breast cancer, comprising: an X-ray device for generating projections of the body volume from different directions;a data processing unit coupled to said X-ray device and adapted to reconstruct, from said projections, sectional images of the body volume at different depths, wherein each sectional image corresponds to a certain depth of the body volume; andan indication device for interactively indicating a structure of interest in a reference image, the reference image being derived from one selected from the group consisting of (i) said projections, (ii) said sectional images, and (iii) both said projections and said sectional images,wherein the data processing unit is further adapted to estimate a depth of said indicated structure of interest in the body volume with use of said sectional images and to reconstruct additional sectional images of the body volume around the estimated depth for a fine-tuning of the depth estimation. 2. The system according to claim 1, wherein further the data processing unit is adapted to: a) determine a reference feature vector of said indicated structure of interest on the reference image;b) determine a feature vector at corresponding locations of said indicated structure of interest in all sectional images and quantify its similarity to the reference feature vector; andc) determine the estimated depth of said indicated structure of interest as a depth at which said similarity has an extreme value. 3. The system according to claim 1, wherein the X-ray device comprises (i) an X-ray source that can be moved along a given trajectory, wherein the trajectory includes an arc or a line, and (ii) a digital X-ray detector. 4. The system according to claim 1, wherein the processing unit is adapted to generate tomosynthetic images of the body volume. 5. The system according to claim 1, further comprising: a display device for displaying one selected from the group consisting of (i) said projections, (ii) said sectional images, and (iii) both said projections and said sectional images of the body volume. 6. The system according to claim 1, further comprising: an input device for inputting a continuous value which controls a choice of sectional image to be processed. 7. The system according to claim 1, further comprising: a biopsy device with a needle for extracting a sample of the structure of interest, said biopsy device being controlled by the data processing unit. 8. System for the investigation of a body volume for diagnosing breast cancer, comprising: an X-ray device for generating projections of the body volume from different directions;a data processing unit coupled to said X-ray device and adapted to reconstruct sectional images of the body volume at different depths; andan indication device for interactively indicating a structure of interest in a reference image which is derived from one selected from the group consisting of (i) said projections,... | Koninklijke Philips Electronics N.V. | A61B005/05 G06K009/00
US07833597 | B2 | 2010-11-16 | Polymeric marker with high radiopacity for use in medical devices | 2007-10-24 | High radiopacity is achieved in a polymeric marker by combining a polymeric resin, a powdered radiopaque agent having uniformly shaped particles of a specific particle size distribution and a wetting agent. The method to produce the marker calls for the blending and pelletization of these materials followed by extrusion onto support beading. The resulting supported tubing is subsequently cut to length with the beading still in place. After ejection of the beading remnant the marker is slipped into place on the device to be marked and attached by melt bonding. Marking of a guidewire allows lesions to be measured while the marking of balloon catheters allow the balloon to be properly positioned relative to a lesion. | 1. An intraluminal medical device, comprising an elongated shaft or core member having a radiopaque marker circular band secured to a surface of the shaft or core member, and the radiopaque marker is formed of a compounded polymer mixture comprising: a) a polymer;b) radiopaque particles having an average diameter of at least about 2 microns and a maximum diameter of about 20 microns, the radiopaque particles comprising about 90.8 to about 93.2 weight percent of the compounded polymer mixture; andc) a wetting agent for facilitating encapsulation of said particles by said polymer, and wherein the marker band has a radiopacity that is substantially constant along the length of the marker band, and a double wall thickness that varies by less than 0.001 inches along the entire length of the marker band. 2. The intraluminal medical device of claim 1 wherein the double wall thickness of the radiopaque marker is about 0.004 to about 0.006 inches. 3. The intraluminal medical device of claim 1 wherein the double wall thickness of the radiopaque marker varies by about ±0.0005 inches along the entire length of the marker band. 4. The intraluminal medical device of claim 1 wherein the radiopaque marker is secured to the shaft or core member with an adhesive. | Advanced Cardiovascular Systems, Inc. | B29C047/00 A61M031/00 A61B005/05
US07850616 | B1 | 2010-12-14 | Determination of diastolic heart failure | 2005-06-15 | An exemplary method includes detecting a change in state of a cardiac valve, detecting elongation of the left ventricle substantially along its major axis, determining a time difference between the change in state of the cardiac valve and the elongation of the left ventricle and, based at least in part on the time difference, deciding whether a diastolic abnormality exists. Other exemplary methods, devices, systems, etc., are also disclosed. | 1. A method, implemented by an implantable device, the method comprising: detecting a change in state of a cardiac valve wherein a state of a cardiac valve comprises a closed state or an open state;detecting elongation of the left ventricle substantially along its major axis due to relaxation of the left ventricle;determining a time difference between the change in state of the cardiac valve and the elongation of the left ventricle;comparing the time difference between the change in state of the cardiac valve and the elongation of the left ventricle to a threshold; andbased at least in part on the comparison, deciding whether a diastolic abnormality exists. 2. The method of claim 1 wherein the change in state of a cardiac valve corresponds to a change in state of the mitral valve. 3. The method of claim 1 wherein the change in state of a cardiac valve corresponds to a change in state of the aortic valve. 4. The method of claim 1 wherein the detecting a change in state comprises sensing pressure using an implantable pressure sensor. 5. The method of claim 1 wherein the detecting a change in state comprises sensing vibration using an implantable vibration sensor. 6. The method of claim 1 wherein the detecting elongation comprises sensing impedance using an implantable impedance sensor. | Pacesetter, Inc. | A61B005/021 A61B005/0225
US07828739 | B2 | 2010-11-09 | Apnea detection system | 2005-12-21 | A method and apparatus are described for utilizing a source of vascular pulse waveform data from a patient for the purpose of measuring pulsus paradoxus. The arterial pulse waveform data source described is a pulse oximeter plethysmograph but can be any similar waveform data source, including an intra-arterial transducer, external blood pressure transducer, or plethysmograph. Through incorporation of measurements of values, such as an area under a pulse waveform curve, that are time-domain functions of a change in amplitude of the pulse waveform over a duration of the waveform, embodiments of the present invention represent a significant improvement upon previously described methods of measuring pulsus paradoxus. | 1. A method, for detecting apnea in a mammal, comprising: obtaining data indicative of pulsatile cardiovascular behavior from the mammal, the data comprising at least a first and a second pulse waveform;determining, by a processor, a first value of a function of a change in amplitude of the first pulse waveform, with respect to time, over at least a partial duration of the first pulse waveform;determining a second value of a function of a change in amplitude of the second pulse waveform, with respect to time, over at least a partial duration of the second pulse waveform; anddetermining at least one of a probability of apnea, a presence of apnea, and an absence of apnea based on a comparison of the first value and the second value;wherein said first pulse waveform is generated during one of either a respiratory inspiration and a respiratory expiration of said mammal, and said second pulse waveform is generated during the other of the inspiration and the expiration of said mammal. 2. The method of claim 1, further comprising measuring pulsus paradoxus from a comparison of said first value and said second value. 3. The method of claim 1, wherein the function of a change in amplitude of the first pulse waveform comprises a slope. 4. The method of claim 1, wherein the function of a change in amplitude of the first pulse waveform comprises an integral. 5. The method of claim 1, wherein the function of a change in amplitude of the first pulse waveform comprises an AUC. 6. The method of claim 1, wherein the comparing comprises correlating a magnitude of the difference between the first value and second value with at least one of a predetermined magnitude and range of pulsus paradoxus. 7. The method of claim 6, wherein the magnitude of pulsus paradoxus is used to determine the presence of apnea. 8. The method of claim 1, wherein the comparison of the first and second values is used to determine the presence of central apnea. 9. The method of claim 1, wherein the comparison of the first and second values is used to determine the presence of obstructive apnea. 10. The method of claim 1, wherein the comparison of the first and second values is used to determine the presence of mixed apnea. 11. The method of claim 1, wherein the comparison of the first and second values is used to monitor an infant. 12. The method of claim 1, further comprising providing an alarm to indicate that the presence of apnea is detected. 13. The method of claim 1, further comprising alerting a healthcare professional when the presence of apnea is detected. 14. A method for measuring pulsus paradoxus in a mammal, the method comprising: obtaining data indicative of pulsatile cardiovascular behavior from the mammal, the data comprising a first plethysmographic or pressure amplitude obtained at a first time interval after an onset of a first pulse, and a second plethysmographic or pressure amplitude obtained at a second time interval after the onset of the first pulse;deter... | Precision Pulsus, Inc. | A61B005/00
US07860552 | B2 | 2010-12-28 | CNS assay for prediction of therapeutic efficacy for neuropathic pain and other functional illnesses | 2005-09-30 | In general, the invention features methods for identifying therapeutic interventions for functional illnesses, e.g., neuropathic pain or depression, and methods for identifying markers for functional illnesses by employing an acute painful stimulus or other stimulus for a functional illness and measuring levels of CNS activity. | 1. A method for identifying a therapeutic intervention for a functional illness, said method comprising the steps of: (a) administering a candidate therapeutic intervention to a subject;(b) administering an acute painful stimulus to said subject; and(c) measuring a level of activation of the central nervous system (CNS) indicative of said functional illness in said subject, wherein a modification of said level of CNS activation relative to a control level of activation characteristic of said functional illness is indicative of the efficacy of said candidate therapeutic intervention for said functional illness, wherein said level of activation is not measured with reference to specific regions of the CNS. 2. The method of claim 1, wherein said functional illness comprises neuropathic pain. 3. The method of claim 2, wherein said neuropathic pain comprises post herpetic neuralgia, HIV neuropathy, diabetic neuropathy, Fabry's disease, peripheral neuropathy, trigeminal neuralgia, post incisional neuropathic pain, phantom limb pain, reflex sympathetic dystrophy, causalgia, anesthesia dolorosa, intercoastal neuralgia, post-traumatic localized pain, complex regional pain syndrome, or neuropathic pain caused by trauma, lead, or chemotherapy. 4. The method of claim 1, wherein said acute painful stimulus comprises a change in temperature, mechanical force, a pin prick, or administration of a compound. 5. The method of claim 1, wherein said level of activation is measured by functional magnetic resonance imaging. 6. The method of claim 1, wherein said therapeutic intervention comprises a compound, a physical stimulus, an electrical stimulus, a thermal stimulus, electromagnetic radiation, counseling, or a surgical, medical, or dental procedure. 7. The method of claim 1, wherein said therapeutic intervention is administered subtherapeutically. 8. A method for identifying a therapeutic intervention for a psychiatric disorder, said method comprising the steps of: (a) administering a candidate therapeutic intervention to a subject;(b) administering an acute painful stimulus for said psychiatric disorder to said subject; and(c) measuring the level of CNS activation in said subject, wherein a modification of said level of CNS activation relative to a control level of CNS activation characteristic of said psychiatric disorder is indicative of the efficacy of said candidate therapeutic intervention for said psychiatric disorder, wherein said level of activation is not measured with reference to specific regions of the CNS. 9. The method of claim 8, wherein said psychiatric disorder comprises an affective disorder, an anxiety disorder, attention deficit hyperactivity disorder, a memory disorder, post traumatic stress disorder, a psychotic disorder, a substance abuse disorder, an addiction, a phobia, or obsessive compulsive disorder. 10. The method of claim 9, wherein said psychiatric disorder comprises depression, bipolar disorder, mania, dysphoria, anxiety, ... | The McLean Hospital Corporation | A61B005/05
US07841989 | B2 | 2010-11-30 | Electromagnetic personnel interdiction control method and system | 2005-10-04 | A non-lethal and non-destructive electromagnetic personnel interdiction control stun type weapon system and method utilizes beamed radio frequency energy in a frequency range and modulated to impose a Lorentz force on the vestibular system or sensory cells of a remote human subject sufficient to disrupt the mechanical transduction process and/or the chemical engine by which sound, position and other sensory input are converted to messages by nerve cells and processed by the brain to produce complete disorientation, confusion, and incapacitation sufficient to temporarily render the subject powerless to resist arrest or subjugation. Removal of the electromagnetic energy leaves the nerve cells and surrounding tissues with no damage and second order effects of severe motion sickness and psychological effects of helplessness remains until the subject's body chemistry returns to normal. | 1. A non-lethal method for immobilizing a human subject by utilizing electromagnetic energy to disrupt the subject's sense of balance and perception of body position and movement, comprising the steps of: providing means for generating and transmitting a beam of radio frequency electromagnetic energy having a selected frequency and modulation correlated to a physical and a chemical structure of a human sensory cell to disrupt proper operation of either of a mechanical transduction process or a chemical process by which sensory input is converted by the sensory cell to messages that are processed by a human brain; anddirecting said beam of radio frequency electromagnetic energy from a remote location toward the human subject to be immobilized to impose a Lorentz force on selected sensory cells of the subject sufficient to disrupt proper operation of the cell's mechanical transduction process or chemical process and alter the sensory input converted to messages processed by the subject's brain to adversely affect the perception of body position and movement so as to produce disorientation, confusion, and incapacitation sufficient to immobilize and render the subject powerless to resist arrest or subjugation. 2. The method according to claim 1, wherein said step of directing said beam of radio frequency electromagnetic energy from a remote location toward the human subject to be immobilized comprises imposing a Lorentz force on sensory receptor hair cells of vestibular organs of the subject's inner ear that are connected by auditory nerves operating on exchange of ionized calcium and potassium through semi-permeable cell membranes to produce electrical impulses through the auditory brainstem to the auditory cortex of the brain which interprets the nerve impulses as sound or position relative to gravity; andsaid Lorentz force causing random excitation or movement of said hair cells so as to disrupt the exchange of ionized calcium and potassium and produce sensations of intense wideband sound and confusion between the subject's visual perception and the vestibular sensations transmitted to the subject's brain that affect the subject's perception of the subject's body position relative to gravity resulting in disorientation sufficient to immobilize and render the subject powerless to resist arrest or subjugation. 3. The method according to claim 1, wherein said step of directing said beam of radio frequency electromagnetic energy from a remote location toward the human subject to be immobilized comprises imposing a Lorentz force on cells of the subject's cochlea and retina tissues such that electromagnetic field and subsequent surface charges on the cochlea and retina tissues changes the impedance of their conduction capabilities to disrupt ion flow that regulates nerve pulse output to the subject's brain and produce confusion between the subject's visual perception and the vestibular sensations transmitted to the subject's brain that affect the subj... | Invocon, Inc. | A61B005/00 A61B005/103 A61B005/117 F42B008/00 F42B010/00 F42B012/00 F42B014/06 F42B030/00 F41C009/00 H04B001/02
US07856882 | B2 | 2010-12-28 | System and method for generating three-dimensional density-based defect map | 2008-07-02 | A system for mapping a condition of a structure that includes a plurality of support members covered by a wall is disclosed and includes a computer processor having a memory, a two-dimensional model of the structure stored in the memory and a three-dimensional model generator operatively associated with the computer. A density sensor that includes an ultrasonic transducer is operatively connected to the computer processor and provided with an ultra wideband transmitter. A position locating system is provided for determining the position of the ultra wideband transmitter, and hence the density sensor, in a frame of reference and communicating the position to the computer processor which displays indications of density on a three-dimensional model of the structure. A method for mapping a condition of a structure is also disclosed. | 1. A system for mapping a condition of a structure comprising a plurality of support members covered by a wall comprising: a computer processor having a memory;a position locating system for determining the position of a marker in a frame of reference and communicating said position to said computer processor;a density sensor in communication with said computer processor and, wherein said marker is associated with said density sensor and wherein said density sensor is operatively connected to a signal processor for processing a signal received from said sensor and determining the density of a material that said signal has traversed relative to a known reference density. 2. The system of claim 1 wherein said position locating system comprises a plurality of receivers located at known positions within said frame of reference. 3. The system of claim 2 wherein said receivers comprise ultra wideband receivers. 4. The system of claim 3 wherein said marker comprises a transmitter. 5. The system of claim 4 wherein said marker comprises an ultra wideband transmitter. 6. The system of claim 4 wherein said density sensor comprises an ultrasonic transducer. 7. The system of claim 1 wherein said position locating system further includes a reference transmitter. 8. The system of claim 1 further including a two-dimensional representation of the structure stored in said memory. 9. The system of claim 8 including a three-dimensional model generator operatively associated with said computer. 10. A system for mapping a condition of a structure comprising a plurality of support members covered by a wall comprising: processor means;position locating means for determining the position of a marker in a frame of reference and communicating said position to said processor means;density sensing means operatively connected to said processor means and, wherein said marker is associated with said density sensing means and wherein said density sensing means comprises a signal processor for processing a signal received from said sensor and determining the density of a material that said signal has traversed relative to a known reference density. 11. The system of claim 10 wherein said position locating means comprises a plurality of receiver means. 12. The system of claim 11 wherein said density sensing means comprises an ultrasonic transducer. 13. The system of claim 12 wherein said marker comprises an ultra wideband transmitter. 14. The system of claim 1, wherein position locating system is a global positioning system. 15. The system of claim 10, position locating means is a global positioning system. 16. The system of claim 1, wherein the density sensor determines the density of a material and the density is an actual density of the material. 17. the system of claim 1, wherein the density sensor determines the density of a material and the density is a relative density of the material. 18. The method of claim 10, wherein the density sensing means determines... | - | G01N029/04 A61B005/05
US07831288 | B1 | 2010-11-09 | Method for mapping potential distribution of a heart chamber | 2000-06-07 | A method of mapping electrical activity of a patient's heart. The system includes a set of electrodes spaced from the heart wall and a set of electrodes in contact with the heart wall. Voltage measurements from the electrodes are used to generate three-dimensional and two-dimensional maps of the electrical activity of the heart. | 1. A method of mapping a volumetric electrical potential distribution of a heart chamber arising from electrical activation in a myocardium comprising the steps of: a) positioning an electrode array within the heart chamber;b) determining heart chamber volume and shape;c) computing the position of said array within said heart chamber;d) measuring electrical potentials on said array;e) computing the three-dimensional volumetric electrical field distribution of said heart chamber volume from a spherical harmonic series expression containing said measured electrical potentials, and said array position; andf) displaying said volumetric electrical field distribution. 2. The method of claim 1, wherein step f) further comprises displaying the volumetric electrical field distribution on a display surface representing the heart chamber volume and shape. 3. The method of claim 2, further comprising displaying the position of the array within said heart chamber on the display generated in step f). 4. The method of claim 1, wherein step f) further comprises displaying the volumetric electrical field distribution on a wire grid display surface representing the heart chamber volume and shape. 5. The method of claim 4, wherein step f) further comprises displaying the volumetric electrical field distribution as a set of continuously filled color scale areas superimposed on the wire grid display surface representing the heart chamber volume and shape. 6. The method of claim 1, wherein step f) further comprises displaying the volumetric electrical field distribution as a set of iso-potential lines on a display surface representing the heart chamber volume and shape. 7. The method of claim 1, wherein step f) further comprises displaying the volumetric electrical field distribution as a set of continuously filled color scale areas on a display surface representing the heart chamber volume and shape. | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/042
US07831294 | B2 | 2010-11-09 | System and method of surgical imagining with anatomical overlay for navigation of surgical devices | 2004-10-07 | A system and method are provided for control of a navigation system for deploying a medical device within a subject, and for enhancement of a display image of anatomical features for viewing the projected location and movement of medical devices, and projected locations of a variety of anatomical features and other spatial markers in the operating region. The display of the X-ray imaging system information is augmented in a manner such that a physician can more easily become oriented in three dimensions with the use of a single-plane X-ray display. The projection of points and geometrical shapes within the subject body onto a known imaging plane can be obtained using associated imaging parameters and projective geometry. | 1. A medical imaging system for imaging an operating region in a subject, the system comprising an imaging beam source, an imaging plate, and a display for displaying an image generated by the imaging beam on the imaging plate together with an overlay of at least one object in the operating region, wherein the system is configured to generate an overlay from the position of the object, the position of the imaging beam source, and the position and orientation of the imaging plate, wherein the location of where an overlay of the projected object is to appear on the display is determined in part based on the location of the imaging beam source and the location of the imaging plate, such that for any object in three-dimensional space a corresponding projected object may be suitably overlayed on the display. 2. The medical imaging system according to claim 1 wherein the object is at least one point in the operating region. 3. The medical imaging system according to claim 1 wherein the object is a loop comprising a plurality of points in the operating region. 4. The medical imaging system according to claim 1 wherein the object is a medical device in the operating region. 5. The medical imaging system according to claim 1 wherein the position of the medical device in the operating region is determined by a localization system. 6. The medical imaging system according to claim 1 wherein the position of the medical device is calculated using a mathematical model of the medical device. 7. The medical imaging system according to claim 1 wherein the imaging beam source and the imaging plate are movable about the operating region, and wherein the display updates both the displayed image and the overlay as the positions of the imaging beam and imaging plane change. 8. A method of providing medical images of an operating region in a subject's body, the method comprising: generating an image of the operating region using an imaging beam source and an imaging plate; and displaying the generated image together with an overlay of at least one object in the operating region, the overlay generated from the position of the object, the position of the imaging beam source, and the position and orientation of the imaging plate, wherein the location of where an overlay of the projected object is to appear on the display is determined in part based on the location of the imaging beam source and the location of the imaging plate, such that for any object in three-dimensional space a corresponding projected object may be suitably overlayed on the display. 9. The medical imaging system according to claim 8 wherein the object is at least one point in the operating region. 10. The medical imaging system according to claim 8 wherein the object is a loop comprising a plurality of points in the operating region. 11. The medical imaging system according to claim 8 wherein the object is a medical device in the operating region. 12. The medical imaging system according to ... | Stereotaxis, Inc. | A61B005/05
US08301236 | B2 | 2012-10-30 | System and method for high resolution wireless full disclosure ECG episode monitoring and analysis | 2010-05-06 | High resolution full disclosure ECG data is transferred from a body sensor device to a handheld device via a wireless protocol. The handheld device transfers the full disclosure ECG data via a cellular network to a data center for analysis. A monitoring center communicating with the data center facilitates technician and physician review and interactions. | 1. A system for use by a patient having access to a cellular network comprising: a body sensor device adapted to be worn by the patient and having a sensor circuit, a sensor processor, a sensor storage memory and a sensor transmitter, wherein the sensor circuit detects an analog full disclosure ECG signal of the patient, wherein the sensor processor stores full disclosure ECG data corresponding to the detected analog signal in the sensor storage memory, and wherein the sensor transmitter transmits a full disclosure ECG signal including the full disclosure ECG data stored in the sensor storage memory;a handheld device adapted to be carried by the patient and having a handheld storage memory, a handheld processor, a handheld receiver and a handheld transmitter, wherein the handheld receiver receives the full disclosure ECG signal, wherein the handheld processor stores the full disclosure ECG data included in the received full disclosure ECG signal in the handheld storage memory, and wherein the handheld transmitter transmits a packet signal including the full disclosure ECG data stored in the handheld storage memory via the cellular network;a data center remote from the patient and having a data center receiver, a data center processor and a data center storage memory, wherein the data center receiver receives the packet signal, wherein the data center processor stores the full disclosure ECG data included in the received packet signal in the data center storage memory, and wherein the data center processor analyzes the full disclosure ECG data stored in the data center storage memory to identify any anomalies in the full disclosure ECG data stored in the data center storage memory;a monitoring center linked to the data center and having a technician port accessible by a technician device for evaluating the full disclosure ECG data stored in the data center storage memory, for considering any identified anomalies and for providing reports, and a physician port accessible by a physician for viewing provided technician reports and for viewing the full disclosure ECG data stored in the data center storage memory,whereby the system has the ability to telemeter full disclosure ECG data from remote locations. 2. The system of claim 1 wherein the sensor transmitter streams the full disclosure ECG signal including the full disclosure ECG data in near real-time almost immediately after the detected full disclosure ECG data is stored in the sensor storage memory, wherein the handheld transmitter streams the packet signal including the full disclosure ECG data almost immediately after the full disclosure ECG data is stored in the handheld storage memory and wherein the data center streams the full disclosure ECG data included in the received packet signal to the technician port for near real-time viewing by the technician and/or wherein the data center streams the full disclosure ECG data included in the received packet signal to the physician port for near ... | Biomedical Systems Corporation | A61B005/0404
US08140156 | B2 | 2012-03-20 | Heart sound sensing to reduce inappropriate tachyarrhythmia therapy | 2009-06-30 | Techniques for detecting heart sounds to reduce inappropriate tachyarrhythmia therapy are described. In some examples, a medical device determines that a cardiac rhythm of the patient is treatable with a therapy, such as a defibrillation pulse, based on a cardiac electrogram (EGM). The medical device analyzes detected heart sounds, and withholds or allows the therapy based on the analysis of the heart sounds. | 1. A method comprising: determining that a cardiac rhythm of a heart of a patient is treatable by delivery of an antitachyarrhythmia therapy from a medical device to the patient based on an analysis of electrical activity of the heart of the patient;receiving a heart sound signal from a heart sound sensor, the heart sound signal representing sounds generated by the heart and bloodflow of the patient; detecting a plurality of heart sounds within the heart sound signal;classifying each of the detected heart sounds as one of a first classification or a second classification based on one or more features of the detected heart sounds;selectively delivering or withholding the antitachyarrhythmia therapy based on the classifications of the detected heart sounds; andgenerating an envelope signal of the heart sound signal, wherein detecting a plurality of heart sounds within the heart sound signal comprises detecting a plurality of heart sounds within the envelope signal, wherein generating the envelope signal comprises:bandpass filtering the heart sound signal;rectifying the filtered signal;smoothing the rectified signal; andapplying a box-car filter to the smoothed signal to generate the envelope signal. 2. The method of claim 1, wherein the first classification comprises a normal classification and the second classification comprises an abnormal classification, and selectively delivering or withholding the antitachyarrhythmia therapy comprises withholding the tachyarrhythmia therapy if at least M of the last N detected heart sounds is classified as normal. 3. The method of claim 1, wherein the antitachyarrhythmia therapy comprises a defibrillation pulse. 4. The method of claim 1, wherein determining that the cardiac rhythm is treatable by delivery of the antitachyarrhythmia therapy from the medical device, receiving the heart sound signal, detecting the plurality of heart sounds, classifying each of the detected heart sounds, and selectively delivering or withholding the antitachyarrhythmia therapy comprises determining that the cardiac rhythm is treatable by delivery of the antitachyarrhythmia therapy, receiving the heart sound signal, detecting the plurality of heart sounds, classifying each of the detected heart sounds, and selectively delivering or withholding the antitachyarrhythmia therapy by the medical device. 5. The method of claim 4, wherein the medical device comprises an implantable cardiac device. 6. The method of claim 1, further comprising delivering a plurality of pacing pulses to the heart, wherein receiving the heart sound signal comprises receiving the heart sound signal during the delivery of the pacing pulses. 7. The method of claim 6, wherein delivering the plurality of pacing pulses comprises delivering antitachyarrhythmia pacing. 8. A method comprising: determining that a cardiac rhythm of a heart of a patient is treatable by delivery of an antitachyarrhythmia therapy from a medical device to the patient based on an analy... | Medtronic, Inc. | A61N001/00
US08016756 | B2 | 2011-09-13 | Medical system and method of switching antenna | 2010-04-28 | A medical system includes a capsule type medical apparatus introduced into a subject and having an antenna having directionality and a communication unit which executes a communication through the antenna, an external antenna, which is disposed outside of the subject and has an antenna having directionality and an antenna disposed together with the antenna and having directionality in a direction different from the antenna, and an external machine having a selection unit which selects any of the antennas in the external antenna and a second communication unit which executes a communication with the first communication unit through the antenna whose polarized wave surface is selected by the selection unit. | 1. A medical system comprising: a capsule type medical apparatus that is introduced into a subject and includes a first antenna which outputs a radio wave having directionality in a first polarized wave direction and a first communication unit which executes a communication through the first antenna;an external antenna that is disposed outside of the subject and includes a second antenna having directionality to a second polarized wave direction and a third antenna that is disposed together with the second antenna and has directionality to a third polarized wave direction different from the second polarized wave direction of the second antenna; andan external machine that includes a selection unit which selects any of the second and third antennas and a second communication unit which executes a communication with the first communication unit through the external antenna whose plane of polarization is selected by the selection unit, whereinthe external antenna includes a board of a monolayer or multilayer structure on which the second and third antennas are formed, andthe second antenna and the third antenna are formed on different planes or different layers in the board. 2. The medical system according to claim 1, wherein the external machine includes a reception strength detection unit that detects a reception strength on the plane of polarization selected by the selection unit, andthe selection unit selects any of the second and third antennas based on the reception strength detected by the reception strength detection unit. 3. The medical system according to claim 1, wherein at least one of the second antenna and the third antenna is a dipole antenna, a loop antenna, or a micro strip antenna. 4. The medical system according to claim 1, wherein the board has a plastic property. 5. The medical system according to claim 1, comprising a plurality of external antennas, wherein the selection unit selects any of the second and third antennas as to each of the plurality of external antennas. 6. The medical system according to claim 1, wherein the external antenna is fixed to an outside surface of the subject. 7. The medical system according to claim 1, wherein a center of gravity of the capsule type medical apparatus is offset from a center. 8. The medical system according to claim 1, wherein the first antenna is a coil antenna. 9. The medical system according to claim 1, further comprising: a sensor that detects a tilt angle of the external antenna to the ground surface;wherein the selection unit selects any of the second and third antennas in response to the tilt angle detected by the sensor, and the second communication unit executes a communication with the first communication unit through the external antenna whose plane of polarization is selected by the selection unit. 10. The medical system according to claim 9, wherein the external antenna and the sensor are formed on the same board. 11. The medical system according to claim 9, whe... | Olympus Medical Systems Corp. | A61B005/07 A61B005/00
US08668656 | B2 | 2014-03-11 | Method and apparatus for improving fluidic flow and sample capture | 2004-12-31 | A system and method is provided for the capture of bodily fluid upon lancing of a patient. In one embodiment, the fluid sample capture aperture mesh (320) ring is placed in the pathway of a finger penetrating member (340). The aperture mesh ring has a center clearance area that allows the penetrating member to pierce the skin unobstructed. The aperture mesh ring may contain a series of fluid sampling meshes as to allow the release bodily fluid to “wick” into the fluid sampling meshes for transport to the respective sensor. The invention may also relate to a method of improving the fluidic flow through a membrane mesh structure for the transportation of bodily fluids from a point of sampling to a point of measurement. | 1. A body fluid sampling device for use on a patient, comprising: a cartridge;at least one penetrating member positioned in the cartridge;a controllable penetrating member driver configured to be coupled to the at least one penetrating member;a processor coupled to the controllable penetrating member and providing a controllable velocity of the at least one penetrating member;a sample port for receiving a body fluid in response to a puncturing event by the at least one penetrating member;a sample chamber;a channel coupled to the sample port and the sample chamber, at least a portion of the channel being a capillary channel with a hydrophilic surface positioned adjacent to the sample chamber;at least one electrode positioned in the sample chamber and configured to determine an amount of an analyte in the body fluid; anda mesh membrane positioned in the channel and at the sample port or adjacent to the sample port to guide a flow of the body fluid through at least a portion of the capillary channel with the mesh providing that the body fluid contacts the channel regardless of an orientation of body fluid sampling device, the mesh membrane and channel configured to provide a transport of the body fluid to the sample chamber and a measurement in less than 10 seconds. 2. The device of claim 1 further comprising a plurality of electrodes each having a sensory material. 3. The device of claim 1 wherein a sensory material is mounted on a plurality of electrodes. 4. The device of claim 1 a plurality of sets of electrodes are associated with each penetrating member. 5. The device of claim 1 wherein the mesh membrane has a lollipop configuration. 6. The device of claim 1 wherein the mesh membrane is oriented perpendicular to a path of the penetrating member. 7. The device of claim 1 wherein the mesh membrane is oriented to intersect a path of the penetrating member. 8. The device of claim 1 wherein topside sections of the capillary portion of the channel include a PET film hydrophobic on an outer most layer and hydrophilic on an inner layer abutting against a hydrophobic double-sided adhesive layer. 9. The device of claim 1 wherein bottom side sections of the capillary section include a PET film hydrophilic on the inner layer abutting against a hydrophobic adhesive and hydrophobic on the outside, wherein an inner fluidic channel region is a sandwich structure of top PET film/fluidic mesh structures/and bottom PET film, wherein PET surfaces abutting the mesh structures are hydrophilic. 10. The device of claim 1 wherein a plurality of wicking materials are positioned in a ring configuration around the cartridge. 11. The device of claim 1 wherein mesh membrane is also positioned in a ring configuration around the cartridge. 12. A method of sampling a body fluid, comprising: providing a fluid sampling device that includes a cartridge, at least one penetrating member positioned in the cartridge, a controllable penetrating driver, a processor, a sample... | Sanofi-Aventis Deutschland GmbH | A61B005/00 B65D081/00 A61B017/14 A61B017/32
US08515537 | B2 | 2013-08-20 | Tissue oxygenation monitoring in heart failure | 2010-06-10 | A medical device system and associated method control the delivery of a therapy to a patient. The system includes an activity sensor and detects a change in activity level of the patient. The system further include an optical sensor to sense signal corresponding to tissue light attenuation. The system computes a tissue oxygenation measurement in response to detecting a change in activity level. A parameter controlling delivery of the therapy is adjusted in response to detecting the decreased tissue oxygenation. | 1. A method for controlling a therapy in a patient, comprising; sensing an activity sensor signal;detecting a change in activity level of the patient from the activity sensor signal;sensing an optical sensor signal corresponding to tissue light attenuation in a tissue volume;computing a tissue oxygenation measurement from the optical sensor signal in response to detecting the change in activity level, the tissue oxygenation measurement comprising contributions from arterial and venous blood in a measurement volume, wherein computing the tissue oxygenation measurement comprises computing a trend;detecting a decrease in tissue oxygenation in response to the tissue oxygenation measurement, wherein detecting a decrease in the tissue oxygenation measurement comprises:defining a plurality of thresholds for each one of a respective plurality of activity levels,selecting one of the plurality of thresholds in response to detecting the change in activity; andcomparing the tissue oxygenation measurement and the selected one of the plurality of thresholds; andadjusting a parameter controlling delivery of the therapy in response to detecting the decreased tissue oxygenation. 2. The method of claim 1, wherein adjusting the parameter comprises adjusting the parameter to cause a change in the trend. 3. The method of claim 1, wherein detecting the change in activity comprises detecting a change that causes an adjustment to a sensor-indicated pacing rate. 4. The method of claim 1, wherein detecting a decrease in the tissue oxygenation measurement comprises: defining a plurality of thresholds for each one of a respective plurality of heart rates;sensing a cardiac signal;measuring a heart rate of the patient in response to the cardiac signal;selecting one of the plurality of thresholds based on the measured heart rate; andcomparing the tissue oxygenation measurement and the selected one of the plurality of thresholds. 5. The method of claim 1, wherein the plurality of thresholds comprises a first threshold corresponding to a first activity level and a second threshold corresponding to a second activity level, the first activity level being a higher level of activity than the second activity level, and the first threshold being greater than the second threshold. 6. The method of claim 1, wherein detecting a decrease in the tissue oxygenation measurement comprises storing a previous tissue oxygenation measurement and comparing the tissue oxygenation measurement and the previous tissue oxygenation measurement. 7. The method of claim 1, further comprising: storing a plurality of parameter settings for each of a plurality of activity levels;adjusting the parameter by selecting one of the plurality of settings based on the detected change in activity. 8. The method of claim 1, further comprising: sensing a cardiac signal;measuring a heart rate of the patient in response to the cardiac signal;storing a plurality of parameter settings for each of a plurality of heart... | Medtronic, Inc. | A61N001/00
US08679026 | B2 | 2014-03-25 | Medical device for detecting pulmonary artery pressure | 2007-11-08 | In a medical device and method to monitor pulmonary artery pressure of a patient, a first parameter related to the right ventricular straight volume of the patient's is detected, and a second parameter related to the right ventricular ejection rate of the patient's heart, or related to the workload of the patient's heart, is also determined. A pulmonary pressure index is determined by combining the first and second parameters, with variations of the pulmonary pressure index indicating variations in the pulmonary artery pressure. Pulmonary artery hypertension can be monitored with such a device and method. | 1. A medical device to monitor pulmonary artery pressure (PAP) of a patient, comprising: a first detector configured to measure and determine a first parameter (SVR) indicative of a right ventricular stroke volume of the patient's heart;a second detector configured to measure and determine a second parameter (ERR) indicative of speed of blood ejected from a right ventricle of the patient's heart; anda processing unit configured to determine a pulmonary pressure index (IP) by combining the first and second parameters, with variations of the pulmonary pressure index indicating variations of PAP. 2. Medical device according to claim 1, wherein said processing unit is configured to use the determined pulmonary pressure index (IP) to detect pulmonary artery hypertension (PAH). 3. Medical device according to claim 1, wherein said processing unit is configured to use the determined pulmonary pressure index (IP) to monitor left ventricular end diastolic pressure. 4. Medical device according to claim 1, wherein said processing unit is configured to determine the pulmonary pressure index as a ratio between the first parameter and the second parameter. 5. Medical device according to claim 1, further comprising a control unit that controls said first and second detectors to respectively measure said first and second parameters during a predetermined time-period. 6. Medical device according to claim 5 wherein the control unit controls said first and second detectors to respectively measure said first and second parameters as consecutive measurements performed at regular intervals. 7. Medical device according to claim 1, further comprising a control unit that controls said first and second detectors to respectively measure said first and second parameters during a predetermined number of heart cycles. 8. A medical device as claimed in claim 7 wherein the control unit controls said first and second detectors to respectively measure said first and second parameters as consecutive measurements performed at regular intervals. 9. Medical device according to claim 1, further comprising a storage unit, in communication with the output of the processing unit, that stores values that are determined for the pulmonary pressure index during a number of heart cycles. 10. Medical device according to claim 1, further comprising a comparator that compares a calculated IP to one or many specified PAP thresholds representing different levels of PAP, and generates a PAP indication signal in result of said comparison. 11. Medical device according to claim 10, wherein the comparator employs, as said one or many specified PAP thresholds, a threshold that represents a dangerous level of PAP, and when the IP value is higher than the PAP threshold an alert function is activated that generates an alarm signal to notify the patient. 12. Medical device according to claim 10, wherein at least one of the specified PAP thresholds represents pulmonary artery hypertension (PAH). 1... | St. Jude Medical, AB | A61B005/0215 A61B005/021
US08473040 | B2 | 2013-06-25 | Electrocardiographic device and method | 2008-12-19 | A hand-held device for acquiring electrocardiographic (ECG) signals from a patient comprises an electrode array with a gripping handle (15) and four electrodes (11-14). The electrodes typically form a quadrilateral array, the array being adapted for placement on the patients thoracic region without the need for upper clothing or underwear to be removed. The device can be used in a method to rapidly acquire ECG data relating to atrial fibrillation and/or flutter without the need for a conventional 12-lead ECG apparatus. The data can be transmitted to a computer for further manipulation and analysis. A kit for putting the method into effect includes the device, a means for communication with a computer and a computer program for running on the computer. | 1. A hand-held device for acquiring electrocardiographic signals from a patient, comprising: a handle constructed and arranged for manual gripping of the device;an electrode array comprising four electrodes adapted for placement on a thoracic region of the patient whereby electrical contact may be made with mid-anterior thoracic skin of the patient for acquiring electrocardiographic signals through the four electrodes; anda first electrode housing carrying a first pair of the four electrodes and a second electrode housing carrying a second pair of the four electrodes,wherein each electrode is spaced from 4 cm to 20 cm from each adjacent electrode,wherein the first and second electrode housings are constructed and arranged to be mutually spaced to provide a gap therebetween, andwherein the gap is arranged to provide a void having a width of at least 1 cm and a depth of at least 1 cm when the four electrodes are positioned on said mid-anterior thoracic region of said patient. 2. The device of claim 1, wherein one or more of the electrodes is oriented or orientable with respect to the hand-held device whereby contact with said mid-anterior thoracic region of said patient may be optimized. 3. The device of claim 1, wherein all of the electrodes are orientable with respect to the device. 4. The device of claim 1, wherein one or more electrode is pivotally mounted. 5. The device of claim 1, wherein one or more electrode is positioned on a surface of a deformable pad. 6. The device of claim 5, wherein the deformable pad comprises a polymeric skin holding a silicone gel. 7. The device of claim 1, wherein the gap is arranged to permit placement of the four electrodes onto said mid-anterior thoracic region of said patient whilst said patient is wearing a brassiere. 8. The device of claim 1, wherein the four electrodes of the electrode array are positioned substantially at the vertices of a quadrilateral. 9. The device of claim 1, wherein the electrodes comprise a gold surface arranged to make electrical contact with said patient's mid-anterior thoracic skin. 10. The device of claim 1, further comprising an actuating means adapted to initiate capture of electrocardiographic signals over a predetermined time period onto a data storage means as an electrocardiographic dataset. 11. The device of claim 10, wherein the electrocardiographic dataset comprises six channels. 12. The device of claim 10, wherein the device comprises the data storage means. 13. The device of claim 1, further comprising a display means adapted to display instantaneous electrocardiographic signals from said patient. 14. The device of claim 1, further comprising a data transfer means for transferring the electrocardiographic signals or an electrocardiographic dataset to a first computing means. 15. The device of claim 14, wherein the data transfer means comprises a power supply means adapted to provide electrical power to the device. 16. The device of claim 14, wherein the ... | Central Manchester University Hospitals NHS Foundation Trust | A61B005/02 A61N001/00
US09173613 | B2 | 2015-11-03 | Method of autonomously monitoring the movement activity of a fetus | 2008-05-07 | The present invention relates to a system (1) and method for automatically monitoring the movement activity of a fetus. In order to provide a simple and reliable technique for monitoring the movement activity of a fetus, a method of automatically monitoring the movement activity of a fetus is suggested, the method comprising the steps of: detecting (200) the overall movements of a pregnant female, said overall movements comprising movements of the pregnant female and movements of the fetus; detecting at least one second physiological signal (26) of the pregnant female; and—determining (300) the movement activity of the fetus by analyzing the detected overall movements of the pregnant female, analyzing the at least one second physiological signal (26) of the pregnant female, and discriminating movements of the pregnant female from movements of the fetus depending on the analyzing results. | 1. A method of automatically monitoring movement activity of a fetus, the method comprising the steps of: detecting overall movements of a pregnant female, said overall movements comprising a combination of movements of the pregnant female and movements of the fetus;detecting at least one second physiological signal of the pregnant female; anddetermining the movement activity of the fetus by analyzing the detected overall movements of the pregnant female, analyzing the at least one second physiological signal of the pregnant female, and discriminating movements of the pregnant female from independent movements of the fetus depending on the analyzing results. 2. The method as claimed in claim 1, wherein the energy of a movement signal of the pregnant female is analyzed. 3. The method as claimed in claim 1, wherein the pregnant female's heart rate and/or heart rate variability is used as second physiological signal. 4. The method as claimed in claim 3, wherein a movement is classified as child movement, if the movement does not correlate with a changing of the pregnant female's heart rate. 5. The method as claimed in claim 1, wherein the movements of the pregnant female are detected during the pregnant female's sleep. 6. A system for automatically monitoring movement activity of a fetus, the system comprising: at least one sensor for detecting overall movements of a pregnant female, said overall movements comprising a combination of movements of the pregnant female and movements of the fetus;at least one sensor for detecting at least one second physiological signal of the pregnant female; andan analyzing device for determining the movement activity of the fetus by analyzing the detected overall movements of the pregnant female and analyzing the at least one second physiological signal of the pregnant female; said analyzing device comprising a discriminator for discriminating movements of the pregnant female from independent movements of the fetus depending on the analyzing results. 7. The system as claimed in claim 6, wherein the at least one sensor for detecting movements and/or the at least one sensor for detecting at least one second physiological signal of the pregnant female is connected to or integrated into a bed. 8. A computer program embodied on a memory and executed by a processor of a computer, for automatically monitoring movement activity of a fetus, said computer program comprising computer instructions to determine the movement activity of a fetus by: analyzing previously detected overall movements of a pregnant female, said overall movements comprising a combination of movements of the pregnant female and movements of the fetus;analyzing at least one previously detected second physiological signal of the pregnant female; anddiscriminating movements of the pregnant female from independent movements of the fetus depending on the analyzing results. | Koninklijke Philips N.V. | A61B005/00 A61B005/11 A61B005/0245
US08672854 | B2 | 2014-03-18 | System for calibrating a PTT-based blood pressure measurement using arm height | 2009-05-20 | The invention provides a system and method for measuring vital signs (e.g. SYS, DIA, SpO2, heart rate, and respiratory rate) and motion (e.g. activity level, posture, degree of motion, and arm height) from a patient. The system features: (i) first and second sensors configured to independently generate time-dependent waveforms indicative of one or more contractile properties of the patient's heart; and (ii) at least three motion-detecting sensors positioned on the forearm, upper arm, and a body location other than the forearm or upper arm of the patient. Each motion-detecting sensor generates at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which it is affixed. A processing component, typically worn on the patient's body and featuring a microprocessor, receives the time-dependent waveforms generated by the different sensors and processes them to determine: (i) a pulse transit time calculated using a time difference between features in two separate time-dependent waveforms, (ii) a blood pressure value calculated from the time difference, and (iii) a motion parameter calculated from at least one motion waveform. | 1. A system for measuring a blood pressure from a patient, comprising: (a) a first sensor configured to generate a first time-dependent waveform indicative of one or more contractile properties of the patient's heart;(b) an electrical sensor comprising: at least two electrodes configured to detect electrical signals from the patient's body, and an electrical circuit operably connected to the electrodes and configured to process the detected electrical signals, the electrical sensor generating a time dependent electrical waveform indicative of one or more contractile properties of the patient's heart;(c) at least two motion-detecting sensors positioned on separate locations on the patient's body, each of said motion-detecting sensors generating at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which the motion detecting sensor is affixed; and,(d) a processing component configured to be worn on the patient's body and comprising a microprocessor, said processing component configured to receive the first time-dependent waveform, the time-dependent electrical waveform, and the at least one time-dependent motion waveform generated by each motion-detecting sensor, and to determine: (i) a first arm position from one or more of the time-dependent motion waveforms acquired in a first arm position;(ii) a first time difference between a feature present in each of the first time-dependent waveform and the first time-dependent electrical waveform acquired in the first arm position;(iii) a second arm position from one or more of the time-dependent motion waveforms acquired in a second arm position;(iv) a second time difference between a feature present in each of the first time dependent waveform and the first time-dependent electrical waveform acquired in the second arm position;(v) a first blood pressure value calculated from the first time difference;(vi) a second blood pressure value calculated from the second time difference;(vii) a blood pressure factor calculated using the first blood pressure value and the first time difference or a property derived therefrom, together with the second blood pressure value and the second time difference or a property derived therefrom. 2. The system of claim 1, wherein the processing component further comprises a display component. 3. The system of claim 1, wherein the display component renders a graphic interface configured to instruct the patient to move an arm to the first arm position and then to the second arm position. 4. The system of claim 1, wherein the processing component further comprises an audio component. 5. The system of claim 4, wherein the audio component renders audible instructions that instruct the patient to move an arm to the first arm position and then to the second arm position. 6. The system of claim 1, wherein the processing component is further configured to detect when the patient's arm is in the first position by processing at leas... | Sotera Wireless, Inc. | A61B005/021
US08491473 | B2 | 2013-07-23 | Medical procedure mat and draping system | 2008-02-08 | A medical procedure mat for supporting a patient. A mat having a cushioned substrate and one or more physiological monitoring lines, the physiological monitoring lines having guided portions along the outer edge of the cushioned substrate, and flexible portions operatively coupled to the guided portions, the flexible portions being adapted to extend toward the patient. A draping system for use during medical procedures, including a top and/or bottom drape. A bottom drape including an absorbent material and/or a privacy shield. A top drape including a workbench surface. A mat and draping system in which a bottom and top drape mate to each other and/or to the mat. A draping system providing openings or holes to allow passage of one or more physiological monitoring lines when used in conjunction with the mat. | 1. A medical procedure mat comprising: a cushioned substrate having a top surface for supporting a patient lying thereon,a bottom surface adapted to be supported by a medical procedure table,an inner portion, andan outer edge portion comprising a head, a foot, and left and right sides; and a first physiological monitoring line for communicating a physiological signal from a physiological sensor associated with the patient, the first physiological monitoring line comprisinga guided portion having a proximal end and a distal end,a flexible portion, andan operative coupling between the distal end of the guided portion and the flexible portion, the operative coupling being disposed proximate the outer edge portion of the cushioned substrate;the guided portion of the physiological monitoring line being disposed generally along the outer edge portion of the cushioned substrate,the proximal end of the guided portion adapted to be connected to a monitoring device, andthe flexible portion being adapted to extend toward the inner portion of the cushioned substrate, further comprising at least a second physiological monitoring line wherein at least the first and second physiological monitoring lines are housed within a conduit, the conduit being disposed along the outer edge portion and within the cushioned substrate. 2. The mat of claim 1 wherein the first physiological monitoring line is adapted to communicate electrical signals. 3. The mat of claim 1 wherein the first physiological monitoring line is adapted to communicate fluid signals. 4. The mat of claim 1 wherein at least one of the first and second physiological monitoring lines is an ECG cable. 5. The mat of claim 1 wherein at least one of the first and second physiological monitoring lines is a blood pressure monitoring line. 6. The mat of claim 1 wherein at least one of the first and second physiological monitoring lines is an oxygen saturation monitoring line. 7. The mat of claim 1 wherein the conduit comprises two or more discontinuous portions. 8. The mat of claim 1 wherein the conduit is a generally annular member. 9. The mat of claim 8 wherein the annular member has a distal opening disposed proximate the operative coupling of one of the physiological monitoring lines. 10. The mat of claim 9 wherein the annular member further includes one or more side openings, each side opening being adapted to allow for operative coupling between the guided and flexible portions of one of the physiological monitoring lines. 11. The mat of claim 10 wherein one or more operative couplings are positioned on an outside surface of the cushioned substrate. 12. The mat of claim 8 wherein the annular member is a tube adapted to bend to conform to the outer edge portion. 13. The mat of claim 1 wherein the conduit comprises a hollow space formed in the cushioned substrate for guiding the guided portion of the physiological monitoring lines. 14. The mat of claim 13 wherein the outer edge portion of the c... | ACIST Medical Systems, Inc. | A61B005/0402
US08603008 | B2 | 2013-12-10 | Apparatus for generating and collecting saliva samples containing enhanced cell concentrations | 2010-04-06 | An apparatus for generating and collecting saliva samples containing enhanced cell concentrations includes a mouthpiece having a mouth contact portion and a flow channel which may include a tapered basin, a conduit in fluid communication with the flow channel, and one or more collectors couplable to the conduit. The mouth contact portion and basin may be textured to enhance cell removal. The mouth contact portion may include a plurality of internal and external channels disposed around the flow channel. Filters may be included between the flow channel and the collectors. A closure member is provided. | 1. A passive saliva sample collection device operable without need of a vacuum source, comprising: a mouthpiece having a toroidal mouth contact portion including an outer contact surface to go against the interior surfaces of a mouth, and a flow channel through said mouthpiece;a conduit coupled to said mouthpiece, said conduit having an interior channel in fluid communication with said flow channel; and,one or more collectors coupled to said conduit. 2. A passive saliva sample collection device operable without need of a vacuum source, comprising: a mouthpiece having a toroidal mouth contact portion including an outer contact surface to go against the interior surfaces of a mouth, and a flow channel through said mouthpiece;a conduit coupled to said mouthpiece, said conduit including an interior channel in fluid communication with said flow channel and one or more mating ports in fluid communication with said conduit interior channel; and,wherein each of said one or more mating ports is also removably couplable to a collector. 3. A device as in claim 2, further comprising one or more collectors corresponding to said one or more mating ports, each of said one or more collectors removably couplable to a corresponding mating port. 4. A device as in claim 2, wherein said flow channel comprises an open concave basin having a surface tapering from said outer contact surface inward to said conduit, the basin surface forming a continuous contact surface with the outer contact surface. 5. A device as in claim 2 or 4, further comprising: wherein said outer contact surface includes texturing. 6. A device as in claim 5, further comprising: wherein said texturing includes a plurality of parallel ridges. 7. A device as in claim 4, further comprising: wherein said basin surface includes texturing. 8. A device as in claim 2 or 4, further comprising: wherein said mouth contact portion outer surface includes at least opposing front and back contact surfaces, and said mouth contact portion further includes a plurality of internal channels extending through said mouth contact portion from said front contact surface to said back contact surface, said internal channels disposed around said flow channel. 9. A device as in claim 2 or 4, further comprising: wherein said mouth contact portion outer surface includes a plurality of external channels disposed around said flow channel. 10. A device as in claim 2 or 4, further comprising: wherein said one or more mating ports comprises a plurality of mating ports, and said conduit includes a plurality of collector channels corresponding to said plurality of mating ports, each of said plurality collector channels extending within at least a portion of said conduit interior channel, each of said plurality of collector channels in fluid communication with said conduit interior channel and one of said plurality of mating ports. 11. A device as in claim 2 or 4, further comprising: a filter disposed between said mouthpiece ... | Oasis Diagnostics Corporation | A61B005/00
US08414500 | B2 | 2013-04-09 | Arteriosclerosis diagnostic device | 2010-05-20 | An arteriosclerosis diagnostic device according to various embodiments is a simple device, resistant to an external factor, such as an error resulting from a skin surface, and capable of measuring the degree of hardness of an artery. The arteriosclerosis diagnostic device detects a heart sound and a pulse wave at least one location of a living body, the pulse wave propagating through an artery in relation to the heart sound, converts detected signals thereof into respective frequency signals, specifies the peak frequency of each of the frequency signals, and determines the degree of arteriosclerosis on the basis of the difference between the peak frequency of the heart sound and the peak frequency of the pulse wave. Accordingly, the degree of arteriosclerosis can be determined by comparison between the frequency signals. | 1. An arteriosclerosis diagnostic device comprising: first detection module to detect a heart sound;second detection module to detect a pulse wave in at least one location of a living body, the pulse wave propagating through an artery in relation to the heart sound;converter module to convert detection signals detected by the first and second detection modules into their respective frequency signals; anddetermining module to determine the degree of arteriosclerosis by comparing the frequency signal of the heart sound and the frequency signal of the pulse wave. 2. The arteriosclerosis diagnostic device according to claim 1, further comprising an extracting module to extract only the first sound component accompanying systole from the detection signals detected by the first and the second detection modules by using a window function. 3. The arteriosclerosis diagnostic device according to claim 1, wherein the determining module to determine the degree of arteriosclerosis includes: peak-frequency determining module to specify a peak frequency or a center frequency of each of the frequency signals; andcomparing module to determine the degree of arteriosclerosis by comparing the peak frequencies or the center frequencies of the frequency signals. 4. The arteriosclerosis diagnostic device according to claim 3, wherein the determining module to determine the degree of arteriosclerosis determines the degree of arteriosclerosis on the basis of a difference between the peak frequency or the center frequency of one of the frequency signals and the peak frequency or the center frequency of another of the frequency signals. 5. The arteriosclerosis diagnostic device according to claim 3, wherein the determining module to determine the degree of arteriosclerosis determines the degree of arteriosclerosis on the basis of a ratio of the peak frequency or the center frequency of one of the frequency signals to the peak frequency or the center frequency of another of the frequency signals. 6. The arteriosclerosis diagnostic device according to claim 3, wherein the determining module to determine the degree of arteriosclerosis determines the degree of arteriosclerosis on the basis of a ratio of a difference between the peak frequency or the center frequency of the heart sound and the peak frequency or the center frequency of the pulse wave at a first location of a living body to a difference between the peak frequency or the center frequency of the heart sound and the peak frequency or the center frequency of the pulse wave at a second location different from the first location. 7. The arteriosclerosis diagnostic device according to claim 1, wherein each of the first and second detection modules is a piezoelectric transducer. 8. The arteriosclerosis diagnostic device according to claim 2, wherein the determining module to determine the degree of arteriosclerosis includes: peak-frequency determining module to specify a peak frequency or a center frequency of eac... | The Doshisha Murata Manufacturing Co., Ltd. | A61B005/0402
US09974465 | B2 | 2018-05-22 | Capnography device and method | 2008-01-03 | There is provided a method for dynamically determining a breath related parameter, which includes averaging a breath related parameter over a first period of time, calculating a breath related value over a second period of time and adapting the duration of the first period of time according to the calculated breath related value. | 1. A medical monitoring device for dynamically determining a breath related parameter, the medical monitoring device comprising: a medical sensor configured to measure the breath related parameter of a patient and generate a plurality of data during breath cycles of the patient over time, wherein the medical sensor comprises a CO2 sensor;a processor communicatively coupled to the medical sensor and configured to: repeatedly receive the plurality of data measured during the breath cycles from the medical sensor; andfor each new breath cycle received:average the plurality of data measured during n number of the breath cycles over time to determine an average breath related parameter; wherein the average breath related parameter is a weighted average of the plurality of data measured during the breath cycles over time, and wherein n is at least 2;determine a breath related value associated with the plurality of data measured for the n number of the breath cycles over time, wherein the breath related value is indicative of the reliability of the plurality of data measured, and wherein the breath related value is a variance, standard deviation, or variability index for the plurality of data measured;compare the breath related value to a predetermined breath related value threshold wherein the processer adjusts a time period to n+1 number of the breath cycles and re-averages the breath related parameter based on data measured for n+1 number of the breath cycles over time in response to the breath related value being above the predetermined breath related value threshold, wherein the n+1 breath cycles correspond to the last n+1 breath cycles of the breath cycles associated with the updated average breath related parameter, andwherein the processer does not adjust the time period of the breath cycles and displays the average breath related parameter determined during the n number of the breath cycles over time in response to the breath related value being below the predetermined breath related value threshold, and wherein the average breath related parameter is determined using the last n number of breath cycles of the breath cycles associated with the plurality of data measured over time; andoutput an alarm associated with a medical condition of the patient in response to the average breath related parameter being outside a range of first predetermined alarm threshold. 2. The device according to claim 1, further comprising a sampler adapted to sample the breath related parameter. 3. The device according to claim 1, wherein the processor is further adapted to trigger the alarm when the averaged breath related parameter is at or below a second predetermined alarm threshold. 4. The device according to claim 1, wherein said processor is further adapted to trigger the alarm when the averaged breath related parameter is at or above a third predetermined alarm threshold. 5. The device of claim 1, wherein the breath related parameter comprises: respiratory ... | ORIDION MEDICAL 1987 LTD. | A61B005/08 A61B005/083 A61B005/00
US20100312092 | A1 | 2010-12-09 | Method and system for vascular elastography | 2009-11-12 | The method for vascular elastography comprises: i) obtaining a sequence of radio-frequency (RF) images including pre-tissue-motion and post-tissue-motion images in digital form of a vessel delimited by a vascular wall; the pre-tissue-motion and post-tissue-motion images being representative of first and second time-delayed configuration, of the whole vessel; ii) partitioning both the pre-tissue-motion and post-tissue-motion images within the vascular wall into corresponding data windows; approximating a trajectory between the pre- and post-tissue-motion for corresponding data windows; and using the trajectory for each data window to compute the full strain tensor in each data window, which allow determining the Von Mises coefficient. The method can be adapted for non-invasive vascular elastography (NIVE), for non-invasive vascular micro-elastography (MicroNlVE) on small vessels, and for endovascular elastography (EVE). | 1. A method for vascular elastography comprising: providing pre-tissue-motion and post-tissue-motion images in digital form of a vessel delimited by a vascular wall; said pre-tissue-motion and post-tissue-motion images being representative of first and second time-delayed configuration of said vessel;partitioning at least portions of both said pre-tissue-motion and post-tissue-motion images into corresponding data windows;approximating a trajectory between said pre-tissue-motion and post-tissue-motion images for corresponding data windows; andusing the trajectory for each data window to compute a strain tensor in each data window. 2. A method as recited in claim 1, further comprising using said strain tensor in each data window to create an elastogram of at least part of said vessel. 3. A method as recited in claim 1, wherein said pre-tissue-motion and post-tissue-motion images are radio-frequency (RF) images. 4. A method as recited in claim 3, wherein said pre-tissue-motion and post-tissue-motion images are part of a sequence of radio-frequency (RF) images. 5. A method as recited in claim 1, wherein said pre-tissue-motion and post-tissue-motion images are issued from magnetic resonance imaging (MRI), optical coherence tomography (OCT), brightness mode (B-mode) or Doppler-based ultrasound modality imaging. 6. A method as recited in claim 1, wherein providing pre-tissue-motion and post-tissue-motion images in digital form of a vessel includes inducing tissue compression or dilatation on said vessel. 7. A method as recited in claim 6, wherein inducing tissue dilatation on said vessel is achieved by cardiac pulsation. 8. A method as recited in claim 1, wherein said strain tensor is the full strain tensor in at least one of said data windows. 9. A method as recited in claim 8, wherein said full strain tensor is computed from three-dimensional or two-dimensional ultrasound data. 10. A method as recited in claim 8, further comprising using said full strain tensor to compute the Von Mises (VM) coefficient in each data window. 11. A method as recited in claim 1, wherein approximating a trajectory for each said data window includes using a Lagrangian speckle model estimator (LSME). 12. A method as recited in claim 1, wherein a trajectory is approximated in each said data window using zero-order and first-order terms of a Taylor-series expansion, yielding:[xyz]=[x(0,0,0,t)y(0,0,0,t)z(0,0,0,t)]Tr+[∂x∂x0∂x∂y0∂x∂z0∂y∂x0∂y∂y0∂y∂z0∂z∂x0∂z∂y0∂z∂z0](0,0,0,t)LT[x0y0z0][Tr] is a translation vector [LT] is a linear geometrical transformation of coordinates (x, y, z) represents the new position of a point (x0, Y0, z0). 13. A method as recited in claim 12, wherein using said trajectory for each said data window to compute a strain tensor in each data window includes performing a non-linear minimization for each data window Wij by computing a transformation [LT] providing the best match between each Wij of said pre-tissue motion image and a correspond... | - | A61B005/055
US20100305633 | A1 | 2010-12-02 | Diagnostic And Therapeutic Chest Casing | 2010-06-01 | A system and method is provided for detection, forewarning, and rapid therapeutic treatment of a patient's ischemic and arrhythmic heart condition, congestive heart failure, respiratory failure, etc. The system comprises a chest casing of a form fitting material. The inner surface of the chest casing comprises electrocardiogram sensor electrodes, auscultation sensors, impedance sensor electrodes, etc. The electrocardiogram sensor electrodes detect abnormal electrocardiogram signals. The auscultation sensors record internal sounds at predefined cardiac and respiratory auscultation sensor points. The impedance sensor electrodes measure thoracic impedance across two or more points on the chest wall. A control unit is connected to the sensors for collecting and processing patient information. The control unit transmits the processed patient information to an external monitoring station. The chest casing comprises therapeutic delivery points for delivering therapeutic electrical dosages and different therapies, for example, cooling therapy, to the patient via therapeutic devices connected on the chest casing. | 1. A system for detection, forewarning, and rapid therapeutic treatment of heart and cardiopulmonary conditions in a patient, comprising: a chest casing made of a form fitting material configured to conform to an upper body of said patient, wherein said chest casing defines an inner surface for establishing adequate contact with skin of said patient and an outer surface opposite to said inner surface, wherein said inner surface comprises: a predefined plurality of electrocardiogram sensor electrodes for detecting abnormal electrocardiogram signals of said patient; anda predefined plurality of auscultation sensors for recording internal sounds at predefined cardiac auscultation sensor points and respiratory auscultation sensor points located on a chest wall of said patient;a control unit attached to said outer surface of said chest casing and connected to said electrocardiogram sensor electrodes and said auscultation sensors for collecting and processing patient information comprising one or more signals from said electrocardiogram sensor electrodes and said auscultation sensors, wherein said control unit communicates with an external monitoring station for transmitting said processed patient information to said external monitoring station; andsaid chest casing comprising one or more therapeutic delivery points for delivering therapeutic electrical dosages to said patient automatically and/or based on feedback from said external monitoring station to said control unit in response to said transmitted patient information. 2. The system of claim 1, wherein said inner surface of said chest casing comprises one or more impedance sensor electrodes for measuring thoracic impedance across two or more points on said chest wall of said patient, wherein said control unit processes said measured thoracic impedance to detect signs of abnormality between said two or more points and to detect signs of congestive heart failure. 3. The system of claim 1, wherein said chest casing comprises one or more embedded coolant circuits for delivering cooling therapy to lower body temperature of said patient. 4. The system of claim 1, wherein said chest casing comprises one or more sleeves made of said form fitting material configured to conform to shoulders and arms of said patient, wherein an inner surface of said sleeves comprises one or more of a temperature sensor, a pulse oxygenation sensor, and a blood pressure sensor. 5. The system of claim 1, wherein said predefined cardiac auscultation sensor points on said chest wall of said patient for recording said internal sounds by said auscultation sensors correspond to apex of heart, aortic valve area, and left sternal border. 6. The system of claim 1, wherein said respiratory auscultation sensor points on said chest wall of said patient for recording said internal sounds by said auscultation sensors correspond to lower lung fields, middle lung fields, and upper lung fields on a back of said patient. 7. The system of... | - | A61N001/00 A61B005/0205 A61F007/00 A61B005/00 A61M016/00
US20100298670 | A1 | 2010-11-25 | ELECTROLYTE MONITORING USING IMPLANTED CARDIAC RHYTHM MANAGEMENT DEVICE | 2009-05-20 | A method for diagnosing an electrolyte level with a cardiac rhythm management device includes recording intra-cardiac electrograms from multiple sites. The method determines the electrolyte level based upon a comparative analysis of intra-cardiac electrograms recorded from at least two of the sites. The electrolyte level can be quantified based upon a general model, or a patient specific model. | 1. A method for diagnosing an electrolyte level with a cardiac rhythm management device, comprising: recording intra-cardiac electrograms from a plurality of sites; anddetermining the electrolyte level based upon a comparative analysis of the intra-cardiac electrograms from at least two of the sites. 2. The method of claim 1, in which the determining further comprises quantifying the electrolyte level using a model including at least one feature of the intra-cardiac electrograms from the at least two sites. 3. The method of claim 2, further comprising transmitting a warning when the diagnosed electrolyte level exceeds a threshold value. 4. The method of claim 2, further comprising deriving weighting factors of the model based upon data observed from a plurality of patients. 5. The method of claim 4, in which the deriving uses normalized features. 6. The method of claim 4, further comprising calibrating the model to a specific patient based upon a single measured electrolyte level for the specific patient and the corresponding intra-cardiac electrograms from the at least two sites of the specific patient. 7. The method of claim 6, in which the calibrating further comprises accounting for a plurality of electrolyte level measurements for the specific patient over time and the corresponding intra-cardiac electrograms from the at least two sites of the specific patient. 8. The method of claim 1, in which the determining the electrolyte level is further based upon a time comparative analysis of the intra-cardiac electrograms from the at least two sites. 9. The method of claim 1, in which the determining further comprises comparing repolarization features of the intra-cardiac electrograms from the at least two sites to determine a potassium level. 10. The method of claim 1, in which the determining further comprises comparing depolarization features of the intra-cardiac electrograms from the at least two sites to determine a sodium level. 11. The method of claim 1, in which the determining further comprises comparing plateau features of the intra-cardiac electrograms from the at least two sites to determine a calcium level. 12. The method of claim 1, further comprising wirelessly transmitting the determined electrolyte level to an external receiver. 13. A medical device for diagnosing an electrolyte level of a patient, comprising: a first electrode adapted to be implanted at a first cardiac site to facilitate generating a first intra-cardiac electrogram;a second electrode adapted to be implanted at a second cardiac site to facilitate generating a second intra-cardiac electrogram; anda processing system that analyzes at least one feature from the first and second intra-cardiac electrograms to quantify the electrolyte level. 14. The medical device of claim 13, in which the device comprises a cardiac rhythm management device. 15. The medical device of claim 13, in which the device comprises an external monitor. 16. A system for diagnosing a... | PACESETTER, INC. | A61B005/145
US07853059 | B2 | 2010-12-14 | Magnetic resonance imaging apparatus, method of making an imaging-plan, and method of imaging | 2006-11-29 | The present invention provides a magnetic resonance imaging apparatus including a unit for specifying a plurality of regions of interest on a plurality of original images acquired by imaging several times while shifting imaging positions, a unit for finding an approximate line involved in the regions of interest on an image formed by combining the original images, and a determining unit for determining conditions required to image cross sections crossing the original images on the basis of the approximate line. | 1. A magnetic resonance imaging apparatus comprising: a unit for specifying a plurality of regions of interest on a plurality of original images acquired by imaging several times while changing imaging positions;a unit for finding an approximate line related to the regions of interest on a composite image formed by combining the original images, said approximate line passing intermediately between said regions of interest; anda determining unit for determining conditions required to image at least one cross sectional image substantially parallel to said approximate line and encompassing plural of said regions of interest across plural of said original images on the basis of the approximate line. 2. The apparatus according to claim 1, wherein the determining unit determines a direction of a surface that is formed along the approximate line and crosses the original images as a cross-sectional direction. 3. The apparatus according to claim 1, wherein the determining unit includes:a unit for defining coverage surrounding all of the regions of interest centered about the approximate line; anda unit for determining the number of slices on the basis of the size of coverage and predetermined thickness and gap of the slices. 4. The apparatus according to claim 1, wherein the determining unit determines a central position of a scan area for one imaging position on the basis of intersecting points of a line passing the center of the original image and the approximate line. 5. The apparatus according to claim 4, wherein the determining unit determines the intersecting points as central positions of the scan areas. 6. The apparatus according to claim 4, wherein the determining unit finds two adjacent intermediate points of the intersection points for each of the original image and determines points, which are spaced at a predetermined distance along the approximate line from the central points corresponding to the half of the width of the scan areas, as central points of the scan areas. 7. The apparatus according to claim 1, wherein the determining unit includes a unit for showing the determined conditions to an operator. 8. The apparatus according to claim 7, wherein the determining unit further includes a unit for changing the determined conditions according to request of the operator. 9. A magnetic resonance imaging apparatus comprising: an imaging unit for generating magnetic resonance signals by applying gradient magnetic fields and high frequency pulses to an object in a static magnetic field;a bed for longitudinally sliding a top board with an object placed;a high frequency coil for detecting the magnetic resonance signals;a display unit for displaying positioning images;a specifying unit for specifying a plurality of regions of interest according to specifying order of an operator on the positioning images;a calculating unit for calculating coverage surrounding a plurality of regions of interest depending on the regions of interest specified b... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | G06K009/00 A61B005/05
US07833165 | B2 | 2010-11-16 | System for monitoring neural shunt function and associated methods | 2006-08-15 | A system and method for evaluating neural shunt functionality is provided. Accordingly, in one aspect a method for monitoring neural shunt functionality may include measuring a first intracranial pressure pulse inside the shunt, measuring a second intracranial pressure pulse outside the shunt, and comparing pulsatile characteristics from the first measurement to the second measurement in order to determine shunt functionality. | 1. A method for monitoring neural shunt functionality, comprising: measuring a first intracranial pressure pulse inside the shunt using a first sensor;measuring a second intracranial pressure pulse outside the shunt using a second sensor; andcomparing pulsatile characteristics from the first measurement to the second measurement using a computational device in order to determine shunt functionality. 2. The method of claim 1, wherein measuring inside the shunt further includes inserting the first sensor into the shunt to measure the first intracranial pressure pulse. 3. The method of claim 2, wherein inserting the first sensor into the shunt further includes inserting the first sensor into a structure that is fluidically coupled to the shunt. 4. The method of claim 1, wherein measuring outside the shunt further includes measuring the second intracranial pressure pulse from a location outside the cranium. 5. The method of claim 4, wherein measuring from a location outside the cranium includes measuring the second intracranial pressure pulse from tympanic membrane displacements. 6. The method of claim 4, wherein measuring from a location outside the cranium includes measuring the second intracranial pressure pulse from a supraorbital artery. 7. The method of claim 1, wherein the pulsatile characteristics include pulse amplitude. 8. The method of claim 1, wherein the pulsatile characteristics include pulse risetime or pulse falltime. 9. The method of claim 1, wherein comparing pulsatile characteristics further includes determining a correlation measurement between the first intracranial pressure pulse and the second intracranial pressure pulse. 10. The method of claim 1, further comprising clearing an occluded portion of the neural shunt. | - | A61B005/00 A61B005/02
US20100317966 | A1 | 2010-12-16 | Hybrid-multimodal magneto-optical contrast markers | 2010-06-04 | A contrast marker for use in imaging applications, wherein the marker is sensitive to an applied magnetic field through the formation of radical pair species and wherein the marker is independently magnetically sensitive, optically sensitive, and magneto-optically sensitive. | 1. A contrast marker for use in imaging applications, wherein said marker is sensitive to an applied magnetic field through the formation of radical pair species and wherein said marker is independently magnetically sensitive, optically sensitive, and magneto-optically sensitive. 2. A contrast marker according to claim 1, wherein said marker is a donor-acceptor complex. 3. A contrast agent according to claim 1, wherein said radical pair species include reactive oxygen species, nitroxyl radicals, radical pairs associated with tryptophan and FADH, radical pairs generated from donor-acceptor porphyrin complexes, and radical pairs generated form enzymatic processes. 4. A contrast marker according to claim 1, wherein said contrast marker is sensitive to said magnetic field through the formation of radical pair species that are paramagnetic in the triplet state. 5. A contrast marker according to claim 4, wherein said radical pair species are produced through a charge transfer process. 6. A contrast marker according to claim 2, wherein said marker contains at least two chemical entities covalently linked to form one compound. 7. A contrast marker according to claim 2, wherein the donor or acceptor part of the compound is endogenous to the organism. 8. A contrast marker according to claim 1, wherein said marker is part of a probe, said probe being further comprised of at least one chemical component in order to increase a multimodality thereof. 9. A contrast marker according to claim 8, wherein said multimodality includes electron-spin resonance, magnetic resonance imaging, optical and near-infrared imaging or spectroscopy, x-ray imaging, ultrasound imaging and nuclear imaging techniques. 10. A probe for use in imaging applications, said probe including a contrast marker according to claim 1, said probe being adapted to target a detection and to maximize a change in luminescence response to a magnetic field. 11. A probe for use in imaging applications, said probe including a contrast marker according to claim 1, said probe being adapted to target a detection site and to maximize a change in magnetic response to an optical field. 12. A contrast marker according to claim 1, wherein a luminescence signal produced by said marker under optical, magnetic, or both activation is measured through luminescence intensity, lifetime, spectral properties, or a combination thereof. 13. A contrast marker according to claim 1, wherein a magnetic signal produced by said marker under optical, magnetic, or both activation is measured through a resonance spectral response induced through a radiofrequency pulse. | Institut National D'Optique | A61B005/05 C07D209/20 C07D487/22 C01B013/00 C01B021/20
US20100280418 | A1 | 2010-11-04 | METHOD AND SYSTEM FOR EVALUATING A MOVEMENT OF A PATIENT | 2010-03-23 | A system for evaluating the movement of a patient having a first display unit for displaying a reference movement, a first detection unit, to be situated in particular in the location of the patient, for detecting the movement of the patient, an evaluation unit for ascertaining a measure characterizing a deviation between the detected movement and the reference movement, which is finally displayed on the first or a second display unit. It is possible to achieve an evaluation of the movement of the patient that is independent of the presence of a therapist. | 1. A system for evaluating a movement of a patient, comprising: at least one display unit for displaying a reference movement;a first detection unit situated in a location of the patient for detecting a movement of the patient; andan evaluation unit for ascertaining a measure characterizing a deviation between the detected movement and the reference movement,wherein the at least one display unit displays the ascertained measure. 2. The system according to claim 1, further comprising a memory unit for storing at least one of the detected movement of the patient and the ascertained measure. 3. The system according to claim 1, further comprising a second detection unit situated in a location of a therapist for detecting the reference movement. 4. The system according to claim 1, further comprising a further display unit situated in a location of a therapist for displaying at least one of the detected movement and the ascertained measure. 5. The system according to claim 3, wherein at least one of the first and second detection units includes a video sensor system unit. 6. The system according to claim 3, wherein at least one of the first and second detection units includes an inertial sensor system unit. 7. The system according to claim 1, further comprising: a first station situated in the location of the patient; anda second station situated in a location of a therapist. 8. A method for evaluating a movement of a patient, comprising: displaying a reference movement;detecting the movement of the patient;ascertaining a measure characterizing a deviation between the detected movement and the reference movement; anddisplaying the ascertained measure. 9. The method according to claim 8, further comprising detecting the reference movement. 10. The method according to claim 8, further comprising storing at least one of the movement of the patient and the ascertained measure. | - | A61B005/103
US20100331644 | A1 | 2010-12-30 | HOUSING FOR AN INTRAVASCULAR SENSOR | 2010-07-01 | An apparatus houses an intravascular sensor and is configured to measure the analyte in a biological sample of a host. The apparatus includes a fluid coupler having a first end configured to mate with a connecting end of a catheter and a second end configured to mate with a tubing assembly including, for example, an infusion pump, and a housing connected to the fluid coupler. The housing is configured to receive a sensor disposed within the fluid coupler such that when the fluid coupler is mated to the catheter, the sensor can be exposed to a biological sample. The housing is also configured to electrically couple the sensor with an external device, such as a processor for receiving and analyzing the sensor output. The housing and the fluid coupler are connected such that a fluidic seal is formed thereby preventing fluid in the fluid coupler from entering the housing. | 1. An apparatus for use with an intravascular sensor, the apparatus comprising: a fluid coupler comprising a first end and a second end, wherein the first end is configured to releasably couple with a connecting end of an intravascular catheter, and wherein the second end is configured to releasably couple with a tubing assembly, said fluid coupler adapted to at least partially house an intravascular sensor; anda housing connected to the fluid coupler, the housing configured to receive a portion of at least one sensor disposed at least partially within the fluid coupler such that when the fluid coupler is mated to an intravascular catheter inserted into a circulatory system of the host, the at least one sensor can be exposed to a biological sample, the housing further configured to electrically couple the at least one sensor with at least one external device. 2. The apparatus of claim 1 wherein a seal is formed around a peripheral of the sensor at or proximate the intersection of the fluid coupler and the housing. 3. The apparatus of claim 1 wherein the connection between the fluid coupler and the housing forms a seal configured to prevent flow of fluid from the fluid coupler into the housing. 4. The apparatus of claim 1 wherein the fluid coupler further comprises a seal proximate the connection between the fluid coupler and the housing, the seal configured to prevent flow of fluid from the fluid coupler into the housing. 5. The apparatus of claim 1 wherein the housing further comprises a seal proximate the connection between the housing and the fluid coupler, the seal configured to prevent flow of fluid from the fluid coupler into the housing. 6. The apparatus of claim 1 wherein the at least one sensor comprises at least one conductive wire with an electrode fainted on a distal end and wherein the housing is further configured to receive the at least one conductive wire of the sensor and electrically couple the at least one conductive wire with the at least one external device. 7. The apparatus of claim 6 wherein the housing further comprises at least one connector configured to electrically couple the at least one conductive wire with the at least one external device. 8. The apparatus of claim 7 wherein the at least one connector comprises at least one elastomeric contact. 9. The apparatus of claim 7 wherein the housing further comprises a printed circuit board configured to electrically couple with the at least one connector and the at least one external device thereby electrically coupling the at least one conductive wire of the sensor with the at least one external device. 10. The apparatus of claim 1 wherein the apparatus is configured to measure at least one analyte in the biological sample of the host and wherein the sensor comprises an analyte sensor. 11. The apparatus of claim 1 wherein the tubing assembly comprises an infusion system configured to supply a fluid to the fluid coupler, the fluid coupler configured to transfer t... | DexCom, Inc. | A61B005/1468
US20100292601 | A1 | 2010-11-18 | APPARATUS FOR AND METHOD OF CONDENSING EXHALED BREATH | 2008-02-15 | A breath condensing apparatus includes a mouthpiece, a condenser having an inlet coupled to the mouthpiece, a first container for collecting condensate communicating with an outlet of the condenser, and a second container for collecting residual, non-condensed exhaled breath coupled to the condenser. A method of condensing exhaled breath includes the features of receiving exhaled vapor into a condenser, condensing at least part of the exhaled vapor, collecting at least part of residual, non-condensed exhaled vapor, recirculating at least part of the collected residual vapor through the condenser, condensing at least part of the recirculated vapor, and collecting condensate. | 1-15. (canceled) 16. A breath condensing apparatus, the apparatus comprising: a mouthpiece,a condenser having an inlet in fluid communication with the mouthpiece,a first container for collecting condensate in fluid communication with an outlet of the condenser, anda second container for collecting residual, non-condensed exhaled breath in fluid communication with the condenser. 17. The apparatus according to claim 16, wherein the second container is a collapsible bag made of a flexible, thin-walled material. 18. The apparatus according to claim 17, wherein the flexible, thin-walled material comprises a film of an inert synthetic material. 19. The apparatus according to claim 16, wherein the second container is connected to both the inlet and the outlet of the condenser thus enabling recirculation of breath collected in the second container through the condenser. 20. The apparatus according to claim 16, wherein the second container has a volume of between approximately 10 liters and approximately 100 liters. 21. The apparatus according to claim 16, further comprising one or more heating elements for maintaining residual, non-condensed exhaled breath gaseous and at a temperature of between approximately 30° C. and approximately 42° C. 22. The apparatus according to claim 21, further comprising a casing, wherein the second container and the heating elements are disposed inside the casing. 23. The apparatus according to claim 16, further comprising a wiper positionable inside the condenser to wipe condensate from the inner wall of the condenser and towards the first container for collecting condensate. 24. The apparatus according to claim 16, further comprising one or more channels located at or below the inlet and having both a tangential and an axial component, relative to the central axis of the condenser, to direct exhaled breath towards the inner wall of the tube. 25. The apparatus according to claim 16, wherein at least one of the mouthpiece or a duct connecting the mouthpiece to the condenser comprises one of: a two-way non-rebreathing valve system,a flow meter, anda saliva-trap. 26. The apparatus according to claim 16, further comprising a sensor for measuring the relative humidity of the contents of the second container. 27. The apparatus according to claim 16, further comprising a ventilator adapted to force recirculation of non-condensed exhaled breath. 28. A method of condensing exhaled breath, the method comprising: receiving exhaled breath from a subject into a condenser,condensing at least part of the exhaled breath,collecting at least part of residual, non-condensed exhaled breath,recirculating at least part of the collected residual breath through the condenser,condensing at least part of the recirculated breath, andcollecting condensate. 29. The method according to claim 28, further comprising collecting the residual, non-condensed air in a collapsible bag. 30. The method according to claim 29, wherein the bag is coll... | UNIVERSITEIT MAASTRICHT | A61B005/097
US20100312136 | A1 | 2010-12-09 | SYSTEMS AND METHODS FOR THERMOGRAPHIC-GUIDED MEDICAL TREATMENT | 2009-12-28 | A system for detecting the risk of breast cancer comprises a thermographic imaging device configured to acquire thermographic images of a female torso; an analysis authority communicatively coupled with the thermographic imaging device, the analysis authority including an input port configured to receive thermographic images from the thermographic imaging device, a database configured to store the thermographic images, a qualitative analysis module configured to automatically perform a qualitative analysis of the thermographic images and generate a qualitative score based on the qualitative analysis, a quantitative analysis module configured to automatically perform a quantitative analysis of the thermographic images and generate a quantitative score based on the quantitative analysis, and a scoring module configured to correlate the qualitative and quantitative scores with a score indicating a risk of breast cancer. | 1. A system for detecting the risk of breast cancer, comprising; a thermographic imaging device configured to acquire thermographic images of a female torso;an analysis authority communicatively coupled with the thermographic imaging device, the analysis authority including: an input port configured to receive thermographic images from the thermographic imaging device,a database configured to store the thermographic images,a qualitative analysis module configured to automatically perform a qualitative analysis of the thermographic images and generate a qualitative score based on the qualitative analysis,a quantitative analysis module configured to automatically perform a quantitative analysis of the thermographic images and generate a quantitative score based on the quantitative analysis, anda scoring module configured to correlate the qualitative and quantitative scores with a score indicating a risk of breast cancer. 2. The system of claim 1, wherein the thermographic imaging device is configured to acquire a first series of images and a second series of images, and wherein the qualitative analysis is performed based on comparisons between the two image series. 3. The system of claim 1, wherein the thermographic imaging device is configured to acquire a first series of images and a second series of images, and further comprising a comparison module configured to compare the two image series and to identify the areas of interest based on the comparison, and wherein the qualitative and quantitative analysis modules are configured to perform the qualitative and quantitative analysis on the areas of interest. 4. The system of claim 2, wherein the second series of images are acquired after an autonomic challenge. 5. The system of claim 3, wherein the comparison module is further configured to determine whether a temperature calibration issue exists based on a comparison between the two series of images. 6. The system of claim 1, wherein the thermographic images include a anterior image of a breast region of interest, a left oblique image of the breast area of interest; and a right oblique image of the breast area of interest. 7. The system of claim 1, wherein the analysis authority is further configured to receive and store photographs of the female torso, and to use the photographs to identify areas of interest in the thermographic images. 8. The system of claim 1, further comprising a report module configured to automatically generate a report based on the out put of the qualitative analysis module, quantitative analysis module, and scoring module. 9. The system of claim 1, further comprising a cancer identification module configured to identify risk of a certain type of cancer based on the output of the qualitative analysis module, quantitative analysis module, and scoring module. 10. The system of claim 1, further comprising a treatment identification module configured to correlate the output of the scoring module with a specific treatm... | ADVANCED FUNCTIONAL SYSTEMS, LLC | A61B005/01
US20100312139 | A1 | 2010-12-09 | Method for Identifying Mammals at Risk for Elevated intracranial Pressure | 2010-06-04 | Disclosed are methods of diagnosing, screening and treating individuals at risk of elevated intracranial pressure based on detecting a transient decrease in serum ceruloplasmin and/or serum copper levels in the individual soon after occurrence of a head injury. This early drop in ceruloplasmin or copper level indicates an impending, potentially life threatening, elevation in the individual's intracranial pressure. Also disclosed are diagnostic applications and processes for identifying and producing diagnostic and prognostic assays useful for predicting elevated intracranial pressure, and test kits for performing such tests. | 1. A process for identifying a patient at risk for elevated intracranial pressure, said patient having experienced a precipitating event that includes brain injury or potential brain injury, the process comprising: measuring a ceruloplasmin level, copper level, or both ceruloplasmin and copper levels in a test sample of body fluid obtained from the patient, wherein said test sample is obtained within 36 hours after the precipitating event; anddetermining from said measurement that the serum ceruloplasmin level, copper level, or both levels, are below the respective normal ranges, wherein a subnormal level of either serum ceruloplasmin or serum copper, or both, indicates increased risk for elevated intracranial pressure. 2. The process of claim 1, wherein the body fluid comprises serum. 3. The process of claim 1, wherein the body fluid is selected from the group consisting of serum, plasma, blood, urine, cerebral spinal fluid, urine and saliva, and said determining comprises correlating said measured levels to a serum ceruloplasmin level, a serum copper level, or to both levels. 4. (canceled) 5. The process of claim 1, wherein said measuring comprises measuring ceruloplasmin level. 6. The process of claim 5, wherein said measuring comprises performing an immunoassay. 7.-8. (canceled) 9. The process of claim 1, wherein said measuring comprises measuring a copper level. 10.-12. (canceled) 13. The process of claim 1 further comprising: identifying said patient as a candidate for more invasive monitoring techniques and closer observation, based on said determination that the serum ceruloplasmin level, the serum copper level or both levels are below the respective normal ranges; andusing an invasive intracranial pressure monitoring technique on the patient. 14. (canceled) 15. The process of claim 1 further comprising: after said precipitating event and prior to said measuring and determining, establishing a first diagnosis of the patient as being theoretically at risk for elevated intracranial pressure; andafter said measuring and determining, establishing a second diagnosis of the patient as being at increased risk for elevated intracranial pressure, based on said determination that the serum ceruloplasmin level, serum copper level or both levels are below the respective normal ranges. 16.-17. (canceled) 18. The method of claim 1, wherein said patient is included in a population of individuals who have experienced a precipitating event that may include or cause brain injury, the method comprising: measuring a ceruloplasmin level, copper level, or both ceruloplasmin and copper levels in said test sample obtained from each said individual, wherein each said test sample is obtained within 36 hours after the respective precipitating event;determining from each said measurement whether the serum ceruloplasmin level, serum copper level or both levels are below the respective normal ranges, wherein a subnormal level of either serum ceruloplasmin... | BOARD OF REGENTS OF THE UNIVERSITY OF TEXAS SYSTEM | A61B005/00 G01N033/20 G01N033/68 C12Q001/00 G01N033/543 G01N033/53
US20100292608 | A1 | 2010-11-18 | ANALYSIS INSTRUMENT | 2008-04-29 | An analysis instrument (2) has a hole (23) for allowing a laser beam, which is applied to a skin to withdraw bodily fluids, to advance through it and also has a cover (25) for closing an opening (23B), into which the laser beam enters, of the hole (23) and allowing the laser beam to pass through it. The analysis instrument (2) has electrodes (20, 21) stacked one upon another while being electrically connected and having through-holes (20A, 21A) defining the hole (23). Preferably, the analysis instrument (2) further has a gas discharge flow path (26) for discharging gas in the hole (23) to the outside. | 1. An analysis tool comprising a hole through which a laser beam to be emitted to skin to withdraw bodily fluid travels, and a cover which closes an opening in the hole from which the laser beam enters and through which the laser beam can pass. 2. The analysis tool according to claim 1, further comprising a plurality of electrodes which are laminated together in a state such that they are electrically insulated from each other, and which have through holes that define the hole. 3. The analysis tool according to claim 2, further comprising a discharge passage through which gas in the hole is discharged to the outside. 4. The analysis tool according to claim 3, wherein the discharge passage is provided by forming, in an insulation layer interposed between the electrode and the cover, a slit which communicates with the hole. 5. The analysis tool according to claim 1, wherein the hole is for applying capillary action to suck bodily fluid from the skin. 6. The analysis tool according to claim 5, further comprising a reagent section formed on an inner surface of the hole. 7. The analysis tool according to claim 5, further comprising a passage through which bodily fluid sucked in the hole moves, and a reagent section formed in the passage. 8. The analysis tool according to claim 7, further comprising a substrate which supports the cover and which is provided with a plurality of electrodes. 9. The analysis tool according to claim 8, further comprising a spacer which is interposed between the substrate and the cover, and which defines the passage. | ARKRAY, Inc. | A61B005/151
US20100331709 | A1 | 2010-12-30 | DETECTION UNIT FOR BLOOD PRESSURE INFORMATION MEASUREMENT DEVICE AND BLOOD PRESSURE INFORMATION MEASUREMENT DEVICE | 2009-01-19 | A blood pressure information measurement device includes a cuff (10A) serving as a detection unit. The cuff (10A) includes an air bladder (40), a photoelectric sensor (50), a belt member (20), and a fixing stand (32), and is attached to a living body by wrapping the belt member (20) around a measuring site. The photoelectric sensor (50) optically detects an intra-arterial volume fluctuation, and includes a light emitting element (51) and a light receiving element (52). The fixing stand (32) includes a base portion (32a) with a sensor attachment surface (32a1) attached with the photoelectric sensor (50), and a guide portion (32b) arranged projecting out from the base portion (32a) toward the sensor attachment surface (32a1) side, and having a distal end placed at a body surface near the measuring site in the attachment state of the cuff (10A). The air bladder (40) is arranged on the sensor attachment surface (32a1) so as to cover the photoelectric sensor (50), and the guide portion (32b) is arranged to surround the photoelectric sensor (50). With such a configuration, the blood pressure information measurement device capable of acquiring a volume pulse wave easily and at high accuracy can be obtained. | 1. A detection unit for blood pressure information measurement device comprising: a compression fluid bag for compressing an artery in a measuring site by compressing the measuring site;a photoelectric sensor, including a light emitting portion and a light receiving portion for radiating detection light toward the measuring site from the light emitting portion, receiving the detection light transmitted through the measuring site with the light receiving portion, and outputting an output signal corresponding to a light quantity of the received detection light; anda fixing unit for fixing the photoelectric sensor with respect to the measuring site; whereinthe fixing unit includes a base portion with a sensor attachment surface attached with the photoelectric sensor, and a guide portion arranged projecting out from the base portion toward the sensor attachment surface side, the guide portion having a distal end directly or indirectly placed to a body surface near the measuring site with the photoelectric sensor fixed with respect to the measuring site by the fixing unit;the compression fluid bag is arranged on the sensor attachment surface so as to cover the photoelectric sensor; andthe guide portion is positioned to surround the photoelectric sensor when the fixing unit is seen from a normal direction of the sensor attachment surface. 2. The detection unit for blood pressure information measurement device according to claim 1, wherein the guide portion has a wall-shape or a columnar shape. 3. The detection unit for blood pressure information measurement device according to claim 1, wherein the fixing unit includes a belt member attached by being wrapped around a living body including the measuring site. 4. The detection unit for blood pressure information measurement device according to claim 3, wherein the light emitting portion and the light receiving portion are arranged in line in a longitudinal direction of the belt member. 5. A blood pressure information measurement device comprising: the detection unit for blood pressure information device according to claim 1;a drive unit for causing the light emitting portion to emit light;a light receiving quantity detector for detecting a fluctuation of light receiving quantity based on an output signal outputted from the photoelectric sensor; anda volume pulse wave acquiring unit for acquiring a volume pulse wave of an artery based on information obtained by the light receiving quantity detector. 6. The blood pressure information measurement device according to claim 5, wherein the drive unit causes the light emitting portion to intermittently emit pulsed light. 7. The blood pressure information measurement device according to claim 5, further comprising a pressure adjustment mechanism for expanding and contracting the compression fluid bag by adjusting an internal pressure of the compression fluid bag. 8. The blood pressure information measurement device according to claim 5, further comprising ... | OMRON HEALTHCARE CO., LTD. | A61B005/022
US20100312131 | A1 | 2010-12-09 | NOISE DETECTION AND RESPONSE FOR USE WHEN MONITORING FOR ARRHYTHMIAS | 2009-06-09 | Methods and systems of noise detection and response for use when monitoring for arrhythmias are described herein. At least two electrodes are used to obtain a signal indicative of cardiac electrical activity. The signal is bandpass filtered to obtain a filtered signal. Ventricular depolarizations are monitored for based on comparisons of the filtered signal to a first threshold. Arrhythmias are monitored for based on ventricular depolarization detections that occur as a result of monitoring for ventricular depolarizations. During one or more noise detection windows, noise is monitored for and a likelihood that monitoring for arrhythmias is adversely affected by noise is determined based on results thereof. Whether and/or how the monitoring for arrhythmias is performed is modified when it is determined that monitoring for arrhythmias is likely adversely affected by noise. | 1. A method of noise detection and response for use when monitoring for arrhythmias, comprising: (a) using at least two electrodes to obtain a signal indicative of cardiac electrical activity;(b) bandpass filtering the signal to obtain a filtered signal;(c) monitoring for ventricular depolarizations based on comparisons of the filtered signal to a first threshold;(d) monitoring for arrhythmias based on ventricular depolarization detections that occur as a result of the monitoring at step (c);(e) monitoring for noise during one or more noise detection windows;(f) determining, based on results of the monitoring for noise, when it is likely that the monitoring for arrhythmias is adversely affected by noise; and(g) modifying whether and/or how the monitoring for arrhythmias is performed, when it is determined that the monitoring for arrhythmias is likely adversely affected by noise. 2. The method of claim 1, wherein steps (d) is performed before step (e), step (e) is performed before step (d), or step (d) and step (e) are performed simultaneously. 3. The method of claim 1, wherein the signal is one of an electrocardiogram (ECG) signal and an intracardiac electrogram (IEGM) signal. 4. The method of claim 3, wherein step (a) comprises: using at least two electrodes implanted subcutaneously to obtain a subcutaneous ECG signal;using at least two non-implanted surface electrodes to obtain a surface ECG signal; orusing at least two electrodes implanted within or on the heart to obtain an IEGM signal. 5. The method of claim 1, wherein step (b) comprises bandpass filtering the signal to achieve an effective frequency range of about 8.5 Hz to about 30 Hz. 6. The method of claim 1, wherein step (c) comprises determining a ventricular depolarization detection when the filtered signal crosses the first threshold. 7. The method of claim 1, wherein step (d) comprises monitoring for at least ventricular tachyarrhythmia (VT) and atrial fibrillation (AF). 8. The method of claim 1, wherein step (e) comprises monitoring for noise based on comparisons of portions of the filtered signal, corresponding to the noise detection windows, to a second threshold having a lower magnitude than the first threshold. 9. The method of claim 8, wherein step (e) includes determining that noise is detected within a noise detection window based on how many times the portion of the filtered signal, corresponding to the noise detection window, crosses the second threshold. 10. The method of claim 8, wherein the second threshold is a specified percentage of the first threshold, with the specified percentage having a lower limit of about 25% and an upper limit of about 85%. 11. The method of claim 8, wherein: step (e) comprises monitoring for noise during a plurality of back-to-back noise detection windows; andeach said noise detection window comprises a length of time having a lower limit of about 90 ms and an upper limit of about 250 ms. 12. The method of claim 11, wherein each s... | - | A61B005/046
US20100324398 | A1 | 2010-12-23 | NON-INVASIVE CHARACTERIZATION OF A PHYSIOLOGICAL PARAMETER | 2008-05-12 | The present invention provides a method and device for characterizing a physiological parameter. The method, in one application, uses one or more non-invasive sensors to collect patient data, and may also collect data on environmental conditions. At least some of the patient data has a direct relationship with the physiological parameter, that is, a change in the physiological parameter is reflected in the data set, although the magnitude of the physiological parameter may masked by noise, interference, or other environmental or patient influences. The direct patient data preferably has a generally linear relationship with the physiological parameter, and if not, the patient data is linearized according to an algorithm, table, or other adjustment process. These linearizing processes may be predefined, and may adaptively learn or adjust. A blind signal source process is applied to the linearized data to generate separated signals, and the signal associated with the physiological parameter is identified. The identified signal is scaled or further processed, and the characterization result is presented. Although the method and device are described for use with a human, they may be advantageously used on animals. | 1-76. (canceled) 77. A method selected from the group consisting of: a. estimating a concentration level of a blood analyte, optionally glucose, comprising: (i) non-invasively measuring a plurality of variables in a patient to obtain a set of input data, optionally by emitting at least one pair of wavelengths (optionally within the range of about 600 to about 1 millimeter) in the from an energy source towards a first selected area of the patient and detecting energy emerging from a second selected area of the patient wherein at least one first variable of the plurality of variables depends on the patient's blood analyte concentration level and optionally comprises a variable selected from electrical impedance variable (optionally an impedance spectroscopy variable), a capacitance variable (optionally a skin capacitance variable), and a current variable, wherein at least one second variable of the plurality of variables does not depends on the patient's blood analyte concentration level, and wherein optionally the at least one first variable depends on the at least one second variable;(ii) nonlinearly filtering at least part of the set of input data to obtain a set of filtered data, wherein optionally the nonlinearly filtering comprises an at least partially adaptive component; and(iii) applying a source separation method to the set of filtered data to obtain a set of output data, wherein optionally the source separation method is at least partially adaptive;b. estimating a blood-analyte concentration level in a patient, wherein the blood-analyte optionally is glucose comprising: (i) receiving a first set of input variables, wherein the first set of input variables do not comprise any invasively-measured variables, wherein at least one first variable of the first set of input variables is influenced by the patient's blood analyte concentration level, and wherein at least one second variable of the first set of input variables is not influenced by the patient's blood analyte concentration level;(ii) pre-processing at least one of the first set of input variables to produce a second set of variables, wherein the pre-processing optionally is at least partially adaptive and optionally comprises nonlinearly transforming at least one of the first set of input variables; and(iii) applying a linear separation method to the second set of variables produce a third set of variables, wherein the linear separation method optionally is at least partially adaptive and optionally comprises a blind source separation method (optionally at least one of an Independent Component Analysis (ICA) and an Independent Vector Analysis (IVA) method); andc. characterizing a target physiological parameter, comprising: (i) collecting a first data set of data from a patient, the first data set having a direct relationship with the target physiological parameter, wherein collecting the first data set optionally further comprises using an optical, electrical, RF, infrared sensor, ... | - | A61B005/145
US07831301 | B2 | 2010-11-09 | Heart failure monitor quicklook summary for patient management systems | 2003-07-03 | Continuous remote monitoring of patients based on data obtained from an implantable hemodynamic monitor provides an interactive patient management system. Using network systems, patients are remotely monitored to continuously diagnose and treat heart-failure conditions. A screen displayable summary provides continuous feedback and information to physicians, patients and authorized third parties. The quick look summary includes various sites and presentation tailored to match the patients' and physicians' needs. The quick look summary further includes intelligent features that understand and retain the user's interests, preferences and use patterns. Patients, physicians and other caregivers are seamlessly connected to monitor and serve the chronic needs of heart-failure patients in a reliable and economic manner. | 1. A system for monitoring hemodynamic status of a patient, the system comprising: an implantable hemodynamic monitor (IHM) recording physiologic and cardiac data in a patient relating to hemodynamic conditions of the patient;an interface to interrogate the IHM and retrieve the patient data;means for transferring the retrieved data to a remote data center location having a web server;a processor to detect an abnormal event from the patient data relating to hemodynamic conditions of the patient; anda portal interface including a browser at a remote clinician station providing access to the retrieved data over a network via the web server at the remote data center location,the portal interface providing access to a viewable summary display page displaying whether an abnormal event was detected. 2. The system of claim 1 wherein said hemodynamic patient data comprises blood pressure values. 3. The system of claim 1 wherein said network comprises a worldwide web, Internet connection. 4. The system of claim 1 wherein said remote location includes a database server accessed through the web server. 5. The system of claim 1 wherein said data includes measurements relating to pulmonary pressure and heart rate. 6. The system of claim 1 wherein said web server is accessible to patients and physicians. 7. The system of claim 1 said web server enables access to various zones including a dedicated public zone, confidential zone and a private zone. 8. The system of claim 1 wherein the portal interface providing access to a viewable summary display page comprising a heart failure status via the web server. 9. The system of claim 8 wherein the summary display comprising a display of trends in the retrieved data. 10. The system of claim 9 wherein the displayed trends comprise a cardiac pressure trend. 11. The system of claim 1 wherein the portal interface providing access to a plurality of pages segregated between public and private pages and comprising: a physician section operating on both the public and the private pages,a public section operating only on the public pages, anda patient section operating on both the public and the private pages. | Medtronic, Inc. | A61B005/04
US07836890 | B2 | 2010-11-23 | Systems for immobilization of a patient for diagnosis and treatment of breast tissue | 2004-10-18 | A breast-board system, components and methods of use therefore. A generally U-shaped wrist support with a flattened bottom and rounded post like tops that may be used as either a cupping-support or as a grasping-post is mounted on a post atop a tiltable breast-board. A second support having a generally U-shaped cross section mounted on a supporting framework extends from a base, which is rotatably mounted to the tiltable breast-board near the edge thereof. The base includes a portion extending over the edge of the breast-board, and, a slidable locking mechanism on the bottom surface, which may be retracted to slide the mechanism out of a recess located on the breast-board. The support may then be rotated to a desired position and the slidable locking mechanism actuated to engage with a desired recess on the breast-board. | 1. A breast-board system sized and adapted to attach to a radiation therapy table comprising: a unitary base board;a planar tilt board having an upper end and a lower end, the lower end being attached to the unitary base board, the planar tilt board being configured to fully support a thoracic portion of a body of a patient, the planar tilt board being rotatable relative to the unitary base board between a first position in which the planar tilt board is oriented generally parallel to the unitary base board and at least one additional position in which the planar tilt board is oriented at an angle to the unitary base board;at least one movable support member extending between the unitary base board and the planar tilt board, the at least one movable support member supporting the planar tilt board relative to the unitary base board when the planar tilt board is in the at least one additional position, and the at least one movable support member having a portion movable in a longitudinal direction relative to the unitary base board;at least one support recess in the unitary base board for receiving the portion of the at least one movable support member therein when the planar tilt board is in the at least one additional position, the portion of the at least one movable support member at least partially in contact with the unitary base board; andat least a first support locking member located on the unitary base board adjacent to the at least one support recess and movable between a first position and a second position, the at least a first support locking member at least partially covering the at least one support recess and preventing removal of the portion of the at least one movable support member from the at least one support recess when the at least a first support locking member is in the first position, the portion of the at least one movable support member being removable from the at least one support recess when the at least a first support locking member is in the second position. 2. The breast-board system of claim 1, further comprising an arm support comprising: an insertion member configured for insertion into an insertion opening in the planar tilt board, the insertion member being rotatable between two or more rotational positions relative to the planar tilt board when the insertion member is inserted into the insertion opening in the planar tilt board; anda locking device configured to selectively lock the arm support in the two or more rotational positions. 3. The breast-board system of claim 2, wherein the arm support further comprises a handle disposed on a top surface of a base and in communication with the locking device, such that actuation of the handle will cause the locking device to withdraw from a retaining recess located in an edge of the planar tilt board, when the base is mounted thereon. 4. The breast-board system of claim 3, wherein the arm support further comprises a supporting framework extending from the base t... | Diacor, Inc. | A61G015/00 A61F005/37 A61F013/00 A47B071/00 A47B013/00 A47B007/00 A47C020/00 A47F005/12 A47B003/00 A61B006/04 A61B005/05
US07837629 | B2 | 2010-11-23 | System and method for generating baseline data for automated management of edema | 2006-12-06 | A system and method for generating baseline data for automated management of edema is disclosed. Collected device measures are accumulated to record raw physiometry for a patient, wherein the patient is regularly monitored by an implantable medical device, beginning with an initial observation period. Derived device measures are generated to provide derivative physiometry determined at least in part from the collected device measures. A patient status indicator is determined by analyzing the collected and derived device measures to diagnose a pathophysiology indicative of an absence, onset, progression, regression, and status quo of edema, wherein the collected and derived device measures and the patient status indicator originating from the initial observation period include baseline data. | 1. A programmer for generating baseline data for automated management of edema, comprising: a memory to store device measures received from an implantable medical device, comprising: collected device measures to record raw physiometry for a patient regularly monitored by the implantable medical device, beginning with an initial observation period; andderived device measures to provide derivative physiometry determined at least in part from the collected device measures; anda processor to determine a wellness indicator through analysis of the collected and derived device measures to diagnose a patient status indicative of at least one of a change and a status quo of edema,wherein the collected and derived device measures and the wellness indicator originating from the initial observation period comprise baseline data. 2. A programmer according to claim 1, wherein the analysis comprises one or more of: a monitor to accumulate collected device measures and to derive derived device measures for use in determining the wellness indicator for edema;a patient well being analyzer to monitor patient well being in conjunction with conditions selected from the group comprising heart failure, myocardial ischemia, and high blood pressure; anda programming interface to apply programming data to effect therapy modification selected from the group comprising device parameters, drug infusion parameters, and neural stimulation parameters. 3. A programmer according to claim 1, further comprising: an evaluator to process physiometric data conjunction with monitoring of the patient, wherein the physiometric data is selected from the group comprising: analytic data recorded, collected, and stored for edema evaluation;alert data generated in response to significant events, wherein the alert data is selected from the group comprising absolute measures, long term average measures, and short term average measures; andaction data executed in response to indications selected from the group comprising the significant events, ECG signals, activity level, temperature, pressure, and oxygen saturation. 4. A programmer according to claim 1, further comprising: physiometric data comprising one or more of the collected and derived device measures sensed through one or more of transthoracic impedance, respiratory rate, diurnal cycles, posture, resting, heartbeat cycle, breathing, movement, ECG data, activity level, body position, temperature, heart rate, heart rate variability, heart events, respiration rate, blood oxygen levels, patient symptoms, and pulmonary or transthoracic edema and local edema-related impedance values. 5. A programmer according to claim 1, further comprising: programming parameters, comprising at least one of: monitoring parameters to monitor the patient while asleep; andexclusion parameters to exclude physiometric data comprising one or more of the collected and derived device measures when the patient exhibits an increased respiratory rate due to exercise... | Cardiac Pacemakers, Inc. | A61B005/02
US07859264 | B2 | 2010-12-28 | Superconducting loop, saddle and birdcage MRI coils capable of simultaneously imaging small nonhuman animals | 2005-01-20 | New MRI coil and resonators are disclosed based solely on superconducting inductive element and built-in capacitive elements as well as hybrid superconducting-metal inductive and capacitive elements having superior SNR. Single and multiple small animal MRI imaging units are also disclosed including one or more resonators of this invention surrounding one or more small animal cavities. Methods for making and using the MRI coils and/or arrays are also disclosed. | 1. An MRI apparatus capable of imaging small nonhuman animals comprising: a vacuum housing including at least one aperture, where each aperture is configured to receive a small nonhuman animal,a coolant assembly including a coolant inlet,a coolant outlet anda cold plate, where a coolant is configured to cool the cold plate,at least one resonator, each resonator comprising: two closed saddle-shaped coils in thermal contact with the cold plate, each saddle-shaped coil comprising: four members, each member including a superconducting layer, where the members are arranged in order to form four overlapping regions,each overlapping region comprises: a capacitor formed from overlapping portions of the superconducting layers of two of the members and an overlapping region dielectric layer interposed therebetween,where two of the members are straight and two of the members are curvilinear in order to form the closed saddle-shaped coils, andwhere a resonator surrounds each at least one aperture. 2. The apparatus of claim 1, wherein each member further comprises: a substrate dielectric layer upon which the superconducting layer was formed, where the dielectric layer of the straight members comprisea rigid dielectric material and the dielectric layer of the curvilinear members comprise a flexible dielectric material. 3. The apparatus of claim 1, wherein the superconducting layer of the curvilinear members comprise a plurality of substantially flat superconducting segments. 4. The apparatus of claim 3, wherein the overlapping regions of the superconducting layer of the curvilinear member comprise one of the substantially flat superconducting segments. 5. The apparatus of claim 1, wherein the overlapping region dielectric layers comprise separate dielectric layers distinct from the substrate dielectric layers. 6. The apparatus of claim 5, wherein the substrate and the overlapping region dielectric layers are composed of the same or different dielectric material. 7. The apparatus of claim 1, wherein the overlapping region dielectric layers comprise portions of the substrate dielectric layers of the members. 8. The apparatus of claim 1, wherein each coil apparatus includes: a metal layer formed on an exposed portion of a dielectric layer or an external dielectric layer formed on an exposed portion of a superconducting layer with a metal layer formed on the outer surface of the external dielectric layer, where the metal layer forms a coupling or decoupling capacitive element with a corresponding portion of the superconducting layer. 9. The apparatus of claim 8, each coil apparatus further includes: a wire bonded to the metal layer, where the wire links a plurality of the apparatuses together to form an array or to connect the apparatus to a pre-amplifier. 10. The apparatus of claim 1, further comprising: a plurality of separated spaced apart apertures, each configured to receive a small nonhuman animal allowing the MRI apparatus to simultaneously receive... | The University of Houston | G01V003/00 A61B005/055
US20100292948 | A1 | 2010-11-18 | Method of Calibrating an Analyte-Measurement Device, and Associated Methods, Devices and Systems | 2010-07-26 | The invention relates to a method for calibrating an analyte-measurement device that is used to evaluate a concentration of analyte in bodily fluid at or from a measurement site in a body. The method involves measuring a concentration, or calibration concentration, of an analyte in blood from an “off-finger” calibration site, and calibrating the analyte-measurement device based on that calibration concentration. The invention also relates to a device, system, or kit for measuring a concentration of an analyte in a body, which employs a calibration device for adjusting analyte concentration measured in bodily fluid based on an analyte concentration measured in blood from an “off-finger” calibration site. | 1-22. (canceled) 23. A method for calibrating a signal from a subcutaneous sensor, comprising: obtaining a calibration measurement from a calibration sensor placed in contact with capillary blood from an off-finger calibration site within a body;evaluating whether the calibration measurement is within a predetermined range of analyte concentration; anddetermining that the calibration measurement is suitable for use in converting an analyte signal from the subcutaneous sensor into an analyte concentration if the calibration measurement is within the predetermined range,wherein the subcutaneous sensor comprises a working electrode, the working electrode comprising an analyte responsive enzyme and a mediator, wherein at least one of the analyte-responsive enzyme and the mediator is chemically bonded to a polymer disposed on the working electrode. 24. The method of claim 23, wherein at least one of the analyte-responsive enzyme and the mediator is covalently bonded to the polymer. 25. The method of claim 24, wherein at least one of the analyte-responsive enzyme and the mediator forms one or more of carbon-carbon, carbon-nitrogen, or metal-carbon covalent bonds with the polymer. 26. The method of claim 24, wherein at least one of the analyte-responsive enzyme and the mediator is covalently bonded to the polymer through a linker moiety. 27. The method of claim 24, wherein the mediator is covalently bonded to the polymer through one or more ligands of the mediator. 28. The method of claim 27, wherein the one or more ligands comprises a substituted or unsubstituted heterocyclic nitrogen-containing moiety. 29. The method of claim 24, wherein one or more ligands of the mediator are substituents of the polymer backbone. 30. The method of claim 24, wherein the analyte-responsive enzyme is covalently bonded to the polymer through one or more side chains of the enzyme. 31. The method of claim 30, wherein the one or more side chains of the analyte-responsive enzyme is selected from the group consisting of substituted or unsubstituted amines, alcohols, thiols, phenols, imidazoles, indols, and carboxylic acids. 32. The method of claim 24, wherein one or more side chains of the analyte-responsive enzyme are substituents of the polymer backbone. 33. The method of claim 23, wherein the polymer is crosslinked. 34. The method of claim 33, wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer. 35. The method of claim 34, wherein at least one of the analyte-responsive enzyme and the mediator is crosslinked with the polymer through one or more crosslinking agents. 36. The method of claim 35, wherein the one or more crosslinking agents is selected from the group consisting of substituted or unsubstituted epoxides, aldehydes, imidoesters, N-hydroxysuccinimide esters, carbodiimides, cyanuric chloride, tetrachlorobenzoquinone, benzoquinone and tetracyanoquinodimethane. 37. The method of claim 34, wherein one or ... | Abbott Diabetes Care Inc. | G06F019/00 A61B005/145
US07840256 | B2 | 2010-11-23 | Image guided tracking array and method | 2005-12-12 | An array for use with a surgical navigation system is provided. The array comprises a frame and first, second and third markers attached to the frame. The first, second and third markers are detectable by a tracking system used in image guided surgery, and the first marker is movable relative to the frame. In one embodiment, at least one marker slides along the frame from a first position where it is held in place to a second position where it is also held in place. In another embodiment, one or more of the markers is removed from the frame at a first position and reattached to the frame at a second position. In still another embodiment, a portion of the frame itself moves with the movable marker. | 1. An array for use with a surgical navigation system, comprising: a frame having a plurality of retainers configured to a marker thereto; andfirst, second and third markers attached to the frame, the first, second and third markers defining a first three-dimensional configuration of the array that is detectable by a tracking system used in a surgical navigation system;wherein the first marker is movable relative to the frame from a first retainer to a second unoccupied retainer to define a second three-dimensional configuration of the array, the second three-dimensional configuration being different from the first three-dimensional configuration. 2. The array of claim 1, wherein the first marker is slidable on the frame between a first position and a second position. 3. The array of claim 1, further comprising a coupling member disposed at one end of the frame and adapted to releasably and interchangeably attach the array to at least two different surgical components to distinguish between the at least two different surgical components through detection of at least one of the first, second and third markers. 4. The array of claim 1, wherein the frame comprises a movable arm member having the first marker attached thereto, wherein the movable arm member and the first marker move relative to the frame. 5. The array of claim 1, wherein the first marker is detachable from the frame at a first position and reattachable to the frame at a second position. 6. A method of performing a surgery using surgical navigation, comprising: (a) providing a tracking system and an array having a frame having a plurality of retainers with first, second and third markers attached thereto, the array being identified and tracked by the tracking system;(b) using a first surgical component to which the array is attached during the surgery while tracking the first surgical component with the tracking system;(c) detaching the array from the first surgical component and reattaching the array to a second surgical component;(d) moving the position of the first marker from a first position to a second position relative to the frame, wherein the tracking system identifies the second surgical component to which the array is attached; and(e) using the second surgical component to which the array is attached during the surgery while tracking the second surgical component with the tracking system. 7. The method of claim 6, wherein step (d) comprises sliding the first marker along the frame from the first position to the second position. 8. The method of claim 6, wherein step (d) comprises moving a movable arm member having the first marker attached thereto, wherein the first marker moves from the first position to the second position. 9. The method of claim 8, wherein step (d) comprises detaching the first marker from the frame at the first position and reattaching the first marker to the frame at the second position. 10. A surgical navigation system, comprising: a computer... | Biomet Manufacturing Corporation | A61B005/05
US07835783 | B1 | 2010-11-16 | Magnetic resonance imaging methods and apparatus for time-series motion tracking with inversion recovery compensation | 2003-04-21 | Image contributions produced by an untagged specimen magnetization component in magnetic resonance imaging are controlled by applying one or more radiofrequency (RF) pulses that invert at least a portion of the untagged specimen magnetization. In an example, a specimen is tagged with a spatially modulated magnetization that is used to produce an image signal that includes a contribution associated with the tagged magnetization and an untagged magnetization. The untagged magnetization is substantially along an axial direction defined by an applied axial magnetic field. The untagged magnetization increases in magnitude because of so-called T1 relaxation. A contribution to the image signal increases for a predetermined time or to a predetermined magnitude, and a 180-degree pulse is applied to invert at least a portion of the untagged magnetization. The untagged magnetization is then antiparallel with respect to the applied axial magnetic field. Additional inversion recovery causes the untagged magnetization to increase from a negative value to zero and then becomes positive. As a result, signal contributions associated with the untagged magnetization are reduced. Additional 180-degree pulses can be applied whenever the untagged magnetization becomes larger than a predetermined value so that image contrast can be maintained. When the tagged magnetization decreases to a predetermined level, an initial specimen magnetization is reestablished for subsequent imaging. | 1. A magnetic resonance imaging method, comprising: establishing a tagged magnetization simultaneously throughout a specimen;repetitively applying a compensation pulse to the specimen, wherein the compensation pulse is configured to invert at least an untagged specimen axial magnetization component associated with decay of the tagged magnetization;applying an imaging pulse sequence after each repetitively applied compensation pulse; andobtaining images of at least a portion of a specimen based on image signals associated with the imaging pulse sequences, wherein the image signals include contributions from a tagged axial magnetization component and an untagged specimen magnetization associated with the decay of the tagged magnetization. 2. The method of claim 1, wherein the compensation pulse is substantially a π-pulse. 3. The method of claim 1, wherein the image signal is based on a spatially varying amplitude of the tagged axial magnetization component. 4. The method of claim 1, wherein the image signal is based on a spatially varying phase of the tagged axial magnetization component. 5. A method for magnetic resonance imaging, comprising: establishing an initial spatially tagged magnetization throughout a specimen volume;repetitively obtaining image signals associated with the initial spatially tagged magnetization in the volume, wherein image signal acquisitions are interleaved with applications of a compensation pulse that is selected to invert at least a portion of the spatially tagged magnetization and an untagged magnetization produced by decay of the spatially tagged magnetization. 6. The method of claim 5, further comprising: establishing a subsequent spatially tagged magnetization simultaneously throughout a specimen volume;repetitively obtaining image signals associated with the subsequent spatially tagged magnetization in the volume, wherein image signal acquisitions are interleaved with applications of a compensation pulse that is selected to invert at least a portion of the spatially tagged magnetization and at least a portion of an untagged magnetization produced by decay of the spatially tagged magnetization. 7. The method of claim 6, wherein the subsequent spatially tagged magnetization is re-established after a time interval associated with a decrease in magnitude of the initial spatially tagged magnetization. 8. The method of claim 6, further comprising displaying images based on the image signals. 9. A method, comprising: (a) situating a specimen in a substantially constant magnetic field that is directed along a z-axis so as to establish a substantially constant magnetization in at least a selected portion of the specimen, the constant magnetization having a magnitude M0;(b) selecting a flip angle α;(c) establishing a stored longitudinal magnetization having an unencoded component with an initial magnitude MU and an encoded component with an initial magnitude MENC in a specimen, wherein the magnitudes MU and MENC are... | The United States of America as represented by the Department of Health and Human Services | A61B005/05
US08295577 | B2 | 2012-10-23 | Method and apparatus for guiding a device in a totally occluded or partly occluded tubular organ | 2006-10-22 | An apparatus and method for detecting, tracking and registering a device (218) within a tubular organ (210) of a subject. The devices include guide wire (216) tip or therapeutic devices, and the detection and tracking uses fluoroscopic images taken prior to or during a catheterization operation. The devices are fused with images or projections of models depicting tile tubular organs. Tile method and apparatus are used for treating chronic total occlusion or near total occlusion situations, by navigating a driller along the tubular organ (210) proximally to the occlusion point, in areas which are-not viewable in an angiogram, and optionally enabling the penetration of the occlusion in a preferred area. | 1. A method for displaying an image of a device within an occluded tubular organ of a subject, for assisting in navigating the device through an occluded area of the tubular organ, the image representing the device, the occluded area of the tubular organ, an at least one part of the tubular organ proximal to the occluded area of the tubular organ, and an at least one part of the tubular organ distal to the occluded area of the tubular organ, the method comprising: registering a current image taken by an imaging modality showing at least two points of the device, with a last injection image, the last injection image imaging an at least one part of the tubular organ proximal to the occluded area, and a part immediately distal to the occluded area appearing insufficiently distinguishable from background;registering the last injection image with a reference image showing the occluded area, an at least one part of the tubular organ proximal to the occluded area, and an at least one part of the tubular organ distal to the occluded area;fusing information from the current image, the last injection image and the reference image, to obtain a combined image comprising the device, the occluded area of the tubular organ, an at least one part proximal to the occluded area of the tubular organ, and an at least one part distal to the occluded area of the tubular organ. 2. The method of claim 1 wherein the reference image is constructed from a reference source, the reference source constructed from images taken by a second imaging modality. 3. The method of claim 2 wherein the second imaging modality is a Computerized Tomography imaging device. 4. The method of claim 1 further comprising a step of navigating the device along the tubular organ using the at least one combined image to penetrate the occluded area. 5. The method of claim 1 wherein the imaging modality is an angiogram. 6. The method of claim 1 further comprising the steps of: receiving at least two images of at least two cross sections of the tubular organ;determining a penetration area; andpenetrating the occluded area at the penetration area. 7. The method of claim 6 wherein the at least two cross sections are constructed from the reference source. 8. The method of claim 1 when used for penetrating a chronic total occlusion. 9. The method of claim 1 when used for penetrating a chronic near total occlusion. 10. An apparatus for aiding the navigation of a device within a an occluded tubular organ of a subject, the apparatus comprising: at least one detection component for detecting on an image the device within the tubular organ of the subject;at least one registration component for: registering a current image taken by an imaging modality showing at least two points of the device, with a last injection image, the last injection image imaging an at least one part of the tubular organ proximal to the occluded area, and an area immediately distal to the occluded area appearing insufficiently d... |  | G06K009/00
US08292807 | B2 | 2012-10-23 | Mobile medical workstation | 2010-09-01 | A medical workstation is defined by a supporting structure having at least one medical diagnostic instrument disposed thereupon. A first display is further disposed on a first side of the supporting structure and a second display is disposed on a second side of the supporting structure in a manner substantially opposite from the first display. Each of the first and second displays are interconnected to the at least one medical diagnostic instrument to permit at least one of the displays in order to display diagnostic results and are tandemly or independently controllable. | 1. An integrated apparatus for use in a patient encounter, said apparatus comprising: at least one medical diagnostic instrument including a vital signs monitoring device for measuring various physiological parameters of a patient, said physiological parameters including blood pressure, glucose, body temperature, pulse rate and blood oxygen saturation; anda computing device connected to said at least one medical diagnostic instrument, wherein at least one of said computing device and said at least one medical diagnostic instrument are programmed to detect changes in a patient condition based on changes in two or more parameters of said patient and in which said computing device is programmed to alert the user that changes in said patient condition have exceeded at least one predetermined threshold, andin which the changes in said at least two physiological parameters are used by said computing device to predict the onset of sepsis, wherein said computing device is programmed to calculate a glycemic control score of a patient based on historical data and to determine whether a patient's blood sugar level is within a predetermined range. 2. An integrated apparatus as recited in claim 1, wherein said computing device is programmed to perform an independent determination as to whether the patient has had an infection. 3. An integrated apparatus as recited in claim 2, wherein the computing device is programmed to make a further determination to whether the patient has at least two signs of Systemic Inflammatory Response Syndrome (SIRS) if the existence of a known or suspected infection has been determined. 4. An integrated apparatus as recited in claim 3, wherein the computing device is programmed to make a determination as to whether at least one organ of the patient is failing or dysfunctional if at least two signs of SIRS have been identified. 5. An integrated apparatus for use in a patient encounter, said apparatus comprising: at least one medical diagnostic instrument including a vital signs monitoring device for measuring various physiological parameters of a patient, said physiological parameters including blood pressure, glucose, body temperature, pulse rate and blood oxygen saturation; anda computing device connected to said at least one medical diagnostic instrument, wherein at least one of said computing device and said at least one medical diagnostic instrument are programmed to detect changes in a patient condition based on changes in two or more parameters of said patient and in which said computing device is programmed to alert the user that changes in said patient condition have exceeded at least one predetermined threshold, andin which the changes in said at least two physiological parameters are used by said computing device to predict the onset of sepsis, wherein said apparatus is programmed to determine a first control score based on a first parameter and a second alert score based on at least a second parameter. 6. An integra... | Welch Allyn, Inc. | A61B005/00 A61B005/05 G09G001/14 G08B005/22 G08B023/00
US09037213 | B2 | 2015-05-19 | Method and apparatus to estimate location and orientation of objects during magnetic resonance imaging | 2009-01-08 | Tracking based on the gradient fields of magnetic resonance imaging (MRI) scanners based on passive operation of the tracking system without any change of the scanner's hardware or mode of operation. To achieve better tracking performance, a technique to create a custom MRI pulse sequence is disclosed. Through this technique any standard pulse sequence of the scanner can be modified to include gradient activations specifically designated for tracking. These tracking gradient activations are added in a way that does not affect the image quality of the native sequence. The scan time may remain the same as with the native sequence or longer due to the additional gradient activations. The tracking system itself can use all the gradient activations (gradient activations for imaging and gradient activations for tracking) or eliminate some of the gradients and lock onto the specific gradient activations that are added to the custom pulse sequence. | 1. A method of determining an instantaneous location and orientation of an object moving through a three-dimensional space within an imaging space of a magnetic resonance imaging apparatus comprising three gradient coils during operation of said magnetic resonance imaging apparatus, the method comprising: producing gradient field maps for said gradient coils prior to a scanning operation of the magnetic resonance imaging apparatus, wherein said gradient field maps are produced using tracking gradient command parameters;applying to said object a sensor for measurement of instantaneous values of magnetic fields which are generated by activation of said gradient coils within an imaging space of said magnetic resonance imaging apparatus;activating gradient coils of said magnetic resonance imaging apparatus during said scanning operation of said magnetic resonance imaging apparatus in imaging mode gradient activation so as to produce a magnetic field used for imaging;activating said gradient coils of said magnetic resonance imaging apparatus during said scanning operation of said magnetic resonance imaging apparatus using said tracking gradient command parameters in tracking mode gradient activation so as to produce a magnetic field used for tracking, wherein said magnetic field used for tracking is produced using only said three gradient coils without the need for RF activation;measuring instantaneous values of magnetic fields which are generated in said sensor by the tracking mode activation of said gradient coils of said magnetic resonance imaging apparatus;calculating predicted values of magnetic fields to be generated in said sensor by the tracking mode activation of said gradient coils by using said gradient field maps; andprocessing said measured instantaneous values of the magnetic fields generated by the tracking mode activation of said gradient coils, together with said calculated predicted values of magnetic fields to compute the instantaneous location and orientation of said object within said imaging space. 2. The method of claim 1, wherein said activating said gradient coils using said tracking gradient command parameters comprises choosing tracking gradient command parameters for bi-modally activating said gradient coils of said magnetic resonance imaging apparatus, wherein said bi-modal activation has a zero net effect on said imaging. 3. The method of claim 1, wherein said activating said gradient coils using said tracking gradient command parameters comprises measuring said command parameters. 4. The method of claim 1, wherein said activating said gradient coils using said tracking gradient command parameters comprises retrieving stored information on said command parameters. 5. The method of claim 1, wherein said activating said gradient coils using said tracking gradient command parameters comprises identifying an amplitude and rate of change of current flow of said tracking mode gradient activation signal. 6. The method of cl... | ROBIN MEDICAL INC. | A61B005/05 G01R033/028 G01R033/28
US08433385 | B2 | 2013-04-30 | Implantable medical device | 2010-06-18 | An analyte measuring system has an implantable medical device having a signal source arranged for generating a current signal and electrodes for applying the current signal to a surrounding tissue in a subject body. The device measures a resulting voltage signal with the electrodes and calculates an impedance signal therefrom. The system comprises a signal processor arranged for generating an estimate of a concentration of an analyte in the tissue based on a spectrum analysis of the determined impedance signal. | 1. An analyte measuring system comprising: an implantable medical device comprising: an electrode connecting arrangement connectable to at least two implantable electrodes;a signal source that generates a current or voltage signal and for applying said current or voltage signal to a tissue using said at least two implantable electrodes;a signal measurer that measures a resulting voltage or current signal from said tissue sensed by said at least two implantable electrodes;an impedance determiner that determines an impedance signal based on said generated current or voltage signal and said measured resulting voltage or current signal; anda signal processor that generates an estimate of a concentration of an analyte in said tissue by making a spectral analysis of said impedance signal determined by said impedance determiner;wherein said signal processor generates said estimate of said analyte concentration based on a ratio of an amplitude of said impedance signal at a first frequency and an amplitude of said impedance signal at a second frequency; andwherein said signal processor generates said estimate of said analyte concentration based on a comparison of said ratio and a corresponding ratio determined by said signal processor based on a spectral analysis of a reference impedance signal determined based on a previously applied current or voltage signal and a previously measured resulting voltage or current signal. 2. The system according to claim 1, wherein said signal source generates a composite current or voltage signal having a defined bandwidth covering a frequency range;said signal measurer measures a resulting composite voltage or current signal from said tissue sensed by said at least two implantable electrodes;said impedance determiner determines said impedance signal based on said generated composite current or voltage signal and said measured composite resulting voltage or current signal; andsaid signal processor generates said estimate of said concentration of said analyte based on an analysis of said impedance signal at multiple selected frequencies in said frequency range obtained by bandpass filtering said impedance signal. 3. The system according to claim 1, wherein said implantable medical device comprises a transmitter with a connected transmitting antenna for transmitting a representation of said impedance signal to a non-implantable device comprising a receiver with a connected receiving antenna that receives said representation, wherein said signal processor is arranged in said non-implantable device. 4. The system according to claim 1, wherein said signal processor is arranged in said implantable medical device. 5. The system according to claim 1, wherein said signal source generates said current or voltage of at least one frequency in a range of 1 MHz to 1 THz. 6. The system according to claim 5, wherein said signal source generates said current or voltage of at least one frequency in a range of 10 MHz to 10 GHz. 7. Th... | St. Jude Medical AB | A61B005/05
US08725233 | B2 | 2014-05-13 | Load detection device and load detection method | 2008-12-04 | According to this load detection method, load (Fi) in insertion of a linear body having flexibility in an inlet of a hollow tube inserted in a bending vessel in a body is detected, an image of the hollow tube or the linear body is taken, a degree of bending of the hollow tube or the linear body is detected based on the image, and load (Fo) at a tip end of the linear body is calculated based on the detected insertion load (Fi) and the degree of bending. Therefore, the load (Fo) at the tip end of the thin linear body can readily be detected. | 1. A load detection device for detecting load applied to a tip end of a linear body having flexibility when said linear body is inserted in a hollow tube inserted in a bending vessel, comprising: a degree of bending detection device configured to detect a degree of bending of said vessel, said hollow tube, or said linear body;an insertion load detection device configured to detect load in insertion of said linear body in an inlet of said hollow tube; anda processing unit communicatively coupled to the degree of bending detection device and to the insertion load detection device and configured to find the load applied to the tip end of said linear body based on a result of detection by the degree of bending detection device and the insertion load detection device, and based on a coefficient of friction between the linear body and the hollow tube,wherein said processing unit is further configured to correct said coefficient of friction based on variation in flexural rigidity depending on a position in said linear body where the flexural rigidity of said linear body varies depending on the position in said linear body. 2. The load detection device according to claim 1, wherein said degree of bending detection device is further configured to detect the degree of bending of said vessel, said hollow tube, or said linear body based on an image of said vessel, said hollow tube, or said linear body. 3. The load detection device according to claim 1, wherein said coefficient of friction is measured in advance. 4. The load detection device according to claim 1, wherein said vessel is a blood vessel,said hollow tube is a catheter, andsaid linear body is a treatment wire to be inserted in said catheter. 5. The load detection device according to claim 4, wherein said treatment wire includes a delivery wire and an embolization coil provided at a tip end portion of the delivery wire, for filling a cerebral aneurysm. 6. The load detection device according to claim 5, wherein said degree of bending detection device is further configured to detect a degree of bending of said treatment wire when a tip end of said embolization coil enters the cerebral aneurysm, andsaid processing unit is further configured to find load applied to said cerebral aneurysm by said embolization coil. 7. The load detection device according to claim 4, wherein said degree of bending detection device is further configured to detect a degree of bending of said blood vessel based on a three-dimensional image of said blood vessel. 8. The load detection device according to claim 4, wherein said degree of bending detection device is further configured to detect a degree of bending of said treatment wire based on a two-dimensional image of said treatment wire taken by using X-rays during a surgical operation and a three-dimensional image of the blood vessel taken in advance, as superimposed on each other. 9. The load detection device according to claim 1, further comprising a drive unit co... | NTN Corporation | A61B005/05 A61B005/103 A61M029/00
US08571647 | B2 | 2013-10-29 | Impedance based anatomy generation | 2009-05-08 | Methods and systems for determining anatomical information based on signals measured by multiple, spatially distributed electrodes on a catheter are disclosed herein. In some examples, methods and systems disclosed herein can include causing current to flow between at least some of the electrodes, measuring an electrical signal at each of one or more measuring electrodes, and determining anatomical information based on the measured signals. | 1. A method comprising: inserting a catheter into a heart, the catheter comprising three or more electrodes;causing the catheter to move to each of multiple, different positions in the heart;for each of the multiple, different catheter positions, causing current to flow between at least some of the electrodes and in response to current flow, measuring an electrical signal at each of one or more of the electrodes; andsubsequent to causing current flow and measuring the electrical signals at the multiple, different catheter positions, determining anatomical information about the heart based on positions of the catheter electrodes at a plurality or the multiple, different catheter positions, the measured signals at the plurality of the multiple, different catheter positions, and a surface location threshold applied to a measure related to a distance between a surface of the heart and the catheter;wherein determining the anatomical information based on the measured signals further compromise:distinguishing electrical signals indicative of cardiac electrical activity from those responsive to the injected current; anddetermining the anatomical information based on the electrical signals responsive to the injected current. 2. The method of claim 1, wherein the determination of the anatomical information accounts for a change in conductivity at the cardiac chamber boundary. 3. The method of claim 2, wherein the determination accounts for a first conductivity inside the cardiac chamber boundary and a second conductivity outside the cardiac chamber boundary. 4. The method of claim 1, wherein determining the anatomical information comprises: generating impedance information based on the measured signals at the different catheter positions; anddetermining the anatomical information based at least in part on the impedance information. 5. The method of claim 4, wherein the impedance information is based on a conductivity contrast between blood and surrounding tissue. 6. The method of claim 4, wherein the impedance information is based on a permittivity contrast between blood and surrounding tissue. 7. The method of claim 1, wherein determining the anatomical information compromises: generating impedance information based on the measured signals at the different catheter positions, the impedance information comprising complex impedance information based on both a permittivity contrast and a conductivity contrast between blood and surrounding tissue; anddetermining the anatomical information based at least in part on the impedance information. 8. The method of claim 1, wherein the anatomical information comprises a representation of at least a portion of a boundary of the heart. 9. The method of claim 1, wherein determining the anatomical information comprises detecting a boundary of the heart and the method further comprises displaying at least a portion of the anatomical information. 10. The method of claim 1, wherein determining the anatomical inform... | Rhythmia Medical, Inc. | A61B005/00
US08369917 | B2 | 2013-02-05 | Finger-clipped oximeter with finger pressed plate | 2008-05-20 | A finger-clipped oximeter includes an upper case (2) in which a measuring element (4), a data processing circuit, batteries (6) and a display (3) are provided, a rectangular through hole being provided in the upper case (2); a finger pressed plate (10) in which a measuring element is also provided, a finger inserting hole (11), a rectangular column (7) with a position corresponding to that of the rectangular through hole of the upper case (2); and a pair of connectors which are capable of ascending or descending, each of the connectors including one U-shaped connecting rod (8). | 1. A finger-clipped oximeter, wherein the finger-clipped oximeter comprises: an upper case in which batteries and a display are provided, a rectangular through hole being provided in the upper case;a finger pressed plate, a finger inserting hole being formed between an upper surface of the finger pressed plate and a lower surface of the upper case, a rectangular column with a position corresponding to that of the rectangular through hole of the upper case being provided on the finger pressed plate; anda pair of connectors which are capable of ascending or descending, each of the connectors including one U-shaped connecting rod, a middle portion of the U-shaped connecting rod being a straight portion which extends through an edge portion on a side of the finger pressed plate, two arm portions of the U-shaped connecting rod being inserted into the upper case to serve as a connecting piece between the upper case and the finger pressed plate. 2. The finger-clipped oximeter according to claim 1, wherein the middle portion of the U-shaped connecting rod is pressure-cast to an edge portion on a side of the finger pressed plate. 3. The finger-clipped oximeter according to claim 2, wherein at least a portion of the middle portion of the U-shaped connecting rod has a structure which can prevent the U-shaped connecting rod from being rotated with respect to the finger pressed plate. 4. The finger-clipped oximeter according to claim 1, wherein at least one column like portion which extends straight up is provided at each of four corners of an upper surface of a bottom wall of the upper case, and positions of the two arm portions of the U-shaped connecting rod are limited by the at least one column like portion. 5. The finger-clipped oximeter according to claim 1, wherein the upper case has a top cover. 6. The finger-clipped oximeter according to claim 1, wherein the finger pressed plate is made of plastics or silica gel. 7. The finger-clipped oximeter according to claim 1, wherein a V-shaped portion which extends toward the middle portion of the U-shaped connecting rod is provided at each end of the two arm portions of the U-shaped connecting rod. 8. The finger-clipped oximeter according to claim 7, wherein the V-shaped portion is opened toward two sides of the U-shaped connecting rod, or is opened toward the straight portion in the middle portion of the U shaped connecting rod. 9. The finger-clipped oximeter according to claim 1, wherein one movable magnet is provided at each end of the two arm portions of the U-shaped connecting rod, while one stationary magnet is provided at each root of the two arm portions of the U-shaped connecting rod, and the movable magnet and the stationary magnet are provided so that magnetic poles with the same polarity are opposed to each other. | Beijing Choice Electronic Technology Co., Ltd. | A61B005/00
US09827388 | B2 | 2017-11-28 | Methods and apparatus for detecting and treating respiratory insufficiency | 2010-03-31 | Devices and systems provide methods of detecting a severity change in respiratory insufficiency (RI) or chronic obstructive pulmonary disease (COPD) condition of a patient. In an example embodiment, a detection monitoring device determines one or more severity change indicators based on a measure of supplied pressure or other representative measure determined by the device. The supplied pressure may optionally be determined during pressure treatment that satisfies a target ventilation. The supplied pressure or representative data may be compared to one or more thresholds that are selected to represent a change in the condition of the RI or COPD patient such as an exacerbation of a prior condition. Results of the comparisons may trigger one or more warnings or messages to notify a patient or physician of a pending change to the patient's RI or COPD condition so that the patient may more immediately seek medical attention to treat the condition. | 1. A device for monitoring a patient to evaluate a respiratory insufficiency condition of the patient comprising: a transducer to generate a signal representative of a pressure delivered by a flow generator; anda processor, coupled with the transducer, the processor configured to control an evaluation of a severity change indicator based on detecting, using a measure of determined pressure from the signal, a change in the measure of determined pressure from a plurality of treatment sessions, the evaluation comprising a detection of a continuous rise in the measure of determined pressure successively through the plurality of treatment sessions, the indicator representing information about a change in severity of a respiratory insufficiency condition of the patient,wherein the respiratory insufficiency condition is attributable to an insufficiency condition of the patient's lungs. 2. The device of claim 1 further comprising the flow generator configured to generate a flow of breathable gas at a pressure above atmospheric pressure to a patient interface, wherein the processor is configured to control the flow generator. 3. The device of claim 2 wherein the flow generator is configured to generate the flow of breathable gas at a pressure above atmospheric pressure to maintain a target ventilation, and wherein the measure of determined pressure comprises a measure of pressure generated by the flow generator in order to maintain the target ventilation. 4. The device of claim 3 wherein the target ventilation is a minute ventilation and the measure of determined pressure is an average pressure. 5. The device of claim 3 wherein the target ventilation is an estimate of alveolar ventilation. 6. The device of claim 3 wherein the evaluation of the severity change indicator comprises a threshold comparison of the continuous rise in the measure of determined pressure successively through the treatment sessions. 7. The device of claim 6 wherein the threshold comparison involves the measure of determined pressure from one of the plurality of treatment sessions and a prior measure of determined pressure used to maintain the target ventilation during a prior treatment session of the plurality of treatment sessions. 8. The device of claim 6 wherein the threshold comparison represents an indication of a repeated rise in the measure of determined pressure over a period of days. 9. The device of claim 6 wherein the threshold comparison involves a difference between the measure of determined pressure from one of the plurality of treatment sessions and a prior measure of determined pressure used to maintain the target ventilation during a prior treatment session of the plurality of treatment sessions. 10. The device of claim 6 wherein the threshold comparison includes a ratio of the measure of determined pressure from one of the plurality of treatment sessions and a prior measure of determined pressure used to maintain the target ventilation during a prior treat... | ResMed Limited | A61B005/08 A61M016/00 A61B005/00
US10020075 | B2 | 2018-07-10 | Systems and methods for monitoring and/or managing patient orientation using a dynamically adjusted relief period | 2010-03-24 | A monitoring system and method tracks a patient's position over time and ensures that proper turning or other manipulation is done within the time prescribed. Preferably, the techniques herein continuously monitor patient position and alert medical or other personnel of the need for turning or other patient manipulation. The system may be implemented within a medical or other care facility, or within a patient's home. | 1. A system for managing patient orientation according to a patient turn protocol, the system comprising: a sensor configured to be physically wearable by a patient and outputting sensor-based orientation information representative of changes in the patient's body orientation with respect to a support surface, during a series of (a) pressure-on periods in which the patient is in a pressure-on orientation in which a particular area of the patient's body experiences pressure from the support surface, wherein the series of pressure-on periods have varying durations, and (b) relief periods in which the patient is in a relief orientation other than the pressure-on orientation;a processor and memory in communication with the sensor and configured to: receive and store the sensor-based orientation information together with temporal information associated with the sensor-based orientation information;generate commands or alerts for changing the orientation of the patient according to the patient turn protocol based on information including: (a) the stored sensor-based orientation information and associated temporal information;(b) a maximum time-on threshold that indicates a maximum time the patient is permitted to remain in the pressure-on orientation during each respective pressure-on period before changing to the relief orientation to help avoid development of a pressure ulcer; and(c) a minimum time-off threshold that indicates a minimum time the patient is required to remain in the relief orientation during each respective relief period before returning to the pressure-on orientation; andfor at least some of the relief periods, dynamically set the minimum time-off threshold for each respective relief period by: measuring a duration of a pressure-on period immediately preceding the respective relief period based on the sensor-based orientation information and associated temporal information; andcalculating a minimum time-off threshold for the respective relief period based on a proportional relationship between the minimum time-off threshold and a respective pressure-on duration including the measured duration of the pressure-on period immediately preceding the respective relief period;wherein according to the proportional relationship, increasing the respective pressure-on duration results in a proportional increase in the calculated minimum time-off threshold, such that numerically different minimum time-off thresholds are calculated for different relief periods, depending on the respective pressure-on durations preceding the different relief periods; anda device configured to output the commands or alerts for changing the orientation of the patient according to the patient turn protocol. 2. The system of claim 1, wherein the monitor comprises a bedside receiving station. 3. The system of claim 2, wherein the bedside receiving station is an intelligent mobile device. 4. The system of claim 1, wherein the monitor comprises an attendant receiving s... | LEAF HEALTHCARE, INC. | G16H040/63 A61B005/11 A61B005/00 G06F019/00
US09304202 | B2 | 2016-04-05 | Multiuse optical sensor | 2010-05-20 | One or more electromagnetic radiation sources, such as a light emitting diode, may emit electromagnetic waves into a volume of space. When an object enters the volume of space, the electromagnetic waves may reflect off the object and strike one or more position sensitive detectors after passing through an imaging optical system such as glass, plastic lens, or a pinhole located at known distances from the sources. Mixed signal electronics may process detected signals at the position sensitive detectors to calculate position information as well as total reflected light intensity, which may be used in medical and other applications. A transparent barrier may separate the sources and detectors from the objects entering the volume of space and reflecting emitted waves. Methods and devices are provided. | 1. A device comprising: a source emitting electromagnetic radiation through a transparent barrier;a position sensitive detector behind a first side of the barrier measuring a position and intensity of the electromagnetic radiation reflected off a body part of an organism placed on a surface of a second side of the barrier, the source modulated to reduce an effect of ambient radiation on the measurement; andelectronics to calculate (a) a position of the body part on the surface of the barrier based on the position of the measured reflected radiation and (b) a photoplethysmograph from changes to the reflected light intensity;wherein the position sensitive detector measures the position of the reflected electromagnetic radiation by generating first and second currents at first and second ends, respectively, of the position sensitive detector, the first and second currents varying depending on a light distribution of the reflected electromagnetic radiation on a surface of the position sensitive detector facing the first side of the barrier. 2. The device of claim 1, wherein the calculated position of the body part updates as the body part is moved. 3. The device of claim 1, wherein the electronics further calculate a heart rate of the organism from the photoplethysmograph. 4. The device of claim 3, wherein the electronics are configured to instruct a user where to move the body part to improve the reflected light intensity. 5. The device of claim 1, further comprising a plurality of sources, each emitting electromagnetic radiation having a unique wavelength, the sources modulated to identify the source of reflected radiation at the detector. 6. The device of claim 5, wherein the wavelength of the radiation from a first source is around 660 nm and the wavelength of the radiation from a second source is around 900 nm. 7. The device of claim 5, wherein the electronics further calculate a measure of oxygenation. 8. The device of claim 1, wherein the device is a computing device. 9. The device of claim 8, wherein the computing device is portable. 10. The device of claim 9, wherein the computing device is a personal digital assistant. 11. The device of claim 1, wherein the electronics determine that higher levels of scattered light correspond to greater pressure applied by the body part. 12. The device of claim 11, further comprising a user interface, and wherein at least one of: a degree of zooming and highlighting performed on the user interface, indicates a level of pressure exerted on the user interface. 13. The device of claim 1, wherein the source emits the electromagnetic radiation in at least two different wavelengths, and the electronics calculate the photoplethysmograph for each one of the different wavelengths. 14. The device of claim 1, wherein the barrier includes one of a focusing lens and an aperture. 15. The device of claim 14, wherein one of a focusing lens and an aperture is located in the barrier and adjacent to the surface... | ANALOG DEVICES, INC. | A61B005/00 G01S017/06 A61B005/1455 G01S003/783 G01S005/16 G01S017/87 A61B005/024
US09775529 | B2 | 2017-10-03 | Body-worn pulse oximeter | 2009-09-14 | The invention provides a body-worn system that continuously measures pulse oximetry and blood pressure, along with motion, posture, and activity level, from an ambulatory patient. The system features an oximetry probe that comfortably clips to the base of the patient's thumb, thereby freeing up their fingers for conventional activities in a hospital, such as reading and eating. The probe secures to the thumb and measures time-dependent signals corresponding to LEDs operating near 660 and 905 nm. Analog versions of these signals pass through a low-profile cable to a wrist-worn transceiver that encloses a processing unit. Also within the wrist-worn transceiver is an accelerometer, a wireless system that sends information through a network to a remote receiver, e.g. a computer located in a central nursing station. | 1. A method for determining an oxygen saturation value from a patient, comprising the following steps: (a) positioning a first radiation source, a second radiation source, and a radiation detector on the base of the thumb such that the first radiation source and second radiation source are positioned proximally to one another and the radiation detector is positioned at an included angle of between 35° and 55°, wherein the included angle is the angle formed between a first plane drawn parallel to a planar surface comprising the emitting surface of the first and second radiation sources and a second plane drawn parallel to a planar surface of a circuit board comprising the receiving surface of the radiation detector, and such that a first radiation emitted by the first radiation source and a second radiation emitted by the second radiation source pass through the thumb tissue and are received by the princeps pollicis artery before reaching the radiation detector;(b) measuring a first time-dependent signal by detecting the first radiation emitted by the first radiation source after the first radiation passes through the thumb tissue;(c) measuring a second time-dependent signal by detecting the second radiation emitted by the second radiation source after the second radiation passes through a portion of the patient;(d) measuring a time-dependent motion signal with at least one motion sensor;(e) measuring a time-dependent electrical signal with an electrical sensor comprising at least two electrodes;(f) processing the time-dependent electrical signal to determine at least one reference point;(g) processing the time-dependent motion signal to determine if a motion-related event that affects at least one of the first and second time-dependent signals has occurred that would render the first and second time-dependent signals unreliable for calculation of SpO2;(h) if the motion event that would render the first and second time-dependent signals unreliable for calculation of SpO2 has not occurred, processing both the first and second time-dependent signals from the reference point to determine processed forms of the first and second time-dependent signals; and(i) determining the oxygen saturation value from the processed forms of the first and second time-dependent signals. 2. The method of claim 1, wherein step (g) further comprises processing the time-dependent motion signal to determine a motion-related event selected from the patient's posture, activity level, and degree of motion. 3. The method of claim 2, wherein the motion-related event is posture, and step (g) further comprises determining a parameter corresponding to orientation of the patient's torso. 4. The method of claim 3, wherein the parameter corresponding to orientation of the patient's torso is a vector. 5. The method of claim 4, wherein step (g) further comprises processing signals corresponding to three axes of a motion sensor to determine the vector. 6. The method of claim 5, wher... | SOTERA WIRELESS, INC. | A61B005/1455 A61B005/0245 A61B005/00 A61B005/021 A61B005/022 A61B005/0452 A61B005/11
US20100332140 | A1 | 2010-12-30 | METHOD OF ASSESSING THE EATING EXPERIENCE OF A COMPANION ANIMAL | 2009-06-30 | A method of assessing the eating experience of a companion animal. Further, a feeding device utilized in assessing the eating experience of a companion animal. Further, a method of modifying at least one attribute of a food product as a result of assessing the eating experience of a companion animal. | 1. A method of assessing the eating experience of a companion animal, the method comprising the steps of: a. providing a feeding device, the feeding device comprising: i. a food product placement area;ii. a weigh scale operatively associated with the food product placement area to weigh a food product placed on the food product placement area;iii. a food product placement area image collector to collect image data of a behavioral response of the companion animal to the food product placed on the food product placement area;b. placing a food product on the food product placement area;c. collecting image data of the companion animal's behavioral response to the food product placed on the food product placement area;d. collecting consumption data; ande. correlating the image data with the consumption data to assess the eating experience of the companion animal. 2. The method of claim 1 wherein the food product is placed into a food receptacle and the food receptacle containing the food product is placed on the food product placement area. 3. The method of claim 2 wherein the food receptacle is transparent. 4. The method of claim 1 wherein the assessment of the eating experience of the companion animal is conducted in an environment selected from the group consisting of a home, a veterinarian clinic, a research laboratory, a breeder environment, or a kennel. 5. The method of claim 4 wherein the assessment of the eating experience of the companion animal is conducted at home. 6. The method of claim 4 wherein the environment comprises an environment image collector. 7. The method of claim 1 further comprising a step of administering an eating experience perception questionnaire to an owner, breeder or caregiver of the companion animal and collecting eating experience perception data from the eating experience perception questionnaire. 8. The method of claim 7 further comprising a step of comparing the eating experience perception conclusion to the assessment of the eating experience of the companion animal. 9. The method of claim 1 wherein the feeding device further comprises a sound collector. 10. The method of claim 1 wherein the feeding device further comprises a detection device. 11. A feeding device comprising: a. a food product placement area;b. a weigh scale operatively associated with the food product placement area to weigh a food product placed on the food product placement area;c. a food product placement area image collector to obtain image data of a behavioral response of a companion animal to the food product placed on the food product placement area; andd. a detection device to indicate the presence of the companion animal at the feeding device. 12. The feeding device of claim 11 further comprising a sound collector. 13. The feeding device of claim 11 further comprising a second food product placement area and a second weigh scale operatively associated with the second food product placement area to weigh a second food produc... | - | G06F019/00 A01K005/00 G06K009/00 A61B005/00 A23K001/18
US20100292553 | A1 | 2010-11-18 | Analyte Monitoring Device and Methods of Use | 2010-07-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A device for monitoring glucose concentration in a biological sample of a host, the device comprising: a substantially continuous glucose sensor that produces a data stream indicative of a glucose concentration in a host, the data stream comprising a plurality of time spaced sensor data points;an integrated receiver that receives the data stream from the substantially continuous glucose sensor, wherein the integrated receiver comprises: an in vitro single point glucose monitor configured to receive a biological sample from the host and to measure the concentration of glucose in the sample;a processor; and a computer readable memory comprising: instructions configured to cause the processor to calibrate the data stream received from the continuous glucose sensor;instructions configured to cause the processor to determine a rate of change of the calibrated data stream from the substantially continuous analyte sensor; andinstructions configured to cause the processor to not re-calibrate and/or update calibration of the data stream using a glucose concentration measured by the single point glucose monitor when the rate of change of the calibrated data stream is determined to be aberrant, wherein the rate of change of the calibrated data stream is determined to be aberrant when the rate of change of the calibrated data stream is greater than +/−0.5 mg/dL/min. 3. The device of claim 2, wherein the rate of change of the calibrated data stream is determined to be aberrant when the rate of change of the calibrated data stream is greater than +/−0.25 mg/dL/min. 4. The device of claim 2, wherein the integrated receiver comprises a user interface, wherein the user interface is configured to request a biological sample for measurement via the single point glucose monitor only when the rate of change of the calibrated data stream is not determined to be aberrant. 5. The device of claim 2, wherein the computer readable memory further comprises instructions configured to cause the processor to receive user information from an external source and process the user information received from the external source. 6. The device of claim 5, wherein the user information comprises information selected from the group consisting of mealtime information, exercise information, insulin administration, therapy recommendations and reference analyte values. 7. The device of claim 2, wherein the computer readable memory further comprises instructions configured to detect at least one condition selected from the group consisting of present hypoglycemia, predicted hypoglycemia, present hypergylcemia and predicted hypoglycemia, wherein the instructions are configured to trigger an alarm or alert in response to the detection. 8. The device of claim 2, wherein the integrated receiver comprises a user interface configured to display continuous glucose sensor data and single point glucose monitor data. 9. The device of claim 2, wherein the computer readable memo... | Abbott Diabetes Care Inc. | A61B005/05
US20100305413 | A1 | 2010-12-02 | SENSING APPARATUS | 2008-08-05 | A process for preparing a tube for insertion into the human or animal body, said process comprising: (a) providing a hollow tube comprising a first material and a second material, the first material being in the form of a coil or a tubular mesh, and the second material coating the first material in order to form a continuous substantially impermeable outer wall of the hollow tube; (b) selectively removing a portion of said second material in order to generate at least one opening in a region of the outer wall, while retaining the first material in said region. | 1. A process for preparing a tube for insertion into the human or animal body, said process comprising: (a) providing a hollow tube comprising a first material and a second material, the first material being in the form of a coil or a tubular mesh, and the second material coating the first material in order to form a continuous substantially impermeable outer wall of the hollow tube;(b) selectively removing a portion of said second material in order to generate at least one opening in a region of the outer wall, while retaining the first material in said region. 2. A process as claimed in claim 1 wherein the second material comprises a polymeric material. 3. A process as claimed in claim 2 wherein the second material comprises a synthetic polymeric material selected from polyamides, polyesters, polyurethanes, polyolefins, fluoropolymers and mixtures thereof. 4. A process as claimed in claim 1 wherein the second material is selectively removed by laser ablation. 5. A process as claimed in claim 1 wherein the first material comprises at least one metal selected from stainless steel, titanium and gold. 6. A process as claimed in claim 1 wherein the first material is in the form of a tubular mesh, said mesh being formed from a plurality of filaments. 7. A process as claimed in claim 6 wherein the filaments are arranged in order to define a mesh structure having multiple open areas of at least 0.3 cm2 per cm2 of the mesh structure. 8. A process as claimed in claim 6 wherein the first material comprises a plurality of helically wound filaments, at least a first said filament extending helically in the opposite sense to at least a second said filament. 9. A process as claimed in claim 1 wherein the first material forms a tubular mesh in which the mesh structure is a braid. 10. A process according to claim 1 wherein the first material is in the form of a coil, and wherein the size of the opening formed by selective removal of the second material can be varied by extending or compressing said coil. 11. A process according to claim 1 wherein the second material is selectively removed in the shape of an arc extending around at least a portion of the circumference of the tube. 12. A process as claimed in claim 11 wherein the arc formed by selective removal of the second material does not extend around the complete circumference of the tube. 13. A process as claimed in claim 1 wherein the second material is selectively removed in the shape of a strip or strips extending longitudinally along at least a portion of the tube. 14. A tube for insertion into the human or animal body which is preparable according to a process comprising: (a) providing a hollow tube comprising a first material and a second material, the first material being in the form of a coil or a tubular mesh, and the second material coating the first material in order to form a continuous substantially impermeable outer wall of the hollow tube;(b) selectively removing a portion of sa... | - | A61B005/145 B29C035/08
US20100286490 | A1 | 2010-11-11 | INTERACTIVE PATIENT MONITORING SYSTEM USING SPEECH RECOGNITION | 2007-04-20 | Systems, methods and techniques are described for monitoring a subject. The subject's safety, health and wellbeing can be monitoring using a system that receives input indicating the subject's status. The system can verbally interact with the subject to obtain information on the subject's status. The words used by the subject or the quality of the subject's response can be used to decide whether to contact emergency services to assist the subject. | 1-229. (canceled) 230. A system for monitoring a subject, comprising: a microphone;a speaker;a speech recognition system;a speech synthesizer;a quality of responsiveness determiner; anda processor configured to: initiate computer generated verbal interaction with the subject, including selecting speech for the speech synthesizer to synthesize and cause to be generated by the speaker, wherein the verbal interaction with the subject is selected to elicit a verbal response from the subject;receive digitized sound representing verbal responses from the subject and captured by the microphone;cause the speech recognition system to perform speech recognition on the digitized sound;cause the quality of responsiveness determiner to determine a quality of responsiveness of the subject to the synthesized speech; anddetermine whether to contact a predetermined contact for the subject after the quality of the responsiveness is determined. 231. The system of claim 230, wherein determining whether to contact a predetermined contact for the subject includes basing the determination on the quality of the responsiveness. 232. The system of claim 231, wherein the quality of responsiveness is one of delayed, valid or invalid. 233. (canceled) 234. The system of claim 232, wherein when a valid response is determined, the processor is further configured to: determine whether the valid response indicates that the subject is experiencing a health problem; andif the subject is experiencing a health problem, cause emergency services to be contacted. 235. (canceled) 236. The system of claim 232, wherein the processor is configured to determine to contact a predetermined contact when the quality of responsiveness is delayed or invalid. 237. The system of claim 230, wherein initiating the computer generated verbal interaction initiates a script of questions related to detecting a heart attack or a stroke. 238. The system of claim 230, wherein the processor is configured to: after determining with a computer the quality of the responsiveness, generate additional synthesized speech to elicit a further verbal response from the subject, wherein the additional synthesized speech poses a question to the subject regarding a safety or health status of the subject;receive a response to the question regarding the safety or health status of subject; andperform speech recognition on the response to generate corresponding subsequent text, wherein determining whether to contact a predetermined contact is based on the subsequent text. 239-252. (canceled) 253. A personal emergency monitoring system, comprising: a) a speaker that outputs audio signal, including computer generated speech, to a subjectb) a microphone that receives sounds, including speech of the subject, including speech related to responses to questions and statements, and speech related to keywords and keyword phrases initiated by the subject, and converts the sounds into an input audio signal; andc) a control unit... | IQ LIFE, INC. | A61B005/00 G10L015/00 G08B023/00
US20100298667 | A1 | 2010-11-25 | LIVING BODY INFORMATION DETECTION APPARATUS AND BLOOD-PRESSURE METER | 2010-07-30 | A living body information detection circuit including a light-emitting element for irradiating a part of a living body with irradiating light. The living body information detection circuit further including a light-receiving element for receiving scattered light of the irradiating light scattered in the part of the living body to detect a pulse waveform. The living body information detection circuit also including a light shielding structure for limiting an angle of light entering the light-receiving element in front of the light-receiving element. | 1. A living body information detection circuit comprising: a light-emitting element for irradiating a part of a living body with irradiating light;a light-receiving element for receiving scattered light of the irradiating light scattered in the part of the living body to detect a pulse waveform; anda light shielding structure for limiting an angle of light entering the light-receiving element in front of the light-receiving element. 2. The living body information detection circuit as claimed in claim 1, wherein the light shielding structure is a hood provided in front of the light-receiving element. 3. The living body information detection circuit as claimed in claim 1, wherein the light shielding structure is a light shielding structure including an aperture provided in front of the light-receiving element. 4. A living body information detection circuit comprising: a light-emitting element for irradiating a part of a living body with irradiating light;a light-receiving element for receiving scattered light of the irradiating light scattered in the part of the living body so as to detect a pulse waveform; anda lens for concentrating scattered light, in the scattered light, from a particular position of the living body onto a light-receiving surface of the light-receiving element. 5. A living body information detection circuit comprising: a light-emitting element for irradiating a part of the living body with irradiating light;a light-receiving element for receiving scattered light of the irradiating light scattered in the part of the living body so as to detect a pulse waveform; anda light shielding structure, provided in front of the light-emitting element, for limiting an angle of outgoing light from the light-emitting element. 6. The living body information detection circuit as claimed in claim 5, wherein the light shielding structure is a hood provided in front of the light-emitting element. 7. The living body information detection circuit as claimed in claim 5, wherein the light shielding structure is a the light shielding structure including an aperture provided in front of the light-emitting element. 8. A living body information detection circuit comprising: a light-emitting element for irradiating a part of a living body with irradiating light;a light-receiving element for receiving scattered light of the irradiating light scattered in the part of the living body so as to detect a pulse waveform; anda lens, in front of the light-receiving element, for concentrating outgoing light from the light-emitting element onto a particular position. 9. A living body information detection circuit comprising: an edge emitting laser for irradiating a part of a living body with irradiating light; anda light-receiving element for receiving scattered light of the irradiating light scattered in the part of the living body so as to detect a pulse waveform. 10. A living body information detection circuit comprising: a vertical cavity surface emittin... | NIPPON TELEGRAPH AND TELEPHONE CORP. | A61B005/00
US07848789 | B2 | 2010-12-07 | Hybrid magnetic-base and impedance-based position sensing | 2009-04-17 | A position sensing system includes a probe adapted to be introduced into a body cavity of a subject. The lasso-shaped probe includes a magnetic field transducer in the rigid section proximal to the flexible lasso-shaped distal section which has at least one probe electrode. A control unit is configured to measure position coordinates of the probe using the magnetic field transducer. The control unit also measures an impedance between the at least one probe electrodes and one or more points on a body surface of the subject. Using the measured position coordinates, the control unit calibrates the measured impedance and provides an indication of the deformation and deflection of the flexible lasso-shaped distal section in relation to the magnetic field transducer in the rigid proximal section. | 1. A position sensing system, comprising: a probe having a substantially rigid section and a flexible lasso-shaped distal section adapted to be introduced into a body cavity of a subject wherein the probe comprises a magnetic field transducer in said rigid section and at least one probe electrode on the flexible lasso-shaped distal section; and,a control unit configured to measure position coordinates of the probe using the magnetic field transducer, and to measure an impedance between the at least one probe electrode and one or more points on a body surface of the subject, and to calibrate the measured impedance using the measured position coordinates. 2. The system according to claim 1, wherein the magnetic field transducer comprises one or more coils. 3. The system according to claim 1, and comprising an external magnetic field generator adapted to generate an external magnetic field, which causes the magnetic field transducer to generate a position signal, wherein the control unit is configured to receive and process the position signal so as to measure the position coordinates of the probe. 4. The system according to claim 1, wherein the magnetic field transducer is adapted to generate a magnetic field, and wherein the system further comprises an external magnetic field sensor adapted, responsively to the magnetic field, to generate a position signal, and wherein the control unit is configured to receive and process the position signal so as to measure the position coordinates of the probe. 5. The system according to claim 1, wherein the control unit is configured to drive an electrical current between the at least one probe electrode and body surface electrodes at the one or more points, and to measure the impedance responsively to the current. 6. The system according to claim 5, wherein the control unit is configured to maintain a constant voltage between the at least one probe electrode and the body surface electrodes, and to measure the current at the constant voltage. 7. The system according to claim 5, wherein the control unit is configured to maintain a constant current between the at least one probe electrode and the body surface electrodes, and to measure the voltage between the at least one probe electrode and the body surface electrodes. 8. The system according to claim 1, wherein the control unit is configured to apply a voltage across at least one pair of body surface electrodes at the one or more points and to measure the impedance by sensing a voltage drop at the at least one probe electrode. 9. The system according to claim 1, wherein the control unit is configured to determine the position coordinates using the magnetic transducer at a plurality of locations in the body cavity and to determine the impedance at the plurality of the locations, and to generate a calibration map determining the impedance at the plurality of the locations. 10. The system according to claim 9, wherein the control unit is configured to cal... | Biosense Webster, Inc. | A61B005/05
US07841993 | B2 | 2010-11-30 | Lancet-integrated sensor, measuring device for lancet-integrated sensor, and cartridge | 2008-11-07 | A sensor and a lancet are integrated with each other. A thin strip-shaped sensor and a lancet are integrated so that the lancet moves in parallel, along a longitudinal direction of the sensor. A measuring device to which an integrated lancet and sensor is attached is provided with a function of driving the attached lancet. | 1. A biosensor cartridge for housing a plurality of biosensors, each biosensor of the plurality of biosensors having a reagent layer for detecting a specific component in a bodily fluid of a subject and an electrode for outputting an electric signal indicating that the reagent layer detects the specific component, said biosensor cartridge comprising: a cartridge body having a plurality of slits that are each shaped so as to conform to a front end of the biosensor, a rear end of the biosensor extending from the front end in a direction of insertion of the biosensor into a measuring device;wherein said plurality of slits are formed in said cartridge body so as to be parallel with each other such that the plurality of biosensors will be arranged in said slits so as to be parallel to each other and so as to extend perpendicularly from said cartridge body upward in the direction of insertion into the measuring device;wherein said plurality of slits are spaced at intervals such that an insertion slot of the measuring device for a target biosensor in one of said plurality of slits can cover the target biosensor and capture and pick up the target biosensor without touching biosensors in said plurality of slits adjacent to the target biosensor;wherein said cartridge body has a lid portion that covers an upper surface of said cartridge body;wherein said cartridge body has a hinge portion supporting a shaft of said lid portion so that said lid portion is rotatable at one end of said upper surface of said cartridge body;wherein said cartridge body is a rectangular parallelepiped in shape;wherein a pair of bearings that project outwardly in a longitudinal direction of said cartridge body from an end of said upper surface of said cartridge body, each of said bearings having an elongate bearing hole;wherein a pair of axial projections at an end of said lid portion engaged with said pair of bearings in said bearing hole thereof, said axial projections serving as fulcra upon rotation of said lid portion relative to said cartridge body;wherein a peripheral edge formed at the periphery of said upper surface of said cartridge body, said peripheral edge extending in a direction that is perpendicular to side walls of said cartridge body, and said peripheral edge having a plurality of notches therein;wherein reverse L-shaped projections that are formed on the periphery of an opening of said lid portion in the same plane as side walls of said lid portion and so as to be in positions that correspond to said plurality of notches, respectively, when said lid portion is rotated to cover said upper surface of said cartridge body; andwherein an elastic sealing member on an upper surface of said peripheral edge for sealing with said lid portion. 2. The biosensor cartridge of claim 1, wherein said lid is a hollow rectangular parallelepiped in shape and has an opening at a portion thereof that is opposed to said upper surface of said cartridge body. 3. The biosensor cartridge ... | Panasonic Corporation | A61B005/00 A61B005/05 B65D081/00
US07848891 | B2 | 2010-12-07 | Modulation ratio determination with accommodation of uncertainty | 2006-09-29 | Embodiments of the present invention relate to a system and method for detecting a blood characteristic in a patient. Embodiments of the present invention may comprise detecting a first modulating signal at a first wavelength, detecting a second modulating signal at a second wavelength, and determining a relative amplitude of the first and second modulating signals. Further, embodiments of the present invention may comprise regressing the first and second modulating signals relative to one another, wherein a first uncertainty value in the first modulating signal and a second uncertainty value in the second modulating signal are accommodated. | 1. A method of detecting a blood characteristic in a patient, comprising: identifying a plurality of first modulating signal values with a processor based on detecting a first modulating signal at a first wavelength with a detector over a time period;identifying a plurality of second modulating signal values with the processor based on detecting a second modulating signal at a second wavelength with the detector over the time period; anddetermining a relative amplitude of the first and second modulating signal values with the processor comprising: weighting the first and second modulating signal values with non-zero weighting factors such that more weighting is given to values identified later in the time period and less weighting is given to values identified earlier in the time period; andregressing the first and second modulating signal values relative to one another, wherein a first uncertainty value in the first modulating signal and a second uncertainty value in the second modulating signal are accommodated and the non-zero weighting factors are incorporated. 2. The method of claim 1, comprising determining a best-fit relationship based on the regressing of the first and second modulating signal values. 3. The method of claim 2, comprising calculating a blood oxygen level, a carboxyhemoglobin level, a methemoglobin level, a total hemoglobin, a bilirubin level, a glucose level, a pH, or CO2 level based on the best-fit relationship. 4. The method of claim 1, wherein the first and second modulating signals comprise photoplethysmographic signals. 5. The method of claim 1, wherein the first uncertainty value is non-zero and the second uncertainty value is non-zero. 6. The method of claim 1, wherein the first and second uncertainty values are not equal. 7. The method of claim 1, wherein regressing the first and second modulating signal values includes performing an orthogonal fit. 8. The method of claim 1, wherein regressing the first and second modulating signals includes performing a Weighted Principal Component regression. 9. The method of claim 1, wherein regressing the first and second modulating signal values includes determining mean values of data points from the first and second modulating signals. 10. The method of claim 9, comprising defining a mean data point from the mean values of the data points and defining vectors from the mean data point to each of the data points. 11. The method of claim 1, comprising calculating a blood oxygen level based on the relative amplitude. 12. A system for detecting a blood characteristic in a patient, comprising: a detector configured to detect a first modulating signal at a first wavelength and a second modulating signal at a second wavelength after dispersion of the first modulating signal and the second modulating signal by tissue; anda monitor configured to: determine a plurality of first modulating signal values over a time period based on detection of the first modulating signal;dete... | Nellcor Puritan Bennett LLC | G06F019/00 G01N033/50 A61B005/0295
US07824347 | B2 | 2010-11-02 | System for marking a location for treatment within the gastrointestinal tract | 2003-05-01 | A device and method for mapping, diagnosing and treating the intestinal tract is provided using a capsule passing through the intestinal tract. Further, a capsule tracking system is provided for tracking a capsule's location along the length of an intestinal tract as various treatment and/or sensing modalities are employed. In one variation, an acoustic signal is used to determine the location of the capsule. A map of sensed information may be derived from the pass of a capsule. Capsules may be subsequently passed through to treat the intestinal tract at a determined location along its length. One variation uses an electrical stimulation capsule to treat and/or diagnose a condition in the intestinal tract. | 1. A system for marking a location for treatment within a gastrointestinal tract comprising: a capsule sized to pass through an intestinal tract of a patient;an acoustic transmitter coupled to the capsule, the acoustic transmitter configured to transmit an acoustic tracking signal between the capsule and a location external to a patient's body, to track a three dimensional location of the capsule within the patient's body as the capsule is passing through the intestinal tract;a reference signal generator coupled to the capsule, the reference signal generator configured to generate and transmit a reference signal for identifying a time of tracking signal origination; andan actuation mechanism comprising a controller configured to actuate a pump within the capsule for marking of the three dimensional location as the capsule is passing through the intestinal tract at a desired treatment location in the intestinal tract based on the capsule location. 2. The capsule of claim 1 wherein the capsule further comprises a container containing a material for marking the location, and wherein the actuation mechanism comprises a release configured to release the material. 3. The system of claim 2 wherein the material is selected from a group comprising a dye, a radiopaque marker, a radioactive material, and a fluorescing compound. 4. The system of claim 1 wherein the capsule further comprises an immobilization mechanism configured to immobilize the capsule at the location, and wherein the actuating mechanism is configured to actuate the immobilizing mechanism to immobilize the capsule at the location. 5. The system of claim 4 wherein the immobilizing mechanism comprises a wall engaging member. 6. The system of claim 5 wherein the wall engaging member comprises an expandable member. 7. The system of claim 5 wherein the wall engaging member comprises an anchor. 8. The system of claim 1 wherein the capsule further comprises an identification mechanism actuable by the actuating mechanism. 9. The system of claim 8 wherein the identification mechanism is selected from a group comprising: a dye, a radiopaque marker, a radioactive material, a fluorescing compound, and a detectable energy emitter. 10. The system of claim 1 wherein the capsule further comprises a telemetry circuit configured to receive a telemetry signal to instruct the actuating mechanism to operate to mark the location. 11. The system of claim 1 further comprising a plurality of acoustic transducers configured to receive the tracking signal. 12. The system of claim 11 wherein the plurality of acoustic transducers are arranged in a cartesian coordinate system. 13. The system of claim 11 wherein the plurality of acoustic transducers are configured to receive and transmit signals between each other. 14. The system of claim 1 further comprising: a sensor coupled to the capsule, configured to sense a parameter corresponding to the location to be marked for treatment. 15. The system of claim 1... | Entrack, Inc. | A61B005/103 A61B005/117 A61B007/00 A61B005/00 A61B005/04 A61B005/05 A61B006/00 A61B008/00 A61B008/14 A61B001/00 A61B001/04 H04R025/00 A61H001/00 A61H001/02 A61H005/00
US07835780 | B1 | 2010-11-16 | MR invasive device and method for active MR guidance of invasive devices with target navigation | 2001-03-30 | An invasive device having an inductive coupling element. One embodiment of the invasive device includes a plurality of receive coils inductively coupled to a communicating coil. The receive coils are selectively tuned and detuned to receive MR signals for providing coordinate information used for device tracking. A second embodiment of the invasive device includes a receive coil having a plurality of winding elements separated from each other by different distances. A method of rapidly acquiring both the invasive device orientation and position information to dynamically adapt MR scan planes to continuously follow the invasive device relative to a target is provided. The target-navigation technique automatically defines the MR scan plane and a time domain multiplexing technique is applied for MR imaging and device tracking. Using these techniques, the acquired MR images shows both the invasive device and the target tissue. | 1. A catheter configured for use in a vasculature of a patient in an MR system, comprising: a first coil having a signal generating winding element for generating a first coil signal induced by the MR system, the first coil being arranged inside the catheter; anda second coil having a communicating winding element inductively coupled with the signal generating winding element for providing a coil signal representative of the first coil signal, the second coil being arranged inside the catheter;where the first coil includes a second signal generating winding element and a third signal generating winding element, where the first signal generating winding element and the second signal generating winding element are ie separated from the third signal generating winding element by a different distance,where the first coil and second coil are configured to provide an MR tracking signal. 2. The catheter of claim 1, where the first signal generating winding element and the second signal generating winding element form a first opposed solenoid and the first signal winding element and the third signal generating winding element form a second opposed solenoid. 3. The catheter of claim 1, comprising a flexible mounting structure on which the first coil and the second coil are secured, the winding elements being collinear with respect to each other when the flexible mounting structure is in a first position, and the winding elements not being collinear with respect to each other when the flexible mounting structure is in a second position. 4. The catheter of claim 1, where at least one of the winding elements is a loop. 5. The catheter of claim 1, where at least one of the winding elements is a saddle. 6. The catheter of claim 1, where the first and second coils are micro coils. 7. The catheter of claim 1, where the device includes a device tip, and where the first and second coils are located within about 5 mm of the tip. 8. The catheter of claim 1, comprising: a third coil including a signal generating winding element for generating a second coil signal induced by the MR system, the second coil having a second communicating winding element inductively coupled with the signal generating winding element of the third coil for providing a signal representative of the second coil signal. 9. The catheter of claim 8, where the communicating winding elements form an opposed solenoid. 10. A catheter configured for use in a vasculature of a patient in an MR system, comprising: a first coil including first, second, and third receiving winding elements for generating respective coil signals induced by the MR system, the first receiving winding element being disposed between the second receiving winding element and the third receiving winding elements and the second receiving winding element and the third receiving winding element being separated from the first receiving winding element by different distances; anda communicating coil disposed laterally adjacent... | - | A61B005/05
US07848794 | B2 | 2010-12-07 | Device for detecting electrical potentials of the forehead region of a patent | 2005-12-19 | A device for detecting electrical potentials in the forehead area of a patient has a transverse member adapted to rest against the forehead of the patient, a vertical member extending crosswise to the transverse member and a mask on the vertical member. Electrodes are provided on the transverse members at least at its ends and at the center at which the vertical bar crosses the transverse member. | 1. An apparatus for detecting electrical potentials in the forehead region of a patient, in particular for determining sleep stages, comprising: a breathing mask adapted to fit over at least a nose of the patient; andan electrode device provided to the breathing mask and having at least two electrode elements directly engagable with the forehead region for determining sleep stages,a measuring circuit arrangement for producing measurement data in accordance with the electrical potentials detected by the electrode device, the measuring circuit arrangement being integrated into a forehead support element provided to said breathing mask such that by virtue of fitting of the mask to the nose, the electrode elements are accurately and repeatedly aligned in a correct position relative to the patient's forehead region in use, anda signal transmission device for transmission of the measurement data produced by the measuring circuit arrangement. 2. The apparatus defined in claim 1 wherein the measuring circuit arrangement has a data compression device for forwarding a compressed data set to the signal transmission device. 3. The apparatus defined in claim 1, further comprising a breathing gas duct connected to the breathing mask, wherein the forehead support element is connected to the breathing gas duct. 4. The apparatus defined in claim 3, wherein the breathing gas duct extends through the forehead support element. 5. The apparatus defined in claim 1, wherein the measuring circuit arrangement is provided in a recess of the forehead support element and contacts of the measuring circuit arrangement contact co-operating contacts in the recess. 6. The apparatus defined in claim 1, comprising three electrode elements, wherein two electrode elements are provided at respective ends of the forehead support element and the third electrode element is provided in a middle portion of the forehead support element. 7. The apparatus defined in claim 1, wherein the measuring circuit arrangement comprises a clock and is configured to determine a time at which measurement data is produced. 8. The apparatus defined in claim 1, wherein the measuring circuit arrangement is provided in a band which is fitted in the forehead support element. 9. The apparatus defined in claim 8, wherein the at least two electrode elements are provided in the band. 10. The device defined in claim 1, wherein the at least two electrode elements are elastically mounted to the forehead support element. 11. A device for detecting electrical potentials in the forehead region of a patient, in particular for determining sleep stages, comprising: a breathing mask adapted to communicate with the upper airways of the patient in use; andan electrode device provided to the breathing mask and having at least two electrode elements directly engagable with the forehead region for determining sleep stages,a measuring circuit arrangement for producing measurement data in accordance with the electrical p... | MAP Medizin-Technologie GmbH | A61B005/04 A62B018/02
US07844316 | B1 | 2010-11-30 | EKG cable | 2006-10-23 | The present invention relates to an EKG cable for connection to an EKG machine. The cable has a root cable with a connector for connection to the EKG machine. The root cable is connected to a cable connector. The cable connector has at least two branch cables connected thereon. Each of the branch cables has a plurality of nodes disposed thereon. Each of the nodes has two connections each being capable of being connected to an electrode disc. Each of the nodes have opposing upper and lower surface. Each of the surfaces may have indicia disposed thereon. Each of the surfaces may be color-coded. The indicia and the color-coding may be used independently of each other or in combination with each other. | 1. An EKG cable for connection to an EKG machine so that the machine can take the EKG of a patient, the EKG cable comprising: a root cable with distal and proximal ends, the proximal end of the root cable being joined to a cable connector, the distal end of the root cable being joined to a connector for connection to an EKG machine, two branch cables being connected to said cable connector, wherein a first of said branch cables has two nodes disposed thereon, a proximal node which is disposed on said first cable proximal the cable connector and a distal node which is disposed on said first cable on or adjacent a distal end, wherein a second of said branch cables has three nodes disposed thereon, wherein each of said nodes has an upper surface and a lower surface disposed opposite the upper surface, wherein each of the surfaces of all of the nodes has a connection for receiving an electrode disc, wherein the upper surface of a first of one of said nodes has indicia for indicating connection to an electrode disc selected from the group consisting of RA, RL, LA, and LL; and said lower surface of the first one of said nodes has indicia which indicates connection to an electrode disc which is different from the indicia on the upper surface of the first one of said nodes. 2. The EKG cable of claim 1, wherein a second one of said nodes has indicia for indicating connection to an electrode disc selected from the group V, a heart symbol, and a picture of a heart. 3. The EKG cable of claim 1, wherein the upper surface of a second one of said nodes has indicia for indicating connection to an electrode disc selected from the group consisting of RA, RL LA, and LL; and the indicia on the upper surface of the first node is different from the indicia on the upper surface of the second node. 4. The EKG cable of claim 3, wherein the upper surface of a third of one of said nodes has indicia for indicating connection to an electrode disc selected from the group consisting of RA, RL LA, and LL; and the indicia on the upper surface of the three nodes are different from each other. 5. The EKG cable of claim 4, wherein the upper surface of a fourth one of said nodes has indicia for indicating connection to an electrode disc selected from the group consisting of RA, RL LA, and LL; and the indicia on the upper surface of the four nodes are different from each other. 6. The EKG cable of claim 5, wherein the lower surface of said fourth node has indicia which is different from the indicia on said upper surface of said fourth node. 7. The EKG cable of claim 4, wherein the lower surface of said third node has indicia which is different from the indicia on said upper surface of said third node. 8. The EKG cable of claim 3, wherein the lower surface of said second node has indicia which is different from the indicia on said upper surface of said second node. 9. An EKG cable which can be deployed on a patient in both and upward and a downward direction, the cable compris... | - | A61B005/0402
US07841982 | B2 | 2010-11-30 | Apparatus and method for imaging objects with wavefields | 2005-09-09 | A transmission wave field imaging method, comprising the transmission of an incident wave field into an object, the incident wave field propagating into the object and, at least, partially scattering. Also includes the measuring of a wave field transmitted, at least in part, through an object to obtain a measured wave field, the measured wave field based, in part, on the incident wave field and the object. Additionally, the processing of the measured wave field utilizing a parabolic approximation reconstruction algorithm to generate an image data set representing at least one image of the object. | 1. A transmission wave field imaging method, comprising: transmitting an incident wave field into an object, the incident wave field propagating into the object and, at least, partially scattering;measuring a wave field transmitted, at least in part, through the object to obtain a measured wave field, the measured wave field being defined, in part, by the incident wave field and the object; andprocessing the measured wave field utilizing a parabolic approximation reconstruction algorithm to generate an image data set representing at least one image of the object. 2. The method of claim 1, wherein the parabolic approximation reconstruction algorithm calculates a predicted total wave field utilizing a single independent variable, the single independent variable representing scattering potential. 3. The method of claim 1, wherein the parabolic approximation reconstruction algorithm utilizes a parabolic binomial approximation. 4. The method of claim 1, wherein the object is tissue, the incident wave field is an acoustic wave field and the parabolic approximation reconstruction algorithm models acoustic wave propagation through the tissue. 5. The method of claiml, wherein the parabolic approximation reconstruction algorithm models forward wave propagation of the incident wave field. 6. The method of claim 1, wherein the object is a breast scanner and the processing generates the image of a breast during a breast scan. 7. The method of claim 1, wherein the image represents quantitative estimates of a speed-of-sound associated with the object. 8. The method of claim 1, wherein the image represents quantitative estimates of an attenuation associated with the object. 9. The method of claim 1, wherein the incident wave field represents an ultrasound wave field and the method further comprises displaying the at least one image as at least two images co-displayed, the at least two images including a speed-of-sound image and an attenuation image. 10. The method of claim 1, wherein the incident wave field represents an ultrasound wave field and the method further comprises displaying the at least one transmissive image as at least two images co-displayed, the at least two images including a speed-of-sound image and an attenuation image. 11. The method of claim 1, wherein the incident wave field comprises at least one of acoustic energy, elastic energy, electromagnetic energy and microwave energy. 12. The method of claim 1, wherein the transmitting and measuring are performed by a transmitter and receiver, respectively, the method further comprising moving at least one of the transmitter and receiver with respect to the object. 13. The method of claim 1, further comprising repeating the transmitting and receiving at multiple slices within the object to generate a 3D volume of data. 14. The method of claim 1, further comprising displaying 2-D images representative of slices through the object. 15. The method of claim 1, further comprising displayin... | Techniscan, Inc. | A61B005/00 G01N023/00
US07857772 | B2 | 2010-12-28 | Stride-monitoring device | 2004-10-07 | A stride monitoring device that can be used particularly for sports applications and that includes a pair of shoes, one of which includes at least one magnetic mass, and the other includes at least a measurement device to make at least one measurement, and a processor for processing of the measurement. The measurement device includes at least one accelerometer and at least one magnetometer capable of outputting signals that can be processed to determine stride parameters. | 1. A stride monitoring device, comprising: a first shoe including at least one magnetic mass;a second shoe including at least one magnetometer configured to measure a magnetic field produced by the magnetic mass in the first shoe and to output magnetic field signals based on the measured magnetic field produced by the magnetic mass in the first shoe, wherein said magnetic field signals can be processed to determine stride parameters and a distance between the shoes, andsaid second shoe further includes at least one accelerometer configured to measure an acceleration and to output acceleration signals based on the measured acceleration, and the accelerometer is further configured to output acceleration signals which are analyzed by a processor to determine instants of impact of said second shoe, and wherein the instants of impact are taken into account for calibrating in time a dynamic measurement of the distance between the shoes. 2. A device according to claim 1, wherein each of the first and second shoes includes the at least one magnetic mass, measurement means for making at least one physical measurement, and electronic means for processing the physical measurement, the measurement means including the at least one accelerometer and the at least one magnetometer configured to output signals that can be processed to determine the stride parameters. 3. A device according to claim 1, wherein the magnetic mass includes at least one permanent magnet. 4. A device according to claim 1, wherein the second shoe includes a plurality of accelerometers. 5. A device according to claim 1, wherein the second shoe includes a plurality of magnetometers. 6. A device according to claim 1, wherein the second shoe comprises said at least one accelerometer and electronic means for processing said magnetic field signals and said acceleration signals, wherein said electronic means comprises means for transmitting a signal output by the electronic means. 7. A device according to claim 6, further comprising portable means for receiving the signal transmitted by the means for transmitting and for displaying data representative of the signal. 8. A device according to claim 7, wherein the portable means comprises: means for receiving data;electronic data processing means for processing data, the electronic data processing means including a memory;means for controlling inputs; andmeans for displaying. 9. A device according to claim 8, wherein the memory includes: a calibration unit configured to calibrate the signal transmitted by the transmission means, as a function of stride length and magnetic characteristics of the shoes,a stride length estimating algorithm,an algorithm to calibrate the signal transmitted by the transmission means as a function of the parameters input by a user, andan algorithm to estimate the stride speed. 10. A device according to claim 9, wherein the calibration unit is configured to determine a mathematical calibration law by a polynomial... | Commissariat a l'Energie Atomique | A61B005/103 A61B005/117 A43D001/00 G01C019/36 G01C019/00 G01C017/18 G01C017/00 G01C017/04 G01C017/30
US07831299 | B2 | 2010-11-09 | ECG rhythm advisory method | 2006-05-09 | A method of automatically determining which type of treatment is most appropriate for a cardiac arrest victim, the method comprising transforming one or more time domain electrocardiogram (ECG) signals into a frequency domain representation comprising a plurality of discrete frequency bands, combining the discrete frequency bands into a plurality of analysis bands, wherein there are fewer analysis bands than discrete frequency bands, determining the content of the analysis bands, and determining the type of treatment based on the content of the analysis bands. | 1. A method of automatically determining which type of treatment is most appropriate for a patient, the method comprising: transforming a time domain electrocardiogram (ECG) signal from the patient into frequency domain data;processing the frequency domain data to determine the variance of the energy distributed across a selected frequency range relative to a reference frequency of the spectrum of the ECG signal; anddetermining the type of treatment based on the determined variance. 2. The method of claim 1 wherein the reference frequency comprises one of the following: mean frequency, median frequency, center frequency, peak frequency. 3. The method of claim 1 further comprising providing a prompt to a user of the result of the determination of which type of treatment is most appropriate. 4. The method of claim 1 wherein determining the type of treatment based on the determined variance comprises comparing the determined variance to a threshold. 5. The method of claim 1 wherein determining the type of treatment based on the determined variance comprises combining the determined variance with one or more of the following in a multidimensional space: a number of peaks in the selected frequency range;relative values of multiple peaks in the selected frequency range; andrelative bandwidths of multiple peaks in the selected frequency range. 6. The method of claim 1 wherein the selected frequency range excludes a low-frequency portion in which a chest compression artifact dominates. 7. The method of claim 1 wherein variance is determined by computing the weighted distance of each frequency component from the reference frequency. 8. The method of claim 7 wherein the weighted distance of a given frequency component from the reference frequency comprises a squared difference between the given frequency component and the reference frequency multiplied by a magnitude of the frequency domain data at the given frequency component. 9. The method of claim 7 wherein the variance is determined by integrating the computed weighted distances over the selected frequency range. | ZOLL Medical Corporation | A61B005/04
US08630694 | B2 | 2014-01-14 | Medical imaging system and fluid-driven injector | 2008-08-13 | Some injectors of the invention may include a fluid drive responsive to pressure of a working fluid (e.g., liquid, pneumatic, or both) to impart a sequence of forces to drive a delivery device (e.g., a syringe) to deliver a medical fluid (e.g., a contrast agent, a radiopharmaceutical, a drug, or a combination thereof). Some injectors may include a multimedia tube configured to pass a working fluid (e.g., air) and a light signal (e.g., infrared). Some injectors may include a peristaltic drive responsive to pressure of a working fluid. | 1. A medical system, comprising: a magnetic resonance imaging system comprising an imaging bore;a fluid-driven injector disposed within the imaging bore and comprising a fluid delivery device, a fluid conduit, a working fluid, and a substantially non-ferrous and substantially non-magnetic fluid drive, wherein the fluid drive is configured to not interfere with magnetic resonance imaging;at least one pump;a flowpath extending between the at least one pump and the fluid drive;a working fluid, wherein the at least one pump provides the working fluid to the fluid drive through the flowpath; anda wall with shielding, wherein the at least one pump and the fluid-driven injector are disposed on opposite sides of the wall;wherein the fluid drive is responsive to pressure of the working fluid from the at least one pump to impart a sequence of forces to drive the delivery device to flow medical fluid from the injector, wherein the fluid drive is configured to sequentially squeeze the fluid conduit, wherein the fluid drive comprises a plurality of pistons disposed in series along the fluid conduit, and wherein the pistons are configured to move sequentially to impart the sequence of forces at different times and axial positions along the fluid conduit. 2. The medical system of claim 1, wherein the fluid drive is configured to not emit or detect radio frequency radiation. 3. The medical system of claim 1, further comprising a control coupled to the at least one pump. 4. The medical system of claim 1, wherein the delivery device has an output disposed generally proximate the fluid drive. 5. The medical system of claim 1, wherein the fluid drive further comprises a separate cylinder for each piston, and wherein the medical system further comprises: a working fluid source coupled to the fluid drive, wherein the working fluid source comprises the working fluid, the at least one pump, and a manifold; anda separate working fluid conduit extending between each of the cylinders of the fluid drive and the manifold, wherein the working fluid source provides sequential pulses one after another to the cylinders of the fluid drive to cause movement of the corresponding pistons, and wherein the flowpath comprises each working fluid conduit. 6. The medical system of claim 5, wherein the at least one pump is on an inlet side of the manifold and the working fluid conduits are on an outlet side of the manifold. 7. The medical system of claim 5, wherein the at least one pump is a single pump, and wherein the working fluid source comprises a plurality of valves corresponding to the number of cylinders and corresponding pistons. 8. The medical system of claim 5, wherein the at least one pump comprises multiple pumps, and wherein the working fluid source comprises a plurality of valves corresponding to the number of cylinders and corresponding pistons. | Mallinckrodt LLC | A61B005/055
US08611976 | B2 | 2013-12-17 | Breathable physiological sensor | 2010-05-13 | A sensor includes a sensor pad that allows air and moisture to diffuse from a patient's skin. A light source is disposed on the sensor pad is configured to generate near-infrared light. A light detector disposed on the sensor pad is configured to detect near-infrared light generated by the light source. | 1. A sensor comprising: a sensor pad configured to allow air and moisture to diffuse from a patient's skin, wherein the sensor pad includes a top layer and a bottom layer, at least one of which is formed from a breathable material, the breathable material having cells arranged in an open cell structure, the breathable material further having a thickness relative to the size of the cells configured to laterally diffuse airflow and moisture from the patient's skin at a rate that is equivalent to at least a lateral diffusion rate achieved by a polyurethane open cell foam with a cell size of 0.2-0.5 mm and a thickness of 1-2 mm;a light source disposed on the sensor pad and configured to generate near-infrared light; anda light detector disposed on the sensor pad and configured to detect near-infrared light generated by the light source. 2. A sensor as set forth in claim 1, wherein the sensor pad includes a breathable adhesive disposed between the top layer and the bottom layer. 3. A sensor as set forth in claim 2, wherein the breathable adhesive includes a porous material that allows a minimum amount of air to flow between the top layer and the bottom layer of the sensor pad. 4. A sensor as set forth in claim 1, wherein the breathable material forming at least one of the top layer and the bottom layer includes an open cell foam. 5. A sensor as set forth in claim 4, wherein the breathable material is opaque to near-infrared light. 6. A sensor as set forth in claim 1, wherein the light source and the light detector are disposed on a circuit board that is at least partially disposed between the top layer and the bottom layer. 7. A sensor as set forth in claim 6, wherein the circuit board includes a flexible printed circuit board. 8. A sensor as set forth in claim 6, wherein the circuit board is at least partially coated with a water protective coating. 9. A sensor as set forth in claim 1, wherein the sensor pad is configured to be adhered to a patient's skin. 10. A sensor comprising: a sensor pad having a top layer disposed on a bottom layer, wherein the top layer is adhered to the bottom layer with a breathable adhesive and wherein the top layer and the bottom layer are at least partially formed from a breathable material to allow air and moisture to diffuse from a patient's skin, the breathable material having cells arranged in an open cell structure, the breathable material further having a thickness relative to the size of the cells configured to laterally diffuse airflow and moisture from the patient's skin at a rate that is equivalent to at least a lateral diffusion rate achieved by a poly urethane open cell foam with a cell size of 0.2-0.5 mm and a thickness of 1-2 mm; anda circuit board disposed on the sensor pad, the circuit board having a light source configured to generate near-infrared light and a light detector configured to receive the near-infrared light generated by the light source after the near-infrared light travels through ... | Covidien LP | A61B005/00
US08706219 | B2 | 2014-04-22 | Method and apparatus for monitoring an organ of a patient | 2010-01-28 | An apparatus for determining tissue versus fluid components of an organ include a detector that generates a detector signal based on electrical signals derived from tissue and fluid. The apparatus includes a signal processor in communication with the detector which subtracts in real time a tissue component from the detector signal and produces a fluid volume signal. A method for monitoring a patient's fluid volume of a patient's organ. An apparatus for monitoring a patient's organ. A method for monitoring a patient's organ. A method to piggyback an admittance system onto a AICD/Bi-ventricular Pacemaker for a heart of a patient, in particular a weakened heart having features consistent with congestive heart failure. An apparatus for monitoring an organ, such as a heart, lungs, brain, skeletal muscle, and bladder of a patient which includes a detector which detects the admittance of the organ. The apparatus includes a transmitter in communication with the detector which transmits a wireless signal indicative of the admittance of the organ. A method for monitoring an organ of a patient includes the steps of detecting with a detector the admittance of the organ. There is the step of transmitting with a transmitter in communication with the detector a wireless signal indicative of admittance of the organ. | 1. An implanted device for a heart of a patient comprising: a housing;a first electrode and a second electrode configured to be disposed in the right ventricular septum (RV) of the heart;a lead having a third and fourth electrode that is configured to extend from the housing into the coronary sinus extending onto the lateral epicardial coronary vein of the heart;a drive circuit disposed in the housing to cause the electrodes to generate emitted electrical signals;a transmitter disposed in the housing that transmit received signals from the electrodes after the emitted signals have passed through the heart;a detector that generates a detector signal based on electrical signals derived from tissue and fluid in the heart; anda signal processor in communication with the detector which subtracts in real time a tissue component from the detector signal and produces a fluid volume signal to determine stroke volume of the heart using admittance as a function of electricity path through the tissue. 2. The implanted device as described in claim 1 wherein the electrodes produce a combined signal that has a tissue component and a fluid component. 3. The implanted device as described in claim 2 wherein the electrodes have a varying distance between them and the processor determines heart admittance based on the varying distance between the electrodes. 4. The implanted device of claim 3 wherein the drive circuit causes the electrodes to generate emitted electrical signals at only a single frequency of about 20 kHz. 5. The implanted device as described in claim 4 wherein the fluid is blood, the tissue is myocardium and the processor determines LV, RV, LA, or RA length, area, and/or volume of the heart in 1, 2 or 3 dimensions. 6. The implanted device as described in claim 5 wherein the myocardial component Rm is determined byRm=-Im⁢{Z_}×(1+(ωɛmσm)2)ωɛmσm,where Z is the complex impedance; ω=2πf where is the frequency; ∈m is permittivity of muscle; and σm is the conductivity of muscle. 7. The implanted device as described in claim 6 wherein the fluid is blood and blood component Rb is determined byRb=Re⁢{Z_}-Rm1+(ω⁢⁢Rm⁢Cm)2,where Z is the complex impedance; ω=2πf where f is the frequency; Rm is the resistance of muscle, and Cm is the capacitance of muscle. 8. The implanted device as described in claim 7 wherein the signal processor subtracts in real time the myocardial component from the detector signal and produces the left and right ventricle, and left and right atrial blood volume signal. 9. An implanted device for a heart of a patient comprising: a housing;a first electrode and a second electrode configured to be disposed in the right ventricular septum (RV) of the heart;a lead having a third and fourth electrode that is configured to extend from the housing into the coronary sinus extending onto the lateral epicardial coronary vein of the heart, the first, second, third and fourth electrodes produce a combined signal that has a tissue component and a ... | Board of Regents, The University of Texas System | A61N001/362
US08862221 | B2 | 2014-10-14 | Monitoring mechanical heart properties | 2007-02-21 | In a method and system for monitoring mechanical properties of a heart in a subject, multiple cardiogenic impedance values reflective of the impedance of the heart in connection with a transition from inhalation to exhalation in the subject are determined. Correspondingly, multiple cardiogenic impedance values reflective of the impedance of the heart in connection with a transition from exhalation to inhalation are determined. The impedance values are collectively processed to form a trend parameter. The value determination and processing is performed over several respiratory cycles spaced apart in time to form a plurality of trend parameters over time. The mechanical properties of the heart are monitored by processing these different trend parameters. The data collection and optionally at least a part of the data processing is performed by an implantable medical device. | 1. A method of monitoring mechanical properties of a heart in a subject, comprising the steps of: a) determining a first set of multiple cardiogenic impedance values reflective of the impedance of said heart during a defined period of a heart cycle occurring in connection with a transition from inhalation to exhalation in said subject;b) determining a second set of multiple cardiogenic impedance values reflective of the impedance of said heart during said defined period of a heart cycle occurring in connection with a transition from exhalation to inhalation in said subject;c) collectively processing with a data processor said multiple cardiogenic impedance values of said first set as a function of said second set to determine a relationship between the first and second set and analyzing the relationship between the first and second sets to form a trend parameter representative of respiratory effect on said cardiogenic impedance values;d) repeating said determining steps a), b) and processing step c) to form a plurality of trend parameters over time; ande) monitoring said mechanical properties of said heart by processing said plurality of trend parameters. 2. The method according to claim 1, wherein step a) comprises: applying a first current signal or a first voltage signal to at least a portion of said heart during a first portion of a respiratory cycle;measuring a first resulting voltage signal or a first resulting current signal over at least a portion of said heart during said first portion of said respiratory cycle; andgenerating said first set of multiple cardiogenic impedance values based on said first current signal and said first resulting voltage signal or said first voltage signal and said first resulting current signal; and wherein step b) comprises: applying a second current signal or a second voltage signal to at least a portion of said heart during a second portion of said respiratory cycle; measuring a second resulting voltage signal or a second resulting current signal over at least a portion of said heart during said second portion of said respiratory cycle; and generating said second set of multiple cardiogenic impedance values based on said second current signal and said second resulting voltage signal or said second voltage signal and said second resulting current signal. 3. The method according to claim 1, wherein step c) comprises the steps of: plotting said multiple cardiogenic impedance values of said second set as a function of said multiple cardiogenic impedance values of said first set to obtain multiple plot points; and generating said trend parameter by processing said multiple plot points. 4. The method according to claim 3, wherein said generating step comprises the steps of: fitting an ellipse to said multiple plot points; and generating said trend parameter as a parameter being representative of said ellipse. 5. The method according to claim 3, wherein said generating step comprises the steps of: fitting a st... | St. Jude Medical AB | A61B005/053 A61B005/08 A61N001/365 A61B005/042 A61B005/145
US09492092 | B2 | 2016-11-15 | Method for continuously monitoring a patient using a body-worn device and associated system for alarms/alerts | 2009-05-20 | The invention provides a body-worn monitor that measures a patient's vital signs (e.g. blood pressure, SpO2, heart rate, respiratory rate, and temperature) while simultaneously characterizing their activity state (e.g. resting, walking, convulsing, falling). The body-worn monitor processes this information to minimize corruption of the vital signs by motion-related artifacts. A software framework generates alarms/alerts based on threshold values that are either preset or determined in real time. The framework additionally includes a series of ‘heuristic’ rules that take the patient's activity state and motion into account, and process the vital signs accordingly. These rules, for example, indicate that a walking patient is likely breathing and has a regular heart rate, even if their motion-corrupted vital signs suggest otherwise. | 1. A method for continuously monitoring a patient, comprising the following steps: detecting a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a first sensor comprising a first detector configured to be worn on the patient's body;detecting a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a second sensor comprising a second detector configured to be worn on the patient's body;detecting a first time-dependent motion waveform with a first motion-detecting sensor adapted to be positioned on the patient's wrist, detecting a second time-dependent motion waveform with a second motion-detecting sensor adapted to be positioned on the patient's upper arm, and detecting a third time-dependent motion waveform with a third motion-detecting sensor adapted to be positioned on the patient's chest, wherein each time-dependent motion waveform is indicative of motion of the location on the patient's body to which the corresponding motion detecting sensor is affixed;processing the first and second time-dependent physiological waveforms to determine at least one vital sign from the patient;analyzing at least a portion of the first, second, and third time-dependent motion waveforms with a motion-determining algorithm;determining the patient's activity state by discriminating between resting, moving, sitting, standing, walking, running, falling, lying down, and convulsing states by using results of said analyzing at least a portion of the first, second, and third time-dependent motion waveforms; andgenerating an alarm by processing the patient's activity state and comparing the at least one vital sign to a predetermined alarm criteria corresponding to the activity state. 2. The method of claim 1, wherein the analyzing step further comprises calculating a mathematical transform of at least one of the first, second, and third time-dependent motion waveforms to determine a frequency-dependent motion waveform. 3. The method of claim 2, wherein the analyzing step further comprises analyzing frequencies in the frequency-dependent motion waveform within a band between 0-3 Hz to determine if the patient is walking. 4. The method of claim 3, wherein the analyzing step further comprises analyzing frequencies in the frequency-dependent motion waveform within a band between 0-10 Hz to determine if the patient is convulsing. 5. The method of claim 4, wherein the analyzing step further comprises analyzing frequencies in frequency-dependent motion waveform within a band between 0-15 Hz to determine if the patient is falling. 6. The method of claim 1, wherein the analyzing step further comprises calculating a time-dependent change in at least two of the first, second, and third time-dependent motion waveforms. 7. The method of claim 6, wherein the analyzing step further comprises calculating a standard deviation of at least one of th... | SOTERA WIRELESS, INC. | A61B005/08 A61B005/0205 A61B005/021 G08B021/04 A61B005/02 A61B005/022 A61B005/0402 A61B005/00 A61B005/01 A61B005/024 A61B005/0245 A61B005/04 A61B005/0404 A61B005/0452 A61B005/0456 A61B005/11 A61B005/145 G06F019/00
US09952296 | B2 | 2018-04-24 | Magnetic resonance safety monitoring systems and methods | 2008-12-22 | A magnetic resonance method comprises applying a radio frequency excitation in an examination region (14), measuring a magnetic resonance signal generated by the applied radio frequency excitation in a subject (16) in the examination region, monitoring a radio frequency parameter during the applying, and evaluating subject safety based on the monitoring. A magnetic resonance safety monitor (40) comprises an analyzer (42, 44, 46, 50) configured to (i) receive a radio frequency signal during magnetic resonance excitation, (ii) extract a radio frequency parameter from the received radio frequency signal, and (iii) evaluate subject safety based on the extracted radio frequency parameter, and a remediation module (54) configured to perform a remediation of the magnetic resonance excitation responsive to the evaluation (iii) indicating a potentially unsafe condition. | 1. A magnetic resonance method comprising: controlling, with an MR controller, radio frequency transmit coil to excite magnetic resonance in a subject in an examination region by applying first radio frequency excitation signals to induce magnetic resonance signals in the subject in the examination region prior to an interventional procedure performed using an interventional instrument, and applying second radio frequency excitation signals to induce magnetic resonance signals in the subject in the examination region during the interventional procedure performed using the interventional instrument;receiving the applied radio frequency excitation signals;evaluating, with a safety monitor, parameters of the received radio frequency excitation signals, wherein the parameters are indicative of electrical coupling with the interventional instrument causing a current to flow in the interventional instrument, wherein a comparison is performed between a parameter acquired from the first radio frequency excitation signals with a parameter acquired from the second radio frequency excitation signals, wherein the comparison is indicative of an overheating condition of the subject in the examination region; andcausing, with the safety monitor, in response to the evaluating of the parameters, the MR controller to automatically at least one of: (1) terminate the applying of the second radio frequency excitation signals or (2) adjust the applying of the second radio frequency excitation signals. 2. The magnetic resonance method as set forth in claim 1, further comprising: disposing the interventional instrument in the examination region; andperforming the interventional procedure using the interventional instrument;wherein the applying of the second radio frequency excitation signals and the measuring of the induced magnetic resonance signals is repeated during the interventional procedure to track movement of the interventional instrument. 3. The magnetic resonance method as set forth in claim 2, wherein the applying of the second radio frequency excitation signals is terminated in response to the evaluating of the parameters indicating coupling with the interventional instrument. 4. The magnetic resonance method as set forth in claim 1, further including: during the applying of the second radio frequency excitation signals, detuning a magnetic resonance receive coil which receives the induced magnetic resonance signal; andreceiving the radio frequency excitation signals using the detuned magnetic resonance receive coil. 5. The magnetic resonance method as set forth in claim 4, wherein the comparison further includes: comparing a strength of the first received radio frequency excitation signals with a strength of the second received radio frequency excitation signals. 6. The magnetic resonance method as set forth in claim 1, wherein the parameters include one or more of a resonance frequency of the radio frequency excitation signals, a phase of the radio f... | KONINKLIJKE PHILIP N.V. | A61B005/055 G08B023/00 G01R033/28
US10973414 | B2 | 2021-04-13 | Vital sign monitoring system featuring 3 accelerometers | 2009-05-20 | The invention provides a system and method for measuring vital signs (e.g. SYS, DIA, SpO2, heart rate, and respiratory rate) and motion (e.g. activity level, posture, degree of motion, and arm height) from a patient. The system features: (i) first and second sensors configured to independently generate time-dependent waveforms indicative of one or more contractile properties of the patient's heart; and (ii) at least three motion-detecting sensors positioned on the forearm, upper arm, and a body location other than the forearm or upper arm of the patient. Each motion-detecting sensor generates at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which it is affixed. A processing component, typically worn on the patient's body and featuring a microprocessor, receives the time-dependent waveforms generated by the different sensors and processes them to determine: (i) a pulse transit time calculated using a time difference between features in two separate time-dependent waveforms, (ii) a blood pressure value calculated from the time difference, and (iii) a motion parameter calculated from at least one motion waveform. | 1. A system for measuring vital signs from a patient, comprising: (a) a first sensor configured to generate a first time-dependent waveform indicative of one or more contractile properties of the patient's heart;(b) a second sensor configured to generate a second time-dependent waveform indicative of one or more contractile properties of the patient's heart;(c) a first motion detecting sensor configured to be worn on the patient's forearm and to generate at least one first time-dependent motion waveform indicative of motion of the patient's forearm, a second motion-detecting sensor configured to be worn on the patient's upper arm and to generate at least one second time-dependent motion waveform indicative of motion of the patient's upper arm, and a third motion-detecting sensor configured to be worn on the patient's body at a location other than the forearm or upper arm of the patient and to generate at least one third time-dependent motion waveform indicative of motion of the location on the patient's body to which the third motion-detecting sensor is affixed;(d) a processing component configured to be worn on the patient's body and comprising a microprocessor, said processing component configured to receive the first time-dependent waveform, the second time-dependent waveform, and the at least one first, second and third time-dependent motion waveforms and determine therefrom: (i) a pulse transit time calculated using a time difference between a time-dependent feature in each of the first and second time-dependent waveforms, (ii) the patient's posture, activity level, and degree of motion calculated using the first, second and third time-dependent motion waveforms, an arm height calculated using first and second time-dependent motion waveforms, and (iii) a blood pressure value calculated using said time difference and the patient's arm height and posture, and an alarm condition determined using the blood pressure value, wherein the alarm condition is modified according to the patient's posture, activity level, and degree of motion; and(e) a wireless communication system configured to be worn on the patient's body and operably connect to the processing component, the wireless system configured to wirelessly transmit the blood pressure value, the patient's alarm condition, and the patient's posture, activity level, and degree of motion determined by the processing component to a remote receiver;wherein the alarm condition indicates a need for medical intervention and is suppressed when the patient's posture, activity level, and degree of motion indicates said patient is ambulatory. 2. A system according to claim 1, wherein: said first sensor is an optical sensor comprising: a source of electromagnetic radiation configured to irradiate tissue of the patient with radiation emitted therefrom, and a detector configured to detect one or more properties of the electromagnetic radiation after it irradiates said tissue, wherein the first time-dependent... | SOTERA WIRELESS, INC. | A61B005/02 A61B005/0205 A61B005/00 A61B005/021 A61B005/022 A61B005/024 A61B005/0285 A61B005/0402 A61B005/0452 A61B005/11 A61B005/0295 A61B005/04 A61B005/0408 A61B005/0428 A61B005/044 A61B005/08 A61B005/1455
US20100331739 | A1 | 2010-12-30 | METHOD AND SYSTEM FOR PREDICTING CALVING | 2008-05-05 | A method and a system for predicting a calving date of an animal comprises monitoring a behavioral pattern of an animal, determining a given change in the behavioral pattern, and predicting the calving date responsive to the given change in the behavioral pattern. | 1. A method for predicting a calving date of an animal, the method comprising: monitoring a behavioral pattern of an animal based on a pattern of leg postures of the animal with a calving sensing device configured and adapted to be attached to a leg of an animal;determining a given change in the behavioral pattern while the calving sensing device is positioned on the leg; andpredicting the calving date responsive to the given change in the behavioral pattern. 2. The method according to claim 1 wherein the behavioral pattern is an activity pattern. 3. (canceled) 4. The method according to claim 2 wherein monitoring the activity pattern of the animal includes monitoring at least one of a laying time of the animal, and a laying frequency of the animal. 5. (canceled) 6. The method according to claim 2 wherein monitoring the activity pattern of the animal includes monitoring at least one of a number of leg movements of the animal, and a number of leg movements per hour of the animal. 7. (canceled) 8. The method according to claim 2 wherein monitoring the activity pattern of the animal includes monitoring a standing time of the animal. 9. The method according to claim 1, wherein monitoring the behavioral pattern of the animal includes monitoring a frequency of passes by a central communication unit. 10. The method according to claim 1, wherein monitoring the behavioral pattern of the animal is performed at a specified time period of a day. 11. The method according to claim 2, wherein the given change in the activity pattern comprises a change in at least one of a laying time of the animal, and a change in a laying frequency of the animal. 12. (canceled) 13. The method according to claim 2, wherein the given change in the activity pattern comprises a change in at least one of a number of leg movements of the animal and a number of leg movements per hour of the animal. 14. (canceled) 15. The method according to claim 2, wherein the given change in the activity pattern comprises a change in a standing time of the animal. 16. The method according to claim 2, wherein the given change in the activity pattern is a change in a ratio of a number of leg movements of the animal and a laying time of the animal. 17. The method according to claim 1, wherein the given change in the behavioral pattern is a change in the behavioral pattern over a day divided by an average of the behavioral pattern over a number of previous days. 18. The method according claim 1, wherein the given change in the behavioral pattern is a change in the behavioral pattern over a day divided by a standard deviation of the behavioral pattern over a number of previous days. 19. The method according claim 1, wherein the given change in the behavioral pattern is a change in the frequency that the animal passes by a central communication unit. 20. The method according to claim 1, wherein the given change is a change over a defined threshold. 21. The method according to claim 1,... | S.A.E Afikim The State of Israel, Ministry of Agriculture & Rur al Development, Agricltural Research Organization | A61B005/103 A01K029/00
US20100286509 | A1 | 2010-11-11 | METHODS AND APPARATUS FOR PERFORMING PROCEDURES ON TARGET LOCATIONS IN THE BODY | 2010-07-21 | A method of performing a medical procedure on target tissue within the body includes immobilizing the target tissue with an immobilization device, positioning an insertion tool (e.g., a needle, a needle and marker wire, a needle gun, etc.) to be inserted into the target tissue on the immobilization device based on at least one reading of the target tissue, and inserting the insertion tool into the target tissue. An immobilization device for performing methods includes a frame assembly, a tissue immobilization assembly coupled to the frame assembly, the tissue immobilization assembly including a structure adapted to immobilize tissue, and a needle assembly releasably connectable to the frame assembly. The needle assembly includes a needle capable of being inserted into the immobilized tissue. | 1. A method of performing a medical procedure on target tissue within the body, the method comprising: immobilizing the target tissue with an immobilization device;positioning an insertion tool to be inserted into the target tissue on the immobilization device based on at least one reading of the target tissue; andinserting the insertion tool into the target tissue. 2. The method of claim 1, further comprising taking at least one reading of a location of target tissue, which comprises taking a first mammography reading. 3. The method of claim 2, wherein taking at least one reading of a location of target tissue comprises taking a second mammography reading. 4. The method of claim 3, further comprising plotting the first and second mammography readings on an image grid, wherein a first distance from a known reference point on the immobilization device is determined from indicia on the image grid to position the needle. 5. The method of claim 4, further comprising correcting for at least one anomaly created by the first and second mammography readings. 6. The method of claim 4, wherein the plotting of the first and second mammography readings on the image grid is generated by a PDA. 7. The method of claim 1, further comprising performing a medical procedure on the target tissue. 8. The method of claim 7, wherein performing a medical procedure is selected from a group comprising one of the following: needle localization; needle aspiration;vacuum-assisted biopsy;virtual guided breast surgery;open surgical biopsy; in-office biopsy;sentinel node biopsy; andimaging correlation and implantation. 9. The method of claim 8, wherein the immobilization device is attached to a unit to perform the medical procedure. 10. The method of claim 1, further comprising placing a marker wire within the needle to mark the target tissue. 11. The method of claim 1, further comprising adjusting the immobilization device to compress tissue surrounding the target tissue. 12. The method of claim 1, further comprising performing at least one diagnostic procedure to obtain the at least one reading of the target tissue. 13. The method of claim 12, wherein the immobilization device is attached to a mammography unit to take the at least one reading. 14. The method of claim 1, further comprising determining a length of a needle to be inserted into the target tissue. 15-27. (canceled) 28. A method for performing a medical procedure on target tissue, the method comprising: immobilizing the target tissue;performing a diagnostic procedure on the immobilized target tissue; andperforming a medical procedure on the immobilized target tissue. 29. The method of claim 28, wherein performing a medical procedure is selected from a group comprising one of the following: needle localization;needle aspiration;vacuum-assisted biopsy;virtual guided breast surgery;open surgical biopsy;in-office biopsy;sentinel node biopsy; andtherapeutic imaging correlation and implantation. 30. The ... | TRAXYZ MEDICAL, INC. | A61B005/05
US20100305458 | A1 | 2010-12-02 | BLOOD PRESSURE METER AND METHOD FOR OPERATING A BLOOD PRESSURE METER | 2008-10-08 | A blood-pressure measuring device includes a resistor having a first end and a second end, a first pressure sensor disposed on the first end and a second pressure sensor disposed on the second side and configured for connection to a blood circuit, wherein the first pressure sensor, the resistor and the second pressure sensor are connected to each other in series. A computing device is connectable to the first pressure sensor and the second pressure sensor and is configured to perform a plausibility check by comparing measurement results of the first pressure sensor and of the second pressure sensor. | 1-19. (canceled) 20. A blood-pressure measuring device comprising: a resistor having a first end and a second end;a first pressure sensor disposed on the first end of the resistor;a second pressure sensor disposed on the second side of the resistor and configured for connection to a blood circuit, wherein the first pressure sensor, the resistor and the second pressure sensor are connected to each other in series;a computing device connectable to the first pressure sensor and the second pressure sensor and configured to perform a plausibility check by comparing measurement results of the first pressure sensor and of the second pressure sensor. 21. The blood-pressure measuring device as recited in claim 20, wherein the plausibility check includes performing at least one of a first check as to whether a first pressure measured by the first pressure sensor is lower than a predetermined maximum pressure, a second check as to whether the first pressure is higher than a predetermined minimum pressure, a third check as to whether a difference between the first pressure and a second pressure measured by the second pressure sensor is more than a predetermined minimum differential pressure, and a fourth check as to whether the difference between the first pressure the second pressure is less than a predetermined maximum differential pressure. 22. The blood-pressure measuring device as recited in claim 21, further comprising an alarm device configured to sound an alarm when the maximum pressure or the maximum differential pressure is exceeded or when the minimum pressure or of the minimum differential pressure is not reached. 23. The blood-pressure measuring device as recited in claim 20, further comprising a display device configured to display the result of the plausibility check. 24. The blood-pressure measuring device as recited in claim 23, wherein the display device is configured to display the results in a first, a second and a third stage, wherein, the first stage includes plausible results, the second stage includes deviating results and the third stage includes implausible results. 25. The blood-pressure measuring device as recited in claim 20, wherein the computing device is configured to calculate a rinsing rate in the resistor taking into account a first measurement result ascertainable by the first pressure sensor and a second measurement result ascertainable by the second pressure sensor. 26. The blood-pressure measuring device as recited in claim 20, wherein the resistor is provided with an encoded version of a flow constant of the resistor. 27. The blood-pressure measuring device as recited in claim 26, wherein the computing device is configured to recognize the code. 28. The blood-pressure measuring device as recited in claim 20, further comprising a joint casing and wherein the resistor, the first pressure sensor and the second pressure sensor are disposed in the joint casing. 29. The blood-pressure measuring device as recited in... | - | A61B005/0215
US20100298692 | A1 | 2010-11-25 | METHOD FOR DETECTING TUMOR CELL INVASION USING SHORT DIFFUSION TIMES | 2008-05-22 | An improved method for detecting tumor cell invasion using diffusion times as short as two (2) msec incorporates diffusion weighing imaging techniques into a standard spin echo (SE) pulse sequence to minimize the effects of compartment boundary restrictions on diffusion values and corresponding MRI imaging data related to glioma invasion. | 1. A method for producing an image of an organ or tissue using a Magnetic Resonance Imaging (MRI) system comprising the steps of: a) acquiring diffusion-weighted images using the MRI system, wherein diffusion-weighted images are acquired, in at least one direction, for a range of diffusion-weightings having at least two b-values using short diffusion times;b) computing image maps of apparent diffusion coefficient (ADC); andc) fitting the data to a stretched exponential model to create maps of an index of tissue heterogeneity (a) and the distributed diffusion coefficient (DDC). 2. The method of claim 1 wherein the short diffusion times of step a) are less than 25 msec. 3. The method of claim 1 wherein step b) is performed by computing the ADC, on a per-voxel basis, from MRI signals collected at a minimum of two different b-values. 4. The method of claim 1 wherein step c) is performed by: a) computing the stretching parameter α, on a per-voxel basis, by fitting to the stretched-exponential model the diffusion-weighted MRI signals as a function of b-value, andb) computing the stretching parameter DDC, on a per-voxel basis, by fitting to the stretched-exponential model the diffusion-weighted MRI signals as a function of b-value. 5. The method of claim 1 in which the pulse sequence is a single or multi-shot spin-echo echo planar imaging (SE-EPI) sequence with one or more pairs of balanced shaped bipolar gradients positioned around the refocusing pulse. 6. The method of claim 1 in which the pulse sequence is a single or multi-shot spin-echo spiral-based sequence with one or more pairs of balanced shaped bipolar gradients positioned around the refocusing pulse. 7. The method of claim 1 in which the diffusion weightings are either oscillating or sinusoidal. 8. The method of claim 1 wherein step a) includes the use of parallel imaging methods to acquire the diffusion-weighted images. 9. The method as recited in claim 1 wherein step a) includes the use of three dimension spiral-based imaging methods to acquire the diffusion-weighted images. 10. A method for producing an image of an organ or tissue using a Magnetic Resonance Imaging (MRI) system comprising the steps of: a) acquiring diffusion-tensor images (DTI) using the MRI system, wherein diffusion-weighted images are acquired in at least 6 different directions for a range of diffusion-weightings having at least two b-values and using short diffusion times;b) computing image maps of apparent diffusion coefficient (ADC), fractional anisotropy (FA), and the diffusion tensor; andc) fitting the data to a stretched exponential model to create maps of an index of tissue heterogeneity, α, and the distributed diffusion coefficient (DDC). 11. The method of claim 10 in which step b) is performed by computing ADC, FA or DTI images from MRI signals collected at a minimum of two different b-values. 12. The method of claim 10 in which step c) is performed by: a) computing the stretching parameter α, on a pe... | IMAGING BIOMETRICS, LLC | A61B005/055
US20100324431 | A1 | 2010-12-23 | Determining Disease State Using An Induced Load | 2009-06-18 | The present disclosure relates to determining a patient's disease state based at least in pail on obtaining or determining certain underlying characteristics, such as vasotone, venous compliance, or ability of the vascular system to drain venous blood, of the patient's vascular system. The characteristics may be obtained by analyzing changes to a patient signal, such as the overall signal change, the rate of change, the shape of the change, changes in signal energy, or changes in the baseline and/or the amplitude of the signal, and/or the time period(s) over which the signal changes, that are caused by inducing a load on the vascular system. In some embodiments, the signal changes may be analyzed by transforming the signal using, for example, a continuous wavelet transform. The patient's health status or disease state may be determined using the one or more vascular system characteristics that influenced the signal change. | 1. A method for communicating a disease state of a patient, the method comprising: obtaining a signal from a patient sensor, wherein the signal changes in response to inducing a load on a vascular system of the patient;analyzing the signal at a first point to obtain a first vascular system characteristic, wherein the first point occurs after the load is induced on the vascular system and wherein the signal is analyzed by a processor coupled to the patient sensor;determining the disease state based at least in part on the first vascular system characteristic; andcommunicating the disease state to an output device. 2. The method of claim 1, wherein the inducing the load comprises changing an elevation of a limb to which the patient sensor is coupled. 3. The method of claim 1, wherein the first vascular system characteristic is obtained based at least in part on a time period between the inducing the load and the first point occurring after the load is induced. 4. The method of claim 1, wherein the first point occurs while the signal is changing in response to inducing the load. 5. The method of claim 1, wherein the signal is a photoplethysmograph signal. 6. The method of claim 5, wherein the first point occurs within a dicrotic notch of a pulse of the photoplethysmograph signal. 7. The method of claim 5, further comprising: transforming using the processor the photoplethysmograph signal into a transformed signal using a continuous wavelet transform;generating using the processor a scalogram based at least in part on the transformed signal;identifying using the processor a band on the scalogram; andobtaining the first vascular system characteristic based at least in part on the band. 8. The method of claim 1, wherein the first vascular system characteristic is one of the group consisting of vasotone, venous compliance, and/or the ability of venous blood to drain, and/or combinations thereof. 9. The method of claim 1, further comprising: analyzing the signal at the first point to obtain a second vascular system characteristic; anddetermining the disease state based at least in part on the first and second vascular system characteristics. 10. The method of claim 1, further comprising: generating an alarm based at least in part on the disease state; andcommunicating the alarm to the output device. 11. A system for communicating a disease state of a patient, the system comprising: an input signal generator for generating a signal, wherein the signal changes based at least in part upon inducing a load on a vascular system of the patient;a processor coupled to the input signal generator, wherein the processor is capable of: analyzing the signal at a first point to obtain a first vascular system characteristic, wherein the first point occurs after the load is induced on the vascular system;determining the disease state based at least in part on the first vascular system characteristic; andan output device coupled to the processor, wherein the out... | Nellcor Puritan Bennett Ireland | A61B005/02
US07856262 | B2 | 2010-12-21 | Volume specific characterization of human skin by electrical immitance | 2003-05-14 | This invention relates to a sensor assembly and a method for measuring characteristics of a surface, preferably skin, comprising a first pair of current supply electrodes coupled to a current source, providing an electrical current to the skin, at least one pickup electrodes at chosen positions relative to the current supply electrodes, at least a first of said pickup electrodes being coupled to an instrument for measuring the voltage between said first pickup electrode and at least one of the pickup or current supply electrodes. | 1. Method for inspecting whether an object consists of living skin, the method comprising the following steps: measuring electrical characteristics of two outer layers of skin, wherein said two outer layers of skin comprise respectively a stratum corneum and living skin,comparing the measured electrical characteristics with corresponding, known characteristics of skin of a living or a dead finger, wherein the measurements of electrical characteristics in the two skin layers is characterized inapplying an alternating current or voltage with at least one frequency to the skin between two current supply electrodes,measuring a complex voltage as a function of frequency between one of said current supply electrodes and a first electrode positioned at a chosen distance from said one current supply electrode for said at least one frequency,measuring a complex voltage between the first electrode and a second electrode, wherein the second electrode is positioned at a larger distance from said one first current supply electrode than said first electrode for said at least one frequency,calculating immittance for the two skin layers based on said voltage measurements, and using said immittance as a basis for comparing with corresponding, known characteristics of skin of a live or a dead finger. 2. Method according to claim 1, wherein said comparison of electrical characteristics comprises comparing dispersion of resistivity for at least one skin layer. 3. Method according to claim 1, wherein said comparing of electrical characteristics comprises comparing dispersion of permittivity for at least one skin layer. 4. Method according to claim 1, wherein said comparing of electrical characteristics comprises comparing resistivity for at least one skin layer. 5. Method according to claim 1, wherein said comparing of electrical characteristics comprises comparing relative permittivity for at least one skin layer. 6. Method according to claim 1 wherein the distance between said one current supply electrode and said first electrode is less than 1 mm. | Idex ASA | A61B005/05
US20100317989 | A1 | 2010-12-16 | Systems and Methods for Performing Neurophysiologic Assesments With Pressure Monitoring | 2010-06-16 | Systems and methods for performing neurophysiologic assessments of neural tissue including nerve pathology monitoring which may or may not be augmented by adding the ability to assess or monitor the pressure being exerted upon a nerve or nerve root before, during and/or after retraction. | 1. A method for assessing the pathology of a nerve, comprising the steps of: retracting a nerve within a patient using a nerve retractor equipped with a pressure sensor;stimulating said nerve with an electrical stimulation signal delivered by at least one stimulation electrode; andusing at least one sensor to detect neuromuscular responses evoked by said stimulation signal. 2. The method of claim 1 wherein said pressure sensor is linked to a processing unit and said processing unit calculates the pressure exerted on said nerve by said nerve retractor. 3. The method of claim 2, wherein said processing unit further includes a display, wherein said display provides at least one of alpha-numeric and graphic indicia relating to said calculated pressure. 4. The method of claim 3, wherein said alpha-numeric and graphic indicia includes at least one of a maximum pressure value, time value, color, 3-dimensional pressure map, 2-dimensional pressure map, chart, and a graph. 5. The method of claim 2, comprising the additional steps of: using a control unit communicatively linked to said stimulation electrode and said sensor to identify a relationship between said stimulation signal and said neuromuscular response. 6. The method of claim 5, wherein said control unit includes a display and said display provides at least one of alpha-numeric and graphic indicia relating to said identified relationship. 7. The method of claim 5, wherein said identified relationship is the lowest stimulation current necessary to evoke a neuromuscular response. 8. The method of claim 7, wherein said control unit is configured to perform a threshold hunting algorithm to identify the lowest stimulation current necessary to evoke a neuromuscular response. 9. The method of claim 1, wherein said stimulation electrode is positioned on said nerve retractor. 10. The method of claim 5, wherein said control unit and said processing unit are a single unit. 11. The method of claim 5, wherein said control unit and said processing unit are distinct units. | NuVasive Inc. | A61B005/05
US07853428 | B2 | 2010-12-14 | Apparatus for displaying fitness exercise condition | 2008-01-11 | An apparatus for displaying a fitness exercise condition is provided. A user (that is, an exerciser) of the apparatus inputs personal basic information on the exerciser. When the apparatus is applied to an arm or wrist of the exerciser and the exerciser carries out fitness exercise, a heart rate is measured during the fitness exercise to record it. Load data indicating hardness of the fitness exercise is calculated on the basis of the personal basic information thus inputted and data on the recorded heart rate. A display section is caused to display notice with a color according to a load on the exerciser for each of predetermined time periods on the basis of the calculated load data. The display section may include first and second display sub sections for displaying first and second notice with first and second colors, respectively. | 1. An apparatus for displaying a fitness exercise condition, the apparatus comprising: a basic information inputting section that inputs personal basic information on an exerciser;a heart rate data acquiring section that acquires data on a heart rate measured during a fitness exercise;a load calculating section that calculates load data indicating hardness of the fitness exercise on the basis of the personal basic information input by the basic information inputting section and the data on the heart rate acquired by the heart rate data acquiring section;a display section; anda display controlling section that causes the display section to display a notice comprising a plurality of color areas each representing a predetermined time interval with a color according to a load on the exerciser for each of the predetermined time intervals on the basis of the load data calculated by the load calculating section,wherein each of the color areas representing the predetermined time intervals that have elapsed remains visible after the respective predetermine time interval has elapsed. 2. The apparatus as claimed in claim 1, wherein the display controlling section causes the display section to display the notice in synchronization with the heart rate. 3. The apparatus as claimed in claim 1, wherein: the notice further comprises a plurality of lines each corresponding to one of the predetermined time intervals,each of the color areas is located at one end or adjacent to the one end of one of the lines, and changes in at least one of a color or a size according to the load on the exerciser, andeach of the color areas corresponding to the predetermined time interval that has elapsed remains visible after the respective predetermined time interval has elapsed. 4. The apparatus as claimed in claim 3, wherein a plurality of lines extend radially from a common central portion. 5. The apparatus as claimed in claim 3, wherein the plurality of lines are parallel to each other. 6. The apparatus as claimed in claim 3, wherein the color area is at one end of each of the lines and changes both in color and size. 7. The apparatus as claimed in claim 3, wherein each of the ends of the lines has a size area that changes in size and the color area corresponding thereto is adjacent to the respective size area. 8. The apparatus as claimed in claim 1, wherein: the display section includes a first display and a second display, andthe display control section causes the first display to display the notice corresponding to each of the time intervals, with a color corresponding to the load data calculated for the time interval and causes the second display to display the notice with the color corresponding to the load of the last previous predetermined time interval. 9. The apparatus as claimed in claim 1, wherein the display control section further causes the display section to display the notice on a date-by-date basis. 10. An apparatus for displaying a fitness exercise co... | Yamaha Corporation | G01C022/00 A61B005/24
US07860546 | B2 | 2010-12-28 | Medical electrode with self-lifting tabs | 2006-06-01 | A medical electrode includes a backing layer having a top face and a bottom face, and a shrinkable layer covering at least a portion of the top face of the backing layer. Shrinkage of the shrinkable layer results in flexing of a portion of the backing layer to aid placement and attachment of the electrode to the patient. | 1. A medical electrode comprising: an electrically conductive layer;a backing layer having a top face and a bottom face, wherein at least a portion of the bottom face is in contact with the electrically conductive layer; anda shrinkable layer in contact with a portion of the top face of the backing layer, the shrinkable layer being (i) solely positioned on at least a portion of a perimeter of the backing layer, and (ii) defining a tab protruding outside a periphery of the conductive layer. 2. The electrode of claim 1, wherein the shrinkable layer has the form of a single contiguous region or a plurality of non-contiguous regions. 3. The electrode of claim 1, wherein the shrinkable layer comprises a material which shrinks when exposed to one or more of: light, heat, or charged particles. 4. The electrode of claim 3, wherein the shrinkable layer comprises a material which shrinks when exposed to ultraviolet light. 5. The electrode of claim 3, wherein the shrinkable layer comprises a material which shrinks when exposed to an electron beam. 6. The electrode of claim 1, wherein the shrinkable layer comprises a room-temperature vulcanizing material. 7. The electrode of claim 1, wherein the shrinkable layer comprises a varnish. 8. The electrode of claim 1, wherein the shrinkable layer comprises a solid film bonded to the backing layer. 9. The electrode of claim 8, wherein the solid film is a polymer film. 10. The electrode of claim 8, wherein the solid film is a metal film. 11. The electrode of claim 10, where the metal film is deposited using chemical vapor disposition. 12. The electrode of claim 1, wherein the backing layer comprises at least two sublayers, including a top sublayer containing the top face and a bottom sublayer containing the bottom face. 13. The electrode of claim 12, wherein at least one sublayer is electrically conductive. 14. The electrode of claim 12, wherein at least one sublayer is electrically non-conducting. 15. The electrode of claim 1, wherein the backing layer is opaque to infrared, visible, and ultra-violet light. 16. The electrode of claim 1, wherein the electrically conductive layer comprises an electrically conductive gel layer. 17. The electrode of claim 16 further comprising a discontinuous pattern of electrically conducting material contacting the bottom face of the backing layer, the electrically conductive gel layer covering the discontinuous pattern and contacting the bottom face of the backing layer. 18. The electrode of claim 1, wherein the electrically conducting layer, the backing layer, and the shrinkable layer are translucent to x-rays. 19. A medical electrode with a self-lifting tab, comprising: an electrically conductive layer comprising a hydrogel, a discontinuous pattern of silver/silver chloride stripes in contact with the hydrogel and arrayed along diagonals with respect to the edges of the electrode,a backing layer comprising at least one electrically non-conductive sublayer and at l... | Tyco Healthcare Group LP | A61B005/04
US07824341 | B2 | 2010-11-02 | Method and apparatus for locating the fossa ovalis and performing transseptal puncture | 2007-09-19 | A method of identifying the fossa ovalis in a patient by positioning one or more electrodes against the tissue of the interatrial septum of the patient and acquiring unipolar and/or bipolar electrograms of the tissue of the interatrial septum while moving the electrodes to a plurality of positions against the tissue of the interatrial septum. The fossa ovalis is identified on the basis of unipolar voltage reduction, signal fractionation, broadened signal, reduced signal slew rate, reduced local myocardial impedance, increased phase angle and/or increased pacing threshold. An apparatus for identify the fossa ovalis is also provided. | 1. A method of identifying the fossa ovalis in a patient, the method comprising: (a) providing an electrophysiology catheter including at least a pair of electrodes coupled to a portion of the catheter;(b) advancing a guidewire through a lumen within the catheter and into a patient's heart;(c) deploying the catheter over the guidewire so as to position the electrodes against the surface tissue of the interatrial septum of the patient;(d) acquiring one of a plurality of a unipolar electrogram and a bipolar electrogram of the surface tissue of the interatrial septum while moving the electrodes to a plurality of positions against the surface tissue of the interatrial septum; and(e) identifying the fossa ovalis on the basis of at least one of the following parameters: a unipolar voltage reduction, a signal fractionation, a broadened signal, a reduced signal slew rate, a reduced local myocardial impedance, an increased phase angle and an increased pacing threshold. 2. The method of claim 1, wherein the fossa ovalis is also identified on the basis of bipolar voltage reduction and further comprising producing a signal indicating that the fossa ovalis has been identified. 3. The method of claim 1, wherein the fossa ovalis is identified on the basis of at least two of the following parameters: a unipolar voltage reduction, a signal fractionation, a broadened signal, a reduced signal slew rate, a reduced local myocardial impedance, an increased phase angle and an increased pacing threshold. 4. The method of claim 1, wherein the fossa ovalis is identified on the basis of unipolar voltage reduction. 5. The method of claim 1, wherein the catheter is configured and dimensioned to receive a medical device therethrough such that a distal end of the medical device emerges from a distal end of the catheter. 6. The method of claim 5, wherein the catheter is a transseptal dilator and the medical device is a transseptal puncture needle. 7. A method of performing a transseptal puncture on a patient, the method comprising: (a) providing an electrophysiology catheter including at least a pair of electrodes and an open proximal end and an open distal end configured to pass a guidewire therethrough;(b) advancing the catheter into a patient's heart over the guidewire;(c) positioning the electrodes against the surface tissue of the interatrial septum of the patient;(d) acquiring one of a unipolar electrogram and a bipolar electrogram of the surface tissue of the interatrial septum, while moving the electrodes to a plurality of positions against the surface tissue of the interatrial septum;(e) identifying the fossa ovalis on the basis of at least one of the following parameters: a unipolar voltage reduction, a signal fractionation, a broadened signal, a reduced signal slew rate, a reduced local myocardial impedance, an increased phase angle and an increased pacing threshold; and(f) penetrating the surface tissue of the interatrial septum through the fossa ovalis in ord... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B005/05
US07831076 | B2 | 2010-11-09 | Coloring electroanatomical maps to indicate ultrasound data acquisition | 2006-12-08 | During acquisition of ultrasound data in a medical imaging procedure, three-dimensional model of a structure being imaged, e.g., an electroanatomical map, is co-displayed and visually marked, to indicate progress of data acquisition. The plane of intersection successive two-dimensional images are marked on the as a line or colored region on the three-dimensional model. This display enables the operator to determine regions where sufficient data have been captured, and guides the operator to areas where additional data collection is still needed. Various color schemes are used to indicate the relative sufficiency of data collection. | 1. A computer-assisted method for producing images of a living subject, comprising the steps of: displaying a three-dimensional model of a surface of a structure in a body of said subject;acquiring a sequence of two-dimensional images of at least a portion of said structure, said two-dimensional images having respective image planes;while acquiring said sequence, marking said three-dimensional model to show respective intersections of said image planes with said surface;wherein said two-dimensional images are acquired as realtime images by acquiring a plurality of two-dimensional ultrasound slices, computed tomographic slices, or realtime magnetic resonance slices, further comprising the steps of constructing three-dimensional images from said realtime images for use in said step of displaying. 2. The method according to claim 1, wherein said two-dimensional images are selected from the group consisting of two-dimensional ultrasound images, computed tomographic images and magnetic resonance images. 3. The method according to claim 1, further comprising the step of displaying said three-dimensional model and said respective intersections of said image planes with said surface on said three-dimensional model. 4. The method according to claim 3, wherein displaying said three-dimensional model comprises applying a pseudocolor to said respective intersections of said image planes with said surface. 5. The method according to claim 3, wherein marking said three-dimensional model comprises the steps of: interpolating areas of said three-dimensional model between said respective intersections;marking said interpolated areas, andsaid step of displaying said three-dimensional model comprises displaying said interpolated areas. 6. The method according to claim 1, further comprising the steps of: reconstructing a three-dimensional anatomic image of said structure from said two-dimensional images; anddisplaying at least a portion of said three-dimensional anatomic image with said three-dimensional model. 7. The method according to claim 6, wherein said portion of said three-dimensional anatomic image does not extend beyond a predefined distance from a surface of said three-dimensional model. 8. The method according to claim 1, wherein said structure is a heart and said three-dimensional model is an anatomical map. 9. An apparatus for producing images of a living subject comprising: a display; a memory for storing a three-dimensional model of a surface of a structure in a body of said subject;a two-dimensional imaging module operative for acquiring a sequence of two-dimensional anatomic images of at least a portion of said structure, said two-dimensional anatomic images having respective image planes;a processor linked to said memory and to said two-dimensional imaging module, said processor operative for marking said three-dimensional model on said display to show respective intersections of said image planes with said surface; andwherein said two-dim... | Biosense Webster, Inc. | G06K009/00 A61B005/05
US07831296 | B2 | 2010-11-09 | X-ray mammography with tomosynthesis | 2003-11-26 | A method and system for producing tomosynthetic images of a patient's breast. An x-ray source that delivers x-rays through a breast immobilized and compressed between a compression paddle and a breast platform and form an image at a digital x-ray receptor panel. Multiple x-ray images are taken as the x-ray source and the receptor move relative to the immobilized breast. In one preferred embodiment, the x-ray source travels from −15° to +15°. The source can travel in an arc around the breast while the receptor travels linearly while remaining parallel and at the same distance from the breast platform. The sets of x-ray image data taken at different angles are combined to form tomosynthetic images that can be viewed in different formats, alone or as an adjunct to conventional mammograms. | 1. A method of imaging a patient's breast with x-rays using a flat panel digital x-ray imager to obtain both two dimensional (2D) image data for a conventional mammogram data and three-dimensional (3D) image data for tomosynthesis images in a single breast compression, using an anti-scatter grid in obtaining some but not all the image data, comprising: immobilizing a patient's breast between an x-ray source and a flat panel digital x-ray imager;energizing the x-ray source to emit x-rays at each of a plurality of different angular positions of the source relative to the breast while the breast remains immobilized, and concurrently using the imager to derive x-ray projection image data for the respective positions;wherein the image data for at least one of said positions, called hereafter a mammogram position, are taken at a view matching that of a conventional CC and/or MLO view used in conventional mammography, and the image data for a plurality of other positions, called hereafter tomosynthesis positions, are taken at other relative angles of the source and breast;wherein the x-ray dose to the patient's breast when taking said image data for each of said tomosynthesis positions is less than the x-ray dose when taking said image data for said mammogram position, and the x-ray dose for said mammogram position is similar to a dose used for a conventional mammogram;using an anti-scatter grid between the patient's breast an the imager for at least the mammogram position but not for at least some of the tomosynthesis positions; andprocessing the image data taken for the mammogram position to form and display a mammogram and using the image data taken for said tomosynthesis positions to form and display tomosynthesis images of the breast. 2. The method of claim 1, wherein the image data for each of said tomosynthesis positions are taken at a patient x-ray dose that is much less than for the image data for mammogram position. 3. The method of claim 1, wherein the step of energizing the x-ray source at different angular positions comprises energizing the source intermittently during a continuous movement thereof relative to the breast covering at least some said positions. 4. The method of claim 3, wherein the energizing step further includes energizing the source while stationary relative to the breast for at least one of said positions. 5. The method of claim 1, wherein the image data for the mammogram position are acquired when the source is stationary relative to the breast. 6. The method of claim 1, wherein the step of energizing the x-ray source at different angular positions comprises keeping the source stationary relative to the breast at a number of said positions. 7. The method of claim 1, wherein at least some of the tomosynthesis images conform to image planes that are generally parallel to that of the mammogram. 8. The method of claim 1, wherein at least some of the tomosynthesis images conform to image plane that are non-parallel to ... | Hologic, Inc. | A61B005/05
US07840046 | B2 | 2010-11-23 | System and method for detection of breast masses and calcifications using the tomosynthesis projection and reconstructed images | 2007-06-25 | A method of detecting breast masses and calcifications in digitized images, includes providing a plurality of 2-dimensional (2D) digital X-ray projectional breast images acquired from different viewing angles, extracting candidate lesions and 2D features from said 2D projectional images, computing spicularity characteristics of said candidate lesions, including location, periodicity, and amplitude, applying learning algorithms to said candidate lesions to predict a probability of malignancy of said lesion, receiving from said learning algorithm a probability map of detections for each breast image, said detections comprising associating pixels with a probability of being associated with a malignancy, creating a synthetic 2D slice for each X-ray image wherein malignant regions are indicated by ellipses on a non-malignant background, and constructing a synthetic 3-dimensional (3D) image volume from said 2D synthetic slices. | 1. A method of detecting breast masses and calcifications in digitized images, comprising the steps of: providing a plurality of 2-dimensional (2D) digital X-ray projectional breast images acquired from different viewing angles, each said X-ray projectional image comprising a set of intensities on a 2D grid of pixels;extracting candidate lesions and 2D features from said 2D projectional images;computing spicularity characteristics of said candidate lesions, said characteristics including, but not limited to, location, periodicity, and amplitude;applying machine learning algorithms to said candidate lesions to predict a probability of malignancy of said lesion, wherein said learning algorithms have been trained on a training set containing known examples of the features for lesions identified by a physician;receiving from said learning algorithm a probability map of detections for each breast image, said detections comprising associating pixels with a probability of being associated with a malignancy;creating a synthetic 2D slice for each X-ray image wherein malignant regions are indicated by ellipses on a non-malignant background; andconstructing a synthetic 3-dimensional (3D) image volume from said 2D synthetic slices. 2. The method of claim 1, wherein said probability map associates each pixel on each 2D X-ray image with a probability of being associated with a malignancy. 3. The method of claim 1, wherein said probability map returns a list of pixel locations and associated malignancy probabilities, wherein only those locations with malignancy probabilities above a threshold are returned. 4. The method of claim 1, further comprising constructing a 3D image from said 2D X-ray breast images, and superimposing said 3D synthetic image on said constructed 3D X-ray image. 5. The method of claim 1, further comprising thresholding said detections in said 3D synthetic image wherein only those detections whose probability is above said threshold are visualized. 6. The method of claim 1, wherein the malignancy probability in said malignant regions are represented by Gaussian functions inside said ellipses, wherein a peak amplitude of each said Gaussian corresponds to the pixel within said ellipse with a maximum probability of being associated with said malignancy, and whose probability amplitude decreases smoothly to a background value. 7. The method of claim 6, wherein constructing said synthetic 3D image comprises seeking those 2D ellipses in those 2D synthetic slices that overlap in said 3D synthetic image, representing those 3D regions formed by overlapping 2D ellipses by 3D ellipsoids, representing the malignancy probability within each 3D ellipsoid by a 3D Gaussian whose probability amplitude decreases smoothly from a peak value within said ellipsoid to a background value. 8. The method of claim 1, wherein said spicularity characteristics are computed using a spherical or polar coordinate transformation on said lesion followed by band pass f... | Siemens Medical Solutions USA, Inc. | G06K009/00 A61B006/04 A61B005/05
US08956294 | B2 | 2015-02-17 | Body-worn system for continuously monitoring a patients BP, HR, SpO2, RR, temperature, and motion; also describes specific monitors for apnea, ASY, VTAC, VFIB, and ‘bed sore’ index | 2009-05-20 | The invention provides a body-worn monitor that measures a patient's vital signs (e.g. blood pressure, SpO2, heart rate, respiratory rate, and temperature) while simultaneously characterizing their activity state (e.g. resting, walking, convulsing, falling). The body-worn monitor processes this information to minimize corruption of the vital signs by motion-related artifacts. A software framework generates alarms/alerts based on threshold values that are either preset or determined in real time. The framework additionally includes a series of ‘heuristic’ rules that take the patient's activity state and motion into account, and process the vital signs accordingly. These rules, for example, indicate that a walking patient is likely breathing and has a regular heart rate, even if their motion-corrupted vital signs suggest otherwise. | 1. A method for continuously monitoring a patient, comprising the following steps: detecting a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a first sensor comprising a first detector configured to be worn on the patient's body;detecting a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a second sensor comprising a second detector configured to be worn on the patient's body;detecting sets of time-dependent motion waveforms with at least two motion-detecting sensors positioned on locations selected from a forearm, upper arm, and a body location other than the forearm or upper arm of the patient, wherein each set of motion waveforms is indicative of motion of the location on the patient's body to which it is affixed;processing the first and second time-dependent physiological waveforms and the sets of time-dependent motion waveforms with one or more microprocessors, the microprocessor(s) configured to determine at least one vital sign from the patient using the first and second time-dependent physiological waveforms;analyze at least a portion of the sets of time-dependent motion waveforms with a motion-determining algorithm to determine a posture and an activity state for the patient, and determine therefrom if the patient is resting, moving, sitting, standing, walking, running, falling, lying down, or convulsing; andgenerate an alarm by processing the patient's motion and comparing the vital sign to a predetermined motion-dependent alarm/alert thresholds and heuristic rules corresponding to the motion. 2. A method according to claim 1, further comprising displaying a graphical icon simultaneously indicating patient posture and activity state on a monitor positioned at a location remote from the patient. | Sotera Wireless, Inc. | A61B005/0205 A61B005/103 G08B019/00 A61B005/00 A61B005/021 A61B005/11 G08B021/04 A61B005/01 A61B005/024 A61B005/0245 A61B005/04 A61B005/0404 A61B005/0452 A61B005/0456 A61B005/08 A61B005/145 G06F019/00
US08460196 | B2 | 2013-06-11 | Method and apparatus for monitoring brain activity | 2009-05-29 | Method for monitoring brain activity. Left and right cardiac pulse signals are detected at bilateral locations on a body for a selected number of cardiac cycles. Computing apparatus computes the standard deviation of the left and right pulse signals for the selected number of cardiac cycles. The standard deviations are normalized on the computing apparatus by dividing the left and right standard deviations by the mean of the left and right pulse signals computed over the selected number of cardiac cycles to produce a left and a right Coefficient of Variation of the pulse signals. A Bilateral Pulse Index is generated from the left and right Coefficients of Variation. The Bilateral Pulse Index relates to brain activity. Cardiac pulse signals from a single location on the body may be utilized in some embodiments. | 1. Method for monitoring brain activity comprising: detecting with a sensor or sensors left and right cardiac pulse signals at bilateral locations on a body for a selected number of cardiac cycles;computing on a computing apparatus the standard deviation of the left and right pulse signals for the selected number of cardiac cycles;normalizing on the computing apparatus the standard deviation of the left and right pulse signals by dividing the left and right standard deviations by the mean of the left and right pulse signals, respectively, computed over the selected number of cardiac cycles to produce a left and a right Coefficient of Variation (CV) of the pulse signals; andgenerating a Bilateral Pulse Index (BPI) from the left and right Coefficients of Variation according to the equation BPI=((Left CV+Right CV)/2)+(Absolute (Left CV−Right CV)),wherein BPI relates to brain activity. 2. The method of claim 1 wherein the bilateral locations are left and right earlobes. 3. The method of claim 1 wherein the bilateral locations are left and right portions of a forehead. 4. The method of 1 wherein the bilateral locations are left and right carotid arteries at the neck. 5. The method of claim 1 wherein the brain activity is the sleep state. 6. The method of claim 1 wherein the brain activity is awareness under anesthesia. 7. The method of claim 1 wherein the pulse signals are detected by flesh penetrating electromagnetic radiation sensors. 8. The method of claim 1 wherein the pulse signals are detected by plethysmographic or oximeter sensors. 9. The method of claim 1 wherein the sensor or sensors are selected from the group consisting of optical sensors, RF displacement sensors, capacitance displacement sensors, pressure sensors, acoustic distance sensors, acoustic doppler flow sensors, optical displacement sensors, and ear canal sensors. 10. The method of claim 1 wherein the sensor or sensors are connected to the computing apparatus by cables that limit the introduction of motion artifacts during movement. 11. The method of claim 1 wherein the sensor or sensors form a two-channel instrument. | Atlantis Limited Partnership | A61B005/02
US08229538 | B2 | 2012-07-24 | Multi-bend steerable mapping catheter | 2010-07-16 | An electrophysiology catheter introduced through the groin and inferior vena cava into the right side of the heart comprises an elongate flexible shaft having a steerable distal section and a prolapsing section located proximally of the distal section. The distal section is inserted into the coronary sinus and a back-steering force is applied to the catheter to anchor the distal section therein, after which the catheter is further advanced to prolapse the prolapsing section against the high right atrium. Electrical pathways in both the coronary sinus and the high right atrium are mapped using respective electrode pairs carried on the distal and prolapsing sections of the catheter. | 1. An electrophysiology catheter, comprising: an elongate flexible shaft having a steerable distal section, a prolapsing section located proximally of the distal section, and a hinge joint located between the prolapsing section and the steerable distal section, with a soft outer tube extending from a proximal end of the prolapsing section to a distal end of the steerable distal section, the flexible shaft defining an inner lumen;a first plurality of electrodes carried on the steerable distal section;a second plurality of electrodes carried on the prolapsing section; anda flexible inner compression coil extending through the inner lumen of the flexible shaft, wherein a distal portion of the compression coil extends at least partially through the prolapsing section. 2. The catheter of claim 1, wherein the elongate flexible shaft has a distal transition region between the prolapsing section and the steerable distal section, and wherein the hinge joint is disposed in the distal transition region. 3. The catheter of claim 1, wherein the elongate flexible shaft further has a main body section located proximally of the prolapsing section, and wherein the elongate shaft comprises a reinforcing braid embedded in the soft outer tube. 4. The catheter of claim 3, wherein a cross-sectional shape of the compression coil changes from substantially circular in the main body section to substantially oval or flattened in the prolapsing section. 5. The catheter of claim 4, wherein the prolapsing section is configured to prolapse in a prolapse plane that is different from a plane of the distal section when the distal section is positioned in a patient's coronary sinus. 6. The catheter of claim 5, wherein the hinge joint is configured to bend in the prolapse plane. 7. A method of mapping conductive pathways in a patient's heart tissue using the catheter of claim 5, comprising: inserting the catheter through an access location in or proximate to the patient's groin and into the patient's venous system;advancing the catheter through the patient's inferior vena cava and into the patient's right atrium;directing the steerable distal section into the patient's coronary sinus; andfurther advancing the catheter to cause the prolapsing section to prolapse in the prolapse plane against the patient's high right atrium and the hinge joint to bend in the prolapse plane. 8. The method of claim 7, wherein further advancing the catheter torques the catheter between the patient's coronary sinus and high right atrium. 9. The method of claim 7, further comprising applying a back-steering force to the catheter to thereby anchor the distal section in the coronary sinus prior to prolapsing the prolapsing section against the high right atrium. 10. The method of claim 7, further comprising mapping electrical pathways in the coronary sinus using at least one pair of the first plurality of electrodes. 11. The method of claim 7, further comprising mapping electrical pathways in the ... | Boston Scientific Scimed, Inc. | A61B005/04
US08529457 | B2 | 2013-09-10 | System and kit for stress and relaxation management | 2009-02-16 | This invention relates to a system and a kit for stress and relaxation management. A cardiac activity sensor (101) is used for measuring the heart rate variability (HRV) signal of the user and a respiration sensor (102) for measuring the respiratory signal of the user. The system contains a user interaction device (103) having an input unit (104) for receiving user specific data and an output unit for providing information output to the user. A processor (106) is used to assess the stress level of the user by determining a user related stress index. The processor is also used to monitor the user during a relaxation exercise by means of determining a relaxation index based on the measured HRV and respiratory signals, the relaxation index being continuously adapted to the incoming measured signals and based thereon the processor instructs the output unit to provide the user with biofeedback and support messages. Finally, the processor uses the user specific data as an input in generating a first set of rules defining an improvement plan for self-management of stress and relaxation. The first set of rules is adapted to trigger commands instructing the output unit to provide the user with motivation related messages. Also, at least a portion of said user specific data is further used to define a second set of rules indicating the user's personal goals. | 1. A system for stress and relaxation management, comprising: at least one cardiac activity sensor for measuring a heart rate variability (HRV) signal of a user,at least one respiration sensor for measuring a respiratory signal the user (109),a user interaction device comprising an input unit for receiving user specific data and an output unit for providing information output to the user, anda processor adapted to assess the stress level of the user by determining a user related stress index, the user related stress index being determined based on a first and a second HRV parameter value associated with a first posture and a second posture, respectively, where a difference between the HRV parameter values is used as an indicator to determine the user related stress index, the stress index being used as an input in triggering commands instructing the output unit to provide the user with a message indicating the stress level of the user,the processor further being adapted to monitor the user during a relaxation exercise by means of determining a relaxation index based on the measured HRV and respiratory signals by means of calculating an index of coherence between the HRV signal and the respiratory signal during the relaxation exercise, the relaxation index being continuously adapted to the incoming measured signals and based thereon the processor instructs the output unit to provide the user with biofeedback and support messages,the processor further being adapted to use the user specific data as an input in generating a first set of rules defining an improvement plan for self-management of stress and relaxation, said first set of rules being adapted to trigger commands instructing the output unit to provide the user with motivation related messages. 2. A system according to claim 1, wherein said user specific data comprises at least one of: input data entered by the user via said input unit,input data obtained from the user via a response to questionnaires presented to the user,input data from an external agent communicating via said input unit, anddata obtained from measured signals from said sensors. 3. A system according to claim 2, wherein at least a portion of said user specific data is further used to define a second set of rules indicating the user's personal goals. 4. A system according to claim 1, wherein said processor is further adapted to compare the user related stress index, initially with baseline measures and subsequently with previous determined stress indexes, wherein said comparing results in triggering commands instructing the output unit to provide the user with messages indicating the stress level of the user and trends. 5. A system according to claim 4, wherein said baseline measures are determined during a initialization phase where the processor uses the measured signals to determine one or more of the following parameters: a personal range for the stress index value, determined from said stress assessment,a personal ... | Koninklijke Philips N.V. | A61B005/00 A61B005/02 A61B005/04 A61B005/08
US09011327 | B2 | 2015-04-21 | Capacitive sensing and communicating | 2008-12-18 | The invention relates to a device being placeable on or close to a human or animal body, wherein the device comprises a sensing unit configured to capacitively sense a physiological signal from the human or animal body and capacitively communicate a body coupled communication signal for a body area network (BAN). With the device a single sensing unit can be used to sense or measure a physiological signal and communicate a body coupled communication signal. Advantageously, the same couplers can simultaneously be used for sensing, transmitting and receiving. The device does not require an extra antenna and RF unit and there is no need for using wires. | 1. A device being placeable on or close to a human or animal body, the device comprising: a sensing unit configured to capacitively sense a physiological signal during predetermined sensing periods from said human or animal body and to capacitively communicate a body coupled communication signal during predetermined transmission periods,a combining unit configured to generate combined data by combining at least one of: data generated by said device at different points in time and data generated by said device and a plurality of other devices, the combining unit being further configured to combine different measured data to decide or derive a further estimate of a physiological quantity,a transmit data generating unit configured to generate transmit data from at least one of stored data, said combined data and other data,a detecting unit configured to detect data conveyed by a received body coupled communication (BCC) RX signal transmitted from another device,wherein said device is configured to communicate in a synchronous manner with said plurality of other devices such that no device communicates with another device during a sensing period by time multiplexing said predetermined measurement periods and said predetermined transmission periods of each device such that a first predetermined number of time slots are allocated to said predetermined sensing periods and a second predetermined number of time slots are allocated to said predetermined transmission periods. 2. The device according to claim 1, further comprising a first filtering unit configured to low-pass filter or band-pass filter said physiological signal. 3. The device according to claim 1, further comprising a processing unit configured to process a signal based on said physiological signal. 4. The device according to claim 1, further comprising a second filtering unit configured to high-pass filter or band-pass filter a received body coupled communication signal. 5. The device according to claim 1, further comprising a storing unit configured to store at least one of data based on said physiological signal, data conveyed by a received body coupled communication signal and data conveyed by another communication signal. 6. The device according to claim 1, further comprising a modulating unit configured to modulate transmit data to obtain a body coupled communication signal to be transmitted. 7. The device according to claim 1, further comprising a third filtering unit configured to high-pass filter a body coupled communication signal to be transmitted. 8. The device according to claim 1, further comprising a communicating unit configured to communicate a communication signal to a device separated from said human or animal body. 9. The device according to claim 1, wherein said device is configured to operate in a non synchronous mode by simultaneously sensing and communicating said physiological signal said body coupled communication signal. 10. The device to claim 1, wherein ... | Koninklijke Philips N.V. | A61B005/00 H04W074/08 H04B007/26 H04B005/00 G06F019/00 H04B013/00 H04B007/00 H04W084/00 H04W074/00 A61B005/0402 A61B005/0476 A61B005/0488
US08883086 | B2 | 2014-11-11 | Apparatus and method for measuring biomedical data and measurement strip | 2010-03-11 | Disclosed herein are an apparatus and method for measuring biomedical data and a measurement strip. The apparatus includes a plurality of detection units arranged within a strip reception area on a plane and spaced apart from each other, a measurement type determination unit for determining whether reactive portions are present in areas of the measurement strip corresponding to the plurality of detection units based on detection results obtained by the detection units and determining a type of measurement based on results of the determination, a biomedical data measurement unit for activating part or all of the detection units according to the type of measurement determined by the measurement type determination unit, and measuring the biomedical data using the activated detection units, and an output unit for outputting the measured biomedical data to an outside. | 1. An apparatus for measuring biomedical data, comprising: a plurality of detection units configured to respectively correspond to predetermined areas of a measurement strip;a measurement type determination unit for determining whether reactive portions are present in each of the predetermined areas of the measurement strip corresponding to the plurality of detection units based on detection results obtained by the detection units and determining a type of measurement based on the presence of the reactive portion in each of the predetermined areas of the measurement strip;a biomedical data measurement unit for activating at least one detection unit corresponding to the type of measurement determined by the measurement type determination unit among the plurality of detection units, and measuring the biomedical data using the activated detection units; andan output unit for outputting the measured biomedical data to an outside. 2. The apparatus as set forth in claim 1, further comprising a strip insertion determination unit for determining whether the measurement strip has been inserted based on the detection results obtained by the plurality of detection units. 3. The apparatus as set forth in claim 1, wherein the output unit is a display unit for displaying the biomedical data on a screen. 4. The apparatus as set forth in claim 1, wherein each of the plurality of detection units comprises: a light-emitting unit for emitting light toward a reactive portion of the measurement strip; anda light-receiving unit for detecting a reflectance of the reactive portion of the measurement strip. 5. The apparatus as set forth in claim 1, wherein the type of measurement comprises measurement of an amount of total cholesterol, measurement of an amount of High Density Lipoprotein (HDL) cholesterol, and/or measurement of an amount of triglyceride. 6. A method of measuring biomedical data from the measurement strip inserted in the apparatus of claim 1, comprising: determining whether the reactive portions are present in the predetermined areas of the measurement strip corresponding to the plurality of respective detection units of the apparatus, based on the detection results obtained by the detection units;determining the type of measurement based on results of the determination; andmeasuring the biomedical data based on the type of measurement. 7. The method as set forth in claim 6, further comprising, before determining the type of measurement, determining whether the measurement strip has been inserted based on the detection results obtained by the detection units. 8. The apparatus as set forth in claim 1, wherein the activating of at least one detection unit among the plurality of detection unit is determined by amount of light received from each detection unit. | Infopia Co., Ltd. | G01N033/52 A61B005/145 G01N021/00 G01N021/17 G01N033/92 G01N033/487 G01N033/49 G01N021/84 A61B005/1455
US08738118 | B2 | 2014-05-27 | Cable system for generating signals for detecting motion and measuring vital signs | 2009-05-20 | The invention provides a system and method for measuring vital signs (e.g. SYS, DIA, SpO2, heart rate, and respiratory rate) and motion (e.g. activity level, posture, degree of motion, and arm height) from a patient. The system features: (i) first and second sensors configured to independently generate time-dependent waveforms indicative of one or more contractile properties of the patient's heart; and (ii) at least three motion-detecting sensors positioned on the forearm, upper arm, and a body location other than the forearm or upper arm of the patient. Each motion-detecting sensor generates at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which it is affixed. A processing component, typically worn on the patient's body and featuring a microprocessor, receives the time-dependent waveforms generated by the different sensors and processes them to determine: (i) a pulse transit time calculated using a time difference between features in two separate time-dependent waveforms, (ii) a blood pressure value calculated from the time difference, and (iii) a motion parameter calculated from at least one motion waveform. | 1. A system for measuring vital signs from a patient, comprising: (a) a sensor configured to detect a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;(b) a cable system comprising:a cable comprising a cable span terminating in a first terminal end and a second terminal end, the first terminal end comprising a housing integral with the cable and enclosing an integrated circuit configured to operably connect to at least two electrodes and to determine from signals obtained therefrom time dependent digitized ECG waveforms, and a first accelerometer configured to generate a first time-dependent digitized motion waveform,the second terminal end comprising a connector configured to operably connect the cable to a processing component,at least two electrodes configured to detect electrical signals from the patient's body and to operably connect to the integrated circuit, and,a second accelerometer positioned along the cable span and integral with the cable, the second accelerometer configured to generate a second time-dependent motion waveform, the cable being operably connected to the integrated circuit, the first accelerometer, and the second accelerometer to transmit the time dependent digitized ECG waveforms and the first and second time-dependent digitized motion waveforms to the processing component; and,(c) the processing component, wherein the processing component is configured to be worn on the patient's body and comprising a microprocessor and a first electrical connector for operably connecting the microprocessor to the sensor and a second electrical connector for receiving the second terminal end to operably connect the microprocessor to the cable, said processing component configured to receive the time-dependent ECG waveforms, the time-dependent physiological waveform, and the first and second time-dependent motion waveforms transmitted via the cable and determine therefrom: (i) at least one vital sign calculated using at least one of the time-dependent physiological waveform and the time-dependent ECG waveforms, and (ii) at least one motion parameter selected from the patient's posture, activity level, arm height, and degree of motion calculated using the first and second time-dependent motion waveforms. 2. The system of claim 1, wherein the first time-dependent motion waveform is transmitted to the processing component through said cable as a first digital data stream, and the second time-dependent motion waveform is transmitted to the processing component through said cable as a second digital data stream, wherein the first digital data stream is separately resolvable from the second digital data stream by the processing component. 3. The system of claim 2, wherein the time-dependent ECG waveforms are transmitted to the processing component through said cable as a third digital data stream separately resolvable from each of the first and second digital data streams. ... | Sotera Wireless, Inc. | A61B005/02
US08437824 | B2 | 2013-05-07 | Body-worn pulse oximeter | 2009-09-14 | The invention provides a body-worn system that continuously measures pulse oximetry and blood pressure, along with motion, posture, and activity level, from an ambulatory patient. The system features an oximetry probe that comfortably clips to the base of the patient's thumb, thereby freeing up their fingers for conventional activities in a hospital, such as reading and eating. The probe secures to the thumb and measures time-dependent signals corresponding to LEDs operating near 660 and 905 nm. Analog versions of these signals pass through a low-profile cable to a wrist-worn transceiver that encloses a processing unit. Also within the wrist-worn transceiver is an accelerometer, a wireless system that sends information through a network to a remote receiver, e.g. a computer located in a central nursing station. | 1. A method for monitoring a patient, comprising: (a) measuring a first time-dependent signal by detecting a first radiation emitted by a first radiation source after the first radiation passes through a portion of the patient;(b) measuring a second time-dependent signal by detecting electrical signals from the patient using at least two electrodes, and processing the electrical signals with an electrical circuit to form the second time-dependent signal;(c) measuring a time-dependent motion signal with at least one motion sensor;(d) processing the first and second time-dependent signals to determine a transit time corresponding to a single heartbeat in a collection of heartbeats;(e) processing multiple transit times, each corresponding to a heartbeat within the collection of heartbeats, with a first algorithm to identify a portion of the first time-dependent signal that may be affected by motion;(f) analyzing the time-dependent motion signal to estimate if the portion of the first time-dependent signal is affected by motion;(g) applying a mathematical filter that processes both the portion of the first time-dependent signal and a numerical template representing a pre-determined portion of a time-dependent signal if it is determined that the portion of the first time-dependent signal is affected by motion;(h) determining a filtered version of at least a portion of the first time-dependent signal with the mathematical filter; and(i) processing the filtered version of the first time-dependent signal to determine a oxygen saturation value for the patient. 2. The method of claim 1, wherein step (e) further comprises determining a transit time that has a high or low value that lies outside a pre-determined range. 3. The method of claim 2, wherein step (e) further comprises applying a statistical filter to the multiple transit times to determine the transit time that lies outside a pre-determined range. 4. The method of claim 3, wherein the statistical filter comprises taking an average and standard deviation of the multiple transit times, and determining transit times that fall outside the average by more than a multiple of one standard deviation. 5. The method of claim 1, wherein step (f) further comprises comparing an amplitude of the time-dependent motion signal to a pre-determined value to estimate if the portion of the first time-dependent signal is affected by motion. 6. The method of claim 1, wherein step (f) further comprises comparing a power value from a frequency-domain spectrum calculated from the time-dependent motion signal to a pre-determined value to estimate if the portion of the first time-dependent signal is affected by motion. 7. The method of claim 1, wherein the mathematical filter of step (g) is a matched filter that processes both the portion of the first time-dependent signal and a numerical template representing a pre-determined portion of a time-dependent signal if it is determined that the portion of the first time-dep... | Sotera Wireless, Inc. | A61B005/1455
US09585573 | B2 | 2017-03-07 | Apparatus and method for estimating blood pressure by using variable characteristic ratio | 2010-03-15 | Provided is a method for estimating blood pressure using a variable characteristic ratio. In the method, a pressure wave detected at a measurement body part of a subject is obtained to calculate a characteristic ratio of the subject based on a correlation between the characteristic ratio used to determine blood pressure at a measurement body part and the shape of the pressure wave detected at the measurement body part. The blood pressure at the measurement body part of the subject is estimated using the calculated characteristic ratio. | 1. A method of estimating blood pressure, the method comprising: obtaining, by at least one sensor, a pressure wave detected at a measurement body part of a subject and a reference pressure wave detected at a reference measurement body part of the subject which is different from the measurement body part;determining, by at least one processor, an index to be a difference between a blood pressure determined by applying a fixed characteristic ratio and a blood pressure having a maximum amplitude of the detected pressure wave which is transmitted by the at least one sensor;calculating, by at least one processor, a reference characteristic ratio using the detected reference pressure wave which is transmitted by the at least one sensor;calculating, by at least one processor, a variable characteristic ratio based on a correlation between the determined index and the reference characteristic ratio;estimating, by at least one processor, blood pressure at the measurement body part of the subject using the calculated variable characteristic ratio; andoutputting, by an output unit, the blood pressure at the measurement body part of the subject,wherein the fixed characteristic ratio comprises a characteristic ratio which is fixed and irrespective of the shape of the pressure wave,wherein the variable characteristic ratio comprises a characteristic ratio which is a calculated characteristic ratio of the subject and variable according to the shape of the obtained pressure wave, andwherein the reference characteristic ratio comprises a characteristic ratio which is used to determine blood pressure at a measurement body part and determined using a reference blood pressure measured at the reference measurement body part. 2. The method of claim 1, wherein the shape of the pressure wave represents a rate of amplitude reduction with reference to a pressure value having a maximum amplitude as the pressure value one of increases and decreases. 3. The method of claim 1, wherein the reference blood pressure comprises at least one selected from the group consisting of a diastolic blood pressure and a systolic blood pressure, which are determined by applying the fixed characteristic ratio used to determine the blood pressure, irrespective of the shape of the pressure wave, to a pressure wave detected while an upper arm of the measurement body part is pressed. 4. The method of claim 1, wherein the reference blood pressure is measured using an invasive method. 5. The method of claim 1, wherein the index represents the shape of the detected pressure wave. 6. The method of claim 5, wherein the calculating of the variable characteristic ratio according to the shape of the obtained pressure wave includes applying the determined index to a relational equation derived using a plurality of data representing the correlation. 7. The method of claim 5, wherein the estimating of the blood pressure includes at least one blood pressure having an amplitude corresponding to the calc... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/02 A61B005/022 A61B005/021
US20100312143 | A1 | 2010-12-09 | HUMAN BODY MEASUREMENT SYSTEM AND INFORMATION PROVISION METHOD USING THE SAME | 2009-09-28 | A human body measurement system and a method for using such a system. The human body measurement system includes a photographing unit that includes a plurality of cameras and a hub which send images, captured by the cameras, to a PC of an integrated controller through the hub, a weight scale configured to measure body weight of a subject of measurement, and a body composition analyzer configured to measure body composition of the subject of measurement. | 1. A human body measurement system comprising: a photographing unit comprising a plurality of cameras and a hub which send images, captured by the cameras, to a computer of an integrated controller through the hub;a weight scale configured to measure body weight of a subject of measurement; anda body composition analyzer configured to measure body composition of the subject of measurement. 2. The human body measurement system as set forth in claim 1, wherein the image data captured by the photographing unit is three-dimensionally modeled to provide measurements of length and volume of each portion of a human body, 3D simulation is provided based on the measured data, and information about exercise and dietary prescriptions is provided on the basis of the measurement of the body weight using the weight scale and the analysis of the body composition using the body composition analyzer; wherein the simulation and the information are provided simultaneously. 3. The human body measurement system as set forth in claim 1, wherein the body weight is measured using the weight scale at the same time that photographing is performed using the photographing unit, height is measured using the captured data, and the measured data is automatically input to the body composition analyzer and used for analysis of the body composition data measured by the body composition analyzer. 4. The human body measurement system as set forth in claim 3, wherein a 3D shape model is formed by approximating shape information using polygons through extraction of colored corresponding points of multiview images captured by the photographing unit, connecting polygons of each section to each other and then creating a complete mesh structure. 5. The human body measurement system as set forth in claim 1, further comprising: a system control unit for testing and maintaining the system and controlling the photographing unit, the weight scale, the body composition analyzer, an internal monitor, an illuminator, a speaker and a microphone through the computer;an integrated controller for analyzing and storing acquired images, body weight data and body composition data through the system control unit;an operating desk for connecting to the integrated controller over a network and remotely managing and operating the cameras and the analyzer; anda power supply for supplying power to the system. 6. The human body measurement system as set forth in claim 5, wherein the measurement apparatus further comprises an internal monitor for enabling the subject of measurement to monitor status of progress. 7. The human body measurement system as set forth in claim 5, wherein the operating desk comprises a computer for referring to and managing measured data, performing simulation analysis and offering prescriptions, and a communication microphone and speaker for transmitting information about a position of measurement and status of progress to the subject of measurement. 8. The human body measurem... | MINIMEDREAM CO., Ltd. | A61B005/103 H04N005/225
US20100317978 | A1 | 2010-12-16 | IMPLANTABLE MEDICAL DEVICE HOUSING MODIFIED FOR PIEZOELECTRIC ENERGY HARVESTING | 2010-05-06 | Methods, systems, and apparatus for powering and/or recharging medical devices implanted within the body are described. An illustrative implantable sensor for sensing one or more physiologic parameters within a body lumen includes a housing having an exterior wall that has an inner surface and an outer surface and that defines an internal cavity. A portion of the housing includes an electrically conductive material that functions as a first electrical conductor. A flexible piezoelectric layer is disposed adjacent to a portion of the exterior wall and a second electrical conductor is disposed adjacent to the piezoelectric layer. The piezoelectric layer is configured to displace in response to periodic pressure pulses within the body lumen and generate a voltage differential between the first and second electrical conductors. | 1. An implantable sensor for sensing one or more physiologic parameters within a body lumen, comprising: a housing including an exterior housing wall having an inner surface and an outer surface, the exterior housing wall defining an internal cavity and including an electrically conductive material forming a first electrical conductor;a flexible piezoelectric layer disposed adjacent to a portion of the exterior wall;a second electrical conductor disposed adjacent to the piezoelectric layer;a sensor module disposed within the interior cavity of the housing, the sensor module configured to sense one or more physiologic parameters within the body lumen; andpower circuitry configured to convert a voltage differential between the first and second conductors into an operating current for powering one or more components of the implantable sensor;wherein the piezoelectric layer is configured to displace in response to periodic pressure pulses within the body lumen and generate a voltage differential between the first and second electrical conductors. 2. The implantable sensor of claim 1, further comprising a rechargeable storage device electrically coupled to the power circuitry. 3. The implantable sensor of claim 1, wherein the sensor module is at least partially powered by the operating current provided by the power circuitry. 4. The implantable sensor of claim 3, wherein the sensor module comprises a pressure sensor. 5. The implantable sensor of claim 1, further comprising a communication module disposed within the interior cavity, and wherein the communication module is at least partially powered by the operating current provided by the power circuitry. 6. The implantable sensor of claim 1, wherein the housing is configured to be attached to an implantable medical device that includes a component at least partially powered by the operating current provided by the power circuitry. 7. The implantable sensor of claim 6, wherein the housing is configured to be mechanically and electrically coupled to the implantable medical device. 8. The implantable sensor of claim 1, wherein the housing comprises a flexible diaphragm forming a portion of the exterior housing wall, and the piezoelectric layer is disposed on the flexible diaphragm. 9. The implantable sensor of claim 1, wherein the exterior housing wall is flexible and the piezoelectric layer is disposed on the inner surface of the exterior housing wall. 10. The implantable sensor of claim 9, wherein the piezoelectric layer comprises a piezoelectric helix spirally disposed about the inner surface of the exterior housing wall, and wherein the second electrical conductor comprises an electrically conductive helix spirally disposed on the piezoelectric helix. 11. The implantable sensor of claim 1, wherein the exterior housing wall is flexible and the piezoelectric layer is disposed adjacent to the outer surface of the exterior housing wall. 12. The implantable sensor of claim 11, wherein the piezo... | - | A61B005/0215 A61B005/00 A61B005/03 H02N002/18
US20100324390 | A1 | 2010-12-23 | MEASUREMENT OF OXYGEN SATURATION OF BLOOD HAEMOGLOBIN | 2008-11-03 | The invention provides a chest-based oximeter (1) for measuring oxygen saturation of haemoglobin in blood of the chest of a subject, comprising at least one radiation source (5,7) adapted to emit radiation onto the chest, at least one radiation detector (9) adapted to detect radiation reflected from the chest, and a pressure device (11) adapted to apply pressure to the oximeter to connect the oximeter to the chest. | 1. A chest-based oximeter for measuring oxygen saturation of haemoglobin in blood of the chest of a subject, comprising at least one radiation source adapted to emit radiation onto the chest,at least one radiation detector adapted to detect radiation reflected from the chest, anda pressure device adapted to apply pressure to the oximeter to connect the oximeter to the chest. 2. A chest-based oximeter according to claim 1 in which the pressure device is applied to skin of the chest of the subject, and comprises a material which has a Young's modulus which is lower than that of the skin of the chest of the subject. 3. A chest-based oximeter according to claim 2 in which the pressure device stresses and applies a pressure on the oximeter towards the skin, to connect the oximeter to the chest. 4. A chest-based oximeter according to claim 1 in which the pressure device comprises any of a biasing device, a finger push device, a belt. 5. A chest-based oximeter according to claim 1 in which the pressure device applies a pressure in the range of approximately 1 Pascal to approximately 100000 Pascal. 6. A chest-based oximeter according to claim 1 in which the pressure device optically couples the radiation from the radiation source to the chest of the subject, and optically couples the radiation reflected from the chest of the subject to the radiation detector. 7. A chest-based oximeter according to claim 1 in which the at least one radiation source and the at least one radiation detector are mounted in the chest-based oximeter such that they are spaced apart by a distance in the range of approximately 0.5 mm to approximately 2 cm. 8. A chest-based oximeter according to claim 1 in which the at least one radiation source and the at least one radiation detector are mounted in the chest-based oximeter such that, in use, they are positioned in the range of approximately 1 cm to approximately 20 cm above the chest of the subject. 9. A chest-based oximeter according to claim 1 in which the at least one radiation source emits radiation having one or more infra red peak wavelengths. 10. A chest-based oximeter according to claim 9 which further comprises a second radiation source which emits radiation having one or more infra red peak wavelengths. 11. A chest-based oximeter according to claim 10 in which the at least one radiation source emits radiation having at least a first infra red peak wavelength, and the second radiation source emits radiation having at least a second, different, infra red peak wavelength. 12. A chest-based oximeter according to claim 11 in which the or each infra red peak wavelength is in the range of 600 nm to 1500 nm, for example 780 nm, 810 nm, 820 nm, 830 nm, 840 nm, 850 nm, 870 nm, 880 nm, 890 nm, 910 nm, 940 nm, 970 nm, 1050 nm, 1070 nm, 1200 nm, 1300 nm, 1350 nm, 1450 nm, 1550 nm. 13. A chest-based oximeter according to claim 1 in which the at least one radiation source emits radiation having one or more visible peak wave... | - | A61B005/1455 A61B005/0402 A61B005/01 A61B005/08
US20100331632 | A1 | 2010-12-30 | WIRELESS POLYSOMNOGRAPHY SYSTEM | 2010-06-28 | A distributed wireless polysomnography (PSG) system is provided. The system includes plural wireless physiological signal acquiring devices and a base station, in which the base station is wirelessly and bi-directionally communicates with the wireless physiological signal acquiring devices. During PSG examination, the wireless physiological signal acquiring devices are worn by a patient and the base station is connected to a remote computer device via a network. Then, each of the wireless physiological signal acquiring devices acquires physiological signals through at least a sensing element connected thereto and/or built therein, and the acquired physiological signals are wirelessly transmitted, in real time, to the base station and then to, via the network, the remote computer device, so as to achieve a real time monitoring of the patient's physiological signals during sleep. Moreover, the base station is capable of executing at least one of configuring the wireless physiological signal acquiring devices, controlling the operations of the wireless physiological signal acquiring devices, displaying the physiological signals acquired by and transmitted from the wireless physiological signal acquiring devices, and indicating the statuses of the wireless physiological signal acquiring devices during the operations. | 1. A wireless polysomnography (PSG) system, comprising plural wireless physiological signal acquiring devices, and a base station, wherein:the base station wirelessly and bi-directionally communicates with the wireless physiological signal acquiring devices;the wireless physiological signal acquiring devices are worn by a patient and the base station is connected to a remote computer device via a network; andeach of the wireless physiological signal acquiring devices has at least a sensing element connected thereto and/or built therein for acquiring physiological signals, and the acquired physiological signals are wirelessly transmitted, in real time, to the base station and then to, via the network, the remote computer device, so as to achieve a real time monitoring of the patient's physiological signals during sleep; andwherein:the base station is capable of executing at least one of:configuring the wireless physiological signal acquiring devices;controlling the operations of the wireless physiological signal acquiring devices;displaying the physiological signals acquired by and transmitted from the wireless physiological signal acquiring devices; andindicating the statuses of the wireless physiological signal acquiring devices during the operations. 2. The system as claimed in claim 1, wherein each of the wireless physiological signal acquiring devices further comprises a memory for storing the acquired physiological signals to be used in an analysis process. 3. The system as claimed in claim 2, wherein the stored data from the wireless physiological signal acquiring devices are combined in the analysis process. 4. The system as claimed in claim 2, wherein the memory for data storage is implemented to be removable. 5. The system as claimed in claim 1, wherein the acquired physiological signals are stored in the base station and/or the remote computer device. 6. The system as claimed in claim 1, wherein at least one of the wireless physiological signal acquiring devices is implemented to include an event marker. 7. The system as claimed in claim 1, wherein the operations of the wireless physiological signal acquiring devices comprise executing an impedance check. 8. The system as claimed in claim 1, wherein the statuses of the wireless physiological signal acquiring devices comprise the attachments to the patient's body surface. 9. The system as claimed in claim 1, further comprising an adapting device, having a network interface and a power interface, connected the base station. 10. The system as claimed in claim 9, wherein the adapting device further comprises at least an additional communication interface for expanding the functionality of the base station. 11. The system as claimed in claim 9, wherein the adapting device is implemented to have a dock structure for receiving the base station. 12. The system as claimed in claim 1, wherein the base station displays the waveforms of the received physiological signals. 13. The system... | - | A61B005/00
US20100324440 | A1 | 2010-12-23 | REAL TIME STIMULUS TRIGGERED BY BRAIN STATE TO ENHANCE PERCEPTION AND COGNITION | 2009-06-19 | An approach is provided for real time stimulus triggered by brain state and includes receiving data that indicates a brain state and a set of one or more stimuli associated with the brain state. Onset of an instance of the brain state is detected in a subject. In response to detecting onset of the instance, application to the subject of a stimulus of the set is initiated before the instance ends. In some embodiments, the brain state is determined based on a range of values for a function of brain signal data, wherein the range of values is associated with desired performance in response to an associated stimulus. The approach can enhance performance, enhance learning or enhance the probing of impact of that state on perception, action or cognition. | 1 A method comprising: receiving data that indicates a brain state and a set of one or more stimuli associated with the brain state;detecting onset in a subject of an instance of the brain state; andin response to detecting onset of the instance, initiating application to the subject of a stimulus of the set before the instance ends. 2. A method of claim 1, wherein detecting the onset in the subject of the instance of the brain state further comprises determining that a value of a function of one or more electrical signals detected at corresponding electrodes placed near the subject falls within a predetermined range of values. 3. A method of claim 1, wherein the set of one or more stimuli includes a gain for a device to assist perception of external phenomenon. 4. A method of claim 1, wherein: the brain state is associated with a superior capacity to perceive a particular sensory input; andthe set of one or more stimuli includes the particular sensory input. 5. A method of claim 1, wherein: the brain state is associated with a superior capacity to perform a particular function; andthe set of one or more stimuli includes an alert to the subject to attempt to perform the particular function. 6. A method of claim 5, wherein: the particular function is memorization of a fact; andthe alert includes a presentation of the fact. 7. A method of claim 5, wherein the particular function is a movement. 8. A method of claim 1, wherein: the brain state is associated with a superior capacity to perform a particular function; andthe set of one or more stimuli includes a gain for a device to assist the subject to perform the particular function. 9. A method of claim 8, wherein: the particular function is a movement; andthe device causes the subject to execute the movement. 10. A method of claim 1, wherein the brain state is associated with a superior capacity to respond to a stimulus of the set based on measurements of performance of the subject's response to the stimulus and simultaneous measurements of one or more electrical signals detected at corresponding electrodes placed near the subject. 11. A method of claim 1, wherein the brain state is associated with a superior capacity to respond to a stimulus of the set based on measurements of performance of a different subject's response to the stimulus and simultaneous measurements of one or more electrical signals detected at corresponding electrodes placed near the different subject. 12. A method of claim 1, wherein the instance of the brain state ends within about ten seconds of the onset of the instance of the brain state. 13. A method of claim 1, wherein the instance of the brain state ends within about one second of the onset of the instance of the brain state. 14. A method of claim 2, wherein initiating application to the subject of the stimulus before the instance ends further comprises initiating application of the stimulus at a particular phase of an oscillation in the electrical signal th... | Massachusetts Institute of Technology | A61B005/0484
US07850621 | B2 | 2010-12-14 | Method and apparatus for body fluid sampling and analyte sensing | 2004-06-07 | A method of controlling a penetrating member is provided. The method comprises providing a lancing device comprising a penetrating member driver having a position sensor and a processor that can determine the relative position and velocity of the penetrating member based on measuring relative position of the penetrating member with respect to time; providing a predetermined velocity control trajectory based on a model of the driver and a model of tissue to be contacted. In some embodiments, a feedforward control to maintain penetrating member velocity along said trajectory. | 1. A method of controlling a penetrating member at a tissue site for receiving a body fluid from a wound created at the tissue site, the method comprising: (a) providing a lancing device that includes a penetrating member driver coupled to a position sensor and a processor, the position sensor and processor determining a relative position and velocity of the penetrating member based on measuring relative position of the penetrating member with respect to time;(b) launching the penetrating member from the lancing device;(c) measuring relative position of the penetrating member with respect to time with the processor and position sensor;(d) accelerating the penetrating member towards a target tissue;(e) in response to measuring relative position of the penetrating with respect to time using an adaptive control algorithm coupled to the penetrating driver to decelerate the penetrating member to follow a desired trajectory, wherein the penetrating member is not decelerated in a sudden manner where the target tissue bounds away from the penetrating member. 2. The method of claim 1 further comprising: using a planned velocity profile to reach a desired depth of penetration in the tissue site. 3. The method of claim 1 wherein a magnitude of a known value of deceleration is determined with the processor by the deceleration of the penetrating member due to intrinsic frictional forces of the lancing device alone. 4. The method of claim 1 further comprising using a look-up table to determine penetrating member settings based on a desired penetration depth. 5. The method of claim 1 wherein a magnitude of a known value of deceleration is determined by the processor by deceleration of the penetrating member due to intrinsic frictional forces of the lancing device alone. 6. The method of claim 5 wherein a magnitude of a known value of deceleration is determined empirically and implemented by the processor by observing a magnitude of deceleration of the penetrating member when the penetrating member is known to be making contact with the target tissue. | Pelikan Technologies, Inc. | A61B005/00 A61B017/32 A61B017/14 B65D081/00
US07835784 | B2 | 2010-11-16 | Method and apparatus for positioning a reference frame | 2005-09-21 | A method and apparatus to perform a procedure that can include a processor assisted surgical procedure. During the procedure patient space and image space can be registered to allow for tracking of various tracking sensors. A dynamic reference frame can be used to maintain localization of the patient space with the image space. The dynamic reference frame can be fixedly interconnected with a bone portion of the anatomy. | 1. A surgical navigation system for use in a surgical procedure on an anatomy including a bone, comprising: a tracking system operable to track a tracking sensor;a surgical instrument operable to be tracked by said tracking system;a dynamic reference frame having the tracking sensor; anda dynamic reference frame positioning member extending along an axis from a first end to a second end;wherein said first end forms a one piece bone engaging portion having a first portion and a second portion extending at a fixed angle relative to one another;wherein the tracking sensor is tracked by the tracking system;wherein the dynamic reference frame positioning member is operable and configured to engage a bone in the anatomy and be axially impacted along the axis of the dynamic reference frame positioning member into the bone in a selectively fixed manner to hold said dynamic reference frame at a selected position; andwherein the tracking system is configured to correlate the surgical instrument relative to the dynamic reference frame. 2. The surgical navigation system of claim 1, wherein said dynamic reference frame positioning member positions said dynamic reference frame substantially immobile in rotational movement, axial movement, translational movement, yaw movement, pitch movement, and roll movement, or combinations thereof relative to the anatomy. 3. The surgical navigation system of claim 1, wherein said tracking system includes an electromagnetic tracking system, an optical tracking system, an infrared tracking system, an acoustic tracking system, or combinations thereof. 4. The surgical navigation system of claim 1, further comprising: an imaging system operable to obtain patient image data. 5. The surgical navigation system of claim 4, wherein said imaging system is selected from a group consisting of a C-arm, a fluoroscope, an MRI, a PET scanner, a computer tomography system, an x-ray system, an ultrasound system, and accommodations thereof. 6. The surgical navigational system of claim 4, further comprising: a monitor;wherein the image data is displayed on the monitor and said dynamic reference frame is tracked relative to said image data. 7. The surgical navigation system of claim 1, further comprise: at least one of a cannula, a dilator, a slap hammer, an impactor, a tap cap or combinations thereof. 8. The surgical navigation system of claim 1, further comprising a cannula operable to be passed through a selected portion of the anatomy relative to the bone in the anatomy: wherein said dynamic reference frame positioning member is operable to be positioned relative to the bone with said cannula. 9. The surgical navigation system of claim 8, further comprising: a dilator;wherein said dilator is operable with said cannula to move the cannula through a selected portion of the anatomy. 10. The surgical navigation system of claim 1, wherein said surgical instrument is at least one of an impactor, a stylet, a reamer, a reamer driver, a prob... | Medtronic Navigation, Inc. | A61B005/05
US07860560 | B2 | 2010-12-28 | Method for evaluating heart rate variability | 2007-08-20 | The heartbeat frequency is not constant but subject to natural variations. This change in heartbeat frequency is called the heart rate variability (HRV). A method is provided for online-displaying a HRV biofeedback signal by which the cardiovascular state of a patient can be purposefully influenced. The parameters HRVi as determined according to the invention are used for said HRV biofeedback signal, wherein, as a function of the display, the patient should alter his behaviour in such a way that a desired cardiovascular state is attained. The inventive HRVi parameters provide a HRV biofeedback signal with high dynamics which allow the patient to react quickly to changes and therewith to reliably achieve and stabilise his own desired cardiovascular state. | 1. Method for online-displaying a HRV biofeedback signal to a patient with the following steps: providing a plurality of sampling sequences fh,i(kTs) where i=1, 2, 3, . . . , m and k=0, 1, 2, . . . , n, wherein each sampling sequence fh,i(kTs) represents the sampling sequence of the heart rate fh of the patient over the period n·Ts,determining a distribution function Fh,i(f) for each sampling sequence fh,i(kTs) as a function of frequency by means of a spectral analysis,determining a first integral fraction Imax,i of the distribution function Fh,i(f) for a first frequency range, wherein the first frequency range includes a maximum of the distribution function Fh,i(f);determining a second integral fraction Iref,i of the distribution function Fh,i(f) for a second frequency range which includes additional or different frequency ranges to the first frequency range,standardising the integral fractions Imax,i and Iref,i using a standardising factor Kref where Ĩmax,i=Kref*Imax,i andĨref,i=Kref*Iref,i in such a way that for the second integral fraction the condition Ĩref,i≧1 is satisfied,determining a parameter HRVi for each sampling sequence fh,i(kTs) with the respective integral fractions Ĩmax,i and Ĩref,i with a function f(Ĩref) according toHRVi=I~max,i*f⁡(I~ref,i),wherein the function f(Ĩref) for each integral fraction Ĩref,i satisfies the conditions1>f⁡(I~ref,i)>1I~ref,iand0>∂f⁡(I~ref)∂I~ref⁢⁢¦I~ref,i⁢>-f⁡(I~ref,i)I~ref,iand wherein the function f(Ĩref), the first and second frequency ranges and the standardising factor Kref are the same for all m evaluated sampling sequences fh,i(kTs), andusing each parameter HRVi for the HRV biofeedback signal for online-displaying to the patient. 2. Method according to claim 1, wherein the function f(Ĩref) is defined as follows:f⁡(I~ref)=1(I~ref)a⁢⁢where⁢⁢⁢0&lt;a&lt;1. 3. Method according to claim 2, wherein a=0.75. 4. Method according to claim 2, wherein a=0.5. 5. Method according to claim 2, wherein a=0.25. 6. Method according to claim 1, wherein the distribution function Fh,i(f) is determined by way of a fast Fourier transform of the sampling sequence fh,i(kTs). 7. Method according to claim 1, wherein the distribution function Fh,i(f) is determined by way of an autoregressive spectral estimate of the sampling sequence fh,i(kTs). | BioSign GmbH | A61B005/04
US20100331649 | A1 | 2010-12-30 | INTERACTIVE ANALYSIS METHOD FOR BIOSIGNALS | 2009-02-05 | An interactive analysis method for biosignals is disclosed. The method includes steps of providing at least a biosignal acquisition device, a user utilizing the biosignal acquisition device to acquire at least one kind of biosignals therefrom, the user uploading the acquired biosignals to a service platform through a network, wherein the service platform provides at least a set of processing means for said kind of biosignals with at least an algorithm included therein, the user selecting at least a desired processing means from the set of processing means, the service platform performing the selected processing means on the uploaded biosignals for producing an analysis result, the user deciding a handling flow for the analysis result, the user deciding a handling flow for the uploaded biosignals, and the service platform performing the handling flows decided by the user. Therefore, through the biosignal interactive analysis method, not only the users can customize their own operation procedure, but the physiological examination and diagnosis also can be performed at any desired time and location once the biosignal acquisition device is connected to the service platform via network. | 1. An interactive analysis method for biosignals, comprising steps of: providing at least a biosignal acquisition device;a user utilizing the biosignal acquisition device to acquire at least one kind of biosignals therefrom;the user uploading the acquired biosignals to a service platform through a network, wherein the service platform provides at least a set of processing means for said kind of biosignals with at least an algorithm included therein;the user selecting at least a desired processing means from the set of processing means;the service platform performing the selected processing means on the uploaded biosignals for producing an analysis result;the user deciding a handling flow for the analysis result;the user deciding a handling flow for the uploaded biosignals; andthe service platform performing the handling flows decided by the user. 2. The method as claimed in claim 1, wherein the algorithm is implemented to be analysis algorithm and/or signal processing algorithm. 3. The method as claimed in claim 2, wherein the signal processing algorithm is related to time domain, frequency domain, parameter/characteristic or wavelet analysis. 4. The method as claimed in claim 2, wherein the parameters and operation conditions of the analysis algorithm and the signal processing algorithm are capable of be modified by the user. 5. The method as claimed in claim 1, wherein the biosignals acquired by the biosignal acquisition device are transmitted to a local device first and then further transmitted to the service platform, 6. The method as claimed in claim 5, wherein the local device has a network interface for connecting to the network. 7. The method as claimed in claim 6, wherein the network interface is a wireless network interface or wired network interface. 8. The method as claimed in claim 5, wherein the local device provides a user interface for the user to control functions provided by the service platform. 9. The method as claimed in claim 8, wherein the user interface is a web-based user interface, an application-based user interface, or a mixed user interface. 10. The method as claimed in claim 5, wherein the local device is a computer, a PDA, a mobile phone, or a device with networking capability. 11. The method as claimed in claim 1, wherein the biosignal acquisition device is integrated with the local device to be one entity. 12. The method as claimed in claim 1, wherein said at least one kind of biosignals is one or more biosignals selected from a group consisting of: ECG, airflow, snore, respiratory effort, EEG, oxygen saturation, limb movement, EMG, EOG, blood pressure, heart rate, blood sugar, and physiological ultrasonics. 13. The method as claimed in claim 1, wherein when the biosignal acquisition device is connecting to the service platform, the service platform performs an authentication to identify the user. 14. The method as claimed in claim 1, wherein the handling flow of the biosignals includes viewing the tr... | - | A61B005/021 A61B005/00 A61B005/0402 A61B005/087 A61B005/0476 A61B005/11 A61B005/0488 A61B005/145 G06F003/048 G06F021/00
US20100310755 | A1 | 2010-12-09 | ULTRA-SENSITIVE DETECTION OF FAINT FINGERPRINTS ON ROUGH SURFACES AND CORRESPONDING FINGERPRINT DETECTION KIT | 2009-06-04 | An ultra-sensitive method for visualizing latent fingerprints using a chemical multiplier to amplify even very faint prints on rough or otherwise difficult surfaces, involving a chemical free radical reaction which produces millions of identifiable molecules from each adsorbed initiator molecule and thereby greatly increases the detection sensitivity. A corresponding fingerprint detection kit may be utilized to carry out such method, in which the kit includes a free radical chain reaction initiator, and a reagent that can react via the free radical chain reaction in the presence of the initiator and form colored, fluorescent, polymeric, IR- or UV-absorbing molecules that are effective for visualization of the fingerprint. | 1. A method useful in visualizing fingerprints, comprising contacting a surface having a potential visualizable fingerprint thereon with a free radical chain reaction initiator, contacting the surface with a reagent capable of reacting via the free radical chain reaction in the presence of the initiator to form a colored, fluorescent, polymeric, infrared radiation-absorbing, or ultraviolet radiation-absorbing reaction product for at least partial visualization of a fingerprint when present on said surface. 2. The method of claim 1 where the initiator comprises a gaseous compound selected from the group consisting of iodine, bromine, chlorine, nitrogen dioxide, hydrogen peroxide and ozone. 3. The method of claim 1, wherein the initiator comprises a solution of a compound selected from the group consisting of organic peroxides, hydroperoxides, peroxy acids, diazo and azo compounds. 4. The method of claim 1, wherein the initiator is applied as a spray or a mist of a solution of the initiator. 5. The method of claim 1, wherein the reagent comprises a solution of a chromogenic compound selected from the group consisting of organic compounds that include aromatic amines with single or multiple rings and/or single or multiple amine groups. 6. The method of claim 1, wherein the reagent comprises a solution of a chromogenic compound selected from the group consisting of 3,3′-dimethoxybenzidine, benzidine, 4-amino-1-naphthalene sulfonic acid sodium salt, N,N′-diphenylbenzidine, o-Tolidine and 3,3′-dimethoxybenzidine dihydrochloride. 7. The method of claim 1, wherein the reagent comprises a solution of a monomer with an olefinic bond that can be polymerized by a free radical chain reaction. 8. The method of claim 1, wherein initiation of the free radical chain reaction is conducted using radiation selected from the group consisting of visible, IR and UV light. 9. The method of claim 1, wherein initiation of the free radical chain reaction is conducted using heating of the surface with hot air and/or infrared radiation. 10. The method of claim 1, further comprising recording said fingerprint using a camera operative to record images in the visible, infrared or ultraviolet range. 11. A method for visualization of fingerprints, comprising: exposing a surface suspected of having one or more fingerprints on it, to a vapor, mist or solution of chemical material capable of initiating a free radical chain reaction; andexposing the surface to a vapor, mist or solution of a reagent that can react via the free radical chain reaction and form colored, fluorescent, polymeric, IR or UV-absorbing molecules,whereby the free radical chain reaction is effected to at least partially visualize the fingerprints. 12. The method of claim 11, wherein the free radical chain reaction is effected with use of thermal heating, ultraviolet radiation exposure, or a combination thereof. 13. The method of claim 11, wherein the free radical chain reaction is effected under ambie... | APPEALING PRODUCTS, INC. | A61B005/117
US07856261 | B2 | 2010-12-21 | Cable with spacer for positioning multiple medical sensors | 2005-12-05 | A subject monitoring apparatus is provided. The apparatus includes a drive unit (10); a sensor unit (50); and a plurality of cable end units (500) for providing electrical connection between electrodes disposed on a subject and cables passing from the cable end units to the drive and sensor units. Each cable end unit includes first and second connectors (340, 380), and a spacer (400) which substantially maintains a pre-determined distance (L) between the first and second connectors. | 1. A subject monitoring apparatus comprising: a drive unit;a sensor unit;a plurality of cable end units for providing electrical connection between electrodes disposed on a subject and cables passing from the cable end units to the drive and sensor units, each cable end unit comprising:first and second connectors; anda spacer which substantially maintains a pre-determined distance between the first and second connectors, said spacer comprising a plurality of links and a plurality of joints which join the links. 2. A subject monitoring apparatus as set forth in claim 1 wherein: the first connector is disposed at a terminal end of a drive cable; andthe second connector is disposed at a terminal end of a sensor cable. 3. A subject monitoring apparatus as set forth in claim 2 wherein: the first and second connectors have respective first and second housings and the first and second housings and the spacer are of unitary construction. 4. A subject monitoring apparatus as set forth in claim 2 wherein: the spacer is flexible in a direction which accommodates shape and/or movement of the subject. 5. A subject monitoring apparatus as set forth in claim 2 wherein: the electrodes include individually fabricated ECG electrodes. 6. A subject monitoring apparatus as set forth in claim 5 wherein: each electrode includes a male snap portion which protrudes from a base portion; andeach connector includes one of a female snap portion and a grabber, for making connection with the male snap portion. 7. A subject monitoring apparatus as set forth in claim 1 wherein: the joints include hinge joints. 8. An apparatus for electrically connecting a subject monitor with a subject, the apparatus comprising: one or more cables including first and second cable branches having first and second terminal ends with respective first and second connectors wherein each connector is adapted to mechanically connect with an electrode secured to a subject to form electrical contact between the electrode and the respective cable branch; anda spacer disposed between the first and second terminal ends which maintains the first and second connectors spaced apart from each other substantially at a pre-determined distance, said spacer comprising a plurality of links and a plurality of joints which join the links. 9. An apparatus as set forth in claim 8 wherein: the spacer is flexible. 10. An apparatus as set forth in claim 8 wherein: the first cable branch is a drive branch of an impedance cardiography subject monitor; andthe second cable branch is a sensor branch of the impedance cardiography subject monitor. 11. A main cable for use with subject monitors, the cable comprising: a plurality of first and second cable branches;a plurality of cable end units disposed at terminal ends of the first and second cable branches;the cable end unit comprising: first and second connectors; anda spacer which spaces the first and second connectors apart from each other substantially at a pre-dete... | Koninklijke Philips Electronics N.V. | A61B005/04
US07850618 | B2 | 2010-12-14 | Method and apparatus for determining conditions of biological tissues | 2008-01-10 | The present invention relates to a method of determining characteristics of biological tissues in humans and animals. In particular, it relates to determining the characteristics of tissues such as the lungs and airways by introducing a sound to the tissue, and recording the sound. The invention further includes an apparatus capable of such measurement. In a first aspect of the present invention there is provided a method of determining characteristics of biological tissue in situ, including: introducing a sound to the tissue at first position; detecting the sound at another position spaced from the first position after it has traveled through the tissue; calculating the velocity and attenuation of sound that has traveled through the tissue from the first position to another position; and correlating the velocity and attenuation of the detected sound to characteristics of the biological tissue. | 1. Apparatus for determining characteristics of biological tissues, the apparatus comprising: a sound generating device which generates a sound signal having audible frequencies;a sound introduction means comprising a nosepiece adaptor for introduction of the generated sound into the biological tissue;a recording device which records the sound after it has traveled from one position of the biological tissue, through the tissue and to another position of the tissue;an analysis device which calculates the velocity and attenuation with which the sound travels through the tissue. 2. Apparatus according to claim 1, wherein the characteristics of biological tissues comprise characteristics of one or both of the lung and airway. 3. Apparatus according to claim 1, wherein the sound introduction means is adapted for introduction of the generated sound into one or both nares of the nose. 4. Apparatus according to claim 1, wherein sound introduction means is adapted for introduction of the generated sound into both the mouth and one or both nares via the nosepiece adaptor, and the mouth. 5. Apparatus according to claim 1, wherein the sound introduction means comprises an electroacoustic transducer. 6. Apparatus according to claim 5, wherein the electroacoustic transducer generates audible sound in bursts of 0.1 second to 20 seconds duration. 7. Apparatus according to claim 1, wherein the audible sound generated is in the range of 20 Hz to 25 kHz. 8. Apparatus according to claim 1, wherein the analysis device further performs spectral analysis on the data recorded. 9. Apparatus according to claim 1, wherein the analysis device comprises one or more filters adapted for removing environmental noise and unwanted signal recorded by the recording device. 10. Apparatus according to claim 1, wherein the recording device comprises a microphone. 11. Apparatus according to claim 2, wherein the characteristics of biological tissues comprise characteristics of the upper airway. 12. Apparatus according to claim 1, adapted for continuous monitoring and determination of characteristics of biological tissues. | Pulmosonix Pty Ltd. | A61B005/08
US07850674 | B2 | 2010-12-14 | Glucose measuring module and insulin pump combination | 2008-09-26 | A combination glucose measuring and insulin pumping device is described. The two major components, a glucose measuring module and an insulin pump, are held together by a quick attach and release mechanism. Communication between the glucose measuring module and the insulin pump is by a wireless modality. The glucose measuring module determines the glucose level in a sample, and wirelessly transmits the data to the insulin pump, where the data are stored in a memory, and are available for visual display on the insulin pump, and for incorporation into selection of appropriate protocols for the rate of insulin infusion by the pump into the patient. | 1. A method of measuring the glucose level of a diabetic using a modular glucose measuring module and insulin pump combination comprising: (a) attaching a glucose measuring module without a display to an insulin pump housing using a mechanical clasp comprising a quick attach and release mechanism, the attaching including engaging a lip of the quick attach and release mechanism with a lip-engaging portion of the pump housing to align the insulin pump housing with the glucose measuring module;(b) inserting a glucose test strip in the module;(c) contacting the glucose test strip to a blood sample;(d) determining data relating to a glucose value based on said sample using a microprocessor in the module;(e) transmitting said data wirelessly from said module to said insulin pump and converting said data to a glucose value using a microprocessor in said pump;(f) displaying said glucose value on a display on said pump; and(g) detaching the clasp and module from the pump housing. 2. The method of claim 1 wherein the insulin pump housing is watertight. 3. The method of claim 1 including activating the glucose measuring module when the glucose test strip is inserted. 4. The method of claim 3 wherein activating the glucose measuring module includes transmission of a signal to the insulin pump. 5. The method of claim 4 wherein the signal is displayed on the display of the pump indicating activation of the glucose measuring module. 6. The method of claim 1 including controlling one or more functions of the glucose measuring module using a control button provided on the insulin pump housing. 7. A method of measuring the glucose level of a diabetic using a modular glucose measuring module and insulin pump combination, comprising: (a) attaching a glucose measuring module without a display to an insulin pump housing with a mechanical clasp comprising a quick attach and release mechanism in a single action;(b) inserting a glucose test strip in the module;(c) contacting the glucose test strip to a blood sample;(d) determining data relating to a glucose value based on said sample using a microprocessor in the module;(e) transmitting said data wirelessly from said module to said insulin pump and converting said data to a glucose value using a microprocessor in said pump;(f) displaying said glucose value on a display on said pump; and(g) detaching the clasp and module from the pump housing in a single action. 8. A method as in claim 7, wherein the attaching comprises engaging a lip of the quick attach and release mechanism with a lip-engaging portion of the pump housing to align the insulin pump housing with the glucose measuring module. 9. The method of claim 7 including activating the glucose measuring module when the glucose test strip is inserted. 10. The method of claim 9 wherein activating the glucose measuring module includes transmission of a signal to the insulin pump. 11. The method of claim 10 wherein the signal is displayed on the display of the p... | Abbott Diabetes Care Inc. Deltec, Inc. | A61M031/00 A61B005/00
US07890170 | B2 | 2011-02-15 | Rate aberrant beat selection and template formation | 2010-07-19 | A cardiac rhythm management system can be used to detect episode beats associated with cardiac events in a subject's body. These events may be monitored and depolarization morphology information can be derived for candidate arrhythmic beats in an arrhythmia episode. An arrhythmic beat morphology template may be formed from selecting at least one of the candidate arrhythmic beats based upon user's labeling according to specific morphologies of one or more candidate episodes. Methods of use are also presented. | 1. A method for discriminating arrhythmias, comprising: selecting one or more arrhythmia episodes of an arrhythmia type from a plurality of arrhythmia episodes;screening beats of the one or more arrhythmia episodes using specified criteria;determining morphology similarity scores each between two beats of the screened beats, including calculating a correlation between the two beats;selecting beats from the screened beats using the morphology similarity scores; andforming an arrhythmic beat morphology template for the arrhythmia type using the selected beats. 2. The method of claim 1, further comprising: determining a similarity between a beat and the arrhythmic beat morphology template after the formation of the arrhythmic beat morphology template; andclassifying the beat using the similarity. 3. The method of claim 1, wherein selecting the beats comprises: grouping the screened beats into groups of morphologically similar beats using the morphology similarity scores; andselecting a group having a largest number of the morphologically similar beats. 4. The method of claim 3, wherein forming the arrhythmic beat morphology template comprises averaging the morphologically similar beats of the selected group. 5. The method of claim 3, wherein selecting the beats comprises: ranking the groups using the morphology similarity scores and one or more additional criteria; andselecting a group using the ranking of the groups. 6. The method of claim 1, further comprising confirming validity of the arrhythmic beat morphology template using a similarity between the arrhythmic beat morphology template and morphologies of the beats selected to form the arrhythmic beat morphology template. 7. The method of claim 6, wherein forming the arrhythmic beat morphology template comprises forming the arrhythmic beat morphology template for the arrhythmia type using the selected beats if the number of the selected beats exceeds a specified minimum number. 8. The method of claim 1, wherein selecting the one or more arrhythmia episodes of the arrhythmia type comprises: performing automatic classifications of episodes of the plurality of arrhythmia episodes;receiving physician classifications of the episodes of the plurality of arrhythmia episodes; andcomparing the automatic classifications to the physician classifications. 9. The method of claim 8, further comprising reviewing and validating the formed arrhythmic beat morphology template by a physician. 10. The method of claim 1, wherein forming the arrhythmic beat morphology template for the arrhythmia type comprises forming an arrhythmic beat morphology template for supraventricular tachyarrhythmia (SVT). 11. A method for discriminating arrhythmias, comprising: classifying at least one arrhythmia episode as an arrhythmia type, the at least one arrhythmia episode including a plurality of beats;determining morphology similarity scores each between a two beats of the plurality of beats;grouping beats of the pluralit... | Cardiac Pacemakers, Inc. | A61N001/00
US08321001 | B2 | 2012-11-27 | Implantable P-wave dispersion detector | 2010-04-16 | An electromedical implant, having a far-field electrocardiogram capturing unit for recording a far-field electrocardiogram signal, which is connected or can be connected to at least two electrodes for recording electric signals reflecting the curve of the far-field electrocardiogram of the right and left atria,a detection unit, which is designed to detect signal features characterizing atrial cardiac actions in an electrocardiogram signal,an averaging unit, which is connected to the recording unit and the detection unit and designed to generate an averaged P-wave signal in that the averaging unit averages a plurality of signal sections of the electrocardiogram signal associated with a particular detected atrial cardiac action, andan evaluation unit, which is connected to the averaging unit and designed to determine the duration of an averaged P-wave in the particular averaged P-wave signal. | 1. An electromedical implant, comprising at least two electrodes configured to record electric signals that reflect a curve of a far-field electrocardiogram signal of a right and left atria of a heart;a far-field electrocardiogram capturing unit configured to record a first far-field electrocardiogram signal and record a second far-field electrocardiogram signal between different electrodes of said at least two electrodes,wherein said far-field electrocardiogram capturing unit is connected or can be connected to the at least two electrodes;a detection unit, which is configured to detect signal features that characterize atrial cardiac actions in an electrocardiogram signal;an averaging unit, which is connected to the far-field electrocardiogram capturing unit and the detection unit and is configured to average a plurality of signal sections of the first far-field electrocardiogram signal associated with a particular detected atrial cardiac action to generate a first averaged P-wave signal associated with the first far-field electrocardiogram signal;average a plurality of signal sections of the second far-field electrocardiogram signal associated with a particular detected atrial cardiac action to generate a second averaged P-wave signal associated with the first far-field electrocardiogram signal;determine the earliest onset of atrial excitation through a search of said first averaged P-wave signal and said second averaged P-wave signal; and,an evaluation unit, which is connected to the averaging unit and which is configured to determine a duration of an averaged P-wave in a particular averaged P-wave signal based on said earliest onset of said atrial excitation in said first averaged P-wave signal and said second averaged P-wave signal. 2. The electromedical implant according to claim 1, wherein the detection unit is connected to the far-field electrocardiogram capturing unit, or is part thereof, and wherein said detection unit is configured to detect the signal features that characterize the atrial cardiac actions in the electrocardiogram signal recorded by the far-field electrocardiogram capturing unit. 3. The electromedical implant according to claim 1, wherein the detection unit is configured to detect the signal features that characterize the atrial cardiac actions in a filtered electrocardiogram signal. 4. The electromedical implant according to claim 1, further comprising a housing having an electrically conductive section, wherein the far-field electrocardiogram capturing unit is connected to the electrically conductive section of the housing of the electromedical implant as a first of the at least two electrodes. 5. The electromedical implant according to claim 4, further comprising a defibrillation electrode configured for intracardiac positioning, wherein the far-field electrocardiogram capturing unit is connected to the defibrillation electrode as the second of the at least two electrodes. 6. The electromedical implant according... | Biotronik CRM Patent AG | A61B005/02 A61B005/04
US08747327 | B2 | 2014-06-10 | Apparatus and method for measuring blood pressure | 2010-01-04 | A blood pressure measuring apparatus includes a sensing unit including a plurality of sensors sensing sphygmus waves at a measurement site, a selection unit selecting one sensor of the plurality of sensors based on the sphygmus waves sensed by the plurality of sensors, and a blood pressure estimation unit estimating blood pressure of the measurement site based on a sphygmus wave sensed by the selected sensor. | 1. A blood pressure measuring apparatus comprising: a sensing unit comprising a plurality of sensors each sensing a sphygmus wave at a same measurement site;an analysis unit which filters each sphygmus wave,generates an envelope for each filtered sphygmus wave, andcalculates a waveform characteristic value for each envelope from slopes or slope variations of the each envelope;a comparison unit comparing the characteristic values of the envelopes to determine the smallest or largest characteristic value, among the characteristic values of the envelopes;a selection unit selecting one sensor of the plurality of sensors which sensed the sphygmus wave, the envelope of which has the smallest or largest characteristic value; anda blood pressure estimation unit estimating blood pressure of the measurement site based on the sphygmus wave which is sensed by the selected one sensor. 2. The blood pressure measuring apparatus of claim 1, wherein the analysis unit filters the each sphygmus wave through a high-pass filter to pass a high-frequency component of the each sphygmus wave, and determines a maximum amplitude of the each filtered sphygmus wave; andthe comparison unit compares the determined maximum amplitudes of the filtered sphygmus waves to determine the largest maximum amplitude. 3. The blood pressure measuring apparatus of claim 2, wherein the selected one sensor sensed the sphygmus wave, of which the filtered sphygmus wave has the largest maximum amplitude. 4. The blood pressure measuring apparatus of claim 1, wherein the analysis unit filters the each sphygmus wave through a low-pass filter to pass a low-frequency component of the each sphygmus wave,generates the envelope from each low-pass filtered sphygmus wave, andcalculates the characteristic value for the each envelope from the slopes or the slope variations of the each envelope. 5. The blood pressure measuring apparatus of claim 4, wherein the analysis unit determines for the each envelope, an envelope inflection zone where a slope of the envelope consecutively increases and decreases, andcalculates for the each envelope, a maximum slope and a difference between the increasing slope and the decreasing slope in the envelope inflection zone, andthe comparison unit compares the maximum slopes and the differences, to determine the largest maximum slope and the largest difference, andwherein the selected one sensor sensed the sphygmus wave, the envelope of which has the largest maximum slope, orthe selected one sensor sensed the sphygmus wave, the envelope of which has the largest difference. 6. The blood pressure measuring apparatus of claim 4, wherein the analysis unit determines for the each envelope, an envelope inflection zone where a slope of the envelope consecutively decrease and increases, andcalculates for the each envelope, a minimum slope and a difference between the decreasing slope and the increasing slope in the envelope inflection zone, andthe comparison unit compares the min... | Samsung Electronics Co., Ltd. | A61B005/02
US08550996 | B2 | 2013-10-08 | Medical vital sign indication tool, system and method | 2008-07-25 | A vital sign evaluation tool (100, 200, 300) for evaluating a measure of severity of illness associated with a patient is provided, where the tool comprises a displayed grid (102) divided into a plurality of coded regions (104 to 106, 206, 306) each assigned to a vital sign associated with the patient. Each of the plurality of coded regions are scaled according to one of a plurality of age ranges (136, 202, 302) associated with the patient and adapted to generate a score value corresponding to a measurement of the vital sign. The score value within the plurality of the coded regions is summed to produce a total score (156), where the total score is adapted to provide an indication of the severity of illness of the patient. | 1. A computer implemented method of determining a measure associated with a severity of illness of a patient, the method comprising: receiving, by the computer, information indicating the age of the patient;receiving, by the computer, measurement data pertaining to the patient, the measurement data comprising a plurality of vital signs including respiratory rate, heart rate, respiratory effort, oxygen therapy, capillary refill time, transcutaneous oxygen saturation, and systolic blood pressure;determining, by the computer, for each vital sign a set of ranges, wherein each range has associated therewith a score value based on a severity of illness, and wherein the upper and lower bounds of each range are determined based on the age of the patient;processing, by the computer, the measurement data to identify, for each vital sign, a corresponding range and a corresponding score value, thereby obtaining a set of score values; andprocessing, by the computer, the set of score values and calculating a total score associated with a severity of illness of the patient. 2. The method according to claim 1 wherein the score value comprises an integer value that increases based on the vital sign deviating from normality. 3. The method according to claim 1 wherein the total score comprises an integer value that increases based on an increase in the severity of illness. 4. The method according to claim 1 further comprising processing the total score to determine one or more care recommendations. 5. The method according to claim 4 wherein the one or more care recommendations comprise admission of the patient to an intensive care unit (ICU) based on the total score exceeding a pre-determined threshold. 6. The method according to claim 4 wherein the one or more care recommendations comprise the scheduling of a patient assessment. 7. The method according to claim 1 further comprising processing the total score to predict the occurrence of an adverse event. 8. The method according to claim 1 further comprising processing the total score to determine a resource allocation recommendation. 9. The method according to claim 8 further comprising processing a plurality of total scores associated with a plurality of patients to determine a resource allocation recommendation. 10. The method according to claim 9 wherein the resource allocation recommendation comprises a determination of medical staffing needs. 11. The method according to claim 10 wherein the determination of medical staffing needs comprises a determination of a number of nurses. 12. The method according to claim 1 further comprising determining, based on the total score, a designated time interval for subsequent documentation of the set of measurements for the patient. 13. The method according to claim 12 wherein the designated time interval is determined, at least in part, based on whether or not the total score is obtained as an initial determination of the total score or as a subsequent determin... | The Hospital for Sick Children | A61B005/00
US08369931 | B2 | 2013-02-05 | Method of determining the position of a deep brain stimulation electrode | 2008-06-23 | The present invention relates to a method of determining the position of a deep brain stimulation (DBS) electrode which finds the position of the DBS electrode with respect to a deep brain target region, by using a first volume data set containing information on the deep brain target region and a second volume data set containing information on the DBS electrode implanted toward the deep brain target region, and which includes: a first step of generating a subvolume of the deep brain target region from the first volume data set, and also generating a subvolume of the DBS electrode from the second volume data set; and a second step of overlapping and displaying the subvolume of the deep brain target region and the subvolume of the DBS electrode. | 1. A method of determining a position of a deep brain stimulation (DBS) electrode, the method locating the position of said DBS electrode with respect to a targeted region in a deep brain of a living subject, by using a first volume data set including information on said targeted region and a second volume data set including information on said DBS electrode implanted toward said targeted region, the method comprising the steps of: generating a subvolume of said targeted region from said first volume data set, and also generating a subvolume of said DBS electrode from said second volume data set;overlapping and displaying said subvolume of said targeted region and said subvolume of said DBS electrode to determine said position of said DBS electrode with respect to said targeted region, with said first volume data set and said second volume data set aligned; anddetermining the position of the DBS electrode with respect to said target region from viewing the overlapping of the subvolumes of said targeted region and said DBS electrode,wherein said first volume data set is acquired before the implantation of said DBS electrode, and said second volume data set is acquired after the implantation of said DBS electrode and,wherein said first volume data set is an MRI image and said second volume data set is a CT image. 2. The method of claim 1, wherein the alignment of said first volume data set with said second volume data set is carried out by means of mutual information between said first volume data set and said second volume data set. 3. The method of claim 1, comprising: a step of reslicing said first volume data set on the basis of inter commissural anterior commissure (AC)-posterior commissural (PC) so as to improve identification of said targeted region, prior to the alignment. 4. The method of claim 3, comprising: a step of reslicing said second volume data set on the basis of AC-PC so as to reduce an alignment time, prior to the alignment. 5. The method of claim 4, wherein the alignment is carried out by means of mutual information between said first volume data set and said second volume data set. 6. The method of claim 1, wherein said subvolume of said targeted region is generated from said first volume data set by means of an atlas. 7. The method of claim 1, wherein said targeted region is a subthalamic nucleus (STN). 8. A method of determining a position of a deep brain stimulation (DBS) electrode, comprising the steps of: acquiring a first volume data set including information on a targeted region in a deep brain of a living subject and a second volume data set including information on said DBS electrode implanted toward said targeted region;determining a correlation between said first volume data set and said second volume data set for alignment by means of mutual information;extracting at least said targeted region from said first volume data set and extracting at least said DBS electrode from said second volume data set;displayi... | Cybermed, Inc. Seoul National University Industry Foundation | A61B005/00
US08527251 | B2 | 2013-09-03 | Method and system for multi-component heart and aorta modeling for decision support in cardiac disease | 2010-04-30 | A method and system for generating a patient specific anatomical heart model is disclosed. Volumetric image data, such as computed tomography (CT), echocardiography, or magnetic resonance (MR) image data of a patient's cardiac region is received. Individual models for multiple heart components, such as the left ventricle (LV) endocardium, LV epicardium, right ventricle (RV), left atrium (LA), right atrium (RA), mitral valve, aortic valve, aorta, and pulmonary trunk, are estimated in said volumetric cardiac image data. A multi-component patient specific anatomical heart model is generated by integrating the individual models for each of the heart components. Fluid Structure Interaction (FSI) simulations are performed on the patient specific anatomical model, and patient specific clinical parameters are extracted based on the patient specific heart model and the FSI simulations. Disease progression modeling and risk stratification are performed based on the patient specific clinical parameters. | 1. A method for multi-component heart modeling and cardiac disease decision support, comprising: receiving at least one sequence of volumetric cardiac imaging data of a patient generated using at least one medical imaging modality;generating a multi-component patient specific 4D geometric model of the heart and aorta estimated from the at least one sequence of volumetric cardiac imaging data of the patient by: estimating individual 4D models for a plurality of heart components in said at least one sequence of volumetric cardiac imaging data, andgenerating a patient specific 4D geometrical heart model by integrating the individual 4D models for each of said plurality of heart components;generating a patient specific 4D computational model based on one or more of personalized geometry, material properties, fluid boundary conditions, and flow velocity measurements in the 4D geometric model;generating a patient specific biomechanical model based on Fluid Structure Interaction (FSI) simulations using the 4D computational model;extracting patient specific clinical parameters based on the 4D geometric model and the FSI simulations; andperforming disease progression modeling and risk stratification based on the patient specific clinical parameters. 2. The method of claim 1, further comprising: determining morphological and dynamic parameters for any of a plurality of heart components based on the 4D geometric heart model. 3. The method of claim 2, further comprising: determining at least one hemodynamic parameter based on the biomechanical model. 4. The method of claim 3, wherein said at least one hemodynamic parameter comprises at least one of wall shear stress, compliance and stiffness, turbulence intensity, dilation pressure, and full heart blood velocity patterns. 5. The method of claim 1, wherein said step of receiving at least one sequence of volumetric cardiac imaging data of the patient generated using at least one medical imaging modality comprises: receiving at least one of a 4D computed tomography (CT) image sequence and a 4D echocardiography image sequence, and a 4D magnetic resonance imaging (MRI) sequence. 6. The method of claim 1, wherein said step of estimating individual 4D models for a plurality of heart components in said at least one sequence of volumetric cardiac imaging data comprises: estimating individual models of a left ventricle (LV) endocardium, LV epicardium, right ventricle (RV), left atrium (LA), right atrium (RA), mitral valve, aortic valve, aorta, and pulmonary trunk. 7. The method of claim 1, wherein said step of generating a patient specific 4D geometrical heart model by integrating the individual 4D models for each of said plurality of heart components comprises: establishing mesh point correspondence between connecting ones of the individual models. 8. The method of claim 1, wherein: said step of estimating individual 4D models for a plurality of heart components in said at least one sequence of volumetric card... | Siemens Aktiengesellschaft | G06G007/58
US08814841 | B2 | 2014-08-26 | Apparatus and method for wound volume measurement | 2008-10-07 | A method and apparatus for measuring a volume of a wound are described, the method comprising the steps of: applying a dressing over a wound, the volume of which is to be measured, the dressing including at least a sealing drape over the wound so as to create a sealed wound cavity; creating a vacuum in said wound cavity by vacuum pump means so as to produce a predetermined vacuum in the wound cavity; measuring a volume of air extracted from said wound cavity in producing said predetermined vacuum; and, calculating a volume of said wound. | 1. A method of measuring a volume of a wound, the method comprising: creating a vacuum in a wound cavity so as to produce a desired level of vacuum in the wound cavity, wherein the wound cavity is created by placing a sealing drape over a wound;measuring a volume of air extracted from said wound cavity in producing said desired level of vacuum;measuring a leak rate and calculating a volume of said wound based at least in part on the measured volume of air and the measured leak rate, wherein calculating the volume of said wound comprises correcting the wound volume based at least in part on the measured leak rate. 2. The method according to claim 1, wherein the volume of air extracted is measured by a flow meter. 3. The method according to claim 1, wherein volumes of one or more components located in a fluid flow path for creating the vacuum in the wound cavity are measured to produce a correction factor for calculating the volume of said wound. 4. The method according to claim 1, wherein the wound volume is computed by a control module configured to execute software operations. 5. The method according to claim 1, wherein creating the vacuum in the wound cavity comprises generating the vacuum until a steady state level of vacuum in the wound cavity is achieved. 6. The method according to claim 1, wherein the wound volume is calculated at a time of sealing drape change. 7. The method according to claim 1, wherein the wound volume is calculated when a waste canister is empty. 8. The method according to claim 1, further comprising placing the sealing drape in a repeatable manner at each sealing drape change and at each wound volume calculation. 9. The method according to claim 3, wherein the correction factor is subtracted from the measured volume of air. 10. The method according to claim 1, wherein the leak rate is measured by measuring flow of air in a fluid flow path for creating the vacuum in the wound cavity after steady state level of vacuum in the wound cavity has been reached. 11. The method according to claim 10, further comprising computing, based at least in part on the leak rate, a volume of air extracted from the wound cavity due to a leak and correcting the calculated wound volume by subtracting the volume of air extracted from the wound cavity due to the leak from the measured volume of air. | Smith & Nephew PLC | A61M001/00
US09585574 | B2 | 2017-03-07 | Magnetic resonance system and method including an automated non-magnetic medical monitor | 2009-02-02 | In magnetic resonance imaging (MRI), the powerful magnetic fields can interfere with, damage, cause premature failure in, and attract certain non-MR safe instruments. Electromagnetic ally sensitive components are eliminated in favor of MR safe components, such as ceramic piezoelectric components and bi-metallic components. Instruments that previously had to be kept a safe distance away from the main magnet (12), e.g., beyond the 5 Gauss line, while a patient was being scanned are now allowed near the patient without fear of damage to the instrument (40) or danger to the patient and medical staff or the MRI device (10). | 1. A magnetic resonance system comprising: a main magnet for generating a substantially uniform main magnetic field in an examination region;a gradient coil assembly for imposing gradient magnetic fields on the main magnetic field in the examination region, spatially encoding the main magnetic field;a radio frequency assembly for inducing magnetic resonance in selected dipoles of a subject in the examination region, and receiving the magnetic resonance; andan automated non-magnetic monitor disposed in the examination region to measure at least one physiological parameter of the subject, the non-magnetic monitor including one of a blood pressure monitor and a gas analysis monitor, wherein the non-magnetic monitor further includes an electrically operated pump disposed in the main magnetic field. 2. The magnetic resonance system of claim 1 wherein the pump includes a piezoelectric ceramic peristaltic pump. 3. The magnetic resonance system as set forth in claim 2, further including: a high voltage driver for the piezoelectric pump disposed in the examination region. 4. The magnetic resonance system as set forth in claim 1, wherein the monitor includes at least one valve disposed in the main magnetic field. 5. The magnetic resonance system as set forth in claim 4, wherein the at least one valve includes: a piezoelectric ceramic diaphragm valve; anda high voltage driver for the piezoelectric diaphragm valve. 6. The magnetic resonance system as set forth in claim 1, wherein the monitor includes a blood pressure monitor, the blood pressure monitor further including: at least one valve disposed in the main magnetic field;a controller for controlling high voltage drivers for driving the pump and the at least one valve;a memory for storing at least blood pressure readings and a time at which they were taken;an electromagnetically compatible enclosure around the controller and the memory to shield from radiate emissions. 7. The magnetic resonance system as set forth in claim 1, wherein the main magnet generates a magnetic field of at least 2500 Gauss (0.25 T) in the examination region. 8. A method of magnetic resonance comprising: generating a main magnetic field with a main magnet of at least 2500 Gauss of a magnetic resonance imaging device;applying gradient magnetic fields to the examination region with a gradient coil system;applying RF pulses to the examination region with an RF assembly;controlling at least one valve and at least one pump located in the at least 2500 Gauss main magnetic field to alternately inflate and deflate a blood pressure cuff. 9. The method as set forth in claim 8, further including: inflating a pressure cuff located in the main magnetic field by closing a first, linear piezoelectric ceramic diaphragm valve, closing a second, piezoelectric ceramic diaphragm dump valve, and actuating a piezoelectric ceramic pump;reading a pressure by releasing air from the pressure cuff by opening the first valve and turning the pump off;... | KONINKLIJKE PHILIPS N.V. | A61B005/055 A61B005/022 A61B005/021 G01R033/28
US09826905 | B2 | 2017-11-28 | Using colored probes in patient monitoring | 2009-05-07 | The present disclosure provides a sensor with color-coded indications that various patient physiological parameters are being monitored, such as blood oxygen saturation, blood pressure, respiration rate, and respiration effort. The sensor may sense a physical characteristic used to monitor the physiological parameter, and a visible light emitter emits visible light of a first color that is color-coded to the physiological parameter, but is not used to sense the physical characteristic. The visible light emitter may emit visibly flashing light in response to the sensor sensing a threshold value of the physical characteristic. The sensor may include a second light emitter that may sense the physical characteristic, and may emit light of a second color that is color-coded to a first or second physiological parameter. In some embodiments, the first and second colors may visibly mix. The first and second visible light emitters may emit light independently, including visibly flashing light. | 1. A sensor for sensing physical characteristics of a patient, the sensor comprising: a support structure; anda sensor component coupled to the support structure for sensing a physical characteristic of the patient, wherein the sensor component is configured to: generate a physiological signal indicative of the physical characteristic;transmit the physiological signal to a monitor external to the sensor, wherein the physiological signal is capable of being used by the external monitor to monitor first and second physiological parameters;receive a first drive signal from the external monitor indicating when the external monitor is using the physiological signal to monitor the first physiological parameter; andreceive a second drive signal from the external monitor indicating when the external monitor is using the physiological signal to monitor the second physiological parameter, the sensor component comprising:an infrared light emitter, wherein the infrared light is used to sense the physical characteristic;a detector for detecting the infrared light after it passes within the tissue of the patient;a first visible light emitter that emits visible light of a first color, wherein: the first visible light emitter is not used to sense the physical characteristic;the first color is color-coded to the first physiological parameter; andthe first visible light emitter emits visible light in response to the sensor component receiving the first drive signal indicating when the external monitor is using the physiological signal to monitor the first physiological parameter; anda second visible light emitter that emits visible light of a second color, wherein: the second visible light emitter is not used to sense the physical characteristic;the second color is color-coded to the second physiological parameter; andthe second visible light emitter emits visible light in response to the sensor component receiving the second drive signal indicating when the external monitor is using the physiological signal to monitor the second physiological parameter. 2. The sensor of claim 1 wherein the first physiological parameter comprises blood pressure, respiration rate, and/or respiration effort. 3. The sensor of claim 1 wherein the first visible light emitter emits visibly flashing light when the external monitor determines that the physical characteristic sensed by the sensor component for the first physiological parameter has reached a threshold. 4. The sensor of claim 1 wherein the first and second colors mix to appear as a third color that is color-coded to the first physiological parameter and the second physiological parameter. 5. The sensor of claim 4 wherein the third color appears to visibly flash when the external monitor determines that the physical characteristic sensed by the sensor component for the first physiological parameter and the second physiological parameter has reached a threshold. 6. The sensor of claim 1 wherein the second visible light em... | Nellcor Puritan Bennett Ireland | A61B005/00 A61B005/1455 A61B005/024 A61B005/08
US09808163 | B2 | 2017-11-07 | Method and apparatus for heart rate monitoring | 2010-04-27 | A method and apparatus for monitoring heart rate. The method includes receiving a digital heart monitoring signal, determining the integrity of said signal with a “hand detect” signal that confirms electrical connection to the subject, dividing the digital heart monitoring signal into at least one frame, generating a folded correlation value for the center sample in the at least one frame, identify the number of peaks folded correlation values based on amplitude and distance parameters of the digital heart monitoring signal in the at least one frame, removing false peaks and collecting peaks corresponding to a length of time, determining the heart rate based on the identified peaks wherein the identified peaks relate to a minimum distance between valid heart beat peaks based on a heart rate estimate, computing the weighted average of heart rate value based on the number of previous heart rate values. | 1. A method for monitoring a heart rate, comprising: receiving ECG (electrocardiogram) signals;converting the ECG signals to a digital heart monitoring signal as a sequence of heart monitoring samples;receiving at a processor the digital heart monitoring signal;dividing the digital heart monitoring signal into successive overlapping frames, each with an odd number N of samples, and each frame encompassing multiple heart beats;setting amplitude and distance parameters, the distance parameter corresponding to a minimum distance between valid heart beat peaks based on a current estimate of heart rate;processing a current frame to determine a current heart rate using the amplitude and distance parameters by: generating folded correlation values based on a folded correlation of (N−1)/2 samples preceding a center sample, and time-reversed (N−1)/2 samples following the center sample;identifying a number of peaks using the folded correlation values, including removing false peaks using the amplitude parameter, andremoving false peaks using the distance parameter corresponding to a length of time between valid heart beat peaks;determining the current heart rate based on the identified peaks using at least the distance parameter;selectively updating at least the distance parameter in preparation for processing a next frame; andgenerating heart rate signals corresponding to the determined current heart rate. 2. The method of claim 1, further comprising: selectively updating the amplitude parameter in preparation for processing the next frame. 3. The method of claim 1, further comprising: computing a weighted average heart rate based on a predetermined number of previous heart rates, each computed from a respective frame. 4. The method of claim 1, further comprising: restricting a change in heart rate from the current frame to the next frame based on a predetermined timer. 5. An apparatus for monitoring a heart rate, comprising: means for receiving ECG (electrocardiogram) signals;means for converting the ECG signals to a digital heart monitoring signal as a sequence of heart monitoring samples;processing means for processing the digital heart monitoring signal, by: dividing the digital heart monitoring signal into successive overlapping frames, each with an odd number N of samples, and each frame encompassing multiple heart beats;setting amplitude and distance parameters, the distance parameter corresponding to a minimum distance between valid heart beat peaks based on a current estimate of heart rate;processing a current frame to determine a current heart rate using the amplitude and distance parameters, including generating folded correlation values based on a folded correlation of (N−1)/2 samples preceding a center sample, and time-reversed (N−1)/2 samples following the center sample;identifying a number of peaks using the folded correlation values, including removing false peaks using the amplitude parameter, andremoving false peaks using the distanc... | TEXAS INSTRUMENTS INCORPORATED | A61B005/04 A61B005/0245 A61B005/00
US10188295 | B2 | 2019-01-29 | Integrated sensor network methods and systems | 2010-06-01 | Methods and systems for an integrated sensor network are described. In one embodiment, sensor data may be accessed from a plurality of motion sensors and a bed sensor deployed in a living unit for a first time period. An activity pattern for the first time period may be identified based on at least a portion of sensor data associated with the first time period. The activity pattern may represent a physical and cognitive health condition of a person residing in the living unit. Additional sensor data may be accessed from the motion sensors and the bed sensor deployed for a second time period. A determination of whether a deviation of the activity pattern of the first time period has occurred for the second time period may be performed. An alert may be generated based on a determination that the derivation has occurred. In some embodiments, user feedback is captured on the significance of the alerts, and the alert method is customized based on this feedback. Additional methods and systems are disclosed. | 1. A method comprising: processing sensor data from a plurality of motion sensors and a bed sensor deployed in a living unit for a first time period to establish a frame of reference for an activity pattern by a person residing in the living unit, wherein the activity pattern frame of reference represents a health condition of a person residing in the living unit, wherein the processed sensor data for the first time period comprises a plurality of data values corresponding to a plurality of subintervals within the first time period, wherein each subinterval data value represents (1) a total number of sensor hits during that subinterval of the first time period or (2) a total time for sensor firings during that subinterval of the first time period, and wherein the activity pattern frame of reference comprises a statistical value about the subinterval data values over the first time period;computing a standard deviation for the processed sensor data over the first time period based on the subinterval data values;defining a statistical parameter threshold, wherein the statistical parameter threshold comprises a defined number of the standard deviations;processing additional sensor data from the plurality of motion sensors and the bed sensor deployed in the living unit for a second time period, the second time period occurring after the first time period;computing a value for the activity pattern for the second time period based on the processed additional sensor data, the computed activity pattern value of the second time period also representing a health condition of the person, wherein the computed activity pattern value comprises (1) data about a total number of sensor hits during the second time period if the activity pattern frame of reference is based on the total number of sensor hits during the first time period or (2) data about a total time for sensor firings during the second time period if the activity pattern frame of reference is based on the total time for sensor firings during the first time period;comparing the computed activity pattern value with the activity pattern frame of reference and the defined statistical parameter threshold to determine whether the computed activity pattern value deviates from the activity pattern frame of reference by more than the defined number of the standard deviations;generating an alert based on a determination that the computed activity pattern value deviates from the activity pattern frame of reference by more than the defined number of standard deviations;receiving a feedback response to the alert, the feedback response including feedback regarding clinical relevance of the alert; andadjusting the defined number of the standard deviations for the statistical parameter threshold to a different defined number of the standard deviations based on the feedback response to the alert. 2. A method of comprising: accessing sensor data from a plurality of motion sensors and a bed sensor deployed in a livi... | The Curators of the University of Missouri | A61B005/00 A61B005/0205 A61B005/11 G08B021/04 G16H050/20 G06F019/00 A61B005/021 A61B005/024 A61B005/08 G16H015/00
US09492240 | B2 | 2016-11-15 | Virtual measurement tool for minimally invasive surgery | 2009-06-16 | Robotic and/or measurement devices, systems, and methods for telesurgical and other applications employ input devices operatively coupled to tools so as to allow a system user to manipulate tissues and other structures being measured. The system may make use of three dimensional position information from stereoscopic images. Two or more discrete points can be designated in three dimensions so as to provide a cumulative length along a straight or curving structure, an area measurement, a volume measurement, or the like. The discrete points may be identified by a single surgical tool or by distances separating two or more surgical tools, with the user optionally measuring a structure longer than a field of view of the stereoscopic image capture device by walking a pair of tools “hand-over-hand” along the structure. By allowing the system user to interact with the tissues while designating the tissue locations, and by employing imaging data to determine the measurements, the measurement accuracy and ease of measurement may be enhanced. | 1. A surgical system for measuring tissue, the system comprising: a first robotic tool having a tool tip;a stereoscopic image capture device orientable toward a surgical site for capturing left and right stereoscopic images of the tissue at the surgical site and capturing left and right stereoscopic images of the first robotic tool, the image capture device in use generating stereoscopic image data corresponding to the left and right stereoscopic images of the tissue at the surgical site and the first robotic tool;a stereoscopic display coupled to the image capture device so as to show the captured left and right stereoscopic images of the tissue at the surgical site and the captured left and right stereoscopic images of the first robotic tool to a system user;an input for designating tissue locations indicated by the first robotic tool, wherein the input is coupled to a master controller coupled with the first robotic tool for receiving movement commands from the system user to guide the first robotic tool to designate a first tissue location, with the first robotic tool tip disposed over the first tissue location as seen from a dominant eye of the system user, without touching the first tissue location, the first tissue location having one or more first points of interest within both of the left and right stereoscopic images of the tissue at the surgical site shown on the stereoscopic display; anda processor configured to track Cartesian position of the first robotic tool tip and to match an image of the one or more points of interest within one of the left and right stereoscopic images of the tissue at the surgical site to an image of the one or more points of interest within the other of the left and right stereoscopic images of the tissue at the surgical site, based upon local characteristics of the one or more points of interest, so as to determine a three-dimensional location of the tissue that corresponds to the matched one or more points of interest, wherein the three-dimensional location corresponds to the first designated tissue location indicated by the first robotic tool, and to superimpose a first graphical indicator on the stereoscopic images of the tissue at the determined three-dimensional location of the tissue that corresponds to the matched points of interest, without touching the first designated tissue location, the processor being further configured for generating a measurement between at least the first graphical indicator and a second designated tissue location;wherein the first graphical indicator is superimposed on the stereoscopic images of the tissue at the determined three-dimensional location that corresponds to the matched one or more points of interest, after the first robotic tool is moved to designate the first designated tissue location. 2. The system of claim 1, wherein the processor superimposes a graphical tool measurement marker on the stereoscopic images of the tissue at the surgical site coincident with ... | Intuitive Surgical Operations, Inc. | A61B017/00 A61B001/04 A61B001/00
US20100298654 | A1 | 2010-11-25 | BLOOD PRESSURE-MONITORING SYSTEM WITH ALARM/ALERT SYSTEM THAT ACCOUNTS FOR PATIENT MOTION | 2009-05-20 | The invention provides a system and method for measuring vital signs (e.g. SYS, DIA, SpO2, heart rate, and respiratory rate) and motion (e.g. activity level, posture, degree of motion, and arm height) from a patient. The system features: (i) first and second sensors configured to independently generate time-dependent waveforms indicative of one or more contractile properties of the patient's heart; and (ii) at least three motion-detecting sensors positioned on the forearm, upper arm, and a body location other than the forearm or upper arm of the patient. Each motion-detecting sensor generates at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which it is affixed. A processing component, typically worn on the patient's body and featuring a microprocessor, receives the time-dependent waveforms generated by the different sensors and processes them to determine: (i) a pulse transit time calculated using a time difference between features in two separate time-dependent waveforms, (ii) a blood pressure value calculated from the time difference, and (iii) a motion parameter calculated from at least one motion waveform. | 1. A system for measuring vital signs from a patient, comprising: (a) a first sensor configured to generate a first time-dependent waveform indicative of one or more contractile properties of the patient's heart;(b) a second sensor configured to generate a second time-dependent waveform indicative of one or more contractile properties of the patient's heart;(c) at least two motion-detecting sensors positioned on a location selected from a forearm, upper arm, and a body location other than the forearm or upper arm of the patient, each of said motion-detecting sensors generating at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which the motion-detecting sensor is affixed; and,(d) a processing component configured to be worn on the patient's body and comprising a microprocessor, said processing component configured to receive the first time-dependent waveform, the second time-dependent waveform, and the at least one time-dependent motion waveform generated by each motion-detecting sensor and determine therefrom: (i) one or more vital signs calculated using the first and second time-dependent waveforms; (ii) at least one motion parameter selected from the patient's posture, activity level, arm height, and degree of motion and calculated using at least one motion waveform generated by each of said motion-detecting sensors; and (iii) at least one alarm condition calculated using one or more of the vital signs and at least one motion parameter, wherein an alarm condition indicated by a variance in one or more of the vital signs is regulated according to one or more motion parameters. 2. A system according to claim 1, wherein: said first time-dependent sensor is an optical sensor comprising: a source of electromagnetic radiation configured to irradiate tissue of the patient with radiation emitted therefrom, and a detector configured to detect one or more properties of the electromagnetic radiation after it irradiates said tissue, wherein the first time-dependent waveform is an optical waveform indicative of volumetric changes in the irradiated tissue caused by ventricular contraction of the patient's heart; andsaid second time-dependent sensor is an electrical sensor comprising: at least two electrodes configured to detect electrical signals from the patient's body, and an electrical circuit operably connected to the electrodes and configured to process the detected electrical signals, wherein said second time-dependent waveform is an electrical waveform indicative of ventricular depolarization of the patient's heart. 3. A system according to claim 2, wherein the optical waveform is generated by a pressure waveform resulting from ejection of blood from the left ventricle. 4. A system according to claim 3, wherein the pulse transit time is a time difference between an electrocardiogram QRS complex and a corresponding inflection point in said optical waveform. 5. A system according to claim 1, wh... | TRIAGE WIRELESS, INC. | A61B005/0205 A61B005/11
US20100324394 | A1 | 2010-12-23 | Analyte Monitoring Device and Methods of Use | 2010-08-12 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for processing sensor data from a continuous analyte sensor, comprising: receiving sensor data from a continuous analyte sensor;processing the sensor data; andproviding output based at least in part on the sensor data, wherein the sensor data is displayed on a graph. 3. The method of claim 2, wherein providing output comprises displaying the output in real-time. 4. The method of claim 2, wherein processing the sensor data comprises prospectively calibrating the sensor data, and wherein the sensor data displayed comprises a prospectively calibrated sensor data point. 5. The method of claim 2, wherein processing the sensor data comprises calibrating the sensor data, and wherein the sensor data displayed on the graph comprises a calibrated sensor data point. 6. The method of claim 2, wherein the continuous analyte sensor is a continuous glucose sensor. 7. A system for processing sensor data from a continuous analyte sensor, comprising: an input module operatively connected to a continuous analyte sensor so as to receive sensor data from the continuous analyte sensor;a processor module that processes the sensor data; andan output module that provides output based on the sensor data, wherein the output comprises sensor data displayed on a graph. 8. The system of claim 7, wherein output is displayed in real-time. 9. The system of claim 7, wherein the processor module prospectively calibrates the sensor data, and wherein the sensor data displayed on the graph is prospectively calibrated sensor data. 10. The system of claim 7, wherein the processor module calibrates the sensor data, and wherein the sensor data displayed on the graph is calibrated sensor data. 11. The system of claim 7, wherein the continuous analyte sensor is a continuous glucose sensor. | Abbott Diabetes Care Inc. | A61B005/145
US07826893 | B2 | 2010-11-02 | Method and apparatus for generating a template for arrhythmia detection and electrogram morphology classification | 2007-06-04 | An implantable medical device and associated method for automatically generating morphology templates during fast cardiac rhythms, confirming a provisional template as a confirmed template, and using the confirmed template to classify subsequent detected arrhythmias. A provisional SVT template may be created during a fast ventricular rate and activated as a confirmed SVT template upon verification that the fast rate was due to an SVT. The confirmed SVT template may be used to discriminate SVT from VT/VF. | 1. A method of discriminating with a medical device among individual ones of a plurality of cardiac complexes of a heart of a patient, comprising the steps of: detecting a fast cardiac rate wherein the plurality of cardiac complexes are occurring at a rate corresponding to an arrhythmia detection interval;sensing the plurality of cardiac complexes during the fast cardiac rate;iteratively comparing each individual one of the plurality of cardiac complexes with each other one of the plurality of cardiac complexes to obtain a plurality of comparisons, each individual one of the plurality of comparisons corresponding to one of the plurality of cardiac complexes;identifying one of the plurality of cardiac complexes as being representative of the plurality of cardiac complexes based, at least in part, on the plurality of comparisons; andpresenting the one of the plurality of cardiac complexes representative of the plurality of cardiac complexes to a user. 2. The method of claim 1, wherein the iteratively comparing step comprises: transforming each of the plurality of cardiac complexes into a plurality of corresponding signal wavelet coefficients;identifying higher amplitude ones of each of the plurality of corresponding signal wavelet coefficients;wherein each of the plurality of comparisons are based, at least in part, on comparisons of the higher amplitude ones of each of the plurality of signal wavelet coefficients. 3. The method of claim 2, wherein transforming each of the plurality of cardiac complexes comprises transforming each of the plurality of cardiac complexes using a wavelet transform to obtain the plurality of corresponding signal wavelet coefficients. 4. The method of claim 1, further comprising: identifying the fast cardiac rate as a supraventricular tachycardia;identifying the one of the plurality of cardiac complexes representative of the plurality of cardiac complexes as a supraventricular tachycardia template based, at least in part, on the fast cardiac rate being a supraventricular tachycardia. 5. A method for monitoring heart rhythms comprised of a plurality of cardiac complexes with a medical device, comprising the steps of: detecting a fast cardiac rate based, at least in part, on the plurality of cardiac complexes occurring at a rate corresponding to an arrhythmia detection interval;sensing first individual ones of the plurality of cardiac complexes during the fast cardiac rate;iteratively comparing each individual one of the first individual ones of the plurality of cardiac complexes with each other one of the first individual ones of the plurality of cardiac complexes to obtain a first plurality of comparisons, each individual one of the first plurality of comparisons corresponding to one of the first individual ones of the plurality of cardiac complexes;identifying one of the first individual ones of the plurality of cardiac complexes as being representative of the first individual ones of the plurality of cardiac comple... | Medtronic, Inc. | A61B005/04 A61B005/0472
US07837633 | B2 | 2010-11-23 | Lancing device and multi-lancet cartridge | 2005-06-30 | A medical lancing device including a replaceable multi-lancet cartridge. The lancing device includes a drive mechanism, an activation mechanism, and an advancing mechanism. The advancing mechanism includes a linear-pull slide that is moved in and out to operate an indexing ratchet mechanism, a cam-guided charger mechanism, and a cam-guided lancet cap displacement mechanism. The indexing ratchet mechanism sequentially advances the lancets in the cartridge to an active position. The cam-guided charger mechanism charges the drive mechanism and separates the cap from the active lancet. And the cam-guided cap displacement mechanism moves the separated cap from the lancing stroke path of the active lancet. The activation mechanism then releases the charged active lancet to traverse the unobstructed lancing stroke path to pierce the subject's skin at a desired lancing site. | 1. A lancing device for use with a cartridge assembly having an array of lancets, each lancet including a detachable cap, the lancing device comprising: a housing for receiving the cartridge assembly;an indexing ratchet mechanism adapted to sequentially advance the lancets into an active lancet position;a drive mechanism having a drive member adapted to engage and retract an active-positioned one of the lancets to detach the corresponding active-positioned cap, and to drive the active-positioned lancet along a lancing stroke;a charger mechanism adapted to retract the drive member;a cap displacement mechanism adapted to engage and move the active-positioned, detached lancet cap from the lancing stroke path of the active-positioned lancet; anda slider that is slidable relative to the housing and adapted to actuate the indexing ratchet mechanism, the charger mechanism, and the cap displacement mechanism as the slider is moved between a first position and a second position. 2. The lancing device of claim 1, wherein the cartridge assembly includes ratchet teeth, and wherein the indexing ratchet mechanism includes a resiliently deflectable pawl extending from the slider that engages the ratchet teeth to advance the lancets in the cartridge to the active position. 3. The lancing device of claim 1, wherein the charger mechanism includes a resiliently deflected cam arm that extends from the slider and is adapted to retract the drive member. 4. The lancing device of claim 3, wherein the drive member has a charging follower and the cam arm is resiliently deflected by the charging follower when the slider is moved from the first position toward the second position. 5. The lancing device of claim 4, wherein the cam arm is moved past the charging follower when the slider reaches the second position, and the cam arm returns to its neutral position. 6. The lancing device of claim 5, wherein the cam arm engages the charging follower and retracts the drive member as the slider is moved from the second position back to the first position. 7. The lancing device of claim 3, wherein the drive member has a locking follower and the slider defines a locking wall that blocks a travel path of the drive member locking follower except when the slider is in the second position. 8. The lancing device of claim 3, wherein the lancet cap is detached from the active lancet when the drive member retracts the lancet. 9. The lancing device of claim 8, wherein the slider includes a cam surface and the cap displacement mechanism includes a lifter with a follower that is guided by the cam surface to move a separated lancet cap from the lancing stroke path of the active lancet when the slider is moved from the second position to the first position. 10. The lancing device of claim 1, wherein the slider includes a cam surface and the cap displacement mechanism includes a lifter with a follower that is guided by the cam surface to move a separated lancet cap from the lancing stroke... | Facet Technologies, LLC | A61B005/00 A61B017/32 A61B017/14 B65D081/00
US20100331733 | A1 | 2010-12-30 | SENSING DEVICE AND METHOD FOR AN ORTHOPEDIC JOINT | 2010-03-26 | One embodiment is directed to a sensor for measuring a skeletal system. A signal path of the system comprises an amplifier (612), a sensor element, and an amplifier (620). The sensor element comprises a transducer (4), a waveguide (5), and a transducer (30). An external condition is applied to the sensor element. For example, the sensor element is placed in an artificial orthopedic joint to measure loading of the joint. Pulsed energy waves are emitted by the transducer (4) into the waveguide (5). The transducer (30) receives each pulsed energy wave after it propagates through the waveguide (5). The transit time of each pulsed energy wave corresponds to the external condition applied to the sensor. The transducer (30) outputs a signal corresponding to each pulsed energy wave. A detection circuit edge detects the signal and outputs a pulse to the transducer (4) to generate a new pulse energy wave. | 1. A pulsed-loop mode measurement system comprising: one or more sensing assemblies;a pulsed system;one or more load surfaces; and electronic circuitry, where the pulsed system maintains positive closed-loop feedback of pulsed energy waves in one or more energy propagating structures of the sensing assembly and where the system measures parameters of the muscular-skeletal system. 2. The system of claim 1, where the pulsed system modulates a time period of pulsed energy waves as a function of changes in distance or velocity through a medium of the one or more energy propagating structures, or a combination of changes in distance and velocity, caused by changes in the one or more energy propagating structures. 3. The system of claim 1, further comprising a pulse shaper to dampen a wave shape for optimal transmission and reception in accordance with a matched network. 4. The system of claim 1, further comprising a digital block for digitizing the frequency of operation of the pulsed system. 5. The system of claim 1 where the pulsed system is configured to operate wireless in pulsed-loop mode according to one or more operational criteria, such as, but not limited to, power level, applied force level, standby mode, application context, temperature, or other parameter level. 6. The system of claim 5, where the system operates to measure changes in propagation time due to changes in the length of one or more waveguides coupled to the one or more load surfaces such that the physical length change under load are in proportion to the applied force. 7. A sensor module comprising one or more sensors for sensing a muscular-skeletal system each sensor comprising: a first transducer;a waveguide having a first surface and a second surface where the first transducer couples to the first surface of the waveguide; anda second transducer coupled to the second surface of the waveguide where pulsed energy waves propagate through the waveguide and where a transit time of a pulsed energy wave through the wave guide corresponds to one or more measured parameters of the muscular-skeletal system. 8. The sensor module of claim 7 where a change in length of the waveguide results in a corresponding change in the transit time of the pulsed energy wave and where the transit time or a change in transit time in conjunction with material properties of the waveguide corresponds to the one or more measured parameters. 9. The sensor module of claim 7 where each pulsed energy wave is detected after propagating through the waveguide, where a pulse is generated when each pulsed energy wave is detected, and where the pulse is coupled to the first transducer. 10. The sensor module of claim 9 further including: a first amplifier having an input and an output coupled to an input of the first transducer; anda second amplifier having an input coupled to an output of the second transducer and an output coupled to the input of the first amplifier. 11. The sensor module of claim 7 wher... | OrthoSensor | A61B005/103
US07841098 | B2 | 2010-11-30 | Linear measurement apparatus and method | 2008-06-30 | A linear measurement device for measuring a wheelchair user for fitting a wheelchair, with a measuring bar having with a scale and a fixed locater attached at a first end. A second locator opposing the first locater moves on the measuring bar to a selected position for determining a distance measurement relative to the fixed locator for fitting the wheelchair user to the wheelchair. A method of measuring for wheelchair fitting with a linear measurement apparatus is disclosed. | 1. A linear measurement device for obtaining measurements for fitting a wheelchair user to a wheelchair, comprising: a measuring bar provided with a scale having a plurality of spaced-apart demarcations and associated numerical indicia;a fixed locater attached at a first end of the measuring bar;a second locator attached to a slider carried on the measuring bar, the second locater opposing the fixed locater and having an indicator thereon that aligns with a portion of the scale;a pair of pressure pads, each received on a respective one of the fixed and second locators, for contacting a selected body surface for determining a distance measurement, the pressure pad selectively attached to the second locater to position facing towards or away from the opposing pad on the fixed locator,whereby the second locator moves relative to a longitudinal axis of the measuring bar to a selected position for determining a distance measurement relative to the fixed locator for fitting the wheelchair user to the wheelchair. 2. The linear measurement device as recited in claim 1, wherein the slider comprises a housing having two opposing pairs of aligned wheels, each pair of aligned wheels rotatably mounted in the housing and bearing on a respective side wall of the measuring bar, the rollers guiding the movement of the slider relative to the measuring bar. 3. The linear measurement device as recited in claim 1, wherein the pressure pads have a planar bearing surface and a means for attaching the pressure pad to the locator. 4. The linear measurement device as recited in claim 3, wherein the means for attaching comprises a channel in the pressure pad for selectively receiving a portion of the locator. 5. The linear measurement device as recited in claim 4, wherein the pressure plate includes a pair of spaced-apart opposing walls that define the channel. 6. The linear measurement device as recited in claim 1, wherein the fixed locater comprises a base for resting on a surface that receives the wheels of the wheelchair, for taking a vertical measurement upon moving the second locator in contact with a selected portion of the wheelchair or the wheelchair user. 7. The linear measurement device as recited in claim 2, wherein the housing for the slider defines a window for observing therein the numerical indicia. 8. The linear measurement device as recited in claim 1, where in the indicator comprises an elongated member having an edge that aligns with the portion of the scale. 9. The linear measurement device as recited in claim 1, wherein the indicator comprises a light-emitting device that emits a light beam for aligning with a selected surface for determining a distance relative to the fixed locator. 10. The linear measurement device as recited in claim 1, further comprising means for a digital display on the slider for the measured value. 11. A method of measuring for fitting a wheelchair user to a wheelchair, comprising the steps of: (a) providing a measur... | Max Mobility, LLC | A61B005/103 G01B001/00 G01B005/14
US07831450 | B2 | 2010-11-09 | Medical order information display system | 2001-09-05 | A medical information system receives patient data and information from various sources and displays such information in a variety of formats for use by member of a medical team in a hospital, clinic, or office. The system receives patient information from doctors, pharmacists, patient monitoring equipment, testing laboratories, and/or computer databases. Access to selected subsets of patient information is provided by user selection of specific data sets identified by job function selection icons. A member of the medical team can record observations about a patient using key words and phrases which can be supplemented with additional text for customized notation. Multiple types of patient data are selectively displayed simultaneously, and to multiple remote users. The system can access stored data according to user-specified formulae to compute a score or metric which reflects a relationship between various factors, each factor being weighted appropriately according to its significance as defined in the formula. A user can selectively display data in graphic form by “clicking” on a row of tabular data in a tabular region of the display and “dragging and dropping” that row to a graphic display region of the display. | 1. A medical order information display system, comprising: A. a display device responsive to signals to generate a medical data image;B. storage device for storing data representative of patient orders,C. order device responsive to user action to generate and store order data in said storage device, said order data being representative of patient orders, said order data for a patient including: i. order data representative of the identity of a prescriber of a prescription,ii. i.d. data representative of the identification of said patient,iii. drug data representative of a drug prescribed for a patient in accordance with the prescription,iv. administration data representative of prescribed dosage and administration times in accordance with said prescription, andv. authentication data for preventing unauthorized modification of said order data and preventing entry of non-complying order data, wherein said non complying order data is data that is not in accordance with said prescription,wherein said order data further includes authentication data representative of an encrypted identifier uniquely associated with said prescription, and said order device further includes means for generating by a user an access request associated with said authentication data for a selected one of said prescriptions,D. a display controller for generating said display signals, including: i. selectively operative means for generating said display signals wherein at least information representative of said i.d. data, said drug data, and said administration data is displayed on said display device,ii. selectively operative administrator device responsive to user action to generate and store in said storage device, action data representative of an action taken with respect to said patient and one of said prescriptions, said action data for one of said prescriptions being representative: a. the delivery of a drug to said patientb. the identity of said drug delivered to said patient, andc. the dose and time of said drug delivered to said patient,iii. monitor means for monitoring said order data for one or more of said patients and for detecting when action data does not correspond to correct fulfillment of said order data and responsive thereto, generating a display signal representative of an alarm,wherein said display controller includes means for confirming that said access request corresponds to said authentication data for said prescription and upon such confirmation, enabling said order device to modify said data for said patient and for preventing modification of said order data otherwise. 2. The system according to claim 1 wherein said monitor means generates in conjunction with said alarm display signal, an additional alarm from the group consisting of audible alarm, visual alarm and data alarm. 3. A medical information calculator display system, comprising: A. a display device responsive to display signals to generate a medical data image,B. a display controller ... | I.M.D. Soft Ltd. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07826891 | B2 | 2010-11-02 | Method and apparatus for trending a physiological cardiac parameter | 2006-09-07 | The present invention relates to an implantable cardioverter-defibrillator or pacemaker whose standard circuitry is used to trend a physiological cardiac parameter using intra-cardiac impedance measurements. The trend information may be used to predict the onset of a sudden cardiac death (SCD) event. By being able to predict the onset of an SCD event, patients and their physicians may be forewarned of a life-threatening event allowing them to respond accordingly. The trend information may also be used to predict the efficacy of cardiac-related medications, monitor progress of congestive heart failure, detect the occurrence of myocardial infarction, or simply track changes in sympathetic tone. | 1. A system for detecting myocardial infarction, comprising: a measuring device that measures intra-cardiac impedance;a derivation module that derives a pre-ejection period from the measured impedance;a trending module that trends the derived pre-ejection period over spaced time intervals to create trended pre-ejection period data; andan analyzing module that analyzes the trended pre-ejection period data to and detects a myocardial infarction from said analysis. 2. The system of claim 1, comprising comparing trended pre-ejection period data and detecting a difference between the trended pre-ejection period data. 3. The system of claim 1, further comprising a reporting module that reports the trended pre-ejection period data to an outside source. 4. The system of claim 3, wherein the reporting module reports trended pre-ejection period data that detects a myocardial infarction. 5. The system of claim 1, wherein the derivation module and the trending module are packaged with an implantable measuring device. 6. The system of claim 5, wherein the package is capable of being implanted in a human body. 7. The system of claim 1, further comprising an external device for storing the trended pre-ejection period data. 8. The system of claim 1, wherein the pre-ejection period correlates to sympathetic and parasympathetic activity. 9. The system of claim 1, wherein the system downloads the trended pre-ejection period data to a separately located storage device. 10. The system of claim 1, wherein the measuring device measures intracardiac impedance by applying a current between two electrodes and measuring a resulting voltage that is used to calculate the cardiac impedance. 11. The system of claim 10, wherein the electrodes are part of at least one unipolar lead and a remote device. 12. The system of claim 10, wherein the electrodes are part of at least one bipolar lead. 13. The system of claim 10, wherein the electrodes are part of at least one unipolar lead and a bipolar lead. 14. The system of claim 10, wherein the electrodes are part of at least one bipolar lead and a remote device. 15. A device for trending a pre-ejection period, comprising: an impedance module that measures an intra-cardiac impedance at spaced time intervals;a derivation module that derives the pre-ejection period using the measured impedance;a trending module that generates trended pre-ejection period data using the derived pre-ejection period; andan analyzing module that analyzes the trended pre-ejection period data and detects a myocardial infarction from said analysis. 16. The system of claim 1, wherein the measuring device is configured to measure intra-cardiac impedance using an implantable cardiac rhythm management device. 17. The system of claim 1, wherein the analyzing module is configured to provide an indication of the detected myocardial infarction to a device or user. 18. The system of claim 8, wherein the pre-ejection period comprises shortened pre-ejectio... | Cardiac Pacemakers, Inc. | A61B005/053
US07857768 | B2 | 2010-12-28 | Integrated tail mounted blood pressure monitor and pulse oximeter system for animal research | 2008-10-10 | An integrated tail mounted blood pressure monitor and pulse oximeter system includes a tail blood flow occlusion member configured to selectively occlude blood flow through the tail; a sensor coupled to the tail blood flow occlusion member detecting a degree of operation thereof; Light sources coupled to the tail closer to the distal end of the tail than the tail blood flow occlusion member, and selectively directing light of two different wavelengths into the tail; a light receiver coupled to the tail and selectively receiving a signal associated with light directed into the tail from the light sources; and a controller configured to selectively determine blood pressure parameters from the data and pulse oximeter parameters from the data. | 1. An integrated tail mounted blood pressure monitor and pulse oximeter system comprising: A tail blood flow occlusion member configured to be secured to a subject's tail and selectively occlude blood flow through the tail;A sensor coupled to the tail blood flow occlusion member configured to detect a degree of operation of the tail blood flow occlusion member;Light sources configured to be coupled to the tail in a position closer to the distal end of the tail than the position of the tail blood flow occlusion member, and configured to selectively direct light of at least two different wavelengths into the tail;At least one Light receiver configured to be coupled to the tail in a position closer to the distal end of the tail than the position of the tail blood flow occlusion member, and configured to selectively receive a signal associated with light that has been directed into the tail from the light sources;Controller coupled to the tail blood flow occlusion member for controlling the tail blood flow occlusion member, and coupled to the sensor and the at least one light receivers for receiving data there from, wherein the controller is configured to selectively determine blood pressure parameters from the data and pulse oximeter parameters from the data, wherein said pulse oximeter parameters include breath rate, pulse distention, and breath distention. 2. The integrated tail mounted blood pressure monitor and pulse oximeter system according to claim 1 wherein the tail blood flow occlusion member includes at least one inflatable tail cuff portion. 3. The integrated tail mounted blood pressure monitor and pulse oximeter system according to claim 1 wherein the tail blood flow occlusion member includes two housing halves that are selectively movable toward and away from each other. 4. The integrated tail mounted blood pressure monitor and pulse oximeter system according to claim 3 wherein a spring bias holds the two housing halves in a relaxed, non-blood flow occluding position. 5. The integrated tail mounted blood pressure monitor and pulse oximeter system according to claim 3 wherein the movement of the housing halves toward each other is configured to selectively cut off blood flow through the tail. 6. The integrated tail mounted blood pressure monitor and pulse oximeter system according to claim 1 wherein the tail blood flow occlusion member includes an inflatable tail cuff portion that is wrapped around the tail. 7. The integrated tail mounted blood pressure monitor and pulse oximeter system according to claim 1 further including an animal holder for containing the animal and wherein the tail blood flow occlusion member is secured to the holder. 8. An integrated blood pressure monitor and pulse oximeter system comprising: A blood flow occlusion member configured to be secured to a subject's appendage and selectively occlude blood flow through the appendage;A sensor coupled to the blood flow occlusion member configured to detect a degre... | Starr Life Sciences Corporation | A61B005/0205 A61B005/1455 A61B005/022
US07833160 | B2 | 2010-11-16 | Blood viscosity measurement device | 2006-01-27 | A blood viscosity measurement device has a flow rate calculating section that calculates a flow rate of blood flowing through a blood vessel of a living body. An inner diameter calculating section calculates an inner diameter of the blood vessel. A blood pressure data acquiring device acquires blood pressure data corresponding to a blood pressure of the blood vessel. A viscosity calculating section calculates a viscosity of the blood flowing through the blood vessel using each of the calculated flow rate, the calculated inner diameter, and the acquired blood pressure data. | 1. A blood viscosity measurement device comprising: flow rate measurement means for measuring a flow rate of blood flowing through a blood vessel of a living body;inner diameter measurement means for measuring an inner diameter of the blood vessel corresponding to the measured flow rate;blood pressure acquiring means for acquiring a blood pressure of the blood vessel at a time different from a time when the flow rate of the blood flowing through the blood vessel and the inner diameter of the blood vessel are measured; andviscosity determining means for determining a viscosity of the blood flowing through the blood vessel using maximum values of the measured flow rate, the measured inner diameter, and the acquired blood pressure of the blood vessel. 2. A blood viscosity measurement device comprising: a flow rate measurement section that measures a flow rate of blood flowing through a blood vessel of a living body;an inner diameter measurement section that measures an inner diameter of the blood vessel corresponding to the measured flow rate;a blood pressure acquiring section that acquires a blood pressure of the blood vessel corresponding to the measured flow rate;a viscosity determining section that determines a viscosity of the blood flowing through the blood vessel using maximum values of the measured flow rate, the measured inner diameter, and the acquired blood pressure of the blood vessel;a surge transmission element that selectively transmits one of continuous waves and pulse waves from a surface of the living body into the living body; anda circuit unit that selectively generates the continuous waves and the pulse waves in the surge transmission element;wherein when the circuit unit switches between the generation of the continuous waves and the pulse waves in the surge transmission element, the surge transmission element transmits the continuous waves to be used by the flow rate measurement section and the pulse waves to be used by the inner diameter measurement section. 3. A blood viscosity measurement device comprising: a flow rate processing section that calculates a flow rate of blood flowing through a blood vessel of a living body;an inner diameter processing section that calculates an inner diameter of the blood vessel;a blood pressure data acquiring device that acquires blood pressure data corresponding to a blood pressure of the blood vessel; anda viscosity processing section that calculates a viscosity of the blood flowing through the blood vessel using a maximum values of the calculated flow rate, the calculated inner diameter, and the acquired blood pressure data;wherein the blood pressure data acquiring device acquires the blood pressure data separately and independently from the calculation of the flow rate of blood by the flow rate processing section and the calculation of the inner diameter of the blood vessel by the inner diameter processing section; andwherein the blood pressure data acquiring device acquires the blood ... | Seiko Instruments Inc. | A61B008/00 A61B005/00
US07857760 | B2 | 2010-12-28 | Analyte sensor | 2006-02-22 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A sensor system for measuring an analyte concentration in a host, the system comprising: at least one electrode configured for implantation in a host and configured to measure an analyte concentration in a tissue of the host;sensor electronics operably connected to the at least one electrode and configured to provide analyte data representative of an analyte concentration in the host, wherein the sensor electronics are configured to provide an identification of sensor insertion into the tissue of the host or an identification of sensor removal from the tissue of the host; andan information tag comprising a code, wherein the code comprises sensor implantation information configured to control or limit an amount of time after sensor insertion into the tissue of the host over which information can be obtained from the at least one electrode, and wherein the code further comprises a number that represents how many sensor insertions into the tissue of the host are allowed using a unique license code. 2. The sensor system of claim 1, wherein the information tag comprises a memory. 3. The sensor system of claim 2, wherein the information tag transmits information using at least one connection selected from the group consisting of a serial connection, a radio frequency connection, an acoustic frequency connection, an infrared frequency connection, and a magnetic induction connection. 4. The sensor system of claim 1, further comprising: a mounting unit configured to maintain the sensor positioned transcutaneously within the tissue of the host. 5. The sensor system of claim 4, wherein the information tag is embedded within the mounting unit. 6. The sensor system of claim 1, further comprising a receiver configured to receive the analyte data from the sensor electronics. 7. The sensor system of claim 6, wherein the receiver is configured to read the code from the information tag. 8. The sensor system of claim 1, further comprising packaging configured to contain at least a portion of the sensor system during transport. 9. The sensor system of claim 8, wherein the information tag is in or on the packaging. 10. The sensor system of claim 1, wherein the code further comprises at least one item selected from the group consisting of manufacturing information, calibration information, identification information, expiration information, and archived data. 11. An analyte sensor assembly, the assembly comprising: a mounting unit adapted for mounting on a skin of a host;an electronics unit configured to releasably mate with the mounting unit, wherein the electronics unit is configured to provide an identification of sensor insertion into a tissue of the host or an identification of sensor removal from the tissue of the host;a sensor configured to measure a concentration of an analyte in the host, wherein the sensor is operably connected to the electronics unit when the electronics unit is mated to the mounting unit; andan information tag comprising a cod... | DexCom, Inc. | A61B005/00
US07828743 | B2 | 2010-11-09 | Medical body core thermometer | 2006-11-29 | A temperature sensing device accurately determines the core temperature of a warm blooded animal or human based on at least one measurement of the temperature of the skin of the warm blooded animal or human. The device includes a housing, and a first contact type temperature sensing element coupled to the housing. The first contact type temperature sensing element includes a first temperature sensor that is operative to measure the temperature of the skin when the first contact type temperature sensing element is in contact with the skin. The first temperature sensor produces at least a first signal representative of the measured skin temperature. An electronic circuit uses the first signal to determine the core temperature of the warm blooded animal or human. An electronic communication device, such as a display, is coupled to the electronic circuit for communicating the core temperature to a user. | 1. A temperature sensing device operative to determine the body temperature of a warm blooded animal or human based on at least one measurement of the temperature of the skin of the warm blooded animal or human, the device comprising: a housing;a first contact type temperature sensing element being adapted for contacting the skin and including a first temperature sensor, the first temperature sensor operative to measure the temperature of the skin when the first contact type temperature sensing element is in contact with the skin and produce at least a first signal;a moveable element coupled to said housing and carrying the first contact type temperature sensing element, the moveable element configured to be moved relative to the housing into at least first and second positions, the first position being a position at which the first contact type temperature sensing element is adapted to not contact the skin and the second position being a position at which the first contact type temperature sensing element is adapted for contact with the skin for consistent contact with the skin;a guard coupled to the housing and being moveable relative to the moveable element, the guard shielding the first contact type temperature sensing element, the guard configured to retract proximally relative to the first contact type temperature sensing element to unshield the first contact type temperature sensing element while the guard is pressed against the skin to bring the first contact type temperature sensing element in thermal contact with the skin;an electronic circuit electrically coupled with the first contact type temperature sensing element and operative to use the first signal to determine the body temperature of the warm blooded animal or human; andan electronic communication device coupled to the electronic circuit and operative to communicate the body temperature to a user. 2. The device of claim 1, further comprising a thermal insulator positioned adjacent the first temperature sensor. 3. The device of claim 2, further comprising: a second temperature sensor coupled to the housing, the thermal insulator positioned generally between the first and second temperature sensors so as to thermally decouple the first and second temperature sensors from each other, and the second temperature sensor further being positioned so as to be thermally decoupled from the skin during measurement of the skin with the first temperature sensor, the second temperature sensor operative to detect a reference temperature represented by at least a second signal, and wherein the electronic circuit uses the first and second signals to accurately determine the body temperature. 4. The device of claim 1, wherein the moveable element further comprises a shaft formed from a thermally insulating material. 5. The device of claim 4, wherein the shaft is spring loaded to normally bias the first contact type temperature sensing element toward the second position. 6. The device of clai... | Advanced Monitors Corporation | A61B005/00
US07841985 | B2 | 2010-11-30 | Sensor identification method and system | 2006-12-22 | An inventive method and system are provided for identifying a compatible photoplethysmographic sensor when interconnected to a given photoplethysmographic monitor. The method and system may further provide for the identification of which of a plurality of compatible sensors is interconnected to allow for selective calibration of the photoplethysmographic monitor. In this regard, the system and method may entail provision of a predetermined drive or test signal to a light source and/or identification element (i.e., sensor elements) of a photoplethysmographic sensor, and the obtainment of a corresponding output signal for use in sensor identification. In one approach, an output signal is obtained from at least one of three sensor elements each of which is interconnected between a different pair of sensor terminals which is a unique combination of two of four sensor terminals on the sensor. In this regard, individual circuits may be formed through each sensor element allowing for enhanced flexibility in signal application. | 1. A method using a photoplethysmographic system, comprising: releasably interconnecting four terminals of a photoplethysmographic sensor to a photoplethysmographic monitor,wherein a first pair of said four terminals on said sensor is connected for receiving first drive signals from said monitor for illuminating a first light source coupled between said first pair of terminals,wherein a second pair of said four terminals on said sensor is connected for receiving second drive signals from said monitor for illuminating a second light source coupled between said second pair of terminals,wherein a third pair of said four terminals on said sensor is connected for receiving an interrogation signal from said monitor for application across an identification element coupled between said third pair of terminals; andwherein said first, second, and third pairs of terminals each utilize a unique combination of said four terminals and wherein each said first, second, and third pairs of terminals share at least one terminal with another of said first, second, and third pairs of terminals;applying a first interrogation signal to a terminal shared by two of said pairs of terminals to obtain a first output signal across one of said two pairs of terminals;applying a second interrogation signal across the other of said two pairs of terminals to obtain a second output signal at said shared terminal; andutilizing at least one of said first and second output signals to identify said sensor. 2. The method of claim 1, wherein at least one of said applying steps comprise applying an electrical signal to one terminal of said third pair of terminals to obtain at least one output signal across said third pair of terminals. 3. The method of claim 2, wherein said identification element is a diode. 4. The method of claim 3, wherein said output comprises a minimal and predeterminable voltage drop across said diode. 5. The method of claim 3, wherein said applying step further comprises; applying a first electrical signal to obtain a first output signal from said diode; andapplying a second electrical signal to obtain a second output signal from said diode. 6. The method of claim 5, wherein said first and second electrical signals have different magnitudes. 7. The method of claim 6, wherein said first and second electrical signals have first and second current levels. 8. The method of claim 7, wherein said first and second current level are selected to produce first and second minimal and predeterminable voltage drops across said diode. 9. The method of claim 8, wherein said first and second voltage drops are substantially equal. 10. The method of claim 7, wherein said first and second electrical signals have a positive polarity and a negative polarity, respectively, as applied to said identification element. 11. The method of claim 1, further comprising: upon obtaining at least one of said first and second output signals, disabling said third pair of terminals on said s... | Datex-Ohmeda, Inc. | A61B005/05 A61B005/02
US08197418 | B2 | 2012-06-12 | Microprobes | 2008-06-09 | Microprobes in various designs are provided including microprobes having a bulk base and a thin probe extending from the bulk base and a Wheatstone-bridge sensor circuit for measuring strain in the thin probe, and microprobes with two thin probes extending from the bulk base with their respective integrated Wheatstone-bridges to eliminate common mode forces experienced by the two thin probes. | 1. A microprobe comprising: a probe body formed from silicon and shaped to include a bulk base and a thin probe integral to the bulk base and extending beyond the bulk base, the probe body comprising: (a) a core shaped to form the bulk base and the thin probe;(b) a lower blade that: is integrally formed with the core; andterminates in a lower blade pointed tip extending beyond a distal end of the thin probe of the core;(c) an upper blade that: is formed from silicon nitride deposited on an upper surface of the core opposite the lower blade; andterminates in an upper blade pointed tip extending beyond the distal end of the thin probe of the core; and(d) a strain gauge that includes two first polysilicon resistors formed in a first region within the bulk base away from the thin probe to have nearly constant resistance values with respect to a strain experienced by the thin probe and two second polysilicon resistors formed in a second region within the bulk base near the thin probe to have resistance values that change with a strain experienced by the thin probe, wherein the first and second polysilicon resistors are connected to form a Wheatstone bridge circuit that measures the strain experienced by the thin probe. 2. The microprobe of claim 1, wherein the first and second polysilicon resistors of the strain gauge are etched on the upper surface. 3. The microprobe of claim 1, wherein the lower blade comprises no strain gauge. 4. The microprobe of claim 1, wherein the upper blade and the lower blade lie on different parallel planes. 5. The microprobe of claim 1, wherein the bulk base tapers along a direction from the bulk base towards the thin probe. 6. The microprobe of claim 5, wherein the bulk base tapers with a catenoid shape. 7. The microprobe of claim 1, further comprising a second thin probe formed in the probe body extending from the bulk base and separate from the thin probe, and a second strain gauge that includes two third polysilicon resistors formed in a third region within the bulk base away from the thin probe to have nearly constant resistance values with respect to a strain experienced by the second thin probe and two fourth polysilicon resistors formed in a fourth region within the bulk base near the second thin probe to have resistance values that change with a strain experienced by the second thin probe, wherein the third and fourth polysilicon resistors are connected to form a second Wheatstone bridge circuit that measures the strain experienced by the second thin probe. 8. The microprobe of claim 1, further comprising a piezoelectric actuator coupled to the microprobe, directly or indirectly, to cause forward-backward movement of the microprobe when actuated. 9. The microprobe of claim 1, wherein the microprobe defines a channel dimensioned for fluid flow. 10. The microprobe of claim 1, wherein the strain gauge is configured to measure a temperature change. 11. The microprobe of claim 1, wherein the strain gauge is c... | Cornell University | A61B005/00 A61B005/103 A61B005/117
US08200323 | B2 | 2012-06-12 | Program products, methods, and systems for providing fitness monitoring services | 2009-05-18 | Program products, methods, and systems for providing fitness monitoring services are disclosed. In an embodiment, a method for providing heart rate information to a user of a portable fitness monitoring service includes: (a) defining a plurality of heart rate zones as ranges of percentages of a maximum heart rate; (b) associating a color with each of said heart rate zones; (c) receiving heart rate information from the user; and (d) providing a graphical display of the heart rate information, wherein a color of a portion of the graphical display corresponds with the color associated with one of said heart rate zones, wherein steps (a)-(d) are executed using at least one processor. | 1. A tangible computer program product comprising a computer readable medium having computer program logic recorded thereon for causing at least one processor to: graphically display a background comprising a plurality of color coded zones, wherein the plurality of color coded zones are associated with numerical ranges of a first performance parameter;graphically display information about the first performance parameter to a user regarding the user's performance during a workout, wherein said first performance parameter information is displayed as a function of an interval parameter and overlaid on top of the color coded zone background; andgraphically display second performance parameter information to the user regarding the user's performance during the workout, wherein said second performance parameter information is displayed as a function of the interval parameter,wherein the second performance parameter information is overlaid on top of the color coded zone background and the first performance parameter information. 2. The tangible computer program product of claim 1, wherein the interval parameter comprises time. 3. The tangible computer program product of claim 2, wherein first performance parameter information associated with a particular point in time is overlaid on top of the appropriate color coded zone based on a correspondence between a particular numerical value of the first performance parameter at the particular point in time and the numerical range associated with the color coded zone. 4. The tangible computer program product of claim 2, further comprising logic for causing the at least one processor to graphically display a pop-up display element for displaying instantaneous performance data about the first performance parameter at a particular point in time in response to the user indicating a particular point in time of interest. 5. The tangible computer program product of claim 1, wherein the interval parameter comprises distance. 6. The tangible computer program product of claim 5, wherein first performance parameter information associated with a particular distance is overlaid on top of the appropriate color coded zone based on a correspondence between a particular numerical value of the first performance parameter at the particular distance and the numerical range associated with the color coded zone. 7. The tangible computer program product of claim 5, further comprising logic for causing the at least one processor to graphically display a pop-up display element for displaying instantaneous performance data about the first performance parameter at a particular distance in response to the user indicating a particular distance of interest. 8. The tangible computer program product of claim 1, wherein the first and second performance parameter information was recorded over a portion of the workout. 9. The tangible computer program product of claim 1, wherein the first performance parameter information comprises heart... | adidas AG | A61B005/04
US08273053 | B2 | 2012-09-25 | Patient status sensor | 2010-05-04 | Embodiments of a patient status sensor can be applied to a patient or trauma victim to provide a quick visual and/or audible indication of the patient's vital signs (e.g., respiration, heart rate, or other vital signs). Certain embodiments are configured as an adhesive patch that includes electrodes for measuring heart rate (and respiration in some implementations), a processor configured to perform calculations for determining one or more vital signs using information from the electrodes, and audible or visual indicators to communicate information about vital signs or patient status to a medical attendant. Certain embodiments include an access opening for providing intraosseous delivery of fluids to bone marrow (e.g., through sternal or long bone) and can be integrated or used with an intraosseous delivery system. Certain embodiments include wired or wireless components to communicate vital signs or patient status to an external monitoring device. | 1. An intraosseous (IO) fluid delivery and patient status system, the IO system comprising: an IO infusion device configured to provide access to an IO space in a bone of a patient, the IO infusion device comprising: a bone portal comprising a fluid delivery channel, the bone portal having a proximal end and a distal end, the distal end configured to be inserted into the bone; anda fluid delivery conduit configured to be coupled to the proximal end of the bone portal so as to provide fluid access to the fluid delivery channel of the bone portal; anda patient status sensor configured to monitor at least heart rate and respiration rate of the patient, the patient status sensor comprising: a substrate comprising an adhesive component configured to adhere the patient status sensor to the patient;a pair of electrodes disposed in or on the substrate the pair of electrodes configured to receive an electrical signal from the body of the patient and in response to provide a body signal comprising a single-channel electrocardiogram (ECG) signal;a visual indicator;a processor configured to (a) receive and process the body signal from the pair of electrodes in order to determine measurements of heart rate and respiration rate, the heart rate determined from the single-channel ECG signal, (b) determine patient status based at least in part on: (b1) the measurement of the heart rate and one or more heart rate limits and (b2) the measurement of the respiration rate and one or more respiration rate limits, and (c) output a patient status signal to the visual indicator, the visual indicator configured to output visual information indicative of the patient status; anda power source configured to be electrically connected to the visual indicator and the processor. 2. The IO system of claim 1, wherein the distal end of the bone portal is configured to be inserted into the sternum of the patient. 3. The IO system of claim 1, wherein the electrodes of the pair of electrodes are disposed in or on the adhesive component. 4. The IO system of claim 1, wherein the visual indicator, the processor, and the power source are disposed in or on the substrate. 5. The IO system of claim 1, wherein the visual indicator comprises one or more light emitting diodes. 6. The IO system of claim 5, wherein at least one of the one or more light emitting diodes is configured to output a time-varying visual signal, the time-varying visual signal indicative of a time-variation of the body signal. 7. The IO system of claim 6, wherein the time-varying visual signal comprises blinking or flashing, and the time-variation of the body signal comprises heartbeat or respiration. 8. The IO system of claim 1, wherein the power source comprises a battery. 9. The IO system of claim 1, further comprising a removable pull tab configured to electrically isolate the power source from the visual indicator and the processor and to electrically connect the power source to the visual indicator and the pr... | Pyng Medical Corp. | A61M031/00
US08298158 | B2 | 2012-10-30 | Integrated devices having extruded electrode structures and methods of using same | 2009-06-30 | The present disclosure provides electrode structures and integrated electrode structures having one or more conductive materials coextruded with one or more dielectric materials. The disclosed electrode structures can be configured for use as analyte sensors. Also provided, are methods of making and using the electrode structures and integrated electrode structures described herein. | 1. An integrated lancet and electrode structure, comprising: a housing, comprising an extruded electrode structure, the extruded electrode structure comprising: a first extruded body comprising a first conductive material coextruded with a first dielectric material, wherein the first extruded body comprises a first channel;a second extruded body comprising a second conductive material coextruded with a second dielectric material, wherein the second extruded body comprises a second channel,wherein the first extruded body and the second extruded body are engaged to provide the extruded electrode structure,wherein the first conductive material is separated from the second conductive material by a space between the first extruded body and the second extruded body, wherein the space provides a reaction chamber, andwherein the first and second channels together define a lumen having a first opening and a second opening, wherein the lumen is adjacent to and in fluid communication with the reaction chamber, and wherein the lumen does not include the first and second conductive materials; anda lancet slidably positioned within the lumen, such that the lancet is capable of being extended through the lumen to expose the lancet and retracted to a position within the lumen such that the lancet is not exposed. 2. The integrated lancet and electrode structure of claim 1, wherein the electrode structure comprises a spring positioned in the housing and engaged with the lancet, such that when the spring is compressed the lancet extends through the lumen to expose the lancet, and when the spring is uncompressed the lancet is retracted to the position within the lumen. 3. The integrated lancet and electrode structure of claim 2, further comprising a lancet carrier coupled to the lancet. 4. The integrated lancet and electrode structure of claim 3, wherein the spring is a molded component of the lancet carrier. 5. The integrated lancet and electrode structure of claim 4, wherein the spring comprises a first spring arm and a second spring arm. 6. The integrated lancet and electrode structure of claim 1, further comprising a bung slidably positioned within the lumen, the bung comprising a second lumen with a first opening and a second opening, the lancet slidably positioned within the second lumen. 7. The integrated lancet and electrode structure of claim 6, wherein the bung comprises an elastomeric material. 8. The integrated lancet and electrode structure of claim 1, wherein the first extruded body and the second extruded body are snapedly engaged in the absence of an adhesive material. 9. The integrated lancet and electrode structure of claim 1, wherein the first extruded body and the second extruded body are engaged by use of an adhesive material disposed between the first extruded body and the second extruded body. 10. The integrated lancet and electrode structure of claim 3, wherein the first extruded body and the second extruded body comprise a third cha... | Abbott Diabetes Care Inc. | A61B005/00 A61B005/05 B65D081/00
US08721514 | B2 | 2014-05-13 | Electromagnetic pose sensing of HDR brachytherapy applicator | 2008-10-20 | An implantable applicator (10) has at least one guide channel (12) for guiding a radioactive seed or source (66) to a target region. A plurality of imageable fiducials (40) are configured for attachment to the patient adjacent the target region. Electromagnetic sensors (18, 20, 22, 20′, 42) are mounted to the applicator and the imageable fiducials. An electromagnetic tracking system determines the relative position of the applicator mounted sensors (18, 20, 22, 20′) and the fiducial mounted sensors (42). A localizing processor (52) generates a map of the sensors which is combined (56) with a high resolution image (46). A comparing processor (70) monitors for changes in sensor positions, generates a movement correction transform (72) which transforms (74) the planning image into a transformed planning image (76). A brachytherapy treatment processor (60) generates a treatment plan (60a) from the planning image or the transformed planning image for controlling an auto-loader (64) which moves the radioactive source or seed (66) through the channels (12) of the applicator (10). | 1. A brachytherapy apparatus, comprising: an implantable applicator which provides at least one guide channel for guiding a radioactive source or seed to a target region; a plurality of electromagnetic sensors mounted to the applicator;a plurality of imageable fiducials which include fiducial electromagnetic sensors, the fiducials being configured for attachment to a patient adjacent the target region;an electromagnetic tracking system that determines a relative position of the electromagnetic sensors and the fiducial electromagnetic sensors; anda treatment processor that plans a brachytherapy treatment session at least partially based on a combination of the relative position of the electromagnetic sensors and the fiducial electromagnetic sensors and on a planning image including at least the implantable applicator implanted within the patient adjacent the target region and the imageable fiducials attached to the patient adjacent the target region. 2. The apparatus according to claim 1, further including: a comparing processor that determines if and in what manner the electromagnetic sensors have moved relative to the fiducial electromagnetic sensors. 3. The apparatus according to claim 1, wherein the planning of the brachytherapy treatment session by the treatment processor includes a calculation of a brachytherapy treatment plan including a positioning, movement, and timing for moving the radioactive source or seed through the at least one guide channel of the applicator. 4. The apparatus according to claim 3, further including: a comparing processor that determines relative location of the electromagnetic sensors and the fiducial electromagnetic sensors. 5. The apparatus according to claim 4, further including: a transform processor that determines how the applicator has moved relative to a surrounding tissue, the determination of applicator motion being conveyed to the treatment processor for recalculating the brachytherapy treatment plan in accordance with the determined movement. 6. The apparatus according to claim 3, further including: an afterloader connected with the applicator for moving the radioactive source or seed through the at least one guide channel in accordance with the calculated brachytherapy treatment plan. 7. The apparatus according to claim 3, wherein the electromagnetic tracking system determines a map of the electromagnetic sensors and the fiducial electromagnetic sensors during the brachytherapy treatment session implementing the brachytherapy treatment plan, and wherein the treatment processor adjusts the brachytherapy treatment plan during the brachytherapy treatment session. 8. The apparatus according to claim 7, wherein a cumulative dose to subregions of the target region is determined during the implementation of the brachytherapy treatment plan, and wherein in response to at least one of (1) the cumulative dose in the subregions deviating from the brachytherapy treatment plan or (2) sensing movement of the... | Koninklijke Philips N.V. | A61N005/00
US08886334 | B2 | 2014-11-11 | Systems, methods, and devices using stretchable or flexible electronics for medical applications | 2009-12-11 | System, devices and methods are presented that integrate stretchable or flexible circuitry, including arrays of active devices for enhanced sensing, diagnostic, and therapeutic capabilities. The invention enables conformal sensing contact with tissues of interest, such as the inner wall of a lumen, a nerve bundle, or the surface of the heart. Such direct, conformal contact increases accuracy of measurement and delivery of therapy. | 1. An apparatus comprising: a stretchable substrate;a stretchable circuitry selectively attached to said stretchable substrate, said stretchable circuitry comprising:an array of recording electrodes receiving signals from a plurality of nerve sources when at least a portion of said array of recording electrodes is in electrical contact with said plurality of nerve sources;a plurality of amplifiers, each amplifier of said plurality of amplifiers being in electronic communication with a recording electrode of said array of recording electrodes;at least one stretchable interconnect capable of being made longer or wider without tearing or breaking, said at least one stretchable interconnect disposed between, and in electrical communication with, at least one recording electrode of the array of recording electrodes and at least one amplifier of the plurality of amplifiers, said at least one recording electrode of the array of recording electrodes and said at least one amplifier of the plurality of amplifiers being selectively attached to said stretchable substrate, and at least a portion of said at least one stretchable interconnect being free of said stretchable substrate; andan array of stimulating electrodes; anda processing facility in electronic communication with said array of recording electrodes and said array of stimulating electrodes that receives said signals from a plurality of nerve sources from said recording electrodes and is configured to determine a pattern of stimulation signals to be effected by said array of stimulating electrodes. 2. The apparatus of claim 1, wherein said electrical contact comprises physical contact. 3. The apparatus of claim 1, further comprising a multiplexer configured to match said signals from said nerve sources and cause at least one of said plurality of stimulating electrodes to dispatch a corresponding signal to a second plurality of nerves. 4. The apparatus of claim 1, further comprising a user interface to allow an operator to adjust said pattern of stimulation signals. 5. The apparatus of claim 1, wherein said stretchable substrate is an inflatable body. 6. The apparatus of claim 5, wherein said inflatable body is a disk. 7. The apparatus of claim 1, wherein said pattern of stimulation signals is dynamically configurable. 8. The apparatus of claim 1, wherein said processing facility is further configured to generate data related to the electrical conductivity of said nerve sources. 9. The apparatus of claim 8, wherein said processing facility is in electronic communication with an output facility and causes said output facility to generate a map, said map being based on said data related to the electrical conductivity of said nerve sources. 10. The apparatus of claim 1, wherein said stretchable circuitry is encapsulated with a thin polymer layer. 11. The apparatus of claim 1, wherein said stretchable circuitry comprises at least one stretchable electrical interconnect, and wherein said stretc... | MC10, Inc. | A61N001/04 A61B005/00 A61B001/00 A61B005/04 A61B001/01 A61B005/01 A61B005/145 H01L027/146
US08588908 | B2 | 2013-11-19 | System, method and computer program product for detection of changes in health status and risk of imminent illness | 2009-02-04 | A method for analysis of cardiac rhythm and RR interval time series based on entropy related data and entropy based measures. The information is related to but distinct from entropy, and is derived from histograms of interval match counts in which the y-axis is the frequency of intervals of length in that have the match count given on the x-axis. The phenotype of the histogram informs on the presence of atrial fibrillation or, in the presence of sinus rhythm, the degree of congestive heart failure. | 1. A method of detecting abnormal cardiac rhythms and clinical status of a subject comprising: obtaining physiological data obtained from a sampling device from the subject comprising one or more series of intervals;calculating entropy related data using a computer processing device configured for processing the physiological data based on the one or more series of intervals, wherein the entropy related data are new entropy based measures of time series comprising: interval match count histograms, andhistogram derived measures in which the y-axis of the histograms is the frequency of intervals of length m that have the match count given on the x-axis; andgenerating a diagnostic output based on the entropy related data. 2. The method of claim 1, wherein generating the diagnostic output comprises using a regression model to combine the numbers of matching intervals. 3. The method of claim 1, wherein the physiological data comprises two or more series of intervals obtained simultaneously from different physiological signals. 4. The method of claim 1 further comprising separating the one or more series of intervals into a plurality of subsets of the one or more series of intervals. 5. The method of claim 4, wherein calculating the entropy related data further comprises averaging entropy related data from each of the plurality of the subsets of the one or more series of intervals. 6. The method of claim 4, wherein the subsets of the one or more series of intervals comprise less than 13 intervals. 7. The method of claim 1, wherein generating the diagnostic output further comprises using coefficient of sample entropy (COSEn). 8. The method of claim 1 further comprising producing an output to a medical care provider to diagnose abnormal cardiac rhythms and clinical status of the subject. 9. The method of claim 1, wherein the one or series of intervals comprises a number of samples during one period of a cardiac rhythm. 10. The method of claim 1, wherein calculating entropy related data based on the interval match count histograms further comprises a tolerance value (r) for determining whether two or more intervals match. 11. The method of claim 9, wherein the data from the cardiac rhythm of a subject comprises samples from an EKG waveform. 12. The method of claim 1, further comprising calculating an absolute entropy measurement. 13. The method of claim 12, wherein the absolute entropy measurement is a coefficient of sample entropy (COSEn). 14. An apparatus for detecting abnormal cardiac rhythms and clinical status of a subject comprising: a sampling device for obtaining physiological data from the subject comprising a one or more series of intervals;a computer processing device configured for processing the physiological data from the subject into entropy related data and producing a diagnostic result, wherein the entropy related data are new entropy based measures of time series comprising: interval match count histograms, andhistogram deri... | University of Virginia Patent Foundation | A61N001/365
US09020572 | B2 | 2015-04-28 | Systems and methods for processing, transmitting and displaying sensor data | 2010-09-10 | Systems and methods for continuous measurement of an analyte in a host are provided. The system generally includes a continuous analyte sensor configured to continuously measure a concentration of analyte in a host and a sensor electronics module physically connected to the continuous analyte sensor during sensor use, wherein the sensor electronics module is further configured to directly wirelessly communicate displayable sensor information to a plurality of different types of display devices. | 1. A portable device sized to fit in a user's hand and configured to operably connect to a continuous analyte sensor electronics unit, wherein the portable device comprises a telemetry module, a data storage device and a user interface, wherein the telemetry module is configured to wirelessly receive sensor data from an electronics unit of a continuous analyte sensor in a plurality of different data transmission formats,wherein the data storage device is configured to store sensor data,wherein the portable device is configured to display sensor data on the user interface,wherein the portable device is configured to receive sensor data in a first, default data transmission format from the continuous analyte sensor while in a default communication mode,wherein the portable device is configured to send a request to the continuous analyte sensor electronics unit comprising an indication to switch to a second data transmission format that is different from the default data transmission format and an indication of requested sensor data, andwherein the portable device is configured to programically switch to a second communication mode to receive responsive sensor data comprising the requested sensor data in the second data transmission format and programically switch back to the default communication mode upon completion of receipt of the responsive sensor data. 2. The portable device of claim 1, further comprising an authentication mechanism. 3. The portable device of claim 1, wherein the sensor data received from the continuous analyte sensor comprises displayable sensor information. 4. The portable device of claim 1, wherein the portable device is configured to process sensor data from the continuous analyte sensor to provide displayable sensor information. 5. The portable device of claim 1, wherein the portable device is configured to receive user input. 6. The portable device of claim 5, wherein the user interface comprises one or more buttons configured to accept the user input. 7. The portable device of claim 5, wherein the user input comprises calibration information. 8. The portable device of claim 5, wherein the portable device is configured to wirelessly transmit the user input to the continuous analyte sensor. 9. The portable device of claim 1, wherein the portable device is configured to display, on the user interface, a number representative of the analyte concentration based on the sensor data. 10. The portable device of claim 9, wherein the portable device is configured to display, on the user interface, an arrow representative of a rate of change of the analyte concentration based on the sensor data. 11. The portable device of claim 1, further comprising a USB connector and a telemetry adapter configured to enable direct wireless access between the sensor electronics module and a personal computer. 12. The portable device of claim 1, wherein the telemetry module is configured to transmit sensor data wirelessly. 13. The port... | DexCom, Inc. | A61B005/05 H04L029/08 A61B005/00 G06F019/00 A61B005/145
US09179985 | B2 | 2015-11-10 | Biopsy guidance by electromagnetic tracking and photonic needle | 2009-02-27 | A system for providing integrated guidance for positioning a needle in a body has two levels of guidance: a coarse guidance and a fine guidance. The system includes a non-invasive tracking system for tracing the biopsy device in the body, for providing the coarse guidance. Furthermore, the system includes an optical detector mounted on the needle for obtaining optical information discriminating tissue in the body, for providing the fine guidance. | 1. A system for integrated guidance for positioning a biopsy device in a body, the system comprising: a tracking device for coarse guidance of the biopsy device to a biopsy location in the body; andan analyze device for fine guidance of the biopsy device to the biopsy location in the body, wherein the fine guidance is a final precise guidance to the biopsy location more precise than the coarse guidance being an initial guidance to the biopsy location, the analyze device comprising an optical detector configured to detect structures of tissues of the body for providing visible images of the structures, and a console configured to perform a spectroscopy to obtain spectral light distributions of the tissues, wherein the visible images from the optical detector include optically obtained image boundaries of the different tissues, wherein the console and the optical detector are connected to each other, wherein the analyze device is configured to provide information for discriminating different tissues including the optically obtained image boundaries of the different tissues of the body from each other based on the visible images of the structures of the tissues combined with the spectral light distributions of the tissues;a processor configured to combine the information for discriminating the different tissues including the optically obtained image boundaries of the different tissues obtained by the analyze device with non-invasive image-determined boundaries from an image of a non-invasive imaging device providing images of the structures to calculate a scale factor between the optically obtained image boundaries and the non-invasive image-determined boundaries, and using the scale factor to enlarge the non-invasive image-determined boundaries to fully overlap the optically obtained image boundaries to obtain new boundaries of a region of interest; andthe biopsy device comprising an elongate element and a tip portion,wherein the biopsy device is configured to be traced by the tracking device,wherein the optical detector is integrated in the tip portion of the biopsy device. 2. The system as claimed in claim 1, wherein the biopsy device is a biopsy needle including a needle base, a needle head and optical fibers without having a lumen, the needle base receiving the optical fibers that extend into the needle head to fiber exit positions at circumferential surface positions of the needle head. 3. The system of claim 2, wherein a first optical fiber of the optical fibers is configured to provide the fine guidance to the region of interest, and a second optical fiber of the optical fibers is configured to provide optical inspection of a portion of the region of interest harvested by the biopsy device. 4. The system of claim 3, wherein the first optical fiber ends at a front surface of the biopsy device and the second optical fiber ends at a side surface of the biopsy device, and wherein the front surface is transverse to side surface. 5. The system... | Koninklijke Philips N.V. | A61B005/06 A61B019/00 A61B010/04 A61B005/00 A61B010/02 A61B017/34 A61B017/00
US08425878 | B2 | 2013-04-23 | Plasma membrane biomarkers preferentially expressed in pancreatic beta cells useful in imaging or targeting beta cells | 2009-02-13 | The present invention is directed to the identification of a biomarker specifically located in the plasma membrane of pancreatic beta cells. It was selected by a Systems Biology approach on Massively Parallel Signal Sequencing datasets obtained in human islets and Affymetrix microarray datasets on human islets, purified rat primary beta and non beta cells and insulinoma cells. Based on a set of specific features the biomarker is a unique candidate for imaging and targeting strategies to study the pancreatic beta cell mass in health and disease (T1 D, T2D, pancreatic cancers, obesity, islet transplantation, beta cell regeneration). The five specific features of the selected biomarkers are: 1) Preferentially expressed in pancreatic islets as compared to surrounding tissues; 2) Higher expression in pancreatic beta cells than in pancreatic alpha cells or than in other islet non-beta cells; 3) Expression levels in pancreatic beta cells are higher or comparable to glucokinase which is an enzyme specifically expressed in the pancreatic beta cell; 4) Located in the membrane and as such targetable with antibodies, peptides or small molecules which allows imaging, targeting and immunohistochemistry; and 5) Expression is not induced during the process of inflammation of the beta cell mass and the protein is not enriched in T-cells and dendritic cells or in other cells participating in the inflammation process. | 1. A method for specifically detecting pancreatic beta cells in a sample of pancreatic tissue, the method comprising detecting an a-isoform of the gamma subunit of the Na, K-ATPase (FXYD2-gamma-a isoform) in the sample. 2. A method for specifically measuring pancreatic beta-cell mass, the method comprising: visualizing pancreatic beta cells in a sample of pancreatic tissue utilizing a labelled molecule specifically binding to FXYD2-gamma a-isoform;quantifying the amount of labelled pancreatic beta cells in the sample; andmeasuring pancreatic beta cell mass determining the percent of pancreatic beta cells in the total amount of pancreatic cells present in the sample. 3. A method for following up the success of the transplantation of pancreatic beta cells in a subject, the method comprising: visualizing pancreatic beta cells in a sample of pancreatic tissue from the subject utilizing a labelled molecule specifically binding to FXYD2-gamma a-isoform;quantifying the amount of labelled pancreatic beta cells in the sample;measuring pancreatic beta cell mass by determining the percent of pancreatic beta cells in the total amount of pancreatic cells present in the sample; anddetermining the success of the transplantation by tracking the pancreatic beta cell mass in the subject over time. 4. A method for specifically detecting pancreatic beta cells in a sample of pancreatic tissue, the method comprising: detecting a marker selected from the group consisting of: FXYD2-gamma a-isoform, Guanine nucleotide-binding G(s) alpha subunit isoform extra-large (GNAS-XLas), Guanine nucleotide-binding G(s) alpha subunit isoform encoded by XL-exon (GNAS-Alex), CDP-diacylglycerol-inositol 3-phosphatidyltransferase (CDIPT), vesicle amine transport protein 1 homolog (VAT1), Calsyntenin 1 (CLSTN1), solute carrier family 7, member 5 (SLC7A5), cortactin (CTTN), BAI1-associated protein 3 (BAIAP3), LY6/PLAUR: domain containing 1 (LLYPD1), Annexin A7 (ANXA7), deleted in malignant brain tumors 1 (DMBT1), RIKEN cDNA 2310057J16 gene (KIAA1543), and solute carrier family 7, member 8 (SLC7A8) in the sample. | Universite Libre de Bruxelles | A61K049/00
US09717501 | B2 | 2017-08-01 | Methods and systems for occluding vessels during cardiac ablation including optional electroanatomical guidance | 2010-05-21 | A method is provided for ablating a portion of the myocardium. The method includes inserting an occlusion catheter into a vessel on a heart, occluding the vessel using the occlusion catheter, inserting an ablation catheter into a chamber of the heart, positioning the ablation catheter against the myocardium, and ablating a portion of the myocardium while the vessel is occluded. The system includes an occlusion catheter having a catheter body including a tubular member having a distal portion and a bend located in the distal portion, a balloon located proximal of the bend and configured to contact an inner surface of the CS when positioned therewithin, a plurality of marker bands positioned on the catheter body, and a plurality of electrodes positioned on the catheter body. One or more electrodes or coils can be used as a reference for an electroanatomical system and can be disposed on the occlusion catheter. | 1. A method for creating a bi-directional mitral isthmus block in a heart, said method comprising: inserting an occlusion catheter having a balloon and at least one of a first electrode and a first coil on a distal portion of the occlusion catheter into a coronary sinus (CS) of the heart, the at least one of the first electrode and the first coil coupled with an electroanatomical localization system comprising a computer system;expanding the balloon to anchor the at least one of the first electrode and the first coil in the CS;inserting an ablation catheter having at least one of a second electrode and a second coil into a left atrium of the heart, the at least one of the second electrode and the second coil is coupled with the electroanatomical localization system and is configured to provide a non-transitory signal for the computer system indicative of a position of the at least one of the second electrode and the second coil;ablating, while the balloon is expanded and anchored in the CS, a portion of a myocardium to create the bi-directional isthmus block with the ablation catheter;selecting, by the computer system, the at least one of the first electrode and the first coil as a reference node; andmonitoring, by the computer system, a location of the ablation catheter by measuring the non-transitory signal with reference to the reference node. 2. A system comprising: a first medical device comprising: a catheter body having a distal portion and a bend located over a part of the distal portion;a balloon located proximal of the bend and configured to contact an inner surface of a coronary sinus (CS) when positioned therewithin;a plurality of marker bands positioned on the catheter body; andat least one of a plurality of electrodes and a coil positioned on the catheter body; andan electroanatomical localization system comprising a computer system coupled with the first medical device, the computer system configured to select one of: (a) the plurality of marker bands, or (b) one of the at least one of a plurality of electrodes, and the coil as a reference node and to determine a position of a second medical device with reference to the reference node,wherein at least one of the balloon and the catheter body includes at least one passageway extending from a first opening to a second opening in the balloon, the at least one passageway having a longitudinal axis centered in the at least one passageway different from an axis centered in a longitudinal lumen of the first medical device, the at least one passageway configured to permit blood to flow through the at least one passageway while the balloon is inflated and contacting an inner surface of the CS and both the first and second opening are within a blood vessel. 3. The system of claim 2 wherein at least one of the plurality of marker bands is located within the balloon. 4. The system of claim 2 further comprising at least one sensor attached to one of an interior surface of the balloon and an ... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B018/12 A61B017/12 A61B018/14 A61B018/00
US10478555 | B2 | 2019-11-19 | Systems and methods for lumbar cerebrospinal fluid access and treatment | 2010-08-16 | A system and method for accessing and treating the cerebrospinal fluid with a multilumen catheter configured for placement along a cerebrospinal fluid pathway and a subcutaneous, dual reservoir/pump. The subcutaneous dual reservoir/pump allows simultaneous, bidirectional cerebrospinal fluid access and cerebrospinal fluid exchange. The two chambers prevent mixing of newly treated and discardable cerebrospinal fluid. The subcutaneous dual reservoir/pump can be used inline with other cerebrospinal fluid devices. The catheter may be coupled with a medical probe that sends a wire to a computational device, which can then send wireless data and receive wireless instructions. A method of assessing cerebrospinal fluid infections is provided, whereby monitoring of the cerebrospinal glucose concentrations with a cerebrospinal glucose sensor and analysis of the data by a computational device can notify a patient or medical provider of an impending infection. This assembly can help better diagnose and treat injury and disease. | 1. An implantable system for chronic access to cerebrospinal fluid, said system comprising: A multilumen bidirectional fluid flow catheter configured for chronic implantation and placement along a cerebrospinal fluid pathway, wherein said multilumen bidirectional fluid flow catheter comprises a plurality of fluid outlets, and a plurality of fluid inletssaid multilumen bidirectional fluid flow catheter configured to connect, via tunneling, to an assembly comprising at least one finger pumpable dome configured for chronic subcutaneous implantation, each said at least one finger pumpable dome having dual functions of fluid reservoir and pump;wherein said multilumen bidirectional fluid flow catheter and said at least one finger pumpable dome are further configured to allow simultaneous, bidirectional, cerebrospinal fluid access and cerebrospinal fluid exchange. 2. The system of claim 1, wherein the multilumen bidirectional fluid flow catheter is configured to be chronically implanted in a lumbar interspace. 3. The system of claim 1, wherein said assembly further comprises a fastener to an anatomic surface that is further configured to be chronically implanted by subcutaneous fastening to a patient's pelvis. 4. The system of claim 1, wherein said multilumen bidirectional fluid flow catheter further comprises a medical sensor. 5. The system of claim 4, wherein said medical sensor is configured to measure intrathecal pressure. 6. The system of claim 1, wherein said assembly further comprises a computational device. 7. The system of claim 6, wherein said computational device is configured to wirelessly transmit physiological data and warning signals to a patient or provider. 8. The system of claim 6, wherein said computational device is further configured to wirelessly accept commands from a medical provider. 9. The system of claim 1, wherein said multilumen bidirectional fluid flow catheter further comprises, between said plurality of fluid outlets and said plurality of fluid inlets, a medical instrument for transmitting or receiving energy; said energy comprising any of acoustic, radiofrequency, ultrasonic, high frequency ultrasound, photoacoustic, and infrared energy. 10. The system of claim 1, wherein said assembly is further configured in a needle impervious base to form a unit configured for subcutaneous implantation, and; wherein said at least one finger pumpable dome is further configured to enable, while positioned subcutaneously, for chronic access by needle. 11. The system of claim 10, wherein said needle impervious base further comprises a wireless transmitter and computational circuitry. 12. A system of cerebrospinal fluid diversion comprising: a multilumen bidirectional fluid flow catheter coupled to a medical sensor for measuring pressure of cerebral spinal fluid;said multilumen bidirectional fluid flow catheter configured to connect, via tunneling, and via an internal subcutaneous valve to an assembly comprising at least one fi... | - | A61M005/172 A61M025/00 A61M027/00 A61B005/03 A61B005/145 A61B005/00 A61M005/142 A61M039/02
US10330513 | B2 | 2019-06-25 | Multi-dynamic-range sensor | 2009-05-27 | A sensing system that produces a multi-dynamic range output is provided. In an illustrative embodiment, a first channel and a second channel receive an analog output signal from a sensing element. The first channel provides a first digital output signal that has a first dynamic range, and the second channel provides a second digital output signal that has a second narrower dynamic range. In some cases, the second narrower dynamic range falls within the first dynamic range, and the first digital output signal may provide a first resolution and the second digital output signal may provide a second greater resolution. The dynamic range and/or resolution of one or more of the first channel and second channel may be dynamically reconfigurable, if desired. | 1. A respiration sensing system for sensing respiration of a patient, the respiration sensing system comprising: a sensing element configured to sense a respiration flow rate, the sensing element configured to provide an analog output signal that is related to the sensed respiration flow rate;a first channel configured to receive the analog output signal from the sensing element, the first channel including a first amplifier having a first gain and a first offset for providing a first amplified output signal to a first Analog-to-Digital Converter (ADC), the first amplifier in combination with the first Analog-to-Digital Converter (ADC) providing a first digital output signal that has a first dynamic range and a first resolution, the first resolution corresponding to a quantization of the first dynamic range using a first quantization level of the first ADC, wherein the first amplified output signal is offset, wherein the first ADC is included within the respiration sensing system;a second channel configured to receive the analog output signal from the sensing element, the second channel including a second amplifier having a second gain different from the first gain for providing a second amplified output signal to a second Analog-to-Digital Converter (ADC), the second amplifier in combination with the second Analog-to-Digital Converter (ADC) providing a second digital output signal that has a second dynamic range and a second resolution, the second resolution corresponding to a quantization of the second dynamic range using a second quantization level of the second ADC, and wherein the second dynamic range is narrower than the first dynamic range, and the second resolution is higher than the first resolution, wherein the second amplified output signal is offset, wherein the second ADC is included within the respiration sensing system;an output selector, formed as an integrated circuit, configured to receive the analog output signal from the sensing element and output the first digital output signal or the second digital output signal according to the analog output signal received by the output selector, wherein the output selector is configured to select and output the second digital output signal if the analog output signal is within the dynamic range of the second ADC and to select and output the first digital output signal if the analog output signal is outside the dynamic range of the second ADC;wherein the output selector is configured to receive the first digital output signal for all analog output signals over the entire first dynamic range, and wherein the output selector is configured to receive the second digital output signal for all analog output signals over the entire second dynamic range; andwherein the first offset and the second offset are adjustable and are configured to cause a level shift and/or a gain shift in the first channel and/or the second channel. 2. The respiration sensing system of claim 1, wherein the second dynami... | Honeywell International Inc. | G01F007/00 A61B005/087 G01F001/696
US20100317954 | A1 | 2010-12-16 | CONDUCTIVE ELASTIC BAND | 2007-12-03 | Provided is a conductive elastic band that can obtain vital information in smart clothes, wearable computers, or bio-shirts, can transfer the obtained vital information externally through a communication network, and can be applied to various kinds of fabrics and clothes while maintaining conductivity and flexibility in shape. The conductive elastic band includes a front surface in contact with the user including a conductive portion having a fabric electrode and a non-conductive portion weaved with general fibers, a non-conductive rear surface weaved with the general fibers and elastic warps located between the front surface and the rear surface. | 1. A conductive elastic band comprising: a front surface in contact with the a human body and including a conductive portion having a fabric electrode and a non-conductive portion weaved with general fibers;a non-conductive rear surface weaved with the general fibers; andelastic warps located between the front surface and the rear surface. 2. The conductive elastic band of claim 1, wherein the front surface is weaved using woofs, and the woofs include conductive woofs on the conductive portion that can sense vital signals from the human body and non-conductive woofs on the non-conductive portion that can be sewed with the general fibers. 3. The conductive elastic band of claim 2, wherein the warps are elastic rubber threads providing elasticity to the conductive elastic band. 4. The conductive elastic band of claim 3, wherein the warps and the woofs are stitched by a chain stitching to fix the front surface and the rear surface simultaneously. 5. The conductive elastic band of claim 4, wherein the rear surface is weaved using non-conductive woofs. 6. The conductive elastic band of claim 5, wherein the non-conductive woofs of the rear surface and the warps are stitched using the chain stitching so as to fix the front surface and the rear surface simultaneously. 7. The conductive elastic band of claim 6, wherein conductive elastic threads are sewed and connected to the conductive elastic band to transfer the user's vital signals externally. | Electronics and Tlecommunications Research Institute | A61B005/04
US20100280395 | A1 | 2010-11-04 | System and Method for Hypertension Management | 2007-01-19 | A system (1) for hypertension management comprises at least one home blood pressure measuring device (10) for blood pressure measurements of an individual at the individual's home. Furthermore, the system (1) comprises at least one point of care measuring system (20) with a point of care blood pressure measuring device (21). This allows blood pressure measurements at a point or care. The system is provided with data communication means for transferring data from the home blood pressure measuring device (10) to a web server (35) or the point of care measuring system (20). | 1-30. (canceled) 31. A system for hypertension management, the system comprising at least one home blood pressure measurement device for blood pressure measurements at an individual's homeat least one point of care measuring system with a point of care blood pressure measuring device for blood pressure measurement at a point of care,wherein the system is provided with data communication means for transferring data from said home blood pressure measuring device to a central data base and/or said point of care measuring system. 32. A system according to claim 31, wherein at least one of said home blood pressure measuring device or said point of care blood pressure measuring device is adapted to determine blood pressure values in accordance with predetermined measurement guidelines. 33. A system according to claim 32, wherein said home blood pressure measuring device has an interface for entering a measurement schedule based on said measurement guidelines, in particular an interface formed by said data communication means. 34. A system according to claim 33, wherein said home blood pressure measuring device is operable in a first, automatic measurement mode and in a second, manual measurement mode, wherein the measurement mode is preferably determined by the measurement schedule. 35. A system for hypertension management, wherein this system is provided with calculating means adapted for determination of at least one of white coat hypertension or masked hypertension on the basis of measurements made with a home blood pressure measuring devices. 36. A system according to claims 31, wherein the home blood pressure measuring devices and the hypertension management system and preferably the home blood pressure measuring device are operable in a diagnostic operating mode, a treatment operating mode and a follow-up operating mode and wherein the operating mode preferably can be set through an interface formed by said data communication means. 37. A system for hypertension management, comprising at least one point of care measuring device for blood pressure measurements at a point of care, wherein said point of care blood pressure measuring device is designed to operate in accordance with predetermined quality standards, in particular in accordance with guidelines of a pay for performance program or in accordance with a medication therapy management program. 38. A system according to claim 31, the system further comprising at least one home based medication compliance monitoring device, said medication compliance monitoring device being designed for data communication with a central database. 39. A system according to claim 31, wherein the system is further provided with a centralised web based software tool and database accessible by at least one of the point of care measuring system and the home blood pressure measuring system. 40. A system according to claim 39, wherein the web based database includes a disease management knowledge database acc... | MICROLIFE INTELLECTUAL PROPERTY GMBH | A61B005/021
US07828752 | B2 | 2010-11-09 | Device and method for locating the anteroposterior femoral axis to determine proper femoral component rotation in knee replacement | 2006-10-05 | A device for indicating the location of the anteroposterior femoral axis has a spike or needle that engages the anterior-most point in the intercondylar notch and a rounded or spherical surface is placed on the trochlear groove of the femur and urged in a posterior direction. That is, the rounded surface or ball is urged toward the bottom of the trochlear groove. This causes the rounded surface or ball to self-locate into the deepest point of the trochlear groove. The anterior-most point in the intercondylar notch and the deepest part of the trochlear groove anteriorly are points that accurately determine the anteroposterior femoral axis, and the device includes a marking device or guide member that is fixed perpendicular to this axis. A surgeon, therefore, has an accurate indication of the correct femoral component rotation for proper implant position in a knee replacement or other surgical procedure. | 1. A device for approximating a location of an anteroposterior femoral axis, or Whiteside's line, wherein the anteroposterior femoral axis is located at a lower end of a human femur, the device comprising: a body;an upper arm, which extends from the body;an upper engagement member that projects from the upper arm;a lower arm, which extends from the body,a lower engagement member that projects from a distal end of the lower arm, whereinthe lower engagement member and the upper engagement member are spaced apart,a distal end of the lower engagement member is adapted to engage an anterior-most point within an intercondylar notch of the femur, when an axis of the lower arm lies generally in an anteroposterior plane, in which the femoral axis also lies,the upper engagement member is adapted to settle substantially in a deepest part of a trochlear groove of the femur, when the upper engagement member is urged toward the lower engagement member and when an axis of the upper arm lies generally in the anteroposterior plane, whereinthe device is constructed and arranged so that when the lower engagement member engages the anterior-most point within the intercondylar notch of the femur and the upper engagement member engages substantially the deepest part of the trochlear groove of the femur, a line extending between the upper engagement member and the lower engagement member approximates the location of Whiteside's line. 2. The device of claim 1, wherein the lower engagement member is a spike or a needle, which is conical and has a single point for engaging the anterior-most point within the intercondylar notch. 3. The device of claim 1, wherein the device includes a guide member connected to the body, wherein the guide member has a longitudinal dimension that extends perpendicularly to the axes of the upper and lower arms and indicates a guide line that is perpendicular to and spaced apart from the line extending between the upper engagement member and the lower engagement member. 4. The device according to claim 3, wherein the guide member includes a drill guide for guiding a drill, and the drill guide includes a drill guide hole, and an axis of the drill guide hole extends in a direction that is perpendicular to and spaced apart from the line extending between the upper engagement member and the lower engagement member. 5. The device according to claim 1, wherein the upper engagement member and the lower engagement member are constructed and arranged so that the line extending between the upper engagement member and the lower engagement member approximates the location of Whiteside's line when the lower engagement member engages the anterior-most point of the intercondylar notch. 6. The device according to claim 1, wherein the upper engagement member is substantially spherical and is sized such that the upper engagement member will settle into the deepest part of the trochlear groove when the upper arm extends parallel to the femoral axis. 7. The ... | Arthroplasty Innovations, LLC | A61B005/103 A61B005/117 A61B017/58 A61B017/60 A61F002/00
US07831289 | B2 | 2010-11-09 | Planning system, method and apparatus for conformal radiation therapy | 2004-10-07 | A system and associated methods to determine an optimal radiation beam arrangement are provided. The system includes a computer planning apparatus which includes a treatment plan optimization computer having a memory and an input device in communication with the treatment plan optimization computer to provide user access to control functions of plan optimization software. An image gathering device is in communication with the treatment plan optimization computer through a communications network to provide an image slice of the tumor target volume and the non-target structure volume. The plan optimization software computationally obtains and then optimizes a proposed radiation beam arrangement iteratively based on constraints to form an optimized radiation beam arrangement. A conformal radiation therapy delivery device in communication with the treatment plan optimization computer through the communications network then applies the optimized radiation beam arrangement to the patient. | 1. A system to determine an optimal radiation beam arrangement for applying radiation to a tumor target volume while minimizing radiation of a non-target structure volume in a patient, the system comprising a computer planning apparatus including: a treatment plan optimization computer having a memory to store data and treatment plan optimization software therein; andan input device in communication with the treatment plan optimization computer to provide user access to control functions of the treatment plan optimization software;the treatment plan optimization software comprising a set of instructions that, when executed by the treatment plan optimization computer, causes the computer to perform operations to computationally obtain a proposed radiation beam arrangement and to computationally iteratively optimize the proposed radiation beam arrangement based on a plurality of constraints to form the optimized radiation beam arrangement, the treatment plan optimization software including a graphical user interface to display an at least two-dimensional image slice of the tumor target volume and the non-target structure volume, graphical objects, and graphical representations of radiation dose distribution calculated for each iteration of the proposed radiation beam arrangement overlaid upon the displayed image slice, the operations comprising: receiving inputs from the input device representing direct graphical user manipulation of a graphical representation of radiation dose distribution for a current iteration of the proposed radiation beam arrangement displayed on the graphical user interface, andresponsive to user manipulation of the displayed graphical representation of radiation dose distribution for the current iteration of the proposed radiation beam arrangement, computationally obtaining and providing data to graphically display a graphical representation of radiation dose distribution for a next iteration of the proposed radiation beam arrangement which substantially realizes the graphically altered radiation dose distribution, to thereby iteratively form the optimized radiation beam arrangement;the graphical representation of the radiation dose distribution for the current iteration of the proposed radiation beam arrangement displayed on the graphical user interface and directly graphically manipulated via the input device, comprising one or more of the following: an isodose plot including a plurality of isodose contours according to the current proposed radiation beam arrangement overlaid upon the image slice of the target tumor volume and the non-target structure volume, at least one of the plurality of isodose contours of the isodose plot directly graphically manipulatable by a user to graphically alter proposed radiation dose to at least one volume of interest to produce the next iteration of the proposed radiation beam arrangement which substantially realizes the graphically altered radiation dose, the at least one volume of inter... | Best Medical International, Inc. | A61B005/05
US07840254 | B2 | 2010-11-23 | Electromagnetically tracked K-wire device | 2006-01-18 | Devices and methods for registering, dynamically referencing, and navigating an anatomical region of interest of a patient are provided using a tracked Kirschner wire (K-wire), where the K-wire includes a position-indicating element. | 1. A device for image guided surgery that includes a k-wire for insertion in an anatomical body, the k-wire comprising: at least one electromagnetic position-indicating element, located in a distal end of the elongated member and capable of providing the position and orientation of the location of the position-indicating element;at least one conductive lead-wire from which signals are carried from the at least one electromagnetic position-indicating element at the distal end along at least a portion of a length of the k-wire to a proximal end;a plurality of connection bands associated with the k-wire, wherein the plurality of connection bands are located at the proximal end and connect the at least one conductive lead-wire to an electromagnetic position sensor system. 2. The device of claim 1, wherein the connection bands are configured to facilitate axial rotation of the k-wire while maintaining contact with the electromagnetic position sensor system. 3. The device of claim 1, wherein the connection bands connect the at least one conductive lead wire to the electromagnetic position sensor system via a rotary slip coupling that permits rotational movement of the k-wire while maintaining electrical contact between the electromagnetic sensor system and the position indicating element. 4. The device of claim 1, wherein a tip of the k-wire includes a sharp point. 5. The device of claim 1, wherein at least part of the k-wire is threaded. 6. The device of claim 1, wherein the at least one position-indicating element contains radio-opaque markings. 7. The device of claim 6, wherein the radio-opaque markings further form a pattern that can be used to encode information relating to the k-wire. 8. The device of claim 7, wherein the information can be decoded using X-rays. 9. The device of claim 1, wherein the k-wire includes an integrated identification system encoding information relating to the elongated member. 10. The device of claim 9, wherein the identification information includes at least one of: a pattern of dark and light regions that effectively constitute barcode or other similar code and that can be optically determined;a plurality of switches that can be used to construct a binary number that encodes information;a memory device; oran RFID device. 11. The device of claim 1, wherein the conductive elements are of the same or smaller diameter than the remainder of the k-wire. 12. The device of claim 1, wherein the k-wire includes recesses, and the plurality of connection bands are located in the recesses. 13. The device of claim 1, wherein the plurality of connection bands is located in recesses of the k-wire and is flush with the k-wire. 14. The device of claim 1, wherein the plurality of connection bands are substantially near each other and are separated by an insulating portion of the k-wire. 15. The device of claim 1, wherein the plurality of connection bands are rings located around the k-wire. 16. The device of claim 1, whe... | Philips Electronics Ltd | A61B005/05
US07848796 | B2 | 2010-12-07 | Division approximation for implantable medical devices | 2009-03-11 | Methods and devices for performing division approximation in implantable and wearable self-powered medical devices. The present invention provides rapid methods for performing an approximation of division on fixed point numbers, where the methods are easily implemented in small, low power consumption devices as may be found in implantable medical devices. One example of use is in rapidly determining the approximate ratio between foreground and background activity in seizure detection algorithms. Some methods approximate the ratio of Numerator (N) to Denominator (D) by raising 2 to the power of the difference in the number of zeros to the left of the Most Significant Set Bit (MSSB) of D vs. N. Some methods may also pad bits to the right of the approximate ratio MSSB using bits from the right of the N MSSB, and/or pre-process the smaller of D or N by rounding the value upward. Methods may be implemented in firmware and/or in discrete logic. | 1. A method for detecting a seizure in an implantable or wearable medical device (IMD), the method comprising: generating a plurality of ratio determinations of foreground seizure energies divided by background seizure energies;generating a majority of the plurality of ratio determinations using a ratio estimation comprising estimating a Numerator (N) divided by a Denominator (D) to generate a ratio result, the estimating including raising 2 raised to the power of a difference between a most significant set bit position (MSSB) of the D and a MSSB of the N; andgenerating an indication of a seizure with the IMD or the wearable medical device based at least in part on the plurality of ratio determinations. 2. The method of claim 1, in which the ratio estimation includes estimating the Numerator. 3. The method of claim 2, in which the ratio estimation includes estimating the Denominator. 4. The method of claim 1, in which the ratio estimation includes estimating the Numerator and the Denominator. 5. The method of claim 1, in which the ratio result of claim 1 has a MSSB, further comprising setting a bit just to the right of the ratio result MSSB to equal a value of a bit just to the right of the N MSSB. 6. The method of claim 1, in which the ratio result of claim 1 has a MSSB, further comprising setting M bits just to the right of the ratio result MSSB to have values of the same number of M bits just to the right of the N MSSB. 7. The method of claim 6, in which M has a value of at least 1. 8. The method of claim 6, in which M is equal to all bits to the right of the MSSB of the ratio result of claim 6. 9. The method of claim 6, further comprising rounding up the smaller of the N or the D. 10. The method of claim 1, further comprising rounding up the smaller of the N or the D. 11. A non-transitory computer readable medium having computer executable instructions for executing the method of claim 1. | Medtronic, Inc. | A61B005/0476
US20100280405 | A1 | 2010-11-04 | Method For Testing Visual Field | 2010-04-29 | The invention relates to a method for testing visual field of a patient with use of a computer system (7) connected to an output-matrix of light points (6) having controlled brightness and to an input device (4) for measurement of patient's reaction. The method comprises the steps of: a) selecting by the computer system (7) the subset of light points from the set comprising all not yet examined light points of the matrix (6), b) defining by the computer system (7) the initial brightness for each point selected in step a), c) projecting by the computer system (7) subsequent light points of the matrix (6) from the subset of light points selected in step a) with different levels of the brightness starting from the initial brightness determined in step b), registering the reaction of the patient indicating whether the projected point is seen by him by the computer system (7) by means of the input device (4), and defining by the computer system (7) the level of the sensitivity of the retina of the patient in the given point, based on registered reactions of the patient, d) defining by the computer system (7) the map of the sensitivity of the retina of the patient. In order to enable shortening of the duration of the test and decreasing its arduousness for a patient as well as to enable to stop the examination before examining all points of the dot matrix without the loss of previously registered information on the quality of the patient's vision, before step d) at least one cycle of steps a)-c) is carried out, wherein in the first cycle of steps a)-c) the initial brightness of a given point projected is defined in step b) at least on the basis of the data concerning the patient, and in the second and each successive cycle of steps a)-c) the initial brightness of a given point projected is defined in step b) on the basis of the brightness determined for the other light points of the matrix (6). | 1. A method for testing visual field of a patient with use of a computer system (7) connected to an output-matrix of light points (6) having controlled brightness and to an input device (4) for measurement of patient reaction, comprising the steps of: a) selecting by the computer system (7) the subset of light points from the set comprising all not yet examined light points of the matrix (6),b) defining by the computer system (7) the initial brightness for each point selected in step a),c) projecting by the computer system (7) subsequent light points of the matrix (6) from the subset of light points selected in step a) with different levels of the brightness starting from the initial brightness determined in step b), registering the reaction of the patient indicating whether the projected point is seen by him by the computer system (7) by means of the input device (4), and defining by the computer system (7) the level of the sensitivity of the retina of the patient in the given point, based on registered reactions of the patient,d) defining by the computer system (7) the map of the sensitivity of the retina of the patient, characterised in that before step d) at least one cycle of steps a)-c) is carried out,wherein in the first cycle of steps a)-c) the initial brightness of a given point projected is defined in step b) at least on the basis of the data concerning the patient, andin the second and each successive cycle of steps a)-c) the initial brightness of a given point projected is defined in step b) on the basis of the brightness determined for the other light points of the matrix (6). 2. The method according to claim 1, characterised in that said data concerning the patient comprise at least the age of the patient and optionally the sex and/or the ethnic group. 3. The method according to claim 1, characterised in that in said second and each successive cycle of steps a)-c) the initial brightness of a given point projected is defined in step b) on the basis of the brightness determined previously for light points of the matrix (6) in the vicinity of a given point. 4. The method according to claim 1, characterised in that in said second and each successive cycle of steps a)-c) the initial brightness of a given point projected is defined in step b) on the basis of approximation or interpolation of a data set comprising levels of brightness determined previously for the other light points of the matrix (6). 5. The method according to claim 1, characterised in that step c) comprises, for each point of the subset from step a) the steps of, determining by the input device (4) whether the patient sees the projected point or does not see the projected point, and alternate decreasing by the computer system (7) the brightness of this point by a given predefined value dlt1 in case the patient has seen the point or increasing the brightness of this point by a given predefined value dlt2, and projecting this point of changed brightness; and subsequent... | OPTOPOL TECHNOLOGY S.A. | A61B005/00
US20100298730 | A1 | 2010-11-25 | METHOD AND APPARATUS FOR MEASURING BREATHING RATE | 2008-06-13 | An example method and apparatus for measuring the breathing rate from the photoplethysmogram signal (PPG) uses auto-regressive modelling of the signal. The PPG signal is windowed in overlapping windows of typically 30 seconds' length, overlapping by 25 seconds, to obtain discrete sections of the signal and each section is low-pass filtered, downsampled and detrended and then AR modelled using an all-pole auto-regressive (AR) model. The AR model allows identification of the dominant frequencies in the signal and the pole corresponding to the breathing rate is identified by considering its magnitude and the breathing rate it represents. Each 30 second window gives a breathing rate estimate and use of successive windows displaced by 5 seconds results in a breathing rate estimate every 5 seconds. The time series of breathing rate estimates can be Kalman filtered to reject measurements which have a large change in magnitude or represent a large change in breathing rate. The measurements may also be fused with measurements from another sensor. | 1. A method of measuring the breathing rate of a subject, comprising obtaining a photoplethysmogram from the subject, and further comprising the steps of: modelling each of a plurality of discrete time periods of the photoplethysmogram using respective autoregressive models, finding poles of the autoregressive models of the discrete time periods of the photoplethysmogram, and calculating from the poles for each period the breathing rate for each time period. 2. A method according to claim 1 wherein the time periods are successive overlapping windows. 3. A method according to claim 2 wherein the successive overlapping windows are displaced from each other by 5 to 15 seconds and are from 20 secs to 1 minute long. 4. A method according to claim 3 wherein the successive overlapping windows are displaced from each other by 5 seconds and are 30 seconds long. 5. A method according to claim 1 further comprising the step of downsampling the photoplethysmogram before modelling. 6. A method according to claim 1 further comprising the step of detrending the photoplethysmogram before modelling. 7. A method according to claim 1 further comprising the step of low-pass filtering the photoplethysmogram before modelling. 8. A method according to claim 1 further comprising the step of identifying as a pole representing the breathing rate a pole which has a magnitude greater than a predetermined threshold. 9. A method according to claim 1 further comprising the step of identifying as a pole representing the breathing rate a pole which corresponds to a predefined allowable range of breathing rates. 10. A method according to claim 1 wherein the order of the model is from 7 to 13. 11. A method according to claim 1 wherein the order of the model is 9, 10 or 11. 12. A method according to claim 1 further comprising the step of filtering the breathing rate measurements for the plurality of discrete time periods to produce a smoothed breathing rate measurement. 13. A method according to claim 12 wherein said filtering step comprises Kalman filtering. 14. A method according to claim 12 further comprising rejecting measurements which represent greater than a predetermined change in breathing rate. 15. A method according to claim 1 further comprising the step of combining the measurement of breathing rate with a breathing rate measurement from another sensor. 16. A method according to claim 15 wherein the combination is a weighted combination with weights derived from confidence in the measurements. 17. Apparatus for measuring the breathing rate of a subject, comprising an input for receiving a photoplethysmogram from the subject, a data processor for executing the steps of modelling each of a plurality of discrete time periods of the photoplethysmogram using respective autoregressive models, finding the poles of the autoregressive models of the discrete time periods of the photoplethysmogram, and calculating from the poles for each period the breathing rate fo... | Oxford Biosignals Limited | A61B005/08 G06F017/10
US20100298674 | A1 | 2010-11-25 | PROTECTIVE SHELL FOR AN IN VIVO SENSOR MADE FROM RESORBABLE POLYMER | 2010-04-21 | An implantable device with in vivo functionality, where the functionality of the device is negatively affected by the inflammation reaction generally associated with tissue injury, encapsulated by a protective coating that prevents damage to the device from any inflammation reactions. The protective coating is designed to persist for a set period of time, generally until after the inflammation reaction of the surrounding in vivo environment in response to the injury caused by the implantation procedure has concluded. The protective coating is further designed to “resorb” (i.e. to dissociate from the device, dissolve, and be absorbed into the surrounding environment) after a set period of time, allowing the device to perform its in vivo functionality unhindered without loss of performance. | 1. A device comprising: (a) an implantable device which has an in vivo functionality and;(b) a layer of protective coating applied onto the implantable device wherein: (1) the protective coating prevents or reduces degradation or interference of the implantable device from inflammation reactions; and(2) the protective coating is designed to resorb over a period of time under in vivo conditions. 2. The device of claim 1, wherein the device is a sensor. 3. The device of claim 2, wherein the sensor is for monitoring blood glucose levels. 4. The device of claim 1, wherein the protective coating is applied to the device by dipping the device in the coating material. 5. The device of claim 1, wherein the protective coating is applied to the device by spraying the coating onto the device. 6. The device of claim 1, wherein the protective coating is applied to the device by thermal extrusion of the coating onto the device. 7. The device of claim 1, wherein the protective coating is a physiologically compatible material comprising one or more polyanhydrides. 8. The device of claim 7, wherein the one or more polyanhydrides comprises poly(sebacic acid) and/or poly(1,3-bis(p-carboxyphenoxy)propane-co-sebacic acid). 9. The device of claim 1, wherein the protective coating is a physiologically compatible material comprising processed collagen. 10. The device of claim 1, wherein the protective coating is a physiologically compatible material comprising one or more polyesters. 11. The device of claim 10, wherein the one or more polyesters comprises polyglycolic acid and/or polylactic acid, or copolymers thereof. 12. The device of claim 1, wherein the layer of protective coating is from about 5 to about 200 microns thick. 13. The device of claim 12, wherein the layer of the protective coating is from about 20 to about 30 microns thick. 14. The device of claim 1, wherein the period of time is from about 4 to about 5 days. 15. The device of claim 2, wherein the sensor comprises a body surrounding a photosensitive detector element and a light source, and further wherein the external surface of the sensor body comprises a matrix layer that comprises one or more indicator molecules. 16. The device of claim 15, wherein the one or more indicator molecules comprises a phenylboronic acid residue. 17. The device of claim 15, wherein the sensor further comprises a power source and a transmitter surrounded by the sensor body. 18. A method for using an implantable device in in vivo applications comprising: (a) providing an implantable device which has an in vivo functionality, and which comprises a layer of a protective coating on the device wherein: (1) the protective coating prevents or reduces degradation or interference of the device from inflammation reactions; and(2) during use the protective coating resorbs into the surrounding environment over a period of time; and(b) implanting the implantable device in a subject body. 19. The method of claim 18, whe... | Sensors for Medicine & Science, Inc. | A61B005/145 A61B005/1459
US20100286562 | A1 | 2010-11-11 | SENSOR FOR BLOOD COMPONENT ANALYSIS | 2010-07-22 | A sensor for blood component analysis, which allows even a trace amount of blood to be led to an analysis portion reliably. The sensor for blood component analysis includes a substrate, a spacer, and a cover. The cover is disposed on the substrate with the spacer intervening between the cover and the substrate, whereby a space that serves as an analysis portion and a channel for leading blood to the analysis portion is formed inside the sensor. Through holes are formed in the substrate and the spacer, respectively, so that a common through hole through which a needle of a lancet can pass is formed when the spacer is disposed on the substrate. The through hole of the spacer communicates with the channel and a top of the through hole of the spacer is covered with the cover, whereby a lancing portion is formed by the common through hole and a portion of the cover covering the top of the through hole of the spacer. In use, the sensor is placed at a position where blood collection is to be performed, the portion of the cover covering the lancing portion is broken through with the needle of the lancet so that the needle is allowed to pass through the common through hole to puncture the position, and blood that has come out is guided by a downwardly protruding burr formed when the cover is broken through so that the blood is led to the channel and flows through the channel to be led to the analysis portion. | 1-9. (canceled) 10. A collector and sensor for blood component analysis comprising: a substantially rectangular substrate having a through hole;a spacer having a through hole;a cover including a through hole as an air vent;a reagent layer; and a lancing portion, wherein the cover is disposed on the substrate, and the spacer is disposed between the cover and the substrate, whereby the through hole of the spacer, the through hole of the substrate, and the cover form a space,the space includes an analysis portion and a channel for leading a blood sample from a blood collector portion to the analysis portion,the reagent layer is disposed in the analysis portion,the thorough hole as an air vent of the cover communicates with the space to serve as an air vent hole,the through hole in the substrate and the through hole in the spacer form a common through hole,the cover covers a top of the through hole of the spacer,the lancing portion is formed by a portion of the cover covering the top of the through hole of the spacer and the common through hole, andthe through hole of the substrate, the through hole of the spacer, and the through hole as an air vent of the cover are located closer to one end in a longitudinal direction with respect to the center of the substrate, whereby the lancing portion, the analysis portion, and the air vent hole are located closer to one end in the longitudinal direction with respect to the center of the collector and sensor. 11. The collector and sensor according to claim 10, wherein the lancing portion, the analysis portion, and the air vent hole are disposed in this order from the center to the end of the sensor. 12. The collector and sensor according to claim 10, wherein the lancing portion, the analysis portion, and the air vent hole are disposed in this order from the center to the end of the sensor, and a plurality of the analysis portions are disposed in series with respect to the longitudinal direction of the substrate between the lancing portion and the air vent hole. 13. The collector and sensor according to claim 10, wherein a plurality of the analysis portions and a plurality of the air vent holes are disposed radially with respect to the one lancing portion, and each of the analysis portions is disposed between the lancing portion and each of the air vent holes. 14. The collector and sensor according to claim 10, wherein a plurality of the analysis portions and a plurality of the air vent holes are disposed so as to surround a plurality of the lancing portions, and each of the analysis portions is disposed between each of the lancing portions and each of the air vent holes. 15. The collector and sensor according to claim 10, wherein the lancing portion and a plurality of the analysis portions disposed in parallel with respect to the longitudinal direction of the substrate are disposed in this order from the center to the end of the sensor, and the channel branches, so that the lancing portion and a plurality... | PANASONIC CORPORATION | A61B005/151
US20100312081 | A1 | 2010-12-09 | Implantable Tissue Ischemia Sensor | 2010-08-17 | An implantable ischemia detection system in which a white LED (105) produces a continuous, visible, broadband light illuminating a target site (125). Light backscattered by the target is collected by a sensor (155), allowing for an index of ischemia to be determined, and subsequently transmitted by a sending unit (167). Power is provided by an internal power source (179). The entire implantable device is encapsulated by a biocompatible shell (102) to add long-term safety with regard to implantation. | 1. An implantable device for invasive implantation internally comprising: a light source for illuminating a target with light;a sensor configured so as to detect light backscattered from local capillary and extravascular tissue within the target site onto said sensor and to provide a tissue oxygenation output signal that is a direct function of the presence or degree of internal tissue ischemia based at least on said light backscattered from said site;a power source configured to supply energy to said sensor and said light source; andan exterior shell enclosing said power source, light source, and sensor configured so as to be biocompatible with respect to implantation internally. 2. A method of monitoring or detecting local tissue ischemia internally in a living tissue comprising the steps of: (a) invasively implanting internally a device having a light emitter and a light detector;(b) illuminating an internal target site with light from the light emitter;(c) detecting light returning from tissue using the light detector, said returning light substantially reflecting capillary and extravascular tissue sites; and(d) determining a measure that is a function of local tissue ischemia internally based upon the detected light, said detected light substantially reflecting said capillary and extravascular tissue internally at the local target site. 3. A device for implantation internally comprising: (a) a white LED for illuminating a target site with light;(b) a wavelength-resolved linear CCD configured so as to detect light substantially reflecting light backscattered from local capillary and extravascular tissue within the target site onto said CCD;(c) a sensor configured so as to determine, based at least upon said reflected light detected by said sensor, an output signal that is a function of the presence or degree of local tissue ischemia internally;(d) a transmitter for receiving and transmitting said output signal;(e) a an inductive coil for receiving power configured to supply energy to said LED, said CCD, and said transmitter; and(f) an exterior shell configured so as to be biocompatible with respect to implantation internally, said shell housing one or more said LED, CCD, sensor, transmitter and inductive coil. 4. The method of claim 2, further including the step of transmitting data before or after said determined measure is determined. | Spectros Corporation | A61B005/1455
US20100331659 | A1 | 2010-12-30 | LOW ARTIFACT DEFIBRILLATION ELECTRODE | 2009-06-29 | An electrode arrangement for low artifact electrocardiogram (ECG) monitoring and defibrillation, comprises at least one conductor print line body having a large silver conductor print line area functioning as a defibrillation area and a small silver-silver chloride (Ag/AgCl) conductor print line area functioning as a sensing area, at least one non halide defibrillation gel layer and halide containing sensing gel layer configured to attach with the defibrillation area and sensing area respectively, and at least one insulated conductive ink line for connecting the conductor print line body to a cable that provides an electrical communication between the conductor print line body and an external equipment. When direct current runs through the conductor print line body, the large silver conductor print line area polarizes and the small Ag/AgCl conductor print line area does not polarize so that small Ag/AgCl conductor print line area senses ECG signal leading to low motion artifact. | 1. An electrode arrangement for low artifact electrocardiogram (ECG) monitoring and defibrillation, comprising: at least one conductor print line body having a large silver conductor print line area functioning as a defibrillation area and a small silver-silver chloride (Ag/AgCl) conductor print line area functioning as a sensing area;at least one non halide defibrillation gel layer configured to attach with the defibrillation area on a bottom side of the conductor print line body;at least one halide containing sensing gel layer configured to attach with the sensing area on the bottom side of the conductor print line body;at least one insulated conductive ink line for connecting the conductor print line body to a cable that provides an electrical communication between the conductor print line body and an external equipment; andwhereby as a direct current (DC) runs through the conductor print line body, the large silver conductor print line area polarizes and the small Ag/AgCl conductor print line area does not polarize so that the small Ag/AgCl conductor print line area is picked up first and senses an ECG signal. 2. The electrode arrangement of claim 1 wherein a motion artifact in the ECG signal may be significantly less due to a small conductive surface area of the sensing area. 3. The electrode arrangement of claim 1 wherein the defibrillation area delivers an electrical impulse to a patient's heart and the sensing area sequentially conveys the ECG signal from the patient's heart to external monitoring equipment. 4. The electrode arrangement of claim 1 wherein the conductor print line body comprises a plurality of top layers including a substrate for conductor print line body, a laminating adhesive, and an insulated backing member. 5. The electrode arrangement of claim 1 wherein the conductor print line body further comprises a plurality of bottom layers including a laminating adhesive, a connection insulation, and a removable cover for skin contacting side. 6. The electrode arrangement of claim 1 wherein the defibrillation area can function as any of a plurality of other cardiac stimulation area including pacing, cardioversion, etc. 7. A method of treating a patient with an electrode arrangement for low artifact electrocardiogram (ECG) monitoring and defibrillation, comprising the steps of: providing the electrode arrangement comprising a conductor print line body having a defibrillation area and a sensing area;removing a cover for a skin contacting side and applying the electrode arrangement that is configured to adhesively secure to the skin of the patient;receiving a first set of electrical impulses from an external stimulating equipment on the defibrillation area and delivering the first set of electrical impulses to a patient's heart; andsequentially receiving a second set of electrical impulses on the sensing area from the patient's heart and conveying the second set of electrical impulses to an external monitoring equipment. 8. ... | - | A61B005/04 A61N001/04
US20100318006 | A1 | 2010-12-16 | POWER REGENERATION IN ACTIVE MUSCLE ASSISTANCE DEVICE AND METHOD | 2010-08-20 | A method for controlling movement using an active powered device including an actuator, joint position sensor, muscle stress sensor, and control system. The device provides primarily muscle support although it is capable of additionally providing joint support (hence the name “active muscle assistance device”). The device is designed for operation in several modes to provide either assistance or resistance to a muscle for the purpose of enhancing mobility, preventing injury, or building muscle strength. The device is designed to operate autonomously or coupled with other like device(s) to provide simultaneous assistance or resistance to multiple muscles. | 1. An apparatus for reducing muscle stress, comprising: a brace having a first portion and a second portion movably coupled to the first portion, the first portion configured to be fastened to a first part of a user's body and the second portion configured to be fastened to a second part of the user's body;a detector operative to detect an intention to exert a muscle by detecting an amount of muscle stress;an actuator coupled to the brace, the actuator operative, when energized, to exert force between the first and second portions; anda controller coupled to the actuator and responsive to the detector, the controller configured to vary an amount of force exerted by the actuator between the first and second portions depending on the amount of muscle stress detected by the detector, wherein varying the amount of force exerted by the actuator reduces stress in the muscle intended to be exerted. 2. The apparatus of claim 1, wherein the detector is a sensor configured to detect varying capacitance, resistance, or pressure. 3. The apparatus of claim 1, wherein the detector is a foot pressure sensor. 4. The apparatus of claim 3, wherein the controller is configured to energize the actuator when a muscle stress is detected by the detector that is over a threshold amount. 5. The apparatus of claim 1, wherein the first portion is configured to be fastened to a leg above a knee, and wherein the second portion is configured to be fasted to the leg below the knee. 6. The apparatus of claim 5, wherein the muscle intended to be exerted is a quadricep of the leg. 7. The apparatus of claim 1, further comprising a battery coupled with the actuator to supply electricity to the actuator. 8. The apparatus of claim 1, further comprising a joint angle sensor configured to detect a joint angle, wherein the controller is configured to vary the amount of force exerted by the actuator depending on the amount of muscles stress and the detected joint angle. 9. The apparatus of claim 1, wherein the apparatus is a portable device. 10. A method of reducing muscle stress, comprising: fastening a first portion of a brace to a first part of a user's body and a second portion of a brace to a second part of the user's body, the first portion and the second portion being movably coupled;detecting an intention to exert a muscle by detecting an amount of muscle stress; andapplying an amount of force between the first and second portions, the amount of force applied varied based on the amount of muscle stress detected such that stress in the muscle intended to be exerted is reduced. 11. The method of claim 10, wherein applying an amount of force comprises applying force with an actuator coupled to the brace. 12. The method of claim 10, wherein detecting an amount of muscle stress comprises detecting an amount of muscle stress with a sensor configured to detect varying capacitance, resistance, or pressure. 13. The method of claim 10, wherein detecting an amount of muscle stre... | - | A61H001/02 A61B005/103
US20100286507 | A1 | 2010-11-11 | DETERMINING THE POSITION OF A NEEDLE | 2008-12-29 | An arrangement for determining the position of a needle in an organ system and a method utilizing the arrangement in cosmetic treatment or cosmetic surgery. The arrangement (1) comprises a needle (2) with a distal tip (3), a measurement circuit (4) measuring the bioimpedance of tissues, a detection means (10) connected to the measurement circuit measuring bioimpedance and arranged to detect the progress of the distal tip (3) of the needle at a real-time basis from the bioimpedance measured from tissues. The arrangement (1) comprises at least two electrodes (16, 18, 21, 22) arranged at the distal tip (3) of the needle and connected to the measurement circuit (4). All the electrodes (16, 18, 21, 22) measuring impedance and needed for monitoring the progress of the distal tip (3) of the needle are arranged to the distal tip (3) of the needle. | 1-23. (canceled) 24. An arrangement for determining the position of a needle in an organ system, the arrangement comprising a needle having a distal tip;a measurement circuit measuring the bioimpedance of tissues;a detection means, the detection means being connected to the measurement circuit measuring the bioimpedance and arranged to detect the progress of the distal tip of the needle in the tissue on a real-time basis from the bioimpedance measured from the tissues,the arrangement further comprising at least two electrodes arranged to the distal tip of the needle, the electrodes being coupled to the measurement circuit; whereinall electrodes measuring bioimpedance and needed for monitoring the progress of the distal tip of the needle are arranged to the distal tip of the needle. 25. An arrangement according to claim 24, wherein the number of the electrodes is two. 26. An arrangement according to claim 24, wherein the number of the electrodes is four. 27. An arrangement according to claim 24, wherein one of the electrodes is the needle frame, which is electrically insulated from the rest of the electrodes arranged to the needle. 28. An arrangement according to claim 24, wherein at least one electrode is made of an electrically conductive material layer which surrounds the needle frame and which is electrically insulated from the rest of the electrodes arranged to the needle. 29. An arrangement according to claim 24, wherein at least one electrode is a discrete piece formed to the outer surface of the needle. 30. An arrangement according to claim 24, wherein that the detection means is a display device. 31. An arrangement according to claim 24, comprising means for identifying the force that resists the progress of the needle in the tissue. 32. An arrangement according to claim 31, wherein the means comprise a pressure sensor and a power sensor. 33. An arrangement according to claim 24, wherein the needle is hollow and comprises a conduit for supplying fluid substances into and/or out of the body. 34. An arrangement according to claim 33, comprising a container connected to the needle conduit. 35. An arrangement according to claim 24, comprising a master station to be arranged to the patient's skin or mucosa, the master station including a frame structure attachable to the patient andat least one power member arranged to move the needle in relation to the frame structure. 36. An arrangement according to claim 35, wherein the master station comprises sounding devices arranged to identify subcutaneous tissue through the patient's skin. 37. An arrangement according to claim 36, wherein the sounding devices comprise an ultrasound transmitter and receiver. 38. An arrangement according to claim 24, comprising memory means for storing an anatomic model created on the basis of anatomical data and means for adjusting the bioimpedance information measured by the electrodes to the model in such a way that the position of the distal tip of th... | - | A61B005/05
US20100312127 | A1 | 2010-12-09 | Blood pressure cuff | 2010-05-28 | A blood pressure cuff is provided with a cuff body 2 and an air supply/exhaust tube 3. The air supply/exhaust tube 3 includes a coupling axle 33 having a communicating hole 35 at one end thereof. The cuff body 2 is provided with a tube coupling part 25 protruded in a radial direction of the cylindrical cuff body from an outer peripheral surface thereof, and the tube coupling part has a coupling axle supporting part 25c that supports the coupling axle rotatably. The coupling axle supporting part 25c aligns a center axis “O” of the coupling axle 33 with a direction approximately perpendicular to a center axis of the tube coupling part, which coincides with the radial direction of the cylindrical cuff body, and approximately perpendicular to an axial direction of the cuff body. The coupling axle 33 can be rotated about the center axis “O”. | 1. A blood pressure cuff, comprising: a cuff body which is wrapped around an arm or wrist of a user to form a substantially cylindrical shape when used;an air bag provided in the cuff body, the air bag being charged with air to occlude blood flow; andan air supply/exhaust tube rotatably coupled to the cuff body at one end thereof to communicate with the air bag, an extension direction of the air supply/exhaust tube from the air bag being changeable,wherein the air supply/exhaust tube is provided with a coupling axle at the one end thereof, the coupling axle having a communicating hole that allows the air supply/exhaust tube to communicates with the air bag,wherein the cuff body is provided with a tube coupling part protruding in a radial direction of the cuff body from an outer peripheral surface thereof, the tube coupling part having a coupling axle supporting part that rotatably supports the coupling axle, andwherein the coupling axle supporting part is adapted to align a center axis of the coupling axle in a direction that is approximately perpendicular to an axial direction of the cuff body and is approximately perpendicular to the protruding direction of the tube coupling part, and to rotate the coupling axle about the center axis of the coupling axle. 2. The blood pressure cuff as set forth in claim 1, further comprising an air leak prevention unit for preventing air leak from between the coupling axle supporting part and the coupling axle. 3. The blood pressure cuff as set forth in claim 1, wherein the air supply/exhaust tube comprises:a tube body; anda coupling member coupling the tube body with the cuff body,wherein the coupling member is formed in an L shape in which respective axial directions of both end portions thereof cross with each other at an approximately right angle and one end of the coupling member is coupled to one end of the tube body,wherein the coupling axle is provided at the other end of the coupling member such that the extension direction of the air supply/exhaust tube is changeable between a first direction and a second direction in the axial direction of the cylindrical cuff body as the coupling axle supported by the coupling axle supporting part is rotated, the first direction and the second direction being in opposite directions to each other,wherein the blood pressure cuff further comprises a locking unit for releasably locking the air supply/exhaust tube to the cuff body,in each of the first direction and the second direction. 4. The blood pressure cuff as set forth in claim 2, wherein the air supply/exhaust tube comprises:a tube body; anda coupling member coupling the tube body with the cuff body,wherein the coupling member is formed in a L shape in which respective axial directions of both end portions thereof cross with each other at an approximately right angle and one end of the coupling member is coupled to one end of the tube body,wherein the coupling axle is provided at the other end of the coupling ... | Panasonic Electric Works Co., Ltd. | A61B005/022
US20100317957 | A1 | 2010-12-16 | THREE-DIMENSIONAL WEARABLE ELECTRODE SET | 2009-06-16 | A three-dimensional (3D) wearable electrode set for a human body is provided. The 3D wearable electrode set comprises an integrative first ring electrode and an integrative second ring electrode. The integrative first ring electrode has a conductive layer and a ring basis, and the integrative second ring electrode also has a conductive layer and a ring basis. The conductive layers of the integrative first and second ring electrodes are adopted to cover around the first and second portions of the human body, respectively. The conductive layers are formed with conductive material. The ring bases are formed with insulating fabric. | 1. A three-dimensional (3D) wearable electrode set for a human body, comprising: an integrative first ring electrode, having a first conductive layer and a first ring basis, being adopted to cover around a first portion of the human body, wherein the first conductive layer is formed with conductive material, the first ring basis is formed with insulating fabric, and the first conductive layer is attached to the first ring basis; andan integrative second ring electrode, having a second conductive layer and a second ring basis, being adopted to cover around a second portion of the human body, wherein the second conductive layer is formed with the conductive material, the second ring basis is formed with the insulating fabric, and the second conductive layer is attached to the second ring basis;wherein the first conductive layer is electrically connected to a first terminal of a processor via a first conductive thread, the second conductive layer is electrically connected to a second terminal of the processor via a second conductive thread. 2. The 3D wearable electrode set of claim 1, wherein the first conductive thread and second conductive thread are yarns blended with the conductive fibers and the insulating fibers. 3. The 3D wearable electrode set of claim 1, wherein the processor generates a first stimulation current and a second stimulation current, the first stimulation current is transmitted to the integrative first ring electrode via the first conductive thread and the first terminal of the processor, and the second stimulation current is transmitted to the integrative second ring electrode via the second conductive thread and the second terminal of the processor. 4. The 3D wearable electrode set of claim 3, wherein the first stimulation current and the second stimulation current are used for transcutaneous electrical nerve stimulation (TENS). 5. The 3D wearable electrode set of claim 1, wherein the first conductive layer receives a first electrical impulse generated from the first portion of the human body and transmits the first electrical impulse to the first terminal of the processor via the first conductive thread. 6. The 3D wearable electrode set of claim 5, wherein the second conductive layer receives a second electrical impulse generated from the second portion of the human body and transmits the second electrical impulse to the second terminal of the processor via the second conductive thread. 7. The 3D wearable electrode set of claim 6, wherein the first electrical impulse and the second electrical impulse forms an electromyography (EMG) signal. 8. The 3D wearable electrode set of claim 6, wherein the first electrical impulse and the second electrical impulse forms an electrocardiography (ECG) signal. 9. The 3D wearable electrode set of claim 1, wherein the conductive material is a fabric woven with conductive fibers and insulating fibers. 10. The 3D wearable electrode set of claim 1, wherein the integrative first ring el... | TEX-RAY INDUSTRIAL CO., LTD. KING'S METAL FIBER TECHNOLOGIES CO., LTD. | A61B005/0408 A61N001/04 A61B005/0492
US07860562 | B2 | 2010-12-28 | External force control method, external force control system and external force control program | 2005-11-29 | A method of controlling an external force applied to an animal to achieve as a target relation the relation between the external force and a motion variable varying with the motion of the animal. A value of an external force f applied to a human leg according to an external force function f(x) on the basis of a measured value of a myoelectric potential x that occurs in the human leg, current I of a motor 220 is controlled according to the set value, and the external force f is applied to the leg through an orthosis 222. A resultant force (the sum of an internal torque and an external torque around a knee joint) F is measured as “a motion variable.” Moreover, a value of a factor γ is set according to a factor function γ(f, F) on the basis of the set value of the external force f and the measured value of the resultant force F. If a deviation δ between the set value of the factor γ and target value γt thereof is equal to or greater than a reference value ε, a new external force function f(x) is set in such a way that the set value of the factor γ approaches the target value γt. | 1. An external force control method for controlling an external force applied to a first living body portion of an animal through an orthosis attached to the first living body portion of the animal that makes a movement along with the activities of animal muscle fibers, the method comprising: a myoelectric potential measurement step of measuring a myoelectric potential x that occurs in a second living body portion of the animal;an external force setting step of setting a value of an external force f applied to the first living body portion of the animal through the orthosis according to an external force function f(x) with the myoelectric potential x as a variable on the basis of the measured value of the myoelectric potential x;a motion variable measurement step of measuring a motion variable y varying with the motion of the animal under the condition of the external force applied through the orthosis;a factor setting step of setting a value of a factor γ according to a factor function γ(f, y) with the external force f and the motion variable y as variables on the basis of the set value of the external force f and the measured value of the motion variable y;a determination step of determining whether a deviation δ between the set value of the factor γ and target value γt thereof is less than a reference value ε; andan external force function setting step of setting a new external force function f(x) in such a way that the set value of the factor γ approaches the target value γt if the deviation δ is determined to be equal to or greater than the reference value ε in the determination step, whereinthe motion variable measurement step comprises measuring the resultant force of an internal force, which is a force exerted on the orthosis by the first living body portion caused by activities of animal muscle fibers in at least the second living body portion of the animal, and an external force, caused by a motion of an actuator and applied through the orthosis to the first living body portion of the animal, as the motion variable y; andthe factor setting step comprises setting the ratio of the external force f to the resultant force of the internal force and the external force of the animal as the factor γ(0≦γ&lt;1). 2. The external force control method according to claim 1, wherein the external force function setting step comprises setting a value of a coefficient α that represents the relation between the myoelectric potential x and the external force f and setting the external force function f(x) according to a basic function F(x, α) with the myoelectric potential x and the coefficient α as variables on the basis of the set value of the coefficient α. 3. The external force control method according to claim 1, wherein the external force function setting step comprises finding the external force target value ft according to the factor function γ(f, y) on the basis of the measured value of the motion variable y and the target value γt of the factor ... | Honda Motor Co., Ltd. | A61B005/04
US07844320 | B2 | 2010-11-30 | Method and apparatus for volumetric image navigation | 2002-08-27 | A surgical navigation system has a computer with a memory and display connected to a surgical instrument or pointer and position tracking system, so that the location and orientation of the pointer are tracked in real time and conveyed to the computer. The computer memory is loaded with data from an MRI, CT, or other volumetric scan of a patient, and this data is utilized to dynamically display 3-dimensional perspective images in real time of the patient's anatomy from the viewpoint of the pointer. The images are segmented and displayed in color to highlight selected anatomical features and to allow the viewer to see beyond obscuring surfaces and structures. The displayed image tracks the movement of the instrument during surgical procedures. The instrument may include an imaging device such as an endoscope or ultrasound transducer, and the system displays also the image for this device from the same viewpoint, and enables the two images to be fused so that a combined image is displayed. The system is adapted for easy and convenient operating room use during surgical procedures. | 1. A system for augmented viewing of a patient target site in a surgical setting by a user, comprising: a computer for storing scan data of the patient target region, and for constructing from the stored data, a three dimensional perspective volumetric image of the target site as seen from a selected orientation;a viewing device, the device having a first ocular instrument and a second ocular instrument in which the patient target site can be viewed, and in which the three dimensional perspective volumetric image can be projected, forming a fused image using the patient target site view and the three dimensional perspective volumetric image;a tracking device coupled to the computer for tracking the selected orientation of the viewing device; anda calibration module in the computer for placing the viewing device in the same coordinate system as the patient in response to tracking information received from the tracking device, wherein the selected orientation from which the volumetric image of the target is constructed by the computer is the selected orientation of the second ocular instrument. 2. The system of claim 1 wherein the patient target site can be viewed in the first ocular instrument and the three dimensional perspective volumetric image can be projected in the second ocular instrument. 3. The system of claim 1 wherein the patient target site can be viewed in the first and second ocular instrument and the three dimensional perspective volumetric image can be projected in the second ocular instrument. 4. The system of claim 1 wherein the patent target site can be viewed in the first and second ocular instrument and the three dimensional perspective volumetric image can be projected in the first and second ocular instrument. 5. The system of claim 1 wherein scan data is selected from a group comprising computed tomography, magnetic resonance imaging data, single photon emission computed tomography, x-ray data, positron emission data, ultrasound data, video data, and computed axial tomography data. 6. The system of claim 1 wherein an orientation of a surgical tool is projected on the three dimensional perspective volumetric image of the target site. 7. The system of claim 1 wherein a position of a surgical tool is projected on the three dimensional perspective volumetric image of the target site. 8. The system of claim 1 further comprising a laser scanner mounted on the viewing device for scanning a laser beam over a surface of the patient target region to create a laser pattern on the patient, and wherein the tracking device included two or more cameras for viewing the laser pattern on the patient from two or more different orientations. 9. The system of claim 8 wherein the calibration module is operative to match surface features of the patient, as determined from the volumetric scan data of the patient, with corresponding surface features, as determined from the laser pattern on the patient, to place the scan data in the same fra... | - | A61B005/05
US07840246 | B1 | 2010-11-23 | Implantable self-calibrating optical sensors | 2005-11-17 | A measurement light detector detects light transmitted by a light source of an implantable system that is scattered back into an implantable housing, and produces a measurement signal indicative of the intensity of the light detected by the measurement light detector. A calibration light detector detects a portion of the transmitted light that has not exited the housing, and produces a calibration signal that is indicative of the intensity of the light detected by the calibration light detector, which is indicative of the intensity of the light transmitted by the light source. Changes in the intensity of the transmitted light are compensated for based on the calibration signal produced by the calibration light detector. This description is not intended to be a complete description of, or limit the scope of, the invention. Other features, aspects, and objects of the invention can be obtained from a review of the specification, the figures, and the claims. | 1. An implantable system, comprising: an implantable housing comprising a window through which light can pass;at least one light source, within said housing, each of which transmits light of a corresponding wavelength, wherein an intensity of the light transmitted by each said light source is controlled by a corresponding drive signal that drives said light source, and wherein a portion of the light of each said wavelength exits said housing through said window;a measurement light detector, within said housing, to detect light of each said wavelength scattered back into said housing through said window, and to produce a measurement signal that is indicative of the intensity of the light of each said wavelength detected by said measurement light detector;an opaque wall that separates said at least one light source from said measurement light detector;a calibration light detector, within said housing on a same side of said opaque wall as said at least one light source, to detect a portion of the light of each said wavelength that has not exited said housing and has been reflected toward said calibration light detector without using a beam splitter, and to produce a calibration signal that is indicative of the intensity of the light of each said wavelength detected by said calibration light detector, which is indicative of the intensity of the light transmitted by each said light source; andcircuitry operative to detect changes in the intensity of the light transmitted by each said light source based on the calibration signal;wherein said housing has an inner surface, a first portion of said inner surface coated with a reflective material to cause a portion of the light of each said wavelength that has not exited said housing to be reflected toward said calibration light detector, and a second portion of said inner surface not coated with the reflective material and corresponding to said window through which a portion of the light of each said wavelength exits said housing. 2. The system of claim 1, wherein said at least one light source comprises a single light source. 3. The system of claim 1, wherein said at least one light source comprises a first light source that transmits light of a first wavelength and a second light source that transmits light of a second wavelength. 4. The system of claim 3, wherein said at least one light source comprises a first, a second and a third light source, that transmit, respectively, light of a first wavelength, light of a second wavelength, and light of a third wavelength. 5. The system of claim 1, wherein said at least one light source comprises first and second spatially separated light sources that transmit light of about 805 nm, wherein said first light source is located closer to said measurement light detector than said second light source. 6. The system of claim 1, wherein the circuitry comprises a processor, state machine or random logic. 7. The system of claim 1, further comprising a processor co... | Pacesetter, Inc. | A61B005/1455
US07853305 | B2 | 2010-12-14 | Trajectory storage apparatus and method for surgical navigation systems | 2005-05-13 | A method and apparatus for determining, calculating, and/or viewing a trajectory. The trajectory can be displayed on a display relative to image data of a patient. A user can use the system to determine relationships between two or more trajectories that have been determined. The relationships can be within two dimensional, three-dimensional, or four dimensional space. The relationships can include distance, angle, etc. The relationships can also be determined between real time trajectories and stored trajectories, or combinations thereof. | 1. A method for determining a relative location of at least two trajectories, the method comprising: providing image data of a patient in a first plane;tracking a three-dimensional trajectory of an object;determining, with a surgical navigation system, a position of a first representation of a first trajectory of the object relative to the provided image data of the patient;determining, with the surgical navigation system, a position of a second trajectory of the object relative to the image data of the patient; andstoring the first trajectory or the second trajectory. 2. The method of claim 1, further comprising: displaying the provided image data;forming a representation of the determined position of the first trajectory of the object; andsuperimposing the formed representation of the first trajectory of the object on the displayed image data. 3. The method of claim 2, further comprising changing the formed representation to a second formed representation and superimposing the second formed representation on the displayed image data. 4. The method of claim 1 further comprising: selecting the first and second trajectory;projecting the first trajectory into a plane of the image data to form a first set of projected points;projecting the second trajectory into the plane of the image data to form a second set of projected points;calculating a first segment connecting the first set of projected points;calculating a second segment connecting the second set of projected points;computing an angle between the first and second segments in the plane of the image data; anddisplaying the computed angle. 5. The method of claim 4 further comprising using the displayed computed angle for performing at least one of bone fragment alignment, bone fracture fixation, implant placement, bone removal and spinal implant placement. 6. The method of claim 1, further comprising storing both of the first trajectory and the second trajectory. 7. The method of claim 6, further comprising: determining which of at least one of the determined first trajectory or the determined second trajectory is a real-time trajectory;wherein the determined real time trajectory is stored contemporaneously with the tracking of the three-dimensional trajectory of the object. 8. The method of claim 1, wherein providing image data of a patient includes providing at least one of one-dimensional image data, two-dimensional image data, three-dimensional image data, four-dimensional image data, or combinations thereof. 9. The method of claim 1, wherein providing image data of the patient in a first plane includes providing image data of a patient in at least a first plane and a second plane. 10. The method of claim 9, wherein tracking a three-dimensional trajectory of an object includes tracking the three-dimensional trajectory of the object relative to at least the first plane or the second plane. 11. The method of claim 9, wherein providing image data of a patient in the first plane incl... | Medtronic Navigation, Inc. | A61B005/05
US07824346 | B2 | 2010-11-02 | Determining condyle displacement utilizing electronic models of dental impressions having a common coordinate system | 2005-09-19 | A system and method for determining condyle displacement during jaw articulation includes a physical model with corresponding reference points. The physical model is positioned and scanned to obtain positional data representing a first and second bite position. This positional data is used to generate a transformation matrix. The position of at least one condyle is determined in reference to positional data scanned from the physical model. The transformation matrix is used to map the position of the condyle with respect to the second bite position. | 1. A method for determining condyle displacement for a patient, the method comprising: mounting a first physical model onto a first scanning plate module to form a first mounting arrangement, the first physical model being representative of a first set of teeth of the patient;mounting a second physical model onto a second scanning plate module to form a second mounting arrangement, the second physical model being representative of a second set of teeth of the patient;positioning the first and second mounting arrangements on a base plate module to arrange the first physical model relative to the second physical model within a coordinate system of a scanning device according to a first bite record;scanning three or more reference sites on at least one of the mounting arrangements with the scanning device to obtain a first set of positional data;positioning the first and second mounting arrangements on the base plate module to arrange the first and second physical models within the coordinate system of the scanning device according to a second bite record;scanning the three or more reference sites with the scanning device to obtain a second set of positional data;determining a transformation matrix from the first and second sets of positional data with a data processing system;determining a first location of a point corresponding to a condyle of the patient in relation to the first set of positional data;determining with the data processing system a second location of the point corresponding to the condyle in relation to the second set of positional data based on the transformation matrix; andgenerating a manipulable electronic model based on the first and second sets of positional data and the transformation matrix to assess condyle displacement of the patient. 2. The method of claim 1, wherein determining the first location of the point corresponding to the condyle includes measuring a distance between the condyle of the patient and at least one point on one of the mandible and the maxilla of the patient using a face bow. 3. The method of claim 1, wherein determining the first location of the point corresponding to the condyle includes: displaying a medical record image showing at least one condyle of the patient;superimposing an electronic model corresponding to one of the mandible and the maxilla of the patient over the medical record image; andselecting a point on the medical record image, wherein the selected point corresponds to the position of the condyle. 4. The method of claim 3, wherein a computer software program selects the condyle point on the medical record image. 5. The method of claim 1, wherein positioning the first and second mounting arrangements on the base plate module further comprises: mounting the first mounting arrangement to the base plate module of the scanning device, the base plate module positioning the first mounting arrangement at a known position within the coordinate system of the scanning device; andmounting t... | Geodigm Corporation | A61B005/103 A61B005/117 A61C019/04 G01C009/00 G01C017/00 G01C019/00
US07833174 | B2 | 2010-11-16 | Method and apparatus for subcutaneously advancing a device between locations | 2003-01-31 | A method and apparatus (10) for advancing a device (12) underneath the skin (102) in a mammalian body comprises a casing (50) made of a magnetic material. The casing (50) has a closed tip (64) at one end and an aperture (65) in an opposite end. The casing (50) includes an annular inner surface (70) that defines a cavity (76) within the casing for receiving a device (12) that projects through the aperture (65). The casing (50) further includes structure for attaching to the device (12). The casing (50) is operable to advance the device (12) through tissue (100) underlying the skin (102) in response to movement of a magnetic field across the skin. | 1. An apparatus for advancing a device underneath the skin in a mammalian body, said apparatus comprising: a casing comprising at least two body portions that are movable relative to each other, said casing made entirely of a magnetic material, said casing having a closed tip at one end and an aperture in an opposite end;said casing including an annular inner surface that defines a cavity within said casing for receiving a device that projects through said aperture;said casing further including means for attaching to the device;said casing being operable to advance the device through tissue underlying the skin in response to movement of a magnetic field across the skin. 2. The apparatus of claim 1 wherein said device comprises an electrical lead. 3. The apparatus of claim 1 wherein said device comprises a shunt catheter. 4. The apparatus of claim 1 further comprising a magnet for providing a magnetic field that is movable across the skin to cause corresponding subcutaneous movement of said casing. 5. The apparatus of claim 1 wherein said at least two body portions of said casing are connected to each other by a hinge. 6. An apparatus for advancing a device underneath the skin in a mammalian body, said apparatus comprising: a casing made entirely of a magnetic material, said casing having a closed tip at one end and an aperture in an opposite end;said casing including an annular inner surface that defines a cavity within said casing for receiving a device that projects through said aperture;said casing further including means for attaching to the device that include at least one support extending from said inner surface of said casing;said casing being operable to advance the device through tissue underlying the skin in response to movement of a magnetic field across the skin. 7. The apparatus of claim 6 wherein said at least one support includes an engagement pad for pressing against the device with an interference fit. 8. The apparatus of claim 1 wherein said means for attaching to the device includes a portion of said inner surface. 9. The apparatus of claim 6 wherein said means for attaching to the device includes a plurality of supports extending from said inner surface of said casing. 10. An apparatus for subcutaneously advancing an elongated member between different locations in a mammalian body, said apparatus comprising: a casing comprising at least two body portions that are movable relative to each other, said casing made entirely of a magnetic material, said casing having oppositely disposed first and second ends, said first end being closed and forming a tip portion of said casing for tunneling through tissue underlying the skin, said second end being open to form a mouth portion of said casing;said casing including annular inner and outer surfaces that extend between said first and second ends and which define a side wall portion of said casing, said inner surface further defining a cavity within said casing for receiving a ... | The Cleveland Clinic Foundation | A61B005/00
US20100331650 | A1 | 2010-12-30 | EPISODIC BLOOD GLUCOSE MONITORING SYSTEM WITH AN INTERACTIVE GRAPHICAL USER INTERFACE AND METHODS THEREOF | 2009-06-25 | A blood glucose monitoring system with a graphical user interface (GUI) and methods of controlling the system and performing episodic blood glucose testing using the GUI are disclosed. The system can comprise the GUI, an output display for displaying the GUI, user interfaces, a memory and a processor. An episodic blood glucose testing protocol can be programmed into the processor using the user interfaces and automatically saved into memory. A user can be alerted to test blood glucose levels based on the inputted testing protocol by the display of an alert icon on the GUI. Blood glucose level results can be received, displayed on the GUI along with a range icon and automatically saved into memory. Statistics of the saved blood glucose results can be calculated by the processor and saved into memory. The statistics can be retrieved for display on the GUI when prompted by the user. | 1. An episodic blood glucose testing method comprising: providing an episodic blood glucose monitoring system having a processor, memory, user interfaces, and an output display supporting a graphical user interface (GUI) which facilitates an episodic blood glucose testing mode;using the user interfaces and the GUI to enter the episodic blood glucose testing mode and to program the processor to execute an episodic blood glucose testing protocol comprised of a plurality of testing alert times;saving into memory the episodic blood glucose testing protocol,responding to a testing alert provided by the processor by inserting a test strip into the system, the testing alert corresponding to one of the testing alert times of the blood glucose testing protocol;conducting a blood glucose test using the inserted test strip;reviewing displayed blood glucose testing results provided along with a range indicator icon on the GUI;saving the blood glucose testing results into the memory; andprompting the processor to display calculated statistics of saved blood glucose testing results on the GUI. 2. The method of claim 1, further comprising toggling between the episodic blood glucose testing mode and a normal blood glucose meter operations mode of the episodic blood glucose monitoring system in which only the displaying of the blood glucose testing results is provided on the output display after conducting the blood glucose test on the inserted test strip. 3. The method of claim 1, further comprising the processor automatically saving into memory the episodic blood glucose testing protocol and the blood glucose results. 4. The method of claim 1, further comprising resetting the testing alert for a set period of time. 5. The method of claim 1, further comprising responding to a re-alerting of the testing alert after elapse of a set period of time. 6. The method of claim 1, further comprising using the user interfaces and GUI to program the episodic blood glucose testing protocol by setting a time for each of the testing alert times and designating each of the testing alert times as a pre-meal testing, a post-meal testing, or a pre-bedtime testing. 7. The method of claim 1, wherein the testing alert comprises an alert icon displayed on the output display along with an icon indicating one of a pre-meal testing, post-meal testing, and pre-bedtime testing. 8. The method of claim 1, further comprising using the user interfaces and the GUI to enter and save in memory at least one of a meal size and a meal speed for the blood glucose test if the alert time was for a post-meal test. 9. The method of claim 1, wherein the calculated statistics comprise the processor calculating what percentage of blood glucose testing results fall within an acceptable blood glucose level range, below an acceptable blood glucose level range, and above an acceptable blood glucose level range. 10. The method of claim 1, wherein the calculated statistics comprise the processor calculat... | ROCHE DIAGNOSTICS OPERATIONS, INC. | A61B005/145
US07833170 | B2 | 2010-11-16 | Needle-insertion device | 2003-12-11 | This invention relates to a lancing apparatus (A1) used for inserting an insertion element (21) into a human skin (Sk) for sampling a body fluid. The lancing apparatus (A1) includes a housing (1) with a cylindrical portion (12) brought into contact with the skin (Sk), a negative pressure generator (6) that generates a negative pressure inside the cylindrical portion (12) to cause the skin (Sk) to swell upward, and a detector (5) that detects that the skin (Sk) has been raised to a predetermined height inside the cylindrical portion (12). | 1. A lancing apparatus configured for sampling a body fluid out of a skin, the apparatus comprising a housing including a cylindrical portion adapted to be brought into contact with the skin, an insertion element movable relative to the housing for sticking the skin, and a negative pressure generator configured to generate a negative pressure inside the cylindrical portion to cause the skin to swell upward, wherein the apparatus further comprises a height detector and a pressure controller configured to cooperate with the height detector, the height detector being configured for detecting that the skin has been raised to a predetermined height inside the cylindrical portion, the height detector being provided separately from the insertion element and including a tapered face that is adapted to come into contact with the skin when the skin swells upward, the pressure controller being configured to execute a control so as to maintain a pressure inside the cylindrical portion within a specific range, the pressure controller being configured to execute the control based on detection by the height detector that the skin has been raised to the predetermined height, the pressure controller being configured to define the specific range by granting a specific tolerance to a reference pressure that the pressure controller is configured to set at a lower value than the pressure inside the cylindrical portion at a time that the height detector has detected that the skin has been raised to the predetermined height. 2. The lancing apparatus according to claim 1, further comprising a pressure detector that detects a pressure inside the cylindrical portion, wherein the controller is configured to execute a control so as to maintain a pressure inside the cylindrical portion within a specific range, based on the pressure detected by the pressure detector. 3. The lancing apparatus according to claim 1, wherein the specific range has an upper limit and a lower limit which the pressure controller is configured to set at a lower value than the pressure inside the cylindrical portion at the time that the height detector has detected that the skin has been raised to the predetermined height. 4. The lancing apparatus according to claim 1, wherein the height detector is capable of detecting a fluctuation of the swelling height of the skin, and wherein the controller is configured to control the pressure inside the cylindrical portion so as to maintain the swelling height of the skin at the predetermined level. 5. The lancing apparatus according to claim 4, wherein the height detector includes a contacting member for contact with the skin when the skin has been raised to the predetermined height, so as to detect a contacting pressure of the skin applied to the contacting member. 6. The lancing apparatus according to claim 5, wherein the controller is configured to control the pressure inside the cylindrical portion so as to maintain the contacting pressure within a se... | ARKRAY, Inc. | A61B005/00 B65D081/00 A61B017/32
US07826885 | B2 | 2010-11-02 | Method and apparatus for calculating index concerning local blood flow circulations | 2007-09-14 | First time-density curves concerning arteries in a specific region and second time-density curves concerning tissues in the specific region are prepared from a plurality of continuous images concerning the specific region of a subject into which a contrast medium is injected. Modulation transfer functions indicating local blood flow circulations in the tissues with respect to the arteries are calculated by curve-fitting so that residual errors of the second time-density curves are minimized with respect to convolution of the modulation transfer functions and first time-density curves. Indices concerning the local blood flow circulations with respect to the respective arteries are calculated from the modulation transfer functions. Maps of the indices for the arteries are prepared, and the maps of these indices are synthesized into one map in accordance with the residual errors for the first time-density curves. | 1. An index calculation method executed by an image processing apparatus, comprising: using, at the image processing apparatus, a weight in accordance with a degree of similarity between pixels to filter a plurality of images in order to reduce noise from the plurality of continuous images concerning a specific region of a subject with a contrast medium injected therein, wherein the weight is a first weight for a first degree of similarity and the weight is a second weight greater than the first weight for a second degree of similarity higher than the first degree of similarity;preparing, at the image processing apparatus, first time-density curves concerning arteries in the specific region and second time-density curves concerning tissues in the specific region from the plurality of images with the reduced noise;calculating, at the image processing apparatus, modulation transfer functions indicating local blood flow circulations in the tissues with respect to the respective arteries by curve-fitting so that residual errors of the second time-density curves are minimized with respect to convolution of the modulation transfer functions and first time-density curves; andcalculating, at the image processing apparatus, indices concerning the local blood flow circulations with respect to the respective arteries from the modulation transfer functions. 2. The method according to claim 1, further comprising: determining the similarity between the pixels based on a difference of time-density curves between the pixels. 3. The method according to claim 1, further comprising: determining the similarity between the pixels based on a difference of time-density curves between the pixels and a dispersal anticipated value of the noise. 4. The method according to claim 1, further comprising: replacing each pixel value of the pixel constituting the image with a weighted average with a peripheral pixel by the weight. 5. The index calculation method according to claim 1, wherein the modulation transfer function is a rectangular function. 6. The method of claim 1, wherein the using comprises: setting, for each of a plurality of pixels, a combination of vector values corresponding to values of said each pixel included in the plurality of images;determining, by statistical testing, degrees of similarity of combinations of vector values corresponding to the plurality of pixels;weighting values of pixels corresponding to combinations of vector values proportionally to the degree of similarity; andcalculating an average of weighted pixel values in each of the plurality of images. 7. An index calculation apparatus comprising: a filter section which uses a weight in accordance with a degree of similarity between pixels to filter a plurality of images in order to reduce noise from the plurality of continuous images concerning a specific region of a subject with a contrast medium injected therein, wherein the weight is a first weight for a first degree of similarity and ... | Kabushiki Kaisha Toshiba | A61B005/05
US07853303 | B2 | 2010-12-14 | Neurological probe and method of using same | 2006-11-16 | A neurological probe has a plurality of stacked electrode elements, each electrode element having stimulation/lesioning and recording electrodes incorporated with a strip of electrically non-conductive substrate. Such a probe is more compact while having a large number of stimulation/lesioning and recording channels. | 1. Neurological probe comprising a plurality of stacked electrode elements, each electrode element comprising a strip of electrically non-conductive substrate having incorporated therewith a first electrode for providing an electrical current and a second electrode for recording electrical activity, wherein the first electrode is a stimulation and/or lesioning electrode and the second electrode is a recording electrode. 2. Probe of claim 1, wherein the stacked electrode elements are staggered to stagger positions of the first and second electrodes along the probe. 3. Probe of claim 2, wherein each electrode element comprises only one stimulation and/or lesioning electrode and only one recording electrode. 4. Probe of claim 3, wherein the first and second electrodes are substantially coplanar in a plane defined by a length and a width of the strip. 5. Probe of claim 4, wherein the electrode elements are stacked in a direction perpendicular to the plane defined by the length and width of the strip. 6. Probe of claim 1, wherein the plurality of stacked electrode elements is four or more stacked electrode elements. 7. Probe of claim 1, wherein the first and second electrodes comprise platinum, titanium, gold, platinum-iridium or tungsten. 8. Probe of claim 1, wherein the first and second electrodes comprise platinum. 9. Probe of claim 1, wherein the non-conductive substrate comprises a biocompatible material. 10. Probe of claim 1, wherein the non-conductive substrate comprises a polyimide. 11. Probe of claim 1, wherein the non-conductive substrate comprises a parylene. 12. Probe of claim 1 further comprising an electrical connector connected to the electrodes. 13. Probe of claim 12 further comprising a protective housing to at least partially protect the electrical connector. 14. Probe of claim 1 further comprising a protective covering to at least partially protect the stacked electrode elements. 15. Probe of claim 1, wherein the plurality of stacked electrode elements form a stack having a cross-sectional area of 500 μm×500 μm or less. 16. Probe of claim 1, wherein the second electrode has an impedance of 100,000 ohms or greater as measured at a frequency of 1000 Hz in a 0.9% NaCI solution, and the first electrode has an impedance of 500 ohms or less as measured at a frequency of 1000 Hz in a 0.9% NaCl solution. 17. Method of modulating and/or measuring neurological activity of a nerve cell or tissue comprising: bringing a probe into proximity of the nerve cell or tissue, the probe comprising a plurality of stacked electrode elements, each electrode element comprising a strip of electrically non-conductive substrate having incorporated therewith a first electrode for providing an electrical current and a second electrode for recording electrical activity, wherein the first electrode is a stimulation and/or lesioning electrode and the second electrode is a recording electrode; and,determining electrical current generated in the secon... | National Research Council of Canada | A61N001/36 A61N001/04 A61B005/04 A61B018/14
US07840253 | B2 | 2010-11-23 | Method and apparatus for surgical navigation | 2005-09-30 | A surgical navigation system for tracking a procedure relative to a patient is disclosed. The system can track a portion of the patient, an instrument, both, and both or either relative to image data, a coordinate system, an atlas, a morphed atlas, or combinations thereof. The system can include a minimally or non-invasive dynamic reference frame and/or fiducial marker, sensor tipped instruments, and isolator circuits. | 1. A system and apparatus for tracking a portion of a patient, an instrument, registering an image, or combinations thereof, the system comprising: a tracking sensor;a tracking system that tracks the tracking sensor;a body operable to surround at least a portion of the tracking sensor;an anchor coupled to the body to hold the body relative to the portion of the patient at a first location by directly engaging the patient at the first location; andan attachment member retractably coupled to the body to hold the body relative to the patient at a second location different from the first location by directly engaging the patient at the second location, the attachment member configured to translate relative to the body from a first position retracted in the body to a second position protruding from the body to directly engage the patient;wherein the body has a dimension to allow percutaneous placement of the body beneath soft tissue of the patient such that the tracking sensor is out of the line of sight of the tracking system, and wherein the body defines a bore to enable the anchor to pass through the body and hold the body relative to the portion of the anatomy. 2. The system and apparatus of claim 1, wherein the dimension is less than about 50 mm. 3. The system and apparatus of claim 1, wherein the dimension is less than about 20 mm. 4. The system and apparatus of claim 1, wherein the dimension is a height of the body. 5. The system and apparatus of claim 1, wherein the body is annular. 6. The system and apparatus of claim 5, wherein the attachment portion comprises a plurality of projections formed radially about a central axis of the body. 7. The system and apparatus of claim 1, wherein the body includes a lower body portion and an upper body portion; wherein said upper body portion includes an upper fixation portion and said lower body portion includes a lower fixation portion to allow for holding the upper body portion and the lower body portion relative to one another. 8. The system and apparatus of claim 7, wherein the upper fixation portion and the lower fixation portion include friction fit portions. 9. The system and apparatus of claim 8, wherein one of the upper fixation portion or the lower fixation portion define a groove and the other of the upper fixation portion or the lower fixation portion define a projection; wherein the projection is operable to fit within the groove to hold the respective portions relative to one another. 10. The system and apparatus of claim 7, wherein the upper fixation portion and lower fixation portion define complementary tapers, complementary threads, adhesives, or combinations thereof. 11. The system and apparatus of claim 7, wherein the anchor further comprises a flange that engages the lower body portion to prevent the rotation of the anchor relative to the body. 12. The system and apparatus of claim 1, further comprising: a tube defining a throughbore; andwherein the body is operable to pas... | Medtronic Navigation, Inc. | A61B005/05 A61B019/00
US07835781 | B2 | 2010-11-16 | Device for superposed MRI and PET imaging | 2007-08-03 | A device is disclosed for superposed MRI and PET imaging and includes an MRI tube that has a first field of view along its longitudinal direction, and a multiplicity of PET detection units arranged opposite one another in pairs about the longitudinal direction. In at least one embodiment the many PET detection units define a second field of view along the longitudinal direction, and their arrangement density along the longitudinal direction is optimized such that the second field of view substantially corresponds to the first field of view. | 1. A device for superposed magnetic resonance tomography and positron emission tomography imaging, comprising: a magnetic resonance tomography tube that defines a first field of view along its longitudinal direction;a multiplicity of positron emission tomography detection units arranged opposite one another in pairs about the longitudinal direction, the multiplicity of positron emission tomography detection units defining a second field of view along the longitudinal direction, whereinthe positron emission tomography detection units respectively include an avalanche photodiode array with an upstream lutetium oxyorthosilicate crystal array and an electrical amplifier circuit,a spacing at least corresponding to an avalanche photodiode array is present between two neighboring avalanche photodiode arrays, andfor each positron emission tomography detection unit, the electrical amplifier circuit is arranged in a plane, parallel to the longitudinal direction, next to at least one of the lutetium oxyorthosilicate crystal array and the avalanche photodiode array, wherein the portion of the total area of said plane which is occupied by the lutetium oxyorthosilicate crystals defines an arrangement density of the positron emission tomography detection units in said plane, andsaid arrangement density of the positron emission tomography detection units being provided along the longitudinal direction such that the second field of view is substantially equal to the first field of view. 2. The device as claimed in claim 1, wherein the positron emission tomography detection units are respectively oriented at the end of the second field of view in the longitudinal direction such that at least one of the lutetium oxyorthosilicate crystal array and the avalanche photodiode array is oriented toward the middle of the magnetic resonance tomography tube, and the electrical amplifier circuit is oriented toward the edge of the magnetic resonance tomography tube. 3. The device as claimed in claim 1, wherein the arrangement density of the positron emission tomography detection units along the longitudinal direction becomes relatively larger from the middle of the second field of view toward the edge thereof. | Siemens Aktiengesellschaft | A61B005/05 A61B006/00
US20100280365 | A1 | 2010-11-04 | GUIDANCE METHOD BASED ON 3D-2D POSE ESTIMATION AND 3D-CT REGISTRATION WITH APPLICATION TO LIVE BRONCHOSCOPY | 2010-07-12 | A method provides guidance to the physician during a live bronchoscopy or other endoscopic procedures. The 3D motion of the bronchoscope is estimated using a fast coarse tracking step followed by a fine registration step. The tracking is based on finding a set of corresponding feature points across a plurality of consecutive bronchoscopic video frames, then estimating for the new pose of the bronchoscope. In the preferred embodiment the pose estimation is based on linearization of the rotation matrix. By giving a set of corresponding points across the current bronchoscopic video image, and the CT-based virtual image as an input, the same method can also be used for manual registration. The fine registration step is preferably a gradient-based Gauss-Newton method that maximizes the correlation between the bronchoscopic video image and the CT-based virtual image. The continuous guidance is provided by estimating the 3D motion of the bronchoscope in a loop. Since depth-map information is available, tracking can be done by solving a 3D-2D pose estimation problem. A 3D-2D pose estimation problem is more constrained than a 2D-2D pose estimation problem and does not suffer from the limitations associated with computing an essential matrix. The use of correlation-based cost, instead of mutual information as a registration cost, makes it simpler to use gradient-based methods for registration. | 1. A video-based method for providing a pose estimate of an endoscope in conjunction with a live endoscopic procedure, the method comprising: acquiring 3D image data of a target structure in advance of a live endoscopic procedure;receiving a frame of live endoscopic video image data including the target structure; andregistering the frame of endoscopic video image data with acquired 3D image data to determine a pose estimate of the endoscope. 2. The method of claim 1, wherein the pose estimate is used to guide the endoscope to a suspect site. 3. The method of claim 2, wherein the suspect site is a mediastinal lymph node. 4. The method of claim 1, wherein the endoscope is a bronchoscope. 5. The method of claim 1, wherein the target structure is the airway tree. 6. The method of claim 5, wherein the registration is performed either near to, or far from, a bifurcation I an airway tree. 7. The method of claim 1, wherein the registration comprises an initial registration step to initiate guidance of the endoscope. 8. The method of claim 1, wherein the registration step includes the steps of: a) estimating a three-dimensional location of the endoscope using (i) known motion information from said live endoscopic video image data, and (ii) local depth information obtained from the previously acquired 3D image data; andb) determining a new pose of the endoscope based on the 3D location estimated in step (a). 9. The method of claim 8, further including the step of performing a fine registration step to minimize errors associated with estimating the 3D motion of the endoscope. 10. The method of claim 9, wherein the fine registration step is performed manually. 11. The method of claim 8, wherein the registration step further includes the steps of: (c) updating the previously acquired 3D image data in accordance with the new pose; and(d) repeating steps (a) through (c) until the guidance is terminated. 12. The method of claim 1, including the step of receiving a sequence of consecutive video image frames as opposed to a single frame. 13. The method of claim 12, wherein the sequence of video image frames is 5 consecutive video frames. 14. A video-based method for registering previously acquired 3D image data and live endoscopic video image data of a patient to obtain a current pose of an endoscope comprising: a) performing an initial registration to register the previously acquired 3D image data and a frame of the live video image data to obtain a current pose and a current depth map;b) selecting a plurality of points associated with the live video image data;c) tracking the points over a plurality of consecutive frames to estimate the two-dimensional (2D) motion of the tracked points;d) deriving a three-dimensional (3D) motion of the endoscope using the 2D motion of the tracked points and the current depth map;e) determining a new pose based on the current depth map and 3D motion of the endoscope; andf) updating the current pose based upon the ne... | The Penn State Research Foundation | A61B005/05 G06K009/00
US07844474 | B2 | 2010-11-30 | System and method for the creation of a virtual observation and access channel in medical 3D images | 2004-07-23 | In a system for the producing and displaying a virtual access channel in medical images, having a computer with a monitor, and a position and orientation-capturing operating unit and a method for producing and displaying such a virtual access channel in medical 3D images, the computer contains a preoperatively obtained 3D data record of a patient to be examined and can display this as a 3D image and is connected with the operating unit such that a virtual optical channel of the 3D data record is excised from the data record based on a registration between the operating unit and the data record as well as based on different channel parameters, and is displayed on the monitor of the computer. | 1. A system for producing a virtual access channel in displayed medical 3D images, comprising: a computer containing a pre-operatively-obtained 3D data record of internal anatomy of a patient said computer being configured to reconstruct a 3D virtual image of said internal anatomy from said 3D data record;a monitor connected to said computer;a position and orientation capturing operating unit connected to said computer that is manually manipulable to simulate a movement and a position of a medical instrument, said operating unit generating and emitting an operating unit output signal that represents said simulated position and simulated movement; andsaid computer being supplied with said operating unit output signal and producing a registration between the simulated position and simulated movement represented by said operating unit output signal and the internal anatomy of the patient in said 3D data record and thus in said 3D virtual image of said internal anatomy, and said computer continuously excising a virtual optical channel in said 3D virtual image of said internal anatomy of the patient in real time and interactively with the manual manipulation of said position and orientation capturing operating unit, based on said registration and based on at least one channel parameter selected from the group consisting of a channel diameter, a channel depth, a shape of a channel wall, and a shape of a channel base and entered into said computer via said operating unit, and displaying, on said monitor, said virtual 3D image of said internal anatomy with said excised virtual optical channel therein as a perspective virtual fly-through of said internal anatomy along said optical channel. 2. A system as claimed in claim 1 wherein said operating unit comprises a unit selected from the group consisting of a mechanically articulated system, a navigation system, a rod-shaped medical instrument with a navigation sensor, an optical marking of a computer input system, and a magnetic marking of a computer input system. 3. A system as claimed in claim 1 wherein said computer leaves selected vessels or organs in said 3D data record undisturbed in said virtual channel for causing said selected vessels or organs to be displayed on said monitor within said channel or extending into said channel. 4. A system as claimed in claim 1 wherein said 3D data record contained in said computer is a data record selected from the group consisting of a magnetic resonance tomography 3D data record, a computed tomography 3D data record, an ultrasound 3D data record, a positron emission tomography 3D data record, and a 3D data record obtained by a nuclear medicine procedure. 5. A system as claimed in claim 1 wherein said computer automatically excises said virtual optical channel by a computation employing an algorithm. 6. A system as claimed in claim 1 wherein said computer additionally causes a representation of said medical instrument to be displayed in said virtual channel. ... | Siemens Aktiengesellschaft | A61B005/00 G06F019/00
US08536866 | B2 | 2013-09-17 | Frequency referencing for chemical exchange saturation transfer (CEST) MRI | 2008-09-26 | A water saturation shift referencing (WASSR) technique performed using a magnetic resonance scanner comprises: acquiring a spatial map of Z spectra that encompass the water center frequency using sufficiently low saturation power and sufficiently short duration selected such that symmetry of the Z spectra is not obscured by magnetization transfer but dominated by direct water saturation effects so that the spectrum is substantially symmetric; and performing a symmetry analysis on the substantially symmetric Z spectra to generate a spatial map of the water center frequency. WASSR-corrected chemical exchange saturation transfer (CEST) imaging is disclosed as an illustrative example. | 1. A magnetic resonance method comprising: acquiring a spatial map of Z-spectra that encompass the water center frequency using saturation power and duration selected such that the Z-spectra are not obscured by magnetization transfer and are dominated by direct water saturation effects so that the Z-spectra are substantially symmetric; andperforming a symmetry analysis on the substantially symmetric Z-spectra to generate a spatial map of the water center frequency. 2. The magnetic resonance method as set forth in claim 1, wherein the saturation duration is selected to be substantially shorter than a duration effective for generating an interfering magnetization transfer effect in the Z-spectra. 3. The magnetic resonance method as set forth in claim 2, wherein the saturation power is selected to be substantially lower than a power effective for generating an interfering magnetization transfer effect in the Z-spectra. 4. The magnetic resonance method as set forth in claim 1, wherein the acquiring a spatial map of substantially symmetric Z-spectra comprises: acquiring images at different saturation frequencies effective to provide a substantially symmetric Z-spectrum at all pixel or voxel locations of the spatial map. 5. The magnetic resonance method as set forth in claim 4, wherein the acquiring of images comprises acquiring images at a sufficient number of different saturation frequencies collectively encompassing the water center frequency. 6. The magnetic resonance method as set forth in claim 4, further comprising: interpolating or fitting the Z-spectra to enhance saturation frequency resolution prior to the symmetry analysis. 7. The magnetic resonance method as set forth in claim 6, wherein the interpolating or fitting comprises: interpolating or fitting the Z-spectra using an interpolation or fitting approach selected from a group consisting of a cubic spline interpolation algorithm, a polynomial interpolation algorithm, an algorithm employing Fourier transform followed by line broadening, zerofilling and inverse Fourier transform. 8. The magnetic resonance method as set forth in claim 1, further comprising: acquiring a spatial map of asymmetric Z-spectra that encompass the water center frequency and that encompass a CEST-related feature using sufficiently high saturation power and sufficiently long duration selected such that the Z-spectra exhibit the CEST-related feature; andcorrecting the CEST-related feature in the asymmetric Z-spectra by removing spectral components that are symmetric about the water center frequency identified by the spatial map of the water center frequency. 9. The magnetic resonance method as set forth in claim 8, further comprising: generating a water saturation shift referencing (WASSR)-corrected CEST map derived from the corrected CEST-related feature in the asymmetric Z-spectra; anddisplaying an image corresponding to the WASSR-corrected CEST map. 10. The magnetic resonance method as set forth in claim 1,... | The Johns Hopkins University | G01V003/00
US09198582 | B2 | 2015-12-01 | Determining a characteristic physiological parameter | 2009-06-30 | The present disclosure relates to monitoring a characteristic physiological parameter of a patient during a suitable time period that either precedes or follows a triggering event, such as a clinician/patient interaction, that may negatively impact the physiological parameter. In some embodiments, physiological parameter values falling between one or more pre-set thresholds may be used to derive the characteristic physiological parameter. In some embodiments, tracking the physiological parameter may provide additional information about the patient's status. In some embodiments, confidence measures may be associated with, or may be used to analyze features of the patient signal to derive information about, the characteristic physiological parameter. The patient signal used to derive a patient's physiological parameter may be of an oscillatory nature or may include oscillatory features that may be analyzed to derive a characteristic blood pressure or a characteristic respiration rate. | 1. A method for determining a characteristic value of a physiological parameter of a patient in a clinical setting, the method comprising: monitoring, using a processor, values of the physiological parameter of the patient across a first time period and a second time period that is subsequent to the first time period based at least in part on a signal obtained from a patient sensor, wherein the first time period does not overlap with the second time period and wherein the first time period and the second time period are defined by the occurrence of a physiological event caused by the clinical setting;identifying, using a processor, the physiological event based on the values of the physiological parameter in the first time period and based on the patient being in the clinical setting;determining, using a processor, the characteristic value of the physiological parameter of the patient based on the values of the physiological parameter in the second time period wherein the characteristic value of the physiological parameter is determined to assess a clinical status of the patient, wherein the physiological parameter is a blood pressure, wherein the physiological event is a blood pressure event, and wherein the characteristic value of the physiological parameter is a characteristic value of blood pressure;determining, using a processor, an initial value of the blood pressure based at least in part on a value of the blood pressure in the first time period; andidentifying, using a processor, an endpoint defined by a relaxation time following the blood pressure event, wherein the determining the characteristic value of the blood pressure is based at least in part on a value of the blood pressure after the endpoint and wherein the relaxation time is identified based at least in part on a slope of the values of the blood pressure across the second time period reaching a threshold value. 2. The method of claim 1 wherein the characteristic value of the blood pressure is continuously determined. 3. The method of claim 1, further comprising generating at least one confidence measure associated with the characteristic value of the blood pressure. 4. The method of claim 1, wherein a difference between the initial value of the blood pressure value and the characteristic value of the blood pressure is analyzed at least in part to indicate a stress of the patient. 5. The method of claim 4, wherein the difference between the initial value of the blood pressure value and the characteristic value of the blood pressure is normalized by one of a group consisting of the initial value of the blood pressure, the characteristic value of the blood pressure, and a weighted mean of the initial value of the blood pressure and the characteristic value of the blood pressure. 6. The method of claim 1, wherein the relaxation time is analyzed at least in part to indicate an ability of the patient to respond to stressful situations. 7. The method of claim 6, wherein the rela... | Nellcor Puritan Bennett Ireland | A61B005/021 A61B005/024 A61B005/08 A61B005/1455 A61B005/00 A61B005/0215 A61B005/022 A61B005/16
US08256414 | B2 | 2012-09-04 | Neonatal colorimetric carbon dioxide detector | 2009-05-26 | An application for a neonatal calorimetric carbon dioxide detector has a calorimetric carbon dioxide detector membrane having a pH-sensitive chemical indicator that undergoes calorimetric change in the presence of carbon dioxide. The detector has a patient orifice in fluid communication with the baby's airway and a respiration equipment orifice connected to a breathing system. The patient orifice is connected to a breathing tube and when the breathing tube is inserted correctly into the trachea, as the baby exhales, carbon dioxide interacts with the calorimetric membrane which changes color based upon the concentration of carbon dioxide. The neonatal calorimetric carbon dioxide detector adds a volume of less than or equal to 1 mL after being attached to a breathing circuit. | 1. A neonatal colorimetric carbon dioxide detector, the neonatal colorimetric carbon dioxide detector formed to insert into an airway circuit, the neonatal colorimetric carbon dioxide detector comprising: an enclosure having a patient orifice in fluid communication with a patient's airway and a respiration equipment orifice, the respiration equipment orifice in fluid communication with the patient orifice, the respiration equipment orifice being partially occluded thereby reducing an internal volume of the respiration equipment orifice; anda colorimetric membrane held within the enclosure and visible from outside of the enclosure, the colorimetric membrane in fluid communication with the patient orifice such that the colorimetric membrane is exposed to exhalation gas from the patient orifice before the exhalation gas exits through the respiration equipment orifice;whereas insertion of the neonatal colorimetric carbon dioxide detector into the airway circuit increases a total internal volume of the airway circuit by less than or equal to 1 mL. 2. The neonatal colorimetric carbon dioxide detector of claim 1, wherein the colorimetric membrane changes color in response to exposure to carbon dioxide gas. 3. The neonatal colorimetric carbon dioxide detector of claim 1, wherein the total internal volume of the enclosure is 3.8 mL and insertion of the neonatal colorimetric carbon dioxide detector into the airway circuit increases the total internal volume of the airway circuit by 1 mL. 4. The neonatal colorimetric carbon dioxide detector of claim 2, wherein the colorimetric membrane changes color from blue to green in response to exposure to exhalation gas having 1% to 2% of carbon dioxide gas. 5. The neonatal colorimetric carbon dioxide detector of claim 2, wherein the colorimetric membrane changes color to yellow/green in response to exposure to exhalation gas having 2% to 5% of carbon dioxide gas. 6. The neonatal colorimetric carbon dioxide detector of claim 2, wherein the colorimetric membrane changes color to yellow in response to exposure to exhalation gas having greater than 5% of carbon dioxide gas. 7. A neonatal colorimetric carbon dioxide detector with carbon dioxide detection, the neonatal colorimetric carbon dioxide detector formed to insert into an airway circuit, the neonatal colorimetric carbon dioxide detector comprising: an enclosure having a patient orifice in communication with a patient's airway and a respiration equipment orifice, the respiration equipment orifice in fluid communication with the patient orifice, the respiration equipment orifice connected to the patient orifice by two pegs on a bottom surface of the respiration equipment orifice mating with two holes on a top surface of the patient orifice, the respiration equipment orifice being partially occluded thereby reducing an internal volume of the respiration equipment orifice; anda colorimetric membrane held within the enclosure and visible through the enclosure, the... | Mercury Enterprises, Inc. | A61M015/00 A61M016/00 A62B007/00 G01J001/48
US08428699 | B2 | 2013-04-23 | Apparatus, method and computer program for measuring properties of an object | 2008-12-16 | An apparatus measures properties of an object such as a heart in a patient (2). The apparatus includes an antenna (3, 4) for sending and/or receiving electromagnetic waves to and/or from the object (2) for measuring a property of the object. The antenna (3, 4) includes an electrically conductive element for functioning as an electrode for measuring a further property of the object. The electromagnetic waves are reflected by the object, causing a Doppler shift. The apparatus uses this shift to provide mechanical information about the object and uses the electrode to provide electrical information about the object, such as an electrocardiogram. | 1. An apparatus for measuring properties of a moving object in a patient, wherein the apparatus comprises: one or more antennas configured to send electromagnetic waves to the moving object and which receive Doppler shifted electromagnetic waves reflected from the moving object; anda movement determination unit configured to determine movement information correlated to a movement of the object based on the Doppler shifted electromagnetic waves, wherein the one or more antennas comprise an electrically conductive element configured to detect an electrical signal having a frequency which differs from a frequency of the electromagnetic waves sent by the antenna and the Doppler shifted electromagnetic waves received from the moving object, wherein the one or more antennas are configured to function as an electrode to measure a second property of the object. 2. The apparatus as defined in claim 1, wherein the electrically conductive element is a metal plate. 3. The apparatus as defined in claim 1, wherein the electrically conductive element is adapted for an electrocardiogram measurement. 4. The apparatus as defined in claim 1, wherein the apparatus comprises a plurality of the antennas, wherein at least one of the plurality of the antennas is configured for at least one of sending and receiving electromagnetic waves at a frequency, which differs from a frequency of electromagnetic waves of another of the plurality of antennas. 5. The apparatus as defined in claim 1, wherein the apparatus comprises a plurality of antennas which are adapted for forming a beam of the electromagnetic waves by using a phased array of the antennas. 6. The apparatus as defined in claim 1 wherein the antenna is configured to be attached to a surface of the patient for preventing a relative movement between the surface and the antenna. 7. The apparatus as defined in claim 1, wherein the second property of the object which is measurable by detecting the electrical signal by the electrically conductive element is an electrocardiac signal. 8. The apparatus as defined in claim 1 wherein the object is a heart of the patient, wherein the apparatus comprises a property determination unit for determining at least one of a blood pressure and an arterial stiffness based on an electrocardiac signal measured by the electrically conductive element functioning as an electrode of at least two of the antennas and based on a mechanical activity of the heart determined from the electromagnetic waves sent to and received from the object by the at least two of the antennas. 9. The apparatus as defined in claim 1 further comprising an electromagnetic wave generating and signal processing unit electrically connected to at least one of the one or more antennas and comprising a capacitor, a low-pass filter, and an amplifier, the low-pass filter connected between and electrically connected to the antenna and amplifier. 10. The apparatus as defined in claim 9, wherein the second property of t... | Koninklijke Philips Electronics N.V. | A61B005/04
US08428683 | B2 | 2013-04-23 | Wearable monitoring apparatus and driving method thereof | 2010-02-08 | Disclosed are a wearable monitoring apparatus and a driving method thereof. The wearable monitoring apparatus comprises: a sensor unit for measuring a biological signal from a human body, wherein the sensor unit is adhered to a skin; and a control unit for searching a location of the sensor unit, supplying power to the sensor unit, and receiving and processing the biological signal from the sensor unit, wherein the control unit is formed to be wearable. | 1. A wearable monitoring apparatus for continuously monitoring a biological signal, comprising: a sensor unit for measuring a biological signal from a human body, wherein the sensor unit is configured to be adhered to a skin;a control unit that transmits a searching signal for searching a location of the sensor unit;wherein upon receiving the searching signal, the sensor unit transmits a response signal to the control unit, the response signal containing information corresponding to a type of the sensor unit;wherein after the sensor unit transmits the response signal, the control unit supplies power to the sensor unit, and the sensor unit measures the biological signal and transmits the biological signal to the control unit; andwherein the control unit is formed to be wearable. 2. The wearable monitoring apparatus according to claim 1, wherein the sensor unit comprises: an adhesive layer having adhesion and flexibility;an inductor that operably connects the sensor unit to the control unit by wireless means using near field coupling;a sensing electrode for sensing the biological signal;a sensor chip for transmitting the response signal containing the information corresponding to the type of the sensor unit to the control unit in response to receiving the searching signal from the control unit, measuring the biological signal using the sensing electrode, and transmitting the measured biological signal to the control unit; anda molding portion formed to cover the sensor chip to prevent the sensor chip from being damaged,wherein at least one of the inductor and the sensing electrode is formed by printing a conductive material on a fiber. 3. The wearable monitoring apparatus according to claim 2, wherein the control unit comprises: a plurality of inductors that operably connect the control unit to the sensor unit by wireless means using near field coupling, wherein the plurality of inductors are arrayed in a coordinate matrix pattern;a power supply for transmitting the searching signal and supplying the power to the sensor unit; anda network controller that; (1) transmits the searching signal in a set coordinate order using the plurality of inductors, (2) receives the response signal from the sensor unit, (3) stores the information corresponding to the type of the sensor unit, (4) stores coordinate information of one inductor of the plurality of inductors that received the response signal, (5) supplies the power to the sensor unit in an amount that meets a set sensing cycle and power transmission strength based on the type of the sensor unit using the one inductor having received the response signal, (6) receives the biological signal measured by the sensor unit, and (7) stores the received biological signal;wherein the network controller is connected to the plurality of inductors; andwherein the plurality of inductors are formed by printing a conductive material on a fiber. 4. The wearable monitoring apparatus according to claim 1, wherein the sen... | Korea Advanced Institute of Science and Technology | A61B005/04
US08352008 | B2 | 2013-01-08 | Active noise cancellation in an optical sensor signal | 2010-06-10 | A medical device system and associated method are used for monitoring tissue oxygenation. An optical sensor produces a signal corresponding to tissue light attenuation. A processor receives the optical sensor signal and computes a first measure of light attenuation at a first light wavelength and a second measure of light attenuation at a second light wavelength. In one embodiment, noise cancellation circuitry receives the first measure and the second measure and generates a guessed ratio of the first and second measures. Using the first measure, the second measure and the guessed ratio, the noise cancellation circuitry provides a peak output power when the guessed ratio corresponds to an actual ratio of the first and second measures. | 1. A medical device system, comprising: an optical sensor for producing a signal corresponding to tissue light attenuation;a processor receiving the optical sensor signal and configured to compute a first measure of light attenuation at a first light wavelength and a second measure of light attenuation at a second light wavelength; andnoise cancellation circuitry for receiving the first measure and the second measure and generating an estimated ratio of the first and second measures, the noise cancellation circuitry providing a peak output power when the guessed ratio corresponds to an actual ratio containing the first measure;the processor determining a tissue oxygenation using the estimated ratio corresponding to the peak output power of the noise cancellation circuitry. 2. The system of claim 1, wherein the optical sensor produces a signal corresponding to tissue light attenuation for at least four wavelengths, and the processor being configured to compute the first measure as a second derivative of light attenuation with respect to a first light wavelength and the second measure as a second derivative of light attenuation with respect to a second light wavelength. 3. The system of claim 1, further comprising a memory to store calibration data relating values of the ratio to tissue oxygen saturation, wherein the processor is configured to compute an absolute tissue oxygen saturation using the estimated ratio corresponding to the peak output power of the noise cancellation circuitry and the calibration data. 4. The system of claim 1, wherein the processor is configured to compute a measure of tissue oxygen saturation using the estimated ratio and compute a measure of total hemoglobin volume fraction using the computed tissue oxygen saturation and a noise cancellation circuitry output signal corresponding to a peak output power. 5. The system of claim 1, further comprising a memory to store a plurality of bins, each bin of the plurality of bins defining a range of possible values for the ratio, wherein the processor is configured to determine the tissue oxygenation by identifying one of the plurality of bins associated with a peak output power of the noise cancellation circuitry. 6. The system of claim 5, wherein determining the tissue oxygenation comprises detecting hypoxia in response to identifying the one of the plurality of bins. 7. The system of claim 1, wherein the noise cancellation circuitry is configured to generate estimated values for the ratio in a binary search mode. 8. The system of claim 1, further comprising a noise detector generating an enable signal for enabling the noise cancellation circuitry. 9. The system of claim 8, further comprising a motion sensor coupled to the noise detector, the noise detector generating the enable signal in response to the motion sensor. 10. The system of claim 8, wherein the noise detector determines a noise measurement of a high frequency content of the optical sensor signal and enables... | Medtronic, Inc. | A61B005/00
US09451910 | B2 | 2016-09-27 | Transcutaneous analyte sensor | 2010-08-27 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. A sensor system for measuring an analyte concentration in a host, the system comprising: a sensor configured to continuously measure an analyte concentration in a host;a sensor electronics module configured to electrically couple to the sensor, wherein the sensor electronics module comprises sensor electronics configured to provide a signal associated with the analyte concentration in the host, a radio frequency identification device, an internal power supply configured to supply power to the electronics module, and a switch operatively connected to the power supply, the switch configured to switch between a first state and a second state, wherein the internal power supply is configured to be in a first power mode when the switch is in the first state and in a second power mode when the switch is in the second state, the first power mode supplying less power to the sensor electronics module than the second power mode, and wherein the internal power supply provides power to the sensor electronics sufficient to provide a bias voltage for the sensor so to generate the signal associated with the analyte concentration; anda receiver device having a power supply configured to inductively power the radio frequency identification device when the receiver is placed in proximity to the sensor electronics module, a processor configured to generate estimated analyte concentration values, and a display configured to display the estimated analyte data values,wherein the radio frequency identification device is configured to transmit the information representative of the measured analyte concentration and a calibration code responsive to the receiver device interrogating the sensor electronics module, wherein the external power source provides the power to the radio frequency identification device so transmission of the information from the radio identification device to the receiver results in little or no impact on the internal power supply, andwherein generating the estimated analyte concentration values is based on the information representative of the measured analyte concentration and the calibration code. 2. The system of claim 1, wherein the internal power source comprises a battery and wherein the switch is a reed switch connected to an output terminal of the battery. 3. The system of claim 1, further comprising a separate unit configured to engage the sensor electronics module, the separate unit having a magnet attached thereto configured to maintain one of the first state or the second state of the switch when the separate unit is engaged with the sensor electronics module. 4. The system of claim 1, further comprising a mounting unit configured to receive the sensor, wherein the mounting unit is adapted for placement adjacent to the host's skin, wherein the sensor electronics module includes a housing configured to be releasably attached to the mounting unit. 5. The system of claim 4, wherein the switch is configured to be in the second state ... | DexCom, Inc. | A61B005/05 A61B005/1473 A61B005/15 A61B005/145 A61B005/1486 A61B005/00
US20100280372 | A1 | 2010-11-04 | Observation device and method | 2010-05-03 | An observation device and method for observing a plurality of physiological and/or anatomical parameters is disclosed. The exemplary embodiment of the observation device is portable and handheld so that it may be used in an emergency situation and is field-deployable. The exemplary embodiment includes an image collector for observing, measuring, and/or recording at least one parameter, which parameter may then be used in analysis and diagnosis of the subject's condition or lack thereof. | 1. An observation device comprising: a. a base member, wherein said base member is portable and includes an image collector; andb. a processing unit, wherein said processing unit is operable to derive at least one metric from an image collected by said image collector. 2. An observation device comprising: a. a base member, said base member comprising: i. at least one image collector;ii. at least one illuminator;iii. a processor unit;b. a handle member, said handle member comprising at least one stimulus generator. 3. The observation device according to claim 2 wherein said handle member further comprises a user interface. 4. The observation device according to claim 3 wherein said handle member further comprises a user input. 5. The observation device according to claim 4 wherein said user interface is further defined as including a user input and a user feedback generator. 6. The observation device according to claim 2 further comprising a storage media positioned in either said handle member or said base member. 7. The observation device according to claim 2 further comprising an electronic communication interface positioned in either said handle member or said base member. 8. The observation device according to claim 2 further comprising a position sensing component for sensing the position of said stimulus generator. 9. The observation device according to claim 2 further comprising a position sensing component for sensing the position of said image collector. 10. The observation device according to claim 1 further comprising a component for observing electrophysiology. 11. The observation device according to claim 1 further comprising a component for performing ultrasonography. 12. A method for observing a subject comprising: a. providing a stimulus to said subject;b. observing the subject's response to said stimulus; andc. correlating the subject's response to a condition of said subject. 13. The method according to claim 12 wherein said stimulus is further defined as a visual stimulus. 14. The method according to claim 13 wherein a handheld observation device is used to provide said stimulus. | - | A61B008/00 A61B005/05
US20100324454 | A1 | 2010-12-23 | DECEPTION DETECTION USING OCULOMOTOR MOVEMENTS | 2010-06-22 | Embodiments of the present invention relate to a rapid, automated, and objective method for using oculomotor measures to determine whether a person is being truthful or deceitful. More specifically, embodiments of the present invention measure a plurality of oculomotor and behavioral dependent measures while a subject reads and responds to written items that are both related and unrelated to a suspected crime. The oculomotor and behavioral measures may include pupil diameter, response times, the number of fixations during reading, the time spent reading and rereading items, and the rate of eye blinks on the current and subsequent item. Several behavioral and oculomotor measures were diagnostic of deception, and a weighted combination of four of those variables correctly classified 84% of guilty and 89% of innocent subjects. | 1. A computer-implemented method for assessing the veracity of a subject comprising: presenting the subject with one or more sets of written test items relating to one or more issues of relevance;presenting the subject with one or more sets of written test items relating to one or more neutral issues;presenting the subject with one or more sets of written test items relating to one or more comparison issues;monitoring, by a computing device, one or more dependent measures while the subject is reading and responding to said test items; andanalyzing, by the computing device, the variance of said dependent measures to assess veracity of said subject. 2. The method for assessing the veracity of a subject of claim 1, wherein the dependent measure is selected from the group consisting of response time, proportion wrong, number of fixations, first past duration, reread duration, pupil diameter, item blink rate and next item blink rate. 3. The method for assessing the veracity of a subject of claim 1, wherein a plurality of oculomotor dependent measures are analyzed to assess the veracity of said subject. 4. The method for assessing the veracity of a subject of claim 1, wherein the variance of a plurality of dependent measures as selected from Table 4 are analyzed to assess the veracity of said subject. 5. The method for assessing the veracity of a subject of claim 1, wherein the test items relating to issues of relevance, neutral issues and comparison issues comprises staggering the test items. 6. The method for assessing the veracity of a subject of claim 2, wherein a plurality of dependent measures are analyzed using a discriminant function and said discriminant function yields a sensitivity of at least 80%. 7. The method for assessing the veracity of a subject of claim 6, wherein a plurality of dependent measures are analyzed using a discriminant function and said discriminant function yields a specificity of at least 80%. 8. A computer-implemented method for assessing the veracity of a subject comprising: presenting the subject with one or more sets of written test items relating to one or more issues of relevance;presenting the subject with one or more sets of written test items relating to one or more neutral issues;presenting the subject with one or more sets of written test items relating to one or more comparison issues;wherein the subject is instructed to respond to the written test items as quickly and accurately as possible;monitoring, by a computing device, one or more dependent measures while the subject is reading and responding to said test item; andanalyzing, by the computing device, the variance of said dependent measures together with the results from one or more independent variables to assess the veracity of said subject. 9. The method for assessing the veracity of a subject of claim 8, wherein the dependent measure is selected from the group consisting of response time, proportion wrong, number of fixations, first past dura... | - | A61B005/103
US20100324827 | A1 | 2010-12-23 | Fluid Responsiveness Measure | 2009-06-18 | A method and system for measuring fluid responsiveness of a patient is disclosed. Information related to fluid responsiveness of a patient may be derived from a PPG signal, for example, by analyzing the PPG signal transformed by a continuous wavelet transform. Other techniques for deriving information related to fluid responsiveness of a patient include, for example, analyzing the amplitude modulation, frequency modulation, and/or baseline changes of a PPG signal. | 1. A method for measuring fluid responsiveness of a patient, the method comprising: using processing equipment for: transforming a physiological signal based at least in part on a wavelet transform to generate a transformed signal,generating a scalogram from the transformed signal, andanalyzing the scalogram to determine information related to the fluid responsiveness; andoutputting to an output device the information. 2. The method of claim 1, wherein the physiological signal comprises a plethysmographic signal, an electrocardiographic signal, an electromyographic signal, and/or an arterial line trace signal, and/or combinations thereof. 3. The method of claim 1, wherein the analyzing comprises identifying carrier wave information from the physiological signal. 4. The method of claim 1, wherein the analyzing comprises identifying respiratory sinus arrhythmia information from the physiological signal. 5. The method of claim 1, wherein the analyzing comprises identifying amplitude modulation information from the physiological signal. 6. The method of claim 5, wherein the physiological signal is a plethysmographic signal and wherein the analyzing further comprises deriving amplitude modulation information of a pulse band ridge of the scalogram. 7. The method of claim 1, wherein the information comprises a relationship between a characteristic of the amplitude modulation of a particular region of the scalogram to a baseline. 8. The method of claim 7, wherein the characteristic of the amplitude modulation comprises a peak-to-peak amplitude modulation, a standard deviation of the amplitude modulation, and/or a median absolute value of the amplitude modulation, and/or combinations thereof. 9. The method of claim 7, wherein the baseline comprises a mean value of the amplitude modulation, a lower bound of the amplitude modulation, and/or an upper bound of the amplitude modulation, and/or combinations thereof. 10. The method of claim 1, further comprising adjusting a fluid delivery mechanism based at least in part on the information related to fluid responsiveness. 11. A system for measuring fluid responsiveness of a patient, the system comprising: a fluid delivery mechanism capable of supplying fluid to the patient;a sensor attached to the patient capable of generating a physiological signal;a memory;a processor coupled to the memory capable of: transforming the physiological signal based at least in part on a wavelet transform stored in the memory to generate a transformed signal,generating a scalogram from the transformed signal, andanalyzing the scalogram to determine information related to the fluid responsiveness; andan output device coupled to the processor. 12. The system of claim 11, wherein the physiological signal comprises a plethysmographic signal, an electrocardiographic signal, an electromyographic signal, and/or an arterial line trace signal, and/or combinations thereof. 13. The system of claim 11, wherein the processor is furt... | Nellcor Puritan Bennett Ireland | A61B005/0452 A61B005/02 A61B005/0488 G06F019/00
US20100286545 | A1 | 2010-11-11 | ACCELEROMETER BASED HEALTH SENSING | 2009-05-06 | Health-sensing and health-action devices and systems are generally described. The health-sensing device may include one or more of a sensor, a filter, and a transmitter. The sensor may be configured to sense one or more factors relating to an indicator of a health related condition or occurrence such as sleep apnea and may include one or more microphone devices, accelerometers, and/or MEMs devices. The filter may be configured to evaluate a signal from the sensor and determine when the indicator has been detected. The transmitter may be arranged for initiating a transmission based on a signal from the filter. The health-action device may be configured for responding to an indicator of a health related condition or occurrence of a user and may include one or more of a receiver, a processor, and a responder. | 1. A health-sensing device to monitor a user for a health related condition, the device comprising: a substrate;a sensor positioned on the substrate and configured to: monitor a factor relating to an indicator of the health related condition or occurrence, andgenerate a sensor signal in response to the monitored factor;a filter positioned on the substrate and configured to: receive the sensor signal,determine when the indicator has been detected, andgenerate a filter signal when the indicator is determined to be detected; anda transmitter positioned on the substrate and configured to: receive the filter signal, andgenerate a transmission responsive to the filter signal. 2. The device of claim 1, wherein the substrate comprises a silicon based element and further includes an enveloping fabric. 3. The device of claim 1, wherein the factor relating to the indicator of the health related condition or occurrence corresponds to motion of the user, and wherein the sensor corresponds to an accelerometer. 4. The device of claim 3, further comprising an affixing element adapted to secure the substrate to the user, the affixing element creating an association between the sensor of the device and the motion of the user. 5. The device of claim 4, wherein the indicator of the health related condition corresponds to a lapse in breathing. 6. The device of claim 5, wherein, to determine when the indicator has been detected, the filter is further configured to time the lapse in breathing to provide a timed result, and compare a timed result to a threshold value. 7. The device of claim 6, wherein the health related condition is sleep apnea and the threshold value is in a range from approximately 10 seconds to approximately 20 seconds. 8. A method for a health-sensing device to monitor a user for a health related condition, the method comprising: sensing motion of the health-sensing device with a sensor to generate a sensor signal, wherein the motion is also associated with motion of the user;analyzing the sensor signal to determine when the motion associated with the health-sensing device identifies an indicator of the health related condition; andinitiating a transmission responsive to the analysis of the sensor signal. 9. The method of claim 8, wherein initiating a transmission includes selectively transmitting the signal when the indicator of the health related condition is identified. 10. The method of claim 9, further comprising encoding the transmission with information including a determination of the indicator of the health related condition as identified. 11. The method of claim 10, further comprising receiving the transmission with a receiver and performing a particular process. 12. The method of claim 11, wherein performing the particular process includes providing a stimulation to the user with a stimulation device. 13. The method of claim 12, wherein performing the particular process includes providing a notification to a third party. 14. ... | - | A61B005/08 A61N001/00
US20100280345 | A1 | 2010-11-04 | Analyte Monitoring Device and Methods of Use | 2010-07-15 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for evaluating a change in a sensitivity of an analyte sensor over a predetermined time period, the method comprising: receiving sensor data from an analyte sensor, wherein the sensor data comprises one or more sensor analyte values measured in a biological sample of a host;intermittently calculating a sensitivity of the analyte sensor based at least in part on reference analyte data; andevaluating a change in sensitivity by evaluating a plurality of time-spaced sensitivity calculations over a predetermined time period. 3. The method of claim 2, wherein the step of evaluating a change in sensitivity comprises evaluating at least two sensitivity measurements during the predetermined time period. 4. The method of claim 3, further comprising averaging or filtering the at least two sensitivity measurements prior to the step of evaluating the at least two sensitivity measurements. 5. The method of claim 2, wherein the step of evaluating a change in sensitivity comprises comparing the change in sensitivity with one or more criteria. 6. The method of claim 5, further comprising using a most recent sensitivity calculation evaluated to calibrate the analyte sensor when the change in sensitivity meets one or more criteria. 7. The method of claim 5, further comprising not using a most recent sensitivity calculation evaluated to calibrate the analyte sensor when the change in sensitivity does not meet one or more criteria. 8. The method of claim 2, wherein the step of intermittently calculating a sensitivity of the analyte sensor comprises intermittently receiving a reference analyte value. 9. The method of claim 2, wherein the step of intermittently receiving a reference analyte value comprises periodically receiving a measured analyte concentration from a reference analyte solution. 10. The method of claim 2, wherein the step of intermittently receiving a reference analyte value comprises receiving a reference analyte value obtained from an in vitro analyte monitor. 11. The method of claim 2, wherein the step of evaluating a sensitivity is iteratively performed on the plurality of time-spaced sensitivity calculations over the predetermined time period. 12. The method of claim 2, wherein the predetermined time period is at least about 30 minutes. 13. The method of claim 12, wherein the predetermined time period is at least about 60 minutes. 14. The method of claim 12, wherein the predetermined time period is at least about 120 minutes. 15. The method of claim 12, wherein the step of evaluating a sensitivity is iteratively performed on the plurality of time-spaced sensitivity calculations over the predetermined time period. 16. The method of claim 12, wherein the step of evaluating a sensitivity comprises evaluating all sensitivity calculations over a sensor session. 17. The method of claim 12, wherein the step of evaluating a sensitivity is based at least in part on a priori sensitivity information. 18. The method ... | Abbott Diabetes Care Inc. | A61B005/05
US07828753 | B2 | 2010-11-09 | Electrode interface system | 2006-10-18 | An electrode interface system for providing a connection between at least one electrode and a maternal-fetal monitor, wherein the interface system converts electrical muscle activity captured by the electrode(s) into uterine activity data signals for use by the maternal-fetal monitor. The electrode interface system of the invention preferably includes a conversion means for converting the signals from the electrode(s) into signals similar to those produced by a tocodynometer. | 1. A method for monitoring progress of labor, comprising: attaching at least one electrode onto a pregnant maternal abdomen;attaching to the electrode(s) an interface cable comprising, a first component that is an electrode interface that connects to the electrode(s) and that receives at least one electrode signal from the electrode(s),a second component that is a signal converter that connects to the electrode interface, wherein said signal converter receives the electrode signal(s) from the electrode interface and processes same said electrode signal(s) into electrical output data that mimics electrical output from a tocodynamometer or intra-uterine pressure catheter, anda third component that is a maternal-fetal monitor port interface that connects to a maternal-fetal monitor, wherein the maternal-fetal monitor port interface connects to the signal converter and receives and subsequently transmits the electrical output data to the maternal-fetal monitor; wherein the components are integrated to form a unitary structure in the cable; andassessing progress of labor by analyzing the electrical output data using the maternal-fetal monitor. 2. The method of claim 1, wherein the analysis of electrical output data generates any one or combination of data selected from the group consisting of: EHG data; uterine contractions during and after labor; uterine atony; intrauterine pressure; and Montevideo units. 3. The method of claim 1, wherein the pregnant maternal abdomen is of an obese patient. 4. The method of claim 1, wherein the signal converter processes same said electrode signals with an electrical analog of Wheatstone bridge configuration into electrical output data that mimics electrical output from said tocodynamometer. 5. An interface cable comprising: a first component that is an electrode interface for operably connecting to at least one electrode and receiving at least one electrode signal from the at least one electrode,a second component that is a signal converter that is connected to the electrode interface, wherein said signal converter receives the at least one electrode signal from the electrode interface and processes same said electrode signal(s) into electrical output data that mimics electrical output from a tocodynamometer or intra-uterine pressure catheter, anda third component that is a maternal-fetal monitor port interface for operably and physically connecting to a maternal-fetal monitor, wherein the maternal-fetal monitor port interface connects to the signal converter and receives and subsequently transmits the electrical output data to the maternal-fetal monitor,wherein the components are integrated to form a unitary structure in the cable. 6. The interface cable of claim 5, further comprising a means for filtering the at least one signal from the at least one electrode. 7. The interface cable of claim 6, wherein the filtering means is any one or combination of items selected from the group consisting of: high pass f... | Convergent Engineering, Inc. | A61B005/00 A61B005/04 A61B005/0488 A61B005/103 H01R011/00
US07854705 | B2 | 2010-12-21 | Ex vivo verification of biopsy tissue samples | 2004-12-16 | An instrument for verification of presence of image enhancing, contrasting agent in a biopsy sample which was obtained by imaging the lesion area with an imaging modality which is sensitive to the contrasting agent. In one embodiment of the invention, an optical spectrometer is used to analyze biopsy samples for the presence of gadolinium enhanced Magnetic Resonance Imaging contrasting agent which typically pools in breast tumors. | 1. A method for indicating a likely presence of abnormal tissue within an ex vivo biopsy sample of tissue of a living patient comprising: providing observable agent to an in vivo region of surgical interest of the living patient;removing from the living patient a biopsy tissue sample from the region of interest to provide a removed tissue sample;minimally-destructively observing the removed tissue samples under conditions that allow observation or detection of the observable agent; andevaluating results of the observation or detection of the observable agent within the removed tissue sample while the patient is still living to confirm that tissue was removed from the region of surgical interest of a still living patient that had been provided with the observable agent in vivo. 2. The method of claim 1 wherein evaluating results comprises assessing a likelihood of the presence of abnormal tissue in the still living patient within a sample and non-destructively observing the removed tissue. 3. The method of claim 1 wherein evaluating results comprises assessing whether the removed tissue sample has been taken from the region of interest in a still living patient and non-destructively observing the removed tissue. 4. The method of claim 1 wherein non-destructively observing is performed by a step comprising spectroscopic analysis while the patient is still living. 5. The method of claim 1 wherein after evaluating results, histopathic examination is performed on the removed tissue sample while the patient is still living. 6. The method of claim 1 wherein non-destructive observation is used and observing comprises at least one procedure selected from the group consisting of colorimetry, sonography, spectroscopy, and magnetic resonance imaging of the removed tissue sample. 7. The method of claim 6 wherein the spectroscopy is selected from the group consisting of scattered radiation spectroscopy, frustrated total internal reflection spectroscopy, Raman spectroscopy, fluorescent spectroscopy, emission spectroscopy, Fourier transform spectroscopy, flame spectroscopy, laser excited emission spectroscopy, mass spectroscopy, Auger spectroscopy, magnetic resonance spectroscopy. 8. The method of claim 6 where intensity changes or density differences from the removed tissue sample provide information that is evaluated. 9. The method of claim 1 wherein the observable agent is known to be absorbed differently by normal tissue and abnormal tissue. 10. The method of claim 1 wherein non-destructive observation is used and the non-destructive observation comprises assessing concentration of the observable agent in the tissue by measuring or detection of absorption of electromagnetic radiation at least one preselected wavelength or one or more spectral bands in at least one of the UV, visible and IR spectral ranges without the damage to the tested tissue. 11. The method of claim 1 wherein a non-destructive observation is used and the non-destructive observat... | - | A61B005/05 A61B005/055
US07853310 | B2 | 2010-12-14 | Automatic analysis in virtual endoscopy | 2006-12-11 | A computer system and a computer-implemented method are provided for interactively displaying a three-dimensional rendering of a structure having a lumen and for indicating regions of abnormal wall structure. A three-dimensional volume of data is formed from a series of two-dimensional images representing at least one physical property associated with the three-dimensional structure. An isosurface of a selected region of interest is created by a computer from the volume of data based on a selected value or values of a physical property representing the selected region of interest. A wireframe model of the isosurface is generated by the computer wherein the wireframe model includes a plurality of vertices. The vertices are then grouped into populations of contiguous vertices having a characteristic indicating abnormal wall structure by the computer. The wireframe model is then rendered by the computer in an interactive three-dimensional display to indicate the populations of abnormal wall structure. | 1. A method for interactively displaying three-dimensional structures comprising the steps of: a. forming a three-dimensional volume of data representing at least one physical property associated with a three-dimensional body;b. isolating a selected region of interest by comparing the value of each of a selected number of voxels in the volume of data to a value range of selected values of the physical property that represent the selected region of interest to remove those voxels having a value outside the selected value range; andc. rendering the isolated region of interest in an interactive three-dimensional display,wherein the step of isolating the selected region of interest comprises a morphological dilation step. 2. The method as recited in claim 1 wherein the body comprises a human body. 3. The method as recited in claim 1 wherein the forming step comprises an image acquisition step wherein a series of two-dimensional images is acquired from which the three-dimensional volume of data is formed. 4. The method as recited in claim 1 wherein the body comprises a colon. 5. The method as recited in claim 4 wherein the rendering step comprises simulating movement along a line which passes along the center of the lumen of the colon. 6. The method as recited in claim 4 wherein the rendering step comprises splitting the colon open in lengthwise sections and displaying the split open colon so that inner surfaces of the split open colon are visible. 7. The method as recited in claim 6 wherein the splitting step includes the steps of: a. defining a cutting plane through the isolated region of interest; andb. rendering all portions of the three-dimensional display on one side of the cutting plane transparent. 8. The method as recited in claim 4 wherein the isolation step comprises a step of determining colon wall thickness. 9. The method as recited in claim 8 wherein the step of determining the colon wall thickness comprises displaying a cross-sectional image through the colon. 10. The method as recited in claim 8 wherein the step of determining colon wall thickness comprises a step of identifying a thickened section of the colon by visually differentiating a thickened section on a displayed image from normal thickness of the colon wall. 11. A method for interactively displaying three-dimensional structures comprising the steps of: a. forming a three-dimensional volume of data representing at least one physical property associated with a three-dimensional body;b. isolating a selected region of interest by comparing the value of each of a selected number of voxels in the volume of data to a value range of selected values of the physical property that represent the selected region of interest to remove those voxels having a value outside the selected value range; andc. rendering the isolated region of interest in an interactive three-dimensional display wherein the body comprises a tracheobronchial airway, a portion of the tracheobronchial airway... | Wake Forest University Health Sciences | A61B005/00
US20100324393 | A1 | 2010-12-23 | CIRCUIT BOARD FOR BODY FLUID COLLECTION, AND BIOSENSOR | 2009-01-21 | A circuit board for body fluid collection includes a measurement unit including a puncture needle and an electrode for making contact with the body fluid collected by the puncturing with the puncture needle, wherein the measurement unit is provided in a plural number and disposed radially on a same plane. | 1. A circuit board for body fluid collection comprising a measurement unit including a puncture needle and an electrode for making contact with the body fluid collected by the puncturing with the puncture needle, wherein the measurement unit is provided in a plural number and disposed radially on a same plane. 2. The circuit board for body fluid collection according to claim 1, wherein the measurement unit comprises a bending portion that is bendable at an upstream side of a distal end of the puncture needle in the puncturing direction. 3. The circuit board for body fluid collection according to claim 1, wherein an opening is provided at a center of the circuit board for body fluid collection, and an engage portion that can be engaged with a driving member for rotating the circuit board for body fluid collection in a circumferential direction is formed at an inner circumferential surface of the opening of the circuit board for body fluid collection. 4. A biosensor comprising a circuit board for body fluid collection, said circuit board for body fluid collection comprising: a measurement unit including a puncture needle and an electrode for making contact with the body fluid collected by the puncturing with the puncture needle,wherein the measurement unit is provided in a plural number and disposed radially on a same plane, anda determination unit that is electrically connected to the electrode and performs a measurement on a component in body fluid. | Nitto Denko Corporation | A61B005/05
US07840249 | B2 | 2010-11-23 | Clinical micro-CT (CMCT) methods, techniques and apparatus | 2005-11-17 | The present invention relates to a method and system (CMCT system) for improving spatial resolution imaging of CT systems. The systems and method can achieve improved spatial resolution while using CT X-ray dosage levels comparable to those currently used in practice. The system and method can be used for micro-tomography and/or micortomosynthesis of a local region and/or volume of interest in a patient head or another body part. | 1. An imaging method comprising: acquiring global image data using a first tomographic imaging modality comprising a standard CT scanner having a first resolution of a head of a subject;acquiring facial surface data from a facial surface scan of the head of the subject;constructing a facial surface model of the head of the subject from the global image data and the facial surface data;positioning the head of the subject according to the facial surface model such that the subject is centered at an origin of a field of view of the first tomographic imaging modality;acquiring local image data using a second tomographic imaging modality comprising a micro-CT scanner having a second resolution through a region of interest of the head of the subject, wherein the region of interest comprises at least one of, a temporal bone or an inner ear; andconstructing an image of the region of interest from the global image data and the local image data, wherein the local image data is higher in resolution than the global image data. 2. The method of claim 1, wherein the global image data is derived from a database or model of the subject. 3. The method of claim 1, wherein the local and global image data are combined based on characteristics of these data and prior knowledge of imaging geometries of the first and second tomographic imaging modalities used. 4. The method of claim 1, wherein the global image data is acquired prior to the acquisition of the local image data and the location of the local image data acquisition is based on the global image data. 5. The method of claim 1, wherein the facial surface scan is performed prior to acquiring the local image data of the subject. 6. The method of claim 1, further comprising acquiring the local image data after the subject has been positioned. 7. The system of claim 1, further comprising positioning or repositioning a cochlear implant in the subject. 8. The method of claim 7, wherein the cochlear implant is positioned or repositioned prior to or during acquisition of the local image data. 9. The method of claim 1, wherein the subject is a small animal. 10. The method of claim 9, wherein the small animal is selected from the group consisting of a mouse, rat and rabbit. 11. The method of claim 1, further comprising post-processing the acquired local image data. 12. The method of claim 11, wherein the post-processing is selected from the group consisting of pre-processing, correction, deblurring, denoising, enhancement, segmentation, registration, and visualization. 13. The method of claim 1, further comprising post-processing the constructed image. 14. The method of claim 13, wherein the post-processing is selected from the group consisting of pre-processing, correction, deblurring, denoising, enhancement, segmentation, registration and visualization. 15. The method of claim 1, wherein resolution of the constructed image is about 100 microns (μm) or lower. 16. The method of claim 1, wherein the local ... | University of Iowa Research Foundation | A61B005/05
US07826894 | B2 | 2010-11-02 | Cognitive control signals for neural prosthetics | 2005-03-22 | In an embodiment, the invention relates to neural prosthetic devices in which control signals are based on the cognitive activity of the prosthetic user. The control signals may be used to control an array of external devices, such as prosthetics, computer systems, and speech synthesizers. Data obtained from monkeys' movement intentions were recorded, decoded with a computer algorithm, and used to position cursors on a computer screen. Not only the intended goals, but also the value of the reward the animals expected to receive at the end of each trial, were decoded from the recordings. The results indicate that brain activity related to cognitive variables can be a viable source of signals for the control of a cognitive-based neural prosthetic. | 1. A method, comprising: implanting at least one electrode in a brain of a subject;receiving electrical signals generated by a plurality of neurons at the at least one implanted electrode, the electrical signals comprising a portion that encodes at least one abstract cognitive concept, the portion of the electrical signals encoding the at least one abstract cognitive concept not encoding either motor movements or intended motor movements, the plurality of neurons having generated the portion of the electrical signals encoding the at least one abstract cognitive concept without having been trained to do so;at least partially decoding the at least one abstract cognitive concept encoded in the portion of the electrical signals by generating a feature set based on the electrical signals and mapping the feature set to the at least one abstract cognitive concept using a transformation rule; andgenerating a control signal to instruct an external device to perform a task associated with the at least partially decoded at least one abstract cognitive concept. 2. The method of claim 1, wherein the at least one abstract cognitive concept encoded in the portion of the electrical signals is an intended goal of the subject. 3. The method of claim 1, wherein the at least one abstract cognitive concept encoded in the portion of the electrical signals is an expected value, speech segment, abstract thought, executive control, attention, decision, motivation, memory, imagery, emotion, or mood of the subject. 4. The method of claim 1, further comprising building a database of neural activity data detected from the subject. 5. The method of claim 4, wherein the database comprises information collected from at least one brain control trial performed by the subject. 6. The method of claim 1, wherein at least a portion of the plurality of neurons are located in at least one area of the brain of the subject that generates motor signals. 7. The method of claim 1, further comprising causing the external device to perform the task. 8. The method of claim 1, wherein the external device is a prosthetic limb. 9. The method of claim 1, wherein the external device is selected from a group consisting of a computer device, a speech synthesizer, a wheelchair, a vehicle, a controller, a touch pad, a keyboard, a robotic hand, and combinations thereof. 10. The method of claim 1, wherein generating a control signal further comprises deriving multiple control signals from a single area in the brain or from a simultaneous recording from multiple brain areas. 11. The method of claim 1, wherein generating a control signal further comprises deriving multiple control signals from local field potential (“LFP”) recordings, spike activity, or combinations thereof. 12. The method of claim 1, wherein the control signal is an expected value signal. 13. The method of claim 1, wherein the task associated with the at least partially decoded at least one abstract cognitive concept comprises... | California Institute of Technology | A61B005/04
US07838726 | B2 | 2010-11-23 | Methods for introducing heterologous cells into fish | 2006-12-21 | The invention provides methods of introducing heterologous cells into fish. After introduction cells remain viable, and in some instances proliferate, for sufficient time to conduct a variety of analyses on the heterologous cells or the fish or both. Such methods are useful for screening potential drugs for toxicity toward introduced cells or for capacity to stimulate differentiation and/or proliferation of introduced cells. Such methods are also useful for diagnosing the presence of small quantities of cancerous cells or pathogens in patient tissue samples. Such methods are also useful for culturing cells for subsequent use in cell or tissue engineering. | 1. A method of cellular analysis, the method comprising: introducing a heterologous cell into a fish;contacting the fish with an agent; andanalyzing a property of the cell to determine whether the property is responsive to the administration of the agent to the fish, wherein a difference in the property in the agent-treated fish compared to the property in an untreated or control fish is indicative of the property being responsive to the administration of the agent to the fish. 2. The method of claim 1, wherein the fish is a teleost. 3. The method of claim 2, wherein the teleost is a zebrafish, medaka, Giant rerio, or puffer fish. 4. The method of claim 1, wherein the fish is an embryo. 5. The method of claim 1, wherein the fish is a larva or an adult. 6. The method of claim 1, wherein the fish is a wild-type strain. 7. The method of claim 1, wherein the fish is a mutant strain. 8. The method of claim 1, wherein a plurality of cells are introduced into the fish. 9. The method of claim 1, wherein a plurality of cells are introduced at multiple sites of the fish. 10. The method of claim 1, wherein the heterologous cell is undifferentiated. 11. The method of claim 1, wherein the heterologous cell is terminally differentiated. 12. The method of claim 1, wherein the agent is from a combinatorial library of peptides or small molecules. 13. The method of claim 1, wherein the agent is a naturally occurring molecule. 14. The method of claim 1, wherein the agent is a hormone, growth factor, or cytokine. 15. The method of claim 1, wherein agent is a polypeptide, beta-turn mimetic, polysaccharide, phospholipid, prostaglandin, steroid, aromatic compound, heterocyclic compound, benzodiazepine, oligomeric N-substituted glycine, or oligocarbamate. 16. The method of claim 1, wherein the agent is administered by adding the agent to media containing the fish. 17. The method of claim 1, wherein the agent is administered by microinjecting the agent into the fish. 18. The method of claim 1, wherein the heterologous cell is a mammalian cell, bacterial cell, plant cell, or bird cell. 19. The method of claim 18, wherein the cell is a mammalian cell. 20. The method of claim 19, wherein the mammalian cell is from a cat, dog, horse, bovine, mouse, rat, rabbit, guinea pig, or primate. 21. The method of claim 20, wherein the mammalian cell is a human cell. 22. The method of claim 1, wherein the heterologous cell is a cancer cell. 23. The method of claim 22, wherein the property is a differentiation marker, movement of the cancer cell relative to an initial site of introduction, death of the cancer cell, or proliferation of the cancer cell. 24. The method of claim 22, which comprises labeling the cancer cell. 25. The method of claim 24, wherein the labeling comprises immunostaining with an antibody specific to a cancer cell antigen. 26. The method of claim 24, wherein the cancer cell is labeled after introducing the cancer cell into the fish. 27. The ... | Phylonix Pharmaceuticals, Inc. | A01K067/027 A61K049/00 A61B005/00 A61B008/00 A61B010/00 C12Q001/00 G01N033/53 G01N033/574 G01N033/00
US09026207 | B2 | 2015-05-05 | Medical device for Atrial fibrillation prediction | 2008-01-28 | In a device and a method for providing correlated measures for predicting potential occurrence of atrial fibrillation, an impedance of the patient is measured to obtain impedance information; cardiogenic data is determined from the information; respiratory data is determined from the information; at least one hemodynamic measure is calculated from the cardiogenic data and at least one apnea measure is calculated from the respiratory data; the hemodynamic and apnea measures are correlated such that the correlated measures can be utilized for predicting potential occurrence of atrial fibrillation. | 1. An implantable medical device including a pulse generator adapted to produce cardiac stimulating pacing pulses, said device being connectable to at least one medical lead comprising electrodes for delivering said pulses to cardiac tissue of a heart of a patient, comprising: an impedance measuring unit, connectable to at least two electrodes of said at least one medical lead, configured to measure an impedance of said patient and to provide impedance information corresponding to said measured impedance;a cardiogenic impedance determining unit configured to receive said impedance information and to determine cardiogenic impedance data from said impedance information;a respiration impedance determining unit configured to receive said impedance information and to determine respiratory impedance data from said respiratory impedance information;a calculation unit configured to determine at least one hemodynamic measure using said cardiogenic impedance data and at least one apnea measure using said respiratory impedance data; anda correlating unit configured to calculate an atrial fibrillation prediction value as a function of a correlation between said hemodynamic and apnea measures, the correlation constituting a relationship, over a time period, between a development of said hemodynamic and apnea measures. 2. The device according to claim 1, wherein said impedance measuring unit measures a cardiogenic impedance substantially corresponding to an impedance of an atrium of said patient. 3. The device according to claim 2, wherein said correlating unit is configured to correlate said hemodynamic and apnea measures such that an increasing apnea activity of said patient coinciding with a decreasing elasticity of an atrial myocardium of said patient can be detected from said correlation. 4. The device according to claim 1, wherein said calculation unit is configured to determine said at least one hemodynamic measure as a time derivative of a cardiogenic impedance curve formed by means of said cardiogenic impedance data, said time derivative corresponding to a change of the cardiogenic impedance during a cardiac cycle. 5. The device according to claim 4, further comprising an ECG-measuring unit, connectable to at least one electrode, configured to receive an ECG signal from said heart,determine an ECG curve, anddetect a QRS complex of said ECG curve, and wherein said calculation unit is adapted to obtain information related to said QRS complex and said ECG curve from said ECG-measuring unit,synchronize said cardiogenic impedance curve with said detected QRS complex, andcalculate said time derivative in a time window following said QRS complex. 6. The device according to claim 1, wherein said calculation unit is configured to compare said hemodynamic measure with a hemodynamic reference value. 7. The device according to claim 1, wherein said apnea measure is a number of apneic events and said calculation unit is configured to detect a first breath fr... | St Jude Medical AB | A61N001/365 A61B005/02 A61B005/0205 A61B005/0452 A61B005/053 A61B005/08 A61B005/00 A61N001/362
US08715200 | B2 | 2014-05-06 | System for determining the blood flow in a coronary artery | 2010-07-20 | The invention relates to a method and a system for determining the blood flow in an individual coronary artery of a patient, wherein the method comprises the steps of positioning a temperature sensor mounted at a distal portion of a guide wire at a distal position in the coronary artery, positioning an infusion catheter in the coronary artery such that the distal end of the infusion catheter is proximally of the temperature sensor, measuring the blood temperature with the temperature sensor, infusing cold indicator fluid with a known infusion rate and known or measurable temperature into the coronary artery by the infusion catheter, measuring the temperature of the mixture of blood and indicator fluid by the temperature sensor, and calculating the coronary blood flow by a formula based on the known and measured quantities. In an extended version, the method comprises steps for relating the calculated coronary flow value to related normal flow values, or related FFR values, or a related flow resistance. | 1. A system for determining a blood flow in an individual coronary artery of a patient, comprising: a sensor guide wire having a distal portion provided with a guide wire sensor and being configured for positioning in the individual coronary artery;an infusion catheter configured for positioning in the individual coronary artery and configured for injecting an indicator fluid into the individual coronary artery;a monitoring device for transforming an output temperature signal from the guide wire sensor into a temperature of medium surrounding the sensor; anda device configured to be connected to the infusion catheter for delivering a cold indicator fluid from a reservoir into the individual coronary artery,wherein the monitoring device is configured to measure a mixture temperature of a mixture of blood and indicator fluid using the guide wire sensor,wherein the monitoring device is configured to calculate the blood flow by a formula,wherein the formula used for calculating the blood flow is:Qb=Qk⁢Tk-TbT-Tbwhere Qb is the blood flow, Qk is a known infusion rate of the indicator fluid, Tk is a known or measurable temperature of the indicator fluid, Tb is a measured blood temperature, and T is the measured mixture temperature. 2. The system according to claim 1, further comprising a pressure sensor being configured for positioning in the individual coronary artery. 3. The system according to claim 2, wherein the pressure sensor and the guide wire sensor are provided on the same guide wire. 4. The system according to claim 1, wherein the guide wire sensor is configured to provide the output temperature signal and an output pressure signal, and wherein the monitoring device is configured to transform the output pressure signal from the guide wire sensor into a pressure of the medium surrounding the guide wire sensor. 5. The system according to claim 1, wherein the monitoring device is configured to measure the indicator fluid temperature at a distal end of the infusion catheter. 6. The system according to claim 1, wherein the guide wire sensor is the only temperature sensor of the system. 7. The system according to claim 1, wherein the device for delivering the cold indicator fluid is configured to continuously deliver the cold indicator fluid from the reservoir into the individual coronary artery. 8. A system for determining a blood flow in an individual coronary artery of a patient, comprising: a sensor guide wire having a distal portion provided with a guide wire sensor and being configured for positioning in the individual coronary artery;an infusion catheter configured for positioning in the individual coronary artery and configured for injecting an indicator fluid into the individual coronary artery;a monitoring device for transforming an output temperature signal from the guide wire sensor into a temperature of medium surrounding the sensor; anda device configured to be connected to the infusion catheter for delivering a cold indicator ... | Radi Medical Systems AB | A61B005/02
US08652057 | B2 | 2014-02-18 | Recyclable or biodegradable blood pressure cuff | 2010-02-12 | A blood pressure cuff is defined by a first sheet, having an opening, and a second sheet. The second sheet is attached to a bottom of the first sheet to form an interior inflatable portion between the first sheet and the second sheet, with the opening of the first sheet fluidly interconnecting the interior inflatable portion with an exterior of the cuff. The blood pressure cuff is made from a single type of material and can be recycled, biodegradable, or composted. | 1. A blood pressure cuff, the cuff comprising: a first sheet having an opening;a second sheet, wherein the second sheet is attached to a bottom of the first sheet;an interior-inflatable portion formed between the bottom of the first sheet and a top of the second sheet, wherein the opening of the first sheet fluidly interconnects the interior-inflatable portion with an exterior of the cuff, the first and second sheets being made from a fluid impermeable material;a socket attached to said cuff through said opening and having a circumferential lip that directly and releasably engages one of a hose adapter and a gage housing, said socket fluidly interconnecting the exterior of said cuff with said interior-inflatable portion;a slotted portion formed in a longitudinally extending section of said first sheet spaced apart from the interior-inflatable portion; andfastener material disposed on the cuff proximate an end of the longitudinally extending section wherein the cuff is configured to wrap around a limb of a patient during blood pressure measurement such that the fastener material adheres to a top surface of the first sheet and the slotted portion receives the socket and wherein the length of slotted portion and positioning of the socket on the cuff are selected to limit the circumference of the cuff in the configuration during blood pressure measurement when the socket is received in the portion. 2. A cuff as recited in claim 1, wherein the cuff further comprises means for indicating that the cuff has already been used. 3. A cuff as recited in claim 2, wherein means for indicating that the cuff has already been used includes alteration of a color of a treated portion of the cuff after exposure to light. 4. A cuff as recited in claim 1, wherein the fastener material is a hook, made from the fluid impermeable material, attached to a bottom of the first sheet and spaced apart from the second sheet. 5. A cuff as recited in claim 4, wherein the hook is attached to an end portion of the first sheet and on a side of the first sheet including the slotted portion. 6. A cuff as recited in claim 4, wherein a top of the first sheet further comprises a loop fastener configured to engage the hook. 7. A cuff as recited in claim 6, wherein the top of the first sheet comprises a non-woven fibrous material. 8. A cuff as recited in claim 6, wherein the non-woven fibrous material has a weight of about 0.5 to 3.0 ounces per square yard. 9. A cuff as recited in claim 6, wherein a film is laminated to the bottom of the first sheet. 10. A cuff as recited in claim 9, wherein the film has a thickness of about 0.002 to 0.008 inches and a weight of about 1.0 to 7.0 ounces per square yard. 11. A cuff as recited in claim 9, wherein the first sheet has a creep elongation of less than about 0.0065 inches. 12. A cuff as recited in claim 6, wherein the bottom of the first sheet has a smooth texture. 13. A cuff as recited in claim 6, wherein the loop fasteners are degrad... | Welch Allyn, Inc. | A61B005/02
US08103338 | B2 | 2012-01-24 | Impedance based anatomy generation | 2009-05-08 | Methods and systems for the determination and representation of anatomical information are disclosed herein. In some examples, methods described herein can include determining a boundary of at least a portion of the heart based on measured electrical signals, determining an area of the boundary corresponding to an endocardial boundary, and displaying a portion of the boundary of the heart with the display including a visual indicia indicative of the determined area of the boundary. | 1. A method comprising: inserting a catheter into a heart, the catheter comprising three or more electrodes;causing current to flow between at least some of the electrodes and in response to current flow, measuring an electrical signal at each of one or more of the electrodes;for each of a plurality of catheter locations, determining a local surface of at least a portion of the heart based on the measured electrical signals and determining valid and invalid areas of the local surface, the valid areas corresponding to an endocardial boundary;determining a representation of a chamber anatomy of at least a portion of the heart based on multiple ones of the valid areas of the local surfaces, the determined representation of a chamber anatomy of at least a portion of the heart representing an area larger than any one of the local surfaces; anddisplaying at least a portion of the determined representation of the chamber anatomy of the at least a portion of the heart with a first visual indicia indicative of the determined valid area corresponding to the endocardial boundary. 2. The method of claim 1, further comprising: measuring cardiac signals at the catheter electrodes in response to electrical activity in the heart; anddetermining physiological information at multiple locations of the surface of the heart based the measured cardiac signals. 3. The method of claim 2, further comprising displaying the physiological information at multiple locations of the surface of the heart. 4. The method of claim 2, further comprising displaying the physiological information at multiple locations of the surface of the heart for only the determined valid areas corresponding to the endocardial boundary. 5. The method of claim 1, wherein displaying the surface representation of the chamber anatomy comprises displaying the surface for a portion of less than all of the heart. 6. The method of claim 1, wherein the determination of the local surfaces accounts for a change in conductivity at the cardiac chamber boundary. 7. The method of claim 6, wherein the determination accounts for a first conductivity inside the cardiac chamber boundary and, a second conductivity outside the cardiac chamber boundary. 8. The method of claim 1, wherein determining the local surfaces comprises determining anatomical information based at least in part on impedance information generated based on the measured signals. 9. The method of claim 8, wherein the impedance information is based on a conductivity contrast between blood and surrounding tissue. 10. The method of claim 8, wherein the impedance information is based on a permittivity contrast between blood and surrounding tissue. 11. The method of claim 8, wherein the impedance information comprises complex impedance information based on both a permittivity contrast and a conductivity contrast between blood and surrounding tissue. 12. The method of claim 1, wherein determining the local surfaces comprises, determining the local... | Rhythmia Medical, Inc. | A61B005/05
US08718742 | B2 | 2014-05-06 | Integrated wireless patch for physiological monitoring | 2008-05-23 | A sensor system in accordance with the present invention comprises a plane member, a plurality of electrodes within the plane member, adapted to contact a human body to detect and monitor human generated voltages. The sensor can be applied to monitor a variety of applications relating to health, disease progression, fitness and wellness. Some of the specific applications include the monitoring of ECG, EEG, EMG, glucose, electrolytes, body hydration, dehydration, tissue state and wounds. Various aspects of the invention are shown by illustrating certain embodiments. Many other embodiments can be used to implement the invented schemes. | 1. A sensor system comprising: a plane member;a plurality of electrodes within a first side of the plane member and adapted to contact a human body to detect and monitor human generated voltages; anda circuit on a second side of the plane member for providing signals over a wireless link, wherein the circuit comprises: a plurality of amplifiers for simultaneously amplifying signals received from the electrodes, wherein each amplifier of said plurality detects differential signals from different pairs of electrodes of said plurality; a plurality of filters for the signals; an analog to digital converter (ADC) for converting the signals to digital signals; a digital signal processor for processing the digital signals; anda radio for wirelessly transmitting the digital signals,wherein the plane member is folded over so the first side overlies the second side. 2. The system of claim 1 wherein the amplifiers detect differential signals from opposite pairs of electrodes. 3. The system of claim 1 which includes a multiplexer for receiving the signals from the plurality of filters and providing the signals to the analog to digital converter. 4. The system of claim 1 wherein a portion of the sensor system is disposable. 5. The system of claim 1 wherein a multiplexer is inserted between the plurality of electrodes and plurality of amplifiers. 6. The system of claim 1, wherein the plurality of electrodes are ECG electrodes. 7. The system of claim 6, wherein the amplifiers simultaneously detect differential signals from at least three pairs of electrodes. 8. The system of claim 6, wherein a stationary patient monitor or mobile device are configured to capture transmitted signals from said radio. | HMicro, Inc. | A61B005/0408 A61B005/0402
US08369939 | B2 | 2013-02-05 | Activation apparatus, method, and computer program for brainwave interface system | 2008-10-21 | An electroencephalogram IF system includes an electroencephalogram measurement section for measuring an electroencephalogram signal, a function control section for analyzing an event-related potential contained in the electroencephalogram signal and outputting a function control signal for controlling a function of a device based on the result of analysis, and an output section for outputting the function control signal. The activation apparatus includes: an activation determination section for, while the electroencephalogram IF system is not functioning, transmitting to the output section a stimulation control signal for controlling presentation and vanishing of a visual stimulation on a single-item on the output section and, within the electroencephalogram signal acquired from the electroencephalogram measurement section, allowing a P200 component value of an event-related potential since the timing of presenting the visual stimulation as a starting point to be compared against a predetermined threshold value, and determining whether or not to output an activation trigger to the function control section based on the result of comparison; and a stimulation attention determination section for determining whether or not the user is paying attention to the visual stimulation based on an N100 component of the event-related potential since the timing of presenting the visual stimulation as a starting point, and causing processing by the activation determination section to begin depending on the determination result. When the activation determination section outputs an activation trigger, the electroencephalogram IF system is activated. | 1. An electroencephalogram interface system comprising: an electroencephalogram measurement section for measuring an electroencephalogram signal from a user,a function control section for analyzing an event-related potential contained in the electroencephalogram signal and outputting a function control signal for controlling a function of a device based on a result of analysis,an output section for outputting the function control signal, andan activation apparatus for activating the electroencephalogram interface system, the activation apparatus comprising:an activation determination section for, while the function control section of the electroencephalogram interface system is inactive, transmitting to the output section a stimulation control signal for controlling presentation and vanishing of a visual stimulation of a single-item on the output section, and, within the electroencephalogram signal acquired from the electroencephalogram measurement section, comparing a P200 component value of an event-related potential since a timing of presenting the visual stimulation as a starting point against a predetermined threshold value, and determining whether or not to output an activation trigger to the function control section based on a result of comparison; anda stimulation attention determination section for determining whether or not the user is paying attention to the visual stimulation based on an N100 component of the event-related potential since the timing of presenting the visual stimulation as a starting point, and causing processing by the activation determination section to begin depending on a determination result, wherein,the activation apparatus activates the electroencephalogram interface system by outputting the activation trigger to the function control section. 2. The electroencephalogram interface system of claim 1, wherein, as the P200 component value, the activation determination section compares a value of the event-related potential in a zone of 200±50 ms since the timing of presenting the visual stimulation as a starting point against the predetermined threshold value. 3. The electroencephalogram interface system of claim 1, wherein, as the P200 component value, the activation determination section compares a local maximum value, a maximum value, or a zone average value of the event-related potential in a zone of 200±50 ms since the timing of presenting the visual stimulation as a starting point against the predetermined threshold value. 4. The electroencephalogram interface system of claim 1, wherein the stimulation attention determination section instructs the activation determination section to change a method of presenting the visual stimulation based on a determination result that the user is paying attention to the visual stimulation. 5. The electroencephalogram interface system of claim 1, further comprising: an imaging device for imaging a video and outputting a video signal; and whereinthe activation apparatus f... | Panasonic Corporation | A61B005/04
US20100331710 | A1 | 2010-12-30 | BLOOD PRESSURE MEASUREMENT CUFF THAT INCLUDES AN ANTIMICROBIAL SUBSTANCE | 2010-03-10 | A blood pressure measurement cuff comprising an inflatable bladder, a source of fluidic pressure in fluidic communication with the inflatable bladder, a manometer in fluidic communication with the inflatable bladder, a cover at least partially covering the inflatable bladder and comprising an antimicrobial substance, and means to place the inflatable bladder sufficiently close to a limb of a living being to allow the measurement of the blood pressure of the living being. The source of fluidic pressure can be a pump that transforms human force into fluid pressure, such as a squeezable ball pump, or a canister of compressed fluid. The manometer can include a dial display or a digital display. The cuff can further comprise a tube providing fluidic communication between the manometer and the inflatable bladder, and the tube can be 1 foot or longer. The cover can be made of a fabric coated with an antimicrobial substance. | 1. A blood pressure measurement cuff comprising: an inflatable bladder;a source of fluidic pressure in fluidic communication with the inflatable bladder;a manometer in fluidic communication with the inflatable bladder;a cover at least partially covering the inflatable bladder;the cover comprising an antimicrobial substance; andmeans to place the inflatable bladder sufficiently close to a limb of a living being to allow the measurement of the blood pressure of the living being. 2. The blood pressure measurement cuff of claim 1, the source of fluidic pressure comprising a pump that transforms human force into fluid pressure. 3. The blood pressure measurement cuff of claim 1, the source of fluidic comprising a canister of compressed fluid. 4. The blood pressure measurement cuff of claim 1, the manometer comprising a dial display. 5. The blood pressure measurement cuff of claim 1, the manometer comprising a digital display. 6. The blood pressure measurement cuff of claim 1 further comprising a tube providing fluidic communication between the manometer and the inflatable bladder, the tube comprising a first length, the first length being greater than 1 foot. 7. The blood pressure measurement cuff of claim 1, the cover comprising a fabric coated with an antimicrobial substance. 8. The blood pressure measurement cuff of claim 7, the antimicrobial substance comprising a silver-ion emitter. 9. The blood pressure measurement cuff of claim 7, the antimicrobial substance comprising any one of 3 trimethoxysilylpropyloctadecyldimethyl ammonium chloride, hyaluronan and its derivatives, triclosan, or a copolymer of chloropropyltrihydroxysilane and octadecylaminodimethyltrihydroxysilylpropyl ammonium chloride. 10. The blood pressure measurement cuff of claim 7, the antimicrobial substance capable of emitting ions that aid in the destruction of a microbe. 11. The blood pressure measurement cuff of claim 1, the cover comprising a monometer receiver and the monometer comprising means to releasably attach the monometer to the monometer receiver. 12. The blood pressure measurement cuff of claim 11, the receiver comprising a loop and the means to releasably attach the monometer to the receiver comprising a spring tension clip. 13. The blood pressure measurement cuff of claim 11, the receiver comprising loops and the means to releasably attach the monometer to the receiver comprising hooks. 14. The blood pressure measurement cuff of claim 1, the cover further comprising a second receiver and an attachment zone, andthe means to place the inflatable bladder sufficiently close to a limb of a living being to allow the measurement of the blood pressure of the living being comprising the second receiver and the attachment zone, wherein second receiver is adapted to be removably attached to the attachment zone. 15. The blood pressure measurement cuff of claim 14, the second receiver comprising loops and the attachment zone comprising hooks. 16. The blood pressur... | - | A61B005/0225 A01N025/34 A01N059/16 A01N043/16 A01N031/14 A01N055/10 A01P001/00
US20100305480 | A1 | 2010-12-02 | Human Motion Classification At Cycle Basis Of Repetitive Joint Movement | 2009-06-01 | Methods and systems for classifying human motion as corresponding to an activity are disclosed. One example method includes sensing motion characteristics associated with the activity to generate a first set of data, identifying a cycle interval in the first set of data; and identifying the activity based on the interval. | 1. A method for classifying motion as corresponding to an activity type, the method comprising: sensing motion characteristics associated with an activity using one or more motion sensors to generate a first set of data;identifying a cycle interval in the first set of data; andidentifying the activity based on the interval. 2. The method as recited in claim 1, wherein the sensed motion characteristics include human limb motion characteristics. 3. The method as recited in claim 2, wherein the human limb is an ankle. 4. The method as recited in claim 1, wherein the one or more motion sensors include a gyroscopic sensor. 5. The method as recited in claim 1, wherein the activity is identified as one of a set of activities comprising: running, walking, rowing, cycling, and elliptical walking. 6. The method as recited in claim 1, further comprising: sensing motion characteristics associated with the activity using the one or more motion sensors to generate a second set of data,wherein the activity is identified based on the second set of data. 7. The method as recited in claim 6, wherein the motion characteristics sensed to generate the second set of data include at least one of trunk and ankle motion characteristics. 8. The method as recited in claim 6, wherein the one or more motion sensors includes an accelerometer, the second set of data being generated using the accelerometer. 9. The method as recited in claim 1, wherein identifying the activity based on the interval includes: calculating a feature of motion data generated by a motion sensor during the interval; andusing the feature to identify the activity. 10. The method as recited in claim 9, wherein the first set of data includes the motion data generated during the interval, and wherein the motion data includes angular velocity data and the calculated feature is an absolute magnitude of the angular velocity data. 11. The method as recited in claim 9, wherein the interval is an interval between consecutive occurrences of a forward swing event performed by a body part, and wherein identifying the activity is based in part on a duration time of the forward swing event. 12. The method as recited in claim 9, wherein the first set of data includes angular velocity data characterizing angular motion of a first ankle, andwherein the feature is an angular velocity feature calculated using at least a portion of the angular velocity data generated during the cycle duration. 13. The method as recited in claim 12, wherein the activity is identified based on whether the angular velocity feature exceeds a first threshold. 14. The method as recited in claim 12, further comprising: sensing vertical acceleration characteristics of at least one of the first and a second ankle using the one or more motion sensors to generate vertical ankle acceleration data; andcalculating a vertical ankle acceleration feature using at least a portion of the vertical ankle acceleration data generated during the inte... | - | A61B005/11
US07843351 | B2 | 2010-11-30 | Back training device | 2004-09-01 | A novel back training device is disclosed. The back training device includes a control box; a detector mounted to the control box, for detecting when a person bends at the waist; and an alert means, operatively connected to the detector for alerting the person when the person bends at the waist. | 1. A back training device, comprising: a portable control box;a detector mounted to the portable control box for detecting when a person bends at the waist beyond a user selectable set point and for measuring a number of degrees associated with how far the person bends at the waist, wherein the detector is configured to be worn by the person such that the detector is positioned upon the person's body to measure bending at the waist;a set button mounted on the portable control box and operatively connected to the detector for setting the user selectable set point, wherein the person sets a user selectable set point while wearing the detector by bending over to a maximum desired angle and by pressing the set button while being bent over to the maximum desired angle, such that the user selectable set point is defined as the maximum desired bending angle of the person;an alert means operatively connected to the detector and mounted to the portable control box for alerting the person each time the person bends at the waist beyond an alert point, wherein the alert point is the same as the user selectable set point, such that the alert means alerts the user each time the user bends beyond the user selectable set point, thereby providing negative feedback to the person each time the person bends at the waist beyond the user selectable set point, thereby training the person to refrain from bending beyond the user selectable set point, wherein the alert means is activated by the detector when it is detected that the person has bent at the waist beyond the user selectable set point;a memory configured to store data corresponding to when the person bends at the waist beyond the user selectable set point, wherein the memory is configured to function as a logic/counter storage device, wherein data is stored each time the person bends at the waist beyond the user selectable set point, the stored data comprising a date, a time, and the number of degrees associated with how far the person bends at the waist;a display operatively connected to the detector configured to display data stored by the memory; anda processor operatively connected to the memory and the display to facilitate the display of data stored by the memory. 2. The device of claim 1, wherein the detector is housed within the portable control box. 3. The device of claim 1, wherein the alert means includes a power source, wherein the power source is electrically connected to the alert means for powering the alert means. 4. The device of claim 3, wherein the alert means is a vibrator. 5. The device of claim 3, wherein the alert means is a buzzer. 6. The device of claim 3, wherein the power source includes at least one battery. 7. The device of claim 3, further comprising a display button for browsing the display output, the display being connected to the power source. 8. The device of claim 1, wherein the display includes a liquid crystal display readout. 9. The device of claim 8, wherein the ... | - | G08B023/00 A61B005/103 A61B005/00
US07846094 | B2 | 2010-12-07 | System and method for neurological injury detection, classification and subsequent injury amelioration | 2005-03-17 | A system (10) and method for detecting and analyzing neurological damage in an injured subject are provided using electrical nerve conduction analyses supported by a computing device (12). An appropriate neurologically protective pharmaceutical for that injury type and circumstances based on the result of said analyses is suggested. | 1. A system for detecting a neurological injury in a subject prior to transport from a site of the injury, said system comprising: a computing device comprising: at least one distal signal emitter attachable to a first position on the subject to emit an electrical signal generated by the computing device into the subject such that the electrical signal is communicated to a nerve in proximity to the first position;at least one signal detector attachable to a second position in electrical communication with a subject central nervous system on the subject to detect the electrical signal transmitted by the nerve as neural conductivity as a detected electrical signal; anda processor for comparing a threshold reference value with the detected electrical signal and indicating the neurological injury when the detected electrical signal is beyond a preselected range of the reference value; anda biochemical analyzer for analyzing fluid samples for the presence of chemical species or concentrations indicative of the neurological injury providing an output signal to said computing device; anda display providing indication of the neurological injury associated with the output signal and the detected electrical signal;said system providing suggestions as to at least one of: immediate interventive neuroprotective pharmaceutical treatments, or physical transport precautions prior to transport from a site of the injury. 2. The system of claim 1 wherein a database is comprised of signal strengths for various positions and muscle groups of the subject. 3. The system of claim 1 wherein the computing device provides a user with instruction for positioning the at least one emitter and the at least one detector on the subject. 4. The system of claim 1 further comprising an ancillary monitoring device providing the computing device with an input relating to a physiological parameter of the subject. 5. The system of claim 1 further comprising a user interface for data input to the computing device. 6. The system of claim 4 further comprising a wireless transmitter coupled to the computing device communicating the indication of neurological injury and input. 7. The system of claim 1 wherein the computing device provides suggested neuroprotective pharmaceutical treatment protocols for the subject. 8. The system of claim 1 in combination with a kit of neurologically active neuroprotective pharmaceuticals and at least one device for introducing a pharmaceutical into the subject. 9. A process for detecting a neurological injury in a subject comprising: attaching a distal emitter at a first position and a detector at a second position to the subject in electrical communication with a subject central nervous system prior to transport of the subject to a trauma center;emitting an electrical signal from a computing device into the subject at the first position via the emitter;detecting the electrical signal transmitted by a nerve at the second position with the detector a... | - | A61B005/00
US07840044 | B2 | 2010-11-23 | Method and system for liver lobe segmentation and pre-operative surgical planning | 2006-11-01 | A method and system is provided for image segmentation for liver objects. Segmentation is performed to obtain a first set of objects relating to liver. More than one types in association with one of the first set of objects are identified. Landmarks are identified based on the segmented first set of objects or the different types of one of the first set of objects. A second set of objects including liver lobes are segmented based on the landmarks. | 1. A computer-implemented method, comprising: performing 3D segmentation to obtain a first set of objects including a 3D liver and one or more 3D vessels;identifying at least one type of vessels in the first set of objects, including a portal vein and a hepatic vein;identifying at least one 3D anatomic landmark point and/or structure based on the 3D liver and/or the at least one type of vessels in accordance with a 3D reference map; andobtaining, via 3D segmentation, a second set of objects which includes one or more 3D liver lobes based on the identified at least one 3D anatomic landmarks point and/or structure. 2. The method of claim 1, further comprising providing a virtual surgical planning setting based on at least one of: the first set of objects;the more than one type of one of the first set of objects;the landmarks; andthe second set of objects. 3. The method of claim 2, wherein in the virtual surgical setting, one or more virtual operations are to be performed. 4. The method of claim 3, wherein an object can be visualized in the virtual surgical setting, wherein the visualized object includes at least one of: the first set of objects; andthe second set of objects. 5. The method of claim 4, wherein the object to be visualized is presented in a manner that preserve spatial relationships with other objects with respect to a certain perspective. 6. The method of claim 4, wherein the object to be visualized can be rotated and translated in a 3D rendering space. 7. The method of claim 4, wherein the object to be visualized is presented with a label. 8. The method of claim 3, wherein the one or more virtual operations can be carried out with respect to a manipulable object. 9. The method of claim 8, wherein the manipulable object includes at least one of: a surface area close to the liver;the liver;a lobe;a vessel; anda nodule. 10. The method of claim 8, wherein the one or more virtual operations include at least one of: a surgical cut;a removal of at least a portion of the manipulable object; andmanipulation of the manipulable object. 11. The method of claim 8, wherein, in carrying out a virtual operation, at least one reaction is generated. 12. The method of claim 11, wherein the at least one reaction includes: a measure made with respect to the virtual operation;an event detected that occurs as a result of the virtual operation;a response generated to indicate a resolution to the event detected; anda recording produced with respect to the virtual operation. 13. The method of claim 12, wherein the measure includes at least one of: a position measure indicating a position that is being affected by the virtual operation;a size measure indicating a size of an area that is being affected by the virtual operation;a distance measure between the position of the area and another object; anda safe margin determined based on the distance measure. 14. The method of claim 12, wherein the event includes at least one of: an opening occurring i... | Edda Technology Inc. | G06K009/00 A61B005/05
US20100317953 | A1 | 2010-12-16 | Medical Devices and Methods of Using the Same | 2010-08-23 | Medical devices having restrictive access, and methods thereof are provided. | 1.-29. (canceled) 30. A health monitoring system, comprising: a data management system, having application software; anda plurality of handheld monitoring devices in bi-directional communication with the data management system,wherein the application software on the data management system interfaces with each of the plurality of handheld monitoring devices to thereby configure operating parameters of the handheld monitoring device and receive data from the handheld monitoring device. 31. The health monitoring system of claim 30, wherein at least one of the plurality of handheld monitoring devices is in wireless communication with the data management system. 32. The health monitoring system of claim 30, wherein the data management system receives instructional commands from a health care provider, and wherein application software on the data management system transmits the instructional commands to at least one of the plurality of handheld monitoring devices. 33. The health monitoring system of claim 32, wherein the transmission is via wireless communication. 34. The health monitoring system of claim 30, wherein at least one of the plurality of handheld monitoring devices is in wired communication with the data management system. 35. The health monitoring system of claim 30, wherein the data management system receives instructional commands from a health care provider, and wherein application software on the data management system transmits the instructional commands to at least one of the plurality of handheld monitoring devices. 36. The health monitoring system of claim 30, wherein the data received from the handheld monitoring device is indicative of an analyte concentration in a blood sample obtained from a user of the handheld monitoring device. 37. The health monitoring system of claim 30, wherein at least one of the plurality of handheld monitoring devices is a blood glucose meter. 38. The health monitoring system of claim 30, wherein at least one of the plurality of handheld monitoring devices is a continuous glucose monitor. 39. The health monitoring system of claim 30, wherein at least one of the plurality of handheld monitoring devices an audible indicator. 40. The health monitoring system of claim 30, wherein at least one of the plurality of handheld monitoring devices is wired to a personal computer. 41. The health monitoring system of claim 30, wherein at least one of the plurality of handheld monitoring devices connects to a cradle, wherein the cradle is in data communication with a personal computer. 42. A health monitoring system, comprising: a data management system, having application software; anda plurality of handheld monitoring devices, wherein each of the plurality of handheld monitoring devices provides for both wireless and wired bi-directional communication with the data management system,wherein the application software on the data management system interfaces with each of the plurality of handheld monitori... | - | A61B005/145
US20100305416 | A1 | 2010-12-02 | NON-INVASIVE PRESSURED PROBING DEVICE | 2007-10-23 | A non-invasive probing device wearable over a skin of a patient for monitoring the patient is provided. The probing device includes a physiological probe having a sensing interface, and a housing for housing the probe. The housing includes a body having a skin-side surface for lying next to the skin of the patient and a fixed or resilient protruding member fixedly or resiliently protruding from the skin-side surface. The protruding member houses the sensing interface. The device also includes attachment means for attaching the device to the skin of the patient. In the case of a resilient protruding member, the device may include a resilient mechanism for biasing the resilient protruding member against the body. Also in the case of a resilient protruding member, the device may include a protective cover for protecting the device against unwanted external influences. | 1. A non-invasive probing device wearable over a skin region of a patient for physiological monitoring thereof, said probing device comprising: a physiological probe having a sensing interface;a housing for housing said probe, said housing comprising a body having a skin-side surface for lying adjacent to the skin region of the patient, said housing further comprising a fixed protruding member protruding from said skin-side surface away from said body and toward the skin region of the patient, said protruding member housing said sensing interface; andattachment means for attaching said housing to the skin region of the patient and holding said skin-side surface and protruding member in pressure contact over the skin region. 2. A non-invasive probing device according to claim 1, wherein said sensing interface comprises at least one output of a light guide. 3. A non-invasive probing device according to claim 1, wherein said sensing interface comprises at least one input of a light guide. 4. A non-invasive probing device according to claim 1, wherein said attachment means comprise an adhesive layer extending over said skin-side surface of said body. 5. A non-invasive probing device according to claim 4, wherein said adhesive layer surrounds the protruding member and has a thickness less than a length of the protruding member. 6. A non-invasive probing device according to claim 1, wherein said attachment means comprise a flexible band or strap for encircling a body part of said patient and holding said housing of said probe in pressure contact over the skin region of the patient. 7. A non-invasive probing device according to claim 6, wherein said strap or said band is elastic. 8. A non-invasive probing device according to claim 1, wherein said attachment means comprise medical-grade adhesive tape for taping over said housing and securing said housing to the skin region of the patient, said medical-grade adhesive tape comprising adhesive strips for adhering to the housing and the skin region of the patient and thereby securing the housing to the skin region of the patient. 9. A non-invasive probing device according to claim 1, wherein said attachment means comprise a medical-grade adhesive patch comprising a non-adhesive portion for covering the housing and an adhesive portion for adhering to the skin region of the patient, thereby securing said housing to the skin region of the patient. 10. A non-invasive optical probing device according to claim 1, further comprising a rigid protective cover enclosing said housing, said cover having an outlet therein for the protruding member. 11. A non-invasive probing device wearable over a skin region of a patient for physiological monitoring thereof, said probing device comprising: a physiological probe having a sensing interface;a housing for housing said probe, said housing comprising a body having a skin-side surface for lying adjacent to the skin region of the patient, said housing further comprisi... | CYBIOCARE, INC. | A61B005/145
US20100286543 | A1 | 2010-11-11 | Automated Cardiac Status Determination System | 2010-03-31 | A system fits a curve to a filtered ECG signal, processes the fitted curve (e.g., by applying transforms such as a KL Transform) and derives parameters for use in classifying heart cycle signal portions (such as an ST segment portion) into particular heart cycle signal portion categories associated with particular segment morphology. A system for heart signal classification includes an interface for receiving an electrical signal waveform comprising an R-wave and including an ST segment portion associated with heart electrical activity of a patient over a heart beat cycle. A signal processor processes data representing the electrical signal waveform by fitting a curve to the ST segment and applying a transform to the fitted curve to derive variance data indicating variance in the fitted curve. A signal classifier classifies the ST segment into one of multiple predetermined categories associated with the fitted ST segment curve geometry in response to the derived variance data. | 1. A system for heart signal classification, comprising: an interface for receiving an electrical signal waveform comprising an R-wave and including an ST segment portion associated with heart electrical activity of a patient over a heart beat cycle;a signal processor for processing data representing said electrical signal waveform by (a) fitting a curve to data representing said ST segment and(b) applying a transform to the fitted curve to derive variance data indicating variance in the fitted curve; anda signal classifier for classifying the ST segment into one of a plurality of predetermined categories associated with the fitted ST segment curve geometry in response to the derived variance data. 2. A system according to claim 1, wherein said signal classifier classifies the ST segment into one of a plurality of predetermined categories associated with characteristics including (a) Concave Elevation and (b) Convex Elevation. 3. A system according to claim 2, wherein said signal classifier classifies the ST segment into one of a plurality of predetermined categories associated with characteristics including (a) Upsloping Depression and (b) Horizontal Depression. 4. A system according to claim 3, wherein said signal classifier classifies the ST segment into one of a plurality of predetermined categories associated with characteristics including Downsloping Depression. 5. A system according to claim 1, wherein said transform comprises a KLT transform or another variance analysis transform. 6. A system according to claim 1, wherein said transform performs Principal Component Analysis (PCA) to transform the data to a new coordinate system such that the greatest variance lies on a first coordinate called the first principal component. 7. A system according to claim 1, wherein said signal processor adaptively fits a first degree curve or a second degree curve selected in response to a determined ST deviation value. 8. A system according to claim 1, wherein said signal processor adaptively fits a curve or a line to an ST segment, selected in response to a determined ST deviation value indicating a positive or negative ST segment slope. 9. A system according to claim 1, wherein said signal classifier classifies the ST segment into one of said plurality of predetermined categories using mapping data associating predetermined ranges of variance data values with corresponding categories of ST segment. 10. A system according to claim 9, wherein said mapping data associates predetermined ranges of variance data values for populations of particular demographic characteristics including at least one of, age, weight, height and gender with corresponding categories of ST segment. 11. A system according to claim 1, wherein said signal processor processes data representing said electrical signal waveform by (a) identifying a J point in said electrical signal waveform,(b) identifying a Ton point in said electrical signal waveform substantially occurring 8... | SIEMENS MEDICAL SOLUTIONS USA, INC. | A61B005/0452
US20100298684 | A1 | 2010-11-25 | ANALYTE SENSOR | 2010-02-26 | Systems and methods of use for continuous analyte measurement of a host's vascular system are provided. In some embodiments, a continuous glucose measurement system includes a vascular access device, a sensor and sensor electronics, the system being configured for insertion into communication with a host's circulatory system. | 1. A system for measuring an analyte, the system comprising: a vascular access device configured to be in communication with a circulatory system of a host;an analyte sensor configured to reside within the vascular access device, wherein the analyte sensor is configured to measure a concentration of an analyte within the circulatory system; anda flow control device configured to draw blood into the vascular access device at a first flow rate of from about 0.001 ml/min to about 0.2 ml/min, wherein the flow control device is further configured to pass an infusion solution through the vascular access device at a second flow rate, wherein the first flow rate is not equal to the second flow rate. 2. The system of claim 1, wherein the analyte is glucose. 3. The system of claim 1, wherein the vascular access device comprises a catheter, and wherein the sensor is located within the catheter. 4. The system of claim 3, wherein the catheter is 21 gauge or smaller. 5. The system of claim 1, wherein the vascular access device comprises a catheter and a fluid coupler connectable to the catheter. 6. The system of claim 1, wherein the vascular access device comprises at least one of a venous catheter or an arterial catheter. 7. The system of claim 1, wherein the vascular access device comprises a Swann Gantz catheter. 8. The system of claim 1, wherein the vascular access device comprises a multi-lumen catheter, and wherein the sensor is located in one of the catheter's lumens. 9. The system of claim 1, wherein the first flow rate is configured to substantially block mixing of the infusion solution and the sample. 10. The system of claim 1, wherein the second flow rate is from about 0.001 ml/min to about 2 ml/min. 11. The system of claim 10, wherein the second flow rate is from about 0.02 ml/min to about 0.35 ml/min. 12. The system of claim 10, wherein the second flow rate is configured to allow a temperature of the infusion solution to substantially equilibrate with a temperature of the host. 13. A system for measuring an analyte, the system comprising: a vascular access device configured to be in communication with a circulatory system of a host;an analyte sensor configured to reside within the vascular access device, wherein the analyte sensor is configured to measure a concentration of an analyte within the circulatory system; anda valve comprising a first position and a second position that is different from the first position, wherein the system is configured such that about 150 microliters or less of a fluid passes through the vascular access device during movement of the valve at a predetermined rate between the first position and the second position during a predetermined amount of time. 14. The system of claim 12, wherein the system is configured to push fluid through the vascular access device during movement of the valve from the first position to the second position. 15. The system of claim 12, wherein the valve is configured such that ... | DexCom, Inc. | A61B005/145
US07835782 | B2 | 2010-11-16 | Integrated PET-MRI scanner | 2006-04-28 | One embodiment of the present invention provides an integrated PET-MRI scanner. This integrated scanner includes a main magnet that generates a magnetic field, wherein images of the subject is generated in a central region of the magnetic field. It also includes a PET scanner which is enclosed by the main magnet. The PET scanner further comprises: (1) at least one ring of scintillators, which is situated in the central region of the magnetic field and, (2) one or more photodetectors, which are coupled to the ring of scintillators, so that the one or more photodetectors are outside the central region of the magnetic field. The integrated scanner also includes radio-frequency (RF) coils which are enclosed by the PET scanner. By keeping the photodetectors and associated circuitry outside the central region of the magnetic field, the integrated scanner reduces the electromagnetic interference (EMI) between the PET scanner and the MRI scanner. | 1. An integrated positron emission tomography (PET)-magnetic resonance imaging (MRI) scanner, comprising: a main magnet that generates a magnetic field during an MRI process for a subject, wherein the image is generated in a central region of the magnetic field;a PET scanner that detects annihilation photons produced by positron decays within a subject during a PET imaging process, wherein the PET scanner is enclosed by the main magnet, wherein the PET scanner further comprises:at least one ring of scintillators that detect positron annihilation photons and output light photons in response to the detected positron annihilation photons, wherein the ring of scintillators is situated in the central region of the magnetic field; andone or more photodetectors that convert the light photons output by the ring of scintillators into electrical signals, wherein the photodetectors are coupled to the ring of scintillators, such that the photodetectors are located outside of the central region of the magnetic field, and wherein the photodetectors are enclosed by the main magnet, such that length of an optical fiber coupling the photodetectors and the ring of scintillators is below a predetermined value; anda set of radio frequency (RF) coils that transmit signals to and receive signals from the subject during the MRI process, wherein the set of RF coils are enclosed by the PET scanner; anda shield configured to shield the photodetectors and associated circuitry of the PET scanner from the RF signals generated by the RF coils of the MRI;wherein keeping the photodetectors and associated circuitry outside of the central region of the magnetic field reduces the electromagnetic interference (EMI) between the PET scanner and the MRI scanner, and wherein keeping the length of the optical fiber below the predetermined value reduces light transmission loss within the optical fiber. 2. The integrated PET-MRI scanner of claim 1, wherein the photodetectors are coupled to the ring of scintillators through optical fibers, wherein one end of the optical fibers is attached to the outputs of the scintillators, while the other end of the optical fibers is attached to the inputs of the photodetectors, wherein the optical fibers transfer the photons outputted by the ring of scintillators to the photodetectors, wherein using the optical fibers facilitates keeping the photodetectors outside of the central region of the magnetic field. 3. The integrated PET-MRI scanner of claim 2, wherein the optical fibers are a fraction of the size of the main magnet, which limits light transmission loss. 4. The integrated PET-MRI scanner of claim 2, wherein the optical fibers include at least one cladding layer which reduces light transmission loss. 5. The integrated PET-MRI scanner of claim 2, wherein the optical fibers include more than one cladding layer which reduces light transmission loss. 6. The integrated PET-MRI scanner of claim 2, wherein the spaces between the optical fibers are... | The Regents of the University of California | A61B005/05
US07860554 | B2 | 2010-12-28 | Visible-near infrared spectroscopy in burn injury assessment | 2001-01-26 | A non-invasive method of characterizing burn injuries using near infrared spectroscopy is described. In the method, a beam of light is emitted into the burnt tissue portion at two or more different tissue depths. The spectra are then compared using multivariate analysis to determine diagnostic regions of the spectra. This information is used to categorize the burn. In some cases, the diagnostic regions correspond to wavelengths related to the hemodynamics of the tissue portion. The spectra can also be repeated over time, thereby allowing trends and changes in the spectra to be measured. This data is in turn used to categorize the burn as either a superficial burn, partial thickness burn, deep partial burn or a full thickness burn. Once the burn has been categorized, the clinician can intervene as needed to treat the burn. | 1. A method of characterizing a burn comprising: emitting a beam of light at a wavelength between 500 to 1100 nm from a source into a burnt tissue portion;collecting and analyzing reflected light from the beam with a detector at a first separation distance from the source, thereby producing a first depth spectrum;emitting a beam of light at a wavelength between 500 to 1100 nm from a source into a burnt tissue portion;collecting and analyzing reflected light from the beam with a detector at a second separation distance from the source, thereby producing a second depth spectrum;investigating depth dependent circulatory alterations by comparing oxyhemoglobin and deoxyhemoglobin balance within the burnt tissue portion from the first depth spectrum and the second depth spectrum; andcategorizing the burn based on said comparison. 2. The method according to claim 1, wherein the first depth spectrum and the second depth spectrum are compared over spectral regions corresponding to at least one of the following: oxygen saturation, total hemoglobin, hydration or oxyhemoglobin levels. 3. A method of characterizing a burn comprising: at a first time point:(a) emitting a beam of light at a wavelength between 500 to 1100 nm from a source into a burnt tissue portion;(b) collecting and analyzing reflected light from the beam with a detector at a first separation distance from the source, thereby producing an early first depth spectrum;(c) emitting a beam of light at a wavelength between 500 to 1100 nm from the source into a burnt tissue portion;(d) collecting and analyzing reflected light from the beam with a detector at a second separation distance from the source, thereby producing an early second depth spectrum;at a second time point:(e) emitting a beam of light at a wavelength between 500 to 1100 nm from the source into a burnt tissue portion;(f) collecting and analyzing reflected light from the beam with a detector at the first separation distance from the source, thereby producing a late first depth spectrum;(g) emitting a beam of light at a wavelength between 500 to 1100 nm from the source into a burnt tissue portion;(h) collecting and analyzing reflected light from the beam with a detector at the second separation distance from the source, thereby producing a late second depth spectrum;(i) investigating depth dependent circulatory alterations over time by comparing oxyhemoglobin and deoxyhemoglobin balance within the burnt tissue potion from the early first depth spectrum to oxyhemoglobin and deoxyhemoglobin balance from the early second depth spectrum and comparing oxyhemoglobin and deoxyhemoglobin balance within the burnt tissue potion from the late first depth spectrum to oxyhemoglobin and deoxyhemoglobin balance from the late second depth spectrum; and(i) determining if burn depth has become worse over time based on said investigation. 4. The method according to claim 3 wherein the first depth spectrum and the second depth spectrum are compared ove... | National Research Council of Canada | A61B005/00
US07846172 | B2 | 2010-12-07 | Device, system and method for interstitial transvascular intervention | 2007-10-26 | Method and apparatus for utilizing the vascular system as a conduit to reach other vascular and extravascular locations within the body. Included are methods for revascularization wherein the extravascular passageways are formed to permit blood flow between vascular locations. Also included are methods for performing transvascular interstitial surgery (TVIS) wherein extravascular passageways are formed from a blood vessel to another vascular or non-vascular intracorporeal location. Also disclosed are devices usable for forming extravascular passageways in accordance with the invention, or for modifying, valving, maintaining or closing such passageways. | 1. A method of creating an arterio-venous fistula in a patient between a starting vessel and a target vessel, comprising: (A) placing a catheter in the starting vessel;(B) advancing a needle from the catheter into the target vessel, thereby creating a fistulous tract between the starting vessel and the target vessel;(C) passing a guidewire through the crossing needle into said target vessel;(D) advancing an anastomotic clip delivery device over said guidewire; and(E) deploying an anastomotic clip device from the delivery device to create an anastomosis between the starting vessel and the target vessel. 2. A method according to claim 1 wherein the starting vessel is an artery and the target vessel is a vein. 3. A method according to claim 1 wherein, after performance of Step C and before performance of Step D, the method further comprises: enlarging the tract between the starting vessel and the target vessel. 4. A method according to claim 3 wherein the enlarging step comprises: advancing a the tract enlarging device over the guidewire; andusing the tract enlarging device to enlarge the tract between the starting vessel and the target vessel. 5. A method according to claim 4 wherein the needle is retracted into the catheter and the catheter is removed prior to advancement of the tract enlarging device over the guidewire. 6. A method according to claim 4 wherein the tract enlarging device comprises a balloon and wherein the using step comprises inflating the balloon to dilate the tract. 7. A method according to claim 4 wherein the tract enlarging device comprises a debulking device and wherein the using step comprises using the debulking device to enlarge the tract. 8. A method according to claim 7 wherein the debulking device emits radiofrequency energy which ablates tissue. 9. A method according to claim 7 wherein the debulking device comprises a cutting element which cuts tissue. 10. A method according to claim 9 wherein the cutting element comprises a pull back cutting element. 11. A method according to claim 1 wherein the anastomotic clip device comprises a stent which stents the fistulous tract. 12. A method according to claim 1 wherein the anastomotic clip device includes a valve for valving flow through the fistulous tract. 13. A method according to claim 1 wherein the catheter further comprises orientation apparatus and wherein the method further comprises the step of using the orientation apparatus to guide adjustment of the rotational orientation of the catheter within the starting vessel prior to advancement of the needle. 14. A method according to claim 13 wherein the orientation apparatus comprises an active orientation apparatus. 15. A method according to claim 13 wherein the orientation apparatus comprises a passive orientation apparatus. | Medtronic Vascular, Inc. | A61B017/22 A61B005/05
US07825667 | B2 | 2010-11-02 | Microwave imaging system and processes, and associated software products | 2005-12-21 | A microwave imaging process, and a system controlled by an associated software product, illuminate a target with microwaves from a transmitting antenna. Receiving antennas receive microwaves scattered by the target, and form microwave data. The illumination and receiving repeat over multiple transmitting antennas and multiple microwave frequencies. The microwave data is processed to form permittivity and conductivity images by selecting a background dispersion model for permittivity and conductivity. Permittivity and conductivity dispersion coefficients are determined, and permittivity and conductivity distributions are calculated, for each of the microwave frequencies. Forward solutions at multiple frequencies are determined from property distributions, and a dispersion coefficient based Jacobian matrix is determined. Dispersion coefficient updates are determined using the microwave data, and the dispersion coefficients are updated. Permittivity and conductivity distributions are recalculated for each of the frequencies, and the forward solutions are determined at multiple frequencies from property distributions. | 1. A microwave imaging process, comprising: illuminating a target with microwaves from a transmitting antenna,receiving microwaves scattered by the target at a plurality of receiving antennas to form microwave data,repeating the illuminating and receiving steps for a plurality of transmitting antennas,repeating the illuminating, receiving, and repeating steps for a plurality of microwave frequencies; andprocessing the microwave data to form permittivity and conductivity images by: (a) selecting a frequency relationship for each of permittivity and conductivity,(b) determining permittivity and conductivity dispersion coefficients for the permittivity and conductivity frequency relationships,(c) calculating permittivity and conductivity distributions for each frequency of the plurality of frequencies,(d) determining forward solutions at multiple frequencies from initial estimates of permittivity and conductivity distributions,(e) determining a dispersion coefficient based Jacobian matrix,(f) determining dispersion coefficient updates using the microwave data,(g) updating the dispersion coefficients,(h) recalculating current permittivity and conductivity distributions for each frequency of the plurality of frequencies, and(i) determining the forward solutions at multiple frequencies from the current permittivity and conductivity distributions. 2. The process of claim 1, further comprising: (j) calculating a multi-frequency difference vector between measured and computed electric field values,(k) computing a tolerance metric, andrepeating steps (e), (f), (g), (h), (i), (j) and (k) until the tolerance metric meets a pre-defined limit. 3. The process of claim 1, wherein selecting the frequency relationship for each of permittivity and conductivity comprises selecting a permittivity relationship equation from the group consisting ofɛr=A⁢⁢ω+B,⁢ln⁡(ɛr)=A⁢⁢ω+B,⁢ɛr=A⁢⁢ln⁡(ω)+B,⁢ln⁡(ɛr)=A⁢⁢ln⁡(ω)+B⁢⁢andɛr=∑i=0N⁢Ai⁢ωi;where εr denotes relative permittivity, ω denotes angular frequency, i is a summation index, N is a polynomial order, and A and B are the permittivity dispersion coefficients. 4. The process of claim 1, further comprising utilizing the permittivity dispersion coefficients to form an image. 5. The process of claim 1, wherein selecting the frequency relationship for each of permittivity and conductivity comprises selecting a conductivity relationship equation from the group consisting ofσ=A⁢⁢ω+B,⁢ln⁡(σ)=A⁢⁢ω+B,⁢σ=A⁢⁢ln⁡(ω)+B,⁢ln⁡(σ)=A⁢⁢ln⁡(ω)+B⁢⁢andσ=∑i=0N⁢Ai⁢ωi;where σ denotes conductivity, ω denotes angular frequency, i is a summation index, N is a polynomial order, and A and B are the conductivity dispersion coefficients. 6. The process of claim 1, further comprising utilizing the conductivity dispersion coefficients to form an image. 7. The process of claim 1, wherein the step of determining the forward solutions at multiple frequencies comprises utilizing electric field values from a previous step of recalculating the forward solutions ... | Microwave Imaging Systems Technologies, Inc. | G01N023/02 A61B005/05
US07824369 | B2 | 2010-11-02 | Balloon catheter having osmotic inflation system | 2005-04-21 | A system and method for inflating and deflating a catheter balloon using osmosis. An elongate catheter shaft has a proximally disposed osmotic membrane in fluid communication with a distally mounted balloon through a fluid-filled catheter lumen. The balloon is inflated by exposing the exterior of the membrane to a solution that is hypotonic relative to the solution within the catheter lumen. The balloon is deflated by exposing the exterior of the membrane to a solution that is hypertonic in relation to the solution within the catheter lumen. | 1. A system comprising: an elongate hollow shaft having a lumen filled with a first fluid solution, the shaft having an outer diameter sized to act as a guidewire;a balloon disposed adjacent a distal end of the shaft and in fluid communication with the lumen;an osmotic membrane disposed adjacent a proximal end of the shaft and in fluid communication with the lumen, wherein the proximal end of the shaft having the osmotic membrane is sized for insertion into a guidewire lumen of a treatment catheter; anda first vial having a first vial opening configured to removably receive the osmotic membrane through the first vial opening into the first vial, the first vial containing a second fluid solution different from the first fluid solution. 2. The system of claim 1 wherein the second solution is hypotonic relative to the first solution. 3. The system of claim 2 further comprising: a second vial different from the first vial containing a third fluid solution different from the second fluid solution, the second vial having a second vial opening configured to removably receive the osmotic membrane through the second vial opening into the second vial. 4. The system of claim 3 wherein the third solution is hypertonic relative to the second solution. 5. The system of claim 3 wherein the first solution, the second solution and the third solution each contain a solute chosen from the group consisting of glucose, lactose, sucrose, sodium chloride, potassium chloride, glycine, albumin, insulin, and combinations thereof. 6. The system of claim 3 wherein the first solution comprises a 0.5 M glucose solution, the second solution comprises a 0.1 M glucose solution and the third solution comprises a 1.0 M glucose solution. 7. The system of claim 3 wherein the first solution comprises a 1 gram NaCl/liter solution, the second solution comprises a 0.1 gram NaCl/liter solution and the third solution comprises a 2 gram NaCl/liter solution. 8. The system of claim 1 wherein the osmotic membrane is a semipermeable membrane. 9. The system of claim 1 wherein the osmotic membrane is a rigid membrane. 10. The system of claim 1 wherein the osmotic membrane is a flexible membrane. 11. The system of claim 1 further comprising a membrane support fixedly attached adjacent the proximal end of the shaft. 12. The system of claim 11 wherein the membrane support comprises an outer membrane support and an inner membrane support, the osmotic membrane being disposed there between. 13. The system of claim 11 wherein the membrane support comprises a framework having a plurality of openings. 14. The system of claim 1 wherein the elongate hollow shaft is a guidewire. 15. A method of using a balloon catheter, the method comprising: providing the balloon catheter comprising:an elongate hollow shaft having a catheter lumen, the shaft sized for insertion into a guide wire lumen of a treatment catheter;a catheter balloon disposed adjacent a distal end of the shaft, the balloon in fluid ... | Medtronic Vascular, Inc. | A61M029/00 A61M025/00 A61B005/00
US07826890 | B1 | 2010-11-02 | Optical detection of intravenous infiltration | 2005-12-06 | An intravenous infiltration detection apparatus for monitoring intravenous failures, which applies an optical method coupled with fiber optics and algorithms for tissue optics to provide a means for noninvasive detection of intravenous infiltration surround the site of IV injection. In the invention, the tissue surrounding the injection site is exposed to a single-wavelength of electromagnetic radiation, and light is collected with only one detector. Changes in the relative intensity of the radiation reflected, scattered, diffused or otherwise emitted provide a means for monitoring infiltration. The invention provides routine, automated, continuous, and real-time monitoring for patients undergoing IV therapy. | 1. Intravenous infiltration detection apparatus comprising: an electronic unita power supply connected to the electronic unit,a continuous single wavelength single light source connected to the power supply for delivering non-encoded light in a certain spectral range,a first disposable light guide comprising an optical fiber having proximal and distal ends, the proximal end being optically connected to the light source for conducting light from the light source, the distal end being configured for mounting near an IV infusion site of a patient,a second disposable light guide comprising an optical fiber having proximal and distal ends, the distal end being configured for mounting near the infusion site for receiving the electromagnetic radiation reflected, scattered, diffused or otherwise emitted from the patient's tissue near an intravenous needle injection site,a detector in the electronics unit at the proximal end of the second light guide for detecting light from the guide,an analyzer in the electronics unit connected to the detector for analyzing the received light, wherein analyzing further comprises comparing intensity of the received light,an indicator connected to the analyzer in the electronics unit for providing information on the occurrence of infiltrations, anda disposable skin contact sensor fixed to and holding the distal ends of the first and second light guides against the patient's tissue. 2. The apparatus of claim 1, wherein the electronics unit is configured to compensate both a baseline light signal value and the light signal during infusion with an ambient light signal value. 3. The apparatus of claim 1, wherein the electronics unit is configured to determine infiltration based on a ratio of a baseline light signal value compensated with an ambient light signal value and the light signal during infusion compensated with the ambient light signal value. 4. The apparatus of claim 1, wherein the distal end of the first light guide is modified for placement on the tissue of a patient. 5. The apparatus of claim 4, wherein the distal end of the second light guide is modified for placement on the tissue of a patient. 6. The apparatus of claim 5, wherein the distal ends of the first and second light guides comprise a skin contact sensor and an adhesive base for attaching the contact sensor to the skin. 7. The apparatus of claim 5, wherein the first and second light guides comprise optical fibers, and wherein the first light guide comprises multiple light-illuminating optical fibers having proximal ends connected to the non-encoded light source of a certain spectral range. 8. The apparatus of claim 5, wherein the first and second light guides comprise optical fibers, and wherein the second light guide comprises multiple light-receiving optical fibers having proximal ends severally connected to a photon detection device. 9. The apparatus of claim 5, wherein the first and second light guides comprise optical fibers, and wherein b... | WinTec, LLC | A61B006/00 A61B005/05
US07824340 | B2 | 2010-11-02 | Pulsimeter sensor using magnetic thin films | 2006-10-19 | The present invention relates to a noninvasive medical pulsimeter sensor using magnetic thin films. By forming a pulse-sensing part array with magnetic sensors such as GMR devices, MTJ devices and the likes, over the skin-contacting part which consists of a magnetic material, the present invention increases the integrity of sensors, minimizes the time for searching the pulse and it is applicable widely to portable pulsimeters and the likes. | 1. A pulsimeter sensor using magnetic thin films, including: a skin-contacting part formed with a magnetic material to be contacted to a skin to examine the pulse;a pulse-sensing part located some distance away from the skin-contacting part and formed as an array type of unit cells with two or more magnetic films; anda spatial part located between the skin-contacting part and the pulse-sensing part;wherein the unit cell of the pulse-sensing part array is a GMR (Giant Magneto Resistance) device; andwherein the pulse-sensing part array is formed by dividing thirty unit cells into three groups corresponding to “Chon”, “Gwan” and “Cheok”. 2. The pulsimeter sensor of claim 1, wherein the magnetic material of the skin-contacting part is made of nano-magnetic particles of magnetic nano-bead type. 3. The pulsimeter sensor of claim 2, wherein the nano-magnetic particles are one selected from Co, Fe3O4 and Fe2O3. 4. The pulsimeter sensor of claim 1, wherein the magnetic material of the skin-contacting part is ultra thin. 5. The pulsimeter sensor of claim 4, wherein the ultra thin magnetic material of the skin-contacting part is a ribbon-type magnetic pad. 6. The pulsimeter sensor of claim 5, wherein the ribbon-type magnetic pad of the skin-contacting part is made of plastic magnets, of which the strength of the magnetic field is 200˜300 Oe at a 3 mm distance. 7. The pulsimeter sensor of claim 1, wherein the GMR device is comprised of a lower layer, a middle layer and an upper layer,wherein the lower layer, facing the skin-contacting part, is a free (variable ferromagnetic) layer,wherein the upper layer is a pinned (fixed ferromagnetic) layer, andwherein the middle layer is a Cu layer as a nonmagnetic layer. 8. The pulsimeter sensor of claim 7, wherein the free layer and the pinned layer are comprised of two or more thin films, which are different from each other, respectively. 9. The pulsimeter sensor of claim 8, wherein the free layer is formed of a CoFe layer and a NiFe layer orderly from one side of the nonmagnetic layer, andthe pinned layer is formed of a CoFe layer and a IrMn layer orderly from the other side of the nonmagnetic layer. 10. The pulsimeter sensor of claim 9, wherein the Cu layer is 2.5˜2.7 nm thick, the CoFe layer of the free layer is 1.8˜2.2 nm thick, the NiFe layer is 3.8˜4.2 nm thick, the CoFe layer of the pinned layer is 3.3˜3.7 nm thick, and the IrMn layer is 13.0˜17.0 nm thick. 11. The pulsimeter sensor of claim 1, wherein the unit cell of the pulse-sensing part is about 1.0 mm×2.0 mm, and the array is arranged as a package of the unit cells with 2×5 matrix structure at each “Chon”, “Gwan” and “Cheok”. 12. The pulsimeter sensor of claim 1, wherein the array of the pulse-sensing part is made by fabricating switching devices and M×N matrix unit cells on one wafer, cutting the M×N matrix unit cells into specific size and forming them as one package. 13. The pulsimeter sensor of claim 1, wherein the spatial part is a consta... | Sangji University Industry Academy Cooperation Foundation | A61B005/02
US07848937 | B2 | 2010-12-07 | Connecting consumers with service providers | 2007-06-15 | A set of answers is received to a series of questions sent to a consumer of services, the set of answers to the series of questions is analyzed in a computer to identify types of service providers that can provide services to the consumer based on the received answers, and an agenda is produced, the agenda including the identified types of service providers and topics for the consumer of services to discuss with a service provider of each type, with the topics determined at least in part based on the answers. | 1. A computer implemented method comprises: receiving a set of answers to a series of questions regarding a health assessment sent to a consumer of services;analyzing in a computer the set of answers to the series of questions by: accessing a database storing process logic and rules data; andapplying the process logic and rules data to the received set of answers from the consumer to determine in the computer one or more medical issues for the consumer;generating in the computer, based at least in part on analysis of the set of answers to the series of questions: a list of types of service providers that can provide services to the consumer, anda list of topics relating to the determined medical issues of the consumer of services;generating a graphical user interface that when rendered on a display device, renders a set of agenda item indicators that were determined based on the medical issues and a description of selected topics from the list of topics, with one or more of the topics in the list of topics juxtaposed to one or more of the agenda item indicators in the set of agenda item indicators. 2. The method of claim 1 further comprising: generating questions tailored to the consumer based at least in part on the received set of answers;receiving answers to the questions tailored to the consumer; andwherein generating, further comprises:generating, based at least in part on analysis of the set of answers to the series of questions and the answers received from the questions tailored to the consumer. 3. The method of claim 1 wherein generating the graphical user interface when rendered on a display device further renders: one or more of the medical issues for the consumer and an agenda associated with the set of agenda item indicators. 4. The method of claim 1 further comprising: receiving a request to transmit an agenda associated with the set of agenda item indicators to a second service provider; andtransmitting the agenda to the second, requested service provider. 5. The method of claim 1 further comprising: generating a list of the determined medical issues;receiving a request to transmit the list of the determined medical issues to a second service provider; andtransmitting the list of the determined medical issues to the second, requested service provider. 6. The method of claim 5 further comprising: receiving a request from a service provider to generate the second series of questions. 7. The method of claim 1 further comprising: applying the process logic and rules data to the received set of answers from the consumer;determining whether additional information is needed to assess the medical issues of the consumer of services; andif additional information is needed, generating a second series of questions;receiving a second set of answers to the second series of questions; andapplying the process logic and rules data to the received second set of answers from the consumer to determine the medical issues of the consumer of servic... | American Well Corporation | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07826884 | B2 | 2010-11-02 | Live fluoroscopic roadmapping including targeted automatic pixel shift for misregistration correction | 2006-10-30 | An X-ray diagnostic imaging system for conducting live fluoroscopic subtraction imaging is described as including an X-ray source for directing X-ray radiation to a patient being examined, an x-ray imaging device positioned for receiving the X-ray radiation and acquiring images in response thereto and a processor arranged in communication with the x-ray source and x-ray imaging device to control acquisition of a contrast-enhanced mask image frame and a live image frame without contrast enhancement, to conduct a pixel shift calculation operation based on a small, user-defined region of interest (ROI), for example, 1/16 of a full frame, to realize a pixel shift vector to correct for motion between live image frames, to shift pixels comprising the mask image frame by pixel shift directions defined by the pixel shift vector, and to subtract the shifted mask image frame from the live image frame to realize a live roadmapping image frame. The system includes a display for displaying the live roadmapping image frame, and a user interface that allows a user to define and capture the small ROI in the displayed image frame for use by the processor conducting the pixel shift calculation. | 1. An X-ray diagnostic imaging system for conducting live fluoroscopic subtraction imaging, comprising: an X-ray source for directing X-ray radiation to a patient being examined;an X-ray imaging device positioned for receiving the X-ray radiation and acquiring images in response thereto;a processor configured to communicate with the X-ray source and X-ray imaging device to control acquisition of a contrast-enhanced mask image frame and a live image frame without contrast enhancement, to conduct a pixel shift vector calculation operation based on a small, user-defined region of interest (ROI) having a location selectable by a user within a displayed image frame during live X-ray imaging and comprising a plurality of adjacent pixels to realize a pixel shift vector derived from the plurality of pixels to correct for motion between an individual live image frame of successive live image frames and a mask image frame, to shift pixels comprising the mask image frame by pixel shift directions defined by the pixel shift vector, and to subtract the shifted mask image frame from the live image frame to realize a live roadmapping image frame;a display for displaying live roadmapping image frames; anda user interface that allows a user to define and capture the small ROI in the displayed image frame for use by the processor conducting the pixel shift calculation. 2. The X-ray diagnostic imaging system as set forth in claim 1, wherein the live roadmapping image is scored for registration quality and said processor calculates a plurality of pixel shift vectors for a corresponding respective plurality of successive live image frames and shifts a mask image frame by said plurality of pixel shift vectors to provide a plurality of different mask image frames corresponding to said plurality of successive live image frames and subtracts the plurality of different mask image frames from the corresponding plurality of successive live image frames. 3. The X-ray diagnostic imaging system as set forth in claim 2, wherein said processor, automatically generates a test pixel shift vector,dynamically modifies the test pixel shift vector in response to the user-defined small ROI, andperforms the subtraction operation over the entire frame using the updated pixel shift vector. 4. The X-ray diagnostic imaging system as set forth in claim 1, wherein the processor is configured to subtract and score all pixels in the user-defined small ROI for registration quality determination. 5. The X-ray diagnostic imaging system as set forth in claim 1, wherein the processor is configured to determine a shift for fractional pixels in the user-captured ROI input when the pixel shift vector is calculated. 6. The X-ray diagnostic imaging system as set forth in claim 1, wherein the size of the small ROI may be modified by user input, and wherein the processor is configured to calculate the pixel shift vector in a time that is proportional to the small ROI dimensions. 7. The X-ray diagnost... | Siemens Medical Solutions USA, Inc. | A61B005/05 A61B006/00 G06K009/00 G01N023/04
US07828741 | B2 | 2010-11-09 | Utilizing lipopolysaccharide in exhaled breath condensate to diagnose gram negative pneumonia | 2005-05-23 | A method for determining whether a subject has Gram negative bacterial pneumonia based on the presence of lipopolysaccharide in exhaled breath condensate collected from the subject. The collection devices utilized to collect exhaled breath condensate from both spontaneously breathing and mechanically ventilated subjects and the devices utilized to determine whether lipopolysaccharide is present in the collected exhaled breath condensate. | 1. A method for diagnosing and monitoring intrapulmonary Gram negative bacterial infection in an air-breathing vertebrate subject, comprising: (a) collecting exhaled breath condensate from an air-breathing vertebrate subject;(b) providing a reaction chamber with an interior surface, wherein said reaction chamber has a material that binds lipopolysaccharide disposed therein;(c) delivering at least a portion of the collected breath condensate to the reaction chamber;(d) washing the reaction chamber with a buffer;(e) delivering a reaction reagent to the reaction chamber;(f) determining whether the subject has an intrapulmonary Gram negative bacterial infection based on a physical change that occurs when the at least a portion of the collected exhaled breath condensate is delivered to the reaction chamber, wherein said physical change is caused by the presence of lipopolysaccharide in the exhaled breath condensate, and(g) identifying the particular strain of the intrapulmonary Gram negative bacteria;wherein identifying includes identifying the particular strain of intrapulmonary Gram negative bacteria based upon a determination of the O-antigen portion of the lipopolysaccharide molecule, wherein the binding material is embedded in a matrix surface disposed within the reaction chamber. 2. A method for diagnosing and monitoring intrapulmonary Gram negative bacterial infection in an air-breathing vertebrate subject, comprising: collecting exhaled breath condensate from the air-breathing vertebrate subject;providing a reaction chamber, wherein said reaction chamber has a reaction reagent disposed therein;delivering at least a portion of the collected exhaled breath condensate to the reaction chamber;determining whether the subject has an intrapulmonary Gram negative bacterial infection based on a physical change that occurs when at least a portion of the collected exhaled breath condensate is delivered to the reaction chamber, wherein said physical change is caused by the presence of lipopolysaccharide in the exhaled breath condensate; andidentifying the particular strain of the intrapulmonary Gram negative bacteria;wherein identifying includes identifying the particular strain of intrapulmonary Gram negative bacteria based upon a determination of the O-antigen portion of the lipopolysaccharide molecule. 3. A method for diagnosing and monitoring intrapulmonary Gram negative bacterial infection in an air-breathing vertebrate subject, comprising: collecting exhaled breath condensate from an air-breathing vertebrate subject;providing a reaction chamber, wherein said reaction chamber has a reaction reagent disposed therein;delivering at least a portion of the collected exhaled breath condensate to the reaction chamber;determining whether the subject has an intrapulmonary Gram negative bacterial infection based on a physical change that occurs when the at least a portion of the collected exhaled breath condensate is delivered to the reaction chamber, wherei... | The Charlotte-Mecklenburg Hospital Authority | A61B005/08
US07834624 | B2 | 2010-11-16 | Method and device for detection of the position of an examination person in a magnetic resonance system | 2008-03-07 | In a method and arrangement for detection of the position of an examination person on a table in a magnetic resonance system, the examination person on the table is moved relative to the magnetic resonance system, RF pulses are radiated while the examination person is moved through the magnetic resonance system, the resulting magnetic resonance signals caused by the RF pulses are detected and the position of the examination person is determined using the acquired magnetic resonance signals. | 1. A method for detecting a position of an examination person in a magnetic resonance scanner, comprising the steps of: placing an examination person on a moveable table and operating the table with computerized system to move the examination person on the table at least partially through the magnetic resonance scanner;irradiating the examination person on the table with at least one RF pulse while the examination person and the table are moving through the magnetic resonance scanner, and thereby causing magnetic resonance signals to be emitted from the examination person that are insufficient by themselves for image generation;optimizing said at least one RF pulse to give a magnetic resonance signal originating from predetermined body tissue, at a known location within the examination person, a differing signal characteristic that differs from magnetic resonance signals originating from a remainder of the examination person;detecting said magnetic resonance signals correlated in said computerized system with moving the table through the magnetic resonance scanner; andautomatically electronically determining a position of the examination person in the magnetic resonance scanner using only the acquired magnetic resonance signal, which has said differing signal characteristic, from the predetermined body tissue correlated with moving the table through said magnetic resonance scanner. 2. A method as claimed in claim 1 wherein said table reaches at least one table position while moving through the magnetic resonance scanner, and comprising: radiating a number of RF pulses in a range between 1 and 4 pulses with said table in said at least one table position, andgenerating a readout gradient and reading out a magnetic resonance signal, as said at least one magnetic resonance signal, generated by said number of RF pulses during said readout gradient. 3. A method as claimed in claim 1 comprising also automatically electronically determining an orientation of the examination person on the table using the detected magnetic resonance signals. 4. A method as claimed in claim 1 comprising optimizing the magnetic resonance signal emitted from the examination person to cause the magnetic resonance signal originating from predetermined body tissue of the examination person to be greater than the magnetic resonance signal originating from a remainder of the body of the examination person. 5. A method as claimed in claim 4 comprising optimizing said magnetic resonance signal to cause the magnetic resonance signal originating from brain tissue, as said predetermined body tissue, to be greater than the magnetic resonance signal that originates from the remainder of the body of the examination person. 6. A method as claimed in claim 5 comprising: automatically electronically determining that said magnetic resonance signal originates from the remainder of the body of the examination person when said magnetic resonance signal is above a first signal level value, a... | Siemens Aktiengesellschaft | G01V003/00 A61B005/05
US07835778 | B2 | 2010-11-16 | Method and apparatus for surgical navigation of a multiple piece construct for implantation | 2003-10-16 | A method and apparatus for percutaneous and/or minimally invasive implantation of a construct. The construct may be implanted using a navigation system for planning and execution of a procedure. A plurality of portions of the construct may be interconnected using locations and paths determined and navigated with the navigation system. | 1. A system for use in navigating an implantation of a human spinal implant, comprising: a first pedicle screw member and a second elongated rod member of the construct adapted to selectively interact with each other after implantation and said first pedicle screw member operable to be fixed into a boney structure;a third pedicle screw member positionable relative to said first pedicle screw member and said second elongated rod member in an anatomical portion;a first localization element selectively and removably fixed to said first pedicle screw member having a depressible member to operate an internal mechanism to engage and disengage from said first pedicle screw member;a second localization element selectively and removably fixed to said second elongated rod member;a detector operable to detect said first localization element and said second localization element when said first localization element and said second localization element is selectively fixed to at least one of said first pedicle screw member or said second elongated rod member; anda processor configured to: assist in navigation of said second elongated rod member relative to said first pedicle screw member;receive position information for all of said first pedicle screw member, said second elongated rod member, and said third pedicle screw member from said detector;determine a relative position between all of said first pedicle screw member, said second elongated rod member, and said third pedicle screw member;assist in navigation of said third pedicle screw member relative to at least one of said first pedicle screw member or said second elongated rod member including generating icons for display on a display device without patient image data displayed on the display device; andsimultaneously track said first localization element while selectively fixed to said first pedicle screw member and said second localization element while selectively fixed to said second elongated rod member. 2. The system of claim 1, wherein said first pedicle screw member includes a fastener operable to be fixed to a selected anatomical portion having a first fixable portion for being fixed to the selected anatomical portion and a second engageable portion to be connected to said second elongated rod member; wherein said second elongated rod member includes a connector to operably interconnect said first pedicle screw member and said third pedicle screw member. 3. The system of claim 2, wherein said third pedicle screw member is selected from a group including a fastener, a connector, an anatomical portion, and combinations thereof. 4. The system of claim 2, wherein said fastener includes a screw with a shank and an adjustable head operable to be positioned into a vertebra; wherein said localization element is operable to be connected to said adjustable head;wherein said processor is operable to determine a position of the adjustable head. 5. The system of claim 1, wherein said first pedicle screw... | Medtronic Navigation, Inc. | A61B005/00
US08630692 | B2 | 2014-01-14 | Method and implantable system for blood-glucose concentration monitoring using parallel methodologies | 2009-04-30 | In an implantable medical device for monitoring glucose concentration in the blood, a blood-glucose concentration analysis is performed using correlations of blood-glucose concentration with measures of metabolic oxygen consumption including oxymetric, and/or temperature. Analysis of electrocardiographic data is used in a parallel method to detect and/or confirm the onset and/or existence and/or extent of hypoglycemia and/or hyperglycemia. Blood-glucose concentration calculation is enhanced by using the combination of the oxygen metabolism analysis and electrocardiographic analysis. | 1. A method, for use by an implantable medical device, for monitoring a level of blood-glucose concentration, the method comprising: (a) generating a first metric of blood-glucose concentration by monitoring one or more property of one or more cardiac electrical signal;(b) generating a second metric of blood-glucose concentration by monitoring one or more property of oxygen metabolism;(c) using one of the first metric of blood-glucose concentration or the second metric of blood-glucose concentration to detect one of hypoglycemia, normoglycemia or hyperglycemia; and(d) using the other one of first metric of blood-glucose concentration or the second metric of blood-glucose concentration to confirm or not confirm the detection at step (c) of the one of hypoglycemia, normoglycemia or hyperglycemia. 2. The method of claim 1, wherein step (a) comprises generating the first metric of blood-glucose concentration by analyzing at least one of: a T-wave amplitude, a Q-T interval, or an S-T segment deviation. 3. The method of claim 1, wherein step (b) comprises generating the second metric of blood-glucose concentration using a metabolic oxygen conformation method which utilizes cardiac output, temperature and blood-oxygen data in combination with stored information regarding a relationship of the cardiac output, temperature and blood-oxygen data to blood-glucose concentration. 4. The method of claim 1, wherein step (b) comprises generating the second metric of blood-glucose concentration by monitoring at least one of: a metric of venous oxygen saturation, a metric of hematocrit, or a metric of arterial oxygen saturation. 5. The method of claim 1, wherein step (d) comprises using the first metric of blood-glucose concentration to confirm a detection at step (c) that used the second metric of blood-glucose concentration if both the first metric of blood-glucose concentration and the second metric of blood-glucose concentration are indicative of the same one of hypoglycemia, normoglycemia or hyperglycemia. 6. The method of claim 1, wherein step (d) comprises using the second metric of blood-glucose concentration to confirm a detection at step (c) that used the first metric of blood-glucose concentration if both the first metric of blood-glucose concentration and the second metric of blood-glucose concentration are indicative of the same one of hypoglycemia, normoglycemia or hyperglycemia. 7. The method of claim 1, wherein step (a) is performed a plurality of times; step (b) is performed in response to generating at step (a) a first metric of blood-glucose concentration that is used at step (c) to detect one of hypoglycemia or hyperglycemia; and step (d) is performed in response to generating at step (b) a second metric of blood-glucose concentration. 8. The method of claim 1, wherein step (b) is performed a plurality of times; step (a) is performed in response to generating at step (b) a second metric of blood-glucose concentration that is used at step ... | Pacesetter, Inc. | A61B005/0205 A61B005/0402 A61B005/042 A61B005/0452 A61B005/145
US08755858 | B2 | 2014-06-17 | Blood sensor, blood testing device and blood analysis method | 2009-02-06 | Provided are a blood sensor that is highly reliable and does not contaminate the interior of the main body of the device, a blood testing device and a blood analysis method. The blood sensor is provided with a plate-shaped base, a blood storage part set roughly in the center of the base, a supply path, one end of which is connected to the storage part and the other end of which is connected to an air pore, multiple detecting electrodes laid on one surface of the supply path connecting electrodes, each coming out of these detection electrodes, and skin detection electrodes in the storage part or near the storage part. The blood detecting device is provided with a skin detecting circuit that detects skin contact by measuring sensor conduction or impedance changes. | 1. A blood sensor that analyzes a component in a blood sample, the blood sensor comprising: a skin detecting section comprising a skin detection electrode arranged on a surface of a blood sensor for analyzing a component in a blood sample, the skin detection electrode being arranged to directly contact skin from which the blood sample is exuded; anda connection electrode that connects with the skin detection electrode comprised in the skin detecting section and outputs a detection result by the skin detecting section to outside the blood sensor,wherein the blood sensor has a planar plate like shape and a hole therein leading to a blood storing section in the blood sensor in which the blood sample is to be stored, and the blood sensor is configured to be attached to a blood sensor apparatus comprising an electrical circuit section, and is further configured to be attached so as to contact the connection electrode with the electrical circuit section. 2. The blood sensor according to claim 1, wherein the skin detecting section has one or more skin detecting electrodes. 3. The blood sensor according to claim 1, wherein the skin detecting section is mounted in a blood supply side. 4. The blood sensor according to claim 1, wherein the skin detecting section is arranged so as not to contact at least one of an oxidation-reduction enzyme and an electron mediator. 5. The blood sensor according to claim 1, wherein the skin detecting section is arranged not to contact a reactive reagent that causes an oxidation-reduction reaction. 6. The blood sensor according to claim 1, wherein the skin detecting section is arranged not to contact a reagent. 7. The blood sensor according to claim 1, further comprising: a blood component measuring section that measures an item related to a concentration of a substance to be analyzed in the blood sample. 8. The blood sensor according to claim 7, wherein: the blood component measuring section is an analysis electrode that measures a blood component and is composed of a working electrode and a counter electrode; andthe skin detecting section is provided as a separate part from the analysis electrode. 9. The blood sensor according to claim 1, further comprising: a plate-like base body;a blood storing section provided in an end part or approximately a center part of the base body;a supply path having one side coupled to the storing section and the other side coupled to an air hole; anda detection electrode that is provided on one surface of the supply path and that is configured to detect a blood which has flowed into the supply path. 10. The blood sensor according to claim 9, wherein skin detecting sections are arranged in two spots sandwiching the storing section to face one another. 11. The blood sensor according to claim 9, wherein skin detecting sections are arranged in four spots sandwiching the storing section to face each other. 12. The blood sensor according to claim 9, wherein the detecting section comprises... | Panasonic Corporation | A61B005/05
US08543179 | B2 | 2013-09-24 | Biomedical sensor device | 2009-06-08 | A biomedical sensor device includes a light source, a probe array, and a photo detector. The light source is configured for emitting infrared radiation. The probe array is contacted to a user's skin to detect an electric wave signal transmitted through the probe array from the skin. The probe array includes a substrate and a plurality of probes mounted on the substrate, wherein the substrate and the probes are non-opaque so that the infrared radiation may be transmitted through the probe array into the skin. The photo detector is configured to detect an infrared signal by measuring the infrared radiation absorption by the skin. | 1. A biomedical sensor device, comprising: a light source, configured for emitting infrared radiation;a probe array, for contacting to a user's skin and for detecting an electric wave signal from the skin, and comprising a substrate and a plurality of probes mounted on the substrate, wherein the substrate and the probes are non-opaque whereby the infrared radiation transmits through the probe array into the skin, wherein each of the probes comprises:a tip portion;a first stalk portion, comprising a top thereof connected to the tip portion, wherein a cross-sectional area of the top is larger than that of a bottom of the first stalk portion; anda second stalk portion comprising a bottom thereof connected to the substrate, and wherein the cross-sectional area of a top of the second stalk portion is smaller than that of the bottom of the second stalk portion; anda photo detector, configured for detecting an infrared signal by measuring the infrared radiation absorption by the skin. 2. The biomedical sensor device as claimed in claim 1, wherein the photo detector further comprises a color filter. 3. The biomedical sensor device as claimed in claim 1, wherein the infrared signal is used for measuring blood oxygen concentration or blood sugar concentration. 4. The biomedical sensor device as claimed in claim 1, wherein the electric wave signal is used for measuring EEG (electroencephalography), ECG (electrocardiography) or EMG (electromyography). 5. The biomedical sensor device as claimed in claim 1, wherein the substrate and the probes are transparent. 6. The biomedical sensor device as claimed in claim 1, wherein the substrate is made of silicon, silicon dioxide or zinc oxide. 7. The biomedical sensor device as claimed in claim 1, wherein the probes are made of TCO (transparent conductive oxides). 8. The biomedical sensor device as claimed in claim 7, wherein the TCO comprises ITO (tin doped Indium oxide), FTO (fluorine doped tin dioxide), AZO (aluminum doped zinc oxide) or GZO (gallium doped zinc oxide). 9. The biomedical sensor device as claimed in claim 1, wherein the tip portion and the first stalk portion are made of silicon, silicon dioxide or zinc oxide. 10. The biomedical sensor device as claimed in claim 1, wherein each of the probes further comprises a conductive layer covering the tip portion and the first stalk portion. 11. The biomedical sensor device as claimed in claim 10, wherein the conductive layer is made of a polymer, a thin film material or a metal. 12. The biomedical sensor device as claimed in claim 1, wherein the tip portion is configured for puncturing stratum corneum and the stratum germinativum of the skin to allow the probe to penetrate into the stratum germinativum. 13. The biomedical sensor device as claimed in claim 1, wherein a length of the tip portion is about 50-150 μm. 14. The biomedical sensor device as claimed in claim 1, wherein the length of the tip portion is about 80-100 μm. 15. The biomedical se... | National Chiao Tung University | A61B005/00
US08447382 | B2 | 2013-05-21 | Positioning system and method of position detecting | 2010-05-26 | A body-insertable device including a resonance circuit for emitting a resonance magnetic field by being induced by a drive magnetic field received from the outside. An external device for deriving position information of the resonance circuit by using a detection signal of detected resonance magnetic field. The external device includes a drive coil driving unit for outputting a drive signal having a predetermined frequency, a drive coil for receiving the output drive signal and generating the drive magnetic field, and a drive magnetic field generation control unit for calculating an amplitude value of the drive signal by using a current amplitude value of the drive signal, detecting a degree of stability indicating whether a sharp change occurs in the drive magnetic field based on the calculated amplitude value, and controlling the deriving of a position information based on the detected degree of stability in the drive magnetic field. | 1. A position detecting system comprising: a body-insertable apparatus disposed in a state where it is introduced in a subject in a detection space; andan external device disposed on the outside of the subject, whereinthe body-insertable apparatus includes a resonance circuit for emitting a resonance magnetic field by being induced by a drive magnetic field received from the outside, andthe external device includes: a drive coil driving circuit configured to output a drive signal having a predetermined frequency;a drive coil configured to receive the drive signal and to generate the drive magnetic field in the detection space based on the drive signal;a sense coil configured to detect the resonance magnetic field and to output a detection signal;a position deriving circuit configured to derive position information of the resonance circuit by using the detection signal;a current detecting circuit configured to detect a current amplitude value of the drive signal input to the drive coil; anda drive magnetic field generation control circuit configured to calculate an amplitude value of a drive signal which is to be output by the drive coil driving circuit based on the current amplitude value,detect a degree of stability of the drive magnetic field on the basis of whether a sharp change occurs in the calculated amplitude value, andcontrol the position deriving circuit to derive position information on the basis of the detected degree of stability of the drive magnetic field. 2. The position detecting system according to claim 1, wherein: the drive magnetic field generation control circuit is configured to detect the degree of stability of the drive magnetic field by calculating a difference between a previously calculated amplitude value and a currently calculated amplitude value,comparing the difference with a preset reference value, andoutputting a result of the comparison as a comparison result signal indicative of the degree of stability to the position deriving circuit, andthe position deriving circuit is further configured to validate or invalidate the detection signal received from the sense coil or the derived position information on the basis of the comparison result signal. 3. The position detecting system according to claim 2, wherein the position deriving circuit is configured to invalidate the detection signal received from the sense coil or the derived position information when the input comparison result signal indicates that the difference is larger than the reference value. 4. The position detecting system according to claim 2, wherein the position deriving circuit includes: a position calculating unit configured to derive the position information on the basis of an average value of a first predetermined number of times of the detection signals received from the sense coil; andan average number increasing/decreasing unit configured to increase/decrease the first predetermined number of times on the basis of the input comparison re... | Olympus Medical Systems Corp. | A61B005/05
US08670943 | B2 | 2014-03-11 | Systems and methods for applying scaling laws of tree structures | 2008-08-12 | Methods are provided for determining flow resistance within at least a portion of a vessel, whereby at least a partial biological tree image is obtained and resistance is calculated therefrom based upon at least the length of a defined vessel portion and a diameter of a stem vessel at a location proximal to a vessel crown. Methods are also provided for determining vessel volume, whereby at least a partial biological tree image is obtained and cumulative crown volume is calculated based on at least the diameter and length of the stem segment. | 1. A method for diagnosing disease in a patient's biological tree, the method comprising the steps of: obtaining a model biological tree;using a processor to produce an image showing a vasculature of at least part of a patient's biological tree, wherein the processor is operably connected to a storage medium capable of receiving and storing the image;identifying a length of a patient vessel portion and a diameter of a patient stem vessel proximal to the vessel crown from the patient's biological tree image;calculating model resistance based upon at least a length of a model vessel portion and a diameter of a model stem vessel proximal to the vessel crown from the model biological tree image;calculating patient resistance based upon at least the length of the patient vessel portion and the diameter of the patient stem vessel proximal to the vessel crown from the patient's biological tree image; andcomparing the calculated model resistance to the calculated patient resistance to determine the extent of vessel and/or organ disease. 2. The method of claim 1, wherein the steps of calculating model resistance and calculating patient resistance are further based upon the use of a constant. 3. The method of claim 1, wherein the patient resistance is a resistance of a stem segment, and wherein the length of the patient vessel portion is a length of a stem. 4. The method of claim 3, wherein the step of calculating the resistance of a stem segment is performed by dividing the length of a stem by the diameter of the patient stem vessel proximal to the vessel crown to the fourth power multiplied by a constant for the stem, or a mathematical equivalent thereof. 5. The method of claim 4, wherein the constant for the stem is equivalent to one hundred and twenty eight multiplied by a fluid viscosity divided by pi, or a mathematical equivalent thereof. 6. The method of claim 1, wherein the patient resistance is a resistance of a crown, and wherein the length of the patient vessel portion is a crown length. 7. The method of claim 6, wherein the step of calculating the resistance of a crown is further based upon the use of a constant for the crown. 8. The method of claim 7, wherein the constant for the crown is equivalent to a maximum resistance multiplied by a maximum diameter to the fourth power divided by a maximum length, or a mathematical equivalent thereof. 9. The method of claim 6, wherein the step of calculating the resistance of a crown is performed by dividing the crown length by the diameter of the patient stem vessel proximal to the vessel crown to the fourth power multiplied by a constant, or a mathematical equivalent thereof. 10. The method of claim 1, wherein the patient resistance is a total resistance of the biological tree, and wherein the length of the patient vessel portion is a cumulative biological tree vessel length. 11. The method of claim 10, wherein the step of calculating the total resistance of the biological tree is further bas... | DTherapeutics, LLC | G06F019/00 A61B005/00
US08241301 | B2 | 2012-08-14 | Guided puncturing needle and puncturing guiding method | 2007-01-23 | A guide puncturing needle which is integratively used with a scanning and detecting device, such as CT, MRI or the like, and provides puncturing, biopsy, injection, implanting, and the physical diagnose and therapy, such as RF, microwave, freezing, laser and the like, and a puncturing guide method for applying the puncturing needle to puncture are provided. Since the guide puncturing needle is provided with a needle-entering angle guide means (2), a puncturing layer levelling component (31), a needle-entering reference line levelling component (32), whether the CT gantry and the MRI scanning layer are at any angle, the actual needle-entering layer is always adjusted to just superpose to the CT scanning layer, the actual needle-entering angle is always just the same as the needle-entering angle needed for the scanning and detecting device, thus allow the puncturing needle to arrive at the focus point exactly accurately. | 1. A puncturing needle guiding device comprising: i) a needle-entering angle guide, adapted to be rotatably connected a needle when placed in said guide via a connector, further including: i.a) an angle reference line (N) vertical to a surface where the needle enters, andi.b) an indicator of said guide, including a needle-entering angle indicating plate, for indicating an intersected angle between an actual enter direction of the needle and the angle reference line (N),ii) a needle-entering layer leveler, sidely affixed to the guide, further comprising ii.a) an indicator of said layer leveler, including a scanning layer leveling benchmark, for indicating a deviatingly rotated angle of a rotating plane of the needle, wherein a deviating angle rotating axis of the indicator of said layer leveler is parallel to the guide (W), andiii) a reference line leveler, making the angle reference line of the guide (W) on the same plane of a hammer weight line, further comprising: iii.a) an indicator of said line leveler, including an angle indicating base plate, whereina deviatingly rotated axis y of the indicator of said line leveler is perpendicular to the deviating angle rotating axis; andiv) a level benchmark, making a horizontal line of the rotating plane parallel to a needle-entering layer, andv) a top open rectangular casing, including a top opening and is affixed sidely to the guide, with the layer leveler and the line leveler (Y) disposed within, wherein the indicator of the layer leveler and the indicator of the line leveler are exposed in the top opening mounted with a transparent cover. 2. The puncturing needle guiding device as claimed in claim 1, wherein a rotating spindle is at the same axis with the angle rotating axis of the layer leveler; said layer leveler and said line leveler being integrally fixed and pivotally installed on said casing with the rotating spindle as a pivotal axis;and said rotating spindle includes a protruding extended out of said casing, with an external surface thereof being installed with a rotating angle scale for the rotating spindle, and a knob with a pointer affixed to said protruding. 3. A puncturing guiding method for targeting a focus point by the needle of a puncturing needle guiding device as claimed in claim 1, comprising the following steps: 1.a) Obtaining 3D coordinates of the focus point by a CT or an MRI scanning device;1.b) Obtaining a puncturing sectional layer (Y0Z) from the CT or MRI scanning device,1.c) marking on a body skin a needle-entering point, and1.d) marking on the body skin a projection line of a scanning layer obtained from the CT or MRI scanning device;1.e) calculating a needle-entering angle, based on an angle of a line connecting the point and the needle-entering point relative to a horizontal line, or relative to a vertical line of the horizontal line;calculating a needle-entering depth, based on a length of the line;2) installing the needle on the needle-entering angle indicating plat... | Huilin Zhang Su Feng Zhejiang Kindly Medical Devices Co., Ltd. | A61B019/00
US08170655 | B2 | 2012-05-01 | Module and device for discerning therapeutic signals from noise in physiological data | 2010-07-20 | A data acquisition module for use in monitoring a plurality of physiological signals is disclosed herein. The data acquisition module may include a first signal processing path for biopotential data, a second signal processing path for therapeutic event data, and a processing unit that receives and processes the data from the first and second signal processing paths. The data acquisition module may further compare identified likely therapeutic events in each of a plurality of psychological signals. | 1. A data acquisition device for use in monitoring a therapeutic event from a plurality of physiological signals each comprising therapeutic event data and biopotential data, the data acquisition module comprising: a signal processing path that isolates the therapeutic event data from each of the plurality of physiological signals;a signal processing module that receives the therapeutic event data from the signal processing path and applies a first signal processing technique to the therapeutic event data to identify likely therapeutic events from each of the plurality of physiological signals;a precordial sweep module that compares the identified likely therapeutic events from each of the plurality of physiological signals; anda therapeutic event determination module that identifies therapeutic events based upon the comparison by the precordial sweep module. 2. The data acquisition device of claim 1, wherein the biopotential data is electrocardiographic data and the therapeutic event data is pacemaker data. 3. The data acquisition device of claim 2, wherein the signal processing path comprises a filter. 4. The data acquisition device of claim 3, wherein the filter passes therapeutic event data with a frequency greater than 250 Hertz. 5. The data acquisition device of claim 1, wherein the signal processing module is computer executable code embodied on a non-transient computer readable medium that is executed by a computer processor. 6. The data acquisition device of claim 1, further comprising: a plurality of biopotential electrodes and each biopotential electrode acquires a physiological signal; anda multiplexer connected to the plurality of biopotential electrodes and the signal processing path, such that the multiplexer directs the plurality of physiological signals through the signal processing path. 7. A data acquisition module for use in monitoring a plurality of physiological signals comprising electrocardiographic data and pacemaker data, the data acquisition module comprising: a first signal processing path for isolating and enhancing the electrocardiographic data of the plurality of physiological signals;a second signal processing path for identifying pacemaker pulse data from the plurality of physiological signals;a processing unit that receives and processes the electrocardiographic data from the first signal processing path and the pacemaker pulse data from the second signal processing path;a signal processing module associated with the processing unit, the signal processing module receiving the pacemaker pulse data, applying a first signal processing technique to identify likely pacemaker pulses;a precordial sweep module that compares the identified likely pacemaker pulses from each of the plurality of physiological signals; anda pacemaker pulse determination module that identifies pacemaker pulses based upon the results from the signal processing module and the precordial sweep module. 8. The data acquisition module of clai... | The General Electric Company | A61B005/04
US08155733 | B2 | 2012-04-10 | Lie detection method and system | 2009-05-15 | The present disclosure provides a heart rate variability (HRV) analysis-based method of lie detection, related computer program product, computer readable storage medium and system. Also provided is a method of HRV analysis using strange entropy. | 1. A method of lie detection, comprising: receiving an input associated with heart rate variability “HRV”; andperforming an HRV analysis based on the input associated with HRV to obtain a lie detection result; wherein said HRV analysis includes nonlinear HRV analysis, wherein the nonlinear HRV analysis includes the following:taking a time series with N elements, u :{u(i): 1 ≦i≦N}, wherein u(i) represents signals carried in the input associated with HRV, and for each u(i), there is a corresponding vector with m elements: X(i)=[u(i),u(i+1), . . . , u(i+m−1)];calculating a base scale BS(i) for each vector X(i);transforming every X(i) to a m-dimensional symbol series Si={s(i), . . . s(i+m−1)},sεA(A=0, 1, 2, 3) based on a scale of a×BS(i);calculating the probability of Si, wherein the probability of each different combination among the entire N−m+1 m-dimensional vectors is:p⁡(π)=#⁢⁢{t¦(ui,…⁢,ut+m-1)⁢⁢has⁢⁢type⁢⁢π}N-m+1,wherein 1 ≦t ≦N−m+1; # is the number of states, and each possible combination π for Si represents a vibration mode for Si;obtaining nonlinear HRV parameters or m-word distribution graph according to the probability of Si; andobtaining a nonlinear HRV parameter PSS, which is defined as sum probabilities of all strange states calculated according to the following equation:P⁢⁢S⁢⁢S=∑Ps∑i=0255⁢Pi;wherein Pi is probability of combination i, Ps is probability of strange state, and ΣPs is probability sum of all strange states. 2. A lie detection system, comprising: a computing unit configured toreceive an input associated with HRV; andperform an HRV analysis based on the input associated with HRV to obtain a lie detection result; wherein said HRV analysis includes nonlinear HRV analysis, wherein the nonlinear HRV analysis includes the following:obtaining a nonlinear HRV parameter PSS, which is defined as sum probabilities of all strange states calculated according to the following equation:P⁢⁢S⁢⁢S=∑Ps∑i=0255⁢Pi;wherein Pi is probability of combination i, Ps is probability of strange state, and ΣPs is probability sum of all strange states. 3. The lie detection system according to claim 2, wherein said nonlinear HRV analysis is performed by using strange entropy method comprising: taking a time series signal with N elements, u :{u(i): 1 ≦i≦N} ,wherein u(i) represents signals carried in the input associated with HRV, and for each u(i), there is a corresponding vector with m elements: X(i)=[u(i),u(i+1), . . . , u(i+m−1)];calculating a base scale BS(i) for each vector X(i);transforming every X(i) to a m-dimensional symbol series Si={s(i), . . . s(i+m−1)},sεA(A=0, 1, 2, 3) based on a scale of a×BS(i);calculating probability of Si such that the probability of each different combination among the entire N−m +1 m-dimensional vectors is:p⁡(π)=#⁢⁢{t¦(ui,…⁢,ut+m-1)⁢⁢has⁢⁢type⁢⁢π}N-m+1,wherein 1 ≦t ≦N −m+1 ,# is the number of states, and each possible combination π for Si represents a vibration mode for Si; andobtaining one or more of nonlinear HRV parame... | Nanjing University | A61B005/0468
US08196694 | B2 | 2012-06-12 | Vehicle immobilizer methods and apparatus based on driver impairment | 2009-05-21 | Embodiments include systems and methods for immobilizing a motor vehicle. A system comprises an impairment sensor system and a control subsystem. The impairment sensor system is adapted to perform an analysis of a sample provided by an operator of a motor vehicle, where the analysis includes determining an impairment-related metric based on the sample. The control subsystem is adapted to control at least one mobility-related apparatus and at least one non-mobility-related apparatus of the motor vehicle. When the result of the analysis indicates that the impairment-related metric does not meet a criteria, the control subsystem is adapted to control the at least one mobility-related apparatus in a manner that disables the motor vehicle from moving and to control the at least one non-mobility-related apparatus in a manner that allows the at least one non-mobility-related apparatus to operate. | 1. A system comprising: an impairment sensor system adapted to perform an analysis of a sample provided by an operator of a motor vehicle, wherein the analysis includes determining an impairment-related metric based on the sample; anda control subsystem, operatively coupled with the impairment sensor system, and adapted to control at least one mobility-related apparatus and at least one non-mobility-related apparatus of the motor vehicle, wherein, in response to a result of the analysis indicating that the impairment-related metric does not meet a criteria thus indicating that the operator is impaired, the control subsystem transitions to a no-mobility operational mode from a different operational mode, wherein transitioning to the no-mobility operational mode includes the control subsystem disabling the motor vehicle from moving through control of the at least one mobility-related apparatus, and simultaneously allowing the at least one non-mobility-related apparatus to operate through control of the at least one non-mobility-related apparatus, and wherein the different operational mode is chosen from a non-operational mode and a full operational mode, wherein the non-operational mode is an operational mode in which an engine of the motor vehicle is off, and wherein the full operational mode is an operational mode in which the engine is enabled and the motor vehicle is capable of being mobilized. 2. The system of claim 1, wherein the at least one mobility-related apparatus comprises: a shift control mechanism of the motor vehicle, which is controlled by locking the shift control mechanism. 3. The system of claim 1, wherein the at least one mobility-related apparatus comprises: a transmission system of the motor vehicle. 4. The system of claim 1, wherein the at least one mobility-related apparatus comprises an apparatus selected from a group of apparatus consisting of a throttle, a fuel delivery system, propulsion battery energy, and an electric park brake system. 5. The system of claim 1, wherein the at least one non-mobility-related apparatus comprises a heating, ventilation, and air conditioning system of the motor vehicle. 6. The system of claim 1, wherein the at least one non-mobility-related apparatus comprises an entertainment system of the motor vehicle. 7. The system of claim 1, wherein the at least one non-mobility-related apparatus comprises the engine of the motor vehicle. 8. The system of claim 1, wherein the sample includes a breath sample, wherein the analysis includes an analysis of the sample to determine the impairment-related metric as a blood alcohol concentration, wherein the result indicates the blood alcohol concentration, and wherein the control subsystem controls the at least one mobility-related apparatus in the manner that disables the motor vehicle when the blood alcohol concentration is above a threshold. 9. A system comprising: an impairment sensor system adapted to perform an analysis of a sample provided by ... | GM Global Technology Operations LLC | B60K028/06
US08496584 | B2 | 2013-07-30 | Machine for rapid identification of infections and/or risk situations related to gastroduodenal pathologies | 2010-07-26 | A machine for performing a method for immediate identification and rapid comparative assessment of indicators of the presence of infections and/or risk situations related to gastroduodenal pathologies, characterized in that it comprises a measurement unit that is constituted by a container supplied by means for aspirating gastric juices in which an agitator of the mixture of gastric juice, reagents and water operates, and at least one probe for sensing the values of the analysis, a hydraulic section that comprises a set of tanks for the reagents required for the analysis and for the water, which are connected to respective apparatuses for feeding said measurement unit and to a pump for feeding water into the stomach of the patient, an apparatus for controlling the temperature of said container, and an electrical section, which comprises a programmable control unit for actuating and controlling said feeder apparatuses and said measurement unit and is suitable to determine the test execution methods. | 1. A machine in combination with an endoscopic unit and a collection vessel, the machine having an inlet and an outlet, wherein the endoscopic unit has a suction duct, the suction duct being connected to the inlet of the machine, and wherein the collection vessel is connected to an exit of the machine and to a suction unit, the machine being able to identify presence/absence of H. pylori infection and acidity condition of a patient, the machine being constituted by: a measurement unit, the measurement unit being constituted bya container, the container being provided with means for aspirating gastric juices and an agitator for a mixture of gastric juice, water and reagents required for analysing the gastric juices, andat least a pH measurement probe and an ammonia measurement probe,a hydraulic section that comprises a set of tanks for the reagents and for the water,means for feeding said measurement unit, said means for feeding being connected to the tanks,a pump for feeding water into the stomach of the patient, the pump being connected the tanksan apparatus for controlling temperature of said container, andan electrical section, the electrical section comprising a programmable control unit, the programmable control unit being able to actuate and control said means for feeding the measurement unit and able to actuate and control said measurement unit, and comprising a microprocessor, wherein the microprocessor is configured to calculate pH and ammonium values of the aspired gastric juice based on measurements of the pH measurement probe and of the ammonia measurement probe, then to compare these pH and ammonium values with preset reference pH and ammonium values and to report verbally to an operator of the endoscope unit, within 2 minutes from the gastric juice aspiration, an acidity condition, a presence of H. pylori infection and a diagnostic procedure to be followed during endoscopy. 2. The machine according to claim 1, wherein said means for aspirating gastric juices comprise at least one pump. 3. The machine according to claim 1, wherein said means for feeding the measurement unit comprise at least one pump for sending said reagents and water and/or air. 4. The machine according to claim 1, wherein said apparatus for controlling the temperature of said container is a generator of heating air at an adjustable temperature. 5. The machine according to claim 1, wherein said container of said measurement unit is of a substantially closed type and is provided with connections for a plurality of ducts for a flow of liquids and/or air and for access of probes. 6. The machine according to claim 5, wherein said container is substantially cylindrical, said container comprising at least one temperature sensor and at least one level sensor suitable to indicate a presence of certain volumes of liquids within said container. 7. The machine according to claim 5, wherein said container comprises, on a side wall thereof, an emptying hole to which an in... | Antonio Tucci | A61B005/00 A61B001/00 B65D081/00
US10238306 | B2 | 2019-03-26 | Method for non-evasively determining an endothelial function and a device for carrying out said method | 2006-04-03 | The invention relates to medicine. The inventive method for non-evasively determining an endothelial function consists in reducing a transmural pressure in a limb, recording the amplitude of plethysmographic signals at different pressures, determining the pressure of a maximum plethysmographic signal amplitude, reducing the pressure to a value corresponding to the specified percentage of the maximum amplitude, carrying out an occlusion sample in the course of which a pressure higher by at least 50 mm Hg than the systolic pressure of a tested patient is produced in a cuff arranged proximally to the located limb area, wherein the occlusion is carried out during at least 5 minutes. The inventive device for carrying out said method comprises a sensory unit, a pressure producing unit, an electronic unit constructed in such a way that it makes it possible to measure the cuff pressure corresponding to the maximum plethysmographic signal amplitude and a unit for controlling the pressure producing unit in such a way that the cuff pressure corresponding to the amplitude of a plethysmographic signal, which represents a specified percentage of the maximum amplitude, is attained. The sensory unit is connected to the electronic unit to the output of which the pressure producing unit is connected. | 1. A method for noninvasive assessment of endothelial function in a patient, which method comprises: a. lowering a transmural pressure in an artery in an extremity by means of a pressure generating cuff;b. measuring, via a monitoring channel, an amplitude of a plethysmographic signal at various transmural pressures generated in step (a);c. determining a pressure in the pressure generating cuff at which the amplitude of the plethysmographic signal is maximal;d. lowering the pressure in the pressure generating cuff to 50% of the pressure at which said maximal amplitude occurs;e. when the pressure in the pressure generating cuff is 50% of the pressure at which said maximal amplitude occurs, running an occlusion test; wherein running an occlusion test comprises generating a pressure in the pressure generating cuff on an observed region of the extremity which exceeds a measured systolic pressure by at least 50 mmHg for at least 5 minutes;f. lowering the pressure in the pressure generating cuff to zero;g. measuring an amplitude of the plethysmographic signal for at least 3 minutes and simultaneously measuring, via a reference channel, an amplitude of a plethysmographic signal detected from an artery in a second extremity where the occlusion test has not been performed;h. using amplitudal and timely analyses to compare data collected through the reference and monitoring channels; andi. providing an assessment of the patient's endothelial function based on the amplitudal and timely analyses. 2. The method according to claim 1, wherein pressure applied to the extremity is increased gradually, with each step being 5 mmHg and lasting 5-10 seconds, at the same time recording the amplitude of the plethysmographic output. 3. The method according to claim 1, wherein (i) during the amplitudal analysis data from the reference channel are compared to those collected from the monitoring channel; (ii) a rate of the increase of the amplitude in the monitoring channel is analyzed; (iii) a relationship of an amplitude of a maximum plethysmographic signal recorded during various transmural pressures is compared to a maximal amplitude of a plethysmographic signal recorded after the running of the occlusion test; and/or (iv) during the temporal analysis, plethysmographic waveforms collected through the reference and monitoring channels are compared, the plethysmographic waveforms are then normalized and a time delay or phase change is determined. 4. An apparatus for noninvasive assessment of endothelial function comprising a double-channeled sensor unit, capable of detecting pulse wave signals from peripheral arteries, a pneumatic unit capable of creating gradually increasing pressure in a pressure generating cuff, a pressure sensor configured to detect a blood pressure of a patient; and an electronic block configured to perform at least the following steps: a. lowering a transmural pressure in an artery in an extremity by means of a pressure generating cuff;b. measuri... | Everist Genomics, Inc. | A61B005/02 A61B005/029 A61B005/0295 A61B005/021 A61B005/024 A61B005/022 A61B005/00
US10085657 | B2 | 2018-10-02 | Body-worn pulse oximeter | 2009-09-14 | The invention provides a body-worn system that continuously measures pulse oximetry and blood pressure, along with motion, posture, and activity level, from an ambulatory patient. The system features an oximetry probe that comfortably clips to the base of the patient's thumb, thereby freeing up their fingers for conventional activities in a hospital, such as reading and eating. The probe secures to the thumb and measures time-dependent signals corresponding to LEDs operating near 660 and 905 nm. Analog versions of these signals pass through a low-profile cable to a wrist-worn transceiver that encloses a processing unit. Also within the wrist-worn transceiver is an accelerometer, a wireless system that sends information through a network to a remote receiver, e.g. a computer located in a central nursing station. | 1. A system for monitoring both blood pressure and oxygen saturation values from a patient, comprising: a pulse oximetry sensor comprising: a first radiation source configured to emit a first optical radiation at a first wavelength;a second radiation source configured to emit a second optical radiation at a second wavelength;a substrate comprising a first substantially planar surface which comprises the emitting surface of the first and second radiation sources;a photodetector configured to detect the first optical radiation emitted by the first radiation source after the first optical radiation passes through a portion of the patient to generate a first signal, and the second optical radiation emitted by the second radiation source after the second optical radiation passes through a portion of the patient to generate a second signal;a substrate comprising a second substantially planar surface which comprises the receiving surface of the photodetector;and a finger-ring housing comprising the first and second radiation sources and the photodetector, the finger-ring housing configured to attach to a base of the patient's thumb such that the first optical radiation emitted by the first radiation source and the second optical radiation emitted by the second radiation source pass through the thumb tissue and are received by the princeps pollicis artery before reaching the photodetector, wherein the first and second radiation sources are proximal to one another and separated from the photodetector by an included angle between 35-55 degrees, wherein the included angle is the angle formed between a first line drawn parallel to the first planar surface and a second line drawn parallel to the second planar surface;a second sensor comprising: at least two electrodes, each configured to attach to the patient and measure electrical signals and an electrical circuit configured to receive and process the electrical signals to determine a time-dependent electrical signal; anda processing unit configured to be worn on the patient's arm and connected to the pulse oximetry sensor to receive the first and second time-dependent signals, and to the second sensor to receive the time-dependent electrical signal, the processing unit comprising a processor configured to process: i) the first and second time-dependent signals to determine AC signals, or signals derived therefrom, to determine the oxygen saturation value; and ii) at least one of the AC signals and the electrical signal, or signals derived therefrom, to determine the blood pressure value. 2. The system of claim 1, wherein the processing unit further comprises a motion sensor. 3. The system of claim 2, wherein the motion sensor is an accelerometer. 4. The system of claim 1, wherein the processing unit further comprises an input port for receiving a cable. 5. The system of claim 4, wherein the cable further comprises an accelerometer. 6. The system of claim 4, wherein the cable further comprises an ECG ci... | SOTERA WIRELESS, INC. | A61B005/1455 A61B005/0245 A61B005/00 A61B005/021 A61B005/022 A61B005/0452 A61B005/11
US09439585 | B2 | 2016-09-13 | Semiconductor based analyte sensors and methods | 2009-02-03 | An analyte sensor is provided that comprises a substrate which includes a semiconductor material. Embodiments may include a core of a conductive material, and a cladding of a semiconductor material, in which the cladding may form at least a portion of a conducting path for a working electrode of the analyte sensor. Method of manufacturing and using the analyte sensor are described, as are numerous other aspects. | 1. An analyte sensor, comprising: a substrate including a core of conductive material and a cladding of a semiconductor material, wherein the cladding is disposed around the core,a protective layer disposed concentrically around, immediately adjacent, and in contact with the substrate, the protective layer including a base layer positioned between a reference electrode and the cladding, wherein an additional layer covers the reference electrode, anda working electrode formed from a portion of the cladding not covered by the protective layer, the working electrode being at least partially coated with an electrochemically-active layer formed from at least one of a noble metal, carbon nanotubes, and a conductive polymer,an active region positioned on the substrate and adapted to be exposed to a fluid environment, the active region including an agent on the electrochemically-active layer on the working electrode. 2. The analyte sensor of claim 1, wherein the cladding has a tensile strength of at least 1000 MPa. 3. The analyte sensor of claim 1, wherein the conductive material comprises carbon. 4. The analyte sensor of claim 1, wherein the substrate has a diameter of 150 microns or less. 5. The analyte sensor of claim 1, wherein the substrate has a diameter of about 75 to 150 microns. 6. The analyte sensor of claim 1, wherein the substrate has a length of about 2.5 mm to 100 mm. 7. The analyte sensor of claim 1, wherein the semiconductor material comprises silicon carbide. 8. The analyte sensor of claim 1, wherein the agent includes a catalyst agent, wherein the catalyst agent is adapted to convert an analyte within the fluid environment into a product from which an electrical current is generated at the working electrode. 9. The analyte sensor of claim 8, wherein the portion of the cladding forming the working electrode is 1% to 50% covered by the electrochemically-active layer. 10. The analyte sensor of claim 8, wherein the portion of the cladding forming the working electrode is 51% to 100% covered by the electrochemically-active layer which has a thickness of less than a micron. 11. The analyte sensor of claim 10, wherein the electrochemically-active layer includes at least one of platinum, gold, palladium, and rhodium. 12. The analyte sensor of claim 8, further comprising: a mediator adapted to provide a transfer of electrons from the catalytic agent to produce an electrical current reading correlative with an analyte concentration level is derived. 13. The analyte sensor of claim 1, wherein the substrate is cylindrical. 14. The analyte sensor of claim 1, wherein the substrate is planar. 15. The analyte sensor of claim 1, wherein the substrate includes a point forming a lancet adapted to facilitate insertion. | Ascensia Diabetes Care Holdings AG | A61B005/05 A61B005/145 A61B005/1486
US20100298719 | A1 | 2010-11-25 | METHOD FOR CALCULATING PRESSURES IN A FLUID STREAM THROUGH A TUBE SECTION, ESPECIALLY A BLOOD VESSEL WITH ATHEROSCLEROTIC PLAQUE | 2008-10-31 | A method for calculating pressures in a fluid streaming through a tube section from an upstream end to a downstream end of the tube section, the method comprising scanning the tube section with a scanner and providing a plurality of 2D scanning images along the tube section with an inlet and at least two arms, by a computer program on the basis of the 2D images automatically N calculating a 3D image of the tube section by using interpolating between the 2D images, by a computer program performing a 2D sectional image cut through the 3D image, the image cut following the fluid stream, calculating in the sectional image cut a fluid pressure distribution along multiple locations inside the tube on the basis of given boundary conditions, the boundary conditions including fluid velocity or fluid pressure at the upstream end. | 1. A method for calculating pressures in a fluid stream through a tube section from an upstream end to a downstream end of the tube section, the tube section comprising at least one flow divider dividing the fluid stream from an inlet to at least two arms, each arm with an outlet, the method comprising scanning the tube section with a scanner and providing a plurality of cross-sectional 2D scanning images along the fluid stream; said cross-sectional 2D scanning images being a 2D slicing of the section along the flow direction in the tube but not necessarily normal to a fluid direction,by a computer program on the basis of the 2D images automatically calculating a 3D image of the tube section by using interpolation between the 2D images,by a computer program performing a 2D sectional image cut through the 3D image, the image cut following the fluid stream in the tube section with the arms,calculating in the sectional image cut a fluid pressure distribution along the tube section on the basis of given boundary conditions, the boundary conditions including fluid velocity or fluid pressure at the upstream end. 2. The method according to claim 1, wherein the method implies calculating stresses and deformations of the tube section in cross-sectional 2D models,providing cross-sectional scans of the tube section depicting the actual deformation in different cross sectional planes forcomparing the calculated cross sectional deformation with the actual deformation for validating the chosen tissue parameters and examining the accuracy of the simulation. 3. The method according to claim 1, wherein the method implies measuring the fluid pressure at the upstream end and measure the fluid velocity at the downstream end and include these measurements in the boundary conditions. 4. The method according to claim 3, wherein the method implies measuring the fluid velocity at the upstream end, and calculating on the basis on the pressure distribution in the tube section the expected fluid velocity at the upstream end, and comparing the measured fluid velocity and the expected fluid velocity at the upstream end as part of an evaluation procedure for the validity of the pressure calculation. 5. The method according to claim 1, wherein the method implies measuring the fluid velocity at the upstream end and wherein the boundary conditions include the measured fluid velocity at the upstream end and an estimated pressure at the downstream end. 6. The method according to claim 5, wherein the method implies measuring the fluid velocity at the downstream end, and calculating on the basis on the pressure distribution in the tube section the expected fluid velocity at the downstream end, and comparing the measured fluid velocity and the expected fluid velocity at the downstream end as part in an evaluation procedure for the validity of the pressure calculation. 7. The method according to claim 1, wherein the tube section comprises a material deposit lump on the inner wall... | - | A61B005/021
US20100331708 | A1 | 2010-12-30 | MONITORING CARDIOVASCULAR CONDITIONS USING SIGNAL TRANSIT TIMES | 2010-06-23 | Methods for monitoring cardiovascular conditions, i.e., hyperdynamic circulation, vasodilation, vasoconstriction, or central-to-peripheral arterial pressure decoupling conditions are described. These methods involve measuring a central signal proportional to or a function of the subject's heart activity and a peripheral signal proportional to or a function of a signal related to the central signal. Then calculating a time difference between features in the central and peripheral signals representing the same heart event. The cardiovascular condition is indicated if the time difference is greater or lower than a threshold value, if the time difference is greater or lower than a threshold value over a specified period of time, or if there is a significant statistical change in the times over the specified time period. These methods can alert a user that a subject is experiencing the cardiovascular condition, which can enable a clinician to appropriately provide treatment to the subject. | 1. A method for monitoring a cardiovascular condition in a subject comprising: measuring a central signal proportional to or a function of the subject's heart activity;measuring a peripheral signal proportional to or a function of a signal related to the central signal; andcalculating a time difference between signal features representing the same heart events for the central signal and the peripheral signal,wherein the cardiovascular condition is indicated if the time difference is greater than a threshold value. 2. A method for monitoring a cardiovascular condition in a subject comprising: measuring a central signal proportional to or a function of the subject's heart activity;measuring a peripheral signal proportional to or a function of a signal related to the central signal; andcalculating a time difference between signal features representing the same heart event for the central signal and the peripheral signal,wherein the cardiovascular condition is indicated if there is a significant statistical change in the time difference over a specified time period. 3. The method of claim 1 or 2, wherein the cardiovascular condition is vasodilation. 4. The method of claim 1 or 2, wherein the cardiovascular condition is vasoconstriction. 5. The method of claim 1 or 2, wherein the cardiovascular condition indicates the occurrence of hyperdynamic circulation. 6. The method of claim 1 or 2, wherein the cardiovascular condition is central-to-peripheral pressure or flow decoupling. 7. The method of any of claim 3, 5, or 6, wherein the threshold value is 150 milliseconds or greater. 8. The method of claim 4, wherein the threshold value is 100 milliseconds or less. 9. The method of claim 2, wherein the statistically significant change is 50 milliseconds or greater. 10. The method of claim 2, wherein the statistically significant change is 0.4 standard deviations or greater. 11. The method of claim 1, wherein the threshold value is 5 minutes or greater. 12. The method of claim 2, wherein the specified time period is 5 minutes or greater. 13. The method of claim 1, wherein the cardiovascular condition is indicated is the time difference is greater than the threshold value for a specified time period. 14. The method of claim 13, wherein the specified time period is 5 minutes or greater. 15. The method of claim 1 or 2, wherein the central signal proportional to or a function of the subject's heart activity is one or more of electrocardiogram, aortic pressure, aortic flow, ultrasound, Doppler, transthoracic bioimpedance, or heart sounds. 16. The method of claim 1 or 2, wherein the peripheral signal proportional to or a function of a peripheral equivalent to the first signal is peripheral pressure, peripheral flow, ultrasound, Doppler, tonometry, or pulse oximetry. 17. The method of claim 1 or 2, wherein the central signal is aortic pressure and the peripheral signal is peripheral pressure. 18. The method of claim 1 or 2, wherein the central signa... | Edwards Lifesciences Corporation | A61B005/02
US20100312123 | A1 | 2010-12-09 | CARDIAC MEASUREMENT SYSTEM AND METHOD | 2007-05-22 | A method of monitoring blood flows through the body includes performing a substantially continuous blood pressure monitoring of a vessel at a first predetermined location to obtain a substantially continuous blood pressure signal; performing blood flow monitoring of a vessel at a first predetermined location to obtain a substantially continuous blood flow signal; providing an updated cardiac measure of said patient by continuously processing the flow signal and the pressure signal. | 1. A method for monitoring blood flows through a patient, the method comprising: (a) performing a substantially continuous blood pressure monitoring of a vessel at a first predetermined location to obtain a substantially continuous blood pressure signal;(b) performing blood flow monitoring of a vessel at a first predetermined location to obtain a substantially continuous blood flow signal;(c) providing an updated cardiac measure of said patient by continuously processing said flow signal and said pressure signal. 2. A method according to claim 1 wherein said updated cardiac measure is a measure of Cardiac Power Output (CPO). 3. A method according to claim 2 wherein said blood pressure signal is processed to provide a measure of Mean Arterial Pressure (MAP), said blood flow signal is processed to provide a measure of Cardiac Output (CO), and said Cardiac Power Output (CPO) is substantially calculated using the formula CPOWatts=(MAPmmHg−CVPmmHg)·CO1.1 min(kg·m2s−3). 4. A method according to claim 1 wherein said updated cardiac measure is a measure of Stroke Work (SW). 5. A method according to claim 4 wherein said blood pressure signal is processed to provide a measure of Mean Arterial Pressure (MAP), said blood flow signal is processed to provide a measure of Stroke Volume (SV), and said Cardiac Work (CW) is substantially calculated using the formulaSWmJ=60450(MAPmmHg-CVPmmHg)·SVml 6. A method according to claim 1 wherein said updated cardiac measure is a measure of Cardiac Work (CW). 7. A method according to claim 1 wherein said updated cardiac measure is a substantially continuous measure of blood flowing through said patient calculated by utilising said blood flow signal to provide an updated cardiac output measurement and calibrating said substantially continuous blood pressure signal. 8. A method according to claim 7, comprises the steps of: (a) calculating a transfer function for transforming said substantially continuous blood pressure signal to produce said substantially continuous measure of blood flowing; and(b) producing said substantially continuous measure of blood flowing by applying said transfer function to said substantially continuous blood flow signal. 9. A method as claimed in claim 8, wherein said transfer function is a scalar value. 10. A method according to claim 1, wherein said flow signal is provided by a continuous wave Doppler ultrasonic transducer. 11. A method according to claim 10 wherein said flow signal is processed to provide beat-to-beat measures of Cardiac Output (CO) or Stroke Volume (SV) 12. A method as claimed in claim 1 wherein said pressure signal is provided by a blood pressure monitor. 13. A method as claimed in claim 1 wherein said cardiac measure is updated repeatedly and automatically. 14. A system for monitoring blood flows through a patient, the system comprising: a transducer element for providing a blood flow signal;a blood pressure monitor for providing a blood pressure signal;... | USCOM LIMITED | A61B005/021
US20100317945 | A1 | 2010-12-16 | CUFF FOR DETERMINING A PHYSIOLOGICAL PARAMETER | 2007-07-20 | The invention relates to a cuff (20) for determining a physiological parameter, said cuff comprising a photoplethysmograph arranged with an emitter (6) for emitting a radiation in a direction of a tissue to be investigated, a detector (5) for detecting the radiation from the tissue and an inflatable bladder (8) for transferring pressure to the tissue, said inflatable bladder comprising a back-layer (8″) and a top-layer (8′), wherein the top-layer is conceived to be brought into contact with the tissue, the top-layer being substantially more flexible than the back-layer. The invention further relates to a measurement system. | 1. A cuff for determining a physiological parameter, comprising: a photoplethysmograph arranged with an emitter for emitting a radiation in a direction of a tissue to be investigated;a detector for detecting the radiation from the tissue; andan inflatable bladder for transferring pressure to the tissue, said inflatable bladder comprising a back-layer and a top-layer,wherein the top-layer is conceived to be brought into contact with the tissue, and wherein the back-layer comprises a flexible substantially non-elastic material which does not undergo deformation a suitable internal cuff pressure and the top layer comprises a flexible substantially elastic material which undergoes deformation on under application of the suitable internal cuff pressure. 2. (canceled) 3. A cuff according to claim 1, wherein the top-layer is arranged to provide an electrical insulation between the tissue and electrical components of the cuff. 4. A cuff according to claim 1, wherein the top-layer is arranged to protect electrical components of the cuff from contamination. 5. A cuff according to claim 1 wherein the top-layer is at least partially manufactured from a biocompatible material. 6. A cuff according to claim 5, wherein the top-layer extends further than an area of the inflatable bladder conceived to come into contact with the tissue in use. 7. A cuff according to clam 1, wherein the back-layer is provided with cut-away areas for transmitting the radiation from the emitter to a tissue or for transmitting a further radiation from the tissue to the detector. 8. (canceled) 9. A cuff according to claim 1, wherein the height of the optical components (LED and photo diode) is adjustable. 10. (canceled) 11. (canceled) 12. (canceled) 13. (canceled) 14. A cuff according to claim 1, wherein the emitter is arranged for emitting a focused beam of radiation. 15. A cuff according to claim 1, wherein the emitter comprises a built-in opaque shield. 16. A cuff according to claim 1, wherein the emitter is arranged to emit radiation with a wavelength in a range of 660 to 1000 nm. 17. A cuff according to claim 1, wherein the detector comprises a photodiode arranged to be sensitive for a radiation with a wavelength matching at least part of a spectrum of the emitter. 18. A cuff according to claim 1, further comprising: an air tube communicating with the inflatable bladder;means for preventing blockage of an airflow in the air tube. 19. A cuff according to claim 18, wherein the air tube is cut at an angle less than 90 degrees and is mounted in the inflatable bladder with the longest projection towards the top-layer. 20. (canceled) 21. A measurement system comprising a cuff according to claim 1 and a data processing unit. 22. A measurement system according to claim 21, wherein the data processing unit is arranged to: compute a compliance or impedance in dependence of at least one measurement of arterial pressure data and patient data using a non-linear model;use sa... | - | A61B005/1455
US07853304 | B2 | 2010-12-14 | Method and device for reconstructing two-dimensional sectional images | 2005-05-13 | The invention relates to a method and device for reconstructing two-dimensional sectional images from a three- or four-dimensional image data set of an object, for example the left ventricle. The method comprises the steps of calculating an anatomical long axis through the object, such that the length and/or shape of the anatomical long axis is adapted to the length and/or shape of the object, and defining a series of sectional planes perpendicular or parallel to the anatomical long axis, and calculating and displaying sectional images corresponding to the sectional planes. | 1. A method for reconstructing two-dimensional sectional images from a four-dimensional dynamic image data set of an object having an interior space surrounded by walls, the four-dimensional image data set being acquired by a medical imaging method and containing several three-dimensional data sets each representing a certain, different time frame within an object cycle, the method comprising: (a) defining a respective anatomical long axis within the image data set through the object such that a length and shape of the respective anatomical long axis is adapted to a length and shape of the object;(b) defining a series of intersection points within the data set distributed along the anatomical long axis;(c) defining a series of sectional planes within the data set, wherein each sectional plane intersects an intersection point on the respective anatomical long axis, wherein the respective anatomical long axis is not contained in said sectional planes with the exception of the intersection point, and wherein the sectional plane and a tangent of the anatomical long axis at the intersection point form approximately a right angle with each other; and(d) calculating sectional images corresponding to the sectional planes from the four-dimensional image data set via a reconstruction device, wherein steps a) to d) are automatically repeated for each different three-dimensional data set within a single object cycle, thereby defining a respective anatomical long axis for each time frame, a length and/or shape of each respective anatomical long axis being variable over time wherein a relative distribution of the series of intersection points along each respective anatomical long axis is kept constant automatically, over the different time frames, such that absolute distances between the series of intersection points vary over time in proportion with the length of each anatomical long axis. 2. The method according to claim 1, wherein the intersection points are either spaced equidistantly along each anatomical long axis, or a position of each intersection point of a respective anatomical long axis in a single time frame are defined by a user and each successive anatomical long axis along with associated intersection points are automatically defined such that each respective anatomical long axis varies from each other respective anatomical long axis over time. 3. The method according to claim 1, wherein the sectional planes are bent so as to intersect the walls of the object at a right angle. 4. The method according to claim 1, for reconstructing two-dimensional images from a four-dimensional dynamic image data set of an object having an interior space surrounded by walls the image data set being acquired by a medical imaging method and containing several three-dimensional data sets each representing a certain time frame, the method further comprising the following steps: (a) defining an anatomical long axis through the object such that the length and/or sha... | Tomtec Imaging Systems GmbH | A61B005/05 A61B008/00 G06K009/36
US07837668 | B2 | 2010-11-23 | Needle assembly for use in delivering precise dosages of proteinaceous pharmaceutical compositions and methods for use of same | 2006-10-10 | The needle assembly of the invention is a quad-partite system for use with an arc-center stereotactic headframe that allows a clinician to deliver recombinant expression vectors through the needle assembly to targeted sites at exact depths in the brain. Methods for use of the needle assembly with a tube insert for the needle cannula that does not bind proteins, as in protein viral capsids, permit delivery of precise volumes of pharmaceutical compositions containing viral recombinant expression vectors for gene therapy. | 1. A method for delivering a precise volume of a proteinaceous pharmaceutical composition comprising an adeno-associated virus (AAV) or lentivirus recombinant expression vector having a polynucleotide-free protein capsid to a body tissue, the method comprising: (a) providing a drug delivery vessel, wherein substantially all portions thereof intended for contact with the pharmaceutical composition consist essentially of a biocompatible polymer that is essentially incapable of binding protein;(b) providing a known quantity of the pharmaceutical composition to the drug delivery vessel, wherein the protein capsid is capable of binding by adsorption to, but the vector is resistant to inactivation by, stainless steel; and,(c) causing the pharmaceutical composition to be expelled from the drug delivery vessel into the body tissue;wherein essentially all of the known quantity of the pharmaceutical composition provided is expelled into the body tissue, andwherein further, the drug delivery vessel is a needle assembly comprising:element (i), which is a stylet having a blunt distal end and a proximal end;element (ii), which is a needle guide having a rounded distal end, a proximal end, and a bore extending therethrough from said proximal end to said distal end, wherein said needle guide bore is slightly larger in inner diameter than the outer diameter of the stylet, so the latter is slidably insertable through the former;element (iii), which is a guide tube having a rounded distal end, a proximal end, and a bore extending therethrough from said proximal end to said distal end, wherein said guide tube bore is slightly larger in inner diameter than the outer diameter of the needle guide, so the latter is slidably insertable through the former, and wherein further the proximal end of the guide tube is fitted with means for attachment to a frame;element (iv), which is a needle cannula having a blunt distal end, a proximal end, and a bore extending therethrough from said proximal end to said distal end, wherein further said needle cannula is slightly smaller in outer diameter than the inner diameter of the needle guide, so the former is slidably insertable through the latter;wherein elements (i), (ii) and (iv) are each fitted with a stop at their proximal ends to prevent forward progress thereof through the bore of another element, and wherein further the length of elements (i), (ii) and (iv) of the needle assembly from their stops to their distal ends is such that the distal ends of element (i) or element (iv) protrude to a predetermined length from the distal end of element (ii) when inserted therethrough, and the distal end of element (ii) protrudes to a pre-determined length from the distal end of element (iii) when inserted therethrough. 2. The method according to claim 1, wherein the recombinant expression vector encodes a transgene. 3. The method according to claim 2, wherein the transgene expresses a neurotrophic factor. 4. The method according to cla... | Ceregene, Inc. | A61M031/00 A61M005/178 A61B005/00
US20100332254 | A1 | 2010-12-30 | IN-VITRO DEVICE SUPPORT FOR X-RAY BASED KIDNEY FUNCTION TEST | 2009-06-26 | A system and method of diagnosing kidney syndromes using an intravenous pyelogram (IVP) is described. The system includes an X-ray imaging device and an in-vitro diagnostic (IVD) device. The IVD device is used to identify risk factors such as sensitivity to contrast agents, or to indicate or rule out other syndromes which may have symptoms similar to kidney malfunction. The level of blood components such as creatinine may be used, in conjunction with other medical data and IVD tests, to determine a risk score for the patient. The risk score may be used to recommend modifications to a baseline procedure such as change in the volume or type of contrast agent, the number of X-rays, or the performance of post-procedure dialysis. | 1. A method of performing a diagnostic procedure on a patient, the method comprising: selecting a diagnostic procedure;retrieving existing patient medical records;determining whether appropriate in-vitro diagnostic tests (IVD) have been performed;performing needed IVD tests;determining a risk score for the patient; andselecting procedure steps and parameters based on the patient risk score. 2. The method of claim 1, wherein the diagnostic procedure is an intravenous pyelogram (IVP). 3. The method of claim 2, wherein a factor in determining the risk score is blood creatinine level as measured by an IVD. 4. The method of claim 1, wherein the needed IVD tests include at least one of an IVD test for diagnosis of: kidney stones, an enlarged prostate, or tumor in the kidney, ureters, or urinary bladder. 5. The method of claim 4, wherein the needed IVD tests include a test for iodine sensitivity. 6. The method of claim 1, wherein the risk score is used to modify at least one of a volume of contrast agent administered, or a period between X-ray exposures. 7. The method of claim 1, wherein when the risk score is very high, a dialysis procedure is performed on the patient after completion of the diagnostic procedure. 8. A computer program product, stored on a computer readable medium, comprising: instructions for configuring a computer to: request medical data of a patient from a data base;receive the medical data from a data base through an interface to a network;accept an operator input identifying the diagnostic procedure to be planned;determine a set of IVD tests needed for the procedure, and whether all of the needed tests have are present in the medical data;request any needed IVD tests that are not present in the medical data and accept the results of the requested IVD tests;determine a risk score to the patient for the diagnostic procedure; andrecommend changes to the diagnostic procedure based on the risk score. 9. The computer program product of claim 8, wherein the risk score includes at least one of the following factors: a probability that a syndrome causing patient symptoms has been diagnosed by the performed IVD tests, and a sensitivity of the patient to the contrast agent as determined from one of the received medical data or the performed IVD tests. 10. The computer program product of claim 9, wherein the diagnostic procedure is an intravenous pyelogram (IVP). 11. The computer program product of claim 10, wherein the risk score includes the blood creatinine level as a factor. 12. The computer program product of claim 11, wherein when the creatinine level is very high a post-IVD dialysis procedure is recommended. 13. The computer program product of claim 11, wherein a volume of contrast agent or a number of X-ray exposures is adjusted based on the risk score. 14. A system for performing an intravenous pyelogram (IVP), comprising: an X-ray device; andan in-vitro diagnostic device (IVD),wherein the IVD is used to perform a test f... | - | G06Q050/00 A61B005/00 G06N005/02 A61B006/00
US20100317932 | A1 | 2010-12-16 | RESPIRATORY FUNCTION MEASURING APPARATUS | 2009-12-15 | A respiratory function measuring apparatus includes: a first sensor configured to detect an invasive blood pressure; a second sensor configured to measure frequency of at least one of a heart beat and a respiration; and a controller configured to extract a respiratory function signal from the invasive blood pressure detected by the first sensor, by using at least one of the frequency measured by the second sensor and a harmonic of the frequency. | 1. A respiratory function measuring apparatus comprising: a first sensor configured to detect an invasive blood pressure;a second sensor configured to measure frequency of at least one of a heart beat and a respiration; anda controller configured to extract a respiratory function signal from the invasive blood pressure detected by the first sensor, by using at least one of the frequency measured by the second sensor and a harmonic of the frequency. 2. The respiratory function measuring apparatus according to claim 1, wherein the controller presumes the extracted respiratory function signal as an intrathoracic pressure. 3. The respiratory function measuring apparatus according to claim 1, wherein when the second sensor measures the frequency of the heart beat, the controller selectively removes components corresponding to the frequency of the heart beat and a harmonic of the frequency from a signal corresponding to the invasive blood pressure. 4. The respiratory function measuring apparatus according to claim 1, wherein when the second sensor measures the frequency of the heart beat, the controller extracts a component corresponding to a frequency which is lower than the frequency of the heart beat from a signal corresponding to the invasive blood pressure. 5. The respiratory function measuring apparatus according to claim 1, wherein when the second sensor measures the frequency of the respiration, the controller extracts components corresponding to the frequency of the respiration and a harmonic of the frequency from a signal corresponding to the invasive blood pressure. 6. The respiratory function measuring apparatus according to claim 1, wherein the frequency of the heart beat is measured from an electrocardiogram, a plethysmogram, or an arterial blood pressure. 7. The respiratory function measuring apparatus according to claim 1, wherein the frequency of the respiration is measured from a capnometry, a respiratory flow, an airway pressure, a transthoracic electrical impedance, or a respiratory temperature. 8. The respiratory function measuring apparatus according to claim 1, wherein the invasive blood pressure is a central venous pressure (CVP) or a peripheral venous pressure. 9. The respiratory function measuring apparatus according to claim 1, further comprising: a respiration variation detector, whereinwhen the second sensor measures the frequency of the respiration, the second sensor presumes an end-tidal, andthe respiration variation detector is configured to obtain a degree of a respiratory signal with respect to the end-tidal. 10. The respiratory function measuring apparatus according to claim 9, wherein the end-tidal is obtained from a flat portion of the invasive blood pressure, a capnometry, a respiratory flow, an airway pressure, a transthoracic electrical impedance, or a respiratory temperature. 11. The respiratory function measuring apparatus according to claim 9, further comprising: a respiration determiner configured t... | Nihon Kohden Corporation | A61B005/0205
US20100326842 | A1 | 2010-12-30 | Extruded Electrode Structures and Methods of Using Same | 2009-06-30 | The present disclosure provides electrode structures and integrated electrode structures having one or more conductive materials coextruded with one or more dielectric materials. The disclosed electrode structures can be configured for use as analyte sensors. Also provided, are methods of making and using the electrode structures and integrated electrode structures described herein. | 1. An electrode structure, comprising: a first conductive material, a second conductive material, and a dielectric material,wherein the first conductive material, the second conductive material, and the dielectric material are coextruded to provide an electrode structure having the first conductive material and the second conductive material electrically isolated by the dielectric material. 2. The electrode structure of claim 1, wherein the electrode structure is configured for in vitro use. 3. The electrode structure of claim 2, wherein the electrode structure is configured to detect an analyte in a sample volume of less than 1 μL. 4. The electrode structure of claim 2, wherein the electrode structure is configured to detect an analyte in a sample volume of less than 500 nL. 5. The electrode structure of claim 2, wherein the electrode structure is configured to detect an analyte in a sample volume of less than 100 nL. 6. The electrode structure of claim 1, wherein the electrode structure is configured for in vivo use. 7. The electrode structure of claim 6, wherein at least a portion of the electrode structure is adapted to be implanted in a subject. 8. The electrode structure of claim 7, wherein at least a portion of the electrode structure is adapted to be subcutaneously positioned in the subject. 9. The electrode structure of claim 1, wherein the first conductive material and the second conductive material comprise conductive polymers. 10. The electrode structure of claim 1, wherein the first conductive material comprises a first color, the second conductive material comprises a second color and the first color and the second color are different. 11. The electrode structure of claim 1, wherein the dielectric material is a biocompatible material. 12. The electrode structure of claim 1, wherein the first conductive material comprises a working electrode and the second conductive material comprises a reference/counter electrode, the working electrode comprising an analyte responsive enzyme disposed on, in or proximate thereto. 13. The electrode structure of claim 12, wherein the first conductive material comprises carbon. 14. The electrode structure of claim 12, wherein the second conductive material comprises Ag/AgCl. 15. The electrode structure of claim 12, wherein the second conductive material comprises platinum and Ag/AgCl. 16. The electrode structure of claim 12, further comprising a membrane disposed over the working electrode. 17. The electrode structure of claim 12, wherein the analyte-responsive enzyme is a glucose-responsive enzyme. 18. The electrode structure of claim 12, wherein the analyte-responsive enzyme is a ketone-responsive enzyme. 19. The electrode structure of claim 12, wherein the working electrode comprises a redox mediator disposed on, in or proximate thereto. 20. The electrode structure of claim 19, wherein the redox mediator comprises a ruthenium-containing complex or an osmium-containing complex. 21... | Abbott Diabetes Care Inc. | G01N027/327 A61B005/1468
US20100331652 | A1 | 2010-12-30 | MODULAR DIABETES MANAGEMENT SYSTEMS | 2009-06-29 | In one embodiment, a modular diabetes management system includes a portable, stand-alone blood glucose meter and a portable docking device that includes an internal receptacle sized and structured to receive and house the blood glucose meter. The docking device is generally operable to interface with the blood glucose meter and perform various diabetes management functions. For example, the docking device may be operable to analyze blood glucose measurement data stored on the blood glucose meter, configure the blood glucose meter, and/or interact with an insulin delivery device, just to name a few possibilities. Other embodiments include unique methods, systems, kits, assemblies, equipment, and/or apparatus which are related to the management of diabetes. | 1. A system, comprising: a portable, stand-alone blood glucose meter including a display and a first connection element;a portable, handheld docking device including a display operable to provide information related to functions performed by said blood glucose meter and a housing defining an external profile of said docking device and an internal receptacle, said internal receptacle including a second connection element engageable with said first connection element; andwherein said blood glucose meter is positionable in said internal receptacle and a communication interface is defined between said blood glucose meter and said docking device when said blood glucose meter is positioned in said internal receptacle and said first connection element is engaged with said second connection element. 2. The system of claim 1, wherein said blood glucose meter includes a meter housing which substantially corresponds in size and shape to said internal receptacle, said meter housing being substantially surrounded by said internal receptacle and positioned within said external profile of said docking device when said blood glucose meter is positioned in said internal receptacle. 3. The system of claim 1, wherein said housing of said docking device includes a sidewall extending around said housing and between a first surface and a second surface positioned opposite of said first surface, and wherein said internal receptacle is positioned between said first and second surfaces. 4. The system of claim 3, wherein said internal receptacle includes a receiving portion extending through said sidewall of said housing. 5. The system of claim 4, wherein said second surface includes an opening in communication with a portion of said internal receptacle adjacent to said receiving portion. 6. The system of claim 4, wherein said second connection element is positioned in said internal receptacle opposite of said receiving portion. 7. The system of claim 1, wherein one of said first connection element and said second connection element defines a plug and the other of said first connection element and said second connection element defines a port structured to receive said plug. 8. The system of claim 1, wherein said blood glucose meter includes a meter housing defining a test element port for receiving a test element. 9. The system of claim 8, wherein said test element comprises a test strip. 10. The system of claim 8, wherein said first connection element is positioned on said meter housing opposite of said test element port. 11. The system of claim 8, wherein said blood glucose meter includes a memory configured to store blood glucose measurement data. 12. The system of claim 11, wherein said blood glucose measurement data is transferred from said blood glucose meter to said docking device when said blood glucose meter is positioned in said internal receptacle and said first connection element is engaged with said second connection element. 13. The system of cl... | Roche Diagnostics Operations, Inc. | A61B005/145
US20100298694 | A1 | 2010-11-25 | STROKE RISK ASSESSMENT | 2010-05-20 | Certain embodiments provide a method of assessing a risk of thromboembolic stroke in a patient. The method comprises acquiring image data and determining an indicator of a degree of fibrosis of a patient's left atrium (LA) based on the image data. The method also comprises estimating a level of risk of thromboembolic stroke of the patient based on the indicator. | 1. A method of assessing a risk of thromboembolic stroke in a patient, the method comprising: determining, from left atrium (LA) tissue image data of a patient, a level of a parameter that is positively proportional or negatively proportional to an amount of abnormal tissue in a wall of the LA of the patient; andoutputting, to an output device, a machine-readable indicator of a comparison between (i) the determined level and (ii) a first threshold level of the parameter, the first threshold level derived from LA tissue image data of at least one person, other than the patient, who has a history of atrial fibrillation (AF);wherein, when the level of the parameter is positively proportional to the amount of abnormal tissue, levels of the parameter equal to or greater than the first threshold level are indicative of a significantly increased risk of thromboembolic stroke relative to levels of the parameter less than the first threshold level;wherein, when the level of the parameter is negatively proportional to the amount of abnormal tissue, levels of the parameter equal to or less than the first threshold level are indicative of a significantly increased risk of thromboembolic stroke relative to levels of the parameter greater than the first threshold level. 2. The method of claim 1, wherein the abnormal tissue present in the LA wall comprises at least one of a fibrotic tissue, a necrotic tissue, a tissue comprising apoptotic cells, a scar tissue, a tissue having impaired electrical conduction, an aberrantly electrically remodeled tissue, and a structurally remodeled tissue. 3. The method of claim 1, wherein the level of the parameter comprises a proportion of fibrotic LA wall tissue relative to a total amount of the LA wall tissue. 4. The method of claim 3, wherein the proportion of fibrotic LA wall tissue relative to a total amount of LA wall tissue comprises a ratio of (i) an imaged volume of fibrotic LA wall tissue to (ii) an imaged total volume of LA wall tissue. 5. The method of claim 1, wherein the tissue image data of the patient is acquired before the patient receives treatment with an AF treatment modality. 6. The method of claim 5, wherein the determining occurs before the patient receives treatment with the AF treatment modality. 7. The method of claim 1, further comprising acquiring the image data by detecting a signal of an agent substantially localized at the abnormal LA tissue. 8. The method of claim 7, wherein the agent comprises a magnetic resonance contrast agent, and wherein the detecting comprises performing magnetic resonance imaging. 9. The method of claim 8, wherein the magnetic resonance imaging comprises delayed enhancement magnetic resonance imaging (DE-MRI). 10. The method of claim 7, further comprising localizing the agent to the abnormal LA tissue by exposing the LA tissue to an antibody or antibody component, coupled to the agent, that binds an epitope present in the abnormal LA tissue and not present in norm... | - | A61B005/055 A61B005/00 A61B006/00
US07833166 | B2 | 2010-11-16 | Breath alcohol-measuring device with an illuminated mouthpiece | 2007-10-24 | A breath alcohol-measuring device is provided, which includes a mouthpiece which can be inserted or attached, with a blow-in opening and with a blow-out opening and with a hole for breathing gas sampling in the breath alcohol-measuring device. The breath alcohol-measuring device (2) is equipped with a radiation source (3) and with a light guide (4) receiving the radiation of the radiation source (3). The light guide (4) is arranged extending from the housing of the breath alcohol-measuring device (2) up to the blow-out opening of the mouthpiece (1) such that the light guide (4) acts as an illuminating means for the mouthpiece (1). | 1. A breath alcohol-measuring device comprising: a breath alcohol-measuring device housing;an alcohol sensor connected to said housing;a mouthpiece with a blow-in opening and with a blow-out opening and with a sample hole for a breathing gas sampling for said alcohol sensor, said mouthpiece being detachably connected to said housing;a radiation source; anda light guide receiving radiation of said radiation source, said light guide being arranged extending from said breath alcohol-measuring device housing up to said blow-out opening of said mouthpiece such that said light guide provides an illuminating means for illuminating an interior of said mouthpiece. 2. A breath alcohol-measuring device in accordance with claim 1, wherein said light guide forms a stop for positioning said mouthpiece, said mouthpiece being detachably connected to said housing which can be inserted or attached. 3. A breath alcohol-measuring device in accordance with claim 1, wherein said radiation source is a light emitting diode (LED) and said light guide is a glass fiber element. 4. A breath alcohol-measuring device in accordance with claim 1, wherein said light guide is an injection-molded element. 5. A breath alcohol-measuring device in accordance with claim 1, wherein: said mouthpiece is at least partially transparent to light or translucent; andsaid radiation source emits radiation in different color ranges illuminating the interior of the mouthpiece in any selected one of the color ranges and for being viewed from an exterior of the mouthpiece. 6. A breath alcohol-measuring device in accordance with claim 5, further comprising a control unit, wherein the different color ranges correspond to different operating or functional states of said breath alcohol-measuring device, said control unit controlling said radiation source to emit predetermined colors of said different color ranges based on a current different operating or functional state of the breath alcohol-measuring device. 7. A breath alcohol-measuring device in accordance with claim 1, wherein said alcohol sensor is an electrochemical sensor. 8. A breath alcohol-measuring device in accordance with claim 1, wherein said radiation source and said light guide are designed as an optical component in one assembly unit. 9. A breath alcohol-measuring device comprising: a breath alcohol-measuring device housing;an alcohol sensor connected to said housing;an at least partially transparent to light or translucent mouthpiece detachably mounted on said housing, said mouthpiece having a blow-in opening, a blow-out opening and a sample hole with an operative connection to said alcohol sensor for providing a breathing gas sampling for said alcohol sensor;a light source; anda light guide receiving radiation of said light source, said light guide being arranged extending from said breath alcohol-measuring device housing up to an operative connection with said mouthpiece blow-out opening for directing light into said mouthp... | Dräger Safety AG & Co. KGaA | A61B005/06
US07833169 | B2 | 2010-11-16 | Device and method for the collection of a urine sample | 2008-01-17 | A device and method for collection of a urine sample is disclosed which comprises a channel, an opening in the channel, and a collector associated with the opening. The channel is formed so that urine can flow through it from a first end to a second end thereof. The opening is provided in the channel at a point between the first and the second ends thereof. The collector is associated with the opening in the channel to be able to receive urine therein. In one embodiment, a member can cover the opening to permit urine to pass through it after contact with urine. Specifically, the member may dissolve or otherwise permit urine to pass through the opening after a period of time to test a urine sample which is not from the initial flow of urine to avoid bacteria contamination. | 1. A device for collecting a urine sample, comprising: an axially extending body having a channel defining a continuous flow path extending from a first end to a second end thereof and through which urine can flow;a generally radial opening in the axially extending body extending from a point in communication with the continuous flow path to a point externally of the axially extending body at a location intermediate the first and the second ends thereof;a collector associated with the generally radial opening externally of the axially extending body for receiving urine; anda member covering the generally radial opening in the axially extending body which permits urine to pass from the continuous flow path defined by the channel, through the generally radial opening in the axially extending body, and into the collector only after the member has been in contact with urine that has passed through the continuous flow path defined by the channel for a period of time;whereby, the collector is only able to receive a urine sample which is not from an initial flow of urine at a location externally of the axially extending body. 2. The device of claim 1 wherein the channel defining the continuous flow path through which urine can flow comprises a lumen associated with a urinary catheter. 3. The device of claim 2 wherein the urinary catheter includes a funnel and the lumen is located within the funnel. 4. The device of claim 2 wherein the urinary catheter includes a catheter tube and the lumen is located within the catheter tube. 5. The device of claim 1 wherein the channel defining the continuous flow path through which urine can flow comprises a lumen of a test device for connection to a urinary catheter. 6. The device of claim 1 including a urine collection bag associated with the second end of the continuous flow path through which urine can flow. 7. The device of claim 1 wherein the collector associated with the generally radial opening externally of the axially extending body is a urine collection cup surrounding the opening. 8. The device of claim 1 wherein the collector associated with the generally radial opening externally of the axially extending body comprises a urine test chamber. 9. The device of claim 1 wherein the member covering the generally radial opening in the axially extending body is a soluble membrane formed of PVOH. 10. The device of claim 1 wherein the member covering the generally radial opening in the axially extending body is a non-soluble membrane fixed in place by an adhesive element which reduces in adhesive effectiveness when in contact with liquid. 11. The device of claim 1 wherein the member covering the generally radial opening in the axially extending body is a non-soluble membrane fixed in place by an adhesive element which reduces in adhesive effectiveness when the adhesive element temperature is raised. 12. The device of claim 1 wherein the member covering the generally radial opening in the axially extendi... | Hollister Incorporated | A61B005/00 B65D081/00 A61M001/00
US07844319 | B2 | 2010-11-30 | Systems and methods for magnetic-resonance-guided interventional procedures | 2002-04-15 | Herein is disclosed a probe, including a first electrode disposed at least partially on the probe surface, a second electrode disposed at least partially on the probe surface, a first conductor electrically coupled to the first electrode, a second conductor electrically coupled to the second electrode, and a reactive element electrically coupling the first conductor and the second conductor. | 1. A probe, comprising: a matching circuit;a first electrode residing on a distal end portion of the probe;a second electrode residing spaced apart from the first electrode on the distal end portion of the probe;a first conductor electrically coupled to the first electrode and extending from the matching circuit;a second conductor electrically coupled to the second electrode and extending from the matching circuit; anda reactive element disposed between the matching circuit and a distal end of the probe, the reactive element electrically coupling the first conductor and the second conductor, wherein the first and second conductors and the reactive element form a loop antenna for a high frequency signal received by the probe. 2. The probe of claim 1, wherein the reactive element is positioned to conduct a high frequency signal between the first conductor and the second conductor. 3. The probe of claim 2, wherein the reactive element is selected to conduct the high frequency signal, wherein the high frequency signal has a frequency higher than about 10 MegaHertz (MHz). 4. The probe of claim 1, wherein the reactive element is positioned to conduct a signal including magnetic resonance imaging frequency energy between the first conductor and the second conductor. 5. The probe of claim 1, wherein the first conductor, second conductor, and reactive element are configured and arranged such that at least one of the first conductor and the second conductor conducts a low frequency signal to or from at least one of the first electrode and the second electrode. 6. The probe of claim 1, wherein the first conductor, second conductor, and reactive element are configured and arranged such that at least one of the first conductor and the second conductor conducts a low frequency signal to or from at least one of the first electrode and the second electrode, wherein the low frequency signal has a frequency of up to about 500 kiloHertz (kHz). 7. The probe of claim 1, wherein the first conductor, second conductor, and reactive element are configured and arranged such that at least one of the first conductor and the second conductor conducts a low frequency signal to or from at least one of the first electrode and the second electrode, wherein the low frequency signal is in the range from about 100 Hertz (Hz) to about 1 kHz. 8. The probe of claim 1, wherein the first conductor, second conductor, and reactive element are configured and arranged such that at least one of the first conductor and the second conductor conducts a low frequency signal to or from at least one of the first electrode and the second electrode, wherein the low frequency signal is about 100 kHz. 9. The probe of claim 1, wherein the reactive element is configured and arranged to conduct a signal including ablation frequency energy to at least one of the first electrode and the second electrode. 10. The probe of claim 1, wherein the reactive element is configured and arranged to conduct a... | - | A61B005/055
US07854703 | B2 | 2010-12-21 | Peripheral neuropathy detection | 2007-09-25 | Briefly, in accordance with one or more embodiments, a peripheral neuropathy monitor may be incorporated into a scale form factor or other device for monitoring a trend towards peripheral neuropathy by a monitored user or patient. The peripheral neuropathy monitor may include a heating element, a cooling element, and/or a vibrating element, and may be capable of performing the sensitivity of the user to hot, cold, and/or vibrations, and monitor such trends over time to facilitate detecting the onset of peripheral neuropathy. The peripheral neuropathy monitor may be capable of communicating with a remote device via a wired or wireless network for operating the monitor, and for storing and/or analyzing the test results. | 1. An apparatus, comprising: a housing;a vibrating element disposed in the housing, the vibrating element being capable of producing a vibration at a selected frequency and amplitude;a temperature element disposed in the housing and being capable of producing thermal output at a selected temperature;a scale element disposed in the housing, the scale element comprising a weight scale form factor wherein the user stands on the housing during operation of the apparatus by the user, and the scale element being capable of detecting a non-zero weight of the at least part of the user when the user stands on the housing during operation of the apparatus,a processor disposed in the housing, the processor being capable of controlling the vibrating element and the temperature element, individually or in combination, to run one or more peripheral neuropathy tests on an extremity of a user that is coupled with the vibrating element or the temperature element, or a combination thereof, in response to a non-zero weight being detected by the scale element;a user actuator to provide user feedback to the processor in response to the one or more neuropathy tests on the extremity of the user; anda memory disposed in the housing, the memory being coupled to the processor to store results of the one or more neuropathy tests and the user feedback received via the user actuator;wherein the processor is further capable of processing the one or more neuropathy tests for detecting a trend towards peripheral neuropathy of the extremity of the user, and in response to a detected trend towards peripheral neuropathy of the extremity of the user, the processor further capable of adaptively learning and selecting one or more neuropathy tests from a group of predetermined stored neuropathy tests and controlling the vibrating element and the temperature element, individually or in combination, to run the one or more selected peripheral neuropathy tests on the extremity of the user, the one or more selected neuropathy tests being selected based at least in part on the results of one or more previous neuropathy tests on the extremity of the user and the user feedback stored in the memory to further identify progression of the peripheral neuropathy of the extremity of the user over time. 2. An apparatus as claimed in claim 1, said temperature element comprising a heating element or a cooling element, or a combination thereof. 3. An apparatus as claimed in claim 2, further comprising a speaker and said user actuator comprising a microphone, said microphone to provide user feedback to the processor in response to the one or more neuropathy tests and the speaker to provide audio information to the user. 4. An apparatus as claimed in claim 3, further comprising a wireless network adapter coupled to the processor, said wireless network adapter being capable of communicating with a remote device, wherein the remote device is capable of storing test results of the one or more neuropathy ... | Intel Corporation | A61B005/00
US07850617 | B2 | 2010-12-14 | Patient monitoring apparatus and method for determining volume responsiveness of a monitored patient | 2007-04-05 | From the global end-diastolic volume GEDV and the global ejection fraction GEF the patient monitor (4) determines a corrected global end-diastolic volume cGEDV according to cGEDV=GEDV/ƒ(GEF) which is used as a novel parameter for volume responsiveness of the patient (3). In the above formula, f(GEF) is a correction function depending on global ejection fraction GEF. Further, from the right ventricular end-diastolic volume RVEDV and the right ventricular ejection fraction RVEF the patient monitor (4) determines a corrected right ventricular end-diastolic volume cRVEDV according to cRVEDV=RVEDV1f(RVEF) which is used as another novel parameter for volume responsiveness of the patient (3). In the above formula, f(RVEF) is a correction function depending on right ventricular ejection fraction RVEF. | 1. A patient monitoring apparatus for determining at least one of volume responsiveness and cardiac preload of a monitored patient, said apparatus comprising a thermodilution measurement assembly for inducing a local temperature change at a first location of the patient's circulatory system and measuring a response signal downstream of said first location andan evaluation unit adapted to determine an end-diastolic volume EDV from said measured response signal wherein said evaluation unit is further adapted to determine, as a parameter for at least one of volume responsiveness and cardiac preload, a corrected end-diastolic volume cEDV as the quotient of said end-diastolic volume EDV and a function ƒ of an ejection fraction EF of said patient according to the formula cEDV=EDV/ƒ(EF). 2. A patient monitoring apparatus according to claim 1, wherein said function ƒ is an exponential function. 3. A patient monitoring apparatus according to claim 2, wherein said function ƒ is of the type ƒ=exp(k1·(k2−EF))wherein k1 is a first constant and k2 is a second constant. 4. A patient monitoring apparatus according to claim 3, wherein said first constant k1 is positive. 5. A patient monitoring apparatus according to claim 4, wherein said first constant k1 is between 2.5 and 3. 6. A patient monitoring apparatus according to claim 3, wherein said second constant k2 represents a normal ejection fraction which has been determined as an average value of the ejection fractions of a plurality of persons. 7. A patient monitoring apparatus according to claim 1, wherein said thermodilution measurement assembly includes injection means (11) for central venous injection of a bolus with a temperature deviating from the patient's blood temperature. 8. A patient monitoring apparatus according to claim 1, wherein said thermodilution measurement assembly includes heating means for locally influencing the patient's blood temperature in a central vein. 9. A patient monitoring apparatus according to claim 1, wherein said thermodilution measurement assembly includes cooling means for locally influencing the patient's blood temperature in a central vein. 10. A patient monitoring apparatus according to claim 1, wherein said thermodilution measurement assembly includes a right heart catheter with sensor means for measuring said response signal in the pulmonary artery of the patient and wherein said end-diastolic volume EDV is a right ventricular end-diastolic volume RVEDV, said ejection fraction EF is a right ventricular ejection fraction RVEF and said evaluation unit is adapted to determine a corrected right ventricular end-diastolic volume cRVEDV according to the formula cRVEDV=RVEDV/ƒ(RVEF). 11. A patient monitoring apparatus according to claim 10, wherein said function ƒ is of the type ƒ=exp(k1·(k2−RVEF))and wherein k1 is a first constant and k2 is a second constant which is between 0 and 1. 12. A patient monitoring apparatus according to claim 1, wherein said thermodilut... | Pulsion Medical Systems AG | A61B005/02
US20100280335 | A1 | 2010-11-04 | PATIENT STATE DETECTION BASED ON SUPERVISED MACHINE LEARNING BASED ALGORITHM | 2010-01-26 | A patient state is detected with at least one classification boundary generated by a supervised machine learning technique, such as a support vector machine. In some examples, the patient state detection is used to at least one of control the delivery of therapy to a patient, to generate a patient notification, to initiate data recording, or to evaluate a patient condition. In addition, an evaluation metric can be determined based on a feature vector, which is determined based on characteristics of a patient parameter signal, and the classification boundary. Example evaluation metrics can be based on a distance between at least one feature vector and the classification boundary and/or a trajectory of a plurality of feature vectors relative to the classification boundary over time. | 1. A method comprising: receiving a signal indicative of a parameter of a patient;applying a first classification algorithm determined based on supervised machine learning to classify a patient state based on the signal; andapplying at least one additional classification algorithm determined based on supervised machine learning to further classify the patient state based on the signal. 2. The method of claim 1, wherein the patient state comprises a posture state. 3. The method of claim 1, wherein the patient state comprises at least one of a seizure state, a movement disorder state, or a mood state. 4. The method of claim 1, wherein the first classification algorithm and the at least one additional classification algorithm classify a severity of the patient state. 5. The method of claim 1, wherein the parameter comprises at least one of patient motion or activity, heart rate, respiratory rate, electrodermal activity, thermal activity or muscle activity. 6. The method of claim 1, wherein applying the first classification algorithm to classify the patient state based on the signal comprises: determining a feature vector based on the signal; anddetermining a first classification of the patient state based on the feature vector and a first classification boundary defined by the first classification algorithm. 7. The method of claim 6, wherein applying at least one additional classification algorithm comprises, after determining the first classification of the patient state, determining a second classification of the patient state based on the feature vector and a second classification boundary defined by a second classification algorithm determined based on supervised machine learning. 8. The method of claim 7, wherein the first classification comprises a determination of whether the signal is indicative of a first posture state or a second posture state of the patient, and the second classification comprises a determination of whether the signal is indicative of a third posture state or a fourth posture state of the patient. 9. The method of claim 8, wherein the first posture state comprises a non-upright posture state, the second and third posture states each comprise an upright posture state, and the fourth posture state comprises an upright and active posture state. 10. The method of claim 6, wherein the first classification comprises a determination of whether the signal is indicative of a seizure state or a non-seizure state of the patient, and the second classification comprises a determination of whether the signal is indicative of a first seizure state comprising a first severity rating or a second seizure state comprising a second severity rating. 11. The method of claim 1, wherein the supervised machine learning comprises at least one of a genetic algorithm or an artificial neural network. 12. The method of claim 11, wherein the artificial neural network comprises at least one of a support vector machine or a Bayesian classifier... | Medtronic, Inc. | A61B005/00 A61B005/11
US08326389 | B2 | 2012-12-04 | System for in vivo biosensing based on the optical response of electronic polymers | 2007-12-05 | A system for continuous in vivo biosensing of specific analyte molecule concentrations based on the dynamic optical properties of electronic polymers is disclosed. The biosensor system includes at least one implant member subcutaneously exposed to the interstitial fluid of the subject, and a reader member at least temporarily positioned over the implant member to probe it with light of specific wavelengths through the skin. The system has many potential applications, including the real-time monitoring of blood glucose levels in diabetics as a method to supplement or replace conventional capillary blood testing. | 1. A system for in vivo biosensing of one or more target analyte molecules in a subject, the system comprising: at least one implant member configured to be completely implanted in the subject's body and to be exposed to the one or more target analyte molecules in the subject's body,the implant member including a reflective substrate having deposited thereon at least one biocompatible polymer comprised of at least one polyaniline or oligomer thereof, and at least one sensitizing agent; the sensitizing agent configured to cause an oxidation-reduction reaction with the polymer in the presence of the target analyte molecule; the biocompatible polymer configured to immobilize the sensitizing agent without covalent bonding and to accept electrons sufficient to alter the biocompatible polymer's optical properties and to provide a detectable signal;at least one reader member separate from the implant member and configured to be at least temporarily positioned external to the subjects' body and adjacent to the implant member,the reader member being configured: i) to emit light having one or more specific wavelengths, and ii) to detect a reflected signal from the implant member; andat least one data system configured to determine a concentration of the target analyte molecule. 2. The biosensor system of claim 1, wherein the reader system is configured to substantially continuously detect the reflected signal from the implant member. 3. The biosensor system of claim 1, wherein the polymer is present as a thin film and/or fiber network material on the reflective substrate. 4. The biosensor system of claim 1, wherein the polymer is an electronic polymer having an optical absorption spectrum in the ultraviolet-visible-near infrared (UV-VIS-NIR) range that is dependent on the polymer's oxidation state. 5. The biosensor system of claim 1, wherein the sensitizing agent comprises one or more enzymes. 6. The biosensor system of claim 1, wherein the sensitizing agent comprises an oxidoreductase and the analyte is the associated substrate. 7. The biosensor system of claim 1, wherein the analyte molecule is glucose, and the sensitizing agent comprises glucose oxidase (GOx). 8. The biosensor system of claim 1, wherein the analyte molecule is cholesterol, and the sensitizing agent comprises cholesterol oxidase. 9. The biosensor system of claim 1, wherein the analyte molecule is ethanol, and the sensitizing agent comprises alcohol oxidase. 10. The biosensor system of claim 1, wherein the polymer incorporates segments of polyaniline that have repeat lengths of at least four aniline rings, or oligomers thereof. 11. The biosensor system of claim 4, wherein the electronic polymer incorporates pendant groups comprised of aniline that has repeat lengths of at least four aniline rings, or oligomers thereof. 12. The biosensor system of claim 1, wherein the polymer incorporates segments of pyridine that have repeat lengths of at least four pyridine rings, or oligomers... | The Ohio State University Research Foundation | A61B005/1455
US08379946 | B2 | 2013-02-19 | Method and control device to operate a magnetic resonance system | 2010-05-05 | In a method of a magnetic resonance system and a method and computer-readable storage medium for the operation thereof to acquire magnetic resonance image data of an examination subject, wherein magnetic resonance system has a number of subsystems and a control device, a number of adjustment measurements to adjust at least one subsystem for making a medical diagnostic data acquisition are implemented through the control device. In these adjustment measurements, an adjustment volume associated with the appertaining adjustment measurement is taken into account that encompasses at least one region of a body containing the examination subject. For this purpose, markings established by the control device within image data of the examination subject and characterizing the spatial occupation (position and orientation) and/or a dimension of the examination subject are determined. Based on these markings the spatial occupation and extent of an adjustment volume for an adjustment measurement are determined automatically by the control device. | 1. A method to operate a magnetic resonance system to acquire magnetic resonance image data of an examination subject, said magnetic resonance system comprising a plurality of subsystems and a control device that operates the subsystems, said method comprising the steps of: prior to operating the subsystems from the control device to acquire diagnostic magnetic resonance data from the examination subject, operating at least one of the subsystems from the control device to implement an adjustment of said at least one of said subsystems for the subsequent diagnostic data acquisition;in the adjustment of said at least one of subsystems, making use of an adjustment volume in said control device that is associated with the adjustment, said adjustment volume comprising at least a region of a body containing the examination subject;through the control device, establishing markings within image data of the examination subject, said markings characterizing at least of a position and orientation of the examination subject, and a dimension of the examination subject; andbased on said markings, automatically determining, in said control device, a spatial occupation and extent of said adjustment volume to implement the adjustment of said at least one of said subsystems. 2. A method as claimed in claim 1 comprising, with said markings, marking anatomically significant features of the examination subject. 3. A method as claimed in claim 2 comprising, with at least some of said markings, marking anatomical landmarks of the examination subject. 4. A method as claimed in claim 1 comprising entering at least some of said markings through a user interface of said control device. 5. A method as claimed in claim 1 comprising, with said markings, marking anatomically significant features of the examination subject and determining said anatomically significant features automatically in said control device. 6. A method as claimed in claim 1 comprising providing said control device with an overview image comprising the examination subject, and making said markings in said overview image. 7. A method as claimed in claim 6 comprising providing at least two overview images of the examination subject to the control device showing respectively different views of slices of the examination subject, and making said markings in said at least two overview images. 8. A method as claimed in claim 1 comprising, in said control device, adapting said adjustment volume to contours of the examination subject. 9. A method as claimed in claim 1 comprising, in said control unit, identifying a standard adjustment volume defined specifically for said examination subject and, in said control device, adapting at least one parameter of said standard adjustment volume to the examination subject based on said markings. 10. A method as claimed in claim 1 comprising, in said control device, determining said at least one of said position and orientation of said examination subject, and said e... | Siemens Aktiengesellschaft Siemens Medical Solutions USA, Inc. | G06K009/00
US08986214 | B2 | 2015-03-24 | System and method for using tissue contact information in an automated mapping of cardiac chambers employing magnetically shaped fields | 2010-02-17 | The invention relates to a method for using tissue contact technology to optimize automated cardiac chamber mapping algorithms to both speed up the mapping process and guarantee the definition of the actual chamber limits. The invention further comprises a method for conveying tissue type information to such automatic mapping algorithms so as to allow them to adapt their point collection density within areas of particular interest. The method is enhanced by the use of a magnetic chamber that employs electromagnetic coils configured as a waveguide that radiate magnetic fields by shaping the necessary flux density axis on and around the catheter distal tip so as to push, pull and rotate the tip on demand and as defined by such automatic mapping algorithms. | 1. A method for automatically mapping the interior surface of a coronary chamber comprising: inserting a catheter with an array of electrodes disposed on its surface into the coronary chamber;generating a shaped magnetic field around the catheter and coronary chamber to selectively move the catheter under closed loop control;receiving a plurality of signals from the electrode array disposed on the catheter;calculating an actual position of the catheter within the coronary chamber from the plurality of received signals;determining a presence of continuous or partial contact between the catheter and the surface of the coronary chamber from the plurality of received signals;generating an automated mapping pattern for the catheter using closed loop control for movement of the catheter to optimize the closed loop automated mapping pattern of the catheter, wherein generating the automated mapping pattern for the catheter further comprises: obtaining a plurality of measurements from the surface of the coronary chamber when continuous contact has been established according to the generated closed loop automated mapping pattern; andobtaining a plurality of measurements from the surface of the coronary chamber when partial contact has been established according to the generated closed loop automated mapping pattern;automatically generating a map of the continuous contact manifold of the coronary chamber based on the obtained continuous contact measurements; andautomatically generating a map of the partial contact zone of the coronary chamber based on the obtained partial contact measurements; andautomatically moving a distal end of the catheter between the generated map of the continuous contact manifold and the generated map of the partial contact zone during generation of the closed loop automated mapping pattern of the catheter, so that travel time and distance through the distal end of the catheter is moved during generation of the closed loop automated mapping pattern is reduced. 2. The method of claim 1 where receiving a plurality of signals from the electrode array disposed on the catheter comprises collecting a plurality of impedance measurements obtained from the coronary chamber and where determining the presence of continuous or partial contact between the catheter and the surface of the coronary chamber comprises determining the presence of continuous or partial contact based on the impedance measurements obtained by the electrode array. 3. The method of claim 1 where obtaining a plurality of measurements from the surface of the coronary chamber when continuous contact has been established according to the generated closed loop automated mapping pattern further comprises: guiding the catheter towards the surface of the coronary chamber until continuous contact between the coronary chamber and the catheter is determined; andmaintaining surface contact between the coronary chamber and the distal tip of the catheter over several systole/diastole... | Magnetecs Inc. | A61B005/02 A61B005/053 A61B005/06 A61B005/107 A61B005/00 A61B005/0402
US09066688 | B2 | 2015-06-30 | Contact activated lancet device | 2010-06-14 | A lancet device including a housing and a lancet structure having a puncturing element. The lancet structure disposed within the housing and adapted for movement between a retaining or pre-actuated position wherein the puncturing element is retained within the housing, and a puncturing position wherein the puncturing element extends through a forward end of the housing. The lancet device includes a drive spring disposed within the housing for biasing the lancet structure toward the puncturing position, and a retaining hub retaining the lancet structure in the retracted position against the bias of the drive spring. The retaining hub includes a pivotal lever in interference engagement with the lancet structure. An actuator within the housing pivots the lever, thereby moving the lancet structure toward the rearward end of the housing to at least partially compress the drive spring, and releasing the lever from interference engagement with the lancet structure. | 1. A lancet device comprising: a housing;an internal contact disposed within the housing;a lancet structure comprising a puncturing element, the lancet structure disposed within the housing and adapted for movement between a pre-actuated position and a puncturing position wherein the puncturing element extends through an end of the housing;a biasing member in cooperation with the lancet structure for biasing the lancet structure toward the puncturing position; anda lever element pivotal about a fulcrum comprising a class 1 lever for providing interference engagement with the lancet structure and retaining the lancet structure in the pre-actuated position and adapted to retain the lancet structure against the bias of the biasing member,wherein movement of the housing toward the lever element causes pivotal movement of the lever element to at least partially compress the biasing member and release the interference engagement between the lever element and the lancet structure, permitting the biasing member to drive the lancet structure through the housing toward the puncturing position, andwherein movement of the housing toward the lever element causes the internal contact within the housing to contact the lever element and pivot the lever element out of engagement with the lancet structure. 2. The lancet device of claim 1, wherein the lever element comprises a wedge-shaped structure including a shoulder in interference engagement with the lancet structure and a contact surface for engagement with the internal contact of the housing. 3. The lancet device of claim 1, wherein the lever element comprises a pair of opposing lever elements which provide interference engagement with the lancet structure and retain the lancet structure in the pre-actuated position and are adapted to retain the lancet structure against the bias of the biasing member. 4. The lancet device of claim 1, further comprising a shield extending through the forward end of the housing and movable with respect to the housing, with the lever element adjacent a rearward end of the shield. 5. The lancet device of claim 4, wherein the lancet structure and the shield include corresponding guiding surfaces for guiding the lancet structure through the shield, and the shield and the housing include corresponding guiding surfaces for guiding the shield through the housing. 6. The lancet device of claim 4, further comprising a retraction spring for retracting the lancet structure within the shield after the drive spring drives the lancet structure through the shield toward the puncturing position, and a locking structure extending between the shield and the housing for maintaining the shield in fixed relation to the housing after the biasing member drives the lancet structure through the shield toward the puncturing position. 7. A lancet device comprising: a housing comprising a forward housing body and a rear housing body, the rear housing body having a rearward end;a lancet structure co... | Becton, Dickinson and Company | A61B017/14 A61B005/15 A61B005/151
US08382705 | B2 | 2013-02-26 | Conveyor apparatus and method for supplying a medicine-solution mixture | 2008-10-01 | A conveyor apparatus and a method for supplying a medicine-solution mixture from a plurality of medicine-active ingredient solutions and additional solution constituents in a body by means of a plurality of feed devices is disclosed. The medicine-active ingredient solutions have different medicine-active ingredient parameters and constituents; the solution constituents have different parameters and constituents. The constituents of at least two medicine-active ingredient solutions and solution constituents are saved in at least one storage device; and are selected by means of one selection device for selection of at least one specific parameter and/or constituent of each medicine-active ingredient solution and/or the solution constituents. By means of a calculating device, at least one additional parameter from all selected parameters is calculated for the control and operation of at least one additional feed device for supplying at least one additional medicine-solution mixture to the body. | 1. A conveyor apparatus for supplying a medicament solution mixture of a plurality of medicament active agent solutions and/or solution constituents to a body, wherein the medicament active agent solutions and/or solution constituents have different medicament active agent parameters and/or ingredients, the conveyor apparatus comprising: a plurality of feed devices;at least one storage unit for storing the medicament active agent parameters and/or ingredients of at least two medicament active agent solutions and/or solution constituents being supplied by a corresponding at least two of the plurality of feed devices;a selection unit connected to the at least one storage unit for selecting at least one of the stored medicament active agent parameters and/or ingredients from each of the at least two medicament active agent solutions and/or solution constituents;a calculation unit for calculating at least one additional parameter based on the selected medicament active agent parameters and/or ingredients from each of the at least two medicament active agent solutions and/or solution constituents;a controller for controlling the operation of at least one other of the plurality of feed devices for supplying at least one additional medicament solution mixture to the body based on the calculated additional parameter. 2. The conveyor apparatus as claimed in claim 1, wherein the at least one additional medicament solution mixture is a solution mixture for an artificial feeding and the medicament active agent solutions are nutrient solutions. 3. The conveyor apparatus as claimed in claim 2, wherein the medicament active agent parameters are the nutrient energy content of the nutrient solution. 4. The conveyor apparatus as claimed in claim 2, wherein the artificial feeding used is parenteral, enteral or a combination of the two. 5. The conveyor apparatus as claimed in claim 1, wherein the solution constituents are carrier solutions and preservative solutions. 6. The conveyor apparatus as claimed in claim 1, wherein the ingredients are carbohydrates, proteins and fats. 7. The conveyor apparatus as claimed in claim 1, wherein the calculation unit is connected to a blood sugar measuring device for measuring the blood sugar level in the body and to an input device for inputting body-specific data. 8. A method for supplying a medicament solution mixture of a plurality of medicament active agent solutions and solution constituents to a body by means of a plurality of feed units, wherein the medicament active agent solutions have various medicament active agent parameters and ingredients and the solution constituents have various parameters and ingredients, comprising the steps of: storing the parameters and/or the ingredients of at least two medicament active agent solutions and solution constituents in at least one storage unit,selecting, at least one specific parameter and/or ingredient from each medicament active agent solution and/or the solution consti... | B. Braun Melsungen AG | A61M037/00
US09078774 | B2 | 2015-07-14 | Systems and methods for processing limb motion | 2010-08-12 | Control systems and methods are disclosed for processing a time series of signals associated with the movement of a device associated with a limb. The time series of motion signals is filtered, such as thorough an autoregressive filter, and compared to stored data sets representing a limb-motion event and/or phase. In certain examples, a plurality of accelerometers generate the time series of motion signals based at least on acceleration measurements in three orthogonal directions and/or planes. The acceleration measurements may relate to the movement of an artificial limb, such as a prosthetic or orthotic device. Upon determining an event and/or phase of limb motion, the control system may trigger an actuator to appropriately adjust one or more prosthetic or orthotic joints. | 1. A method of controlling operation of an actuatable prosthetic device being moveable about a joint, the method comprising: receiving with a processor located on the actuatable prosthetic device a plurality of data points of a sensor signal over a period of time, wherein the plurality of data points are indicative of a motion of the prosthetic device;processing with said processor the plurality of data points with a time series analysis that compares the plurality of data points received over the period of time to a plurality of sample signals so as to associate the plurality of data points with at least one sample signal of the plurality of sample signals, wherein the at least one sample signal comprises a plurality of data points within a time period that is approximately equal in length to the period of time in which the plurality of data points of the sensor signal are received;determining with said processor a control signal based at least on the at least one sample signal associated with the time series of signals; andcontrolling movement of the actuatable prosthetic device based on the control signal determined by said processor. 2. The method of claim 1, wherein said processing includes filtering the plurality of data points with a predictive filter. 3. The method of claim 1, wherein said processing includes filtering the plurality of data points with an autoregressive filter. 4. The method of claim 3, wherein said processing includes filtering the plurality of data points with a Kalman filter. 5. The method of claim 1, wherein said processing with said processor further comprises identifying at least one limb-motion event. 6. The method of claim 5, wherein the at least one limb-motion event comprises at least one of toe down, heel off, midswing, end of swing, toe-off, and heel-strike events. 7. The method of claim 1, wherein the actuatable prosthetic device comprises an actively-adjustable prosthetic ankle device. 8. The method of claim 7, wherein the sensor signal comprises an accelerometer sensor signal from an accelerometer, wherein the accelerometer is attached to the actively-adjustable prosthetic ankle device. 9. The method of claim 1, wherein the plurality of data points are spaced apart at uniform or substantially uniform time intervals. 10. The method of claim 7, wherein the actuatable prosthetic device comprises a prosthetic ankle comprising a foot unit, a lower limb member and an actuator configured to actively control an angle between the foot unit and lower limb member, and controlling movement of the actuatable prosthetic device comprises controlling the actuator. 11. The method of claim 1, further comprising using one or more accelerometers to produce the plurality of data points. 12. The method of claim 1, further comprising using a sensor module that outputs to the processor sensor signals at 100 Hz or more. 13. The method of claim 1, wherein associating the plurality of data points with the at least one sam... | Össur hf | A61F002/24 A61F002/66 A61F002/68 G01C021/16 A61F002/76 A61F002/72 A61F002/60 A61F002/70
US08496600 | B2 | 2013-07-30 | Non-reusable collection device for bodily fluids | 2010-07-29 | A non-reusable device for collecting bodily fluids such as vascular blood from a patient, the device being configured for example to receive a blood collection tube and having a retractable needle attached to a rearwardly biased needle holder that is constrained prior to needle retraction by a rotatably mounted lug ring and that is released during retraction by depressing a trigger pivotably connected to the body of the device to rotate the lug ring, whereby the needle holder is driven into a retraction cavity disposed inside the trigger and the front tip of the needle is retained inside the body of the device. | 1. A device useful for collecting bodily fluid from a patient, the device comprising: a body having a forwardly projecting, rearwardly biased, releasably constrained, selectively retractable needle, the body being configured to receive a fluid collection tube into selectively releasable fluid communication with a rear end of the retractable needle;an elongate trigger comprising a retraction cavity, a rear portion hinged to the body, an open front portion permitting access to the retraction cavity, and a lug-contacting surface selectively engageable with a lug ring rotatably mounted inside the body;whereby pivotable movement of the trigger relative to the body following release of the fluid collection tube from the body rotates the lug ring and thereby releases the rearwardly biased needle to retract into the retraction cavity. 2. The device of claim 1 wherein the bodily fluid is blood. 3. The device of claim 1 wherein the fluid collection tube is a blood collection tube. 4. The device of claim 1 wherein the retractable needle has sharp front and rear needle tips. 5. The device of claim 4 wherein the retractable needle has an upwardly facing bevel on the front needle tip and a downwardly facing bevel on the rear needle tip. 6. The device of claim 4 wherein the rear needle tip is surrounded by a flexible elastomeric sheath. 7. The device of claim 6 wherein the sheath is attached to a needle holder seated inside the body. 8. The device of claim 1 wherein the retractable needle is attached to a needle holder seated inside the body. 9. The device of claim 8 wherein the needle holder has a transverse flange. 10. The device of claim 8 wherein a portion of the needle holder projects forwardly of the body. 11. The device of claim 8 wherein the needle holder is constrained by the lug ring. 12. The device of claim 11 wherein the lug ring comprises a centrally disposed aperture alignable with a transverse flange of the needle holder. 13. The device of claim 1 wherein the needle is rearwardly biased by a compressed retraction spring. 14. The device of claim 1 wherein retractable needle is constrained by the lug ring. 15. The device of claim 1 wherein the body comprises an open back end into which a fluid collection tube can be releasably inserted. 16. The device of claim 1 wherein the body comprises an upwardly facing slot. 17. The device of claim 16 wherein the trigger is depressed inside the slot to rotate the lug ring. 18. The device of claim 1 wherein the body comprises at least one transversely projecting flange member. 19. The device of claim 18 wherein the at least one transversely projecting flange member has a substantially flat bottom. 20. The device of claim 1 wherein the activation member of the trigger is a surface engageable with a lug projecting from the lug ring. 21. The device of claim 1 wherein the body, trigger and lug ring are made of molded plastic. | Retractable Technologies, Inc. | A61B005/00
US09579456 | B2 | 2017-02-28 | Methods for reducing false hypoglycemia alarm occurrence | 2010-05-21 | A system and method for reducing the number of hypoglycemic alarms presented to a user is presented. The system and methods include use of model based state estimation and variable-delayed threshold values to balance the risk of not presenting an alarm caused by an actual hypoglycemic state with the presentation of alarms caused by artifacts in the signals produced by a continuous glucose monitor. | 1. An electronic analyte monitoring system comprising: an electronic analyte monitor comprising an implantable analyte sensor;a processor configured to receive a signal from the electronic analyte monitor indicative of an analyte level obtained from the implantable analyte sensor, wherein the signal can indicate the analyte level as falsely low, the processor also adapted to recursively analyze the signal in accordance with one or more software commands, the one or more software commands programming at least a portion of the processor to operate as a state machine with states comprising: a no hypoglycemia confirm state during which the analyte level is compared to a first threshold, the no hypoglycemia confirm state having a predetermined time delay for activation of an alarm where a duration of the predetermined time delay is dependent on a value of the analyte level, wherein the predetermined time delay is initiated if the analyte level is below the first threshold and, after being initiated, the predetermined time delay is reset if the analyte level rises above the first threshold and the predetermined time delay has not yet expired; anda confirm hypoglycemia state that is entered if the analyte level is below the first threshold after expiration of the predetermined time delay, the confirm hypoglycemia state causing generation of the alarm; andwherein the processor is further coupled to a memory on which instructions are stored, wherein the instructions, when executed by the processor, cause the processor to transmit a command to cause adjustment of a medication delivery rate when there is a confirm hypoglycemia state. 2. The system of claim 1, wherein the duration of the predetermined time delay is a first length when the analyte level is below the first threshold and above a second threshold. 3. The system of claim 2, wherein the duration of the predetermined time delay is a second length when the analyte level is below the second threshold, the second length being less than the first length. 4. The system of claim 3, wherein the state machine further comprises: a confirm severe hypoglycemia state that is entered when the analyte level is below a third threshold that is less than the second threshold and indicative of severe hypoglycemia, the confirm severe hypoglycemia state causing generation of the alarm. 5. The system of claim 2, wherein the predetermined time delay is a first predetermined time delay and, if the duration of the first predetermined time delay is set to the first length and the analyte level falls below the second threshold, the no hypoglycemia confirm state initiates a second predetermined time delay having a second length that is less than the first length, and wherein the confirm hypoglycemia state is also entered if the analyte level is below the second threshold after expiration of the second predetermined time delay, the confirm hypoglycemia state causing generation of the alarm. 6. The system of claim 5, wher... | ABBOTT DIABETES CARE INC. CAMBRIDGE ENTERPRISE LIMITED | A61B005/00 A61M005/172 A61B005/145 A61M005/142 G06F019/00
US09480461 | B2 | 2016-11-01 | Methods for extracting chemicals from nasal cavities and breath | 2010-04-30 | A chemical extraction apparatus has absorbent material for absorbing or adsorbing volatile chemicals from a breath of a user. The chemical extraction apparatus is positioned within such that breaths from the user flow over the absorbent material thereby causing volatile chemicals in the breaths to absorb or adsorb into the absorbent material. The absorbent material is then analyzed to determine the amount of chemicals absorbed or adsorbed into the absorbent material. Based on such analysis, various diseases or conditions can be detected or diagnosed. | 1. A method for diagnosing a medical condition of a user, comprising the steps of: inserting a first chemical extraction apparatus into a first nasal cavity of a user such that trace levels of chemicals from a plurality of breaths are concentrated in absorbent material of the first chemical extraction apparatus, wherein the first chemical extraction apparatus comprises: a hollow support element and a first tubular sample element that is positioned within a cavity of the hollow support element and is dimensioned with a length of less than about two centimeters and a width of less than about one-eighth of an inch,wherein the first tubular sample element comprises the absorbent material and is coupled to the support element via a first at least one arm, and wherein the breaths of the user flow through the cavity and contact the first tubular sample element,wherein the hollow support element has a first opened end and a second opened end, and wherein the breaths flow through the hollow support element from the first opened end to the second opened end, and wherein the first tubular sample element extends longitudinally within the cavity from a first end of the first tubular sample element to a second end of the first tubular sample element in a direction of the flow of the breaths,wherein the first tubular sample element is positioned within the hollow support element such that the first end of the first tubular sample element faces the first opened end and the second end of the first tubular sample element faces the second opened end,wherein the first chemical extraction apparatus comprises a second tubular sample element that is positioned within the cavity of the first chemical extraction apparatus,wherein the second tubular sample element comprises absorbent material and is coupled to the support element via a second at least one arm, and wherein the breaths of the user flow through the cavity and contact the second tubular sample element,wherein the second tubular sample element extends longitudinally within the cavity from a third end of the second tubular sample element to a fourth end of the second tubular sample element in the direction of the flow of the breaths,wherein the second tubular sample element is positioned within the hollow support element such that the third end of the second tubular sample element faces the first opened end and the fourth end of the second tubular sample element faces the second opened end;removing the first chemical extraction apparatus from the first nasal cavity after the plurality of breaths are applied to the first chemical extraction apparatus;separating the first and second tubular sample elements from the support element at the first and second at least one arm subsequent to the removing;thermally desorbing at least one chemical from the absorbent material via a thermal desorption unit subsequent to the separating;determining an amount of the at least one chemical concentrated in the absorbent material ... | Volatile Analysis Corporation | A61B005/08 A61B010/00 A61B005/097 A61J007/00 A61J017/00 A61J017/02 G01N001/40 G01N001/22
US10445463 | B2 | 2019-10-15 | Health booth | 2007-09-11 | A health booth (10) includes a shell (20), at least one chair (30), and at least one measurement arrangement (31, 32, 33, 34) for measuring data relative to the health of a user. The booth includes: a determination apparatus (41-50) for determining, at the time of taking a measurement, at least one condition under which said measurement is taken; and a storage device for storing the measurement results in a data structure together with said condition. | 1. A health booth configured to allow a patient to conduct self-measurement of health data of the patient, the health booth comprising: a shell;a chair arranged inside the shell and configured to support the patient in a seated position;a first measurement device, arranged inside the shell, configured for the patient to measure by him/herself the health data;a second measurement device, arranged on the inside or outside of the shell, configured to measure, at the time of measurement of the health data, environmental condition data that affects the measured health data and interpretation of the measured-health data;a data storage that stores the health data and the environmental condition data together in a computer file; anda transmission system configured to send the computer file to a remote center for evaluation, wherein the second measurement device comprises a patient position sensor configured to measure a portion of the environmental condition data, the patient position sensor further configured to detect a standing position of the patient associated with measurement of health data that requires the patient to be standing during measurement of the health data. 2. The health booth according to claim 1, wherein the second measurement device further comprises at least one of: a thermometer to measure the temperature inside and/or outside the shell; a microphone to measure the sound level inside and/or outside the shell; a hygrometer to measure the humidity content inside and/or outside said shell; and a photo-electric cell to measure a brightness level inside and/or outside said shell. 3. The health booth according claim 1, wherein the shell further comprises at least one opening for connecting the health booth to at least one other health booth. 4. The health booth according to claim 1, wherein the first measurement device comprises a scale configured to measure the weight of the patient when the patient is seated on the chair. 5. The health booth according to claim 4, wherein the second measurement device includes at least one pressure sensor selected from: a first pressure sensor placed on the seat of the chair,a second pressure sensor placed on the backrest of the chair,a third pressure sensor placed on an armrest of the chair, anda fourth pressure sensor placed on a footrest attached to the chair,said at least one pressure sensor being configured to determine whether the patient is present in the chair and properly positioned for the weight measurement. 6. The health booth according to claim 1, wherein the first measurement device comprises an inflatable cuff configured to measure blood pressure of the patient, the cuff further including heart rate sensors and being configured to take measurements when the patient is standing or seated in the chair. 7. The health booth according to claim 1, wherein the first measurement device comprises a height gauge configured to measure the height of the patient in a standing position, and the s... | H4D INTERNATIONAL S.A. | A61B005/00 G06F019/00 G16H050/80
US20100312129 | A1 | 2010-12-09 | CARDIOVASCULAR HAPTIC HANDLE SYSTEM | 2010-07-15 | Cardiac tissue motion characteristics acquired by novel cardiac sensors are analyzed and processed for the derivation of physiological indices. The indices are output to a hand held local or remote volumetric haptic display and enable an operator to obtain motion related dynamic characteristics of cardiac tissues. The ability to tactually sense the motion of cardiac tissue and the affect on such motion from inserted cardiovascular instrumentation enhances the operator's performance of procedures including the positioning and placement of implanted catheters/sensors, extraction of permanently implanted leads and delivery of cardiovascular therapies. Optimal haptic rendering is achieved by using computational techniques to reconstruct the physically and perceptually relevant aspects of acquired signals and bridge the gap between the inserted catheter and operator's hand/catheter handle. | 1. A haptic system providing tactile sensations simulating the dynamic characteristics of a moving biological organ or tissue of a patient comprising: a sensor implanted within the patient and in contact with moving tissue or organ and generating sensor signals indicative of dynamic characteristics of the tissue or organ,a processor receiving the sensor signals and converting them into tactile signals; anda haptic handle having at least one tactile element and an actuator coupled to said tactile element, said actuator receiving said tactile signals and, in response causing said tactile element to move in a matter selected to render said parameter to the user when the user holds the handle. 2. The system of claim 1 wherein said processor is hard wired to said sensor. 3. The system of claim 1 wherein said processor is communicating with said sensor wirelessly to receive said sensor signals, 4. The system of claim 1 further includes a catheter coupled to said haptic handle and having a body extending to said tissue or organ, with said sensor is in contact with said tissue. 6. The system of claim 1 further including a localization device adapted to provide information indicating the location of the sensors. 7. The system of claim 6 wherein said localization device is an external device and includes at least one of a navigational system, an electromagnetic system, an electrical system, magnetic system, impedance based system, an optical system and a thermal system. 8. The system of claim 1 wherein said wherein said handle includes a plurality of tactile elements operated with respective actuators to render at least one of a multidimensional motion. 9. The system of claim 1 wherein said sensor is adapted to sense at least one mechanical property of a moving tissue or organ including one of a texture, temperature, elasticity, thickness, deformability and vibrotactile effects. 10. A haptic handle system providing tactile rendering of the dynamic characteristics of a cardiac tissue comprising: a sensor adapted to be selectively coupled to the cardiac tissue and generating sensor signals indicative of the motion of the cardiac tissue;a processor receiving said sensor signals and being adapted to analyze said sensor signals for various artifacts associated with said sensor signals that degrade the quality of information contained in said sensor signals, said processor being further adapted to compensate for said artifacts and generate corresponding tactile signals;a handle having at least one tactile element and a tactile element actuator receiving said tactile signals and causing said tactile element to move in a manner selected to provide a tactile rendering of the motion of said cardiac tissue. 12. The system of claim 10 wherein at least said handle is disposed remotely from said sensor, said system further including a communication device providing a communication path for signals from said sensor, whereby the handle is adapted to provide virtu... | - | A61B005/02 A61B005/11
US20100298649 | A1 | 2010-11-25 | SYSTEM AND METHODS FOR ASSESSMENT OF THE AGING BRAIN AND ITS BRAIN DISEASE INDUCED BRAIN DYSFUNCTIONS BY SPEECH ANALYSIS | 2008-11-03 | A system and method for assessment of a brain status of a subject are disclosed. The brain status comprises a brain disease induced brain dysfunction. An occurrence and/or stage of the brain disease induced brain dysfunction in the subject is determined. The system comprises an apparatus that is adapted to determine the occurrence and/or stage of the brain disease induced brain dysfunction in the subject from random speech of the subject. The apparatus (200) comprises units that are operatively connected to each other, which comprises a unit (205) for registering the speech of the subject over a period of time; a unit (206) devised for analyzing the registered speech and configured to determine a pause component of the speech; and a unit that is adapted to determine the occurrence and/or stage of the brain disease induced brain dysfunction from the pause component, wherein said pause component is an accumulated pause time of a total time of said speech correlated to said occurrence and/or stage of said brain dysfunction. | 1-22. (canceled) 23. A method of diagnosing a brain disease induced brain dysfunction of a subject, comprising registering speech of said subject over a period of time;determining a pause component of said registered speech; anddetermining an occurrence and/or stage of said brain disease induced brain dysfunction from said pause component, wherein said pause component is an accumulated pause time of a total time of said speech correlated to said occurrence and/or stage of said brain dysfunction, andcomparing said pause component with a pre-determined normal pause-component for said diagnosis. 24. A method for assessment of a brain status of a subject, wherein said assessment is performed internally in a system, wherein said brain status comprises a brain disease induced brain dysfunction, wherein said method comprises analyzing speech of said subject and determining a pause component of said speech; anddetermining an occurrence and/or stage of said brain disease induced brain dysfunction in said subject based on said pause component, wherein said pause component is an accumulated pause time of a total time of said speech correlated to said occurrence and/or stage of said brain dysfunction. 25. The method according to claim 24, comprising registering said speech and/or recording said speech of said subject over a period of time; and wherein said analyzing said speech comprises analyzing said registered speech and/or said recorded speech for determining said pause component of said speech. 26. The method according to claim 24, wherein said analyzing said speech of said subject is performed irrespective of a language of said speech. 27. The method according to claim 24, comprising applying a compensation factor for a specific language of said speech for said assessment. 28. The method according to claim 24, comprising applying a compensation factor related to an age of said subject. 29. The method according to claim 24, wherein said assessment is a cognitive test based assessment, comprising the subject freely defining parameters of said cognitive test, for producing said speech. 30. The method according to claim 24 comprising providing a basis for medical personal for deciding if a subject has signs of a brain disease induced brain dysfunction or not, based on said pause component. 31. The method according to claim 30, further comprising directing primary health care resources to those subjects who are at high risk for having a brain disease induced brain dysfunction, and who need further assessment for their diagnosis, while saving financial costs for unnecessary evaluations of patients with negative test results. 32. The method according to claim 24, comprising basing said occurrence and/or stage of said brain disease induced brain dysfunction on a threshold value of said pause time component. 33. The method according to claim 32, wherein said threshold value comprises different ranges for said occurrence and/or stage of said brain dis... | - | A61B005/00 G01N033/00
US20100331729 | A1 | 2010-12-30 | CIRCUIT BOARD FOR BODY FLUID COLLECTION, AND BIOSENSOR | 2009-01-21 | A circuit board for body fluid collection include a first substrate including a puncture needle; a second substrate disposed at an upstream side of the first substrate in a direction of puncturing with the puncture needle with a space provided therebetween; a flexible insulation layer that extends between the first substrate and the second substrate; and a conductive pattern provided on the flexible insulation layer, the conductive pattern integrally including an electrode, a terminal, and a wiring that electrically connects the electrode and the terminal. | 1. A circuit board for body fluid collection comprising: a first substrate including a puncture needle,a second substrate that is disposed at an upstream side of the first substrate in a direction of puncturing with the puncture needle with a space provided therebetween,a flexible insulation layer that extends between the first substrate and the second substrate, anda conductive pattern provided on the flexible insulation layer, the conductive pattern integrally including an electrode, a terminal, and a wiring that electrically connects the electrode and the terminal. 2. The circuit board for body fluid collection according to claim 1, wherein the terminal is disposed so as to overlap with the second substrate when projected in the thickness direction. 3. The circuit board for body fluid collection according to claim 1, wherein the electrode is disposed so as to overlap with the first substrate when projected in the thickness direction. 4. The circuit board for body fluid collection according to claim 1, wherein the electrode is disposed so as to overlap with a portion between the first substrate and the second substrate when projected in the thickness direction. 5. The circuit board for body fluid collection according to claim 1, wherein the electrode is disposed so as to overlap with the second substrate when projected in the thickness direction. 6. A biosensor comprising a housing and a circuit board for body fluid collection comprising: a first substrate including a puncture needle,a second substrate that is disposed at an upstream side of the first substrate in a direction of puncturing with the puncture needle with a space provided therebetween,a flexible insulation layer that extends between the first substrate and the second substrate, anda conductive pattern provided on the flexible insulation layer, the conductive pattern integrally including an electrode, a terminal, and a wiring that electrically connects the electrode and the terminal,wherein the circuit board for body fluid collection is accommodated in the housing, the second substrate is fixed, and the first substrate is supported so as to be capable of advancing from inside the housing to outside the housing. | Nitto Denko Corporation | A61B005/151
US20100298669 | A1 | 2010-11-25 | BAN-SENSOR RADIO COMMUNICATION DEVICE AND METHOD | 2010-05-17 | A radio communication device for communicating sensor information includes an electromagnetic wave communication unit using electromagnetic waves using the air as a propagation path, an ultrasonic communication unit using ultrasonic waves using an inside of a living body as a propagation path, and a midair/living-body switching control unit for switching between a communication performed by the electromagnetic wave communication unit and a communication performed by the ultrasonic wave communication unit. | 1. A radio communication device for communicating sensor information, comprising: an electromagnetic wave communication unit using electromagnetic waves using the air as a propagation path;an ultrasonic communication unit using ultrasonic waves using an inside of a living body as a propagation path; anda midair/living-body switching control unit for switching between a communication performed by the electromagnetic wave communication unit and a communication performed by the ultrasonic wave communication unit. 2. The radio communication device according to claim 1, wherein when the communication performed by the electromagnetic wave communication unit is conducted, the midair/living-body switching control unit switches from the communication performed by the electromagnetic wave communication unit over to the communication performed by the ultrasonic wave communication unit according to a receiving state of a reception confirm signal from a radio communication device on an opposite party. 3. The radio communication device according to claim 1, wherein when the communication performed by the ultrasonic wave communication unit is conducted, the midair/living-body switching control unit determines a receiving state of a midair electromagnetic wave communication by attempting to conduct the communication performed by the electromagnetic wave communication unit with a radio communication device on an opposite party at a predetermined timing, and when it is determined that the receiving state has become a predetermined state, the midair/living-body switching control unit switches from the communication performed by the ultrasonic wave communication unit over to the communication performed by the electromagnetic wave communication unit. 4. The radio communication device according to claim 1, wherein when a high-security data transfer is required, the midair/living-body switching control unit switches the communication over to the communication performed by the ultrasonic wave communication unit so as to conduct an encryption key data communication with a radio communication device on an opposite party, then switches the communication over to the communication performed by the electromagnetic wave communication unit so as to conduct an encryption communication using the encryption key data with a radio communication device on an opposite party. 5. The radio communication device according to claim 1, wherein when a high-security data transfer which does not requires a high transfer rate is required, the midair/living-body switching control unit switches the communication over to the communication performed by the ultrasonic wave communication unit so as to perform the high-security data transfer. 6. A radio communication method for communicating sensor information, comprising switching between an electromagnetic wave communication performed by an electromagnetic wave communication unit using the air as a propagation path and an ultrasonic communicati... | FUJITSU LIMITED | A61B005/07
US20100312097 | A1 | 2010-12-09 | Hybridized optical-MRI method and device for molecular dynamic monitoring of in vivo response to disease treatment | 2010-06-04 | An apparatus for providing physiological information from an organism in disease diagnosis and treatment monitoring, for use in an MRI instrument. The apparatus operates on the concept of hybridized magneto-optical sensitivity. The MRI includes an MRI scanner and a controller for controlling the MRI scanner. The MRI scanner provides a magnetic field of at least 0.5T. The apparatus further includes a front end built of non-magnetic components, comprising light guides for illuminating a region of interest (ROI) and for collecting light emitted at said ROI; and a back-end comprising a light source for injecting light into said light guides; a light detector for receiving light collected at said ROI; and a processing and control unit for processing said light collected at said ROI. | 1. A method of hybridizing magnetic and optical fields for providing physiological imaging of an organism, said method comprising the steps of: (a) generating a the magnetic field with an MRI device;(b) generating an optical field with an optical device integrated within the MRI device;(c) providing a magneto-optically sensitive contrast marker, wherein the hybridization of both magnetic and optical fields is provided by said magneto-optically sensitive contrast marker injected into the organismwherein said hybridization is based on the local production of paramagnetic radical pairs from the contrast marker interacting with the organism's tissue; and (d) detecting a magnetic resonance response from standard MRI techniques; and(e) detecting at least one of absorbance, luminescence, fluorescence or phosphorescence generated by the interaction of the contrast marker with the organism's tissues. 2. The method of claim 1 wherein the contrast marker interacting with the tissues within the magnetic and optical fields generates a different and specific magneto-optical response for an optical parameter for at least two different values of the magnetic field, such as to generate a magneto-optical response curve. 3. The method of claim 1 wherein a change of the magneto-optical response curve is linked to a change in a physiological parameter of the organism tissues. 4. The method of claim 1 wherein the optical device includes an image generator to generate images. 5. The method of claim 4, wherein an image pixel represents the value of a physiological parameter of the organism tissues, based on the magneto-optical response. 6. The method of claim 1 where in the optical device is integrated in such a way to provide multiple image projections, enabling 3D tomographic imaging. 7. The method of claim 2, wherein said optical parameter is at least one of intensity, spectral properties or lifetime of the detected optical signal. 8. An apparatus for providing physiological information from an organism in disease diagnosis and treatment monitoring, for use in an MRI instrument, said apparatus operating on the concept of hybridized magneto-optical sensitivity; said MRI including an MRI scanner and a controller for controlling said MRI scanner, said MRI scanner providing a magnetic field of at least 0.5T;said apparatus comprising:a front end built of non-magnetic components, comprising light guides for illuminating a region of interest (ROI) and for collecting light emitted at said ROI;a back-end comprising:a light source for injecting light into said light guides;a light detector for receiving light collected at said ROI; anda processing and control unit for processing said light collected at said ROI. 9. An apparatus according to claim 8, wherein said front end is adapted to observe said ROI with no contact. 10. An apparatus according to claim 8, wherein said front end is further provided with bulk optics. 11. An apparatus according to claim 10, wherein sa... | Institut National D'Optique | A61B005/055
US07860561 | B1 | 2010-12-28 | Method of quantifying a subject's wake or sleep state and system for measuring | 2004-06-04 | The present invention relates to a method of quantifying a subject's wake or sleep state, and more particularly to a higher resolution method of determining when a subject is awake or asleep. The present invention additionally relates more specific applications of this method, and a system for making these quantitative measurements. There are numerous embodiments of the present invention, which are envisioned, with a few of those listed below. The present invention relates to a method of analyzing a subject, and preferably a human subject to determine quantitatively when the subject is awake and when the subject is sleeping. The present invention has application for use with alertness devices, for use to better analyze sleeping disorders, anesthesia monitors, psychiatric and hypnosis medication monitoring, and for looking at the effect of various other therapies. | 1. A method of quantifying wake-sleep states comprising the steps of: a) using at least one sensor to measure a subject's brain wave signals over a measurement time period, the measurement time period comprising a number of time segments;b) analyzing with a computer processor the subject's brain wave signals to estimate or determine a dominant frequency profile; dominant frequency indices or dominant frequency index for each time segment; andc) performing with the computer processor a wavelet decomposition on the dominant frequency profile, indices or index over the measurement time period. 2. The method of claim 1, further comprising the steps (d) analyzing the subject's brain wave signals to estimate or determine a power spectrum profile, power spectrum indices or power spectrum index for each time segment; and (e) performing a wavelet decomposition on the power spectrum profile, indices or index over the measurement time period. 3. The method of claim 2, further comprising the step of determining a wake-sleep transition based in part on the results of performing the wavelet decomposition on the power spectrum in the previous step (c) and in part on the results on the results of performing a wavelet decomposition on the dominant frequency from step (e). 4. The method of claim 1, wherein a wake-sleep transition can be determined in real time. 5. The method of claim 1, further comprising the step of quantifying the subject's overall sleep over the measurement time period. 6. The method in claim 1, wherein the method is used to trigger an alert to the subject. 7. A method of quantifying wake-sleep states comprising the steps of: a) using at least one sensor to measure a subject's brain wave signals over a measurement time period;b) analyzing with a computer processor the subject's brain wave signals to estimate or determine time-varying frequency spectral characteristics; andc) performing with the computer processor a wavelet decomposition on the time-varying frequency spectral characteristics over the measurement time period. 8. The method of claim 7, wherein a wake-sleep transition can be determined in real time real time. 9. The method of claim 7, further comprising the step of quantifying the subject's overall sleep over the measurement time period. 10. The method in claim 7, wherein the method is used to trigger an alert to the subject. 11. The method in claim 10, wherein the alert is either an alarm or a physical stimulus. 12. The method in claim 1, wherein the method is used to adjust the subject's sleep medication. 13. The method in claim 1, wherein the method is used to adjust the subject's sleep therapy. 14. The method in claim 1, wherein the method is used to research medications which can be used for sleeping disorders. | Cleveland Medical Devices Inc. | A61B005/04
US07840255 | B2 | 2010-11-23 | X-ray CT apparatus and myocardial perfusion image generating system | 2005-03-10 | An X-ray CT apparatus 1 for radiating X-ray to a subject P so as to scan the subject P and reconstructing a image within the subject P based on an obtained projection data comprises an image generating unit and a blood-flow image generating unit 24e. The image generating unit generates the image based on the projection data in a state in which a concentration of a contrast medium in a myocardial portion of a subject P continuously injected the contrast medium can be considered to be constant. The blood-flow image generating unit 24e generates a blood flow image by removing components of a myocardial tissue from a portion of a myocardial region of the image generated by the image generating unit. | 1. An X-ray CT apparatus comprising: an X-ray tube for radiating X-ray to a subject so as to scan the subject;an X-ray detector for detecting the radiated X-ray;a reconstruction unit for reconstructing an image within the subject based on an obtained projection data;a contrast-medium injector controlling unit configured to control an injector so that the injector continuously injects a contrast medium into a subject by controlling the injector to inject the contrast medium at a first speed for a first time interval and to inject the contrast medium at a second speed for a second time interval following the first time interval, to obtain a state in which a concentration of the contrast medium in a myocardial portion of the subject is considered to be at a preset desired constant value;an image generating unit configured to generate the image based on the projection data once in the state in which the concentration of the contrast medium in the myocardial portion of the subject is considered to be at the preset desired constant value; anda blood-flow image generating unit configured to generate a blood flow image representing a blood flow rate in myocardium by removing components of a myocardial tissue from a portion of a myocardial region of the image generated by the image generating unit. 2. An X-ray CT apparatus according to claim 1, further comprising a determining region unit configured to determine the portion of the myocardial region of the image generated by the image generating unit, wherein the blood-flow image generating unit is configured so as to remove the components of the myocardial tissue from the portion of the myocardial region determined by the determining region unit. 3. An X-ray CT apparatus according to claim 1, wherein the image generating unit is configured to generate the blood flow image by subtracting a CT value equivalent to the myocardium from a CT image of the portion of the myocardial region. 4. An X-ray CT apparatus according to claim 1, further comprising a display processing unit configured to set a value of a window level at the time of performing at least either gradation conversion processing and color tone conversion processing based on a value corresponding to the myocardial portion contrasted by the contrast medium and a value corresponding to the myocardium. 5. An X-ray CT apparatus according to claim 1, further comprising a display processing unit configured to set a value of a window level at a time of performing at least either gradation conversion processing and color tone conversion processing so as to emphasize pixels having a value between a value corresponding to the myocardial portion contrasted by the contrast medium and a value corresponding to the myocardium. 6. An X-ray CT apparatus comprising: an X-ray tube for radiating X-ray to a subject;an X-ray detector for detecting the radiated X-ray;a reconstruction unit for reconstructing an image within the subject based on an obtained projectio... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | A61B005/05
US07857761 | B2 | 2010-12-28 | Acoustic blood analyzer for assessing blood properties | 2004-04-16 | An acoustic blood analyzer with a transducer section of acoustic biosensors for measurement of blood properties including, but not limited to blood coagulation, platelet function and various blood disorders in a blood sample is provided. Methods for use of the blood analyzer in measurement of blood properties are also provided. | 1. A device for analyzing blood, comprising: a transducer element, the transducer element including at least one of piezoelectric, electrostrictive, magnetostrictive, acousto-optic, or thermo-acoustic sensors, or a combination thereof;a biologically active substance in communication with the transducer element, the biologically active substance promoting interactions between the blood and the transducer element;a signal driver in communication with the transducer element, the signal driver applying a signal to the transducer element, and the signal driver varying a value of the signal;an inlet port configured to direct blood to the transducer element; anda signal processor in communication with the transducer element, wherein the signal processor measures a response of the blood to the signal and determines one or more characteristics of the blood as a function of the measured response. 2. The device of claim 1, wherein the transducer element includes an array of sensors. 3. The device of claim 1, wherein the biologically active substance facilitates determination of a characteristic of the blood. 4. The device of claim 1, wherein the value is at least one of the following: a frequency, an amplitude, or a phase. 5. The device of claim 4, wherein the frequency values ranges from 1 KHz to 10 GHz. 6. The device of claim 4, wherein the frequency values are provided to the transducer in at least one of the following ways: individually, sequentially, or simultaneously at the available frequencies. 7. The device of claim 4, wherein the frequency values include at least one of resonant, antiresonant, harmonic, or anharmonic frequencies of a first and higher orders. 8. The device of claim 1, wherein a depth of penetration into the blood by an effect created by the transducer element is in the range of 1 nanometer to 1 centimeter from a surface of the transducer element. 9. The device of claim 1, further comprising a catheter in communication with the transducer. 10. The device of claim 1, wherein the device is self-administered. 11. The device of claim 1, further comprising a first and second acoustic sensor, wherein the first sensor analyzes the blood, and wherein the second sensor compares the blood to a reference fluid. 12. The device of claim 1, wherein the biologically active substance is collagen, wherein the value of the signal that is varied is a frequency and the one or more characteristics including at least one of a platelet adhesion or a coagulation, and wherein a higher frequency permits detection of the platelet adhesion, and wherein a lower frequency permits detection of the coagulation. 13. The device of claim 1, wherein the biologically active substance is thromboplastin, wherein the value of the signal that is varied is a frequency and the one or more characteristics including at least one of a blood coagulation, a plasma coagulation factor concentration or a plasma coagulation factor activation, and wherein a lower frequency... | Drexel University | A61B005/00
US07837634 | B2 | 2010-11-23 | Fiberoptic tissue motion sensor | 2006-03-03 | Devices and methods for evaluating motion of a tissue location, such as cardiac tissue location, e.g., heart wall, are provided. Devices employed in the subject methods are fiberoptic tissue motion sensors that include a light guide having a reflective element at its distal end and a light emitter/detector at its proximal end. In embodiments of the methods, a signal obtained from reflective element stably associated with the tissue location of interest is employed to evaluate movement of the tissue location. The subject methods and devices find use in a variety of different applications, including cardiac resynchronization therapy. | 1. An implantable tissue motion sensor, the tissue motion sensor comprising: a light guide having a proximal end and a distal end;a mirror located at the distal end, wherein the mirror is configured such that a movement of a tissue located in a subject and stably associated with the mirror causes the mirror to move; anda light transmitter and a light detector both optically coupled to the proximal end, wherein the mirror is hinged through a first set of torsion beams to a first frame. 2. The implantable tissue motion sensor according to claim 1, wherein the mirror has a center of mass that is located outside an axis of rotation corresponding to the torsion beams such that motion of a chip containing the mirror in a direction substantially orthogonal to the axis causes the mirror to move. 3. The implantable tissue motion sensor according to claim 1, wherein the first frame is hinged to a second frame through a second set of torsion beams. 4. The implantable tissue motion sensor according to claim 3, wherein said mirror has a center of mass that is located outside both a first axis of rotation corresponding to the first set of torsion beams and a second axis of rotation corresponding to the second set of torsion beams such that motion of a chip containing said mirror in a direction substantially orthogonal to either the first axis or the second axis causes said mirror to move. 5. The implantable tissue motion sensor according to claim 1, wherein the light guide is an optical fiber. 6. The implantable tissue motion sensor according to claim 5, wherein the optical fiber is a single-mode optical fiber. 7. The implantable tissue motion sensor according to claim 5, wherein the optical fiber is a multi-mode optical fiber. 8. The implantable tissue motion sensor according to claim 1, wherein the sensor further comprises a collimating element. 9. The implantable tissue motion sensor according to claim 8, wherein the collimating element is positioned at the distal end of the light guide. 10. The implantable tissue motion sensor according to claim 9, wherein the collimating element is selected from the group consisting of a convex lens, a piano-convex lens, and a graded index (GRIN) lens. 11. The implantable tissue motion sensor according to claim 8, wherein said collimating element is integral with said distal end of said light guide. 12. The implantable tissue motion sensor according to claim 11, wherein said collimating element comprises an expanded core of said light guide. 13. The implantable tissue motion sensor according to claim 1, wherein the distal end of the light guide and the mirror are present in an elongated structure that is dimensioned to be placed within a vascular lead. 14. The implantable tissue motion sensor according to claim 13, wherein the vascular lead is a cardiac pacing lead. 15. The implantable tissue motion sensor according to claim 13, wherein the elongated structure comprises a metal coil. 16. The implantable tis... | Proteus Biomedical, Inc. | A61B005/103 A61B005/117 F21V009/00 G02B006/26 G02B006/42
US07857770 | B2 | 2010-12-28 | Guide wire advancer assembly | 2006-02-01 | A guide wire advancer (102) for advancing a guide wire (107) into a patient and including a body having a longitudinal axis. The body includes a first portion (110) adapted to allow a user to advance a guide wire along the longitudinal axis using a finger and a thumb, and a second portion (108) connected to the first portion. The second portion (108) is adapted to allow a user to advance a guide wire (107) along the longitudinal axis using only one finger or the thumb to bias the guide wire against a platform (164) of the second portion when advancing the guide wire (107). Multiple methods of advancing a guide wire using the advancer are also disclosed. | 1. A guide wire advancer assembly comprising: a guide wire;a guide wire storage tube in which said guide wire is at least partially disposed; andan advancer body comprising a longitudinal axis which is co-linear with said guide wire, wherein the body comprises: a first portion adapted to allow a user to advance a guide wire along the longitudinal axis using a finger and a thumb to directly engage the guide wire, with the first portion at least supporting the guide wire at two axially spaced locations while permitting access between the two axially spaced locations to the guide wire for guide wire advancement; anda second portion having a longitudinally extending guide wire platform spaced from the first portion and along which the guide wire is exposed and which is adapted to allow a user to advance the guide wire along the longitudinal axis using one of the finger or the thumb to engage the guide wire, said guide wire being urgeable by said finger or thumb along the longitudinal axis when being pressed by said finger or thumb against the platform. 2. The guide wire advancer assembly according to claim 1, wherein the first portion and the second portion are relatively rotatable about the longitudinal axis. 3. The guide wire advancer according to claim 1, wherein a more distal of the first and second portions concludes in a distal tip adapted to cooperate with an introducer to enable the guide wire to directly enter the introducer and then into a patient when advanced. 4. The guide wire advancer according to claim 1, wherein the distal tip is frustoconical. 5. The guide wire advancer assembly according to claim 1, wherein the first portion is disposed distally of the second portion. 6. The guide wire advancer assembly according to claim 1, wherein the first portion comprises: a distal portion having a first passageway extending longitudinally therethrough;a proximal portion having a second passageway extending longitudinally therethrough; anda central portion having first and second connecting members extending between the distal portion and the proximal portion, wherein the first and second connecting members define an opening therebetween. 7. The guide wire advancer assembly according to claim 6, wherein the first passageway is adapted to frictionally engage the guide wire. 8. The guide wire advancer assembly according to claim 6, wherein the first and second members are arcuate. 9. The guide wire advancer assembly according to claim 6, wherein the opening is sufficiently large to enable the finger and the thumb to grasp the guide wire extending through the opening and advance the guide wire along the first and second passageways. 10. The guide wire advancer assembly according to claim 1, wherein the second portion comprises a distal portion having a first passageway extending longitudinally therethrough, a proximal portion having a second passageway extending longitudinally therethrough, and wherein the platform is elongated and connec... | Medical Components Inc. | A61B005/00 A61M025/00
US07846975 | B2 | 2010-12-07 | Molecular transporters based on sugar and its analogues and processes for the preparation thereof | 2005-06-29 | The inventive molecular transporter compound shows significantly high permeability through a biological membrane such as a plasma membrane, nuclear membrane and blood-brain barrier, and accordingly, can be effectively used in delivering various biologically active molecules. | 1. A sugar alcohol of formula (1):wherein, R1 and R2 are each independently H, alkyl, arylalkyl, cycloalkyl, heteroalkyl, —(CH2)mNHR′, —(CH2)lCO2R″, —COR′″, —SO2R″″ or a physiologically active molecule; R′, R″, R′″ and R″″ being each independently H, alkyl, arylalkyl, cycloalkyl, heteroalkyl or a physiologically active molecule; m being an integer in the range of 2 to 5; and l being an integer in the range of 1 to 5;R3 isn being an integer in the range of 1 to 12. 2. The sugar alcohol of claim 1, which is an alditol having the skeleton of sorbitol, mannitol, galactitol, or a salt thereof. 3. A sugar alcohol of formula (4):wherein,R isn being an integer in the range of 1 to 12. 4. The sugar alcohol of claim 3, which is an alditol having the skeleton of sorbitol, mannitol, galactitol, or a salt thereof. 5. A composition for delivering a biologically active molecule across a biological membrane into a cell, wherein the biologically active molecule is selected from the group consisting of a peptide, a nucleic acid, a therapeutic molecule, and a diagnostic agent, comprising the compound of formula (1) or formula (4):wherein, R1, R2 and R3 are the same as defined in claim 1, orwherein, R1 is the same as defined in claim 3,and the biologically active agent. 6. The composition of claim 5, wherein the biologically active molecule is an organic compound having a molecular weight ranging from 100 to 1500 g/mol. 7. The composition of claim 5, wherein the biologically active molecule is a peptide and a nucleic acid. 8. The composition of claim 5, wherein the compound of formula 1 forms a conjugate with the biologically active molecule through a covalent bond. 9. The composition of claim 5, wherein the compound of formula 1 forms an ionic complex with the biologically active molecule through ionic bonds. 10. A method for delivering a biologically active molecule across a biological membrane into a cell, wherein the biologically active molecule is one selected from the group consisting of a peptide, a nucleic acid, a therapeutic molecule, and a diagnostic agent, which comprises the step of employing a compound selected from the group consisting of the compounds of formula (1) or (4) as a molecular transporter:wherein, R1, R2 and R3 are the same as defined in claim 1, orWherein, R1 is the same as defined in claim 3,for delivery with the biologically active molecule. | Postech Foundation Postech Academy-Industry Foundation | A01N031/00 A01N033/18 A01N033/24 A61K031/045 A61K031/04 A61B005/055
US07846105 | B2 | 2010-12-07 | Non-invasive blood pressure measurement safety protection method and apparatus | 2009-07-22 | The present invention discloses a non-invasive blood pressure measurement apparatus and a safety protection method. In addition to a main pressure measurement circuit and a main microprocessor circuit, the measurement apparatus also includes an independently disposed assist pressure measurement circuit and an independently disposed assist microprocessor circuit. In normal measurement, the assist microprocessor circuit periodically samples a cuff pressure via the assist pressure measurement circuit, and compares the measured cuff pressure with a specified overpressure protection value, and if the cuff pressure exceeds the specified overpressure protection value, the assist microprocessor circuit outputs a control signal to open a deflation valve until the pressure falls to below the safety pressure. Compared with the prior art, the method for overpressure protection of the present invention is more direct and has higher accuracy of decision, leading to higher reliability of the overpressure protection and higher safety of the measurement apparatus. | 1. A non-invasive blood pressure measurement apparatus, comprising: a main measurement circuit; anda safety protection circuit comprising at least one overpressure protection value and at least one pressurization timing value;wherein the main measurement circuit comprises: a main microprocessor circuit for receiving a first cuff pressure signal of a cuff and calculating a pressure value of the cuff;an air pump control circuit for receiving a first control signal from the main microprocessor circuit for controlling operation of an air pump based on the first cuff pressure signal; anda deflation valve control circuit for receiving a second control signal from the main microprocessor circuit for controlling operation of a deflation valve based on the first cuff pressure signal; andwherein the safety protection circuit comprises an assist microprocessor circuit for collecting a second cuff pressure signal of the cuff in real-time and determining a pressurization time under various modes of operation, for outputting a third control signal to the deflation valve control circuit depending on the second cuff pressure signal received from the assist pressure measurement circuit and the determined pressurization time, and for controlling operation of the deflation valve via the deflation valve control circuit based on a combination of the second cuff pressure signal and the determined pressurization time for a selected mode of operation;wherein the assist microprocessor circuit comprises: a signal input connected to a watchdog signal output of the main microprocessor circuit; anda control signal output connected to a reset signal input of the main microprocessor circuit;wherein when a program of the main microprocessor is dead, is running out of control, or fails to respond, the assist microprocessor circuit forcibly resets the main microprocessor by outputting a reset signal via the control signal output, and after being reset, the main microprocessor circuit outputs the second control signal to open the deflation valve and the first control signal to turn off the air pump. 2. The non-invasive blood pressure measurement apparatus of claim 1, wherein the assist microprocessor circuit comprises an air pump state monitoring port connected to an air pump control signal input of the air pump control circuit. 3. The non-invasive blood pressure measurement apparatus of claim 1, wherein the assist microprocessor circuit comprises a voltage sample signal input connected to a power supply for monitoring a voltage of the power supply. 4. The non-invasive blood pressure measurement apparatus of claim 1, further comprising a key module connected to the main microprocessor circuit and the assist microprocessor circuit, respectively, for inputting commands to the main microprocessor circuit and the assist microprocessor circuit. 5. The non-invasive blood pressure measurement apparatus of claim 1, further comprising a display module connected to the main microprocess... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/00
US07857759 | B2 | 2010-12-28 | Extraction device, analyzer, extraction method, and analysis method | 2004-03-25 | Extraction device for extracting an analyte from living tissue through skin are described that include: (a) a first electrode part having a contact area with the skin of less than about 50 mm2; (b) a through-current electrode part; and (c) a power supply part for supplying electrical energy to the first electrode part and the through-current electrode part, and for extracting an analyte in the first electrode part. Analyzers, extraction methods, and analysis methods are also described. | 1. An extraction device for extracting an analyte in living tissue through skin, the device comprising: a first path-forming electrode part;a first extraction electrode part for extracting an analyte;a through-current electrode part; anda power supply part for supplying electrical energy to the first path-forming electrode part, the first extraction electrode part, and the through-current electrode part, for forming analyte transmission paths in the skin for the passage of the analyte, and for extracting the analyte at the first extraction electrode part;wherein the first path-forming electrode part comprises a first path-forming electrode connected to the power supply part, and a first chamber comprising purified water/ion-conductive material, wherein the purified water/ion-conductive material contacts the first path-forming electrode;wherein the first chamber is configured such that the purified water/ion-conductive material; is configured to contact the skin over an area of less than 25 mm2; andwherein, during formation of the analyte transmission paths, the first path-forming electrode part is connected to the power supply part; andwherein, during analyte extraction, the first path-forming electrode part is disconnected from the power supply part and the first extraction electrode part is connected to the power supply part. 2. The device of claim 1, wherein the power supply part supplies a current of less than about 300 μA. 3. The device of claim 1, wherein the purified water/ion-conductive material is configured to contact the skin-over an area of between 0.01 and 25 mm2. 4. The device of claim 1, wherein the power supply part comprises a constant-current power supply. 5. The device of claim 1, wherein the power supply part comprises a constant-voltage power supply. 6. The device of claim 1, wherein the power supply part outputs a voltage of less than about 10 V. 7. The device of claim 1, further comprising an extraction accelerator part for promoting the extraction of the analyte. 8. The device of claim 7, wherein the extraction accelerator part comprises an ultrasonic irradiation part for irradiating the skin with ultrasonic waves. 9. The device of claim 1, wherein the analyte is glucose. 10. An analyzer for analyzing an analyte extracted through skin, the analyzer comprising: the extraction device of claim 1;an assay part for assaying the analyte extracted in the first electrode part, and for outputting a signal corresponding to an amount of the analyte;an analysis part for analyzing the signal output by the assay part to obtain an analysis result; andan output part for outputting the analysis result obtained by the analysis part. 11. The device of claim 1, wherein the contact area is between 0.01 and 25 mm2. 12. The device of claim 1 further comprising: a second path-forming electrode part; anda second extraction electrode part for extracting an analyte;wherein the power supply part comprises:a first power supply for supplying... | Sysmex Corporation | A61B005/00
US08216130 | B2 | 2012-07-10 | Device, system and method for selective activation of in vivo sensors | 2010-08-11 | A device, system and method for selectively activating or altering the operational mode of an autonomous in vivo device in response to in vivo conditions. They system may include an in vivo sensing device with a condition tester, and a controller. The in vivo sensing device may be in communications with an external receiver. The condition tester may include at least one layer of a dissolvable material coated on at least a portion of the autonomous device. The layer may dissolve when exposed to a specific material of a specified site along the GI tract. A sensor may be exposed when the layer dissolves or a switch autonomously activated with the layer dissolves. | 1. A system for in-vivo sensing at a specified site along the GI tract, the system comprising: an autonomous in-vivo capsule comprising:a condition tester configured to test an in-vivo condition, the condition tester comprising at least one layer of dissolvable material coated on at least a portion of the capsule, wherein the at least one layer dissolves when exposed to a specified material of a specified site along the GI tract; anda component selected from the group consisting of: a sensor which is exposed when the layer dissolves and a switch which is activated when the layer dissolves;wherein said capsule further comprises a controller configured to delay issuing operational mode change orders until more than one electrical signal from said condition tester has been received. 2. The system according to claim 1, wherein the at least one layer comprises pH sensitive dissolvable material. 3. The system according to claim 1, comprising a second sensor and a second layer of dissolvable material, wherein the at least one layer of dissolvable material coats the first sensor, and the second layer of dissolvable material coats the second sensor. 4. The system according to claim 3, wherein the first layer dissolves when exposed to a different material than the material in response to which the second layer dissolves. 5. The system according to claim 3, wherein the first layer dissolves at a different specified site along the GI tract than the specified site at which the second layer dissolves. 6. The system according to claim 3, wherein the first and second layers comprise different substances that dissolve at different rates. 7. The system according to claim 1, comprising an image sensor, wherein said component is a switch, and the switch turns on the image sensor. 8. The system according to claim 1, wherein the system comprises at least one component selected from the group consisting of an encapsulated drug compartment and a sampling inlet, the component covered by the at least one layer of dissolvable material. 9. A method for in-vivo sensing at a specified site along the GI tract, the method comprising: testing an in-vivo condition using a condition tester, the condition tester comprising at least one layer of dissolvable material coated on at least a portion of a capsule;dissolving the at least one layer by exposing the layer to a specific material of a specified site along the GI tract;exposing a sensor when the layer dissolves; anddelaying the issuance of operational mode change orders until more than one electrical signal from said condition tester has been received. 10. The method according to claim 9, wherein dissolving the at least one layer comprises exposing the layer to an environment with a specific pH. 11. The method according to claim 9, comprising coating the first sensor with the at least one layer of dissolvable material and coating a second sensor with a second layer of dissolvable material. 12. The method according to ... | Given Imaging Ltd. | A61B001/04
US07993280 | B2 | 2011-08-09 | Rapid shallow breathing detection for use in congestive heart failure status determination | 2010-07-30 | Systems and methods involve use of a medical device comprising sensing circuitry. One or more respiratory parameters are detected using the device. Patient baseline weight is provided, and an output signal indicative of a patient's congestive heart failure status is generated based on a change in the one or more respiratory parameters and a change in the patient's measured weight or predicted weight relative to the patient baseline weight. The respiratory parameters may include one or more of respiration rate, relative tidal volume, an index indicative of rapid shallow breathing by the patient, an index derived by computing a respiration rate and a tidal volume for each patient breath, and an index indicative of dyspnea, for example. | 1. A system, comprising: a medical device comprising sensing circuitry;detection circuitry disposed in the medical device and coupled to the sensing circuitry, the detection circuitry configured to detect one or more respiratory parameters;a memory configured to store a baseline weight of a patient; anda processor coupled to the memory and the detection circuitry, the processor configured to compute a change in the one or more respiratory parameters relative to a first threshold, compute a prediction of change in patient weight based on the computed change in the one or more respiratory parameters, and generate an output signal in response to the patient weight change prediction exceeding a second threshold, the processor further configured to update the first and second thresholds in response to the change in the one or more respiratory parameters exceeding the first threshold and the patient weight change prediction exceeding the second threshold. 2. The system of claim 1, wherein the change in the one or more respiratory parameters represents a cumulative change or slope of change in the one or more respiratory parameters. 3. The system of claim 1, wherein the processor is configured to compute the patient weight change prediction by deriving coefficients based on a relationship between predicted patient weight change relative to change in the one or more respiratory parameters. 4. The system of claim 3, wherein the processor is configured to derive the coefficients from a linear regression of the patient weight change prediction relative to the change in the one or more respiratory parameters. 5. The system of claim 1, wherein the one or more respiratory parameters comprises respiration rate, and wherein the processor is configured to generate a baseline relationship between the patient weight change prediction relative to change in the respiration rate for the patient, and update the baseline relationship in response to a change in therapy delivered to the patient or a change in patient condition. 6. The system of claim 1, wherein the medical device is at least partially implantable in the patient, and the processor configured to generate the output signal comprises a processor of the at least partially implantable medical device. 7. The system of claim 1, wherein the medical device is configured for implantation in the patient, and the processor is disposed in the implantable medical device. 8. The system of claim 1, wherein the processor configured to generate the output signal comprises at least a processor of a patient-external system. 9. The system of claim 1, wherein at least the sensing circuitry comprises one or more implantable sensing components. 10. The system of claim 1, wherein the one or more respiratory parameters comprises respiration rate. 11. The system of claim 1, wherein the one or more respiratory parameters comprises an index indicative of rapid shallow breathing by the patient. 12. The system of claim 11, wh... | Cardiac Pacemakers, Inc. | A61B005/08
US08920341 | B2 | 2014-12-30 | Test strip device and method for analyzing a body fluid | 2010-04-09 | The invention concerns a test tape device for analyzing a body fluid having a carrier tape, a plurality of lancing elements arranged on the carrier tape which are provided with a tip that can puncture a body part and a collecting structure that takes up the body fluid obtained during the puncture, and test fields mounted on the carrier tape each being associated with a lancing element and can have body fluid applied thereto. According to the invention it is proposed that the lancing elements are each movably attached to the carrier tape by a coupling member and that a used lancing element can be brought into contact with a test field by a transfer movement from a usage position distant from a test field into a contact position so that body fluid can be transferred from the collecting structure onto the test field. | 1. A test tape device for analyzing a body fluid for blood sugar tests having a carrier tape that can be wound onto or is wound onto a spool, a plurality of lancing elements arranged on the carrier tape which are provided with a tip that can puncture a body part and a collecting structure that takes up the body fluid obtained by the puncture, and test fields mounted on the carrier tape each being associated with a lancing element and can have body fluid applied thereto, characterized in that the lancing elements are each movably attached to the carrier tape by a coupling member, and that a used lancing element can be brought into fluidic connection with a test field by a transfer movement executed by the lancing element from a usage position distant from a test field into a contact position, so that body fluid is transferred directly from the collecting structure onto the test field, wherein each lancing element is covered in a sterile manner by a cover foil connected to the carrier tape to form a foil bag around each lancing element, wherein the lancing element is configured to puncture the cover foil, characterized in that the test fields have a reagent layer which is designed to detect an analyte in the body fluid. 2. The test tape device according to claim 1, characterized in that the orientation of the lancing elements relative to the carrier tape can be reversed by the transfer movement. 3. The test tape device according to claim 1, characterized in that the elongate lancing elements can be oriented in the longitudinal direction of the carrier tape at least in the contact position. 4. The test tape device according to claim 1, characterized in that the lancing elements can be contacted with an associated test field by a pivoting movement and in particular a folding movement. 5. The test tape device according to claim 1, characterized in that the lancing elements remain permanently connected to the tape by means of the coupling member during the transfer movement. 6. The test tape device according to claim 1, characterized in that the lancing elements are hinged on the carrier tape in a proximal region that is spaced apart from the tip by means of a joint or hinge as a coupling member. 7. The test tape device according to claim 1, characterized in that the lancing elements are pivotally connected to the carrier tape by means of a bending joint and in particular a film joint. 8. The test tape device according to claim 1, characterized in that the transfer movement of a lancing element can be triggered by advancing the carrier tape. 9. The test tape device according to claim 1, characterized in that the carrier tape is guided laterally past a retaining member and that the retaining member forms a stop obstacle for a lancing element which starts up during the tape transport preferably for folding down the lancing element. 10. The test tape device according to claim 9, characterized in that the retaining member is located in the area of... | Roche Diagnostics Operations, Inc. | A61B005/00 A61B005/151 A61B005/15
US08374675 | B2 | 2013-02-12 | Magnetic resonance angiography with flow-compensated and flow-sensitive imaging | 2010-01-07 | In a magnetic resonance angiography method with flow-compensated and flow-sensitive imaging and a magnetic resonance apparatus for implementing such a method, a first MR data set of the examination region is acquired with an imaging sequence in which vessels in the examination region are shown with high signal intensity, a second MR data set of the examination region with an imaging sequence in which the vessels in the examination region are shown with low signal intensity, and the angiographic magnetic resonance image is calculated in a processor by taking the difference of the first and second data set. The first data set is acquired with an imaging sequence with reduced flow sensitivity and the second data set is acquired with an imaging sequence with an increased flow sensitivity compared to the initial imaging sequence. | 1. A method to generate a magnetic resonance angiography image of an examination region of a subject, comprising the steps of: operating a magnetic resonance data acquisition unit in which an examination subject is located according to an imaging sequence with reduced flow sensitivity to acquire a first magnetic resonance dataset from an examination region of the subject in which vessels in the examination region are shown with high signal intensity;operating said magnetic resonance data acquisition unit according to an imaging sequence with a flow sensitivity that is increased compared to the reduced flow sensitivity with which the first magnetic resonance data set was acquired, to obtain a second set of magnetic resonance data from the examination region in which the vessels in the examination region are shown with a low signal intensity;synchronizing acquisition of each of said first and second magnetic resonance data sets with a cardiac cycle of the examination subject;acquiring all of said first magnetic resonance data set during a diastole of said cardiac cycle, in which blood flow in the vessels in the examination region is slow, and acquiring all of said second magnetic resonance data set during systole of said cardiac cycle, in which blood flow in the vessels in the examination region is faster compared to blood flow during acquisition of said first magnetic resonance data set; andin a processor, reconstructing an angiographic magnetic resonance image of the examination region by mathematically combining said first and second magnetic resonance data sets. 2. A method as claimed in claim 1 comprising mathematically combining said first and second magnetic resonance data sets by generating a single, temporarily inverse data set from said first and second magnetic resonance data sets and transforming said single, temporarily inverse data set into image space. 3. A method as claimed in claim 1 comprising mathematically combining said first and second magnetic resonance data sets by subtracting one of said first and second magnetic resonance data sets from the other of said first and second magnetic resonance data sets. 4. A method as claimed in claim 1 comprising acquiring said first and second magnetic resonance data sets respectively with an imaging sequence that is identical except for said flow sensitivity. 5. A method as claimed in claim 4 comprising employing, as said same imaging sequence for acquiring each of said first and second magnetic resonance data sets, an imaging sequence having a gradient field exhibiting a gradient moment that is activated between successively radiated RF pulses, and reducing the flow sensitivity during acquisition of said first magnetic resonance data set by reducing said gradient moment, at least in a primary direction of said flow, when acquiring said first magnetic resonance data set compared to when acquiring said second magnetic resonance data set. 6. A method as claimed in claim 5 comprising red... | Siemens Medical Solutions USA, Inc. Siemens Aktiengesellschaft | A61B005/055
US08900155 | B2 | 2014-12-02 | Non-invasive blood pressure monitoring system | 2007-11-09 | A non-invasive blood pressure monitoring system, includes an air channel having one or more inlet and exhaust ports to ambient air. An electronically controlled pump inflates and deflates a cuff applied to a patient limb with air via the air channel. A water tight sealed housing containing electronic circuitry for processing signals used in deriving a measurement of patient blood pressure in conjunction with inflating and deflating the cuff. The water tight sealed housing also is sealed from the air channel. | 1. A non-invasive blood pressure monitoring system, comprising: an air channel having one or more inlet and exhaust ports to ambient air;an electronically controlled pump for inflating a cuff applied to a patient limb with air acquired through the one or more inlet and exhaust ports of said air channel; anda water tight sealed housing containing the electronically controlled pump and electronic circuitry inside the housing for processing signals used in deriving a measurement of patient blood pressure in conjunction with inflating and deflating said cuff, the inside of said water tight sealed housing also being sealed from said air channel preventing air and liquid acquired or expelled through said air channel from leaking into the inside of the housing and contacting said electronic circuitry;wherein the one or more inlet and exhaust ports of said air channel and said electronically controlled pump are isolated from said electronic circuitry inside said water tight sealed housing and the seal between one or more inlet and exhaust ports of said air channel and said water tight sealed housing results in said air channel providing an exclusive pathway for air and liquid being passed to and expelled from the electronically controlled pump without air or liquid leakage into the housing. 2. The system according to claim 1, wherein said air channel is capable of being flushed with at least one of (a) air and (b) liquid. 3. The system according to claim 1, wherein said system is capable of achieving an IPX7 international protection rating for said housing. 4. The system according to claim 1, wherein said air channel exclusively acquires and expels air via said air channel substantially without air leakage into said housing. 5. The system according to claim 1, wherein said air channel is a flattened channel having one cross-sectional dimension substantially greater than an other cross-sectional dimension. 6. The system according to claim 1, wherein said water tight sealed housing contains electronic circuitry for controlling said pump. 7. The system according to claim 1, wherein said water tight sealed housing is sealed from said air channel using a gasket. 8. The system according to claim 1, wherein said air channel is selectively removeable. 9. A non-invasive blood pressure monitoring system, comprising: an electronically controlled pump for inflating and deflating a cuff applied to a patient limb with air via an air channel having one or more inlet and exhaust ports to ambient air;a water tight sealed housing containing the electronically controlled pump and electronic circuitry inside the housing for processing signals used in deriving a measurement of patient blood pressure in conjunction with inflating and deflating said cuff, the inside of said water tight sealed housing also being sealed from said air channel preventing air and liquid acquired or expelled through said air channel from leaking into the inside of the housing and contacting ... | Draeger Medical Systems, Inc. | A61B005/02 A61B005/022 A61B005/021
US20100331735 | A1 | 2010-12-30 | WIRELESS POWER MODULATION TELEMETRY FOR MEASURING A PARAMETER OF THE MUSCULAR-SKELETAL SYSTEM | 2010-06-29 | A sensing insert device (100) is disclosed for measuring a parameter of the muscular-skeletal system. The sensing insert device (100) can be temporary or permanent. Used intra-operatively, the sensing insert device (100) comprises an insert dock 202 and a sensing module 200. The sensing module (200) is a self-contained encapsulated measurement device having at least one contacting surface that couples to the muscular-skeletal system. The sensing module (200) comprises one or more sensors (303), electronic circuitry (307), a capacitor (1410), and communication circuitry (320). The electronic circuitry (307) operatively couples to the one or more sensors (303) to measure the parameter. A transmitter (309) transmits parameter measurements. The capacitor (1410) can store sufficient charge to power the sensing module (200) for a total joint reconstruction operation. Energy is wirelessly coupled to an induction coil (1404) thereby generating a power signal that provides charge to the capacitor (1410). | 1. A system to measure a parameter of the muscular-skeletal system comprising: a sensing module including: one or more sensors for coupling to the muscular-skeletal system;electronic circuitry operatively coupled to the sensors; anda capacitor coupled to power the sensing module for a predetermined time. 2. The system of claim 1 where the capacitor is charge wirelessly. 3. The system of claim 2 where the sensing module further includes: an induction coil;a rectifier coupled to the induction coil; anda regulator coupled to the rectifier to regulate a voltage output by the capacitor. 4. The system of claim 3 further including a wireless energy source placed in proximity to the sensing module where the wireless energy source electromagnetically couples to the induction coil and where a power signal is generated in the induction coil that charges the capacitor. 5. The system of claim 4 where the wireless energy source is placed within 20 centimeters of the sensing module. 6. The system of claim 4 where data is modulated on the power signal, where the data is demodulated from the power signal, and where the data is provided to the electronic circuitry. 7. The system of claim 1 where the sensing module is stored in a zero charge state prior to use. 8. The system of claim 7 where the capacitor is fully charged in less than 3 minutes from the zero charge state and where the capacitor has a solid dielectric. 9. The system of claim 8 where the sensing module is in a final insert to measure the parameter of the muscular-skeletal system where the final insert has at least one bearing surface for allowing movement of at least one bone of the muscular-skeletal system. 10. The system of claim 8 where the sensing module is used intra-operatively with a dock as a trial insert to measure the parameter of the muscular-skeletal system. 11. A method of enabling a sensing module for measuring a parameter of the muscular-skeletal system comprising the steps of: placing a wireless energy source in proximity to the sensing module;coupling electromagnetically to an induction coil in the sensing module;generating a power signal in the induction coil comprising energy from the wireless energy source;charging a capacitor with the power signal;enabling the sensing module when the capacitor has stored sufficient charge to power the sensing module for a predetermined time; andmeasuring at least one parameter of the muscular-skeletal system. 12. The method of claim 11 further including the steps of: rectifying the power signal; andregulating a voltage on the capacitor. 13. The method of claim 12 where the step of regulating further includes discharging the capacitor when the voltage on the capacitor exceeds a predetermined threshold. 14. The method of claim 13 further including the steps of: demodulating input data from the power signal;coupling the input data to operational circuitry of the sensing module; andcontrolling at least one operation of the sensing module... | OrthoSensor | A61B005/103
US20100298693 | A1 | 2010-11-25 | BREAST COMPRESSION ASSEMBLY FOR USE IN MRI BIOPSY PROCEDURE | 2010-08-05 | A compression assembly is operable to localize a patient's breast. The compression assembly comprises a frame and a plurality of slats. The frame defines a plurality of tracks. The frame is configured to engage a breast localization fixture. The slats are coupled with the frame. Portions of the slats are disposed in the tracks. One or more of the slats are movable relative to the frame to provide adjustable access to a patient's breast engaged by the compression assembly. The slats may be slid and/or rotated relative to the frame. In some versions, the tracks and slats are provided in two sets. Each set lies along a respective plane, with the two planes being parallel to each other. The slats of one set may ratchetingly engage the slats of the other set to restrain slat movement. In some versions, the slats are removable from the frame for increased access. | 1. A compression assembly operable to localize a patient's breast, the compression assembly comprising: (a) a frame, wherein the frame defines a plurality of tracks, wherein the frame is configured to engage a breast localization fixture; and(b) a plurality of slats coupled with the frame, wherein portions of the slats are disposed in the tracks of the frame, wherein one or more of the slats are movable relative to the frame to provide adjustable access to a patient's breast engaged by the compression assembly. 2. The compression assembly of claim 1, wherein the frame is substantially rectangular. 3. The compression assembly of claim 1, wherein one or more of the slats are slidable relative to the frame. 4. The compression assembly of claim 1, wherein a first slat of the slats includes a ratcheting surface configured to restrain movement of the first slat. 5. The compression assembly of claim 4, wherein a second slat of the slats includes a ratcheting surface configured to restrain movement of the second slat. 6. The compression assembly of claim 5, wherein the first slat and the second slat are adjacent to each other. 7. The compression assembly of claim 6, wherein the ratcheting surface of the first slat and the ratcheting surface of the second slat are configured and positioned to selectively engage each other. 8. The compression assembly of claim 1, wherein the plurality of tracks comprise a first set of tracks and a second set of tracks, wherein the plurality of slats comprise a first set of slats associated with the first set of tracks and a second set of slats associated with the second set of tracks. 9. The compression assembly of claim 8, wherein a first plane passes through the first set of tracks and the first set of slats, wherein a second plane passes through the second set of tracks and the second set of slats. 10. The compression assembly of claim 9, wherein the first plane and the second plane are parallel to each other. 11. The compression assembly of claim 10, wherein the frame defines horizontal and vertical dimensions, wherein the first set of slats are oriented to extend along the horizontal dimension, wherein the first set of slats are slidable along the vertical dimension. 12. The compression assembly of claim 11, wherein each slat of the first set of slats is slidable independently relative to the other slats of the first set of slats. 13. The compression assembly of claim 11, wherein the second set of slats extend along the horizontal dimension, wherein the second set of slats are slidable along the vertical dimension. 14. The compression assembly of claim 1, wherein the tracks and the portions of the slats that are disposed in the tracks together form a tongue and groove configuration. 15. The compression assembly of claim 1, wherein the slats are substantially rectangular. 16. The compression assembly of claim 1, wherein the slats are rotatable relative to the frame. 17. The compression assembly of claim... | DEVICOR MEDICAL PRODUCTS, INC. | A61B005/055
US07860553 | B2 | 2010-12-28 | Two-stage calibration of medical probes | 2006-02-09 | A probe for insertion into the body of a subject includes a sensor, a first microcircuit, which stores first calibration data with respect to the sensor, and a first connector at the proximal end of the probe. A probe adapter includes a second connector, which mates with the first connector, a signal processing circuit, which is processes the sensor signal, and a second microcircuit, which stores second calibration data with respect to the signal processing circuit. A microcontroller in the adapter receives the first and second calibration data and computes combined calibration data. The adapter includes a third connector, which mates with a fourth connector on a console. The console includes signal analysis circuitry, which analyzes the processed signal using the combined calibration data provided by the probe adapter. | 1. Medical apparatus comprising: a probe, having a proximal end and a distal end, which is adapted for insertion into the body of a subject, said probe comprising: a sensor, which outputs a sensor signal;a first microcircuit, which stores first calibration data with respect to the sensor; anda first connector at the proximal end of the probe, electrically coupled at least to the sensor;a probe adapter configured to be detachably connected to the probe, said probe adapter comprising: a second connector, which is arranged to mate with the first connector of the probe;a signal processing circuit, which is coupled to process the sensor signal so as to output a processed signal;a second microcircuit, which stores second calibration data with respect to the signal processing circuit;a microcontroller, which is arranged to receive the first calibration data from the first microcircuit of the probe and second calibration data from second microcircuit of the probe adapter, respectively, and to compute combined calibration data based on the first calibration data and the second calibration data; anda third connector, electrically coupled at least to the signal processing circuit; anda console, comprising: a fourth connector, which is arranged to mate with the third connector; andsignal analysis circuitry, which is coupled to receive at least the processed signal from the fourth connector and is arranged to analyze the processed signal using the combined calibration data provided by the probe adapter. 2. The apparatus according to claim 1, wherein the sensor comprises a position sensor, and wherein the signal analysis circuitry is operative to determine coordinates of the distal end of the probe by analyzing the processed signal. 3. The apparatus according to claim 2, wherein the position sensor is operative to generate the sensor signal responsively to a magnetic field applied externally to the body. 4. The apparatus according to claim 1, wherein the probe comprises a catheter for insertion into a heart of the subject. 5. The apparatus according to claim 1, wherein the probe is a first type of probe, and the processed signal is a first processed signal, and wherein the console is further operable in conjunction with a second type of probe, which is adapted to mate with the fourth connector and to convey at least a second processed signal to the fourth connector, and which includes a memory circuit, which stores third calibration data in a predetermined address space, which is accessed by the signal analysis circuitry in analyzing the second processed signal, andwherein the microcontroller is operative to place the combined calibration data in the predetermined address space for reading by the processing circuitry. 6. The apparatus according to claim 5, wherein the console is operable in conjunction with both of the first and second types of probe without hardware or software modification according to probe type. 7. The apparatus according to claim 1,... | Biosense Webster, Inc. | A61B005/05
US07853318 | B2 | 2010-12-14 | Cardiac sensing by implantable medical devices during magnetic resonance imaging | 2007-03-14 | Embodiments of the invention can include a method for reducing MRI interference from a physiological electrical signal received in an implantable medical device. The method can include the steps of amplifying the physiological electrical signal with a high bandwidth amplifier and sampling the amplified physiological electrical signal at a sampling frequency of at least 8 kHz to obtain a first high-frequency sequence of samples. The method can further includes the steps of processing the first high-frequency sequence of samples to identify a signal artifact that is characteristic of MRI interference and creating a second high-frequency sequence of samples by reducing the signal artifact characteristic of MRI interference from the first sequence of samples. Embodiments can also include a method of determining a heart rate from a physiological electrical signal received in an implantable medical device. Embodiments can also includes an implantable medical device that receives and processes a physiological electrical signal. Other embodiments are also described herein. | 1. A method of determining a heart rate from a physiological electrical signal received in an implantable medical device, the method comprising: (i) receiving an input associated with the presence of an environment that tends to cause corruption of a physiological electrical signal; and(ii) initiating a mode of operation for reducing corruption of the physiological electrical signal, the mode of operation comprising (a) providing a high and a low bandwidth amplifier in an implantable medical device, wherein the high bandwidth amplifier has a bandwidth of at least 4 kHz and the low bandwidth amplifier has a bandwidth equal no greater than 1 kHz;(b) providing a switch configured to select between the high bandwidth and low bandwidth amplifiers, wherein the switch is switched to utilize the high bandwidth amplifier if corruption of the physiological electrical signal is detected;(c) amplifying the physiological electrical signal with the selected amplifier;(d) sampling the amplified physiological electrical signal at a sampling frequency of at least 8 kHz to obtain a first sequence of samples;(e) processing the first sequence of samples to identify a superposed signal artifact characteristic of MRI interference;(f) reducing the superposed signal artifact from the first sequence of samples to form a second sequence of samples; and(g) processing the second sequence of samples to identify the timing of heart contractions. 2. The method of claim 1, where the environment that tends to cause corruption of a physiological electrical signal comprises magnetic resonance imaging. 3. The method of claim 1, where the input comprises receiving a signal by telemetry. 4. The method of claim 3, where the telemetric signal is initiated by a person who is aware of the presence of an environment that tends to cause corruption of an electrical signal. 5. The method of claim 1, where the input comprises receiving a signal from a magnetic field sensor. 6. The method of claim 1, where the input comprises receiving a signal from a sensor configured to measure induced currents within a conductive loop. 7. A method for reducing MRI interference from a physiological electrical signal received in an implantable medical device, the method comprising: providing a high and a low bandwidth amplifier in an implantable medical device, wherein the high bandwidth amplifier has a bandwidth of at least 4 kHz and the low bandwidth amplifier has a bandwidth equal no greater than 1 kHz;providing a switch configured to select between the high bandwidth and low bandwidth amplifiers, wherein the switch is switched to utilize the high bandwidth amplifier if corruption of the physiological electrical signal is detected;amplifying the physiological electrical signal with the selected amplifier;sampling the amplified physiological electrical signal at a sampling frequency of at least 8 kHz to obtain a first sequence of samples;processing the first sequence of samples to identify a signal ar... | Cardiac Pacemakers, Inc. | A61B005/04
US20100298663 | A1 | 2010-11-25 | SYSTEM AND METHOD FOR DETECTION AND TREATMENT OF IRREGULAR METABOLIC FUNCTION | 2009-05-20 | A system and method for detecting and treating irregular metabolic function. The systems and methods generally include an implantable medical device that measures various physiological parameters. Such physiological parameters generally provide an indication of the patient's metabolic rate, respiratory rate, and activity level. For example, the medical device may measure the patient's relative movement, or may measure the patient's oxygen consumption. These measures are then recorded, and the recorded data is used to determine the normal relationship between metabolic rate and other physiological parameters for that patient. The baseline relationship is then used to determine if the patient experiences an increase or decrease in metabolic rate that is not explained by changes in other physiological parameters. Further testing, and treatment if necessary, can be performed if the physiological parameters indicate irregular metabolic function. | 1. A method for detecting irregularities in metabolic function, comprising: implanting a medical device into a patient, wherein the medical device measures one or more physiological parameters indicative of metabolic rate and activity level;collecting data sets, wherein each data set includes a patient's metabolic rate and corresponding activity level;calculating a normal range for metabolic rate based on corresponding activity level;monitoring subsequent data sets;providing an indication if metabolic rate in a subsequent data set falls outside of the normal range for a corresponding activity level. 2. The method of claim 1, wherein said indication is an audible sound that notifies the patient. 3. The method of claim 1, wherein said indication is a vibration that notifies the patient. 4. The method of claim 1, wherein said indication notifies a remote healthcare provider. 5. The method of claim 1, wherein the physiological parameter indicative of metabolic rate includes a measure of the patient's mixed venous saturation, cardiac output, and arterial oxygen saturation. 6. The method of claim 5, wherein the metabolic rate is calculated using a Fick equation. 7. The method of claim 1, further comprising: performing diagnostic testing on the patient to determine if the patient is experiencing irregular metabolic function if the patient's metabolic rate falls outside of the normal range for a measured activity level. 8. The method of claim 7, further comprising: preparing a treatment plan if the diagnostic testing indicates that the patient is experiencing irregular metabolic function. 9. The method of claim 8, further comprising: adjusting the patient's medication to resolve the irregular metabolic function. 10. The method of claim 8, further comprising: prescribing a thyroid hormone to treat the irregular metabolic function. 11. The method of claim 8, further comprising: performing radioactive iodine therapy on the patient to treat the irregular metabolic function. 12. The method of claim 8, further comprising: positioning a lead in an area of the patient's brain to pace the patient's autonomic nervous system. 13. The method of claim 8, further comprising: positioning a lead in an area of the patient's brain to control release of hormones. 14. The method of claim 1, further comprising: measuring data indicative of the patient's respiratory rate. 15. The method of claim 14, wherein the step of calculating a normal range for metabolic rate further comprises determining whether a change in respiratory rate is due to a change in the patient's activity level. 16. The method of claim 1, wherein the step of calculating a normal range is conducted by a processor within the medical device. 17. The method of claim 1, wherein the step of calculating a normal range is conducted by an external receiver. 18. A implantable medical device that detects irregularities in metabolic function, comprising: one or more sensors adapted to measure one or mo... | PACESETTER, INC. | A61B005/00
US20100280331 | A1 | 2010-11-04 | Method and apparatus for non-invasively estimating body core temperature | 2007-11-28 | Non-invasive methods and devices are disclosed to derive estimates of body Core Temperature from external sensors that provide electrocardiograph (ECG) data, and Mean Skin Temperature data. The ECG and Mean Skin Temperature are input to a model that provides estimates of Core Temperature temperature. The model is derived regressively from ECG, Mean Skin Temperature and Core Temperature data obtained from a number of test subjects. A monitoring device may be used, for example, to trigger an alarm, display Core Temperature data to the device wearer or to a remote monitoring station, or to activate an emergency temperature control system or device. | 1. A method of estimating body Core Temperature, comprising: determining Heart Rate Variability;determining Mean Skin Temperature; andcalculating body Core Temperature based on data comprising the determined Heart Rate Variability and Mean Skin Temperature. 2. The method of estimating body Core Temperature according to claim 1, wherein estimating body Core Temperature based on the Heart Rate Variability and Mean Skin Temperature comprises employing an empirically derived model. 3. The method of estimating body Core Temperature according to claim 2, wherein the model comprises one or more parameters and the one or more parameters are derived by a regression analysis performed on data collected from test subjects. 4. The method of estimating body Core Temperature according to claim 3 wherein the data collected from test subjects comprises Heart Rate Variability, Mean Skin Temperature and Core Temperature. 5. The method of estimating body Core Temperature according to claim 1, wherein a spectral analysis is performed on the Heart Rate Variability data. 6. The method of estimating body Core Temperature according to claim 5, wherein the spectral analysis comprises extracting a frequency range from the Heart Rate Variability data spectrum. 7. The method of estimating body Core Temperature according to claim 6, wherein the frequency range extracted from the Heart Rate Variability data spectrum comprises a Very Low Frequency range. 8. The method of estimating body Core Temperature according to claim 1 wherein the data for calculating body Core Temperature comprises a Heart Rate. 9. The method of estimating body Core Temperature according to claim 1 wherein the data for calculating body Core Temperature comprises a body size. 10. The method of estimating body Core Temperature according to claim 3 wherein the body size comprises a body height and body weight. 11. The method of estimating body Core Temperature according to claim 5, wherein the model comprises substantially the following calculation: CT=a0+a1*HR+a2*MST+a3*ambient temp+a4*height*weight+a5*HR*MST+a6*HR*ambient temp+a7*HR*height*weight+a8*VLF*MST+a9*VLF*ambient temp+a10*VLF*height*weight+a11*MST*ambient temp+a12*HR*VLF,where CT is Core Temperature, HR is Heart Rate, MST is Mean Skin Temperature, VLF is a very low frequency spectral component of Heart Rate Variability and a0, a1, a2 a12 are model parameters. 12. An apparatus for estimating the Core Temperature of a body, comprising: a non-invasive sensor to detect cardiac data from the body;a non-invasive sensor to detect temperature data from the body,at least one processor programmed to execute instructions to:derive Heart Rate Variability data from the cardiac data and to derive Mean Skin Temperature data from the temperature data detected from the body, and further to derive an estimate of body Core Temperature from data comprising the Heart Rate Variability and Mean Skin Temperature data. 13. The apparatus according to claim 12,... | Department of the Navy | A61B005/00
US20100292583 | A1 | 2010-11-18 | METHOD AND SYSTEM FOR SYNCHRONIZING BLOOD COMPONENT OR TRACE ANALYTE MEASUREMENT WITH HEART PULSE RATE | 2010-05-13 | A system and method for measuring at least one blood component or trace analyte in the blood of a subject, the system including: a heart pulse rate measuring device for measuring a heart pulse rate of the subject, and a blood component or trace analyte measuring device for measuring the at least one blood component or trace analyte, wherein the blood component or trace analyte measuring device is synchronized with the heart pulse rate of the subject to perform a blood component or trace analyte measurement at a point along a measurement interval determined by the heart pulse rate. | 1. A system for measuring at least one blood component or trace analyte in the blood of a subject, the system comprising: a heart pulse rate measuring device for measuring a heart pulse rate of the subject; anda blood component or trace analyte measuring device for measuring the at least one blood component or trace analyte, wherein the blood component or trace analyte measuring device is synchronized with the heart pulse rate of the subject to perform a blood component or trace analyte measurement at a point in a measurement interval determined by the heart pulse rate. 2. The system of claim 1, wherein the heart pulse rate measuring device comprises a pulse oximeter. 3. The system of claim 1, wherein the heart pulse rate measuring device comprises a pressure sensor. 4. The system of claim 1, wherein the heart pulse rate measuring device comprises an electrocardiogram. 5. The system of claim 1, wherein a plurality of blood component or trace analyte measurements are taken at the same point during a plurality of measurement intervals. 6. A method for measuring at least one blood component or trace analyte of a subject, the method comprising: measuring a heart pulse rate of the subject with a heart pulse rate measuring device, andmeasuring the at least one blood component or trace analyte with a blood component or trace analyte measuring device, wherein a blood component or trace analyte measurement is synchronized with the heart pulse rate of the subject at a point along a measurement interval determined by the heart pulse rate. 7. The method of claim 6, wherein the heart pulse rate measuring device comprises a pulse oximeter. 8. The method of claim 6, wherein the heart pulse rate measuring device comprises a pressure sensor. 9. The method of claim 6, wherein the heart pulse rate measuring device comprises an electrocardiogram. 10. The method of claim 6, wherein a plurality of blood component or trace analyte measurements are taken at the same point during a plurality of measurement intervals. 11. An apparatus for measuring at least one blood component or trace analyte in the blood of a subject, the apparatus comprising: a heart pulse rate measuring device for measuring a heart pulse rate of the subject;a probe comprising a blood component or trace analyte measuring device for measuring the at least one blood component or trace analyte, wherein the blood component or trace analyte measuring device is synchronized with the heart pulse rate of the subject to perform a blood component or trace analyte measurement at a point along a measurement interval determined by the heart pulse rate;a mounting support for the probe that defines at least one path of translational movement along a surface of the mounting support; anda biasing element that exerts substantially constant force to provide translational movement of the probe,wherein the probe is translationally movable along the at least one path in a direction away from an initial point, and ... | - | A61B005/0205
US07853307 | B2 | 2010-12-14 | Methods, apparatuses, and systems useful in conducting image guided interventions | 2008-06-26 | Methods, apparatuses, and systems relating to image guided interventions on dynamic tissue. One embodiment is a method that includes creating a dataset that includes images, one of the images depicting a non-tissue internal reference marker, being linked to non-tissue internal reference marker positional information, and being at least 2-dimensional. Another embodiment is a method that includes receiving a position of an instrument reference marker coupled to an instrument; transforming the position into image space using a position of a non-tissue internal reference marker implanted in a patient; and superimposing a representation of the instrument on an image in which the non-tissue internal reference marker appears. Computer readable media that include machine readable instructions for carrying out the steps of the disclosed methods. Apparatuses, such as integrated circuits, configured to carry out the steps of the disclosed methods. Systems that include devices configured to carry out steps of the disclosed methods. | 1. A non-transitory processor-readable medium storing code representing instructions to cause a processor to perform a process, the code comprising code to: create a dataset that includes a plurality of images associated with a body part, at least one image from the plurality of images depicting a non-tissue internal reference marker, the at least one image from the plurality of images being associated with a non-tissue internal reference marker position information, the at least one image from the plurality of images being at least 2-dimensional;calculate for each image in the dataset a dataset vector based on a position of the non-tissue internal reference marker and the position of an external reference marker;receive data associated with a current position of the external reference marker and a current position of the non-tissue internal reference marker;calculate a current vector using the data associated with the current position of the external reference marker and the data associated with a current position of the non-tissue internal reference marker;identify a match dataset vector (MDV), the MDV being the dataset vector most similar to the current vector, the MDV being associated with an image from the dataset;receive data associated with a current position of an instrument reference marker coupled to an instrument;apply a transformation associated with the image associated with the MDV to the current position of the instrument reference marker, thus transforming the current position of the instrument reference marker into image space; andsuperimpose a representation of the instrument on the image associated with the MDV. 2. The non-transitory processor-readable medium of claim 1, wherein the code to create the dataset includes code to create a dataset that includes at least one image taken via at least one of fluoroscopy, CINE fluoroscopy, computed tomography (CT), magnetic resonance imaging (MRI), ultrasound, positron emission tomography (PET), or functional magnetic resonance imaging (fMRI). 3. The processor-readable medium of claim 1, wherein the code to create the dataset includes code to create a dataset that includes at least one 3-dimensional image. 4. The processor-readable medium of claim 1, wherein code to create the dataset includes code to create a dataset that includes a plurality of images including gated information. 5. The processor-readable medium of claim 1, further comprising code to: cause threshold segmentation of each image of the plurality of images, the threshold segmentation allowing the code to create the dataset. 6. The processor-readable medium of claim 1, further comprising code to: receive a periodic human characteristic signal. 7. The processor-readable medium of claim 6, wherein the code to receive the periodic human characteristic signal includes code to receive a human characteristic signal calculated using the position of the non-tissue internal reference marker. 8. The processor-readable medium... | Veran Medical Technologies, Inc. | A61B005/05
US07860543 | B2 | 2010-12-28 | Analyte detection system with reduced sample volume | 2006-08-15 | In certain embodiments, a method samples a body fluid of a patient. The method includes providing a fluid handling system having one or more fluid passageways. The method further includes infusing an infusion fluid by way of a fluid handling system into a patient through one or more fluid passageways. The method further includes obtaining a sample of body fluid by way of a fluid handling system from a patient through one or more fluid passageways. The obtained sample is no more than 5 milliliters in volume. The method further includes analyzing at least an analyzed portion of the obtained sample by way of an analyte detection system operatively associated with the fluid handling system to determine a concentration of at least one analyte. | 1. A method of sampling a body fluid of a patient and reducing sample dilution, said method comprising: providing a fluid handling system configured to provide access to the body fluid of the patient, the fluid handling system comprising a patient end, a junction in fluid communication with the patient end, a first fluid passageway connected to the junction, and a second fluid passageway connected to the junction;providing a body fluid detection sensor configured to detect the presence of substantially undiluted body fluid at a location in the fluid handling system and generate a signal;with said fluid handling system, infusing an infusion fluid into said patient through the patient end;positioning an air bubble at the entrance to the second fluid passageway adjacent the junction between the first and second fluid passageways;with said fluid handling system, drawing at least a portion of a sample of said body fluid from said patient, through the patient end, through the junction, and into the first fluid passageway, wherein the body fluid detection sensor identifies substantially undiluted body fluid at the location in the fluid handling network and generates a signal to facilitate the introduction of substantially undiluted body fluid into the first fluid passageway such that substantially undiluted body fluid is positioned at the junction between the first and second fluid passageways;with said fluid handling system, drawing at least a portion of the sample of said body fluid through the junction and into the second fluid passageway after the air bubble is positioned such that the leading portion of the sample that enters the second fluid passageway immediately following the air bubble is substantially undiluted blood, and wherein the air bubble is transported through the second fluid passageway forward of the sample such that the air bubble provides a barrier to reduce dilution of the leading portion of the sample that may result from contact with other fluids;with said fluid handling system, delivering at least an analysis portion of the sample to an analyte detection system operatively associated with the fluid handling system;with the analyte detection system, analyzing at least an analyzed portion of said sample to determine a concentration of at least one analyte. 2. The method of claim 1, wherein at least a portion of the sample that is drawn into the second fluid passageway is transferred from the first fluid passageway, through the junction, and into the second fluid passageway. 3. A body fluid handling system configured to reduce sample dilution, the system comprising: a fluid handling network configured to provide access to a body fluid in a patient;a body fluid detection sensor configured to detect the presence of body fluid at a location in the fluid handling network and generate a signal;one or more air injectors configured to insert air into the fluid handling network;a pump system coupled to the fluid handling network, the pum... | OptiScan Biomedical Corporation | A61B005/1455
US07844324 | B2 | 2010-11-30 | Measurement of EEG reactivity | 2007-02-14 | The invention relates to a method and apparatus for assessing the reactivity observable in a certain physiological signal, especially the EEG signal, of a comatose subject. In order to obtain an objective and a reliable measure of the reactivity automatically and without the presence of a trained EEG specialist, a valid signal model is constructed for an EEG signal obtained from the subject. A time reference corresponding to a stimulus is applied and further signal data is obtained from the time series, the further signal data being subsequent to the time reference. By employing the further signal data, the method tests whether the signal model remains to be a valid signal model for the EEG signal also after the stimulus, and indicates, based on the test, whether reactivity is present in the physiological signal. | 1. A method for assessing physiological signal reactivity of a subject in a single-trial analysis, the method comprising: obtaining a single continuous physiological signal from the subject;supplying a stimulus to the subject while the physiological signal is being obtained;constructing, based only on a time series of the physiological signal obtained from the subject, a signal model for the physiological signal using a processor;applying a single time reference to the time series corresponding to the time the stimulus is supplied;obtaining further signal data from the time series of the physiological signal, the further signal data being obtained from the subject subsequent to the time reference;testing whether the signal model is a valid signal model for the physiological signal, wherein the testing includes employing the signal model constructed from the physiological signal obtained at least partially prior to the single time reference and the further signal data obtained subsequent to the single time reference; andindicating, based on the testing, whether reactivity is present in the physiological signal. 2. A method according to claim 1, wherein the constructing includes constructing the signal model based on the time series of the physiological signal, in which the physiological signal is an EEG signal. 3. A method according to claim 1, wherein the constructing includes constructing the signal model based on the time series of the physiological signal, in which the signal model is an autoregressive model. 4. A method according to claim 1, wherein the testing includes calculating a prediction value based on the signal model; anddefining a prediction error based on the further signal data and the prediction value. 5. A method according to claim 4, wherein the indicating includes comparing the prediction error with a predetermined threshold value and deciding, based on the comparing, whether reactivity is present in the physiological signal. 6. A method according to claim 4, wherein the defining includes defining the prediction error, in which the prediction error is dependent on the difference between the prediction value and a value obtained in the further signal data. 7. A method according to claim 1, wherein the constructing includes deriving the signal model based on a first time window of the time series and first data based on the signal model; andthe testing includes deriving (i) a further signal model based on a second time window of the time series and second data based on the further signal model, wherein the second time window includes the further signal data, and (ii) an error value dependent on the difference of the first and second data. 8. A method according to claim 1, wherein the constructing includes deriving first data based on the signal model and the testing includes deriving second data based on the further signal model, in which the signal model and the further signal model are dependent on an autocorrelation fu... | The General Electric Company | A61B005/04
US07831293 | B2 | 2010-11-09 | Method of defining a biological target for treatment | 2006-05-03 | A method of defining a biological target for treatment in which image detectable markers are placed at and correlated to the location or locations of tissue biopsy. By obtaining pathological analysis of the biopsy tissue specimens and correlating the analysis to the corresponding marker locations, the definitive pathological analysis of the target tissue at the marker locations can be correlated to corresponding locations on a functional image of the target tissue. The correlation of the marker locations and pathology to the corresponding locations on a functional image can then be used to prescribe and apply modulated therapy to the target tissue. | 1. A method of correlating a functional image of a target tissue to a pathology of said target tissue comprising: a) obtaining at least one biopsy specimen from said target tissue at a corresponding biopsy location;b) placing a detectable marker in contact with said target tissue substantially at said biopsy location, and correlating said marker to said biopsy location;c) obtaining a pathological analysis of said biopsy specimen and correlating said pathological analysis of said specimen to said biopsy location;d) obtaining at least one anatomic image of said target tissue which detects a location of at least one said detectable marker;e) obtaining at least one functional image of said target tissue; and,f) correlating the location of said at least one detectable marker with a corresponding location on said functional image,whereby said pathological analysis of said target tissue corresponding to said location of said at least one detectable marker can be correlated to a diagnosis of said target tissue produced by said functional image at said location of said at least one detectable marker. 2. The method of claim 1 wherein said anatomic image is selected from CT, MR, Xray, US, Fluoroscopy or a combination or hybrid thereof. 3. The method of claim 1 wherein said functional image is selected from SPECT, PET, MRSI, Optical Biopsy or a combination or hybrid thereof. 4. The method of claim 1 wherein said anatomic and functional images are acquired substantially simultaneously. 5. The method of claim 4 wherein said images are acquired as a hybrid or combination image. 6. The method of claim 5 wherein said hybrid or combination image is selected from SPECT/CT, PET/CT, SPECT/MR, PET/MR, SPECT/MR/CT, SPECT/MR/US, SPECT/CT/US, SPECT/CT/MRSI/US, SPECT/X-ray, and PET/X-ray. 7. The method of claim 1 wherein said anatomic and functional images are acquired separately. 8. The method of claim 7 wherein said separately acquired images are co-registered or fused using image fusion software. 9. The method of claim 1 wherein said target tissue is a prostate. 10. The method of claim 1 comprising obtaining biopsy specimens and placing said detectable markers substantially at multiple said biopsy locations, and wherein a plurality of said markers are visually distinguishable from each other by at least one said anatomic image, functional image or a combination or hybrid thereof. 11. The method of claim 10 wherein said markers are visually distinguishable based on their size shape, image intensity acoustical impedance, digital marker transmission or wavelength. 12. The method of claim 1 comprising correlating said pathological analysis of said target tissue corresponding to said location of said at least one detectable marker with a diagnosis of said target tissue produced by said functional image at said location of said at least one detectable marker. 13. The method of claim 12, further comprising prescribing therapy to said target tissue based on said co... | Advanced Clinical Solutions, Inc. | A61B005/05
US07831292 | B2 | 2010-11-09 | Guidance system and method for surgical procedures with improved feedback | 2003-07-16 | It is often desirable to define objects with respect to images of an anatomy displayed using an image guided surgery system. For non-trivial objects, or those with complicated two or three dimensional forms, it may be difficult to present information in a manner that is simple for a user to understand. The local distance to a surface of interest, such as the surface of the defined object, or to a desired position, the local penetration distance of the surface of interest, or haptic repulsion force, often provides the most useful information for augmenting the interaction of the user with the image guided surgery system. The scalar value of the local distance may be conveyed to the user by visual, audio, tactile, haptic, or other means. | 1. A method for use of a computer-assisted surgery system during a medical procedure, comprising: receiving information on an object of interest;tracking a position of a manually manipulated tool mechanically coupled to a haptic device as a surgeon manually holds and manipulates the tool;determining a scalar distance between a current position of said tool and said object of interest; andproviding an indication of said scalar distance to a user of said tool,wherein the object of interest includes at least one haptic virtual object that represents a virtual cutting boundary for the tool and that is defined at least in part by: a mapping between a pose of the tool and an output wrench of the haptic device, and,wherein providing the indication of the scalar distance includes activating at least one actuator of the haptic device to generate the output wrench when the tool intrudes on the virtual cutting boundary. 2. The method of claim 1, wherein the haptic virtual object includes a definition of a target shape for a surface of an anatomical structure of a patient and the scalar distance represents the current distance between the tool and the target shape. 3. The method of claim 1, wherein said haptic virtual object includes a definition of a curve, a point, a surface, a volume, and a set of desired positions. 4. The method of claim 1, wherein said providing step further comprises causing vibration of a device that is in contact with said user. 5. The method of claim 1, further comprising selecting a type of indication based at least in part on said scalar distance. 6. The method of claim 1, wherein said cutting tool is a bone removal tool, the virtual cutting boundary defines a curved bone surface that is to receive a surgical implant, and the providing step comprises indicating said scalar distance during the bone removal such that during the bone removal, the scalar distance indicates a thickness of bone to be removed. 7. A method for use of a computer-assisted surgery system during a surgical procedure, the method comprising: displaying information about an anatomical target region of a patient including an anatomical object on which the surgical procedure is to he performed to remove tissue by cutting with a surgical cutting tool;defining a virtual cutting boundary with sharp edges for the surgical cutting tool relative to the anatomical object, the anatomical object having a sufficiently high curvature that the surgical cutting tool tends to slip off the sharp edges of the virtual cutting boundary during cutting, the virtual cutting boundary being defined at least in part by a mapping between a pose of the surgical cutting tool and an output wrench of a haptic device to which the surgical cutting tool is attached;tracking a position of the surgical cutting tool as the surgical cutting tool is manually held and moved by a user in performing the surgical procedure;displaying an anatomical image of anatomy of the patient including a represe... | Mako Surgical Corp. | A61B005/05
US07831291 | B2 | 2010-11-09 | Subject moving apparatus and imaging apparatus | 2005-08-10 | A cradle position sensing mechanism that includes a scale unit having a scale, which is used to detect the position of the cradle, formed thereon and being extended to cover a sliding range within which the cradle can be slid; and a sensor unit that moves in the directions, in which the scale unit is extended, along with the slide of the cradle so as to detect the scale of the scale unit. The cradle position sensing mechanism serves as a digital linear encoder. Based on the reading on the scale detected by the sensor, the cradle position sensing mechanism calculates the position into which the cradle is slid. | 1. Subject moving apparatus included in an imaging apparatus, which images a subject within an imaging space, for moving the subject to the imaging space, comprising: a cradle on which the subject lies down;a cradle support that supports the cradle so that the cradle will be slid into the imaging space; anda cradle position sensing mechanism that senses a position into which the cradle is slid, wherein:the cradle position sensing mechanism comprises: a scale unit that has a scale, which is used to detect the position of the cradle, formed thereon and that is extended to cover a sliding range within which the cradle can be slid; and a sensor unit that is moved relatively to the scale unit in directions, in which the scale unit is extended, along with the slide of the cradle, and that detects the scale of the scale unit; andbased on the reading on the scale detected by the sensor unit, the cradle position sensing mechanism calculates a position into which the cradle is slid. 2. The subject moving apparatus according to claim 1, wherein: the cradle support comprises a carriage that supports the cradle, and guide rails that extend in the directions, in which the cradle is slid, and support the carriage; and the sensor unit is mounted on the carriage, and the scale unit is mounted on the cradle support and extended in the directions in which the guide rails are extended. 3. The subject moving apparatus according to claim 2, wherein the cradle position sensing mechanism comprises: a sensor unit control unit that controls the sensor unit; a cable linking the sensor unit and the sensor unit control unit; and a cable guide that sheathes the cable and restricts the movable range of the cable. 4. The subject moving apparatus according to claim 1, wherein the scale unit has the scale thereof realized with a magnetic scale, and the sensor unit includes a magnetic sensor that detects the magnetic scale. 5. The subject moving apparatus according to claim 1, wherein the scale unit has the scale thereof realized with an optical scale, and the sensor unit includes an optical sensor that detects the optical scale. 6. The subject moving apparatus according to claim 1, wherein the sensor unit is separated from the scale unit with a certain space between them so that the sensor unit will be in non-contact with the scale unit. 7. The subject moving apparatus according to claim 1, wherein the sensor unit is disposed in contact with the scale unit. 8. An imaging apparatus for imaging a subject within an imaging space, comprising: a cradle on which the subject lies down;a cradle support that supports the cradle so that the cradle will be slid into the imaging space;a cradle position sensing mechanism that senses a position into which the cradle is slid;a scanner that scans the subject, who is moved into the imaging space by sliding the cradle, so as to acquire raw data; andan image production unit that produces an image of the subject according to the raw data acqu... | GE Medical Systems Global Technology Company, LLC | A47B071/00 A61B005/05
US07826382 | B2 | 2010-11-02 | Close proximity communication device and methods | 2008-05-30 | Disclosed herein are methods and systems for receiving an encoded data packet including one or more error detection bits, one or more close proximity commands, and a communication identifier, decoding the received data packet, performing error detection based on the one or more error detection bits, validating the decoded received data packet, and executing one or more routines associated with the respective one or more close proximity commands when the decoded received data packet is validated. | 1. A method, comprising: establishing a communication range between a data processing unit and a control unit by placing the control unit within a predetermined distance from the data processing unit;receiving an encoded data packet at the data processing unit when the control unit is within the predetermined distance, wherein the received encoded data packet includes one or more error detection bits, one or more close proximity commands, and a communication identifier;decoding the received data packet;performing error detection based on the one or more error detection bits;validating the decoded received data packet; andexecuting one or more routines associated with the respective one or more close proximity commands when the decoded received data packet is validated; wherein the executed one or more routines includes transmitting analyte related data. 2. The method of claim 1 wherein the received data packet is Manchester encoded. 3. The method of claim 1 wherein the one or more error detection bits include a predetermined bit pattern. 4. The method of claim 3 wherein the predetermined bit pattern includes a dotting pattern. 5. The method of claim 1 wherein decoding the received data packet includes performing Manchester decoding. 6. The method of claim 1 wherein validating the decoded received data packet includes comparing the received communication identifier in the data packet with a stored value. 7. The method of claim 1 wherein the communication identifier includes a device identification information. 8. The method of claim 1 wherein the one or more routines are associated with the operation of an analyte monitoring device. 9. The method of claim 1 wherein the executed one or more routines includes a power on routine, a power off routine, a data transfer initiation routine, or a data transfer disable routine. 10. The method of claim 1 wherein the analyte related data includes a monitored analyte level. 11. The method of claim 10 wherein the analyte is glucose. 12. The method of claim 1 including storing the received data packet. 13. A method, comprising: establishing a communication range between a data processing unit and a control unit by placing the control unit within a predetermined distance from the data processing unit;receiving an encoded data packet at the data processing unit when the control unit is within the predetermined distance, wherein the received encoded data packet includes a close proximity command and a communication identifier;decoding the received data packet;validating the decoded received data packet; andexecuting one or more routines associated with the respective one or more close proximity commands when the decoded received data packet is validated, wherein the one or more close proximity commands are configured to control the transmission of analyte related data. 14. The method of claim 13 wherein the encoded data packet is Manchester encoded. 15. The method of claim 13 wherein decoding the rece... | Abbott Diabetes Care Inc. | G08B001/08 A61B005/145
US07840258 | B2 | 2010-11-23 | Imaging device, method thereof, and program | 2004-12-22 | To realize an imaging device, imaging method and program capable of improving image quality. This invention controls light sources so as to irradiate a body with irradiation light of a luminance level higher than that in the air arriving at the body, adjusts the imaging sensitivity of a solid imaging element for performing photoelectric conversion on pattern scattered light obtained through the inside of the body, and masks the luminance level deterioration region of a pattern image obtained from the solid imaging element to treat the masked image as an image to be composed with a reference image. | 1. An imaging device comprising: ambient light measuring means for determining a luminance level of air arriving at a body;light source control means for controlling light sources so as to irradiate the body with irradiation light of a luminance level higher than the luminance level of the air arriving at the body based on the luminance level determined by the ambient light measuring means;sensitivity adjustment means for adjusting imaging sensitivity of a solid imaging element for performing photoelectric conversion on pattern scattered light obtained through an inside of the body; andmasking means for masking a luminance level deterioration region of a pattern image obtained from the solid imaging element to treat the pattern image as a masked image to be composed with a reference image. 2. The imaging device according to claim 1, wherein: the light source control means controls the light sources so as to sequentially irradiate the body with irradiation light every unit time; andthe masking means masks the luminance level deterioration region of the pattern image obtained from the solid imaging element every unit time and composes the pattern image with the reference image. 3. The imaging device according to claim 1, wherein: the light source control means controls the light sources so as to sequentially irradiate the body every unit time at luminance levels different from each other; andthe masking means masks the luminance level deterioration region of the pattern image obtained from the solid imaging element every unit time and composes the pattern image with the reference image. 4. The imaging device according to claim 1, wherein: the sensitivity adjustment means adjusts the imaging sensitivity by sequentially changing an amount being accumulated in the solid imaging element as a result of the photoelectric conversion per unit time; andthe masking means masks the luminance level deterioration region of the pattern image based on the amount obtained from the solid imaging element and composes the pattern image with the reference image. 5. An imaging method comprising: determining a luminance level of air arriving at a body;controlling, with a processor, light sources so as to irradiate the body with irradiation light of a luminance level higher than the luminance level of the air arriving at the body;adjusting imaging sensitivity of a solid imaging element for performing photoelectric conversion on pattern scattered light obtained through an inside of the body; andmasking a luminance level deterioration region of a pattern image obtained from the solid imaging element to treat an image masked as an image to be composed with a reference image. 6. The imaging method according to claim 5, wherein, the controlling includes controlling the light sources so as to sequentially irradiate the body with the irradiation light every unit time; andthe masking includes masking the luminance level deterioration region of the pattern image obtained fro... | Sony Corporation | A61B006/00 A61B005/00
US07835787 | B2 | 2010-11-16 | Single trial detection in encephalography | 2004-10-15 | An EEG cap (8) having 64 or 128 electrodes (10) is placed on the head of the subject (11) who is viewing CRT monitor (14). The signals on each channel are amplified by amplifier (17) and sent to an analog-to-digital converter (20). PC (23) captures and records the amplified signals and the signals are processed by signal processing PC (26) performing linear signal processing. The resulting signal is sent back to a feedback/display PC (29) having monitor (14). | 1. A method for interpreting brain activity comprising the steps of: attaching a high-density array of brain activity sensors to a subject;prompting said subject with an external stimulus;generating a response signal from each of said sensors at a predetermined point in time with reference to the occurrence of said prompting;computing a respective spatial weighting coefficient corresponding to each of said sensors by applying a linear classifier to said predetermined number of response signals;weighting each response signal in accordance with its corresponding weighting coefficient; andlinearly integrating the weighted response signals whereby a single output signal is generated, said output signal having a relatively high correlation to the brain activity of said subject in response to said external stimulus. 2. The method of claim 1 wherein said weighting coefficients are computed such that the weighted responses are maximally discriminating between first and second predetermined times corresponding to two stimuli where a maximum difference in brain activity is expected. 3. The method of claim 2 wherein the linear integration is a summation. 4. The method of claim 2 wherein said brain activity sensors are EEG sensors. 5. The method of claim 2 wherein said brain activity sensors are MEG sensors. 6. The method of claim 2 wherein said stimulus is visual. 7. The method of claim 1 further comprising the steps of: identifying an expected scalp locality for discriminating components of the output signal relative to the external stimulus;computing a sensor projection for each of said sensors;using said sensor projections to identify the scalp locality for the actual discriminating components for said external stimulus; andevaluating the validity of said output signal by determining whether the identified locality for the actual discriminating components is substantially the same as the expected locality. 8. A system for interpreting brain activity comprising: a high-density array of brain activity sensors for application to the head of a subject; anda computer for receiving signals from said sensors, said computer programmed to: compute a respective spatial weighting coefficient corresponding to each of said sensors by applying a linear classifier to signals generated by said sensors in response to an external stimulus,weight each of said signals in accordance with its corresponding weighting coefficient, andlinearly integrate the weighted response signals,whereby a single output signal is generated, said output signal having a relatively high correlation to the brain activity of said subject in response to said external stimulus. 9. The system of claim 8 further comprising a feedback display. 10. The system of claim 9 wherein said feedback display provides said external stimulus. 11. The system of claim 8 wherein said weighting coefficients are computed such that the weighted responses are maximally discriminating between first and second pr... | The Trustees of Columbia University in the City of New York | A61B005/04
US08855770 | B2 | 2014-10-07 | Duodenal eating sensor | 2008-07-24 | Apparatus is described including a sensing electrode that detects electrical activity of a duodenum of a subject. The apparatus includes a control unit configured to facilitate a treatment of the subject responsively to the detected electrical activity. Other embodiments are also described. | 1. Apparatus, comprising: a sensing electrode configured to detect electrical activity of a duodenum of a subject; anda control unit configured, responsively to the detected electrical activity, to perform an action selected from the group consisting of: facilitating a treatment of the subject, and stopping a treatment of the subject,wherein the control unit is configured to calculate a pattern based on changes in the detected electrical activity over time, wherein the control unit comprises a pattern recognition unit configured to compare the calculated pattern to a predetermined characteristic pattern, and wherein the control unit is configured to perform the selected action responsively to the comparing. 2. The apparatus according to claim 1, wherein the pattern recognition unit is configured to compare the calculated pattern to the characteristic pattern irrespective of relative magnitudes of the calculated pattern and the characteristic pattern. 3. The apparatus according to claim 1, wherein the characteristic pattern includes a v-shaped pattern, and wherein the pattern recognition unit is configured to compare the calculated pattern to the v-shaped pattern. 4. The apparatus according to claim 1, wherein the control unit is configured to calculate the pattern based on changes in activation energy of the detected electrical activity over time. 5. The apparatus according to claim 1, wherein the control unit is configured to calculate the pattern based on changes in number of activations of the detected electrical activity per unit time. 6. The apparatus according to claim 1, wherein the control unit is configured to calculate the pattern based on changes over time of a parameter of the detected electrical activity selected from the group consisting of: slow-wave energy, time frequency distribution, and number of activations of slow-waves per unit time. 7. The apparatus according to claim 1, wherein the characteristic pattern includes an eating pattern that is characteristic of a person who is eating, and wherein the pattern recognition unit is configured to compare the calculated pattern to the eating pattern. 8. The apparatus according to claim 7, wherein the control unit comprises a template assessment unit configured to determine a subject-eating pattern that is characteristic of the subject when the subject is eating, and wherein the pattern recognition unit is configured to compare the calculated pattern to the subject-eating pattern. 9. The apparatus according to claim 1, wherein the pattern recognition unit is configured to compare the calculated pattern to the characteristic pattern by comparing a calculated pattern based on at least 2 minutes of the detected electrical activity to the characteristic pattern. 10. The apparatus according to claim 9, wherein the pattern recognition unit is configured to compare the calculated pattern to the characteristic pattern by comparing a calculated pattern based on up to 30 minutes of the... | Betastim, Ltd. | A61N001/00 A61N001/36
US08615287 | B2 | 2013-12-24 | Catheter tracking and endocardium representation generation | 2010-08-30 | Methods and systems are disclosed for determining information about a position of an object within a distribution of materials having different complex conductivities. The method includes: (i) causing current to flow in the distribution; (ii) measuring an electrical signal at each of multiple locations in the distribution of materials in response to the current flow; (iii) providing spatial information about the distribution of materials with respect to a first reference frame, the spatial information indicative of regions of different complex conductivity in the distribution of materials; and (iv) determining the position of the object with respect to the spatial information about the distribution of materials based on measured electrical signals and the spatial information. In certain embodiments, the object is a catheter inserted into a patients heart cavity for cardiac mapping. | 1. A method for determining information about a location and orientation of a catheter in a patient's heart, the method comprising: determining electrical signals as a function of relative position based on spatial information indicative of regions of different complex conductivity in the patient's heart cavity and conductivity values;causing current to flow between each of three or more sets of current-injecting electrode source/sink pairs on a catheter inserted into the patient's heart;measuring by multiple measuring electrodes located on the catheter, electrical signals in response to the current flow caused by each set of current injecting electrode source/sink pairs on the catheter;determining the location and orientation of the catheter in the patient's heart based on the electrical signals measured by the multiple measuring electrodes located on the catheter in response to the current flow caused by the current injecting electrode source/sink pairs on the catheter and thedetermined electrical signals based on the spatial information;measuring, by at least some of the multiple measuring electrodes on the catheter, electrical signals indicative of cardiac electrical activity; anddetermining information about electrical activity of the heart based on the measured electrical signals indicative of cardiac electrical activity and the determined location and orientation of the catheter in the patient's heart;wherein causing the current to flow comprises injecting the current at frequencies spaced from those corresponding to the cardiac electrical activity and the determining the location and orientation of the catheter comprises frequency processing the measured electrical signal responsive to the injected current to distinguish electrical signals responsive to the injected current from those corresponding to the cardiac electrical activity. 2. The method of claim 1, further comprising: displaying the information about the electrical activity of the heart on a representation of the patient's heart. 3. The method of claim 1, wherein determining the location and orientation of the catheter further comprises using an optimization algorithm that determines a conductivity value for each of one or more of the materials in the distribution of materials; determines electrical signals based on the determined conductivity values and the spatial information about the distribution of materials as a function of the relative position; and minimizes differences between the measured electrical signals responsive to the injected current and the determined electrical signals. 4. The method of claim 1, wherein the spatial information about the distribution of materials is determined based on one or more of: a computed tomography (CT) image; a magnetic resonance imaging (MRI) image; a fluoroscopic rotational angiography image; and an ultrasound image. 5. The method of claim 1, wherein determining the location and orientation of the catheter comprises frequency p... | Rhythmia Medical, Inc. | A61B005/05
US08536469 | B2 | 2013-09-17 | Body composition monitor determining a degree of change relative to a past body composition value wherein the favorable stage on the display is obtained even with a slight change at the start of dieting, so that a motivation level of the user is increased | 2009-03-16 | A body composition monitor of the present invention includes means for inputting a reference value of a body composition component, means for inputting a measurement value of the body composition component, body composition comparison means for comparing the reference value and the measurement value so as to determine a body composition component change amount of the measurement value relative to the reference value, display means provided with a plurality of stages for displaying one of the stages corresponding to the body composition component change amount, and stage determination means provided with a plurality of determination widths respectively corresponding to the plurality of stages for determining one of the stages corresponding to the body composition component change amount with using the plurality of determination widths, wherein the plurality of determination widths are not identical to each other. | 1. A body composition monitor, comprising: means for inputting a reference value of a body composition component;means for inputting a measurement value of the body composition component;body composition comparison means for comparing the reference value and the measurement value so as to determine a body composition component change amount of the measurement value relative to the reference value;display means provided with a plurality of stages for displaying one of the stages corresponding to the body composition component change amount; andstage determination means provided with a plurality of determination widths respectively corresponding to the plurality of stages for determining one of the stages corresponding to the body composition component change amount with using the plurality of determination widths, wherein each of the plurality of determination widths is a range defined by the body composition component change amount divided by one of a plurality of threshold values based on the plurality of stages, andthe plurality of determination widths are not identical to each other. 2. The body composition monitor according to claim 1, wherein the determination width of the stage corresponding to the body composition component change amount smaller than a predetermined amount is narrower than the determination widths of the other stages. 3. The body composition monitor according to claim 1, wherein the determination widths are differentiated in accordance with magnitude of the body composition component change amount. 4. The body composition monitor according to claim 1, wherein the determination widths are differentiated in accordance with a positive and a negative change in the body composition component change amount. 5. The body composition monitor according to claim 1, further comprising: means for storing the measurement value, whereina plurality of past measurement values each serve as the reference value, and the stages are displayed in a graph in chronological order. 6. The body composition monitor according to claim 1, comprising: means for inputting a reference value of a body weight;means for inputting a measurement value of the body weight; andbody weight comparison means for comparing the reference value and the measurement value so as to determine a body weight change amount of the measurement value relative to the reference value, whereinthe stage determination means determines the stage based on the body composition component change amount and the body weight change amount. 7. The body composition monitor according to claim 1, wherein the body composition component indicates a body fat amount, a body fat percentage, a subcutaneous fat amount, a subcutaneous fat percentage, a visceral fat amount, a visceral fat area, a visceral fat level, a skeletal muscle amount, or a skeletal muscle percentage. 8. The body composition monitor according to claim 1, wherein the body composition component is calculated from a plurality o... | Omron Healthcare Co., Ltd | A61B005/053 A61B005/05 G01G019/50
US09023292 | B2 | 2015-05-05 | Blood sampling device comprising at least one filter | 2008-09-26 | Disclosed herein is a device for sampling blood by capillarity for separating a plasma phase of the blood from the cell phase. The device is a strip and includes a filter having a membrane with no clumping agents, a prefilter for spreading blood over the surface of the membrane; and an absorbent paper downstream from the filter. The membrane is dimensioned to cover only a part of the surface of the prefilter and the migration of blood by capillarity within the prefilter takes place essentially lengthwise relative to the strip before blood comes into contact with the membrane. | 1. A device for sampling blood by capillarity for separating a plasma phase of the blood from a cell phase, wherein said device is in the form of a strip that comprises: a filter comprising:a membrane having pores of dimension lying in the range 0.1 μm to 5 μm, said membrane having no clumping agent; anda prefilter, on which said membrane is mounted, for spreading the blood over a surface of said membrane;said membrane being further dimensioned so as to have a contact surface with the prefilter that covers only part of a surface of said prefilter, so as to ensure that the migration of blood by capillarity within said prefilter, between an edge face of said prefilter and the contact surface of the membrane with the prefilter, takes place essentially lengthwise relative to the prefilter before blood comes into contact with said membrane, wherein the membrane does not cover the edge face; andan absorbent paper mounted on the membrane. 2. A device according to claim 1, wherein the dimension of the pores lies in the range 0.5 μm to 5 μm. 3. A device according to claim 1, wherein the density of the pores is such that said pores occupy 10% to 90% of the surface area of the membrane of the filter. 4. A device according to claim 3, wherein the density of the pores is such that said pores occupy 40% to 80% of the surface area of the membrane. 5. A device according to claim 1, wherein said prefilter is an absorbent or filtering paper that includes fiber. 6. A device according to claim 1, wherein said device includes a stiff support on which the prefilter is mounted. 7. A device according to claim 1, wherein the device includes a stiff support on which the prefilter is mounted, said support being hydrophobic and non-absorbent. 8. A device according to claim 1, wherein said prefilter is an absorbent or filtering paper that includes glass fibers. 9. A device according to claim 3, wherein the density of the pores is such that said pores occupy 50% to 80% of the surface area of the membrane. 10. A device according to claim 1, wherein the device comprises a stiff support on which the prefilter is mounted, a peel-off adhesive which is peelable from the support, the prefilter, the membrane and the absorbent paper without tearing them away. | Commissariat a l'Energie Atomique et aux Energies Alternatives | A61B005/151 G01N033/49 A61B005/15
US08708905 | B2 | 2014-04-29 | Method, device and computer product for EEG monitoring, analysis and display | 2009-06-12 | A method, device, and computer program product for monitoring the physiological state of a subject is disclosed. A first plurality of physiological signals are acquired from the subject and a second plurality of signal parameters are determined based on the first plurality of physiological signals. Relevant signal information is selected from among the signal information produced and at least one screen page is displayed at a time, thereby to make at least part of the signal information visible to a user. The signal information visible to the user on the at least one screen page may then be revised based on the selection and without user interaction, thereby to accelerate clinical decision-making and to diminish the possibility of undetected changes in the state of the subject. | 1. A method for monitoring the physiological state of a subject, the method comprising: acquiring EEG signals from a subject, wherein a first plurality of EEG signals are acquired from a left hemisphere and a second plurality of EEG signals are acquired from a right hemisphere;determining a third plurality of EEG signal parameters and a fourth plurality of EEG signal parameters, wherein the third plurality of EEG signal parameters are based on the first plurality of EEG signals and the fourth plurality of EEG signal parameters are based on the second plurality of EEG signals;producing signal information comprising at least one type of signal information selected from a group including (1) time series representatives for each of the third plurality of EEG signal parameters and for each of the fourth plurality of EEG signal parameters; and (2) signal waveforms of the first plurality of EEG signals and of the second plurality of EEG signals;selecting relevant signal information from the signal information by comparing the third plurality of EEG signal parameters with each other and by comparing the fourth plurality of EEG signal parameters with each other, wherein at least one EEG signal parameter is selected as the relevant signal information, wherein the selected at least one EEG signal parameter belongs to the third plurality of EEG signal parameters or to the fourth plurality of EEG signal parameters;defining a set of optimal parameters from the third plurality of EEG signal parameters and from the fourth plurality of EEG signal parameters based on subject-specific patient data;examining whether the set of optimal parameters is consistent with the at least one EEG signal parameter selected as the relevant signal information;displaying a primary screen page, thereby to make at least part of the signal information visible to a user; andrevising the signal information visible to the user on the primary screen page to display the selected relevant signal information, wherein the revising is performed in response to the examining and without user interaction. 2. The method according to claim 1, wherein the revising includes changing at least one time series currently visible to the user. 3. The method according to claim 1, wherein the revising includes making at least one signal waveform visible to the user. 4. The method according to claim 1, wherein the revising includes changing the primary screen page currently displayed. 5. The method according to claim 1, wherein the selecting includes comparing at least some of the third plurality of EEG signal parameters and at least some of the fourth plurality of EEG signal parameters with predetermined threshold values. 6. The method according to claim 1, wherein the revising includes changing appearance of the signal information currently visible to the user on the primary screen page. 7. The method according to claim 1, wherein the selecting further includes selecting relevant signal information by... | General Electric Company | A61B005/00 A61B005/0476 G06F019/00
US10555676 | B2 | 2020-02-11 | Method for generating alarms/alerts based on a patient's posture and vital signs | 2009-05-20 | The invention provides a body-worn monitor that measures a patient's vital signs (e.g. blood pressure, SpO2, heart rate, respiratory rate, and temperature) while simultaneously characterizing their activity state (e.g. resting, walking, convulsing, falling). The body-worn monitor processes this information to minimize corruption of the vital signs by motion-related artifacts. A software framework generates alarms/alerts based on threshold values that are either preset or determined in real time. The framework additionally includes a series of ‘heuristic’ rules that take the patient's activity state and motion into account, and process the vital signs accordingly. These rules, for example, indicate that a walking patient is likely breathing and has a regular heart rate, even if their motion-corrupted vital signs suggest otherwise. | 1. A method for generating an alarm while monitoring vital signs and posture of a patient, the method comprising the following steps: (a) detecting a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a first sensor comprising a first detector configured to be worn on the patient's body;(b) detecting a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a second sensor comprising a second detector configured to be worn on the patient's body;(c) detecting a set of time-dependent motion waveforms with at least one motion-detecting sensor positioned on the patient's torso, wherein the set of time-dependent motion waveforms are indicative of motion of the patient's torso;(d) receiving and processing the first and second time-dependent physiological waveforms and the set of time-dependent motion waveforms, and generating an alarm therefrom indicative of a need for medical attention, using one or more processing components comprising a microprocessor, the one or more processing components configured to (i) process the first and second time-dependent physiological waveforms to determine at least one vital sign from the patient;(ii) analyze at least a portion of the set of time-dependent motion waveforms, or a mathematical derivative thereof, to determine a vector corresponding to motion of the patient's torso;(iii) compare the vector to a coordinate space representative of how the motion-detecting sensor is oriented on the patient to determine a posture parameter;(iv) compare the posture parameter to a threshold value to estimate the patient's posture;(v) compare the at least one vital sign from the patient to a predetermined alarm criteria, wherein an alarm is indicated by a variance of the vital sign relative to the predetermined alarm criteria to determine an alarm parameter; and(vi) generate the alarm by collectively processing the patient's posture and the alarm parameter, wherein the alarm indicated by a variance of the vital sign relative to the predetermined alarm criteria is regulated according to the patient's posture. 2. The method of claim 1, wherein the set of time-dependent motion waveforms comprises three time-dependent motion waveforms, each corresponding to a unique axis, and wherein the vector corresponding to motion of the patient's torso is determined from the three time-dependent motion waveforms. 3. The method of claim 2, wherein the coordinate space representative of how the motion-detecting sensor is oriented on the patient is defined by three positional vectors, each indicating an orientation of the motion-detecting sensor on the patient. 4. The method of claim 3, wherein a first positional vector of the three positional vectors corresponds to a vertical axis in the coordinate space, a second positional vector of the three positional vectors corresponds to a horizontal axis in the coordin... | SOTERA WIRELESS, INC. | A61B005/0205 G16H050/50 G16H015/00 A61B005/021 A61B005/11 A61B005/00 G08B021/04 A61B005/02 A61B005/022 A61B005/0402 G16H050/20 A61B005/01 A61B005/024 A61B005/0245 A61B005/04 A61B005/0404 A61B005/0452 A61B005/0456 A61B005/08 A61B005/145
US09386929 | B2 | 2016-07-12 | Aural heart monitoring apparatus and method | 2008-11-18 | A heart monitor is disclosed including an electroacoustic transducer such as an earphone coupled to a controller. The transducer is positioned in a person's ear in acoustic communication with the tympanum. Signals from the transducer are processed to determine the presence of pulsatile blood flow. The heart monitor may be incorporated into a defibrillator to sense the presence of blood flow for use in a shock delivery decision. | 1. An automatic external defibrillator (AED) comprising: a pair of electrodes adapted to contact a subject;an ECG processor, coupled to the electrodes, which acts to analyze an ECG signal;an energy source;a high voltage circuit coupled to the energy source;a shock delivery circuit coupled to the ECG processor, the high voltage circuit, and the electrodes;an earphone adapted to detect acoustic signals of the ear of the subject; andan earphone signal processor, coupled to the earphone, which operates to sense the presence of pulsatile blood flow from signals received by the earphone,wherein the signal processor is coupled to the shock delivery circuit,wherein the shock delivery circuit makes a decision to shock based upon an absence of a pulsatile ECG signal from the ECG processor and an absence of pulsatile blood flow sensed by the earphone signal processor. 2. The automatic external defibrillator of claim 1, wherein the earphone signal processor further comprises a signal conditioning circuit responsive to signals produced by the earphone. 3. The automatic external defibrillator of claim 1, wherein the shock delivery circuit contraindicates a decision to shock in the presence of pulsatile blood flow sensed by the earphone signal processor. 4. The automatic external defibrillator of claim 1, wherein a condition of pulseless electrical activity is indicated by the detection of a pulsatile ECG signal by the ECG processor and the absence of pulsatile blood flow as sensed by the earphone. | KONINKLIJKE PHILIPS N.V. | A61N001/39 A61B005/024 A61B005/00 A61B007/04
US20100305414 | A1 | 2010-12-02 | METHOD AND APPARATUS FOR TRANSMITTING BIOLOGICAL INFORMATION OF USER | 2010-03-15 | A method of transmitting biological information includes selecting a transmission mode from a plurality of transmission modes and transmitting biological information of a user to an external device using the selected transmission mode. Each of the transmission modes transmits different quantities of data to the external device based on a state of the biological information of the user. | 1. A method of transmitting biological information, the method comprising: selecting a transmission mode from a plurality of transmission modes, each of the transmission modes transmitting different quantities of data to an external device based on a state of biological information of a user; andtransmitting the biological information of the user to the external device using the selected transmission mode. 2. The method of claim 1, wherein the selecting the transmission mode comprises selecting a transmission mode, from the plurality of transmission modes, in which a first quantity of data is transmitted to the external device, when a normal state of the biological information is maintained for a predetermined amount of time. 3. The method of claim 2, wherein the transmitting the biological information comprises transmitting a portion of the biological information to the external device according to a priority, which is based on characteristics of all portions of the biological information, using the selected transmission mode. 4. The method of claim 3, wherein the characteristics comprise at least one selected from a group consisting of data size and an importance of each portion of the biological information. 5. The method of claim 2, wherein the transmitting the biological information comprises transmitting portions of the biological information to the external device, each portion including a transmission frequency which varies according to a priority, which is based on characteristics of all portions of the biological information, using the selected transmission mode. 6. The method of claim 5, wherein the characteristics comprise at least one selected from a group consisting of data size and an importance of each portion of the biological information. 7. The method of claim 2, wherein the selecting the transmission mode further comprises selecting a second transmission mode, from the plurality of transmission modes, in which a second quantity of data is transmitted to the external device, when an abnormal state of the biological information is detected, andthe first quantity of data is smaller than the second quantity of data. 8. The method of claim 7, wherein the transmitting the biologic information comprises transmitting all of the biological information to the external device using the selected second transmission mode. 9. The method of claim 7, wherein the transmitting the biologic information comprises transmitting the biological information to the external device using at least one of the first transmission mode and second transmission mode while equalizing a transmission frequency of each portion of the biological information. 10. The method of claim 1, wherein the transmitting the biological information comprises transmitting the biological information to the external device by controlling a transmission frequency of at least one portion of the biological information based on a state of the biological information. 11. The ... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/00 H04B001/04
US20100286533 | A1 | 2010-11-11 | Biosignal-detecting apparatus and method thereof | 2007-10-19 | A method of detecting a biosignal includes; measuring a biosignal from a user; band pass filtering the biosignal to each of a first frequency band and/or a second frequency band, and detecting a first frequency band signal and/or a second frequency band signal; calculating a PSD of the first frequency band signal and detecting a frequency band signal with a maximum PSD in the first frequency band as a heartbeat signal of the user; and calculating a PSD of the second frequency band signal and/or detecting a frequency band signal with a maximum PSD in the second frequency band as a respiration signal of the user. | 1. A method of detecting a biosignal, the method comprising: measuring a biosignal from a user;band pass filtering the biosignal to a frequency band; andcalculating a power spectrum density (PSD) of the band pass filtered signal and detecting a frequency band signal with a maximum PSD. 2. The method of claim 1, wherein the biosignal is detected from the user using either a Doppler radar sensor, or a piezoelectric sensor and a radio frequency (RF) sensor. 3. The method of claim 1, wherein the frequency band is within a range of about 1.0 Hz to about 1.5 Hz. 4. The method of claim 1, wherein the frequency band is set to a predetermined band based on a heart rate that is measured from the user for a predetermined period of time when the user is at rest. 5. The method of claim 1, further comprising: comparing a PSD value of the heartbeat signal with a predetermined threshold value; andproviding an alarm signal when the PSD value of the heartbeat signal is less than the predetermined threshold value as a result of the comparison. 6. The method of claim 5, wherein the predetermined threshold value is set to a value corresponding to about 80% of a previously measured PSD value of the user's heartbeat signal. 7. A method of detecting a biosignal, the method comprising: measuring a biosignal from a user;band pass filtering the biosignal to a frequency band; andcalculating a PSD of the band pass filtered signal and detecting a frequency band signal with a maximum PSD as a respiration signal of the user. 8. The method of claim 7, wherein the biosignal is detected from the user using either a Doppler radar sensor, or a piezoelectric sensor and an RF sensor. 9. The method of claim 7, wherein the frequency band is within the range of about 0.2 Hz to about 0.35 Hz. 10. The method of claim 7, wherein the frequency band is set to a predetermined band based on a respiration rate that is measured from the user for a predetermined period of time when the user is at rest. 11. The method of claim 7, further comprising: comparing a PSD value of the respiration signal with a predetermined threshold value; anddetermining the user is in a breathless state when the PSD value of the respiration signal is less than the predetermined threshold value as a result of the comparison. 12. The method of claim 11, wherein the predetermined threshold value is set to a value corresponding to about 80% of a previously measured PSD value of the user's respiration signal. 13. A method of detecting a biosignal, the method comprising: measuring a biosignal from a user;band pass filtering the biosignal to each of a first frequency band and a second frequency band, and detecting a first frequency band signal and a second frequency band signal;calculating a PSD of the first frequency band signal and detecting a frequency band signal with a maximum PSD in the first frequency band as a heartbeat signal of the user; andcalculating a PSD of the second frequency band signal and detecti... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/0205
US20100298718 | A1 | 2010-11-25 | Multiple Mode, Portable Patient Monitoring System | 2010-04-27 | The specification discloses a number of embodiments directed to patient monitoring systems. The system determines a physiological parameter of a patient and includes an input device for receiving a plurality of physical inputs from a user, such as finger taps, and communicating to a memory interval data indicative of the frequency at which the plurality of physical inputs are received from the user within a time period and a processor in data communication with said memory wherein said processor determines at least one physiological parameter of a patient based upon said interval data. | 1. A system for determining a physiological parameter of a patient, the system comprising: an input device for receiving a plurality of physical inputs from a user and communicating to a memory data indicative of the frequency at which said plurality of physical inputs are received from the user within a time period; anda processor in data communication with said memory wherein said processor determines the value of a physiological parameter based upon said data. 2. The system of claim 1, wherein said data collected by the device represents the time of each physical input by said user. 3. The system of claim 1 wherein said physiological parameter is respiratory rate. 4. The system of claim 1 wherein said physiological parameter is at least one of pulse rate or respiratory rate. 5. The system of claim 1 wherein the input device is a touch screen display. 6. The system of claim 5 wherein the plurality of physical inputs received from the user are finger-taps on said touch screen display. 7. The system of claim 1 wherein the input device is at least one of a touch screen display, mouse pad, or keyboard. 8. The system of claim 7 wherein the plurality of physical inputs received from the user are finger-taps on at least one of said touch screen display, mouse pad, or keyboard. 9. The system of claim 1 wherein the time period is fixed. 10. The system of claim 9 wherein the fixed time period is system-defined. 11. The system of claim 9 wherein the fixed time period is user-defined. 12. The system of claim 1 wherein the time period is variable and is dynamically determined by the system when the user starts and stops the physical input. 13. The system of claim 1 wherein the processor determines the value of the physiological parameter by averaging all of interval data generated by said physical inputs. 14. The system of claim 13 wherein the average interval data is derived by calculating a straight arithmetic average of the measured intervals. 15. The system of claim 13 wherein the average interval data is derived by rejecting the longest and shortest interval data generated by the physical inputs and then averaging the remaining data. 16. The system of claim 13 wherein the average interval data is derived by calculating a median value of the physical inputs. 17. The system of claim 13 wherein the average interval data is derived by analyzing clusters of intervals and using a clustering algorithm to reject long or short intervals that should not be included. 18. The system of claim 13 wherein the average interval data is displayed in real time and continuously updated with each physical input. 19. A system for determining a physiological parameter of a patient, the system comprising: an input device for receiving a plurality of physical inputs from a user, wherein said physical inputs comprise finger taps, and communicating to a memory data indicative of the frequency at which said plurality of physical inputs are received from the user... | - | A61B005/0205
US20100278752 | A1 | 2010-11-04 | Methods for Detection and Prevention of Tick Infestation and Pathogen Transmission | 2008-07-25 | The invention generally features methods for the prevention and detection of a tick infestation. The present invention also features methods for decreasing the ability of a tick to feed on a subject. | 1-71. (canceled) 72. A method for the detecting a tick infestation in a subject comprising: administering to the subject an immunogenic composition comprising one or more cystatin antigens, wherein the cystatin corresponds to a polypeptide comprising an amino acid sequence which is at least 80% homologous to the amino acid sequence of SEQ. ID. NO. 2 and the subject was previously exposed to the immunogenic composition; and wherein the immunogenic composition is administered in an amount that stimulates a desired immune response to the immunogenic composition in the subject; thereby detecting tick infestation in a subject. 73. The method of claim 72, further comprising the step of monitoring the subject for an inflammatory response or a pain response, wherein an inflammatory response or a pain response indicates a tick infestation. 74. The method of claim 73, wherein an inflammatory response is indicated by red skin or swollen skin. 75. The method of claim 72 where ticks do not remain attached to the subject as a result of detecting the tick infestation. 76. The method of claim 72, where tick feeding is reduced or impaired as a result of detecting the tick infestation. 77. The method of claim 72, where transmission of tick borne pathogens is reduced or impaired as a result of detecting the tick infestation. 78. The method of any one of claims 72-74, wherein detecting tick infestation occurs less than 72 hours after tick infestation. 79. The method of any one of claims 72-74, wherein a pain response at a site of tick infestation is detected by the subject. 80. A method of decreasing the ability of a tick to feed on a subject comprising: administering to the subject an immunogenic composition comprising one or more cystatin antigens, wherein the cystatin corresponds to a polypeptide comprising an amino acid sequence which is at least 80% homologous to the amino acid sequence of SEQ. ID. NO. 2; thereby decreasing the ability of a tick to feed on a subject. 81. A method for preventing a tick infestation in a subject comprising: administering to the subject an immunogenic composition comprising one or more cystatin antigens, wherein the cystatin corresponds to a polypeptide comprising an amino acid sequence which is at least 80% homologous to the amino acid sequence of SEQ. ID, NO. 2; thereby preventing a tick infestation in a subject. 82. A method for decreasing a risk for a tick borne disease in a subject comprising: administering to the subject an immunogenic composition comprising one or more cystatin antigens, wherein the cystatin corresponds to a polypeptide comprising an amino acid sequence which is at least 80% homologous to the amino acid sequence of SEQ. ID. NO. 2; andmonitoring the subject for an inflammatory response or a pain response wherein an inflammatory response or a pain response indicates a risk for tick-borne disease in a subject. 83. The method of claim 82, wherein the disease is a member selected from the group consis... | Rhode Island Board of Governors for Higher Education | A61B005/00 A61K039/00
US20100278730 | A1 | 2010-11-04 | Non-Invasive Molecular Imaging of Cellular Histone Deacetylase Substrate Using Magnetic Resonance Spectroscopy (MRS) or Positron Emission Tomography (PET) | 2008-04-08 | We disclose methods of detecting histone deacetylase activity in a mammal by administering to the mammal a compound comprising at least one atom having a nucleus detectable by magnetic resonance spectroscopy, wherein the compound is a substrate of histone deacetylase; and observing the compound or a cleavage product thereof in at least a portion of the body of the mammal by magnetic resonance spectroscopy (MRS). We also disclose methods of detecting histone deacetylase activity in a mammal by administering to the mammal a compound comprising at least one positron-emission-decaying radioisotope, wherein the compound is a substrate of histone deacetylase; and observing the compound or a cleavage product thereof in at least a portion of the body of the mammal by positron emission tomography (PET). We also disclose compounds useful as histone deacetylase substrates. | 1. A method of detecting a histone deacetylase activity in a mammal, comprising: administering to the mammal a compound comprising at least one atom having a nucleus detectable by magnetic resonance spectroscopy, wherein the compound is a substrate of a histone deacetylase; andobserving the compound or a cleavage product thereof in at least a portion of the body of the mammal by magnetic resonance spectroscopy (MRS). 2. The method of claim 1, wherein the compound is selected from the group consisting of Boc-lysine trifluoroacetic acid (BLT); compounds having structure I:wherein R1 is selected from the group consisting of —CH3, —CH2X, —CHX2, —CX3, —(CH2)mCH3, —CH2(CH3)2, and -Ph; R2 is selected from the group consisting of -Ph, -PhX, -PhN(CH3)2, -PhN+(CH3)3, -PhNO2, -PhC≡C, -triazolyl-(CH2)mX, -Ph-triazolyl-(CH2)mX, -Ph(CH2)m, -Ph(CH2)mX, -Ph(CH2)mN(CH3)2, -Ph(CH2)m, -PhN+(CH3)3, -Ph(CH2)mNO2, -Ph(CH2)m, -PhC≡C, -triazolyl-(CH2)mX, -Ph-triazolyl-(CH2)mX, -Ph(CH2)m(CH3CONH), -Ph(CH2)m(CH2XCONH), -Ph(CH2)m(CHX2CONH), -Ph(CH2)m(CX3CONH), —(CH═CH)(CH2)mNH2, -piperidinyl, and -piperidinyl-X; and R3 is selected from the group consisting of —Cnalkyl, -aryl, and -aryl-Cnalkyl; wherein each X is selected from the group consisting of —F and —Br; m is an integer from 1 to 5, inclusive; and n is an integer from 0 to 10, inclusive; and salts and esters thereof. 3. The method of claim 1, wherein the mammal suffers a tumor in the portion of the body, and further comprising administering to the mammal a histone deacetylase substrate. 4. The method of claim 3, wherein the histone deacetylase substrate is selected from the group consisting of compounds having structure I:wherein R1 is selected from the group consisting of —CH3, —CH2X, —CHX2, —CX3, —(CH2)mCH3, —CH2(CH3)2, and -Ph; R2 is selected from the group consisting of -Ph, -PhX, -PhN(CH3)2, -PhN+(CH3)3, -PhNO2, -PhC≡C, -triazolyl-(CH2)mX, -Ph-triazolyl-(CH2)mX, -Ph(CH2)m(CH2)mX, -Ph(CH2)mN(CH3)2, -Ph(CH2)m, -PhN+(CH3)3, -Ph(CH2)mNO2, -Ph(CH2)m, -PhC≡C, -triazolyl-(CH2)mX, -Ph-triazolyl-(CH2)mX, -Ph(CH2)m(CH3CONH), -Ph(CH2)m(CH2XCONH), -Ph(CH2)m(CHX2CONH), -Ph(CH2)m(CX3CONH), —(CH═CH)(CH2)mNH2, -piperidinyl, and -piperidinyl-X; and R3 is selected from the group consisting of —Cnalkyl, -aryl, and -aryl-Cn alkyl; wherein each X is selected from the group consisting of —F and —Br; m is an integer from 1 to 5, inclusive; and n is an integer from 0 to 10, inclusive; and salts and esters thereof. 5. The method of claim 4, wherein the pairing of the histone deacetylase and the histone deacetylase substrate is selected from the group consisting of HDAC 1 and 6-(fluoroacetamido)-1-hexanoicanilide (FAHA); HDAC-4 and FAHA; HDAC-5 and FAHA; HDAC-6 and FAHA; HDAC-8 and FAHA; HDAC-9 and FAHA; HDAC-4 and 6-(trifluoroacetamido)-1-hexanoicanilide (3FAHA); HDAC-5 and 3FAHA; HDAC-7 and 3FAHA; HDAC-8 and 3FAHA; HDAC-9 and 3FAHA; HDAC-11 and 3FAHA; HDAC-3 and 6-(methylacetamido)-1-hexanoicanilide (EAHA); sirtuin-2 and EAHA; HDA... | - | A61K051/00 A61B005/055 C07C237/28
US07850625 | B2 | 2010-12-14 | Method and apparatus for monitoring labor parameter | 2004-08-06 | A method for early detection of a pregnancy complication, including touching a position sensor to a point on a fetal presenting part of a fetus in a mother, and capturing a position of the position sensor, touching the position sensor to a set of points on the mother and capturing the position of the position sensor at each point, and detecting a pregnancy complication sign based upon a predefined criterion for said pregnancy complication. Methods and apparatus are provided for identifying the BPD pattern in an ultrasound image, for determining characteristics of body parts outside of a pelvic region, for BPD reconstruction and for an adapter for a position sensor are also described among other embodiments. | 1. A method for early detection of a pregnancy complication, the method comprising: touching a position sensor to a point on a fetal presenting part of a fetus in a mother, and capturing a position of the position sensor;touching the position sensor to a set of points on the mother and capturing the position of the position sensor at each point;obtaining a correlation of the position of the position sensor on said fetal presenting part to the positions of the position sensor at each point on the mother; anddetecting a pregnancy complication sign based upon a predefined criterion for said pregnancy complication, wherein said pregnancy complication is related to said correlation. 2. The method according to claim 1, wherein said predefined criterion comprises at least one of contractions with a predetermined frequency, cramping, pelvic pressure, excessive vaginal discharge, back pain, premature rupture of membrane (PROM), cervical dilation greater than a predefined amount, and effacement greater than a predefined amount. 3. The method according to claim 1, wherein said pregnancy complication comprises preterm labor or threatened abortion. 4. The method according to claim 1, further comprising after detecting the pregnancy complication sign, performing therapy to delay delivery and to improve fetal survival and outcome. 5. The method according to claim 4, wherein performing therapy comprises tocolytic therapy. 6. The method according to claim 4, wherein performing therapy comprises corticosteroids therapy. 7. The method according to claim 4, further comprising monitoring an effect of the therapy. | Trig Medical Ltd. | A61B005/103
US07828728 | B2 | 2010-11-09 | Analyte sensor | 2007-02-14 | The present invention relates generally to membranes utilized with implantable devices, such as devices for the detection of analyte concentrations in a biological sample. More particularly, the invention relates to novel silicone-hydrophilic polymer blend membranes, and to devices and implantable devices including these membranes. | 1. An implantable glucose sensor, the sensor comprising: an electroactive surface configured for insertion into a host's body; anda membrane system comprising a cell impermeable domain disposed between the electroactive surface and the host's body when implanted, wherein the cell impermeable domain comprises a blend of a silicone polymer and a hydrophilic-hydrophobic polymer, wherein the blend is configured such that the cell impermeable domain allows transport of glucose therethrough, wherein the membrane system further comprises a diffusion resistance domain, and wherein glucose permeability of the cell impermeable domain is greater than a glucose permeability of the diffusion resistance domain. 2. The sensor of claim 1, wherein the membrane system comprises an enzyme domain positioned between the electroactive surface and the cell impermeable domain, wherein the enzyme domain comprises an enzyme. 3. The sensor of claim 1, wherein the diffusion resistance domain is positioned between the electroactive surface and the cell impermeable domain. 4. The sensor of claim 1, wherein the diffusion resistance domain comprises a silicone material configured to control flux of glucose the analyte therethrough. 5. The sensor of claim 1, wherein the hydrophilic-hydrophobic polymer comprises a hydrophilic copolymer. 6. The sensor of claim 5, wherein the hydrophilic copolymer comprises hydroxy substituents. 7. The sensor of claim 5, wherein the hydrophilic copolymer comprises a poly(ethylene oxide)-poly(propylene oxide) copolymer. 8. The sensor of claim 5, wherein the hydrophilic copolymer comprises a poly(ethylene oxide)-poly(propylene oxide)-poly(ethylene oxide) triblock polymer. 9. The sensor of claim 5, wherein at least a portion of the hydrophilic copolymer is at least partially cross-linked. 10. The sensor of claim 5, wherein from about 1% w/w to about 50% w/w of the cell impermeable domain is the hydrophilic copolymer. 11. The sensor of claim 5, wherein from about 5% w/w to about 30% w/w of the cell impermeable domain is the hydrophilic copolymer. 12. The sensor of claim 1, wherein the sensor is configured to be subcutaneously implanted. 13. The sensor of claim 1, wherein the sensor is configured to be intravascularly implanted. 14. A device configured for continuously measuring a glucose concentration in a host, the device comprising: an electrode configured for insertion into a host's body; anda membrane comprising a silicone material comprising a micellar jacket structure configured to attenuate or substantially block non-constant noise-producing interferents, wherein the micellar jacket structure comprises a macromolecular self-organization of silicone globules coated with a hydrophilic polymer. 15. The sensor of claim 14, wherein the silicone material comprises a blend of a silicone elastomer and a hydrophilic copolymer. 16. The sensor of claim 14, wherein the membrane comprises an enzyme domain positioned between the electroactive s... | DexCom, Inc. | A61B005/00 A61B005/05
US20100280382 | A1 | 2010-11-04 | BONE FRACTURE DETECTION | 2009-01-04 | Apparatus for detecting a fracture in a bone, the apparatus comprising: an electric field generator, configured to generate an electric field in proximity of the bone, at a predefined frequency range, the generated electric field being substantially parallel to longitudinal axis of the bone; an electric field measurer, associated with the electric field generator, and configured to measure intensity values of the generated electric field over the predefined frequency range, and an analyzer, associated with the electric field measurer, and configured to analyze the intensity values measured over the predefined frequency range, thereby to detect the fracture in the bone. | 1. Apparatus for detecting a fracture in a bone, the apparatus comprising: an electric field generator, configured to generate an electric field in proximity of the bone, at a predefined frequency range, said generated electric field being substantially parallel to longitudinal axis of the bone;an electric field measurer, associated with said electric field generator, and configured to measure intensity values of said generated electric field over said predefined frequency range, andan analyzer, associated with said electric field measurer, and configured to analyze said intensity values measured over said predefined frequency range, thereby to detect the fracture in the bone. 2. The apparatus of claim 1, wherein said frequency range is between 100 MHz and 20 GHz. 3. The apparatus of claim 1, wherein said frequency range is between 200 MHz and 10 GHz. 4. The apparatus of claim 1, wherein said electric field generator is a Radio-Frequency (RF) signal generator. 5. The apparatus of claim 1, wherein said apparatus further comprises at least two electrodes, deployable along said longitudinal axis of the bone and associated with said electric field generator. 6. The apparatus of claim 5, wherein said two electrodes are ring electrodes. 7. The apparatus of claim 1, wherein said apparatus further comprises at least one electrode deployable along said longitudinal axis of the bone and associated with said electric field measurer. 8. The apparatus of claim 7, wherein said electrode associated with said electric field measurer is movable along said longitudinal axis of the bone. 9. The apparatus of claim 7, wherein said electrode associated with said electric field measurer is a ring electrode. 10. The apparatus of claim 1, wherein said analyzer is further configured to control said electric field generator. 11. The apparatus of claim 1, wherein said analyzer is further configured to control said electric field measurer. 12. The apparatus of claim 1, further comprising a storage device, associated with said analyzer, wherein said analyzer is further configured to store said intensity values measured over said predefined frequency range on said storage device. 13. The apparatus of claim 1, further comprising a storage device, associated with said analyzer, wherein said analyzer is further configured to compare said intensity values measured over said predefined frequency range with reference intensity values stored on said storage device in advance of said generation of said electric field, thereby to detect fracture in the bone. 14. The apparatus of claim 13, wherein the bone is a first limb of a person, and said reference intensity values are intensity values previously obtained from a second limb of said person. 15. The apparatus of claim 13, wherein the bone is a limb of a first person, and said reference intensity values are intensity values previously obtained from a limb of a second person. 16. Method for detecting a fracture in a bone... | - | A61B008/00 A61B005/05
US20100324388 | A1 | 2010-12-23 | BODY-WORN PULSE OXIMETER | 2009-09-14 | The invention provides a body-worn system that continuously measures pulse oximetry and blood pressure, along with motion, posture, and activity level, from an ambulatory patient. The system features an oximetry probe that comfortably clips to the base of the patient's thumb, thereby freeing up their fingers for conventional activities in a hospital, such as reading and eating. The probe secures to the thumb and measures time-dependent signals corresponding to LEDs operating near 660 and 905 nm. Analog versions of these signals pass through a low-profile cable to a wrist-worn transceiver that encloses a processing unit. Also within the wrist-worn transceiver is an accelerometer, a wireless system that sends information through a network to a remote receiver, e.g. a computer located in a central nursing station. | 1. A system for monitoring a patient, comprising: an oximetry sensor comprising a photodetector and a first and second light source, the photodetector configured to generate a first signal by detecting radiation from the first light source after it irradiates the patient, and a second signal by detecting radiation from the second light source after it irradiates the patient, the oximetry sensor comprising a first cable;an ECG system comprising an electrical circuit connected to at least two electrodes attached to the patient to measure an ECG signal, the ECG system configured to be positioned on a location selected from the patient's arm and chest and comprising a second cable;an oscillometric blood pressure monitoring system comprising a pump and a cuff attached to the patient to measure an oscillometric blood pressure value, the oscillometric blood pressure monitoring system configured to be positioned on the patient's arm and comprising a third cable;a motion sensor configured to generate a motion signal representing motion of a part of the patient's body it is attached to;a processing unit configured to be worn on the patient's body and connected to the oximetry sensor through the first cable to receive the first and second signals, to the ECG system through the second cable to receive the ECG signal, to the oscillometric blood pressure monitoring system through the third cable to receive an oscillometric blood pressure value, and to the motion sensor to receive the motion signal, the processing unit comprising a processor configured to process: i) the first and second signals to determine a value for oxygen saturation; ii) the ECG waveform, oscillometric blood pressure value, and at least one of the first and second signals to determine a blood pressure value; and iii) the motion signal to determine one of the patient's posture and activity level. 2. A system for monitoring oxygen saturation from a patient, comprising: an optical sensor configured to generate an optical signal;an ECG sensor configured to be worn on the patient's chest and generate a digital ECG signal;a motion sensor configured to generate a motion signal; anda processing unit configured to be worn on the patient's body and comprising separate input ports that connect to the optical sensor, the ECG sensor and motion sensor, the processing unit configured to determine at least two unique vital signs by collectively processing the optical signal, ECG signal, and motion signal. 3. The system of claim 1, wherein the processing unit is configured to be worn on the patient's wrist. 4. The system of claim 3, wherein the processing unit comprises a first input port to receive the first cable, a second input port to receive the second cable, and a third input port to receive the third cable. 5. The system of claim 4, wherein the second and third input ports are both serial ports configured to receive digital signals. 6. The system of claim 5, wherein the second and third input p... | - | A61B005/1455 A61B005/0402
US07831443 | B2 | 2010-11-09 | Method and computer program product for measuring and utilizing efficiency of medical resource and services providers | 2005-12-16 | A method and computer program product is provided that, in various embodiments, may provide a measure of efficiency of resources used in medical treatment plan options. A normalized Relative Resource Value (RRV) may be obtained that measures the level of resources required to satisfy a “normal” case for a particular aspect of medical care. Applying this value against actual costs allows ranking and/or comparison of health care providers relative to one another on the basis of the price or resources applied or utilized. Moreover, applying the RRV against the quality of outcome allows comparison and ranking of providers on overall value. The invention may be applied to the components of medical care comprising, inter alia, acute inpatient, non-acute inpatient, outpatient, professional office-based care including scheduled outpatient, and pharmacy/prescriptions. The normalized RRV obtained under various embodiments of the invention may measure resource use for medical treatment plan options/resources independent of price. The RRV may be relative across, as well as within, each of the components of medical care. Thus, provider profiling generally as well as profiling across treatment plan options based on efficiency generally is within the scope of various embodiments of the invention. Moreover, providers may be compared based on efficiency for specific conditions, as well as across conditions, across treatment plan options and across components of care using various embodiments of the invention. | 1. A method of measuring and utilizing efficiency of resources utilized by health care providers in the care of medical conditions, comprising: providing a programmed digital computer for compiling a patient database comprising data from more than one component of care, wherein the component of care data further comprises data from at least one provider;providing at least one component of care analysis module;selecting at least one component of care to analyze;providing a programmed digital computer for extracting data corresponding to the at least one selected component of care from the patient database;providing a programmed digital computer for providing the extracted data to the appropriate component of care analysis module;providing a programmed digital computer for identifying and treating outliers from the data;providing a programmed digital computer for generating a modular Relative Resource Value (RRV) using the selected component of care analysis module;providing a programmed digital computer for creating an adjusted modular RRV from the modular RRV to allow inter-modular and intra-modular comparison by provider;providing a programmed digital computer for compensating for preventive care and chronic care medical resources;providing a programmed digital computer for performing inter-modular and/or intra-modular comparison of the adjusted modular RRV by provider;performing the inter-modular and intra-modular comparison of the adjusted modular RRV;providing a programmed digital computer for ranking providers within at least one component of care and/or across all components of care based on the adjusted RRV;wherein the at least one component of care analysis module further comprises: an acute inpatient analysis module; a non-acute inpatient analysis module; a hospital outpatient analysis module; a professional and scheduled outpatient analysis module; and a pharmacy analysis module;wherein the acute inpatient analysis module further comprises: extracting acute inpatient data from the patient database;summing the extracted data to admission level;extracting DRG weights from the patient database;merging DRG weights with extracted inpatient data;identifying outliers, minimum volume data points and ungroupable DRG's;creating a conversion factor using non-outlier and ungroupable data;creating the RRV for outliers, non-outliers, minimum volume data points and ungroupable DRGs; andadjusting the admission level modular RRV for the acute inpatient and non-acute inpatient components of care to an encounter level RRV by allocating the admission level modular RRV to the encounter level. 2. The method of claim 1, wherein the more than one component of care further comprises: acute inpatient; non-acute inpatient; hospital outpatient; professional and scheduled outpatient; and pharmacy, wherein the acute inpatient and non-acute inpatient components of care adjusted RRVs further comprise an adjusted modular RRV, wherein an admission level RRV is adjusted... | Group Health Plan, Inc. | G06Q100/00 G06Q050/00 A61B005/00 G06F019/00
US07831297 | B2 | 2010-11-09 | Guide wire torque device | 2003-05-24 | A device for gripping a guide wire includes a body portion having an opening therethrough, and has an open position and a closed position. When the device is in the open position the guide wire may be passed through the opening. When the device is in the closed position the guide wire is gripped at more than one location thereby allowing force or torque to be applied to the guide wire. A device according to the invention may optionally, or instead, include (a) a snout to act as a guide wire director, (b) a structure, such as an opening, configured to receive an end of a guide wire package, (c) a structure for securing a body member to the device, and/or (d) a mechanism for advancing or retracting a guide wire. If the device includes one of these optional features, it need not be designed to grip a guide wire at two or more locations. | 1. A device for gripping a guide wire at more than one location, the device comprising: (a) a body portion including: (i) a first end, a second end and a substantially uniform opening therethrough;(ii) a first collet at the first end of the body portion, the first collet having a plurality of compressible fingers and a first opening therebetween; and(iii) a second collet at the second end of the body portion, the second end having a plurality of compressible fingers and a second opening therebetween, the second opening being substantially the same size as the first opening when the plurality of compressible fingers of the first collet is fully compressed and the plurality of compressible fingers of the second collet is fully compressed; and(b) a first cap for engaging the first collet and compressing the fingers of the first collet when the first cap is tightened to the body portion, the first cap configured to receive an end of a guide wire package and including an opening through which a guide wire can pass; and(c) a second cap for engaging the second collet and compressing the fingers of the second collet when the second cap is tightened to the body portion, the second cap including (i) a snout, (ii) a structure for securing a body part thereto and (iii) an opening through which a guide wire can pass. 2. A device for gripping a guide wire at more than one location, the device comprising: (a) a body portion including: (i) a first end and a second end;(ii) a first collet at the first end of the body portion, the first collet having a plurality of compressible fingers and a first opening therebetween; and(iii) a second collet at the second end of the body portion, the second end having a plurality of compressable fingers and a second opening therebetween, the second opening being substantially the same size as the first opening when the plurality of compressible fingers of the first collet is fully compressed and the plurality of compressible fingers of the second collet is fully compressed; and(b) a first cap for engaging the first collet and compressing the fingers of the first collet when the first cap is tightened to the body portion and including an opening through which a guide wire can pass; and(c) a second cap for engaging the second collet and compressing the fingers of the second collet when the second cap is tightened to the body portion, the second cap including an opening through which a guide wire can pass. 3. The device of claim 2 wherein one of the caps includes an opening dimensioned to receive an end of a guide wire package. 4. The device of claim 2 that further comprises an elongated snout. 5. The device of claim 4 wherein the snout is formed in one of the caps. 6. The device of claim 2 that further includes a structure for securing a body member to the device. 7. The device of claim 2 that is comprised of plastic. 8. The device of claim 2 wherein the body portion further comprises threads on the first end juxtaposed the... | Scottsdale Medical Devices, Inc. | A61B005/00 A61M025/00
US07831287 | B2 | 2010-11-09 | Dual electrode system for a continuous analyte sensor | 2008-04-28 | Disclosed herein are systems and methods for a continuous analyte sensor, such as a continuous glucose sensor. One such system utilizes first and second working electrodes to measure additional analyte or non-analyte related signal. Such measurements may provide a background and/or sensitivity measurement(s) for use in processing sensor data and may be used to trigger events such as digital filtering of data or suspending display of data. | 1. A continuous glucose sensor configured for insertion into a host and for detecting glucose in the host, the sensor comprising: a first working electrode comprising a first electroactive surface disposed beneath an active enzymatic portion of a membrane system, wherein the first electroactive surface is configured to measure a measurable species and a noise-causing species; anda second working electrode comprising a second electroactive surface disposed beneath at least one of an inactive enzymatic portion of the membrane system or a non-enzymatic portion of the membrane system, wherein the second electroactive surface is configured to measure a noise-causing species, and wherein the membrane system comprises a membrane diffusion barrier that separates the first working electrode and the second working electrode such that diffusion of the measurable species between the first working electrode and the second working electrode is substantially blocked. 2. The sensor of claim 1, wherein the noise-causing species comprises at least one member selected from the group consisting of externally produced H2O2, urea, lactic acid, phosphates, citrates, peroxides, amino acids, amino acid precursors, amino acid break-down products, nitric oxide, NO-donors, NO-precursors, reactive oxygen species, compounds having electroactive acidic, amine or sulfhydryl groups, acetaminophen, ascorbic acid, dopamine, ephedrine, ibuprofen, L-dopa, methyldopa, salicylate, tetracycline, tolazamide, tolbutamide, and triglycerides. 3. The sensor of claim 1, wherein the noise-causing species is non-constant. 4. The sensor of claim 1, wherein the measurable species is H2O2 produced in an active enzymatic portion of the membrane system. 5. The sensor of claim 1, wherein the first electroactive surface and the second electroactive surface are each configured to integrate noise detected about a circumference of the sensor. 6. The sensor of claim 1, wherein the membrane diffusion barrier comprises a discontinuous portion of the membrane disposed between the first electroactive surface and the second electroactive surface. 7. The sensor of claim 1, wherein the membrane diffusion barrier comprises a first resistance domain formed on the first working electrode and a second resistance domain formed on the second working electrode, wherein the first resistance domain and the second resistance are independently formed. 8. The sensor of claim 7, wherein the first resistance domain and the second resistance domain are configured to attenuate diffusion of the measurable species. 9. The sensor of claim 1, further comprising an insulator configured to insulate the first working electrode from the second working electrode. 10. The sensor of claim 9, wherein the sensor membrane is the insulator. 11. The sensor of claim 1, wherein the first working electrode and the second working electrode are spaced apart by less than about 0.12 inches. 12. The sensor of claim 1, wherein the first wor... | DexCom, Inc. | A61B005/00
US07833171 | B2 | 2010-11-16 | Method and apparatus for penetrating tissue | 2007-02-13 | These and other objects of the present invention are achieved in a body fluid sampling system for use on a tissue site that includes an electrically powered drive force generator. A penetrating member is operatively coupled to the force generator. The force generator moves the member along a path out of a housing having a penetrating member exit, into the tissue site, stops in the tissue site, and withdraws out of the tissue site. A cartridge houses the penetrating member. The cartridge has first and second seals coupled to the penetrating member to maintain a sterile environment around a portion of the penetration member prior to penetrating member actuation. A user interface is configured to relay at least one of, penetrating member performance or a penetrating member setting. | 1. A body fluid sampling system for use on a tissue site, the system comprising: an electrically powered drive force generator;a penetrating member operatively coupled to said electrically powered drive force generator, said electrically powered drive force generator moving said member along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;a disposable housing said penetrating member, said disposable having a first and a second seal formed at a distal end and a proximal end of said disposable respectively to maintain a sterile environment around a portion of the penetration member prior to penetrating member actuation;a user interface including a sensor for detecting and responding to audio commands from a user; anda processor coupled to and receiving a signal from said user interface, said processor monitoring and controlling velocity of said penetrating member via said electrically power drive force generator in response to said signal. 2. The system of claim 1, wherein the user interface relays information to a user via an audio device. 3. The system of claim 1, wherein the user interface relays information to a user via a wireless device. 4. The system of claim 1, further comprising: the processor coupled to the electrically powered drive force generator and configured to change a direction and magnitude of force exerted on the penetrating member as it penetrates or advances through the target tissue. 5. The system of claim 1, further comprising: a penetrating member sensor configured to detect a penetrating member position during target tissue penetration. 6. The system of claim 1, wherein the at least one output is selected from, a penetration event of a penetrating member, number of penetrating members remaining, time of day, alarm, penetrating member trajectory waveform profile information, force for last penetration event, the last penetration event, how or low battery status, analyte status, time to change support status, jamming malfunction, and system status. 7. The system of claim 1, wherein the user interface is selected from an LED, an LED digital display, an LCD display, a sound generator, a buzzer, and a vibrating device. 8. The system of claim 1, wherein the housing is selected from at least one of, a telephone, a watch, a PDA, electronic device, medical device, point of care device and a decentralized diagnostic device. 9. The system of claim 1, further comprising: an input device coupled to the housing, the input device selected from one or more pushbuttons, a touch pad independent of the display device, or a touch sensitive screen on a visual display. 10. The system of claim 1, further comprising: a data exchange device for coupling the tissue penetrating system to support equipment. 11. The system of claim 10, further comprising: a data interface configured to couple the tissue penetrating system to support equipme... | Pelikan Technologies, Inc. | A61B005/00
US07850622 | B2 | 2010-12-14 | Tissue penetration device | 2005-12-22 | A moving coil tissue penetration device includes a magnetic source that produces a magnetic field in a magnetically active region. A cylindrical coil is secured to a translation substrate and disposed at least partially within the magnetically active region. A sharpened member is configured to penetrate tissue and mechanically coupled to the translation substrate. | 1. A lancing system comprising a controllable driver that includes a driver coil pack, the controllable driver having a proximal end and a distal end with a distal end of a coupler shaft disposed at the distal end of the controllable driver;a ratchet drive mechanism operably disposed at the distal end of the controllable driver; anda plurality of lancets configured to be sequentially advanced by the ratchet drive mechanism into a coupled configuration with the coupler shaft of the controllable driver, each of a lancet being configured to be coupled to the controllable driver to be advanced from a housing to penetrating a skin surface and create an aperture from which a body fluid flows to an interior of the housing that includes a plurality of sample chambers. 2. The system of claim 1, wherein the plurality of lancets are disposed within a lancet belt. 3. The system of claim 2, wherein the lancet belt is disposed within a lancet belt cartridge having a supply canister for storage of the unused lancets of the lancet belt and a receptacle canister for acceptance and storage of the used lancets of the lancet belt. 4. The system of claim 1, wherein the plurality of lancets are disposed within a lancet module belt. 5. The system of claim 4, wherein the lancet module belt is disposed within a lancet module cartridge having supply canister for the unused lancet modules and a receptacle canister for acceptance and storage of used lancet modules. 6. The system of claim 1, wherein the plurality of lancets are disposed within a sampling module belt comprising a plurality of sampling modules interconnected by a flexible belt member. 7. The system of claim 6, wherein the sampling module belt is disposed within a sampling module cartridge having supply canister for the unused sampling modules and a receptacle canister for acceptance and storage of used sampling modules. 8. The system of claim 1, wherein the plurality of lancets further comprise a drive head secured to a proximal end of each lancet and configured to be engaged by a drive coupler secured to a distal end of the coupler shaft. 9. The system of claim 8, wherein the drive heads are substantially cylindrical in shape and the drive coupler comprises a T-slot drive coupler. 10. A lancing system comprising a controllable driver having a proximal end and a distal end with a distal end of a coupler shaft disposed at the distal end of the controllable driver;a ratchet drive mechanism operably disposed at the distal end of the controllable driver; anda plurality of lancets configured to be sequentially advanced by the ratchet drive mechanism into a coupled configuration with the coupler shaft of the controllable driver, the plurality of lancets are disposed within a sampling module belt comprising a plurality of sampling modules interconnected by a flexible belt member, andwherein the sampling modules of the sampling module belt further comprise an analytical site and sensor contacts and the ratche... | Pelikan Technologies, Inc. | A61B005/00
US08768425 | B2 | 2014-07-01 | Device and method for identification of meconium in amniotic fluid | 2008-10-05 | Disclosed are methods and devices comprising an elongated probe for the in vivo detection of meconium in amniotic fluid held in an amniotic sac based on detecting the presence of Zinc-coproporphyrin I (ZnCP). ZnCP is excited at approximately 405 nm and emits characteristic fluorescence, centered at about 580 nm and less intensely at 630 nm. | 1. A device useful for in vivo identification of the presence of zinc-coproporphyrin I (ZnCP) in amniotic fluid, comprising: a) an excitation unit comprising a light source that, when turned on, produces an excitation light including a wavelength of about 405 nm,b) a detection unit capable of detecting an intensity of light emitted from in vivo amniotic fluid indicative of the presence of ZnCP, said indicative emitted light including a wavelength of about 580 nm, andc) at least one elongated probe having a distal end and a proximal end, said probe comprising an excitation light guide capable of guiding said excitation light from said light source out through the distal end of said probe and thereby causing the excitation light to pass through the wall of an amniotic sac and thereby illuminate the amniotic fluid in vivo,the distal end of said probe being so dimensioned as to pass into or through an endocervical canal, andd) a processor configured to correlate said intensity with the concentration of ZnCP in said amniotic fluid. 2. The device of claim 1, wherein said light source, when turned on, produces an excitation light that consists essentially of monochromatic light of about 405 nm. 3. The device of claim 1, wherein the diameter of said distal end of said elongated probe is no more than about 6 mm (18 French). 4. The device of claim 1, said probe including a detection light guide component of said detection unit configured to guide light from a distal tip of said probe to a light detector component of said detection unit. 5. The device of claim 4, wherein said detection light guide and said excitation light guide are discrete. 6. The device of claim 4, wherein said detection light guide and said excitation light guide are substantially the same. 7. The device of claim 1, comprising a device body containing said light source wherein a distal end of said probe is configured to reversibly couple to said device body so that when said probe is coupled to said device body, said probe is configured to direct excitation light through said excitation light guide and out through said distal end of said probe. 8. The device of claim 7, at least one said probe is sterile and packaged in a sterility preserving package. 9. The device of claim 1, wherein the diameter of said distal end of said elongated probe is no more than 4 mm (12 French). 10. The device of claim 1, wherein the diameter of said distal end of said elongated probe is no more than 3 mm (10 French). 11. The device of claim 1, wherein the diameter of said distal end of said elongated probe is no more than 2 mm (6 French). 12. The device of claim 1, wherein the diameter of said distal end of said elongated probe is no more than 1 mm (3 French). 13. The device of claim 1 wherein said probe comprises said light source. 14. The device of claim 1 wherein said light source is external to said probe. 15. A method for quantifying the concentration or identifying the presence of zinc-co... | Tel Hashomer Medical Research Infrastructure and Services Ltd. | A61B005/1455
US08764676 | B2 | 2014-07-01 | Signal processing in physiological noise | 2010-05-07 | The invention relates to systems and methods for estimating a response of at least a part of a physiological system to a first deterministic stimulus signal. The methods include separating a measured first signal into a plurality of segments, each segment representing a response of the physiological system to a corresponding portion of the first stimulus signal and generating a template signal representing the plurality of segments. The methods further include calculating a measure of similarity of each segment in the plurality of segments to the template signal to provide a set of scalar quantities, and determining a metric representing a characteristic of the response of at least a part of the physiological system to the first stimulus signal. | 1. A method for estimating a response of at least a part of a physiological system to a first deterministic stimulus signal, the method comprising: receiving, at a computing device, a first signal measured from the physiological system in response to the first stimulus signal;separating, using a processor, the measured first signal into a plurality of segments, each segment representing a response of the physiological system to a corresponding portion of the first stimulus signal;generating, using a processor, a template signal representing the plurality of segments;calculating, using a processor, a measure of similarity of each segment in the plurality of segments to the template signal to provide a set of scalar quantities, wherein calculating the measure of similarity comprises calculating a correlation between the segment and the template signal; anddetermining, using a processor, a metric representing a characteristic of the response of at least a part of the physiological system to the first stimulus signal, wherein the metric is a ratio of a measure of central tendency to a measure of dispersion, wherein the measure of central tendency and the measure of dispersion are computed from the set of scalar quantities. 2. The method of claim 1, further comprising measuring the first signal from the physiological system using one or more electrodes. 3. The method of claim 1, wherein each segment is substantially equal in duration to other segments in the plurality of segments. 4. The method of claim 1, wherein each segment is substantially equal in length to a corresponding portion of the stimulus signal. 5. The method of claim 1, further comprising generating the template signal by averaging the segments. 6. The method of claim 1, further comprising: receiving, at a computing device, a second signal measured from the physiological system in response to a second deterministic stimulus signal, an intensity of which is greater than the intensity of the first stimulus signal;separating, using a processor, the measured second signal into a plurality of segments, each segment representing a response of the physiological system to a corresponding portion of the second stimulus signal; andgenerating the template signal based on the plurality of segments from the second measured signal. 7. The method of claim 6, wherein the second signal is measured under substantially the same measurement conditions as for the first signal. 8. The method of claim 1, further comprising: receiving, at a computing device, a plurality of third signals measured from a plurality of physiological systems of substantially same type in a population of subjects in response to a third deterministic stimulus signal;separating, using a processor, each of the plurality of third signals into a plurality of segments, each segment representing a response of the corresponding physiological system to a corresponding portion of the third stimulus signal; andgenerating the template si... | Massachusetts Eye & Ear Infirmary | A61B005/05
US09008750 | B2 | 2015-04-14 | Method for determining the ventilation of a lung | 2010-06-22 | In a method for determining the ventilation of a lung of an object under examination by magnetic resonance imaging, at least two first lung-representing image data sets are acquired at different intervals of the breathing phase. The density change of the lung tissue is automatically determined from the signal difference between the image signals of the first image data sets in at least one corresponding region of the first lung-representing image data set. The lung or the thorax volume is automatically determined using at least two of the first image data sets, or using at least an additional second lung-representing image data set in a breathing phase that corresponds with the breathing phase of a first image data set. The localized, quantitative ventilation of the lung is automatically calculated depending on the density change of the lung tissue and the change of the lung or thorax volume. | 1. For a subject respirating in successive respiratory cycles, each comprising inspiration comprising mechanical intake of fresh gas and expiration comprising mechanical expulsion of breathed gas, a method for generating a quantitative indicator of ventilation that indicates a relation between said intake and said expulsion, said method comprising: operating a magnetic resonance (MR) data acquisition unit to acquire at least two lung-representing first MR image data sets from the subject, with one of said first MR data sets being acquired in an inspiratory phase of said respiratory cycles and another of said first MR data being acquired in expiratory phase of said respiratory cycles, each of said first MR image data sets comprising pixels having respective intensity values produced by MR signals originating from lung tissue in a lung of the subject;in a processor, automatically determining a difference in respective intensities between pixels in a region in said one of said first MR image data sets and pixels in a same region in said another of said first MR image data sets, said difference representing a density change of said lung tissue in said region due to said respirating;in said processor, also automatically determining a volume, selected from the group consisting of a volume of a lung of the subject and a volume of the thorax of the subject, from two of said first MR image data sets or from at least one second image data set acquired from the subject; andin said processor, automatically calculating a quantitative indicator of ventilation for said region, that is spatially-resolved by the pixels in said region, as a predetermined mathematical relationship of said density change of said lung tissue in said region and said volume of a lung, and making said spatially-resolved quantitative indicator of ventilation available as an output from said processor in a humanly-perceptible form. 2. A method according to claim 1, comprising operating said magnetic resonance data acquisition unit to acquire at least one of the first image data sets in three spatial dimensions. 3. A method according to claim 1, comprising operating said magnetic resonance data acquisition unit to acquire at least one of the first image data sets in two spatial dimensions. 4. A method according to claim 3, comprising operating said magnetic resonance data acquisition unit to acquire at least two of the first image data sets with said two spatial dimensions and in multislice operation. 5. A method according to claim 1, comprising operating said magnetic resonance data acquisition unit to acquire at least a second image data set representing the lung in three spatial dimensions. 6. A method according to claim 5, comprising generating a 3D model of an outline of a lung from the second image data set, and determining said volume from said outline. 7. A method according to claim 1, comprising operating said magnetic resonance data acquisition unit to acquire said at least... | Siemens Aktiengesellschaft | A61B005/05 A61B005/055 A61B005/08 A61B005/00
US08311610 | B2 | 2012-11-13 | Biopsy tissue marker | 2009-01-22 | A biopsy site marker is disclosed. The biopsy site marker includes a first marker element and a second marker element. The first marker element is configured for detection by a first imaging modality. The second marker element is configured for detection by a second imaging modality different from the first imaging modality. The second marker element may be a non-absorbable wire having a predetermined shape and is substantially engaged with the first marker element. | 1. A biopsy site marker comprising: a first marker element configured for detection by a first imaging modality, wherein the first marker element is a non-absorbable cylinder;a second marker element configured for detection by a second imaging modality different from the first imaging modality, wherein the second marker element is a non-absorbable wire having a first helical shape positioned around a longitudinal extent of the non-absorbable cylinder of the first marker element; anda third marker element having a second helical shape overlapped with the first helical shape and positioned around a longitudinal extent of the non-absorbable cylinder of the first marker element. 2. The marker of claim 1, wherein the first marker element is a polymer. 3. The marker of claim 1, wherein the first marker element is formed entirely from non-magnetic material. 4. The marker of claim 1, wherein the non-absorbable wire has a twisted configuration. 5. The marker of claim 1, wherein the second marker element is a shape-memory material. 6. The marker of claim 1, wherein the third non-absorbable marker element is detectable by a third imaging modality and not detectable by the first imaging modality and not detectable by the second imaging modality. 7. The marker of claim 6, wherein the third marker element is substantially engaged with each of the first marker element and the second marker element. 8. The marker of claim 1, wherein the third marker element is a third marker wire. 9. The marker of claim 8, wherein the third marker wire is substantially engaged with the non-absorbable wire. 10. The marker of claim 8, wherein the third marker wire has a twisted configuration. 11. The marker of claim 8, wherein the non-absorbable wire and third marker wire are braided. 12. The marker of claim 7, wherein the third marker element is a polymer coating. 13. The marker of claim 6, wherein the third imaging modality is one of ultrasonography, fluoroscopy, magnetic resonance imaging, X-ray, and computed tomography. 14. The marker of claim 1, wherein the first imaging modality is one of ultrasonography, fluoroscopy, magnetic resonance imaging, X-ray, and computed tomography. 15. The marker of claim 1, wherein the second imaging modality is one of ultrasonography, fluoroscopy, magnetic resonance imaging, X-ray, and computed tomography. 16. The marker of claim 1, further comprising a bioabsorbable cover. 17. A biopsy site marker comprising: a first marker element configured for detection by a first imaging modality, wherein the first marker element is a non-absorbable cylinder having a cylindrical side surface having at least one hole; anda second marker element configured for detection by a second imaging modality different from the first imaging modality, wherein the second marker element is a non-absorbable peg substantially engaged with the hole of the first marker element, the non-absorbable peg being configured to extend distally outwardly from the hole... | C. R. Bard, Inc. | A61B019/00
US07949382 | B2 | 2011-05-24 | Devices, systems, methods and tools for continuous glucose monitoring | 2010-07-27 | One aspect of the invention provides a glucose monitor having a plurality of tissue piercing elements, each tissue piercing element having a distal opening, a proximal opening and interior space extending between the distal and proximal openings; a sensing area in fluid communication with the proximal openings of the tissue piercing elements; sensing fluid extending from the sensing area into substantially the entire interior space of the tissue piercing elements; and a glucose sensor adapted to detect a concentration of glucose in the sensing fluid within the sensing area. Another aspect of the invention provides a method of in vivo monitoring of an individual's interstitial fluid glucose concentration including the following steps: inserting distal ends of a plurality of tissue piercing elements through a stratum corneum area of the individual's skin, the tissue piercing elements each having a distal opening, a proximal opening, an interior space extending between the distal and proximal openings, and a sensing fluid filling substantially the entire interior space; and sensing a glucose concentration of the sensing fluid. | 1. A method of in vivo monitoring of an individual's interstitial fluid comprising: creating a plurality of fluid paths through a stratum corneum area of the individual's skin, the fluid paths each comprising a distal end in fluid communication with interstitial fluid of the individual, a proximal end in analyte communication with a sensing area located outside of the patient's body, and an interior space extending between the distal and proximal ends of the fluid path; moving sensing fluid out of a sensing fluid reservoir and into the sensing area;holding the sensing fluid in the sensing area in a non-flowing manner;allowing at least one analyte to passively diffuse from the patient's interstitial fluid through the fluid paths and into the sensing area;sensing a concentration of the at least one analyte in the sensing fluid using a sensor located at least partially in the sensing area during the holding step. 2. The method of claim 1 wherein fluid in the fluid paths is not moving when the at least one analyte is passively diffusing there-through. 3. The method of claim 1 wherein there is no fluid movement between the distal ends of the fluid paths and the sensor when the at least one analyte is passively diffusing through the fluid paths and into the sensing area. 4. The method of claim 1 wherein no interstitial fluid is extracted from the individual during the diffusion and sensing steps. 5. The method of claim 1 further comprising displaying analyte concentration information remote from the stratum corneum area of the individual's skin. 6. The method of claim 5 further comprising wirelessly communicating analyte concentration information to a display. 7. The method of claim 1 wherein the sensing step is performed by a sensor in fluid communication with the sensing area and the interior spaces of the fluid paths, the method further comprising calibrating the sensor by moving sensing fluid into the sensing area. 8. The method of claim 7 further comprising moving sensing fluid out of the sensing area as sensing fluid is moved into the sensing area. 9. The method of claim 1 further comprising attaching a device comprising the sensor area and the sensor to the individual with adhesive. 10. The method of claim 1 further comprising providing a membrane located between the distal ends of the fluid paths and the sensor. 11. The method of claim 1 wherein creating the plurality of fluid paths through the stratum corneum area of the individual's skin comprises applying a microneedle array to the stratum corneum area. 12. The method of claim 1 wherein the proximal ends of the fluid paths are in fluid communication with the sensing area. 13. The method of claim 1 wherein the step of moving sensing fluid comprises operating an actuator. 14. The method of claim 1 wherein the step of moving sensing fluid comprises moving sensing fluid having a glucose concentration of between about 0 mg/dl and about 400 mg/dl. 15. The method of claim 1 wherein th... | ArKal Medical, Inc. | A61B005/05
US08115487 | B2 | 2012-02-14 | Apparatus and method for cryogenically cooling a coil on a magnetic resonance imaging system | 2010-07-19 | A magnetic resonance imaging system includes a primary magnet and a secondary magnet operable to produce magnetic fields within a sample being imaged. The MRI system further includes at least one RF coil that is operable to receive electromagnetic frequencies from the sample. The RF coil is formed from tubing that serves as a cooling conduit through which flows a cooling fluid provided by a cooling source. The cooling fluid cools the RF coils to improve imaging of the sample. | 1. An RF receive coil for use with an MRI imaging system, comprising: at least first and second magnetic resonance radio frequency coils operable to receive magnetic resonance signals from an object under examination, wherein each of the coils is formed from tubing which defines a path for flow of a cryogen; andat least one component that connects the tubing of the first and second magnetic resonance radio frequency coils in series so that the cryogen serially circulates through the tubing of the first and second magnetic resonance radio frequency coils;wherein the RF receive coil is configured to interface with the MRI imaging system and convey the received magnetic resonance signals to an RF detector of the MRI imaging system. 2. The apparatus of claim 1, wherein the RF receive coil and the MRI imaging system are different systems. 3. The apparatus of claim 1, wherein the RF coil is a closed system. 4. The apparatus of claim 3, wherein the MRI imaging system is operable to scan a patient when the RF coil is disconnected therefrom. 5. The apparatus of claim 1 wherein the first and second coils form a volume coil. 6. The apparatus of claim 5 wherein the volume coil is a head coil. 7. The apparatus of claim 1 further including: a third magnetic resonance radio frequency coil operable to receive magnetic resonance signals from the object under examination, wherein the third coil is formed from tubing which defines a path for the flow of the cryogen;a fourth magnetic resonance radio frequency coil operable to receive magnetic resonance signals from the object under examination, wherein the fourth coil is formed from tubing which defines a path for the flow of the cryogen;wherein the tubing of the third and fourth coils is connected fluidly in series, and wherein the first, second, third, and fourth coils are electrically insulated from each other. 8. The apparatus of claim 7 wherein the tubing of the third and fourth coils is connected fluidly in series with the tubing of the first and second coils. 9. The apparatus of claim 8 further including an electrical insulator disposed electrically between the first, second, third, and fourth coils, wherein the electrical insulator includes passages which provide series fluid connections between the tubing of the first, second, third, and fourth coils. 10. The apparatus of claim 1 further including: at least a third magnetic resonance radio frequency coil operable to receive magnetic resonance signals from the object under examination, wherein the third coil is formed from a solid material. 11. The apparatus of claim 10, wherein the further including: at least a third magnetic resonance radio frequency coil operable to receive magnetic resonance signals from the object under examination, wherein the third coil is formed from a solid material. 12. The apparatus of claim 10, wherein the solid material includes a wire. 13. The apparatus of claim 10, further comprising: a cooling conduit positioned a... | m2m Imaging Corporation | G01V003/00
US08131353 | B2 | 2012-03-06 | Methods and apparatuses for navigating the subarachnoid space | 2010-08-31 | Disclosed is a method of navigating a spinal subarchnoid space in a living being, that includes percutaneously introducing a device into the spinal subarachnoid space at an entry location. The device has a first passageway that is sized to slidably receive, and work with, at least a guidewire. The device can be a catheter or a sheath. The method can also include advancing the device within the spinal subarachnoid space at least more than 10 centimeters from the entry location. Alternatively, the method can include advancing the device within the spinal subarachnoid space to facilitate intracranial access with a second device introduced through the first passageway. Also disclosed is a device suited for attachment to a patient's skin, such as a sheath, that includes an elongated member, a skin-attachment apparatus having a flexible skin-attachment flap, and a valve apparatus. | 1. A method of achieving intracranial access from an entry location in a spinal subarachnoid space, comprising: introducing a guidewire into the spinal subarachnoid space;introducing a device over the guidewire and into the spinal subarachnoid space;advancing the device from the spinal subarachnoid space into an intracranial subarachnoid space;introducing a penetration apparatus through the device;puncturing pia mater using the penetration apparatus;advancing a detector through the device and through the pia mater; andplacing the detector on or in brain tissue. 2. The method of claim 1, where the detector is an electroencephalography electrode and the method further comprises: monitoring electrical activity using the electroencephalography electrode. 3. The method of claim 1, where the detector is a sensor and the method further comprises: monitoring a biochemical property using the sensor. 4. The method of claim 3, where the biochemical property comprises pH, glucose concentration, oxygen tension, carbon dioxide concentration, or sodium concentration. 5. The method of claim 1, where the detector is a thermal sensor and the method further comprises: monitoring temperature using the sensor. 6. The method of claim 1, further comprising: monitoring neurotransmitter concentration using the detector. 7. The method of claim 1, where the detector includes an anchoring mechanism. 8. A method of achieving intracranial access from an entry location in a spinal subarachnoid space, comprising: introducing a guidewire into the spinal subarachnoid space;introducing a first device over the guidewire and into the spinal subarachnoid space;advancing the first device from the spinal subarachnoid space into an intracranial subarachnoid space;introducing a penetration apparatus through the first device;puncturing pia mater using the penetration apparatus;advancing a second device through the first device and through the pia mater; andpositioning the second device proximate a brain, a ventricular system, or a cranial nerve. | Board of Regents, The University of Texas System | A61B005/0476
US08603003 | B2 | 2013-12-10 | Trachea pressure determination method and device | 2009-06-03 | According to various embodiments, methods and systems for determining pressure in the lungs may employ intracuff pressure measurements. The intracuff pressure measurements may calibrated or adjusted based on a set of calibration coefficients or a set of calibration curves, which may reflect patient parameters and cuff/tube geometry factors. The resulting calibration may be used to determine a more accurate estimate of lung pressure, which in turn may be used to control a ventilator and provide breathing assistance to a patient. Also provided are tracheal tubes with couplers or other memory devices for storing calibration information. Such tubes may allow calibration at the level of an individual tracheal tube to account for changes in tube geometry and individual patient factors. | 1. A method for determining trachea pressure comprising: using a processor:determining a monitored cuff pressure in an inflated cuff of a tracheal tube disposed in the trachea of a subject;selecting a relationship from a plurality of relationships between the monitored cuff pressure and the trachea pressure; anddetermining the trachea pressure based upon the selected relationship; anddetermining a lung pressure from the tracheal pressure and an offset, wherein the offset is a reference cuff pressure. 2. The method of claim 1, wherein the selected relationship is a function of the monitored cuff pressure, and wherein the method comprises monitoring a current cuff pressure during respiration of the subject, and determining the trachea pressure based upon the relationship and the current cuff pressure. 3. The method of claim 2, wherein the trachea pressure is determined by multiplying a sensed cuff pressure by a ratio of the empirically determined lung pressure to the cuff pressure determined by the selected relationship. 4. The method of claim 3, wherein the lung pressure is determined by subtracting a reference cuff pressure from the product of the sensed cuff pressure and the ratio. 5. The method of claim 1, wherein the lung pressure is determined by an algorithm that implements an equation generally of the form: PT=rPCs−PCo,wherein PT is the trachea pressure, PCs is a cuff pressure monitored during respiration of the subject, PCo is the reference cuff pressure, and r is a ratio of the lung pressure to a sensed cuff pressure. 6. The method of claim 5, wherein the ratio is a non-constant value. 7. The method of claim 6, wherein the ratio varies with the monitored cuff pressure, and wherein the method includes monitoring the current cuff pressure during respiration of the subject and determining the ratio based upon the monitored cuff pressure. 8. The method of claim 6, wherein the ratio is a function of a relationship between the tracheal tube and the trachea of the subject. 9. The method of claim 1, wherein the reference cuff pressure is the initial pressure of the cuff or the lowest trachea pressure seen in the breathing cycle after ventilation has begun. | Covidien LP | A61B005/08
US08827929 | B2 | 2014-09-09 | Method and apparatus for determining the distensibility of a vessel, lumen or a sphincter | 2010-05-27 | An apparatus for measuring distensibility of a vessel, lumen or sphincter includes a balloon catheter having a balloon located adjacent a distal end of a catheter for inserting in, for example, the sphincter, the distensibility of which is to be determined. The balloon is inflated to respective first and second pressures to distend the sphincter to respective first and second transverse cross-sectional areas which are measured. The first and second pressures are selected so that the relationship between the pressure and the cross-sectional area of the sphincter is linear. The distensibility of the sphincter is determined as a distensibility index, which is the value of c from the equation of a line y=mx+c containing the values of the first and second pressures and areas, where x represents pressure and y represents area. | 1. A method for determining the distensibility of one of a vessel wall, a lumen wall and a sphincter, the method comprising: providing a balloon catheter comprising an elongated catheter extending between a proximal end and a distal end, and an inflatable element defining a hollow interior region when inflated located on the catheter, the inflatable element being adapted for locating in the one of the vessel, the lumen and the sphincter,locating the inflatable element of the balloon catheter in the one of the vessel, the lumen and the sphincter,inflating the inflatable element to a first pressure sufficient to distend the one of the vessel wall, the lumen wall and the sphincter,determining a first transverse cross-sectional area adjacent a portion of the one of the vessel wall, the lumen wall and the sphincter, at which the distensibility is to be determined when the inflatable element is inflated to the first pressure,inflating the inflatable element to a second pressure sufficient to distend the one of the vessel wall, the lumen wall and the sphincter,determining a second transverse cross-sectional area adjacent the portion of the one of the vessel wall, the lumen wall and the sphincter, at which the distensibility is to be determined when the inflatable element is inflated to the second pressure sufficient to distend the one of the vessel wall, the lumen wall and the sphincter,the values of the first and second pressures being different and being values at which the relationship between the transverse cross-sectional area of the inflatable element adjacent the portion of the one of the vessel wall, the lumen wall and the sphincter, at which the distensibility is to be determined and the pressure of an inflating medium in the inflatable element is a linear relationship,determining a linear equation of a line containing the values of the first and second transverse cross-sectional areas of the inflatable element adjacent the portion of the one of the vessel wall, the lumen wall and the sphincter, at which the distensibility is to be determined and the values of the first and second pressures of the inflating medium in the inflatable element corresponding to the first and second values of the transverse cross-sectional area thereof, the line being a plot of the values of the transverse cross-sectional area of the inflatable element adjacent the portion of the one of the vessel wall, the lumen wall and the sphincter, at which the distensibility is to be determined against the values of the corresponding pressure of the inflating medium in the inflatable element, anddetermining the distensibility of the one of the vessel wall, the lumen wall and the sphincter as a distensibility index, the distensibility index being the value of a point of intersection of the axis on which one of the transverse cross-sectional area and the pressure is plotted by the line, the equation of which contains the values of the first and second transverse cross-sectional a... | Flip Technologies Limited | A61B005/103 A61B005/117
US08903488 | B2 | 2014-12-02 | System and method for synchronizing energy delivery to the cardiac rhythm | 2010-05-28 | A system for synchronizing application of treatment signals with a cardiac rhythm is provided. The system includes a memory that receives and stores a synchronization signal indicating that a predetermined phase such as R-wave of a cardiac rhythm of a patient has started. A synchronization module analyzes whether the stored synchronization signal is erroneous and if so, prevents a medical treatment device from applying a treatment energy signal such as an IRE pulse to a patient to take into account an irregular heart beat and noise in the synchronization signal in order to maximize safety of the patient. | 1. A system for synchronizing application of treatment signals with a cardiac rhythm, comprising: a memory that receives and stores a synchronization signal indicating that a predetermined phase of a cardiac rhythm of a patient has started; anda synchronization module coupled to the memory and operable to automatically extend a blanking period when the stored synchronization signal is determined to be received during the blanking period. 2. The system of claim 1, wherein the synchronization module is operable to dynamically extend the blanking period when a new synchronization signal is received before the blanking period expires. 3. The system of claim 1, wherein the synchronization module is operable to: start a blanking period after receiving the synchronization signal;control the medical treatment device to apply a first treatment signal; andonce the first treatment signal has been applied, prevent the medical treatment device from applying a subsequent treatment signal within the blanking period. 4. The system of claim 1, wherein the synchronization module is operable to start a new blanking period when a subsequent synchronization signal is received by the memory after the blanking period. 5. The system of claim 1, wherein the synchronization module is operable to: start a new blanking period when a subsequent synchronization signal is received after the blanking period;control the medical treatment device to apply a first treatment signal; andprevent the medical treatment device from applying a subsequent treatment signal within the new blanking period. 6. The system of claim 1, wherein the synchronization module, after receiving the synchronization signal, controls the medical treatment device to apply a first treatment signal after waiting a predetermined time period from the receipt of the synchronization signal. 7. The system of claim 6, wherein the synchronization module controls the medical treatment device to apply at least one IRE pulse as the first treatment signal. 8. The system of claim 1, wherein the synchronization module is operable to: determine a synchronization error based on the received synchronization signal; andupon re-establishing the synchronization, waits for a subsequent synchronization signal to start a new blanking period. 9. The system of claim 1, wherein the memory receives the synchronization signal for each occurrence of an R-wave from a cardiac device. 10. The system of claim 1, wherein the synchronization module is operable to: start a blanking period after receiving the synchronization signal;control the medical treatment device to apply a first treatment signal;once the first treatment signal has been applied, prevent the medical treatment device from applying a subsequent treatment signal within the blanking period;dynamically extend the blanking period when a new synchronization signal is received within the blanking period. 11. A system for synchronizing application of treatment signals with ... | AngioDynamics, Inc. | A61B005/00 A61N001/37 A61N001/32
US09387096 | B2 | 2016-07-12 | Feedback control systems and methods for prosthetic or orthotic devices | 2010-06-16 | Methods and systems are used for monitoring a global position or location of a prosthetic or orthotic device and to provide feedback control of the device. Certain methods may employ remote transmitting devices and receivers to recognize when a prosthetic or orthotic device user is in a moving vehicle and, therefore, initiate automatic shut-off, driving mode, or relaxed mode. Other methods may employ remote transmitting devices and receivers to identify the global position of the prosthetic or orthotic device, compare the global position to a stored terrain mapped database and output feedback control instructions and/or alerts to the prosthetic or orthotic device based at least in part on the stored terrain mapping information. | 1. A method for automatically adjusting a prosthetic or orthotic device, the method comprising: detecting a stationary position of a user of a prosthetic or orthotic device attached to the user;identifying with circuitry operatively coupled to the prosthetic or orthotic device a first geographic location identifying at least one of a longitude or latitude of the prosthetic or orthotic device based on information received at a first time from a transmitting source located remotely to the prosthetic or orthotic device;identifying with the circuitry a second geographic location identifying at least one of a longitude or latitude of the prosthetic or orthotic device based on information received at a second time from the transmitting source;calculating with circuitry a velocity of the prosthetic or orthotic device based on a distance between the first and second geographic locations over a duration of time between the first time and the second time while the user is in the stationary position;comparing with circuitry the calculated velocity to a threshold velocity; andwhen the calculated velocity of the prosthetic or orthotic device exceeds the threshold velocity while the user is in the stationary position, automatically adjusting with circuitry a state of operation of the prosthetic or orthotic device. 2. The method of claim 1, wherein the stationary position is a seated position. 3. The method of claim 2, wherein the threshold velocity is greater than 5 miles per hour. 4. The method of claim 3, wherein the threshold velocity is greater than 20 miles per hour. 5. The method of claim 1, wherein automatically adjusting a state of operation of the prosthetic or orthotic device comprises initiating at least one of an automatic shut-off, a driving mode, or a relaxed mode of the prosthetic or orthotic device. 6. The method of claim 5, wherein initiating the driving mode comprises automatically restricting movement by a knee joint of the prosthetic or orthotic device. 7. The method of claim 5, wherein initiating the relaxed mode comprises automatically restricting movement of both a knee joint and an ankle joint of the prosthetic or orthotic device. 8. The method of claim 1, wherein said identifying the first and second geographic locations with the circuitry comprises detecting the first and second geographic locations with at least one global positioning system (GPS) sensor. 9. The method of claim 1, wherein identifying the first geographic location comprises identifying an elevation of the prosthetic or orthotic device. 10. A method for actively adjusting a lower limb prosthetic or orthotic device, the method comprising: determining a global location of the prosthetic or orthotic device user based on data wirelessly received from a transmitter remotely located to the prosthetic or orthotic device; comparing the determined global location to environmental data related to the determined global location that was previously stored in a memory prio... | Ossur HF | A61F002/68 A61F002/58 A61B005/103 A61B005/00 A61F002/60 A61F002/64 A61F002/70 A61F002/76
US10832813 | B2 | 2020-11-10 | Monitoring device and method | 2010-05-25 | The present invention relates to a method and an apparatus for providing information relating to disturbances and/or disease pattern based on a patient's breathing parameters. Preferably, the apparatus has multiple functionalities and provides information relating to different disturbances and/or disease patterns. In particular, the present invention relates to an apparatus for identifying different indications of a patient, the apparatus comprising a sensor, for sensing signals indicative of a patient's breathing, a processing unit, for processing and/or analyzing the signals, and display means for displaying the outcome of the processing and/or analysis. The corresponding method for identifying different indications of a patient comprises the steps of sensing signals indicative of a patient's breathing, processing and/or analyzing the signals, and displaying the outcome of the processing and/or analysis. | 1. An apparatus for identifying different indications of a patient, the apparatus comprising: a sensor configured to sense a parameter of the patient's breathing and generate a signal based on the sensed parameter;a first A/D converter in communication with the sensor via a first channel and the first A/D converter configured to generate a flow channel signal based on the signal from the sensor;a second A/D converter in communication with the sensor via a second channel and the second A/D converter configured to generate a snore channel signal based on the signal from the sensor;a processor in communication with the first A/D converter and the second A/D converter and said processor configured to process the flow channel signal and/or the snore channel signal to calculate at least two from the group consisting of:a snoring index based on the flow channel signal from the first A/D converter and the snore channel signal from the second A/D converter, the snoring index being calculated by comparing a number of snoring events to a preset internal time base or threshold with a timer or a counter,an apnea/hypopnea index (AHI) based on a start and/or a stop of an apnea, hypopnea, and/or hyperventilation that is based on an envelope of the flow channel signal from the first A/D converter, anda periodic breathing index comprising: a periodicity index based on an average hypopnea length, an average hypopnea pause, and/or a hyperventilation length, and/ora Cheyne Stokes (CS) index based on the envelope of the flow channel signal from the first A/D converter; anda display in communication with the processor, said display configured to display results of calculating at least two of the group consisting of: the snoring index, the AHI, and the periodic breathing index, wherein the apparatus is a handheld and/or a pocket size device configured for at home use without patient performed calibration of the apparatus. 2. The apparatus according to claim 1, wherein the display is further configured to display results of calculating at least two of the group consisting of: the snoring index, the AHI, and the periodic breathing index as logical YES or NO indicators, the logical YES or NO indicators corresponding to a presence or an absence of patient snore, OS A, periodic breathing, and/or CS, respectively. 3. The apparatus according to claim 2, wherein the indicators comprise LEDs (light emitting diode). 4. The apparatus according to claim 1, wherein the apparatus is a closed unit. 5. The apparatus according to claim 1, wherein the apparatus comprises an energy source. 6. The apparatus according to claim 5, wherein the energy source is an accumulator or a rechargeable battery. 7. The apparatus according to claim 1, further comprising a filter and an envelope generator associated with the second channel. 8. The apparatus according to claim 7, wherein the filter has a lower frequency cutoff of about 10 Hz. 9. The apparatus according to claim 8, wherein the filte... | RESMED HUMIDIFICATION TECHNOLOGIES GmbH | A61B005/08 G16H040/63 A61B005/00 A61B005/087
US20100331734 | A1 | 2010-12-30 | SYSTEM AND METHOD FOR ORTHOPEDIC LOAD SENSING INSERT DEVICE | 2010-06-29 | A sensing insert device (100) is disclosed for measuring a parameter of the muscular-skeletal system. The sensing insert device (100) can be temporary or permanent. Used intra-operatively, the sensing insert device (100) comprises an insert dock (202) and a sensing module (200). The sensing module (200) is a self-contained encapsulated measurement device having at least one contacting surface that couples to the muscular-skeletal system. The sensing module (200) is included in a final insert for taking parameter measurements post-operatively. The sensing module (200) comprises one or more sensing assemblages (1802), electronic circuitry (307), an antenna (2302), and communication circuitry (320). The sensing insert device (100) is placed between a final femoral prosthetic component and a final tibial prosthetic component. Measurements are taken at more than one position of the range of motion. The load, balance, and applied location of the parameter can be determined by measurements from sensing module (200). | 1. A measurement system for measuring a parameter of the muscular-skeletal system comprising: a sensing platform comprising: a contacting surface; andat least three sensing assemblages coupled to the contacting surface at predetermined locations and where a location where the parameter is applied to the contacting surface is determined from the parameter measurements of the at least three sensing assemblages. 2. The measurement system of claim 1 where the sensing platform is in a sensing module and where the sensing module comprises: at least one printed circuit board with electronic circuitry mounted thereto;a flexible interconnect coupling the electronic circuitry and the at least three sensing assemblages; andan encapsulating enclosure enclosing the sensing platform, printed circuit board, and flexible interconnect. 3. The measurement system of claim 2 where the printed circuit board underlies the sensing platform in the encapsulating enclosure. 4. The measurement system of claim 2 where the encapsulating enclosure comprises: a housing; anda cap overlying an opening of the housing where the cap is sealed to the housing, where the cap couples to the contacting surface of the sensing platform, and where the cap moves when a force, pressure, or load is applied thereto. 5. The measurement system of claim 4 where the sensing platform is supported by at least one support structure of the housing. 6. The measurement system of claim 1 further including: electronic circuitry comprising: operational circuitry operatively coupled to the at least three sensing assemblages;power management circuitry operatively coupled to the operational circuitry;communication circuitry coupled to the operational circuitry to wirelessly transmit measured parameter data; anda power source operatively coupled to the electronic circuitry. 7. The measurement system of claim 6 where the electronic circuitry is an application specific integrated circuit. 8. The measurement system of claim 1 where a sensing assemblage comprises: a waveguide; andat least one transducer to emit energy waves in the wave guide and to detect propagated energy waves where the measurement system measures a frequency, transit time, or phase of the propagated energy waves. 9. The measurement system of claim 1 further including an insert that includes the sensing platform where the insert is used intra-operatively and where the insert is substantially equal in dimensions to a final insert. 10. The measurement system of claim 9 where the insert comprises a dock and the sensing platform. 11. The measurement system of claim 1 further including a final insert that includes the sensing platform. 12. A measurement system for measuring a force, pressure, or load applied by the muscular-skeletal system where a prosthetic component includes a sensing module and where the sensing module includes an elastic member to return a sensing platform to a predetermined position under quiescent conditions. 13. Th... | OrthoSensor | A61B005/103
US20100312078 | A1 | 2010-12-09 | Analyte Monitoring Device and Methods of Use | 2010-08-18 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1-120. (canceled) 121. An analyte monitoring device comprising: a housing having a sensor port to receive an analyte sensor;a sensor circuit that receives a signal from the analyte sensor when the analyte sensor is received in the sensor port;a measurement circuit coupled to the sensor circuit, wherein the measurement circuit receives a signal from the sensor circuit and provides output data indicative of an analyte concentration of a sample on the analyte sensor; anda transmitter unit that receives the output data from the measurement circuit and wirelessly transmits the data to a remote computer system, and wherein the transmitter wirelessly receives input from the remote computer system to initiate one or more functions. 122. The analyte monitoring device of claim 121, wherein the one or more functions are selected from a group consisting of: initiating a data transfer, changing a transmission frequency or a frequency band of the transmitter unit, deactivating an alarm system, inputting calibration data, and indicating an event and activating storage of data representative of the event. 123. The analyte monitoring device of claim 121, wherein the one or more functions includes initiation of a deactivation mode. 124. The analyte monitoring device of claim 123, wherein initiation of the deactivation mode deactivates an alarm system. 125. The analyte monitoring device of claim 123, wherein a user is prompted to confirm the deactivation mode. 126. The analyte monitoring device of claim 123, wherein the analyte monitoring device continues to process data and store the data for later retrieval despite being in the deactivation mode. 127. The analyte monitoring device of claim 121, further comprising a touch screen display. 128. The analyte monitoring device of claim 121, further comprising a password protection module to prevent an unauthorized transmission of data. 129. The analyte monitoring device of claim 121, further comprising a password protection module to prevent an unauthorized changing of settings. 130. The analyte monitoring device of claim 121, wherein an audible alarm is provided if transmission of data between the transmitter and the remote computer system is interrupted. 131. The analyte monitoring device of claim 121, wherein the transmitter wirelessly transmits the data to a drug administration system. 132. The analyte monitoring device of claim 121, wherein the transmitter uses RF communication to transmit the data to the remote computer system. 133. The analyte monitoring device of claim 121, further comprising an input device. 134. The analyte monitoring device of claim 133, wherein the input device is a keypad or keyboard. 135. The analyte monitoring device of claim 133, wherein the input device includes one or more buttons. 136. An analyte monitoring device comprising: a housing having a sensor port to receive an analyte sensor;a sensor circuit that receives a signal from the analyte sensor when the analyte sen... | Abbott Diabetes Care Inc. | A61B005/00 G06F021/00
US20100292618 | A1 | 2010-11-18 | Motion analysis device, program and method thereof and motion analysis system | 2010-04-27 | The present invention relates to a motion analysis system that includes a sensor device for measuring a motion amount related to an object which repeats a certain motion and a motion analysis device for analyzing the certain motion of the object based on the measured result by the sensor device. The motion analysis device includes means for storing time series data of the motion amount measured by the sensor device; means for analyzing a spectral intensity in each frequency in time series of waveforms shown by the time series data by using a certain function and analyzing by applying different parameters for each frequency to the function; and means for comparing spectral intensity analysis result in each frequency with different frequencies and analyzing a certain physical amount related to the certain motion of the object based on the comparison result. | 1. A motion analysis device constituting a motion analysis system that includes a sensor device for measuring a motion amount related to motion of an object which repeats a certain motion and the motion analysis device for analyzing the certain motion of the object based on the measured result of the sensor device, the motion analysis device comprising: time series data storing section configured to store time series data of the motion amount measured by the sensor device;waveform analyzing section configured to analyze a spectral intensity in each frequency by applying a certain function to a waveform shown by the time series data and analyze by applying different parameters for each frequency to the function; andmotion analysis section configured to compare, for different frequencies, the spectral intensity analysis result in each frequency analyzed by the waveform analyzing section and analyze a certain physical amount of the certain motion of the object by comparing based on the comparison result. 2. The motion analysis device according to claim 1, further comprising motion frequency region storing section configured to store a range of motion frequency which can be obtained regarding the certain motion of the object,wherein the waveform analyzing section analyzes the certain physical amount by using the region information of motion frequency stored in the motion frequency region storing section. 3. The motion analysis device according to claim 2, wherein the motion analysis section uses the spectral intensity analysis result in a range of the motion frequency among the spectral intensity analysis result in each frequency analyzed by the waveform analyzing section as a subject to be compared and analyzes the motion pace of the certain motion of the object based on the comparison result. 4. The motion analysis device according to claim 3, wherein the motion analysis section analyzes a physical amount related to a load of the certain motion of the object by using the spectral intensity analysis result of the frequency corresponding to the motion pace. 5. The motion analysis device according to claim 3, wherein the motion analysis section determines whether an abnormal action faster than a certain speed in the certain motion of the object occurs by using at least the spectral intensity analysis result in each frequency analyzed by the waveform analyzing section. 6. The motion analysis device according to claim 1, wherein the certain function used by the waveform analyzing section is a function of a mother wavelet in wavelet transform and a parameter related to a scale corresponding to the frequency is applied to the function of the mother wavelet. 7. The motion analysis device according to claim 2, wherein the certain function used by the waveform analyzing section is a function of a mother wavelet in wavelet transform and a parameter related to a scale corresponding to the frequency is applied to the function of the mother wavelet. 8. Th... | OKI ELECTRIC INDUSTRY CO., LTD. | A61B005/11
US20100286506 | A1 | 2010-11-11 | METHOD FOR DETERMINING A POSITION OF AN INSTRUMENT WITHIN A BODY CAVITY | 2010-04-08 | A method for determining the depth, direction, and/or angle of an instrument inserted into a cavity in which a user compares the light output of a light source assembly at a distal end of the instrument with known light outputs for that light source assembly. The instrument may include a plurality of light source assemblies having light outputs of different colors and intensities. The instrument may be an introducer that enables a port, e.g., a foam port, to be accurately positioned within an incision. | 1. A method for determining a position of an instrument within a body cavity, comprising the steps of: inserting an instrument into a body cavity, the instrument including at least one light source at a distal end thereof;repositioning the instrument;observing an output of the at least one light source; anddetermining a position of the instrument by comparing the observed output of the at least one light source to a known output of the at least one light source. 2. The method of claim 1, wherein the at least one light source emits light at a constant rate of output. 3. The method of claim 1, wherein the at least one light source includes a plurality of light sources each emitting light of a different color or intensity. 4. The method of claim 1, wherein the instrument is an access device. 5. The method of claim 1, wherein a template is provided having the known light outputs at given positions within a body cavity. 6. The method of claim 1, wherein the instrument further comprises a light sensor. 7. The method of claim 1, wherein the instrument is an introducer device. 8. The method of claim 7, further comprising the step of: prior to the inserting step, mounting a port to the introducer such that the introducer and the port are inserted into the body cavity together. 9. The method of claim 8, wherein the port is a foam port. 10. The method of claim 8, further comprising the step of: after the determining step, removing the introducer from the body cavity while maintaining the port within the body cavity. | Tyco Healthcare Group LP | A61B005/05
US20100312149 | A1 | 2010-12-09 | HARNESS SYSTEM FOR KINEMATIC ANALYSIS OF THE KNEE | 2008-11-26 | A harness for attachment about a knee femur of a subject, said harness comprising two abutment members, said abutment members being oriented against a skin outer surface at predetermined medial and lateral sites relative to a femur so as not to limit motion of the knee, and a strap operatively interconnecting the two abutment members such that the harness is adapted to be used on different knee sizes with the strap surrounding the knee and with the abutment members being urged against the skin outer surface at the predetermined medial and lateral sites by the strap, at least one of the abutment members supporting at least one femoral trackable member adapted to be tracked. The harness may be used for the pivot shift test. | 1. A harness for attachment about a knee of a subject, said harness comprising two abutment members, said abutment members being oriented against a skin outer surface at predetermined medial and lateral sites relative to a femur so as not to limit motion of the knee, and a strap operatively interconnecting the two abutment members such that the harness is adapted to be used on different knee sizes with the strap surrounding the knee and with the abutment members being urged against the skin outer surface at the predetermined medial and lateral sites by the strap, at least one of the abutment members supporting at least one femoral trackable member adapted to be tracked. 2. The harness according to claim 1, wherein each abutment member has at least one plate and an abutment projecting from the plate, the abutment made of a rigid material and being adapted to contact the skin outer surface at the predetermined site, with the plate contacting the skin outer surface in the periphery of the predetermined site. 3. The harness according to claim 2, wherein each abutment member comprises two of said plate, with a first plate supporting the abutment, and a second plate connected to the first plate to define a gap therewith to accommodate the strap. 4. The harness according to claim 2, wherein the plate supporting the abutment is sized 5 cm×6 cm. 5. The harness according to claim 1, wherein the femoral trackable member is connected to a lateral one of the abutment members. 6. The harness according to claim 5, wherein the femoral trackable member is at least one of a passive detectable device, an active detectable device, an accelerometer and a gyroscope. 7. The harness according to claim 1, wherein the strap is made of an elastic material. 8. The harness according to claim 3, wherein a first end of the strap has a loop surrounding the second plate of the medial abutment member so as to be connected thereto, with a second end of the strap passing through the gap of the lateral abutment member to be overlaid on the first end of the strap to surround the knee. 9. The harness according to claim 8, wherein complementary Velcro™ strips are respectively provided on the first end and the second end of the strap to secure the ends of the strap to one another. 10. A harness system for tracking knee movements for a leg, comprising a harness for attachment about a knee of a subject, said harness comprising two abutment members, said abutment members being oriented against a skin outer surface at predetermined medial and lateral sites relative to a femur so as not to limit motion of the knee, and a strap operatively interconnecting the two abutment members such that the harness is adapted to be used on different knee sizes with the strap surrounding the knee and with the abutment members being urged against the skin outer surface at the predetermined medial and lateral sites by the strap, at least one of the abutment members supporting at least one femoral tra... | - | A61B005/103
US20100286570 | A1 | 2010-11-11 | WEIGHT DISTRIBUTION DETERMINATION APPARATUS | 2008-05-01 | An apparatus (10) for determining the weight distribution between front and rear portions of a foot (24) of an individual standing on a surface (26). The apparatus (10) comprises a compressible elongate body member (11) which has an outer resilient housing (12) and at least one inner chamber (14). The at least one inner chamber (14) includes inflation means (22) located therein and the body member (11) is configurable to be positioned between the arch of an individual's foot (24) and the surface (26). The inflation means (22) is inflatable such that the body member (11) provides a pivot point between the front and rear portions (28, 30) of the individual's foot (24). | 1. An apparatus for determining the weight distribution between front and rear portions of a foot of an individual standing on a surface, the apparatus comprising: a compressible elongate body member having an outer resilient housing and at least one inner chamber; andinflation means located within the at least one inner chamber, wherein the body member is configurable to be positioned between the arch of an individual's foot and the surface and the inflation means is inflatable such that the body member provides a pivot point between the front and rear portions of the individual's foot. 2. An apparatus for determining weight distribution as claimed in claim 1, wherein the apparatus further comprises weight distribution measuring means for measuring the weight distribution between the front and rear portions of the foot. 3. An apparatus for determining weight distribution as claimed in claim 2, wherein the weight distribution measuring means includes display means for displaying the weight distribution between the front and rear portions of the foot. 4. An apparatus for determining weight distribution as claimed in claim 2, wherein the weight distribution measuring means includes signal generation means for generating a signal when the weight distribution between the front and rear portions of the individual's foot falls outside a predetermined range. 5. An apparatus for determining weight distribution as claimed in claim 1, wherein the inflation means comprises an inflatable bladder. 6. An apparatus for determining weight distribution as claimed in claim 1, wherein the body member comprises two separate inner chambers, wherein each chamber includes inflation means located therein. 7. An apparatus for determining weight distribution as claimed in claim 6, wherein the body member is of sufficient length to allow the apparatus to be placed under both feet of the individual. 8. An apparatus for determining weight distribution as claimed in claim 6, wherein the apparatus comprises weight distribution measuring means for measuring the weight distribution between the front and rear portions of each foot. 9. An article of footwear including the apparatus as claimed in claim 1, wherein the apparatus is located between the sole of the article of footwear and the surface. 10. An article of footwear as claimed in claim 9, wherein the apparatus is located beneath the arch of the individual's foot. 11. An article of footwear as claimed in claim 9, wherein the apparatus is integrally formed with the article of footwear. | - | A61B005/103 A43B013/00
US20100324457 | A1 | 2010-12-23 | SKELETAL-MUSCULAR POSITION MONITORING DEVICE | 2009-12-10 | A system is disclosed that records position data for portions of a body as a function of time. The position data can be collected from one or more sensors secured to the body either individually or using a patch. The sensors, in some embodiments, can include stretch sensors that produce a change in electrical resistance as the stretch sensors are stretched. A data logger can be used to record the data. Various other elements such as a feedback mechanism or a manual pain indicator can also be included. | 1. A system comprising: a plurality of stretch sensors configured to be attached to a body, wherein each stretch sensor is also configured to translate a longitudinal stretch into a change in an electrical property of the stretch sensor; anda data logger coupled with the plurality of stretch sensors and configured to measure the electrical property of the stretch sensors and store electrical property data as a function of time. 2. The system according to claim 1, wherein the electrical property comprises resistance. 3. The system according to claim 1, further comprising a tilt sensor coupled with the data logger, wherein the data logger is configured to record tilt data as a function of time. 4. The system according to claim 3, wherein the tilt sensor comprises an accelerometer or an inclinometer. 5. The system according to claim 1 further comprising a user interface, wherein the data logger is configured to receive and record an indication from the user through the user interface that a painful event has occurred. 6. The system according to claim 5, wherein the user interface comprises a button. 7. The system according to claim 5, wherein the data logger is configured to determine threshold values from the electrical property data received from the stretch sensors and the indication received from the user. 8. The system according to claim 7 further comprising a feedback device coupled with the data logger, wherein the feedback device is configured to provide a user feedback when electrical property data approaches the threshold value. 9. The system according to claim 1 further comprising a patch configured to adhere to a portion of a body, wherein at least one of the plurality of stretch sensors is coupled with the patch and attached to the body using the patch. 10. A system comprising: a patch configured to couple with a body;a plurality of sensors coupled with the patch, wherein the plurality of sensors are configured to sense a position of the sensors relative to one another;memory; anda controller communicatively coupled with the plurality of sensors and the memory, wherein the controller is configured to receive sensor data and write the sensor data to the memory. 11. The system according to claim 10 further comprising a tilt sensor coupled with the patch and communicatively coupled with the controller, wherein the controller is configured to write tilt data to the memory. 12. The system according to claim 10 further comprising a feedback mechanism communicatively coupled with the controller, wherein the feedback mechanism is configured to provide feedback to a user in response to a signal from the controller, wherein the controller provides a signal to the feedback mechanism when sensor data reaches a threshold value. 13. The system according to claim 12 further comprising a tilt sensor coupled with the patch and communicatively coupled with the controller, wherein the controller provides a signal to the feedback mechanism when... | Jacob Bean | A61B005/103
US07853319 | B2 | 2010-12-14 | Method and system for detecting electrophysiological changes in pre-cancerous and cancerous tissue and epithelium | 2006-04-21 | A method and system for determining a condition of a selected region of epithelial and stromal tissue, for example in the human breast. A plurality of electrodes are used to measure surface and transepithelial electropotential of breast tissue as well as surface electropotential and impedance at one or more locations and at several defined frequencies, particularly very low frequencies. An agent may be introduced into the region of tissue to enhance electrophysiological characteristics. Measurements made at ambient and varying suction and/or positive pressure applied to the epithelial tissue and/or positive pressure conditions applied to the breast are also used as a diagnostic tool. Tissue condition is determined based on the electropotential and impedance profile at different depths of the epithelium, stroma, tissue, or organ, together with an estimate of the functional changes in the epithelium due to altered ion transport and electrophysiological properties of the tissue. Devices for practicing the disclosed methods are also provided. | 1. A method for determining a condition of a region of epithelial breast tissue comprising: (A) establishing a connection between a first electrode and the epithelial tissue of the nipple of a breast with a ductal probe, an electroconductive media or both;(B) placing a second electrode in contact with the surface of the breast;(C) establishing a signal between the first and second electrodes;(D) establishing that the nipple ducts of the breast are open by: (1) measuring the impedance between the first and second electrode at about 5 different frequencies in the range of about 200 Hz to about 60,000 Hz sufficient to establish an impedance curve;(2) treating the nipple using at least one method selected from the group consisting of (a) applying suction and release of suction to the nipple; (b) applying alcohol; and (c) applying a dekeratinizing agent;(3) again measuring the impedance between the first and second electrode at about 5 different frequencies in the range of about 200 Hz to about 60,000 Hz sufficient to establish an impedance curve and comparing the impedance curve obtained to that obtained in (D) (1) above;(4) repeating steps (2) and (3) until the impedance curve obtained in step (D) (3) is substantially unchanged in order to confirm that the ducts are open;(E) measuring between the first and second electrode: (1) a DC potential; and(2) impedance at about 5 different frequencies in the range of about 10 Hz to about 200 Hz and impedance at from about 5 to about 50 different frequencies in the range of about 0.1 Hz to about 10 Hz; and(F) determining the condition of the region of epithelial tissue based on the DC potential and impedance measurements between the first and second electrode. 2. A method for determining a condition of a region of epithelial breast tissue comprising: (A) placing over the nipple of a breast a cup having an interior, and first and second openings, and an electrode disposed within the interior, the cup having a source of suction in communication with the first opening, the second opening having been placed over the nipple;(B) establishing a connection between the electrode and the epithelial tissue of the nipple of a breast with a ductal probe, an electroconductive media or both;(C) placing a second electrode in contact with the surface of the breast;(D) establishing a signal between the first and second electrodes;(E) establishing that the nipple ducts of the breast are open by: (1) measuring the impedance between the first and second electrode at about 5 different frequencies in the range of about 200 Hz to about 60,000 Hz sufficient to establish an impedance curve;(2) treating the nipple using at least one method selected from the group consisting of (a) applying suction and release of suction to the nipple; (b) applying alcohol; and (c) applying a dekeratinizing agent;(3) again measuring the impedance between the first and second electrode at about 5 different frequencies in the range of about 200 Hz to abo... | Epi-Sci, LLC | A61B005/05
US07853457 | B2 | 2010-12-14 | Automatic identification for spot measurements | 2006-05-01 | A patient monitoring system (8) monitors physiological functions of a plurality of patients (12). Each clinician (10) and each patient (12) has an associated identification device (20, 22), which each includes an identification code corresponding to respective clinician and patient. Each identification device (20, 22) includes a respective body-coupled communication device (24, 26) for communicating the identification code to a medical device (14). The clinician (10) activates the medical device (14) which includes a body-coupled communication device (40) including an ID reader (42). The ID reader (42) scans the area to detect whether the clinician identification is present. After the clinician's identification code is read, the medical device (14) is ready to take measurements. The clinician (10) takes the medical device (14) to the patient (12). The ID reader (42) scans the area to detect the patient identification code. The result of the measurement along with the patient identification code is stored in a medical device memory (56) and later transferred to a patient record in a hospital database (70). | 1. A patient monitoring system for monitoring physiological functions of each of a plurality of patients, the system comprising: a plurality of identification devices, each associated with one of a corresponding one of a plurality of patients including: an electronic identification code memory which stores a patient identification code which identifies the corresponding to one of the patients, anda patient body-coupled communication device capacitively coupled with the corresponding patient which communicates the identification code via near field communications; anda medical device including: a sensor which spot measures a physiological function of each of a plurality of patients;a medical device body-coupled communication device which capacitively couples with each patient during spot measuring and automatically reads the respective patient identification codes via near field communications and automatically associates physiological function measurements with the patient identification code of the corresponding patient;anda memory which stores the associated physiological functional measurements and patient identification codes. 2. The system as set forth in claim 1, further including: a clinician identification device including: an electronic clinician identification code memory which stores a clinician identification code, anda clinician body-coupled communication device which capacitively couples to a clinician using the medical device to communicate the clinician identification code to the medical device using the near field communications; andwherein the medical device body coupled communication device automatically associates the clinician identification code of the clinician who uses the medical device with the corresponding generated physiological function measurement and patient identification code. 3. The system as set forth in claim 2, wherein the medical device further includes: a medical device activator which triggers the clinician body-coupled communication device to transmit the electronic clinician identification code. 4. The system as set forth in claim 2, further including: a measurement activator which activates the medical device in response to receiving the electronic clinician identification code using the near field communication and triggers the patient body-coupled communication device to transmit the electronic patient identification code using the near field communications. 5. The system as set forth in claim 4, further including: a hospital database which downloads the physiological function measurements along with the respective patient identification code and clinician identification code from the medical device memory. 6. The system as set forth in claim 1, further including: a clinician identification device including: an electronic clinician identification code memory which stores a clinician identification code, anda clinician body-coupled communication device which capacitively couples to the clinician a... | Koninklijke Philips Electronics N.V. | G06Q010/00 G06F019/00 A61B005/00
US07831290 | B2 | 2010-11-09 | MRI biopsy device localization fixture | 2005-04-12 | A localization mechanism, or fixture, is used in conjunction with a breast coil for breast compression and for guiding a core biopsy instrument during prone biopsy procedures in both open and closed Magnetic Resonance Imaging (MRI) machines. The localization fixture includes a three-dimensional Cartesian positionable guide for supporting and orienting an MRI-compatible biopsy instrument, and, in particular, a sleeve to a biopsy site of suspicious tissues or lesions. A z-stop enhances accurate insertion, prevents over-insertion or inadvertent retraction of the sleeve. The sleeve receives a probe of the MRI-compatible biopsy instrument and may contain various features to enhance its imagability, to enhance vacuum and pressure assist therethrough, etc. | 1. An apparatus for performing a minimally invasive medical procedure with reference to a diagnostic image taken of a patient's breast, the apparatus comprising: a medial compression member;a lateral compression member movable relative to the medial compression member to compress the patient's breast; anda biopsy probe support positionable relative to the lateral compression member, wherein the biopsy probe support further comprises a monocle positionable proximate to the patient's breast operatively configured to guide marking and scoring of an insertion point, wherein the monocle includes a reticule positionable in a first position along an axis of penetration, wherein the monocle is operable to rotate the reticule to a second position not along the axis of penetration after marking and scoring the insertion point. 2. The apparatus of claim 1, further comprising: a medial plate supporting the medial compression member; anda lateral plate supporting the lateral compression member, wherein the medial plate is operable to translate in opposition to the lateral plate;wherein the medial plate includes a proximally extending handle, the lateral plate including a handle extending proximate to the medial plate handle, wherein moving the handles in unison centers the medial and lateral compression members upon a patient's breast, and squeezing the two handles together effects compression of the patient's breast. 3. The apparatus of claim 1, further comprising: a medial plate supporting the medial compression member;a lateral plate supporting the lateral compression member;a base plate including a lateral guide surface, wherein the lateral plate is operable to translate in the lateral guide surface of the base plate;a medial guide surface defined in a selected one of the base plate and the lateral plate;a lateral plate locking member coupled between the base and lateral plates; anda medial plate locking member coupled between the medial plate and the medial guide surface;wherein the medial guide surface is defined in the base plate distal to the lateral plate guide surface, the lateral plate locking member comprising a cam lock locking the lateral plate to the base plate and the medial plate locking member comprising a sliding attachment between the medial and lateral plates movable to effect locking binding therebetween. 4. The apparatus of claim 1, wherein the biopsy probe support further comprises a cradle positioned to support a proximal lower portion of a biopsy device. 5. The apparatus of claim 1, further comprising a proximally biased guard to mitigate inadvertent contact with the biopsy probe support. 6. The apparatus of claim 5, wherein the proximally biased guard comprises a gas charged cylinder extending a guard rail proximally. 7. The apparatus of claim 6, wherein the proximally biased guard comprises a pivotally attached arm and a spring positioned to rotate the arm from a lateral to a proximal direction. 8. The apparatus of claim 1, ... | Devicor Medical Products, Inc. | A61B005/05 A61B017/32
US07853317 | B2 | 2010-12-14 | Method and system for cardiac signal decomposition | 2005-10-13 | A method and system decomposes a cardiac signal, such as an electrocardiogram (ECG) signal, into components. The components are then usable to assist in the detection of an abnormal heart condition. More particularly, a single lead sensor is used to generate a single lead cardiac signal. The cardiac signal is segmented into a set of cycle segments according to detected heart waveforms. The cycle segments are aligned and used to generate a set of cross-sectional signals. The cross-sectional signals are aligned and presented as inputs to a signal separation process, which separates the cardiac signal into a set of components. The components may be grouped according to predefined criteria. The components or groups may be analyzed or displayed to assist in the detection of an abnormal cardiac signal, which may be indicative of an abnormal heart condition. In one example, the signal separation process is a non-orthogonal transformation method such as independent component analysis (ICA). | 1. A method of analyzing an electrocardiogram (ECG) signal, the method comprising: receiving the ECG signal;segmenting the received ECG signal into a set of cycle segments;aligning the set of cycle segments along a relative time axis to obtain aligned cycle segments;generating a set of cross-sectional signals from the aligned cycle segments, wherein each cross-sectional signal is a time-slice signal representing a cross-section of the aligned cycle segments along a relative time of the aligned cycle segments, said each cross-sectional signal corresponding to a different relative time;separating, using a signal separation process, the set of cross-sectional signals into a set of statistically substantially independent components; anddetecting, using the set of components, an abnormal cardiac signal;wherein the steps of receiving, segmenting, aligning, generating, separating, and detecting are performed by a processor-based cardiac signal detection system. 2. The method of claim 1, further comprising the steps of selecting a plurality of the independent components that match a predefined criteria; andarranging the selected independent components into a component group. 3. The method of claim 2, wherein the detecting step comprises detecting the abnormal cardiac signal in the component group. 4. The method of claim 2, wherein the selecting step comprises choosing in accordance with one or more predefined criteria, the one or more predefined criteria comprising one or more frequency criteria, and one of the component groups is arranged to have independent components matching the one or more frequency criteria. 5. The method of claim 2, further comprising displaying at least one of the arranged component groups and a name identifying the predefined criteria. 6. The method of claim 2, further comprising displaying the component group. 7. The method of claim 1, further comprising the steps of: displaying at least some of the independent components to an operator;receiving from the operator instructions to select a plurality of the independent components; andarranging the selected independent components into a component group. 8. The method of claim 1, wherein segmenting the received ECG signal comprises isolating the received ECG signal into a set of QRS complex wave cycle segments. 9. The method of claim 1, wherein aligning the cycle segments comprises detecting an R-wave peak point of each of the cycle segments. 10. The method of claim 1, wherein aligning the cycle segments comprises detecting a P-wave peak point of each of the cycle segments. 11. The method of claim 1, wherein aligning the cycle segments comprises detecting a Q-wave peak point of each of the cycle segments. 12. The method of claim 1, wherein aligning the cycle segments comprises detecting a R-wave peak point of each of the cycle segments and calculating time intervals between adjacent R-wave peak points. 13. The method of claim 1, wherein segmenting the received ECG signa... | Sigmed, Inc. | A61B005/04
US07831300 | B2 | 2010-11-09 | Extrapolating ICA knowledge from one epoch to another for improved fetal ECG separation | 2008-02-20 | A method of utilizing a maternal-fetal monitoring system to monitor the physiological properties of both a maternal patient and a fetus. A series of ECG electrodes are placed across the maternal patient's abdomen and receives ECG input waveforms across sixteen separate channels. The sixteen channels of information are processed using an ICA algorithm to generate a series of ICA output waveforms and a transfer matrix. Following the current epoch, the transfer matrix is applied to the input waveforms on a continuous basis. The conditioned input waveforms are displayed immediately following the first epoch and prior to the expiration of a subsequent epoch. The transfer matrix for the second epoch is combined with the transfer matrix for the first epoch to generate an updated transfer matrix. Various filtering operations on the transfer matrix coefficients may be used to find the updated transfer matrix before the end of the current epoch. | 1. A method of identifying a fetal ECG signal from a plurality of input waveforms from a maternal patient, comprising the steps of: positioning a plurality of ECG electrodes on an abdomen of the maternal patient;obtaining a set of input waveforms from the maternal patient through the ECG electrodes for a plurality of epochs;storing the plurality of input waveforms for a first epoch;applying an independent component analysis (ICA) algorithm to the stored input waveforms for the first epoch to generate a plurality of ICA output waveforms, wherein the ICA algorithm generates a first transfer matrix for the first epoch;continuously applying the first transfer matrix to the plurality of input waveforms as the plurality of input waveforms are obtained following the first epoch to generate a plurality of conditioned waveforms; andcontinuously displaying at least one of the conditioned waveforms following the first epoch in real time. 2. The method of claim 1 wherein the conditioned waveforms are continuously displayed. 3. The method of claim 2 wherein only selected waveforms of the conditioned waveforms are displayed. 4. The method of claim 1 wherein the first epoch has a determined time length. 5. The method of claim 1 wherein the first epoch has a variable time length. 6. The method of claim 1 further comprising the steps of: storing the plurality of input waveforms for a second epoch following the first epoch;applying the ICA algorithm to the stored input waveforms for the second epoch to generate a plurality of ICA output waveforms for the second epoch and a second transfer matrix for the second epoch;combining the first transfer matrix and the second transfer matrix;generating an updated transfer matrix based upon the combination of the first transfer matrix and the second transfer matrix; andcontinuously applying the updated transfer matrix to the plurality of input waveforms following the second epoch. 7. The method of claim 6 wherein the updated transfer matrix is generated based upon an average of corresponding coefficients within the first transfer matrix and the second transfer matrix. 8. The method of claim 6 wherein the updated transfer matrix is generated based upon the most current epoch and the previous epoch. 9. The method of claim 1 further comprising the steps of: determining the conditioned waveform that contains a fetal ECG source; anddisplaying only the conditioned waveform as the fetal ECG signal. 10. A method of separating a fetal ECG signal from a plurality of input waveforms received from a plurality of ECG electrodes attached to an abdomen of a maternal patient, the method comprising the steps of: receiving the plurality of input waveforms from the electrodes, each input waveform being received as one of a plurality of channels each associated with a set of the plurality of electrodes;storing the plurality of input waveforms for a first epoch;applying an independent component analysis (ICA) algorithm to the plurality ... | The General Electric Company | A61B005/0245
US07885699 | B2 | 2011-02-08 | Analyte monitoring device and methods of use | 2010-08-06 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. An in vivo glucose monitoring system comprising: two in vivo working electrodes; anda receiver unit external to a user in signal communication with the two in vivo working electrodes, the receiver unit comprising: one or more processors; anda computer-readable storage medium including instructions which, when executed by the one or more processors, causes the one or more processors to: receive signals from both of the working electrodes indicative of a glucose level of the user;process the signals received from both of the working electrodes; anddetermine the acceleration of the rate of change of the glucose level. 2. The system of claim 1, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to determine whether the acceleration of the rate of change of the glucose level exceeds or meets one or more threshold values. 3. The system of claim 2, wherein the receiver unit further comprises an alarm, and wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the one or more processors to activate the alarm when the acceleration of the rate of change of the glucose level exceeds or meets the one or more threshold values. 4. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 5. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a single datapoint indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 6. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when a predetermined number of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 7. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automatically deactivate when an average of datapoints obtained over a period of time indicates that the acceleration of the rate of change of the glucose level no longer exceeds or meets the one or more threshold values. 8. The system of claim 3, wherein the computer-readable storage medium further includes instructions which, when executed by the one or more processors, causes the alarm to automaticall... | Abbott Diabetes Care Inc. | A61B005/00
US08414579 | B2 | 2013-04-09 | Map and ablate open irrigated hybrid catheter | 2010-06-23 | An embodiment of an open-irrigated catheter system comprises a tip section, a distal insert, and mapping electrodes. The tip section has an exterior wall that defines an open interior region within the tip section. The exterior wall includes mapping electrode openings and irrigation ports. The exterior wall is conductive for delivering radio frequency (RF) energy for an RF ablation procedure. The irrigation ports are in fluid communication with the open interior region to allow fluid to flow from the open interior region through the irrigation ports. The distal insert is positioned within the tip section to separate the open region into a distal fluid reservoir and a proximal fluid reservoir. The mapping electrodes are positioned in the mapping electrode openings in the tip section. | 1. An open-irrigated catheter system for performing mapping and ablation functions, comprising: a tip section having an exterior wall that defines an open interior region within the tip section, wherein the exterior wall includes mapping electrode openings and irrigation ports, and wherein the exterior wall is conductive for delivering radio frequency (RF) energy for an RF ablation procedure, wherein the irrigation ports are in fluid communication with the open interior region to allow fluid to flow from the open interior region through the irrigation ports;a distal insert positioned within the tip section to separate the open region into a distal fluid reservoir and a proximal fluid reservoir;mapping electrodes positioned in the mapping electrode openings in the tip section; andnoise artifact isolators fabricated from an electrical insulator, wherein the mapping electrodes are electrically insulated from the exterior wall of the tip section by the noise artifact isolators, wherein the noise artifact isolators have a lip extending out from the exterior wall of the tip section past the outside surface of the mapping electrodes. 2. The system of claim 1, wherein the distal insert has a peripheral surface with openings therein sized to receive the mapping electrodes. 3. The system of claim 1, further comprising a thermocouple, wherein the distal insert defines a thermocouple opening extending from the distal fluid reservoir to the proximal fluid reservoir, the thermocouple opening configured to receive the thermocouple. 4. The system of claim 1, wherein the distal insert defines a fluid conduit extending from the distal fluid reservoir to the proximal fluid reservoir. 5. The system of claim 1, wherein the noise artifact isolators have a generally annular shape. 6. The system of claim 1, wherein the mapping electrodes include three mapping electrodes approximately equally spaced from each other about a circumference of the tip section. 7. The system of claim 1, wherein the mapping electrodes include four mapping electrodes approximately equally spaced from each other spaced about a circumference of the tip section. 8. The system of claim 1, wherein the irrigation ports in fluid communication with the open region include distal irrigation ports in fluid communication with the distal fluid reservoir. 9. The system of claim 1, wherein the irrigation ports in fluid communication with the open region include proximal irrigation ports in fluid communication with the proximal fluid reservoir. 10. The system of claim 1, wherein the irrigation ports in fluid communication with the open region include both distal irrigation ports in fluid communication with the distal fluid reservoir and proximal irrigation ports in fluid communication with the proximal fluid reservoir. 11. The system of claim 1, further comprising a fluid cooling lumen extending through the distal insert to deliver fluid to the distal fluid reservoir, the fluid cooling lumen having ... | Boston Scientific Scimed, Inc. | A61B018/18 A61B018/04 A61N001/00
US08634891 | B2 | 2014-01-21 | Method and system for self regulation of sensor component contact pressure | 2009-05-20 | A method and system for regulating contact pressure of a sensor is provided. In accordance with an embodiment, a sensor includes a collapsible material capable of regulating the localized component contact pressure exerted on a tissue between a threshold which ensures proper contact between the sensing component and monitoring site but which minimizes incidents of excessive localized pressures which could result in tissue damage over prolonged use. This localized contact pressure range is maintained over a wide range of sensor application pressures. In one embodiment, a window is provided in the sensor and a colorimetric indication of the compression or expansion of the collapsible material may be determined through the window. | 1. A sensor, comprising: a light emitting component comprising a first patient-contacting surface and a first opposing surface;a light detecting component comprising a second patient-contacting surface and a second opposing surface and configured to detect light energy emitted by the light emitting component incident upon the second patient-contacting surface;a backing layer; anda collapsible material positioned between the backing layer and at least one of the first opposing surface or the second opposing surface, wherein the collapsible material, when applied to a tissue, is configured to maintain an expanded configuration under a first pressure below a first threshold, is configured to maintain a partially collapsed configuration under a second pressure above the first threshold and below a second threshold, and is configured to maintain a collapsed configuration under a third pressure above the first threshold and the second threshold, wherein the first threshold is in a range of 10 mmHg to 35 mmHg, and wherein the second threshold is at least 80 mmHg. 2. The sensor of claim 1, wherein the collapsible material comprises a woven material. 3. The sensor of claim 1, wherein the collapsible material comprises a foam material comprising open cells, closed cells, and/or a combination thereof. 4. The sensor of claim 1, wherein the collapsible material comprises a gel. 5. The sensor of claim 1, wherein the collapsible material is a different color based on the degree of expansion or compression of the collapsible material. 6. The sensor of claim 1, comprising a window situated above the collapsible material such that a color of the collapsible material may be viewed. 7. The sensor of claim 1, wherein the backing layer comprises a window situated above the collapsible material through which the thickness of the collapsible material may be visually determined. 8. The sensor of claim 1, wherein the first threshold comprises a pressure which maintains intimate contact between the sensing component and the monitoring site and the second threshold comprises a pressure level which minimizes excessive localized pressure which could result in tissue damage over prolonged periods of use. 9. The sensor of claim 1, wherein the collapsible material comprises a mesh material. 10. The sensor of claim 1, wherein the collapsible material comprises a material having honeycombed cells. 11. The sensor of claim 1, wherein a tissue pressure exerted on a patient tissue by one or more of the first patient-contacting surface or the second patient-contacting surface is 10-35 mmHg in the expanded configuration. 12. A system, comprising: a monitor; anda sensor in electrical communication with the monitor, the sensor comprising: an emitter;a detector configured to generate an electrical signal in response to light emitted by the emitter;a backing layer; anda collapsible material situated between the backing layer and at least one of the emitter or the detector, wherei... | Covidien LP | A61B005/00
US08768432 | B2 | 2014-07-01 | Method for absorption correction of PET data and MR-PET system | 2010-06-24 | An apparatus is disclosed for combined magnetic resonance tomography and positron emission tomography imaging which is designed for recording PET image data of a person under examination from an examination area. In at least one embodiment, the apparatus includes a scanning unit, embodied to scan a prespecified area of the person under examination and based on the scanning, to determine a contour of the person under examination for the prespecified area; and a processing unit, embodied, based on the contour determined, to carry out an absorption correction of PET data which has been recorded from the prespecified area of the person under examination. A method is also disclosed. | 1. An apparatus for combined magnetic resonance tomography and positron emission tomography imaging, comprising: a structurally integrated imaging unit including, a single patient table,a magnetic resonance imaging scanner configured to record, from an examination region, MRT image data of a person under examination,a positron emission tomography scanner configured to record, from the examination region, PET image data of the person under examination,an additional scanning unit, configured to scan an area of the person under examination and based on the scanning, to determine a contour of the person under examination for the area, anda processor configured to, determine a location of structures that attenuate a PET signal based on the determined contour and the MRT image data, andcarry out an absorption correction of the recorded PET image data from the examination region of the person under examination based on the determined location of the structures. 2. The apparatus as claimed in claim 1, wherein the additional scanning unit is arranged before the examination region in a direction in which the person under examination is moved into in the examination region, the additional scanning unit being configured to scan the area, the area differing from the examination region of the apparatus. 3. The apparatus as claimed in claim 1, wherein the additional scanning unit is configured to scan a surface of a slice of the person under examination. 4. The apparatus as claimed in claim 1, wherein the patient table is configured to move the person under examination through the examination region, the additional scanning unit being configured to scan the area of the person under examination slice-by-slice by moving the patient table. 5. The apparatus as claimed in claim 1, wherein the apparatus is configured to simultaneously carry out a PET measurement and a scan of the person under examination by way of the additional scanning unit. 6. The apparatus as claimed in claim 1, wherein the processor is configured to create an attenuation correction map of the scanned area based on the determined contour, the absorption correction of the PET data being undertaken based upon the attenuation correction map. 7. The apparatus as claimed in claim 6, wherein the processor is configured to allocate a constant attenuation coefficient to areas lying within the determined contour. 8. The apparatus as claimed in claim 7, wherein the processor is configured to create the attenuation correction map taking into account information from MRT image data recorded by way of the apparatus in an MRT imaging of the person under examination. 9. The apparatus as claimed in claim 6, wherein the processor is configured to create the attenuation correction map taking into account information from MRT image data recorded by way of the apparatus in an MRT imaging of the person under examination. 10. The apparatus as claimed in claim 1, wherein the additional scanning unit comprises ... | Siemens Aktiengesellschaft | A61B005/055 A61B006/03 A61B006/00 A61B005/107 G01R033/48 G06F019/00
US09592010 | B2 | 2017-03-14 | Dual mode closed-loop system and method for measuring a parameter of the muscular-skeletal system | 2010-06-29 | A dual-mode closed-loop measurement system (100) for capturing a transit time, phase, or frequency of energy waves propagating through a medium (122) is disclosed. A first device comprises an inductor drive circuit (102), an inductor (104), a transducer (106), and a filter (110). A second circuit comprises an inductor (114) and a transducer (116). A parameter to be measured is applied to the medium (122). The medium (122) is coupled between the first device and the second device. The first device initiates the transmit inductor (104) to query via inductive coupling to a receive inductor (114) on the second device via a first path. The inductor (114) triggers a transducer (116) on the second device to emit an energy wave that is propagated in the medium (122) and detected by the first device. The transit time of energy waves is affected by the parameter by known relationship. | 1. A medical measurement system comprising: a first device comprising: a transmit inductor;a first transducer wherein the first transducer is configured to couple to a first region of a muscular-skeletal system;an inductor driver circuit coupled to the transmit inductor;a buffer circuit coupled between the first transducer and the inductor driver circuit; anda second device comprising: a receive inductor configured to inductively couple to the transmit inductor;a second transducer coupled to the receive inductor wherein the second transducer is configured to couple to a second region of the muscular-skeletal system. 2. The system of claim 1 wherein the first and second transducers are configured to be acoustically coupled. 3. The system of claim 1 further including a filter coupled between the buffer circuit and the inductor driver circuit. 4. The system of claim 3 wherein the buffer circuit is configured to output a digital signal. 5. The measurement system of claim 3 wherein the filter is a low pass filter. 6. The measurement system of claim 1 wherein the second device is housed in a screw configured to couple to the muscular-skeletal system. 7. The measurement system of claim 6 wherein the second transducer and the receive inductor are coupled to a flexible interconnect that is placed in a cavity of the screw. 8. The measurement system of claim 7 wherein the flexible interconnect is rolled and placed in a cavity of the screw. 9. The measurement system of claim 8 wherein the second transducer is coupled to the screw. 10. The measurement system of claim 9 wherein the second transducer is configured to transmit an energy wave through the screw. 11. The measurement system of claim 7 wherein the receive inductor is formed in the flexible interconnect. 12. The measurement system of claim 1 wherein the first transducer is configured to propogate at least one energy wave through the muscular-skeletal system, the measurement system further including a receiver station configured to receive measurement data wherein the receiver station is configured to process and display characteristics of the at least one energy wave propagation through the muscular-skeletal system including transit time, frequency, amplitude, or phase. 13. The measurement system of claim 12 wherein the frequency of the energy wave corresponds to bone density. 14. The measurement system of claim 12 wherein the first device is housed in a prosthetic component. 15. The measurement system of claim 14 wherein the frequency of the energy wave corresponds to glue joint integrity. 16. The measurement system of claim 1 where the system propagates an energy wave from the first device to the second device and where the energy wave can be a continuous wave energy wave or a pulsed energy wave. 17. The measurement system of claim 1 wherein a first electrical signal is used to generate an electromagnetic signal which is propagated from the first device to the second device. 18. The... | ORTHOSENSOR INC. | A61B005/05 A61B005/00 A61B008/15
US20100331737 | A1 | 2010-12-30 | ENCAPSULATED FORCE SENSOR FOR MEASURING A PARAMETER OF THE MUSCULAR-SKELETAL SYSTEM | 2010-06-29 | A sensing insert device (100) is disclosed for measuring a parameter of the muscular-skeletal system. The sensing insert device (100) can be temporary or permanent. Used intra-operatively, the sensing insert device (100) comprises an insert dock (202) and a sensing module (200). The sensing module (200) is a self-contained encapsulated measurement device having at least one contacting surface that couples to the muscular-skeletal system. The insert dock (202) is a passive component made for different prosthetic component manufacturers as well as for different size prosthetic components. The sensing module (200) fits in an opening or cavity of the insert dock (202). The intra-operative insert device is substantially equal in dimension to an implanted final insert. The sensing insert device (100) is also a permanent prosthetic component. The sensing module (200) residing within the sensing insert device and coupling to a bearing surface of the insert. | 1. A sensing system for measurement of a parameter of the muscular-skeletal system comprising: a trial insert having dimensions substantially equal to the dimensions of a final insert;a sensing module in the trial insert to measure the parameter. 2. The sensing system of claim 1 where the sensing module is placed in a cavity of the trial insert. 3. The sensing system of claim 1 where the sensing module is encapsulated to be hermetically sealed. 4. The sensing system of claim 2 where the trial insert comprises a dock having a feature that operatively couples to a prosthetic component where the dock includes an opening to receive the sensing module. 5. The sensing system of claim 1 where the trial insert interfaces with permanent prosthetic components. 6. The sensing system of claim 1 where the trial insert includes a cover overlying a contacting surface of the sensing module. 7. The sensing system of claim 1 where the trial insert measures a force, pressure, or load that is substantially equal to the force, pressure, or load applied by the muscular-skeletal system to the final insert when installed. 8. The sensing system of claim 1 where the sensing platform comprises one or more sensing assemblages;electronic circuitry operatively coupled to the one or more sensing assemblages to measure the parameter;a power source coupled to the electronic circuitry;a transmitter coupled to the electronic circuitry;an antenna coupled to the transmitter; andan enclosure that encapsulates the electronic circuitry, the power source, antenna, transmitter, and the one or more sensing assemblageswhere a transit time, frequency, or phase is measured by the sensing platform. 8. (canceled) 9. The sensing system of claim 1 where the sensing module is disposed of after surgery. 10. A prosthetic component with sensing capability for in-situ measurement of the muscular-skeletal system comprising a final insert having a bearing surface where a sensing module resides within the final insert and where the sensing module includes a contacting surface that couples to the bearing surface to measure loading thereon. 11. The prosthetic component 10 where a capacitor in the sensing module is inductively charged and where the capacitor powers the sensing module to measure a parameter of the muscular-skeletal system. 12. The prosthetic component of claim 11 where a transit time, frequency, or phase is measured by the sensing module corresponding to the parameter applied to the bearing surface of the final insert. 13. The sensing system of claim 12 where the sensing platform comprises one or more sensing assemblages;electronic circuitry operatively coupled to the one or more sensing assemblages to measure the parameter;a power source coupled to the electronic circuitry;a transmitter coupled to the electronic circuitry;an antenna coupled to the transmitter; andan enclosure that encapsulates the electronic circuitry, the power source, and the one or more sensing assemblages.... | OrthoSensor | A61B005/103
US20100286534 | A1 | 2010-11-11 | METHODS AND SYSTEMS FOR ASSESSING METABOLIC TRANSITION POINTS | 2007-06-29 | The invention relates to methods and systems for determining one or more metabolic transition points in a subject during a work task. The methods and systems comprise determining the respiration rate (RR) and heart rate (HR) in the subject that is undergoing a work task and calculating the ratio of RR/HR at more than one time point during the work task. The metabolic transition points are identifiable points in time of the RR/HR ratio in the subject that is undergoing a work task. | 1. A method for determining one or more metabolic transition points in a subject during a work task, said method comprising a) determining the respiration rate (RR) or any other related parameter in the subject that is undergoing the work task;b) determining the heart rate (HR) of the subject undergoing the work task;c) calculating the ratio of RR/HR at more than one time point during the work task; andd) identifying at least one distinct point in time of the RR/HR ratio,wherein the identified distinct point(s) correlates with the metabolic transition point of the subject undergoing a work task. 2. The method of claim 1, wherein at least two distinct points in time are identified. 3. The method of claim 2, wherein the metabolic transition points are selected from the group consisting of aerobic metabolism (AM), anaerobic threshold (AT) and respiratory compensation point (RCP). 4. The method of claim 3, wherein AT or RCP are determined. 5. The method of claim 4, wherein the ratio of RR/HR is calculated at least once per minute during the work task. 6. The method of claim 5, wherein the ratio is calculated at least 5 times per minute during the work task. 7. The method of claim 6, wherein identifying at least one distinct point comprises plotting a curve of the RR/HR ratio versus time and identifying specific points along the plotted curve. 8. The method of claim 7, wherein the curve plotting is performed by a computer. 9. The method of claim 7, wherein the curve plotting is performed by hand. 10. The method of claim 9, wherein the AT is determined. 11. The method of claim 10, wherein the AT is the lowest point on the plotted curve. 12. The method of claim 6, wherein identifying at least one distinct point comprises recording each value of the RR/HR ratio onto recordable medium and identifying the specific values of the ratio in relation to all recorded ratio values. 13. The method of claim 12, wherein the AT is determined. 14. The method of claim 13, wherein the AT is the lowest ratio value in relation to all recorded ratio values. 15. A system for determining one or more metabolic transition points in a subject during a work task, said system comprising: a) means for determining the respiration rate (RR) or any other related parameter in the subject that is undergoing the work task;b) means for determining the heart rate (HR) of the subject undergoing the work task;c) means for calculating the ratio of RR/HR at more than one time point during the work task; andd) means for identifying at least one distinct point in time of the RR/HR ratio,wherein the identified distinct points correlate with the metabolic transition point of the subject undergoing a work task. 16. The system of claim 15, wherein the system is capable of identifying at least two distinct points in time. 17. The system of claim 16, wherein the metabolic transition points are selected from the group consisting of aerobic metabolism (AM), anaerobic threshold (AT) and ... | ALTERNA, LIMITED | A61B005/0205
US07826892 | B2 | 2010-11-02 | Method for quantitative assessment of the autonomic nervous system based on heart rate variability analysis | 2007-06-29 | A method is provided for characterizing autonomic nervous system activity of a patient based on heart rate variability analysis using electrocardiographic data from the patient. In addition, a computer-readable medium is provided tangibly embodying a program of instructions executable by a computer to perform a method for characterizing autonomic nervous system activity of a patient based on heart rate variability analysis using electrocardiographic data from the patient. Also, a system is provided for characterizing autonomic nervous system activity of a patient comprising a processor and a machine-readable medium tangibly embodying a program of instructions executable by the processor. | 1. A method for characterizing autonomic nervous system activity of a patient comprising: (a) acquiring electrocardiographic data from the patient with an electrocardiographic device;(b) determining from the electrocardiographic data obtained in step (a) a plurality of time intervals between successive R-waves;(c) analyzing the plurality of time intervals by frequency domain analysis so as to quantitate the following parameters:i) power of high frequency (PHF);ii) power of low frequency 1 (PLF1) (0.07 to 0.15 Hz);iii) power of low frequency 2 (PLF2) (0.033 to 0.07 Hz);iv) difference between power of low frequency 1 and high frequency (PLF1 - PHF);v) maximum amplitude of spectral function in a low frequency band (SLFmax);vi) frequency at a point of maximum amplitude of spectral function in the low frequency band (FLFmax);vii) maximum amplitude of spectral function in a high frequency band (SHFmax);viii) frequency at a point of maximum amplitude of spectral function in the high frequency band (FHFmax);ix) ratio of power of high frequency to power of low frequency 1, (PHF/PLF1) ;x) ratio of power of high frequency to power of low frequency (sum of powers of low frequency 1 and low frequency 2), (PHF/(PLF1 +PLF2));xi) heart rate (HR);and, if a spectral peak falls in the range of a middle frequency band,xii) (SMF) maximum amplitude of spectral function in a middle frequency band; and(d) applying a clusterization algorithm to the parameter values quantitated in step (c) so as to determine two further parameters, one of which correlates with sympathetic nervous system activity and the other of which correlates with parasympathetic nervous system activity, wherein parasympathetic nervous system activity is determined by applying the clusterization algorithm to parameters (c)(i) to (c)(iv) and wherein sympathetic nervous system activity is determined by applying the clusterization algorithm to parameters (c)(ii), (c)(iii), and (c)(v) to (c)(xii), so as to thereby characterize the autonomic nervous system activity of the patient. 2. The method of claim 1, wherein the high frequency is 0.15 to 0.5 Hz. 3. The method of claim 1, wherein the value of PHF, PLF1 or PLF2 is one of 30 values. 4. The method of claim 3, wherein each of the 30values is an integer representing the sub-range of the power spectrum falling between each of the following numbers, in msec2: 99999, 12000, 11000, 9000, 7800, 6400, 5200, 4200, 3200, 2500, 2100, 1700, 1300, 1000, 900, 820, 730, 640, 570, 490, 420, 360, 310, 255, 200, 145, 95, 55, 25, 15 and 0 for PHF;99999, 9000, 7800, 6400, 5200, 4200, 3400, 3000, 2600, 2200, 1850, 1550, 1300, 1100, 1000, 900, 810, 730, 640, 570, 490, 420, 350, 300, 250, 200, 150, 120, 70, 25 and 0for PLF1; and99999, 9000, 7800, 6400, 5200, 4200, 3400, 3000, 2600, 2200, 1850, 1550, 1300, 1100, 1000, 900, 810, 730, 640, 570, 490, 420, 350, 300, 250, 200, 150, 120, 70, 25 and 0for PLF2. 5. The method of claim 1, further comprising replacing an artifactual pa... | Intelwave, LLC | A61B005/04
US20100292588 | A1 | 2010-11-18 | LIVING BODY INFORMATION DETECTION APPARATUS AND BLOOD-PRESSURE METER | 2010-07-30 | A blood-pressure meter includes a frame part including a first arm and a second arm that are opposed to each other. The blood-pressure meter further including a pressure applying part provided on at least one of a side of the first arm opposed to the second arm and a side of the second arm opposed to the first arm. The blood-pressure meter also including a detection part for detecting a pulse wave. | 1. A blood-pressure meter comprising: a frame part including a first arm and a second arm that are opposed to each other;a pressure applying part provided on at least one of a side of the first arm opposed to the second arm and a side of the second arm opposed to the first arm; anda detection part for detecting a pulse wave. 2. The blood-pressure meter as claimed in claim 1, wherein the pressure applying part includes a mechanism using an air pressure or an actuator. 3. The blood-pressure meter as claimed in claim 1, wherein the detection part includes at least one of a photoelectric sensor including a light-emitting element and a light-receiving element, a pressure detection mechanism, a vibration sensor, and a Korotkov sound detection mechanism. 4. The blood-pressure meter as claimed in claim 1, wherein the detection part includes at least a pair of a light-emitting element and a light-receiving element for measuring transmittance of light transmitted between the first arm and the second arm. 5. The blood-pressure meter as claimed in claim 1, wherein the detection part includes at least a pair of a light-emitting element and a light-receiving element for measuring reflectance of light, wherein the pair of the light-emitting element and the light-receiving element is provided on one of the side of the first arm opposed to the second arm or the side of the second arm opposed to the first arm. 6. A blood-pressure meter comprising: a frame part including a first arm and a second arm for pinching a part of an ear part;a pressure applying part for applying a pressure on the part of the ear part, the pressure applying part being provided on a side of the first arm opposed to the second arm; anda detection part for detecting a pulse wave at the part of the ear part. 7. The blood-pressure meter as claimed in claim 6, wherein the pressure applying part includes a mechanism using an air pressure or an actuator. 8. The blood-pressure meter as claimed in claim 6, wherein the detection part includes at least one of a photoelectric sensor including a light-emitting element and a light-receiving element, a pressure detection mechanism, a vibration sensor, and a Korotkov sound detection mechanism. 9. The blood-pressure meter as claimed in claim 6, wherein the detection part includes at least a pair of a light-emitting element and a light-receiving element for measuring transmittance of light transmitted between the pressure applying part and the second arm. 10. The blood-pressure meter as claimed in claim 6, further comprising: a fixing part to be fixed to the part of the ear part, the fixing part being provided on the inside of the second arm; anda fixing adjustment part that is provided with the fixing part at its top and that pushes the fixing part to the part of the ear part. 11. A blood-pressure meter comprising: a frame part including a first arm and a second arm for pinching a part of an ear part;a first pressure applying part for applying a pr... | NIPPON TELEGRAPH AND TELEPHONE CORP. | A61B005/022
US20100312130 | A1 | 2010-12-09 | GRADED RESPONSE TO MYOCARDIAL ISCHEMIA | 2010-08-20 | Severity and confidence level of a myocardial ischemia event can be determined, such as including using an ambulatory medical device, and such information can be used to determine a graded response to the myocardial ischemia event. | 1. An apparatus comprising: an ambulatory medical device, including a ischemia detector circuit configured to detect an indication of a myocardial ischemia event; anda processor circuit, configured to be communicatively coupled to the ischemia detector and to: receive the indication of the myocardial ischemia event;determine a confidence level of the myocardial ischemia event having occurred; andrespond using the confidence level, the responding including at least one of initiating, selecting, or adjusting a response. 2. The apparatus of claim 1, wherein the processor circuit is configured to determine a severity indicator value of the indication of the myocardial ischemia event. 3. The apparatus of claim 2, wherein the processor circuit is configured such that the responding comprises responding using both the severity indicator value and the confidence level. 4. The apparatus of claim 3, wherein the processor circuit is configured such that the severity indicator is multi-valued and the confidence level is multi-valued. 5. The apparatus of claim 1, wherein the processor circuit is configured such that the determining the confidence level comprises using a regression model. 6. The apparatus of claim 5, wherein the processor circuit is configured such that: the detecting the indication of the myocardial ischemia event comprises using one or more sensor measurements from the ischemia detector circuit to detect the indication of the myocardial ischemia event; andthe determining the confidence level comprises computing a probability according to:P=11+-z,wherein z=b0+b1X1+b2X2+ . . . +bmXm, P is a probability of a myocardial ischemia event occurring, z is a measure of a total contribution of all of the one or more sensor measurements used, b0 is a logistic regression intercept, and b1, b2, . . . bm are the logistic regression coefficients of the one or more sensor measurements X1, X2, . . . Xm respectively. 7. The apparatus of claim 1, wherein the ambulatory medical device comprises an implantable medical device; and wherein the implantable medical device includes the processor circuit. 8. The apparatus of claim 1, wherein the processor circuit is configured such that the confidence level is determined using a time-wise sequence of multiple indications of the myocardial ischemia event. 9. A device-readable medium including instructions that, when performed by the device, comprise: detecting an indication of a myocardial ischemia event;determining a confidence level of the myocardial ischemia event having occurred; andresponding using the confidence level, the responding including at least one of initiating, selecting, or adjusting a response. 10. The device-readable medium of claim 9, wherein the instructions that, when performed by the device, comprise determining a severity indicator value of the indication of the myocardial ischemia event. 11. The device-readable medium of claim 10, wherein the responding comprises responding using both... | - | A61B005/02
US20100312090 | A1 | 2010-12-09 | ATHEROSCLEROSIS RISK ASSESSMENT BY PROJECTED VOLUMES AND AREAS OF PLAQUE COMPONENTS | 2010-05-19 | A novel technique directed toward risk assessment of a patient's plaque vulnerability, wherein clinical events may be caused by internal plaque components affecting a lumen within an artery. A surface area projection or shadow of one or more plaque components onto a lumen can be measured and assessed. Optionally, a total volume projection onto the lumen can also be measured and assessed to refine the determination of risk to a patient and to monitor the progression of atherosclerosis over time. | 1. A method for assessing a risk of atherosclerosis in an artery, based on a plurality of images of the artery, comprising the steps of: (a) selecting a plaque component in the plurality of images;(b) projecting a dimension of the plaque component onto a circumference of a lumen within the artery;(c) projecting a surface area of the plaque component along a longitudinal axis of the lumen; and(d) assessing the risk of atherosclerosis in the artery as a function of the dimension of the plaque component projected onto the circumference of the lumen, and the surface area of the plaque component projected along the longitudinal axis of the lumen. 2. The method of claim 1, further comprising the steps of: (a) determining a region of the plaque component;(b) projecting the region of the plaque component onto a surface of the lumen to define a corresponding volume; and(c) carrying out the step of assessing the risk of atherosclerosis also based on the corresponding volume of the region of the plaque component projected onto the surface of the lumen. 3. The method of claim 1, wherein the step of projecting the dimension of the plaque component onto the circumference of the lumen comprises the steps of: (a) identifying boundaries of the artery;(b) identifying lines of thickness between an internal boundary of a wall of the artery and an outer boundary of the wall of the artery;(c) determining specific lines of thickness that intersect the plaque component;(d) marking points where each specific line of thickness intersects the circumference of the lumen; and(e) determining a length along the circumference resulting from the step of projecting the dimension of the plaque component to encompass the points marked on the circumference of the lumen. 4. The method of claim 3, wherein the step of projecting the surface area of the plaque component along a longitudinal axis of the lumen comprises the steps of: (a) determining a thickness of each cross-sectional image in which the lumen is intersected by the lines of thickness;(b) determining a projected area of the plaque component for each cross-sectional image intersected by the lines of thickness;(c) for each cross-sectional image intersected by the lines of thickness, multiplying a projected length of the plaque component by the thickness of the cross-sectional image to determine a projected area for each such cross-sectional image; and(d) determining a sum of all projected areas for the cross-sectional images. 5. The method of claim 4, wherein the step of projecting the region of the plaque component onto the surface of the lumen to define the corresponding volume comprises the steps of: (a) determining a maximum line of thickness for each cross-sectional image in which the plaque component is disposed;(b) determining an average maximum line of thickness based on the maximum line of thickness determined and set equal to a maximum thickness; and(c) determining a projected volume of the region of the plaque ... | University of Washington Center for Commercialization | A61B005/055 A61B008/00 A61B005/00 G06K009/00
US20100286567 | A1 | 2010-11-11 | ELDERLY FALL DETECTION | 2009-05-06 | Health-sensing and health-action devices and systems are generally described. The health-sensing device may include one or more of a sensor, a filter, and a transmitter. The sensor may be configured to sense when a user is likely to have fallen down and may include a sensor in the form of one or more microphone devices, accelerometers, and/or MEMs devices. The filter may be configured to evaluate a signal from the sensor and determine when a fall has occurred and analyze the magnitude of the fall. The transmitter may be arranged for initiating a transmission based on a signal from the filter. The health-action device may be configured for responding to a fall and may include one or more of a receiver, a processor, and a responder. A notifying element may also be included for interrupting the transmission when assistance is not needed. | 1. A health-sensing device for placement on a user to monitor a likelihood of a fall by the user, the device comprising: a sensor configured to: sense a motion associated with the health-sensing device, andgenerate a sensor signal indicative of the sensed motion;a filter configured to: receive the sensor signal,determine when the received sensor signal indicates the likelihood of the fall by the user, andgenerate a filter signal when the likelihood of the fall by the user is determined; anda transmitter configured to: receive the filter signal, andgenerate a transmission responsive to the filter signal. 2. The device of claim 1, wherein the sensor comprises an accelerometer. 3. The device of claim 1, wherein to determine when the received sensor signal indicates the likelihood of the fall by the user, the filter is configured to compare the sensed motion from the received sensor signal to a range of threshold accelerations. 4. The device of claim 3, wherein the range of threshold accelerations comprises a preset response strategy relating to the threshold accelerations within the range. 5. The device of claim 1, further comprising a notifying element positioned on the health-sensing device, wherein the filter is configured to trigger the notifying element. 6. The device of claim 5, further comprising an actuatable element positioned on the health-sensing device and in communication with the filter, the actuatable element and filter being configured to interrupt the transmission. 7. The device of claim 5, further comprising an actuatable element positioned on the health-sensing device and in communication with the filter, the actuatable element and filter being configured to generate an overriding filter signal for transmission by the transmitter. 8. A method for a health-sensing device to monitor a likelihood of a fall by a user, the method comprising: sensing a motion associated with the user to generate a sensor signal indicative of the sensed motion;analyzing the sensor signal to determine when the sensed motion associated with the user identifies the likelihood of the fall by the user; andinitiating a transmission responsive to the analysis of the sensor signal. 9. The method of claim 8, wherein analyzing the sensor signal further comprises comparing the sensed motion from the sensor signal to a threshold for acceleration, and identifying the likelihood of the fall by the user when the sensed motion from the sensor signal exceeds the threshold for acceleration. 10. The method of claim 9, further comprising encoding the transmission with information indicating the sensed motion from the sensor signal exceeded the threshold for acceleration. 11. The method of claim 10, wherein the threshold acceleration comprises a high threshold, an intermediate threshold, and a low threshold. 12. The method of claim 11, further comprising driving a notifying element when the low threshold has been exceeded. 13. The method of claim 12, further comp... | - | A61B005/103
US20100298720 | A1 | 2010-11-25 | In Situ Energy Harvesting Systems for Implanted Medical Devices | 2010-04-16 | This invention concerns miniature implantable power sources that harvest or scavenge energy from the expansion and contraction of biological tissues, for example, an artery or a bundle of muscle fiber. Such power sources employ an energy harvesting element that converts mechanical or thermal energy existing or generated in or from a pulsatile tissue into a form of electrical energy that can be used or stored by an implanted medical device, such as a blood pressure sensor, a flow meter, or the like. Preferred energy harvesting element embodiments utilize a piezoelectric thin film embedded within a flexible, self-curling medical-grade polymer or coating. Such power sources can be used to produce self-powered implanted microsystems with continuous or near-continuous operation, increased lifetimes, reduced need for surgical replacement, and minimized or eliminated external interface requirements. | 1. An in situ biological energy harvesting device comprising an energy harvesting element disposed in a resilient biocompatible insulator and configured as a cuff adapted for energy-transferring association with a pulsatile tissue. 2. An in situ biological energy harvesting device according to claim 1 wherein the pulsatile tissue is a blood vessel, optionally an artery or arterial graft. 3. An in situ biological energy harvesting device according to claim 1 wherein the energy harvesting element comprises a piezoelectric thin film. 4. An in situ biological energy harvesting device according to claim 3 wherein the piezoelectric thin film is embedded in a resilient cuff, optionally a self-curling medical-grade silicone cuff. 5. An in situ biological energy harvesting device according to claim 1 configured as a power source for an implantable medical device, optionally a cardiac stimulation device, a neurostimulator, an implantable drug pump, and device for monitoring or sensing a physiological parameter. 6. An in situ biological energy harvesting device according to claim 1 wherein the implantable medical device is a cardiac stimulation device selected from the group consisting of a pacemaker, a defibrillator, a cardioverter, and a device that includes two or more thereof. 7. An in situ biological energy harvesting device according to claim 1 wherein the implantable medical device that monitors a physiological parameter selected from the group consisting of blood pressure, fluid flow rate, temperature, and electrical activity is a tissue or organ. 8. An implantable medical device that comprises an in situ biological energy harvesting device according to claim 1. 9. A method for monitoring a physiological parameter in vivo, comprising implanting in a patient an implantable medical device according to claim 8 configured to monitor a physiological parameter. 10. A method of treatment, comprising a implanting in a patient an implantable medical device according to claim 8 configured to effect a desired treatment, wherein the implantable medical device optionally is selected from the group consisting of a cardiac stimulation device, a neurostimulator, and a drug pump. | - | A61N001/378 A61B005/021 A61B005/026 A61B005/01
US20100331631 | A1 | 2010-12-30 | OXYGEN SATURATION EAR SENSOR DESIGN THAT OPTIMIZES BOTH ATTACHMENT METHOD AND SIGNAL QUALITY | 2009-06-30 | A system is provided that includes an ear sensor and an external device. The ear sensor includes a sensing component with sensors for sensing various physiological parameters. The ear sensor also includes a retaining component configured to retain the ear sensor to the ear of a wearer. As the retaining component retains the ear sensor to the ear, the sensing component may be configured to have an optimal surface contact between the sensors and the ear tissue, such that an improved physiological signal may be obtained. In some embodiments, the improved physiological signal may result in physiological data, which may be displayed and organized in the external device. | 1. An ear sensor comprising: a retaining component configured to retain the ear sensor to an ear; anda sensing component coupled to the retaining component, wherein the sensing component is configured to retain a sensor to the ear and receive a physiological signal. 2. The ear sensor, as set forth in claim 1, wherein the sensing component comprises: an emitter adapted to transmit light into tissue; anda light detector configured to receive the transmitted light from the tissue, wherein the physiological signal comprises the transmitted light received by the light detector. 3. The ear sensor, as set forth in claim 2, wherein the sensing component comprises opposing members configured to apply force against opposing sides of a portion of the ear, wherein one member of the opposing members comprises the emitter, and the other member of the opposing members comprises the light detector. 4. The ear sensor, as set forth in claim 1, wherein the sensing component is configured to retain the sensor to the ear with a force that optimizes the physiological signal. 5. The ear sensor, as set forth in claim 4, wherein the force is based substantially on optimizing comfort for a user of the ear sensor. 6. The ear sensor, as set forth in claim 1, comprising a processing component configured to process the physiological signal to produce digitized physiological data. 7. The ear sensor, as set forth in claim 6, comprising an output configured to transmit the digitized physiological data to an external device. 8. The ear sensor, as set forth in claim 1, comprising an output configured to transmit the physiological signal to an external device. 9. The ear sensor, as set forth in claim 8, wherein the external device comprises a processor configured to process the physiological signal to produce digitized physiological data. 10. The ear sensor, as set forth in claim 1, wherein the physiological signal is processed to produce data including one or more of oxygen saturation level, pulse rate, temperature, and blood pressure. 11. The ear sensor, as set forth in claim 10, wherein the data is displayed by an external device configured to enable a user to organize and analyze the displayed data. 12. The ear sensor, as set forth in claim 11, wherein the external device is a cell phone, a personal digital assistant, or any other portable device. 13. A physiological monitoring system comprising: an ear sensor comprising: an emitter configured to transmit light into tissue;a detector configured to receive the light from the tissue; anda retaining component configured to retain the ear sensor to a wearer's ear; andan external device coupled to the ear sensor, wherein the external device is configured to display physiological data based on the light received at the detector of the ear sensor. 14. The physiological monitoring system of claim 13, wherein the ear sensor comprises a processor configured to produce physiological data from the light received at the detecto... | Nellcor Puritan Bennett LLC | A61B005/1455 A61B005/00
US20100331628 | A1 | 2010-12-30 | Wireless Monitoring and Control of Medical Devices | 2009-06-29 | In described embodiments, a data collection device receives signals from one or more patient monitoring devices, the signals representing patient data and other vital signs measured at a patient. The data collection device employs statistical quality algorithms to track irregular behavior and out-of-bound events, the behavior and events either being pre-set, adaptively set, or otherwise defined within pre-determined limits. The data collection device communicates alerting signals to a caregiver's handheld device when the irregular behavior and out-of-bound events occur. The alerting signals contain information related to the patient data, irregular behavior and out-of-bound events, thereby allowing a caregiver to take appropriate action. | 1. A patient monitoring system comprising: a data collection device (DCD) coupled to one or more monitoring devices through corresponding links, the DCD configured to i) collect signals from one or more monitoring devices, wherein each signal represents patient data collected by one or more sensors coupled to a patient, and ii) process each signal in accordance with one or more algorithms, wherein said DCD generates at least one of a vital sign (VS) signal and an alerting signal for each signal based on the one or more algorithms for irregular behavior or an out-of-bound event with the corresponding signal; anda portable communication device (PCD), in wireless communication with the DCD, configured to i) receive the VS signal and the alerting signal and ii) provide information associated with the VS signal and the alerting signal to a caregiver. 2. The invention of claim 1, further comprising: a medical database, coupled to the DCD, having medical data relating to at least one of vital sign, patient care, and medical treatment statistical information and adapted to receive patient data from the DCD;a rule and procedure algorithm processor, coupled to at least one of the DCD and the medical database, configured to i) receive a portion of the medical data from the medical database associated with a patient, ii) receive a portion of the patient data, and iii) process the portion of medical data and the portion of patient data in accordance with a statistical monitoring algorithm having a control methodology. 3. The invention of claim 2, wherein the rule and procedure algorithm processor implements a control methodology as a control chart based on a set of rules. 4. The invention of claim 3, wherein the set of rules is a Western Electric Rules of Statistical Quality Control (WE-SQC). 5. The invention of claim 1, wherein the PCD comprises: a processor;a display;a speaker module, anda wireless communication module,wherein the processor is: i) coupled to the display, the speaker module, and the wireless communication module,ii) configured to receive at least one of a vital sign (VS) signal and an alerting signal from the wireless communication module, andiii) configured to provide information associated with the VS signal and the alerting signal to a caregiver through at least one of the display and the speaker module. 6. The invention of claim 5, further comprising an entry device, wherein: the entry device is at least one of a keypad and touch display; andthe processor receives caregiver queries from the entry device. 7. The invention of claim 6, wherein: the processor is configured to i) communicate the caregiver queries from the entry device to the DCD, receive answers to the caregiver queries from the DCD, and communicate the answers to the caregiver through at least one of the speaker module and the display. 8. The invention of claim 5, wherein the wireless communication module implements a wireless communication link operating in accordanc... | LSI Corporation | A61B005/00 G08B023/00
US20100292603 | A1 | 2010-11-18 | Electrical Impedance Myography | 2006-09-21 | Electrical impedance myography (EIM) can be used for the assessment and diagnosis of muscular disorders. EIM includes applying an electrical signal to a region of tissue and measuring a resulting signal. A characteristic of the region of tissue is determined based on the measurement. Performing EIM at different frequencies and/or different angular orientations with respect to a muscle can aid in the assessment and diagnosis. Devices are described that facilitate assessment and diagnosis using EIM. | 1. A method of determining at least one characteristic of a region of tissue, the method comprising acts of: applying a plurality of first electrical signals to the region of tissue, each of the plurality of first electrical signals being applied at a respective one of a plurality of orientations so as to propagate through at least a portion of the region of tissue in a respectively different direction with respect to a first axis of the region of tissue;obtaining a plurality of measurements from the region of tissue, each of the plurality of measurements indicative of a respective one of a plurality of second electrical signals, each of the plurality of second electrical signals resulting from applying a respective one of the plurality of first electrical signals; anddetermining the at least one characteristic based, at least in part, on the plurality of measurements. 2. The method of claim 1, wherein the act of applying the plurality of first electrical signals to the region of tissue includes an act of applying at least one first electrical signal in a direction substantially parallel to the first axis and applying at least one first electrical signal in a direction substantially perpendicular to the first axis. 3. The method of claim 2, wherein the act of applying the plurality of electrical signals to the region of tissue includes an act of applying at least one first electrical signal in a direction substantially parallel to the first axis and applying at least one electrical signal at least every approximately 30° rotation from the first axis over a range of approximately 180°. 4. The method of claim 2, wherein the first axis is substantially aligned in a direction of muscle fibers of the region of tissue. 5. The method of claim 1, wherein the plurality of first electrical signals include a plurality of current signals, and wherein the plurality of second electrical signals include a plurality of voltage signals, each resulting from a respective one of the plurality of current signals. 6. The method of claim 1, wherein the act of determining the at least one characteristic includes an act of determining at least one property of the region of tissue at each of the plurality of orientations, based on the plurality of measurements. 7. The method of claim 6, further comprising an act of generating a first signature of the region of tissue, the first signature comprising the at least one property as a function of orientation. 8. The method of claim 7, wherein the at least one property includes a resistance of the region of tissue at the respective orientation. 9. The method of claim 7, wherein the at least one property includes a reactance of the region of tissue at the respective orientation. 10. The method of claim 7, wherein the at least one property includes a phase derived from a resistance and a reactance of the region of tissue at the respective orientation. 11. The method of claim 7, wherein the act of determining the at least... | Beth Israel Deaconess Medical Center, Inc. Northeastern University | A61B005/053
US20100292543 | A1 | 2010-11-18 | Analysis of Endogenous Fluorescence Images to Extract Morphological/Organization Information About Living Samples | 2008-10-30 | Methods and computer program products for analyzing tissue are provided. The tissue is exposed to light at the appropriate wavelengths for inducing fluorescence. Images of the fluorescing tissue are taken at two or more depths within the tissue. The PSD function is determined for each image at a different depth within the tissue. A characteristic of each PSD function determined is compared, and it is determined whether or not the tissue exhibits a pre-cancerous characteristic. | 1. A method for analyzing tissue having at least two layers of cells comprising: (a) determining a power spectral density (PSD) function at a designated range of spatial frequencies for each of two or more images of the tissue, wherein each image is obtained at a different depth within the tissue, the PSD function being determined using high resolution fluorescence images of the tissue, and(b) comparing the PSD functions obtained. 2. The method of claim 1, wherein the two or more depths correspond to different cellular layers within the tissue. 3. The method of claim 1, wherein the high resolution fluorescence is multi-photon excited fluorescence. 4. The method of claim 1, wherein the high resolution fluorescence is two-photon excited fluorescence. 5. The method of claim 1, wherein the high resolution of fluorescence is 1 micron or less. 6. The method of claim 1, wherein the fluorescence is autofluorescence. 7. The method of claim 1, wherein comparing the PSD functions comprises comparing a characteristic of each PSD function. 8. The method of claim 7, wherein the characteristic of the PSD function is a power exponent α. 9. The method of claim 7, wherein the characteristic of the PSD function is variance across depth. 10. The method of claim 1, wherein the spatial frequency range corresponds to tissue matter. 11. The method of claim 1, wherein the spatial frequency corresponds to nuclear or intercellular matter. 12. The method of claim 1, wherein the spatial frequency range corresponds to subcellular matter. 13. The method of claim 12, wherein the subcellular matter is mitochondria. 14. The method of claim 1, wherein the images are obtained by exposing the tissue to light having a wavelength between 700 and 900 nm. 15. The method of claim 12, wherein the images are obtained by exposing the tissue to light having a wavelength of about 740 nm. 16. The method of claim 12, wherein the images are obtained by exposing the tissue to light having a wavelength of about 800 nm. 17. A method for determining whether tissue exhibits a pre-cancerous characteristic comprising: (a) determining a power spectral density (PSD) function at a designated range of spatial frequencies for each of two or more images of the tissue, wherein each image is obtained at a different depth within the tissue, the PSD function being determined using high resolution fluorescence images of the tissue, and(b) comparing the PSD functions obtained. 18. The method of claim 17, wherein the two or more depths correspond to different cellular layers within the tissue. 19. The method of claim 17, wherein the high resolution fluorescence is multi-photon excited fluorescence. 20. The method of claim 17, wherein the high resolution fluorescence is two-photon excited fluorescence. 21. The method of claim 17, wherein the high resolution of fluorescence is 1 micron or less. 22. The method of claim 17, wherein the fluorescence is autofluorescence. 23. The method of claim 17, ... | Tufts University | A61B005/00 G06K009/00
US20100332250 | A1 | 2010-12-30 | PRE-EXAMINATION MEDICAL DATA ACQUISITION SYSTEM | 2008-12-10 | A pre-examination patient information gathering system comprises an electronic user interface (30, 130) including a display (32) and at least one user input device (34, 36), and an electronic processor (50) configured to present an initial set of questions (54) to a patient via the electronic user interface, receive responses to the initial set of questions from the patient via the electronic user interface, construct or select follow-up questions (68) based on the received responses, present the constructed or selected follow up questions to the patient via the electronic user interface, and receive responses to the constructed or selected follow up questions from the patient via the electronic user interface. A physiological sensor (70, 72, 74, 76, 78, 80) may be configured to autonomously sense a patient physiological parameter as the patient interacts with the electronic user interface. | 1. A pre-examination patient information gathering system comprising: an electronic user interface (30, 130) including a display (32) and at least one user input device (34, 36); andan electronic processor (50) configured to present an initial set of questions (54) to a patient via the electronic user interface, receive responses to the initial set of questions from the patient via the electronic user interface, construct or select follow-up questions (68) based on the received responses, present the constructed or selected follow-up questions to the patient via the electronic user interface, and receive responses to the constructed or selected follow-up questions from the patient via the electronic user interface. 2. The pre-examination patient information gathering system as set forth in claim 1, further comprising: a link (94) with an incentives provider (96) configured to cause the incentives provider to provide an incentive responsive to a selected patient interaction with the pre-examination patient information gathering system. 3. The pre-examination patient information gathering system as set forth in claim 2, wherein the link (94) with the incentives provider (96) includes at least one of: a link with a company configured to cause the company to provide a pecuniary discount responsive to the patient completing an information gathering session using the pre-examination patient information gathering system;a link with a marking engine configured to cause the marking engine to generate a printed coupon redeemable for a pecuniary benefit responsive to the patient completing an information gathering session using the pre-examination patient information gathering system; andexpedite scheduling to visit a medical provider. 4. The pre-examination patient information gathering system as set forth in claim 1, further comprising: a kiosk (22) including the electronic user interface (30) disposed in a medical office (10). 5. The pre-examination patient information gathering system as set forth in claim 4, wherein the kiosk (22) includes a privacy enclosure or curtain (40). 6. The pre-examination patient information gathering system as set forth in claim 4, further comprising: an alarm (28) configured to alert the medical office (10) conditional upon a received response or a combination of received responses satisfying an alarm condition. 7. The pre-examination patient information gathering system as set forth in claim 4, wherein the electronic user interface (30) further comprises: a physiological sensor (70, 72, 74, 76, 78, 80) configured to autonomously sense a patient physiological parameter, the physiological sensor being arranged in or with the kiosk (22) to sense the patient physiological parameter as the patient interacts with the electronic user interface. 8. The pre-examination patient information gathering system as set forth in claim 1, further comprising: an Internet server (136) configured to operatively connect the electronic us... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | G06Q050/00 G06Q030/00 G06Q090/00 G06F017/00 G06F015/16 G06F017/30 A61B005/00
US20100324376 | A1 | 2010-12-23 | Diagnostics Data Collection and Analysis Method and Apparatus | 2010-07-06 | A medical diagnostic data collector/analyzer compiles historical vehicle diagnostic data, including measured vital signs from a number of different people from a variety of population types, and performs statistical analyses on various vital sign combinations to establish ranges corresponding to healthy individuals and various disease conditions. | 1. A computer-implemented method of analyzing medical data, comprising: compiling, via a processor of a diagnostic device, a collection of historical test data points from a plurality of population types;correlating, via the processor, each historical test data point with a population type to produce entries of a diagnostic case history;grouping, via the processor, the entries of the diagnostic case history by population type;defining, via the processor, a range corresponding to the population type of a patient based on the collection of test data points of the entries of the diagnostic case history grouped by population type; anddiagnosing, via the processor, a disease condition of the patient based on the range corresponding to the disease condition for the population type. 2. The computer-implemented method of claim 1, wherein the test data points correspond to a plurality of patients having a plurality of discrete population types. 3. The computer-implemented method of claim 1, wherein the population type corresponds to a state of wellness of one or more patients of the population type. 4. The computer-implemented method of claim 1, wherein the population type corresponds to a particular disease existing in one or more patients of the population type. 5. The computer-implemented method of claim 1, wherein each of the test data points comprises a single set of vital signs recorded at a moment in time. 6. The computer-implemented method of claim 1, wherein each of the test data points comprises a sequence of vital signs recorded over a period of time. 7. The computer-implemented method of claim 1, wherein the step of defining comprises a method of automated reasoning. 8. The computer-implemented method of claim 1, further comprising: representing each of the test data points as a point in a multidimensional vector space; andstatistically analyzing a set of the test data points corresponding to the population type to define a parameter space corresponding to the population type in the multidimensional vector space, wherein the parameter space comprises the range. 9. The computer-implemented method of claim 8, wherein the step of statistically analyzing further comprises: associating with the set a multidimensional probability distribution having a mean value and a multidimensional variable variance vector; andoptimizing the parameter space by identifying an optimal variance vector based on the set. 10. The computer-implemented method of claim 8, wherein the step of statistically analyzing further comprises mapping the disease condition to the parameter space. 11. The computer-implemented method of claim 8, further comprising: performing a dimensionality reduction procedure on the set of test data points and a correlated set of diagnoses corresponding to at least some of the population types. 12. The computer-implemented method of claim 8, further comprising: performing a dimensionality reduction procedure on a plurality of disease con... | SPX CORPORATION | A61B005/00 G06Q050/00 G06F007/00 G06F015/00
US07828751 | B2 | 2010-11-09 | Method and apparatus for retaining a fixation pin to a cannula | 2007-05-07 | A medical device for placement in an anatomy and for use with a guide wire that generally includes a cannula member configured to be inserted into the anatomy. The cannula member has a wall that includes an outer wall surface and an inner wall surface. The inner wall surface defines an inner passage that extends along a longitudinal axis of the medical device. A guide wire holding member is formed in the wall of the cannula member. The guide wire holding member defines a cavity in the wall of the cannula member and an aperture in the inner wall surface in communication with the cavity that can receive the guide wire from the inner passage of the cannula member. | 1. A medical device for placement in an anatomy and for use with a guide wire, the medical device comprising: a cannula member configured to be inserted into the anatomy, said cannula member having a terminal proximal end, a terminal distal end, and a wall that includes an outer wall surface and a cylindrical inner wall surface, said inner wall surface defining an inner passage that extends along a longitudinal axis of the medical device, said outer wall surface remaining circular in cross section from said terminal proximal end to said terminal distal end; anda guide wire holding member formed in said wall of said cannula member and extending continuously from said terminal proximal end to said terminal distal end of said cannula member, said guide wire holding member defining a cavity in said wall of said cannula member between said outer wall surface and said inner wall surface andwherein said inner wall surface comprises an aperture extending continuously from said terminal proximal end to said terminal distal end of said cannula member, said aperture in communication with said cavity that permits the guide wire to move between said inner passage and said cavity of said cannula member in a direction that is generally perpendicular to said longitudinal axis of the medical device when a longitudinal axis of the guide wire is generally parallel to said longitudinal axis of the medical device. 2. The medical device of claim 1, wherein said cavity has a shape that is selected from a group consisting of a partially circular-shape and a partially oval-shape. 3. The medical device of claim 1, wherein said guide wire holding member is bendable. 4. The medical device of claim 1, wherein said guide wire holding member is steerable in response to motion of the guide wire. 5. The medical device of claim 1, wherein said aperture permits the guide wire to move between said inner passage and said cavity of said cannula member at any location along said longitudinal axis of the medical device. 6. The medical device of claim 1, wherein said aperture permits the guide wire to move between said inner passage and said cavity of said cannula member without moving the guide wire in a direction along said longitudinal axis of the cannula member. 7. A medical device for placement in an anatomy and for use with a guide wire, the medical device comprising: a cannula member configured to be inserted into the anatomy, said cannula member having a terminal proximal end, a terminal distal end, and a wall that includes an outer wall surface and a cylindrical inner wall surface, said inner wall surface defining an inner passage that extends along a longitudinal axis of the medical device; anda guide wire holding member formed in said wall of said cannula member and extending continuously from said terminal proximal end to said terminal distal end of said cannula member, said guide wire holding member defining a cavity in said wall of said cannula member between said ou... | Biomet Sports Medicine, LLC | A61B005/00 A61M025/00
US07848553 | B2 | 2010-12-07 | Method and apparatus of multi-modality image fusion | 2007-10-31 | The present invention is directed to a method and apparatus for fusing or combining functional image data and anatomical image data. The invention, which may be carried out through user interaction or automatically, enables composite and clinically valuable images to be generated that display functional and anatomical data acquired with different imaging systems. By identifying fiducial markers on a functional data image and correlating the fiducial markers with anatomical markers or indicia on the anatomical data image, the respective images may be aligned with one another before a composite image is generated. | 1. A method of medical image overlap comprising the steps of: determining at least two anatomical fiducial markers on a functional image;determining corresponding points to the at least two anatomical fiducial markers on an anatomical image;aligning the at least two anatomical fiducial markers with the corresponding points on the anatomical image;warping the functional image to fit constraints of the anatomical image while maintaining alignment of the at least two anatomical fiducial markers and the corresponding points on the anatomical image; andenforcing anatomical constraints during warping by projecting a nearest point on the functional image onto the anatomical image. 2. The method of claim 1 further comprising the step of accessing a model of functional data prior to determining the at least two anatomical fiducial markers. 3. The method of claim 1 wherein the functional image includes perfusion data and the anatomical image includes anatomical data of a coronary artery. 4. The method of claim 3 wherein the at least two anatomical fiducial markers and the corresponding points on the anatomical image correspond to ventricle grooves between ventricles of a medical patient. 5. The method of claim 4 wherein each of the perfusion data and the anatomical data include gated images acquired at different times. 6. The method of claim 3 wherein the perfusion data includes data acquired with positron emission tomography (PET) and the anatomical data includes data acquired with computed tomography (CT). 7. The method of claim 3 further comprising the step of determining anatomical constraints of the functional image based on cardiac motion. 8. The method of claim 1 wherein the step of determining the at least two anatomical fiducial markers includes the step of locating the at least two anatomical fiducial markers in a three-dimensional image. 9. The method of claim 1 wherein the step of aligning includes registering the functional image and the anatomical image by at least one of translating, scaling, and rotating the functional image and the anatomical image with respect to one another. 10. A diagnostic image generation system comprising: at least one database containing functional and anatomical image data; anda computer programmed to: determine at least a pair of fiducial markers on a functional image;locate corresponding anatomical indicia on an anatomical image; andgenerate a composite image of the functional image and the anatomical image such that the fiducial markers and the anatomical indicia are aligned and anatomical constraints are considered, wherein in being programmed to generate the composite image, the computer is programmed to: warp the functional image such that functional image data is fit to anatomical constraints of the anatomical image; andenforce anatomical constraints during warping by projecting a nearest point on the functional image onto the anatomical image. 11. The system of claim 10 wherein the computer is fur... | General Electric Company | A61B005/05 G06K009/36
US07827041 | B2 | 2010-11-02 | System and methods of providing pharmacy services | 2004-03-31 | System and methods for processing a drug prescription transaction are configured to identify and resolve any issue or problem associated with the transaction during one or more early stages of processing. System and methods of processing the transaction handle prescription and customer data entry, pharmacy inventory check, prescription refill authorization check and/or insurance adjudication review well in advance of production and quality assurance stages by identifying and resolving any issue or problem. Prescription fulfillment is prioritized and estimated prescription pick-up times are predicted based on one or more outcomes of such early stage processing. Customers can be provided with realistic and relatively accurate prescription pick up times that a pharmacy can efficiently and consistently meet as a result of fulfillment prioritization. Further, system and methods for determining a staffing schedule for a workflow process requiring different skill levels during different operaton times is provided. | 1. A workflow management method for processing a drug prescription in a pharmacy, the method being performed by execution of a computer readable program by a computer system, the method comprising: receiving a drug prescription transaction including one or more prescriptions;entering data related to the drug prescription transaction into the computer system to initiate the drug prescription transaction;estimating and recording in the computer system a date and a time by which the drug prescription transaction will be fulfilled and available to a customer for pick up, wherein estimating a date and a time for pick up includes entering an estimated pick up time, the estimated pick up time being determined relative to an estimated pick up time for one or more of a plurality of other prescription transactions requiring fulfillment and arranged in at least one work queue stored in the computer system in accordance with their pick up times;checking the pharmacy drug inventory to confirm whether the one or more prescriptions of the prescription transaction can be fulfilled;receiving data related to an insurance adjudication review or a refill authorization of the prescription transaction;resetting in the computer the estimated pick up time of the prescription transaction in response to receiving data related to at least one of: (i) insufficient drug inventory for fulfillment of the one or more prescriptions-, (ii) problems associated with the insurance adjudication review, and (iii) a number of prescription transactions in the work queue,and, if the prescription transaction cannot be fulfilled within a maximum period of time identified by the computer system, recording a reset estimated pick up time of the prescription transaction in the computer system;entering the prescription transaction in the work queue for fulfillment; andprioritizing in real time the position of the prescription transaction in the work queue in accordance with its estimated pick up or reset estimated pick up time relative to the pick up times of the other prescription transactions in the work queue. 2. The workflow management method of claim 1 further comprising relaying the prescription transaction pick up time to the customer. 3. The workflow management method of claim 1 wherein initiating the prescription transaction includes initiating the prescription transaction while in communication with the customer and reporting an outcome of initiating the prescription transaction to the customer. 4. The workflow management method of claim 1 wherein checking the pharmacy drug inventory includes checking the pharmacy drug invention while in communication with the customer and reporting an outcome of checking the pharmacy drug inventory to the customer. 5. The workflow management method of claim 1 wherein receiving data related to the insurance adjudication review includes receiving data related to the insurance adjudication review or refill authorization while in communication with t... | CVS Pharmacy, Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07823296 | B2 | 2010-11-02 | Depth gauge | 2009-07-13 | A depth gauge for measuring the depth of blind holes in bones and/or nails inserted into bones. The measurement tool comprises: a probe with an abutment at its end to engage against the distal edge of the blind hole; a tubular body; and a grooved insert. After the probe, lying into the tubular body, is positioned through the blind hole to be measured, the grooved insert is slid forward inside the tubular body, surrounding the probe, and pushing the probe sideward, towards the blind hole wall. When probe and groove insert are retracted together, they firmly engage against the distal bone wall. The tubular housing is then pushed forward until it abuts against the proximal bone wall, allowing even in small diameter holes, an accurate reading of hole depth on the scale located at the rear end of the probe. | 1. An instrument for measuring blind holes in one of (a) bones and (b) implants inserted within bones, comprising: a. an elongated probe having a proximal end, a distal end, a longitudinal axis, and a scale on the proximal end of the probe;b. a housing slidably mounted relative to the scale and having a length that is less than a length of the elongated probe;c. an abutment formed at the distal end of the probe, the abutment configured and dimensioned to fit through a hole to be measured; andd. an elongated insert having a longitudinal groove for receiving the probe in a sliding engagement with respect thereto,wherein a retraction of the insert in a direction away from the hole causes movement of the abutment toward a distal wall adjacent the hole. 2. The instrument of claim 1, wherein a distal end of the insert includes a tapered portion. 3. The instrument of claim 1, wherein forward movement of the insert along the probe prevents simultaneous retraction of the probe and the groove from the hole. 4. The instrument of claim 1, wherein indicia extend along the scale on the proximal end of the probe to indicate the depth of the hole. 5. The instrument of claim 1, wherein the proximal end of the probe includes a threaded portion, and the instrument further comprises a threaded element that mates with the threaded portion at the proximal end of the probe. 6. The instrument of claim 5, wherein the threaded element is configured and dimensioned to push the housing against a proximal wall adjacent the hole. 7. The instrument of any claim 1, wherein both the probe and the insert slide within the housing. 8. The instrument of claim 1, wherein the housing is tubular. 9. The instrument of claim 2, wherein the distal end of the insert tapers to a sharp edge. 10. A method for measuring a depth of a blind hole in one of (a) bone and (b) implant inserted into a bone, comprising: a. providing a measurement tool including i. an elongated probe having a proximal end, a distal end, a longitudinal axis, and a scale on the proximal end of the probe;ii. a housing slidably mounted relative to the scale and having a length that is less than a length of the elongated probe;iii. an abutment formed at the distal end of the probe, the abutment configured and dimensioned to fit through a hole to be measured; andiv. an elongated insert having a longitudinal groove for receiving the probe in a sliding engagement with respect theretob. inserting the distal end of the probe through the hole;c. sliding the insert toward the distal end of the probe such that the abutment moves toward a distal wall adjacent the hole;d. simultaneously retracting the probe and the grooved insert until the abutment at the distal end of the probe engages the distal wall adjacent the hole;e. moving the tubular housing against a proximal wall adjacent the hole; andf. reading a depth measurement from the scale located on the probe. 11. The method of claim 10, wherein a distal end of the insert i... | Synthes USA, LLC | A61B005/107
US07856366 | B2 | 2010-12-21 | Multiple accounts for health record bank | 2005-09-30 | A health record databank configured to associate electronic health records with multiple-accounts is provided. A network-accessible storage repository is used to store a plurality of electronic medical records for a plurality of individuals. In turn, individuals are provided with an electronic health records account (EHR account), wherein the EHR account is available to store electronic health records associated with the individual, in the storage repository. The individual may authorize access to the electronic records in a given sub-accounts. Requests for access may include both requests to deposit records into the sub-account, and requests to retrieve and view the records in a given sub-account. | 1. A method for providing controlled access to electronic health records, comprising: configuring a network-accessible storage repository to store a plurality of electronic medical records for a plurality of individuals;providing each of the plurality of individuals with an electronic health care records account (EHR account), wherein each EHR account is available to store, in the storage repository, a portion of the electronic health records associated with each of the respective individuals and wherein electronic records associated with at least one of the respective individuals are further associated with at least one sub-account, wherein the respective individual may specify which sub-accounts to associate with a particular electronic health record and wherein at least one sub-account is a restricted-control sub-account, and wherein the respective individual is prevented from limiting access to selected electronic health records associated with the restricted-control sub-account once the respective individual has authorized an entity to access the restricted-control sub-account;in response to a request for access to a particular sub-account received over a data-communications medium, determining whether the respective individual has authorized the entity making the request to access to the particular sub-account; andif so, granting access to the electronic records associated with the particular sub-account. 2. The method of claim 1, wherein at least one sub-account is configured to be associated with electronic health records based on which entity created the electronic health record. 3. The method of claim 1, wherein at least one sub-account is configured to be associated with electronic health records based on a type of the electronic health record. 4. The method of claim 1, wherein the request for access to the particular sub-account is made in order to deposit an electronic health record into the particular sub-account. 5. The method of claim 4, further comprising: in response to the request, storing an electronic health record included with the request in the network-accessible storage repository; andassociating the electronic record with the particular sub-account. 6. The method of claim 1, wherein the request for access to the particular sub-account is made to retrieve electronic records from the particular the sub-account. 7. The method of claim 6, further comprising: retrieving the at least one of the plurality of electronic health records; andreturning the retrieved electronic health records to the entity. 8. A computer-readable medium, containing a program, which when executed by a computer system performs operations for providing patient-controlled access to electronic health records, the operations comprising: configuring a network-accessible storage repository to store a plurality of electronic medical records for a plurality of individuals;providing each of the plurality of individuals with an electronic health care reco... | International Business Machines Corporation | A61B005/00 G06F019/00
US08506495 | B2 | 2013-08-13 | Implantable medical devices with piezoelectric anchoring member | 2010-05-06 | Methods, systems, and apparatus for powering and/or recharging medical devices implanted within the body are described. An illustrative implantable medical device includes a housing having an internal cavity and a flexible anchor assembly that is coupled to the housing. The flexible anchor assembly includes a first electrical conductor, a second electrical conductor, and a piezoelectric layer that is disposed between the first and second electrical conductors and that is configured to displace in response to a physiologic force applied to the flexible anchor assembly and generate a voltage differential between the first and second electrical conductors. The implantable medical device includes power circuitry that converts the voltage differential between the first and second electrical conductors into an operating current for powering one or more components within the implantable medical device and/or for recharging a rechargeable power supply within the device. | 1. An implantable medical device, comprising: a housing including an internal cavity;a flexible anchor assembly coupled to the housing and configured to secure the implantable medical device within a body lumen, the flexible anchor assembly comprising a helical structure including: a helical piezoelectric layer having a first surface and an opposing second surface;a first electrically conductive layer disposed in contact with the first surface of the helical piezoelectric layer;a second electrically conductive layer disposed in contact with the second surface of the helical piezoelectric layer, the helical piezoelectric layer configured to displace in response to periodic diameter fluctuations of the body lumen and generate a voltage differential between the first and second electrically conductive layers; andpower circuitry configured to convert the voltage differential between the first and second electrically conductive layers into an operating current for powering one or more components of the implantable medical device. 2. The implantable medical device of claim 1, further comprising a sensor module disposed within the interior cavity of the housing, the sensor module configured to sense one or more physiologic parameters within the body lumen. 3. The implantable medical device of claim 1, further comprising a wireless communication module disposed within the internal cavity. 4. The implantable medical device of claim 1, further comprising a rechargeable power supply disposed within the internal cavity. | Cardiac Pacemakers, Inc. | A61B005/02
US09364164 | B2 | 2016-06-14 | Non-invasive device and method for locating a structure such as a nerve | 2008-11-06 | Non-invasive device for locating a structure (N2), such as a nerve, in a region of a body having an external surface (Se), the device comprising: a focused ultrasonic transducer (1), advantageously of the HIFU type, designed to produce an ultrasound beam (Fu) through the external surface (Se) in the region, the beam (Fu) having a focal point that can be positioned on the structure (N2); monitoring means (3) for detecting a response of the structure when this is subjected to a stimulation; and stimulation means (1; 4) for stimulating the structure (N2), which delivers a response detected by the monitoring means (3), the stimulation means comprising the ultrasonic transducer (1) capable of delivering an ultrasound beam (Fu) to said structure (N2), so as to disturb it, characterized in that the device includes: an imager (2, 21, 22) mechanically coupled to the ultrasonic transducer (1) so that the imager follows the focal point of the transducer so as to display, on an image, the position of the focal point of the ultrasound beam (Fu) in said region of the body; and marking means, for marking the position of the structure (N2) in the image produced by the imager. | 1. A device for noninvasively locating a structure, such as a nerve, in a zone of a body having an external surface, the device comprising: a focused ultrasonic transducer of a HIFU type, adapted to produce an ultrasound beam through the external surface into the zone, the beam having a focal point which can be positioned on the structure;at least one electrode which can deliver an electrical pulse to said structure for stimulating the structure; anda detector for detecting a response of the structure when subjected to a stimulation by said at least one electrode;wherein the device further comprises:an imaging apparatus coupled mechanically to the ultrasonic transducer so that the imaging apparatus follows the focal point of the transducer in order to visualize on an image the position of the focal point of the ultrasound beam in the zone of the body, said imaging apparatus being unable to see the structure,the imaging apparatus is further adapted to mark a position of the structure by marking the focal point on the image produced if the detector detects that a response of the structure to said stimulation is perturbed by the ultrasound beam, thereby determining the position of the structure, andan emitter for emitting synchronization pulses, the emitter being coupled both to the transducer and to the electrode adapted for synchronizing the ultrasound beam and the stimulation so that they travel through the structure at the same time or with a predetermined time offset, and displacement control for changing the position of the transducer if no perturbance of the response of the structure to said stimulation is detected. 2. The locating device as claimed in claim 1, wherein the emitter comprise offset means adapted to generate an offset in time between the ultrasound beam and the stimulation, so that the ultrasound beam arrives at the structure before the stimulation. 3. The locating device as claimed in claim 1, further comprising calculation means for determining the depth of the structure on the basis of the acoustic time of flight of the ultrasound beam in the zone and the propagation speed of the stimulation. 4. The locating device as claimed in claim 1, wherein the detector comprises an apparatus for monitoring nerve response or an apparatus for monitoring muscle response, such as an electromyograph, an endoscope ora pressure sensor. 5. A method for noninvasively locating a structure, such as a nerve, in a zone of a body having an external surface wherein said method comprises: insonifying the structure from the external surface,stimulating said structure with electrical pulses,monitoring whether the response of the structure to said stimulation is perturbed by the insonification,visualizing the position of the insonification of the structure by marking a position of the structure if a perturbation of the response of the structure is monitored, thereby determining the position of the structure,synchronizing the insonification and the sti... | Theraclion SAS | A61B008/00 A61B005/0488 H02K047/04 H02K053/00 A61B005/05 A61B005/04 A61B005/00 A61B008/14
US20100317960 | A1 | 2010-12-16 | THERMOTHERAPY DEVICE AND METHOD TO IMPLEMENT THERMOTHERAPY | 2010-06-10 | A thermotherapy device has a transmitter, a receiver and a processing unit. The transmitter is designed to emit a high-energy radiation in a treatment region of a patient. The high-energy radiation exhibits a power that is suitable for thermotherapy. The receiver is designed to detect a sound signal that is generated by the treatment region depending on the high-energy radiation radiated into said treatment region. The processing unit is coupled with the transmitter) and the receiver. The processing unit automatically determines information about the treatment region depending on the detected sound signal. | 1. A thermotherapy device comprising: a radiation transmitter that emits high-energy radiation into a treatment region of a patient, said high-energy radiation having a thermotherapy-implementing power;a receiver that detects a sound signal generated by said treatment region dependent on the high-energy radiation radiated into the treatment region, said receiver emitting a receiver output corresponding to the detected sound signal; anda processing unit connected to said transmitter and said receiver that automatically determines information characterizing said treatment region from said receiver output signal. 2. A device as claimed in claim 1 wherein said processing unit is configured to control said transmitter to emit said high-energy radiation as a high-energy radiation pulse having a length for thermoacoustic imaging, and to operate said receiver to detect said sound signal synchronized with emission of said high-energy radiation pulse. 3. A device as claimed in claim 2 wherein said processing unit is configured to control said transmitter to emit additional high-energy radiation pulses dependent on said information. 4. A device as claimed in claim 1 wherein said receiver detects an ultrasound signal as said sound signal. 5. A device as claimed in claim 1 comprising an image processor connected to said processing unit, said image processor being configured to generate image information characterizing said treatment region from said information characterizing said treatment region generated by said processing unit. 6. A device as claimed in claim 5 wherein said image processor is configured to determine at least one of a target structure of said treatment region and a prominent structure in an environment of said treatment region. 7. A device as claimed in claim 5 wherein said image processor is configured to determine an energy deposition in said treatment region. 8. A device as claimed in claim 5 wherein said image processor is configured to determine a temperature in said treatment region. 9. A device as claimed in claim 1 wherein said transmitter is a transmitter is a transmitter selected from the group consisting of high-energy focused ultrasound wave transmitters, radio-frequency wave transmitters, and laser light wave transmitters. 10. A device as claimed in claim 1 wherein said processing unit is configured to determine information characterizing said treatment region selected from the group consisting of a temperature of the treatment region, a temperature change of the treatment region, an anatomical image of the treatment region, a structure of the treatment region, a structural change of the treatment region, a change in physiology of the treatment region, a change in cellular markers in the treatment region, and a change in molecular markers in the treatment region. 11. A device as claimed in claim 1 wherein said transmitter is configured to also emit a low-energy radiation pulse into said treatment region, said low-ene... | - | A61N007/02 A61B008/14 A61B005/055
US20100305412 | A1 | 2010-12-02 | Device and system for wireless monitoring of the vital signs of patients | 2010-03-22 | The present invention relates generally to the field of vital signs monitors, and particularly to a device and system for wireless monitoring of the vital signs of a plurality of patients simultaneously yet independently. The present invention is particularly adaptable for use by medical emergency personnel or medics in any setting, such as road accident, disaster sites, combat zones, or hospitals. The systems and methods of the present invention allows the monitoring of a plurality of patients by a single health care professional providing more effective monitoring and care for a large number of patients at any one time, and retaining all the information for more effective later triage and decision making by health care providers. | 1. An apparatus for wireless monitoring of the vital signs of one or more patients simultaneously, comprising: a plurality of patient mounted vital signs monitors, each comprising: a plurality of sensors for detecting and measuring the vital signs of a patient;a first display for displaying the patient's vital signs, the display being operatively connected to the plurality of sensors, wherein the display can receive and display the patient vital signs;a first transceiver operatively connected to the patient mounted vital signs monitor, and operable to transmit the patient's vital signs data to a remote transceiver; anda first processing means for processing the detected or measured vital signs and for controlling the operation of at least the first display and the first transceiver;a user mounted or carried processing system and monitor comprising: a second transceiver for wireless connection to the first transceivers of the plurality of the patient mounted vital signs monitors;a second display configured to separately display each of the patients' vital signs data received from the plurality of the patient mounted vital signs monitors; anda second processing means for processing the received vital signs data from the plurality of the patient mounted vital signs monitors and for controlling the operation of at least the second display and the second transceiver;whereby, users can monitor and record current and history of the vital signs of a plurality of patients simultaneously by way of the patient mounted vital signs monitor or by way of the user mounted display. 2. The apparatus of claim 1, further comprising a communications module co-operable with at least the user mounted processing system and monitor to receive the vital signs data and configured to transmit the vital signs data to a remote location. 3. The apparatus of claim 1, wherein the plurality of sensors comprises sensors for electrocardiogram, non-invasive blood pressure, heart rate, pulse waveform, respiration rate, and saturation of blood oxygen. 4. The apparatus of claim 1, wherein the plurality of sensor data are processed to calculate additional vital signs data such as shock index, pulse integrity, pulse wave delay, cardiac rhythm complexity, and heart rate variability as an indicator of changes in non-invasive blood pressure, autonomic nervous system integrity, and a summary alarm feature that indicates overall vital signs state of the patient. 5. The apparatus of claim 1, wherein the first display of the patient mounted vital signs monitor comprises a flexible or fixed display disposed to be placed or stuck on a patient and flex with the patient's body. 6. The apparatus of claim 1, wherein the first and the second transceivers wirelessly communicate by way of IEEE 802.11 b or 802.11 g, or 802.15.4 Zigbee, Bluetooth 802.15.1, or a smart phone. 7. The apparatus of claim 1, wherein the second display of the user mounted processing system and monitor is configured to meas... | - | A61B005/00
US20100298677 | A1 | 2010-11-25 | WIRELESS RING-TYPE PHYSICAL DETECTOR | 2009-05-22 | A wireless ring-type physical detector includes a ring, a sensor unit, an amplifier unit, a demultiplexer unit, a processor unit and a wireless transmission unit. The sensor unit uses a light signal to detect the blood oxygen saturation, the heartbeat and continuous blood pressure. The detected light signal is processed by each unit to get a physical parameter which is valuable for a clinic test. | 1. A wireless ring-type physical detector, comprising: a ring, the ring comprising a chamber, an inner surface, and a display screen;a sensor unit, the sensor unit comprising two transmitters and a receiver disposed on the inner surface of the ring, the transmitters and the receiver using a light signal to detect the blood oxygen saturation, the heartbeat and continuous blood pressure and transforming the light signal into a current signal;an amplifier unit, the amplifier unit being disposed in the chamber of the ring and connected to the sensor unit for transforming the current signal into a voltage signal;a demultiplexer unit, the demultiplexer unit being disposed in the chamber of the ring and connected to the amplifier unit for outputting a demultiplex signal after selection and allotment;a processor unit, the processor unit being disposed in the chamber of the ring and connected with the sensor unit, the amplifier unit, and the demultiplexer unit for controlling each unit, the processor unit receiving and calculating the demultiplex signal to come up with a physical parameter and outputting the physical parameter to the display screen; anda wireless transmission unit, the wireless transmission unit being disposed in the chamber of the ring and connected to the processor unit, the wireless transmission unit receiving and transmitting the physical parameter. 2. The wireless ring-type physical detector as claimed in claim 1, wherein the ring comprises a power supply unit therein. 3. The wireless ring-type physical detector as claimed in claim 1, wherein the two transmitters of the sensor unit are disposed on a first side of the inner surface of the ring, the light signal from the transmitters pointing at a first position at a second side of the inner surface of the ring, the receiver being located at the first position. 4. The wireless ring-type physical detector as claimed in claim 1, wherein the two transmitters of the sensor unit are disposed on a first side of the inner surface of the ring, the light signal from the transmitters pointing at a second side of the inner surface of the ring and returning back to a second position at the first side of the inner surface, the receiver being located at the second position. 5. The wireless ring-type physical detector as claimed in claim 1, wherein the amplifier unit is a transimpedence amplifier. 6. The wireless ring-type physical detector as claimed in claim 1, wherein the processor unit comprises an analog digital transformer for transforming a signal to be transmitted to the processor unit into a digital signal. 7. The wireless ring-type physical detector as claimed in claim 1, wherein the processor unit comprises a digital analog transformer for transforming an output signal into an analog signal before transmission. 8. The wireless ring-type physical detector as claimed in claim 1, wherein the wireless transmission unit is a wireless transmission module selected from one of infrared ray,... | ASTEK TECHNOLOGY LTD. | A61B005/02 A61B005/1455
US20100298726 | A1 | 2010-11-25 | LOW-PRESSURIZATION BLOOD PRESSURE MONITORING APPARATUS AND METHOD | 2010-05-12 | A blood pressure monitoring apparatus including a pressing unit which presses a measurement body part of a subject, a sensing unit which senses a sphygmus wave at the measurement body part while the measurement body part is being pressed, a control unit controlling a point of time of stopping the pressing based on an amplitude of the sensed sphygmus wave, and an estimation unit which estimates a blood pressure of the subject based on the sphygmus wave sensed before the pressing performed by the pressing unit is stopped. | 1. A blood pressure monitoring apparatus comprising: a pressing unit which presses a measurement body part of a subject;a sensing unit which senses a sphygmus wave at the measurement body part;a control unit which stops the pressing performed by the pressing unit based on an amplitude of the sensed sphygmus wave; andan estimation unit which estimates a blood pressure of the subject, based on the sphygmus wave sensed before the pressing performed by the pressing unit is stopped. 2. The blood pressure monitoring apparatus of claim 1, wherein the control unit stops the pressing performed by the pressing unit after the amplitude of the sensed sphygmus wave has reached a maximum. 3. The blood pressure monitoring apparatus of claim 1, further comprising a monitoring unit which monitors the amplitude of the sensed sphygmus wave, wherein the control unit stops the pressing performed by the pressing unit when the amplitude of the sensed sphygmus wave decreases, as a result of the monitoring of the monitoring unit. 4. The blood pressure monitoring apparatus of claim 1, wherein the estimation unit comprises a determination unit which determines a first pressure having a maximum amplitude of the sensed sphygmus wave and a second pressure having a ratio of amplitude at which blood pressure is to be estimated, with respect to the maximum amplitude of the sensed sphygmus wave, and the estimation unit estimates the determined first and second pressures from the determination unit, as the blood pressure of the subject. 5. The blood pressure monitoring apparatus of claim 4, wherein the estimation unit further comprises a calculation unit which calculates a systolic blood pressure based on the first pressure and the second pressure determined by the determination unit, and the estimation unit estimates the calculated systolic blood pressure from the calculation unit, as the blood pressure of the subject. 6. The blood pressure monitoring apparatus of claim 5, wherein the calculation unit calculates the blood pressure of the subject by using a relational equation representing a correlation between the first pressure, the second pressure, and the systolic blood pressure. 7. The blood pressure monitoring apparatus of claim 1, wherein the control unit stops the pressing performed by the pressing unit when a pressure applied by the pressing unit reaches a pressure at which blood vessels are occluded. 8. The blood pressure monitoring apparatus of claim 7, further comprising an input unit which obtains an input signal for selecting a pressurization mode which determines a degree of pressing by the pressing unit, wherein the control unit determines a point of time when the pressing performed by the pressing unit is stopped according to the input signal. 9. The blood pressure monitoring apparatus of claim 7, wherein the estimation unit further comprises a determination unit which determines at least one selected from the group consisting of mean arterial pressure, di... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/0225
US20100292565 | A1 | 2010-11-18 | MEDICAL IMAGING MEDICAL DEVICE NAVIGATION FROM AT LEAST TWO 2D PROJECTIONS FROM DIFFERENT ANGLES | 2009-05-18 | In a method or system for performing an image-assisted medical procedure involving placing a medical device in a 3D human subject, at least first and second 2D projected images are obtained of the 3D subject at a respective first angle and at a respective second angle, the 3D subject having at least one of the points selected from the group consisting of a start point and an end point for guidance of the medical device. At least one of the start point and the end point are identified in each of the first and second 2D images. 3D positions of at least one of the start point and the end point are calculated. At least one of these start and end point 3D positions are overlaid as a back-projection to a live 2D image. The procedure is performed utilizing the live 2D image with at least one of the start and end points as a guide to place the medical device. The medical device may comprise, for example, a wire or a needle. | 1. A method for performing an image-assisted medical procedure involving placing a medical device in a 3D human subject, comprising the steps of: obtaining at least a first and a second 2D projected images of the 3D subject at a respective first angle and at a respective second angle, said 3D subject having at least one of the points selected from the group consisting of a start point and an end point for guidance of said medical device;identifying at least one of said start point and said end point in each of the first and second 2D images;calculating 3D positions of at least one of the start point and the end point;overlaying at least one of said start and end point 3D positions as a back-projection to a live 2D image; andperforming the procedure utilizing the live 2D image with at least one of the start and end points as a guide to place said medical device. 2. The method of claim 1 wherein said medical device comprises a wire. 3. The method of claim 1 wherein said medical device comprises a needle. 4. The method of claim 1 wherein said medical device comprises a wire, both said start point and said end point are identified, and said wire is guided from said start point to said end point. 5. The method of claim 1 wherein said medical device comprises a needle, both said start point and said end point are identified, and said needle is guided from said start point to said end point. 6. The method of claim 1 wherein the live 2D image comprises a live 2D fluoro image. 7. The method of claim 1 including the step of calculating said 3D positions of at least one of the start point and the end point based on at least one of the identified start and end points in each of the first and second 2D images and the first and second angles. 8. The method of claim 1 wherein both said start and end points are identified, and a 3D line is calculated between the start and end points. 9. The method of claim 1 wherein said live 2D image includes a recalculated 3D position of at least one of the start and the end points. 10. The method of claim 1 wherein said medical device comprises a wire and said medical procedure involves placing said wire in said 3D human subject to open up a stenosis. 11. The method of claim 10 wherein said stenosis comprises a chronic total occlusion blood vessel. 12. The method of claim 1 wherein said medical device comprises a wire and said medical procedure comprises placing said wire as a guide wire in said 3D human subject to guide a stenting. 13. The method of claim 12 wherein said stenting comprises an intracranial stenting. 14. The method of claim 12 wherein said stenting comprises stenting of an abdominal aortic aneurysm. 15. The method according to claim 1, wherein the step of identifying at least one of the start point and the end point comprises manually selecting the at least one point on an image display with a pointing device. 16. The method according to claim 1, further comprising providing a 3D dataset of a fe... | - | A61B005/05
US07840269 | B2 | 2010-11-23 | Analysis of eating habits | 2007-08-31 | Diet evaluation gastric apparatus (18) is provided, which detects when a patient (10) swallows, and detects the type and amount of matter ingested. The apparatus includes electrodes (74, 100) adapted to be coupled to the fundus and antrum of the patient and to measure electrical and mechanical activity therein, and a control unit (90) to analyze such electrical and mechanical activity and optionally apply electrical energy to modify the activity of tissue of the patient. | 1. A gastric apparatus, comprising: a set of one or more implantable antral electrodes, configured to be coupled to an antral site of an antrum of a subject and to generate an antral electrode signal responsive to a property of the antrum; anda control unit, configured to: receive and analyze the antral electrode signal,determine a time delay between an onset of eating by the subject and a change in electrical activity in the antrum, the change in electrical activity determined responsive to the antral electrode signal, andresponsively to the delay, determine whether food ingested by the subject is predominantly solid or liquid. 2. An apparatus according to claim 1, wherein the control unit is adapted to determine, responsive to the time delay, an extent to which the ingested food includes solid food matter. 3. An apparatus according to claim 1, further comprising a set of one or more implantable fundic electrodes, configured to be coupled to a fundic site of a fundus of a stomach of the subject, and wherein the control unit is configured to identify the onset of eating using the set of fundic electrodes. 4. An apparatus according to claim 3, wherein the control unit is configured to identify, as indicative of the onset of eating, an increase in an amplitude of an electrical potential change generated responsive to a contraction of a muscle of the fundus, detected using the set of fundic electrodes. 5. An apparatus according to claim 1, wherein the change in the electrical activity in the antrum is an onset of a decreased rate of electrical events in the antrum, and wherein the control unit is configured to determine the time delay between the onset of eating and the onset of the decreased rate of the electrical events in the antrum. 6. An apparatus according to claim 5, wherein the control unit is configured to determine, responsive to the time delay, an extent to which the ingested food includes solid food matter. 7. An apparatus according to claim 5, wherein the control unit is configured to determine, responsive to the time delay, an extent to which the ingested food includes liquid food matter. 8. An apparatus according to claim 1, wherein the change in the electrical activity in the antrum is an onset of increased electrical activity in the antrum, and wherein the control unit is configured to determine the time delay between the onset of eating and the onset of the increased electrical activity in the antrum. 9. An apparatus according to claim 8, wherein the control unit is configured to determine, responsive to the time delay, an extent to which the ingested food includes solid food matter. 10. An apparatus according to claim 8, wherein the control unit is configured to determine, responsive to the time delay, an extent to which the ingested food includes liquid food matter. 11. An apparatus according to claim 1, and comprising a set of one or more current-application electrodes, configured to be coupled to a tissue of the subje... | Metacure Limited | A61B005/053
US07824344 | B1 | 2010-11-02 | Method and apparatus for non-invasive analysis of saliva | 2008-07-18 | A method and apparatus for non-invasive analysis of saliva includes a user friendly lolli-pop like apparatus that is self contained and provides the user with a method to stimulate saliva by sucking on the pop head. The saliva activates a color change reaction on the apparatus that results in a visual response that can be compared to a user gauge for analysis in monitoring health conditions and concerns. | 1. A self-contained apparatus for conducting a non-invasive analysis of saliva to determine a level of a certain constituent present in the saliva, the apparatus comprising: a flat card-like device having a flat configuration resembling a relatively thin card;a head having a lolli-pop-like configuration for receiving on the head saliva from a user;a means on the head capable of reacting with the certain constituent to display a visible color change to the head indicative of a level of the certain constituent in the saliva received on the head;a stem connected to the head to provide the user with a simple means to hold the apparatus; anda gauge including visible areas of color representing different levels of the certain constituent, the areas of color being displayed on the card-like device, placed relative to the head and the stem for enabling juxtaposition of the head with the gauge for a visual comparison of the color change at the head with the areas of color of the gauge, thereby providing a direct visual determination of the level of the certain constituent in the saliva received on the head, the head and the stem being die-cut within the card-like device so that the stem is pivoted out of the flat configuration for placement of the head in juxtaposition with a user's mouth to facilitate reception of the saliva on the head, and is pivoted subsequently back into the card-like device to insert the head within the card-like device confined in side-by-side contiguity with the gauge to facilitate a direct visual comparison of the color change of the head with the areas of color of the gauge. 2. The invention of claim 1 wherein the certain constituent is associated with pregnancy and the analysis of the saliva is for pregnancy determination. 3. The invention of claim 1 wherein the certain constituent is alcohol and the analysis of the saliva is of alcohol levels. 4. The invention of claim 1 wherein the certain constituent is keto and the analysis of the saliva is of keto levels. 5. The invention of claim 1 wherein the certain constituent is a drug and the analysis of the saliva is for drug presence. 6. The invention of claim 1 wherein the gauge includes specific numbers and ranges of color to reflect the results of the particular saliva analysis. 7. A method for conducting, within a self-contained apparatus, a non-invasive analysis of saliva to determine a level of a certain constituent present in the saliva, the method comprising: providing a flat card-like device having a flat configuration resembling a relatively thin card;providing a head die-cut within the card-like device, the head having a lolli-pop-like configuration and a means capable of reacting with the certain constituent for displaying a visible color change to the head indicative of a level of the certain constituent in the saliva;providing a stem die-cut within the card-like device and connected to the head to provide the user with a simple means to hold the apparatus;provid... | Pop Test LLC | A61B005/00 B65D081/00 G01N021/00 G01N031/22
US20100331636 | A1 | 2010-12-30 | METHOD FOR THE CONTINUOUS NON-INVASIVE DETERMINATION OF THE CONCENTRATION OF BLOOD CONSTITUENTS | 2008-03-20 | The invention relates to a method for the non-invasive determination of the concentration of blood constituents, in which a radiation source (12) emits several radiation beams (14), each with a different wavelength. A first photo detector (18) receives the measurement radiation (14) of each wavelength that is reflected by a body part (16) to be examined. A second photo detector (22) receives the measurement radiation (24) of each wavelength that is transmitted by the body part (16) to be examined. The measurement radiation (14) of each wavelength that is absorbed by the body part (16) to be examined is then determined on the basis of the measurement of the reflected radiation (20) by the first radiation receiver (18) and the measurement of the transmitted radiation (24) by the second radiation receiver (22). The concentration of the different constituents is calculated from the absorption of the measurement radiation (14) that has been determined for each wavelength. | 1. A method for the non-invasive determination of the concentration of blood constituents, comprising: radiating a plurality of measurement radiations with respectively different wavelengths by means of a radiation source;receiving the measurement radiation of a plurality of said wavelengths, as reflected by the body part to be examined, by means of a first light receiver;receiving the measurement radiation of a plurality of said wavelengths, as transmitted through the body part to be examined, by means of a second light receiver; anddetermining the absorption of the measurement radiation of each wavelength, as caused by the body part to be examined, on the basis of the measurement of said reflected radiation by means of the first radiation receiver and the measurement of said transmitted radiation by means of the second radiation receiverwherein that an absorption AgR is determined by the first radiation receiver on the basis of the measurement of the reflected radiation of a first portion r to ri of the wavelengths, and an absorption AgT is determined by the second radiation receiver on the basis of the measurement of the transmitted radiation of a second portion 1 to r of the wavelengths, wherein, for at least one wavelength r, the reflected as well as the transmitted radiation fraction is measured, comprising the further step of:determining the n concentrations Cb1 to Cbn of blood constituents on the basis of said determined absorptions AgR and respectively AgT of the measurement radiation at each wavelength 1 to n, the computation of the concentrations being performed according to the following equation system: AT(λ1)=E(λ1,b1)*Cb1+ . . . +E(λ1,bn)*Cbn AgT(λT)=E(λr,b1)*Cb1+ . . . +E(λr,bn)*Cbn AgR(λr)=(KT/KR)*{E(λr,b1)*Cb1+ . . . E(λr,bn)*Cbn}AgR(λn)=(KT/KR)*{E(λn,b1)*Cb1+ . . . +E(λn,bn)*Cbn} 2. The method according to claim 1, wherein said method steps are repeated a plurality of times and the respective absorption values for each wavelength of the measurement radiation are stored for each repetition cycle. 3. The method according to claim 2, wherein the stored absorption values for each wavelength are combined into a representation of the temporal development of the absorption for each wavelength of the measurement radiation. 4-7. (canceled) 8. The method according to claim 1, wherein the preceding step: determining a normalization factor for the light intensity of the radiation source. 9. The method according to claim 1, wherein, at least once per measurement, the constant component of the obtained development of the absorption is determined for each wavelength. 10. The method according to claim 9, wherein the determined alternating component of the development of the absorption for each wavelength is normalized in dependence on the determined constant component for each wavelength. 11. The method according to claim 1, wherein analog and/or digital filters are used for separating the DC and alternating components of the detected de... | Enverdis GmbH | A61B005/1455
US20100286539 | A1 | 2010-11-11 | BLOOD PRESSURE MEASUREMENT DEVICE FOR PERFORMING PROCESS TAKING CHANGE OF MEASUREMENT ENVIRONMENT INTO CONSIDERATION | 2009-01-15 | A blood pressure measurement device includes a measurement fluid bag, a sensor for measuring a change in internal pressure of the measurement fluid bag, a blood pressure measurement unit for calculating a blood pressure value based on the change in internal pressure of the measurement fluid bag obtained by the sensor, an environmental information measurement unit for measuring environmental information in association with a calculation process of the blood pressure value, a storage unit for storing a measurement result, a recording processing unit for storing the calculated blood pressure value and the environmental information in the storage unit in association with each other and a notification unit for notifying a blood pressure fluctuation with respect to an environmental fluctuation based on the blood pressure value and the environmental information. | 1. A blood pressure measurement device comprising: a measurement fluid bag;a sensor for measuring a change in internal pressure of the measurement fluid bag;a blood pressure measurement unit for calculating a blood pressure value based on the change in internal pressure of the measurement fluid bag obtained by the sensor;an environmental information measurement unit for measuring environmental information in association with a calculation process of the blood pressure value;a storage unit for storing a measurement result;a recording processing unit for storing the calculated blood pressure value and the environmental information in the storage unit in association with each other; anda notification unit for notifying a blood pressure fluctuation with respect to an environmental fluctuation based on the blood pressure value and the environmental information. 2. The blood pressure measurement device according to claim 1, further comprising: a display unit, whereinthe notification is made by displaying on the display unit a numerical value indicating a proportion of the blood pressure fluctuation with respect to the environmental fluctuation. 3. The blood pressure measurement device according to claim 1, further comprising: a display unit, whereinthe notification is made by displaying on the display unit information on a range of an environment having the blood pressure fluctuation possibly outside a predetermined allowable range. 4. The blood pressure measurement device according to claim 1, further comprising: a classifying unit for classifying the blood pressure value and the environmental information into a predetermined region of the environmental information, whereinthe notification indicates that the blood pressure value in a region of interest of the environmental information is included in a predetermined region. 5. The blood pressure measurement device according to claim 1, further comprising: a display unit, whereinthe notification is made by displaying on the display unit that a proportion of the blood pressure fluctuation with respect to the environmental fluctuation is classified into a predetermined rank. 6. The blood pressure measurement device according to claim 2, further comprising a recommended notification unit for notifying to urge a recommended action to a user. 7. The blood pressure measurement device according to claim 1, further comprising: an approximation unit for calculating an approximate estimation curve of the blood pressure value with respect to the environmental information based on the measured environmental information and the calculated blood pressure value; andan estimation unit for calculating a blood pressure value at a non-measured value of environmental information based on the approximate estimation curve. 8. The blood pressure measurement device according to claim 1, further comprising a monitoring region calculating unit for calculating a monitoring range of the environmental information where the b... | OMRON HEALTHCARE CO., LTD. | A61B005/0225
US07824403 | B2 | 2010-11-02 | Methods and devices for ablation | 2006-04-11 | An ablating device has a cover which holds an interface material such as a gel. The cover contains the interface material during initial placement of the device. The ablating device may also have a removable tip or a membrane filled with fluid. In still another aspect, the ablating device may be submerged in liquid during operation. | 1. A sizing device for sizing an area around pulmonary veins along an epicardial surface, the sizing device comprising an elongate body having a substantially straight proximal portion and a distal portion preshaped into a partial loop when not subjected to any deforming force; anda plurality of sizing indicators along a length of the elongate body,wherein the plurality of sizing indicators is used to measure a length of a perimeter of the epicardial surface when the sizing device is placed around the pulmonary veins along an epicardial surface. 2. The sizing device of claim 1, wherein the distal portion terminates in a tip, and wherein the tip is offset with respect to the elongated elongate body. 3. The sizing device of claim 1, further comprising a connector at a proximal end of the body for connecting the sizing device to an ablation device. 4. The sizing device of claim 3, wherein the connector is a snap-fit connector. 5. The sizing device of claim 1, wherein a cross section of the elongate body at a proximal end thereof is larger than a cross section of the elongate body at a distal end thereof. 6. A sizing device for sizing an area around pulmonary veins along an epicardial surface, the sizing device comprising an elongate body having a distal end and a proximal end, the elongate body being preshaped into a partial loop when in a resting, unbiased state; a connector at the proximal end for connecting the sizing device to an ablation device having at least one ablating element; anda plurality of sizing indicators along a length of the elongate body, wherein the plurality of sizing indicators is used to identify an ablation device that is sized to fit around the pulmonary veins along an epicardial surface. 7. The sizing device of claim 6, wherein the connector is a snap-fit connector. 8. The sizing device of claim 6, wherein the distal end further comprises a tip, and wherein the tip is offset with respect to the elongate body. 9. The sizing device of claim 6, wherein a cross section of the elongate body at the proximal end is larger than a cross section of the elongate body at the distal end. 10. A system for ablating cardiac tissue, the system comprising a sizing device having an elongate body with a distal end and a proximal end, the elongate body being configured to form a partial loop when in a resting, unbiased shape; a connector at the proximal end; and a plurality of sizing indicators along a length of the elongate body; and a plurality of ablating devices, each having a different length, the plurality of ablating devices corresponding to the plurality of sizing indicators,wherein the plurality of sizing indicators is used to select one of the plurality of ablating devices such that the selected ablating device is sized to fit around pulmonary veins along an epicardial surface. 11. The system of claim 10, wherein the ablating devices further comprise an opening for receiving the connector at the proximal end of the sizing d... | St. Jude Medical, Atrial Fibrillation Division, Inc. | A61B018/04 A61B005/103
US07826878 | B2 | 2010-11-02 | Optical tomography using independent component analysis for detection and localization of targets in turbid media | 2005-12-07 | Disclosed is a system and a method for detecting the presence of one or more objects in a turbid medium, the method including: illuminating at least a portion of the turbid medium with incident light having at least one wavelength which interacts with the one or more objects contained in the turbid medium differently than the incident light interacts with the turbid medium; measuring light that emerges from the turbid medium; and detecting and locating the one or more objects using Independent Component Analysis (ICA) of the emergent light from the turbid medium. The present invention is useful for medical applications, such as for finding and locating, a tumor(s) in body organs, or excised tissues. Moreover, the present invention can be used to locate objects in obscuring medium, such as, mines in shallow coastal water, a plane in fog, military targets under fog, smoke or cloud cover. | 1. A method for detecting a presence of one or more objects in a turbid medium, comprising: illuminating, by a light source, at least a portion of the turbid medium with incident light having at least one predetermined wavelength which interacts with the one or more objects contained in the turbid medium differently than the incident light interacts with the turbid medium;measuring, by one or more light detectors, light that emerges from the turbid medium;detecting and locating, by a processor, the one or more objects located within the turbid medium based on an Optical Imaging Independent Component Analysis (OPTICA) of the emergent light from the turbid medium; andextracting three-dimensional location information about the objects from the processor,wherein the incident light includes at least two different wavelengths, measurements are made at different wavelengths to obtain at least two signals Iλ1, and Iλ2, which are compared and correlated, and information about the one or more objects inside the turbid medium is extracted by processing the results of the comparison and correlation. 2. The method as claimed in claim 1, wherein, the light that emerges from the turbid medium has a same wavelength as the incident light, is detected using a light detector, and is analyzed using the OPTICA to determine independent components, and wherein a location of the one or more objects is obtained based on knowledge of the independent components. 3. The method as claimed in claim 1, wherein the light that emerges includes a plurality of wavelengths, at least one of the wavelengths being different from the at least one wavelength of the incident light. 4. The method as claimed in claim 2, wherein the light that emerges includes a plurality of wavelengths, at least one of the wavelengths being different from a wavelength of the incident light. 5. The method as claimed in claim 1, wherein the light source includes at least one of a light pulse, a continuous-wave light, an amplitude modulated light, and a laser light. 6. The method as claimed in claim 1, wherein the at least one wavelength includes wavelengths between 700 nm and 1500 nm. 7. The method as claimed in claim 1, wherein the at least one wavelength of the illuminating light includes wavelengths between 750 nm and 950 nm, and is produced by the laser selected from one of a Ti:sapphire laser, a dye laser, a semiconductor laser, and a solid-state laser. 8. The method as claimed in claim 1 wherein the at least one wavelength of the illuminating light includes wavelengths between 950 nm and 1150 nm, and is produced by the laser selected from one of a Nd:YAG laser, a semiconductor laser, and a solid-state laser. 9. The method as claimed in claim 1 wherein at least one wavelength of the illuminating light includes wavelengths between 1150 nm and 1500 nm, and is produced by the laser selected from one of a Cr4+-based laser, a semiconductor laser, and a color-center laser. 10. The method as claimed i... | Research Foundation of City University of New York | A61B005/00
US07828740 | B2 | 2010-11-09 | Noninvasive apparatus and method for estimating blood pressure | 2006-08-01 | Methods and apparatus for noninvasively estimating a blood pressure are provided. A pulmonic (P) component extracted from a second heart sound (S2) signal is analyzed to obtain a number of oscillations in the P component. A predetermined relationship between the number of oscillations and PAP is used to generate a blood pressure estimate. A PCG sensor including a mechanical filter for receiving heart sounds from a chest wall is also provided. The pressure of the PCG sensor on the chest wall is adjustable. | 1. A method for noninvasively estimating a blood pressure, the method comprising the steps of: a) extracting a pulmonic (P) component from a second heart sound (S2) signal;b) analyzing the extracted P component to obtain a number of oscillations in the P component;c) determining a value that corresponds to the number of oscillations; andd) applying a predetermined relationship between the value and blood pressure to generate a blood pressure estimate. 2. The method according to claim 1, wherein the blood pressure includes a pulmonary artery pressure (PAP) and step (d) includes applying the predetermined relationship between the value and PAP to generate a PAP estimate. 3. The method according to claim 1, the method further including the steps of, prior to step (a): receiving a phonocardiogram (PCG) signal; andextracting the S2 signal from the received PCG signal,wherein step (a) extracts the P component from the extracted S2 signal. 4. The method according to claim 1, wherein the S2 signal is a time-domain S2 signal and step (a) extracts the P component using the time-domain S2 signal. 5. The method according to claim 1, wherein step (b) determines the number of oscillations using a derivative with respect to time of the P component. 6. The method according to claim 1, wherein the predetermined relationship is determined by a linear regression between a plurality of measured numbers of oscillations and a respective plurality of measured blood pressures corresponding to the measured numbers of oscillations. 7. The method according to claim 1, step (a) includes the steps of: determining a P component onset time according to a maximum absolute value of a time derivative of the S2 signal; andselecting a predetermined portion of the S2 signal as the P component using the determined P component onset time. 8. The method according to claim 1, wherein the S2 signal includes a plurality of S2 signals, and step (a) includes extracting a corresponding plurality of P components from the plurality of S2 signals,wherein the blood pressure estimate is generated according to the extracted plurality of P components. 9. The method according to claim 8, further including the step of determining a confidence interval measure associated with the blood pressure estimate. 10. The method according to claim 8, further including the steps of: repeating steps (b), (c) and (d) over the plurality of S2 signals to generate a corresponding plurality of blood pressure estimates; anddetermining a mode of the plurality of blood pressure estimates to generate an average blood pressure estimate. 11. The method according to claim 8, further including the steps of: determining a mean S2 signal from the plurality of S2 signals; andrepeating steps (b), (c) and (d) using the mean S2 signal to generate the blood pressure estimate. 12. The method according to claim 11, wherein the step of determining the mean S2 signal further includes the steps of: determining a plurality of d... | Luca Longhini | A61B005/02
US07837622 | B2 | 2010-11-23 | Recording medium and blood glucose monitoring system using the recording medium | 2002-08-06 | A recording medium includes a sensor circuit unit generating a current in accordance with a blood glucose level in blood, and a control unit calculating blood glucose level data from the data obtained by digital conversion of the generated current at signal processing unit, writing the calculated data to an EEPROM as a memory unit, reading the blood glucose level data from EEPROM and transmitting the same to a portable terminal on which recording medium is mounted, through a communication control unit. The recording medium further includes a power supply control unit receiving power from the portable terminal on which the recording medium is mounted, and supplying power to various portions in the recording medium. As a result, the recording medium can be utilized, when mounted on the portable terminal, a blood glucose monitoring apparatus, and the recording medium can be used in a system for providing bi-directional service between the portable terminal and a server managing the blood glucose information received from the portable terminal connected by a network. | 1. A recording medium operating mounted on a portable terminal capable of supplying power, comprising: a communication control unit exchanging data with said portable terminal;an inlet port for inserting a test paper configured for mounting said test paper on which taken blood is applieda sensor circuit unit generating a current from a blood glucose level of blood applied to said test paper mounted on said inlet port;a signal processing unit converting said generated current to a digital data and outputting the data;an electrically erasable and programmable non-volatile memory unit storing blood glucose level data converted based on the data output from said signal processing unit;a control unit converting the data output from said signal processing unit to said blood glucose level data, writing the converted data to said memory unit, reading said blood glucose level data from said memory unit and transmitting the data to said portable terminal through said communication control unit; anda power supply control unit receiving power from said portable terminal and supplying power to said sensor circuit unit, said signal processing unit, said memory unit, said communication control unit and said control unit; whereinsaid communication control unit, said inlet port, said sensor circuit unit, said signal processing unit, said memory unit, said control unit and said power supply control unit are integrally configured. 2. The recording medium according to claim 1, wherein said portable terminal transmits said measured blood glucose level data to a server managing blood glucose level of a customer, through a communication network. 3. The recording medium according to claim 1, wherein said portable terminal displays said measured blood glucose level data. 4. The recording medium according to claim 1, wherein said portable terminal is one of a portable telephone, a personal digital assistant and a portable personal computer. 5. A blood glucose monitoring system including a portable terminal, anda recording medium operating mounted on said portable terminal, whereinsaid portable terminal includesa first communication unit exchanging data with said recording medium,a power supply unit supplying power to said recording medium anda first control unit;said recording medium includesa second communication control unit exchanging data with said portable terminal,an inlet port for inserting a test paper configured for mounting said test paper on which taken blood is applied,a sensor circuit unit generating a current from blood glucose level of blood applied to said test paper mounted on said inlet port,a signal processing unit converting said generated current to digital data and outputting the digital data,an electrically erasable and programmable non-volatile memory unit storing blood glucose level data converted based on the data output from said signal processing unit,a second control unit converting the data output from said signal processing unit to said ... | Nipro Corporation | A61B005/00 G01N031/00 G05B021/00
US07848793 | B1 | 2010-12-07 | Monitoring for mitral valve regurgitation | 2006-09-29 | Implantable systems, and methods for use therein, for monitoring for mitral valve regurgitation (MR) are provided. An electrogram (EGM) signal and a corresponding pressure signal are obtained, where the EGM signal is representative of electrical functioning of the patient's heart during a plurality of cardiac cycles, and the corresponding pressure signal is representative of pressure within the left atrium the patient's heart during the cardiac cycles. Windows of the pressure signal are defined, based on events detected in the EGM signal, and measurements from the windows are used to monitor for MR. | 1. In an implantable system, a method for monitoring a patient's heart for mitral valve regurgitation (MR), comprising: (a) obtaining an electrogram (EGM) signal and a corresponding pressure signal, the EGM signal representative of electrical functioning of the patient's heart during a plurality of cardiac cycles, the corresponding pressure signal representative of pressure within the left atrium the patient's heart during the cardiac cycles;(b) for each of a plurality of cardiac cycles represented in the EGM and pressure signals, (b.1) determining a maximum peak within a first window of the pressure signal,wherein a start of the first window is defined relative to an event detected in the EGM signal, andwherein a length of the first window is defined to include at least one of an a-wave and a c-wave, but not a v-wave, of the cardiac cycle represented in the pressure signal; and(b.2) determining a maximum peak within a second window of the pressure signal,wherein a start of the second window is defined relative to an event detected in the EGM signal or relative to the first window, andwherein a length of the second window is defined to include the v-wave of the cardiac cycle represented in the pressure signal;(c) comparing the maximum peak detected within one or more said first window to the maximum peak detected within one or more said second window; and(d) monitoring for MR based on results of the comparing. 2. The method of claim 1, wherein: step (c) includes determining a ratio based on the maximum peak detected within one or more said first window and the maximum peak detected within one or more said second window, or vice versa; andstep (d) includes detecting MR by comparing the ratio to a threshold. 3. The method of claim 1, wherein: step (c) includes determining a difference between the maximum peak detected within one or more said first window and the maximum peak detected within one or more said second window; andstep (d) includes detecting MR when the difference exceeds a threshold. 4. An implantable system for monitoring a patient's heart for mitral valve regurgitation (MR), comprising: one or more electrode to obtain an electrogram (EGM) signal representative of electrical functioning of the patient's heart during a plurality of cardiac cycles;one or more sensor to obtain a corresponding pressure signal representative of pressure within the left atrium the patient's heart during the cardiac cycles; andone or more processor to determine, for each of a plurality of cardiac cycles represented in the EGM and pressure signals, a maximum peak within of a first window of the pressure signal, wherein a start of the first window is defined relative to an event detected in the EGM signal, and wherein a length of the first window is defined to include at least one of an a-wave and a c-wave, but not a v-wave, of the cardiac cycle represented in the pressure signal; anda maximum peak within a second window of the pressure signal, wherein a sta... | Pacesetter, Inc. | A61B005/04
US07857769 | B2 | 2010-12-28 | System and method for non-invasively determining a left ventricular end-diastolic pressure | 2005-06-24 | A system and method for determining a left ventricular end-diastolic pressure of blood within a left ventricle of a patient includes non-invasively measuring a plurality of cardiac parameters of the patient. The cardiac parameters may include a mitral valve area, and at least one of a stroke volume of the left ventricle, a left ventricular diastolic filling time, and a maximum velocity of blood entering the left ventricle. The systems and methods further include a step of calculating the left ventricular end-diastolic pressure using the measured cardiac parameters. | 1. A cardiac pressure monitoring system comprising: a sensing unit configured to non-invasively measure a plurality of cardiac parameters of a patient, the cardiac parameters including a mitral valve area, and at least one of a stroke volume of a left ventricle of the patient, a left ventricular diastolic filling time, and a change in velocity of blood entering the left ventricle over time; anda processor configured to receive the plurality of cardiac parameters and calculate a left ventricular pressure using the measured cardiac parameters. 2. The system according to claim 1 further comprising a memory unit configured to store at least one of the cardiac parameters. 3. The system according to claim 1 further comprising a display unit configured to display the left ventricular pressure. 4. The system according to claim 1, wherein: the left ventricular pressure is a left ventricular end-diastolic pressure; andthe processor calculates the left ventricular end-diastolic pressure by using a formula:L⁢⁢V⁢⁢E⁢⁢D⁢⁢P=36⁢ρ⁡(S⁢⁢VAmitral⁡(Δ⁢⁢tD))2,where LVEDP is the left ventricular end-diastolic pressure, ρ is a density of blood; SV is the stroke volume, Amitral is the mitral valve area, and ΔtD is the left ventricular diastolic filling time. 5. The system according to claim 1, wherein: the left ventricular filling time is measured; andthe plurality of cardiac parameters further includes a maximum velocity of the blood entering the left ventricle. 6. The system according to claim 5, wherein: the left ventricular pressure is a left ventricular end-diastolic pressure; andthe processor calculates the left ventricular end-diastolic pressure by using a formula:L⁢⁢V⁢⁢E⁢⁢D⁢⁢P=3⁢πρ⁡(S⁢⁢V)⁢vmaxAmitral⁢⁢Δ⁢⁢tD,where LVEDP is the left ventricular end-diastolic pressure, ρ is a density of blood, SV is the stroke volume, νmax is the maximum velocity of blood entering the left ventricle, Amitral is the mitral valve area, and ΔtD is the left ventricular diastolic filling time. 7. The system according to claim 1, wherein the processor calculates the left ventricular pressure by using a formula:L⁢⁢V⁢⁢P=ρ⁢⁢VAmitral⁢ⅆvⅆt+ρ⁢⁢v2,where LVP is the left ventricular pressure, ρ is a density of blood; V is a volume of the left ventricle, Amitral is the mitral valve area, ν is a velocity of blood entering the left ventricle, andⅆvⅆtis the change in velocity over time. 8. The system according to claim 1, wherein the sensing unit is a sonic sensor associated with an echocardiogram apparatus. 9. The system according to claim 1, wherein the sensing unit is a magnetic resonance imaging (MRI) apparatus. 10. The system according to claim 1, wherein the sensing unit is a computer aided tomography (CAT) scan apparatus. 11. In an electronic system, in which at least one computer monitors cardiac pressures within a heart of patient, a memory medium having software programmed to: receive a plurality of non-invasively measured cardiac parameters associated with a left ventricle of the pa... | - | A61B005/00
US07834884 | B2 | 2010-11-16 | Method and apparatus for displaying information | 2009-02-20 | A method for displaying large amounts of information. The method includes the steps of forming a spatial layout of tiles each corresponding to a representative reference element; mapping observed elements onto the spatial layout of tiles of representative reference elements; assigning a respective value to each respective tile of the spatial layout of the representative elements; and displaying an image of the spatial layout of tiles of representative elements. Each tile includes atomic attributes of representative elements. The invention also relates to an apparatus for displaying large amounts of information. The apparatus includes a tiler forming a spatial layout of tiles, each corresponding to a representative reference element; a comparator mapping observed elements onto said spatial layout of tiles of representative reference elements; an assigner assigning a respective value to each respective tile of said spatial layout of representative reference elements; and a display displaying an image of the spatial layout of tiles of representative reference elements. | 1. A method of diagnosing prostate cancer in a prostate tissue sample comprising: forming a spatial layout of tiles, each tile corresponding to a representative reference element in a mosaic;mapping observed elements of a prostate tissue sample onto the spatial layout of tiles of representative reference elements, the observed elements comprising gene expression levels in a prostate tissue sample;assigning a respective value to each respective tile of the spatial layout of representative reference elements; anddisplaying, using a graphics engine, an image of the spatial layout of tiles of representative reference elements,wherein each tile comprises atomic attributes of representative elements. 2. The method of claim 1, wherein the prostate tissue sample comprises a prostate tumor. 3. The method of claim 2, comprising classifying the stage of the prostate tumor based on the displayed image. 4. A method of diagnosing prostate cancer comprising: forming a spatial layout of tiles, each tile corresponding to a representative reference element in a mosaic;mapping observed elements onto the spatial layout of tiles of representative reference elements, the observed elements comprising gene expression levels in a prostate tissue sample;assigning a respective value to each respective tile of the spatial layout of representative reference elements; anddisplaying, using a graphics engine, an image of the spatial layout of tiles of representative reference elements,wherein each tile comprises a visual representation of the respective value of the number of observed elements. 5. The method claim 4 wherein forming the spatial layout of tiles, each tile corresponding to the representative reference element in the mosaic, comprises forming a self-organizing map. 6. The method of claim 5 wherein forming a self-organizing map comprises: a) forming representative reference elements;b) matching an observed element to a respective representative reference element to select a best unit matching centroid;c) adjusting the best unit matching centroid to form an adjusted best unit matching centroid; andd) iterating steps b and c a plurality of times for each observed element. 7. The method of claim 6 wherein matching the observed element to the respective representative reference element comprises measuring the Euclidean distance between the observed element and each representative reference element; and selecting the representative reference element having the smallest Euclidean distance. 8. The method of claim 6 wherein matching the observed element to the respective representative reference element comprises measuring a correlation between the observed element and each representative reference element; and selecting the representative reference element having the highest correlation. 9. The method of claim 6 wherein adjusting the best unit matching centroid comprises adjusting the value of the representative reference element to form an adjusted best unit match... | Children's Medical Center Corporation | G06T011/20 G09G005/22 G09G005/00 A01N001/00 G06F017/00 A61B005/00
US07828749 | B2 | 2010-11-09 | Blood and interstitial fluid sampling device | 2006-11-22 | A device and method for lancing a patient, virtually simultaneously producing and collecting a small fluid sample from a body. The device comprises a blood collection system including a lancing needle (16), drive mechanism (11), kneading or vibration mechanism (25), optional suction system (7), and sample ejection mechanism. The device is preferably sized to be hand-held in one hand and operable with one hand. The device can optionally contain integral testing or analysis component (83) for receiving the sample and providing testing or analysis indication or readout for the user. A method involves piercing the skin at a rapid rate, kneading the surrounding area by ultrasonic action, piezoelectric or mechanical oscillation to stimulate the blood flow from the wound, drawing the fluid using a pumping system. | 1. An integrated device, comprising: a main body having an opening where the main body is configured to be placed against tissue;a lancing member configured to extend through the opening of the main body to cut an incision in the tissue;a test strip disposed at the opening of the main body to directly collect and analyze the body fluid from the incision;the main body including a slot; andthe test strip being fitted in the slot in the main body. 2. The device of claim 1, further comprising: a stimulation ring received around the opening of the main body, the stimulator ring being configured to oscillate against the tissue for pumping the body fluid from the incision. 3. The device of claim 2, further comprising: a motor configured to oscillate the stimulation ring, wherein the motor is configured to oscillate the stimulation ring at a frequency from 5 to 1000 cycles per minute. 4. The device of claim 3, wherein the stimulation ring is configured to oscillate in a range of 1 to 5 mm. 5. The device of claim 2, wherein the stimulation ring and the main body have a slot in which the test strip is received. 6. The device of claim 5, further comprising: a secondary motion spring; anda secondary motion trigger configured to release the secondary motion spring to lower the test strip onto the body fluid. 7. The device of claim 1, wherein the test strip is configured to electrochemically analyze the sample. 8. An integrated device, comprising: a main body having an opening where the main body is configured to be placed against tissue;a lancing member configured to extend through the opening of the main body to cut an incision in the tissue;a test strip disposed at the opening of the main body to directly collect and analyze the body fluid from the incision; anda stimulation ring received around the opening of the main body, the stimulator ring being configured to oscillate against the tissue for pumping the body fluid from the incision. 9. The device of claim 8, further comprising: a motor configured to oscillate the stimulation ring, wherein the motor is configured to oscillate the stimulation ring at a frequency from 5 to 1000 cycles per minute. 10. The device of claim 9, wherein the stimulation ring is configured to oscillate in a range of 1 to 5 mm. 11. The device of claim 8, wherein the stimulation ring and the main body have a slot in which the test strip is received. 12. The device of claim 11, further comprising: a secondary motion spring; anda secondary motion trigger configured to release the secondary motion spring to lower the test strip onto the body fluid. 13. The device of claim 8, wherein the test strip is configured to electrochemically analyze the sample. | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00
US07835777 | B2 | 2010-11-16 | Device and method for determining analyte levels | 2009-12-22 | Devices and methods for determining analyte levels are described. The devices and methods allow for the implantation of analyte-monitoring devices, such as glucose monitoring devices, that result in the delivery of a dependable flow of blood to deliver sample to the implanted device. The devices comprise a unique microarchitectural arrangement in the sensor region that allows accurate data to be obtained over long periods of time. | 1. A system for processing sensor data from a continuous analyte sensor, comprising: a sensor comprising one or more electrodes configured for implantation in a subcutaneous tissue of a host, wherein the sensor is configured to continuously measure an analyte concentration in the host, whereby sensor data indicative of the subcutaneous analyte concentration is obtained; andan electronic system operably connected to the sensor and configured to calibrate the sensor data in real-time using a reference value determined by analysis of blood, wherein the electronic system is configured to calibrate the sensor data by correlating the reference value with a time-spaced sensor data value of the sensor data. 2. The system of claim 1, wherein the electronic system is configured to determine trends in analyte concentration. 3. The system of claim 1, wherein the system is configured to provide a substantially stable measurement of the analyte concentration in the host, wherein at least 95% of calibrated sensor data values are within 25% of corresponding values determined by analysis of blood over a useful life of the device. 4. The system of claim 1, wherein the analyte is glucose, and wherein the calibration is substantially linear at a glucose concentration of from about 40 mg/dL to about 500 mg/dL. 5. The system of claim 1, wherein the sensor is configured to continuously measure the analyte concentration in the host for at least about 3 days. 6. The system of claim 1, wherein the sensor is configured to continuously measure the analyte concentration in the host for at least about 5 days. 7. The system of claim 1, wherein the sensor is configured to continuously measure the analyte concentration in the host for at least about 7 days. 8. The system of claim 1, wherein the electronic system is physically coupled to the sensor. 9. The system of claim 1, wherein the electronic system is operably connected via radio telemetry to the sensor. 10. The system of claim 1, wherein an offset current defined by a current at 0 mg/dL is configured to remain stable during a period of time of continuous measurement of the analyte concentration in the host. 11. The system of claim 1, wherein the analyte is glucose. 12. A method for processing real-time sensor data from a continuous analyte sensor, comprising: calibrating, with an electronic system, data from a sensor configured to continuously measure a subcutaneous analyte concentration in a host, wherein calibrating comprises correlating a reference value determined by analysis of blood with a time-spaced sensor data value of the sensor data, and wherein the calibrating is performed during a period of time of continuous subcutaneous analyte concentration measurement. 13. The method of claim 12, wherein the electronic system is configured to determine trends in analyte concentration. 14. The method of claim 12, wherein the method is configured to provide a substantially stable measurement of the analyte conce... | DexCom, Inc. | A61B005/00
US08073533 | B2 | 2011-12-06 | Heart rate monitor with cross talk reduction | 2010-09-01 | A heart rate monitor determines a heart rate for a primary subject by reducing noise and cross-talk from unwanted signal transmission sources. A set of pulse peaks having about the same amplitude and an amplitude greater than any other set of peaks are processed to determine a subject's heart rate, while data having an amplitude below that of the set of peaks is ignored. To retrieve enough data to determine a peak amplitude, the signal having heart rate information is sampled at a frequency higher than a normal sampling frequency. In some embodiments, if no data is received for a set time period which is in or above the range, the threshold resets to zero under an assumption that the primary subject is no longer within transmitting range. | 1. A heart rate monitor, comprising: a receiver which receives wireless signals which include a plurality of pulses, the plurality of pulses include at least pulses from a transmitter worn by a first user which represent heart beats of the first user, and pulses from a transmitter worn by a second user which represent heart beats of the second user, where the pulses from the transmitter worn by the second user are unwanted cross-talk;at least one micro-controller associated with the receiver, the at least one micro-controller associates a peak amplitude with each of the plurality of pulses, determines an amplitude range based on at least one of the plurality of pulses, includes pulses of the plurality of pulses whose associated peak amplitude is within the amplitude range as being candidate heart beats of the first user, determines a heart rate of the first user based on one or more of the candidate heart beats, and excludes pulses of the plurality of pulses whose associated peak amplitude is below the amplitude range from being used to determine the heart rate of the first user; andat least one display device associated with the at least one micro-controller, the at least one display device displays the heart rate of the first user. 2. The heart rate monitor of claim 1, wherein: the at least one micro-controller determines the amplitude range based on an associated peak amplitude of a single pulse of the plurality of pulses. 3. The heart rate monitor of claim 1, wherein: the at least one micro-controller determines the amplitude range based on detecting matching associated peak amplitudes of a specified minimum number of multiple pulses of the plurality of pulses. 4. The heart rate monitor of claim 3, wherein: the specified minimum number of multiple pulses is user-adjustable. 5. The heart rate monitor of claim 1, wherein: the at least one micro-controller sets a new amplitude range when a specified minimum number of multiple new pulses are received which are associated with peak amplitudes which are above the amplitude range. 6. The heart rate monitor of claim 5, wherein: the at least one micro-controller considers the specified minimum number of multiple new pulses and subsequent pulses having associated peak amplitudes within the new amplitude range to be associated with a new user. 7. The heart rate monitor of claim 1, wherein: the at least one micro-controller determines the amplitude range by identifying a highest peak amplitude of the at least one of the plurality of pulses, sets a top of the amplitude range by adding a specified percentage of the highest peak amplitude to the highest peak amplitude, and sets a bottom of the amplitude range by subtracting a specified percentage of the highest peak amplitude from the highest peak amplitude. 8. The heart rate monitor of claim 7, wherein: the amplitude range is +/−3 to 5% of the highest peak amplitude. 9. The heart rate monitor of claim 1, wherein: the one or more of the candidate he... | Salutron, Inc. | A61B005/04
US08435447 | B2 | 2013-05-07 | Integrated meter for analyzing biological samples | 2010-07-26 | Analyte monitoring devices and methods therefore are provided. The devices integrate various functions of analyte monitoring, e.g., sample acquisition and testing. | 1. An analyte testing meter, comprising: a housing, the housing including an outer wall, the outer wall having an interior surface and an exterior surface;a cartridge containing a plurality of analyte testing structures, the cartridge held within an interior space within the housing; anda set of mechanical components including: a pusher that advances an analyte testing structure from the cartridge;a slot that holds the analyte testing structure when the analyte testing structure is advanced into a first end of the slot by the pusher; andan electrical connector that electrically couples the analyte testing structure to the meter when the analyte testing structure is held in the slot, the electrical connector enabling a measurement of an analyte level for a fluid sample in contact with the analyte testing structure held in the slot;wherein the slot permits the removal of the analyte testing structure out a second end of the slot for disposal from the meter after a measurement is performed; anda curved inner wall surface within the housing, wherein the pusher is adjacent the curved inner wall surface and contacting the curved inner wall surface such that the pusher is curved along the curved inner wall surface, and wherein the pusher is disposed between the curved inner wall surface and the interior surface of the outer wall when advancing the analyte testing structure. 2. The analyte test meter of claim 1, wherein the pusher comprises a flexible piece of material. 3. The analyte test meter of claim 1, wherein the pusher comprises a chain and an end piece for contacting the analyte testing structure. 4. The analyte test meter of claim 1, wherein the pusher extends along a dimension of the housing and bends around one or more corners of the housing when advancing the analyte testing structure. 5. The analyte test meter of claim 1, wherein the pusher winds and unwinds from a coil. 6. The analyte test meter of claim 1, wherein the pusher moves along a curved track within the housing. 7. The analyte test meter of claim 1, wherein the cartridge comprises at least one guide rail for relative positioning within the housing with respect to the set of components. 8. The analyte test meter of claim 1, further comprising: a seal which generally maintains the plurality of analyte testing structures within the cartridge free from exposure to ambient air, and wherein the analyte test meter is configured to release the seal temporarily to permit advancement of the analyte testing structure to the slot. 9. The analyte test meter of claim 8, wherein the seal is between the cartridge and a tub, and wherein the releasing of the seal comprises moving the tub away from the cartridge. 10. The analyte test meter of claim 1, further comprising: a biasing member at a loading end of the cartridge for biasing the plurality of analyte testing structures toward a dispensing end of the cartridge. 11. The analyte test meter of claim 1, further comprising: a mechanical c... | Abbott Diabetes Care Inc. | C12Q001/68 G01N033/48 G01N033/00 A61B005/00
US08523788 | B2 | 2013-09-03 | Device, system and method for examining a body lumen | 2010-08-05 | An in-vivo examining device, system and method for identifying the presence of strictures in the small bowel are provided. The in-vivo examining device includes a monitoring mechanism that becomes deactivated when exposed to in-vivo substances native to the small bowel or the colon, and a degradable device body that includes at least a first body portion which degrades at a slow rate when exposed to in-vivo substances native to the small bowel and at a fast rate when exposed to in-vivo substances native to the colon. The degradation of the degradable device body exposes the monitoring mechanism to substances native to the small bowel or the colon and thus indicates whether the examining device has safely passed through the small bowel or whether it is retained in the small bowel due to strictures in the small bowel. | 1. An in-vivo examining device to identify the presence of strictures in the small bowel comprising: a monitoring mechanism configured to be deactivated upon exposure to in-vivo substances native to the small bowel or the colon; anda degradable device body comprising at least a first body portion configured to degrade at a first rate when exposed to in-vivo substances native to the small bowel and at a second rate when exposed to in-vivo substances native to the colon, said first rate being slower than said second rate,wherein degradation of said degradable device body exposes said monitoring mechanism to substances native to the small bowel or the colon. 2. The device according to claim 1, further comprising a non-degradable or slowly-degradable coating surrounding said degradable device body, wherein said coating comprises at least one opening. 3. The device according to claim 2, wherein said coating includes parylene C. 4. The device according to claim 1, wherein the monitoring mechanism is a radio-frequency identification (RFID) tag. 5. The device according to claim 1, wherein the monitoring mechanism is composed of a material selected from the group consisting of: paper, degradable plastic and degradable rubber. 6. The device according to according to claim 1, wherein said RFID tag is rolled around itself within the device body. 7. The device according to according to claim 1, wherein said RFID tag is wrapped around the device body and covered by coating. 8. The device according to according to claim 2, wherein said device body comprises filler disposed within said coating, the filler being dissolvable in in-vivo fluids. 9. The device according to according to claim 1, wherein said device body further comprises a second body portion which is configured to degrade at a third rate when exposed to in-vivo substances native to the small bowel and at a fourth rate when exposed to in-vivo substances native to the colon, said first rate being slower than said second rate, wherein said first and second portions are positioned at opposite sides along the longitudinal axis of said device. 10. An in-vivo examining device to identify the presence of strictures in the small bowel comprising: a monitoring mechanism configured to be deactivated upon exposure to in-vivo substances native to the small bowel or the colon; anda degradable device body comprising a first body portion configured to degrade at a first rate when exposed to in-vivo substances native to the small bowel and a second body portion which degrades at a second rate when exposed to in-vivo substances native to the colon, said first rate being slower than said second rate,wherein degradation of said degradable device body exposes said monitoring mechanism to substances native to the small bowel or the colon. 11. The device according to claim 10, wherein said first body portion and said second body portions are plugs positioned at opposite sides along the longitudinal axis of said d... | Given Imaging Ltd. | A61B005/117 A61B005/103
US08430095 | B2 | 2013-04-30 | Bite block | 2009-12-10 | A bite block includes: a first wall defining a hole into which a conduit pipe is to be inserted; a second wall surrounding the first wall to define a flow path with the first wall; and a sample port communicating with the flow path. | 1. A bite block comprising: a first wall defining a hole into which a conduit pipe is to be inserted;a second wall surrounding the first wall to define a flow path with the first wall; anda sample port communicating with the flow path, whereinthe flow path includes a slope circumferentially formed between the first wall and the second wall and extending from the sample port while sloping downward toward peripheral edges of the first and second walls. 2. The bite block according to claim 1, further comprising: an attaching portion to which a respiration information collecting adaptor is to be attached. 3. The bite block according to claim 1, wherein the flow path includes a slope connected from the sample port to peripheral edges of the first and second walls. 4. The bite block according to claim 1, further comprising: an adjustor provided at the sample port and coupled to a respiration information collecting adaptor, the adjustor configured to adjust an attitude of the respiration information collecting adaptor and a distance between the sample port and the respiration information collecting adaptor. 5. The bite block according to claim 1, wherein the sample port includes a divider partitioning an interior into a first chamber communicating with the flow path and a second chamber communicating with an outside. 6. The bite block according to claim 1, further comprising: an attaching portion; anda respiration information collecting adaptor attached to the attaching portion and configured to collect respiration information. 7. The bite block according to claim 6, wherein the respiration information collecting adaptor includes a holder holding a prong which supplies oxygen, andwhen the bite block is attached to a subject, a tip end of the prong is positioned near a nasal cavity or oral cavity of the subject. 8. The bite block according to claim 1, wherein the first wall has a cylindrical shape,the conduit pipe is connected to an endoscope or a rigid scope,the flow path is defined by a space between the first wall and the second wall,and a gas in the oral cavity is directed to the sample port through the flow path. 9. The bite block according to claim 1, wherein an end of the first wall to be opposed to an oral cavity of a subject is in plane with an end of the second wall to be opposed to the oral cavity of the subject. 10. The bite block according to claim 1, wherein the hole defined by the first wall is to be communicated with an outside and an oral cavity of a subject,the flow path defined by the first wall and the second wall is to be communicated with the sample port and the oral cavity of the subject,and the hole and the flow path are separated with each other. | Nihon Kohden Corporation | A61M016/00
US09259290 | B2 | 2016-02-16 | MRI-guided surgical systems with proximity alerts | 2010-06-08 | Systems to facilitate MRI-guided procedures are configured to generate proximity alerts for an MRI-guided procedure associated with at least one target site and/or at least one avoid zone. | 1. A system to facilitate MRI-guided procedures, comprising: a circuit comprising at least one processor and/or at least one Application Specific Integrated Circuit (ASIC);an intrabody device comprising a tip with a deflectable distal end portion which has variable curvature, wherein the intrabody device further comprises a plurality of spaced apart tracking coils residing on the deflectable distal end portion; anda display in communication with the circuit, the display having a User Interface (UI),wherein the circuit is configured to provide at least one volumetric planning map of relevant patient anatomy to the display, wherein the UI is configured to allow a user to select at least one target site and at least one avoid zone on the at least one volumetric planning map,wherein the circuit is further configured to define proximity alert locations for an MRI-guided procedure associated with the user selected at least one target site and/or the user selected at least one avoid zone, track the intrabody device by calculating a tip location and orientation of the tracked intrabody device using known physical data regarding the tracking coils and a predictable deflectable behavior of the deflectable distal end portion, and generate at least one audible and/or visual alert during the MRI-guided procedure when the tracked intrabody device is identified as approaching or being in one of the defined proximity alert locations, wherein the predictable deflectable behavior is based on forces applied to the intrabody device;wherein the circuit is further configured to register the at least one volumetric planning map to MRI 3-D image space associated with the MRI-guided procedure, determine preset scan planes that cover the at least one selected target site and/or avoid zone based on the user selected at least one target site and/or the user selected at least one avoid zone, and wherein, during the MRI-guided procedure, the circuit is configured to determine a location of the tracked intrabody device and automatically select a relevant preset scan plane from one of the determined preset scan planes to obtain an MRI signal. 2. The system of claim 1, wherein the circuit is configured to define the proximity alert locations for the selected at least one target site and the selected at least one avoid zone, wherein the circuit is configured to define boundary limits for the at least one target site for the at least one audible and/or visual alerts, and wherein the boundary limits include boundary limits of different ablation sites for ablation by the tracked intrabody device. 3. The system of claim 2, wherein the UI is configured to carry out at least the following: (a) allow the user to select the boundary limits for the at least one target site by marking locations on the at least one volumetric planning map or by selecting dimensional spacing about a treatment site; (b) provide defined boundary limits for the selected at least one target site based on pre-d... | MRI Interventions, Inc. | A61B005/055 A61B019/00 A61B005/00 A61B018/14 A61B017/00 A61B018/00
US20100305438 | A1 | 2010-12-02 | SYSTEM AND METHOD FOR SCALING STRAIN IMAGE DATA | 2009-05-29 | A computer readable storage medium has a computer program stored thereon, which includes a set of instructions that when executed by a computer causes the computer to access positional data acquired from a material at a plurality of states of compression. The set of instructions also causes the computer to estimate a plurality of strain data sets from the positional data, each of the plurality of strain data sets corresponding to a different state of compression, and compare a first pair of strain data sets of the plurality of strain data sets with each other. The set of set of instructions further causes the computer to calculate a first measure of correlation from the comparison, scale a first strain data set of the first pair of strain data sets if the first measure of correlation is above a correlation threshold, and display the scaled first strain data set to a user. | 1. A computer readable storage medium having a computer program stored thereon, the computer program comprising a set of instructions that when executed by a computer causes the computer to: access positional data acquired from a material at a plurality of states of compression;estimate a plurality of strain data sets from the positional data, each of the plurality of strain data sets corresponding to a different state of compression;compare a first pair of strain data sets of the plurality of strain data sets with each other;calculate a first measure of correlation from the comparison;scale a first strain data set of the first pair of strain data sets if the first measure of correlation is above a correlation threshold; anddisplay the scaled first strain data set to a user. 2. The computer readable storage medium of claim 1 wherein the set of instructions further cause the computer to identify the first strain data set as a last-scaled data set. 3. The computer readable storage medium of claim 1 wherein the set of instructions further cause the computer to scale the first strain data set to a second strain data set of the first pair of strain data sets using a correlation algorithm if the second strain data set is a scaled data set, wherein the correlation algorithm causes the computer to: generate a pairwise association of data from the first and second strain data sets;generate an approximate linear fit to the pairwise association of data;determine a slope and an offset from the linear fit; andscale the first strain data set based on the slope and the offset. 4. The computer readable storage medium of claim 3 wherein the set of instructions that cause the computer to compare a first pair of strain data sets cause the computer to compare a first pair of strain data sets corresponding to a pair of consecutively acquired frames of positional data. 5. The computer readable storage medium of claim 1 wherein the set of instructions further cause the computer to: scale the first strain data set using a histogram algorithm to produce a histogram-scaled first data set if the second strain data set is not scaled, wherein the histogram algorithm causes the computer to: generate a cumulative distribution histogram from the first strain data set;select a subset of the cumulative distribution histogram;select a subset of a display range; andscale the first strain data set based on an upper bound and a lower bound of the subset of the cumulative distribution histogram and an upper bound and a lower bound of the subset of the display range; andscale the second strain data set to the histogram-scaled first strain data set using a correlation algorithm, wherein the correlation algorithm causes the computer to: generate a pairwise association of data from the first and second strain data sets;generate an approximate linear fit to the pairwise association of data;determine a slope and an offset from the linear fit; andscale the second strain data set based on... | - | A61B008/00 A61B005/103
US20100286558 | A1 | 2010-11-11 | FLUID COLLECTION/INJECTION DEVICE HAVING QUICK RELEASE/REMOVABLE DOUBLE-ENDED NEEDLE AND SAFETY SYSTEM | 2010-04-07 | Fluid collection/injection device includes a body having a front end, a back end, and a main hollow section arranged between the front and back ends. A needle hub securing section is arranged on the front end and is structured and arranged to receive therein a needle member. The needle hub securing section includes a fixed part and a movable part. The fixed part is integrally formed with the front end and the movable part is arranged on a member that has one end which is one of fixed to a portion of the main hollow section, connected to a portion of the main hollow section via a living hinge, removably connected to a portion of the main hollow section, and integrally formed with the main hollow section. | 1. A fluid collection/injection device comprising: a body having a front end, a back end, and a main hollow section arranged between the front and back ends;a needle hub securing section arranged on the front end and being structured and arranged to receive therein a needle member; andthe needle hub securing section comprises a fixed part and a movable part,wherein the fixed part is integrally formed with the front end and the movable part is arranged on a member that has one end which is one of: fixed to a portion of the main hollow section;connected to a portion of the main hollow section via a living hinge;removably connected to a portion of the main hollow section; andintegrally formed with the main hollow section. 2. The device of claim 1, wherein the body is one of generally cylindrical and generally tubular. 3. The device of claim 1, wherein the fixed part and the movable part form a generally cylindrical section which the movable part is in an initial position. 4. The device of claim 1, wherein the fixed part and the movable part each generally comprise one-half of an internal locking thread structured and arranged to engage with an external thread of the needle member. 5. The device of claim 1, wherein the member has one end which is fixed to the portion of the main hollow section. 6. The device of claim 1, wherein the member has one end which is removably connected to the portion of the main hollow section. 7. The device of claim 1, wherein the member has one end which is integrally formed with the main hollow section. 8. The device of claim 1, wherein the member has one end which is connected to the portion of the main hollow section via the living hinge. 9. The device of claim 1, wherein the body and the needle hub securing section comprise a one-piece member. 10. The device of claim 1, wherein the device is a one-piece member. 11. The device of claim 1, further comprising guide projections arranged on the front end of the body. 12. The device of claim 1, wherein the member comprises a projection for facilitating removal of the needle member. 13. The device of claim 1, wherein the member comprises an opening sized to receive therein the fixed part. 14. The device of claim 1, wherein the member comprises a finger engaging section which is arranged on a side of the fixed part opposite the movable part. 15. The device of claim 1, wherein the member comprises a finger engaging section which is movable towards the body when it experiences a force and which moves back to an initial position away from the body when the force is removed. 16. The device of claim 1, wherein the member comprises a finger engaging section which is movable towards the body when it experiences a force and which automatically moves back to an initial position away from the body when the force is removed. 17. The device of claim 1, wherein the front end comprises a through opening sized to receive an inner needle end of the needle member. 18. The dev... | STAT MEDICAL DEVICES, INC. | A61M005/34 A61M005/50 A61B005/15
US20100332445 | A1 | 2010-12-30 | ANALYTE TESTING METHOD AND SYSTEM | 2010-06-29 | A system and method of detecting a flagged glucose concentration pattern with the use of medians having a common type of flag collected over discrete time periods so that whenever significant differences between the medians arise, the user or a caretaker of a diabetic user is notified. | 1. A diabetes management system comprising: a plurality of glucose test strips, each test strip configured to receive a physiological sample;a test strip port connector configured to receive the plurality of test strips; anda diabetes data management device comprising: a housing;a microprocessor coupled to a memory, display, and power supply disposed proximate the housing, the microprocessor coupled to the test strip sensor to provide data representative of a first group and second group of blood glucose values of the user over respective first and second time periods so that respective first and second medians of the first and second group are evaluated by the microprocessor to determine whether one of the first and second medians is significantly different enough to inform the user of the same on the display of the device. 2. The diabetes management system of claim 1, in which the first and second medians are calculated by the microprocessor with glucose values including a common type of flag. 3. The diabetes management system of claim 2, in which the common type of flag comprises at least one of a fasting flag or a bedtime flag. 4. The diabetes management system of claim 1, in which the diabetes data management device comprises a blood glucose meter. 5. The diabetes management system of claim 1, in which the diabetes data management device comprises a combination of a blood glucose meter and mobile phone electronically coupled to each other and in which the blood glucose meter includes the test strip port and a microprocessor to provide blood glucose data to a microprocessor of the mobile phone. 6. A method of detecting a fasting glucose concentration pattern with an analyte testing device having a microprocessor coupled to a memory, the method comprising: obtaining from the memory of the analyte testing device a first group and second group of glucose measurements over a first time period and a second time period, respectively;determining whether the fasting glucose concentrations of the first group are significantly different than the fasting glucose concentrations of the second group;calculating a first median and a second median of the fasting glucose measurements over a first time period and a second time period, respectively;displaying a message indicating that the second group has a significantly higher fasting glucose concentration than the first group where the second median is greater than the first median, and the first group and second group are significantly different; anddisplaying a message indicating that the second group has a significantly lower fasting glucose concentration than the first group where the second median is less than the first median, and the first group and second group are significantly different. 7. A method of detecting a fasting glucose concentration pattern for a day of the week with an analyte testing device having a microprocessor coupled to a memory, the method comprising: obtaining from the memo... | LifeScan, Inc. | G06N005/02 A61B005/15 G06F007/08
US20100298724 | A1 | 2010-11-25 | RECYCLABLE OR BIODEGRADABLE BLOOD PRESSURE CUFF | 2009-05-19 | A blood pressure cuff is defined by a sleeve member, a portion of which is inflatable and made from a fluid impermeable material. An exterior port of the sleeve directly and fluidly interconnects with the interior of said inflatable portion, wherein a slotted portion is aligned with the port to receive the port when the cuff is wrapped about the limb of a patient. The slotted portion is sized to provide a range of cuff sizes wherein the cuff can be entirely made from a biodegradable or recyclable material. A plurality of cuffs can be packaged for sequential dispensing thereof. | 1. A disposable blood pressure cuff, said cuff comprising: a sleeve having an interior; andan inflatable portion within at least a portion of said sleeve interior, said cuff including a port that directly and fluidly interconnects with the interior of said inflatable portion, said cuff being entirely made from a single material that enables recyclability of said cuff. 2. A cuff as recited in claim 1, wherein said sleeve is made from a cellulosic material. 3. A cuff as recited in claim 2, wherein said sleeve is made from polypropylene. 4. A cuff as recited in claim 1, wherein said sleeve is made from polyethylene. 5. A cuff as recited in claim 1, said sleeve further including a slotted portion on a side opposite said port. 6. A cuff as recited in claim 5, wherein said sleeve is adapted to be attached onto the limb of a patient in overlapping fashion wherein said port is fitted in said slotted position. 7. A cuff as recited in claim 6, wherein said slotted portion is sized to limit the limb sizes that said cuff can be attached to said patient. 8. A cuff as recited in claim 1, wherein said port is a socket. 9. A cuff as recited in claim 8, wherein said socket is configured to receive at least one of a gage housing and a hose adapter. 10. A cuff as recited in claim 9, wherein said socket is configured to permit said at least one of a gage housing and a hose adapter engaged in said socket to be rotatable about an axis of said socket. 11. A cuff as recited in claim 1, wherein said cuff includes means for preventing said cuff from being reattached to a patient when said cuff is removed therefrom. 12. A cuff as recited in claim 11, wherein said cuff includes a portion that is destructively removed when said cuff is removed from said patient. 13. A cuff as recited in claim 1, wherein said cuff is prevented from being reattached to a patient after a predetermined number of uses. 14. A cuff as recited in claim 1, wherein said sleeve includes closure means for releasably securing said sleeve in a wrapped configuration. 15. A cuff as recited in claim 14, wherein said closure means includes hook and loop fasteners and in which the exterior surface of said sleeve is made from a material that provides adhesion to a hook fastener portion of said sleeve, thereby serving as a loop fastener section. 16. A cuff as recited in claim 13, wherein the adhesion of said sleeve material is degraded upon attachment and subsequent removal from said hook fastener portion. 17. A cuff as recited in claim 5, wherein said slotted portion and said port are aligned in substantially the center of the width of said sleeve. 18. A cuff as recited in claim 15, wherein said port is used as an arterial marker. 19. A cuff as recited in claim 16, wherein said sleeve can be alternatively wrapped in at least two orientations about a patient limb. 20. A cuff as recited in claim 1, wherein said port is made from the same material as said sleeve. 21. A disposable blood pressure... | Welch Allyn, Inc. | A61B005/02 B65D083/08
US20100317967 | A1 | 2010-12-16 | COMPUTER ASSISTED THERAPY MONITORING | 2007-12-05 | A computer assisted therapy apparatus (100) uses data from functional medical imaging examinations of a patient to evaluate the response of a patient to an applied therapy. A lesion tracker (112) tracks lesion(s) detected in the medical imaging examinations, and a lesion quantifier (114) generates quantitative information indicative of functional characteristic of the lesion(s). A trend analyzer (116) uses the quantitative information to determine trends in the functional characteristic. | 1. An apparatus comprising: a lesion detector which detects lesions in medial image data from an imaging examination of a patient;a lesion quantifier in operative communication with the lesion detector, wherein the lesion quantifier uses first functional image data from a first imaging examination of the patient conducted before the application of a therapy to the patient to generate first lesion functional data for a first detected lesion, and wherein the lesion quantifier uses second functional image data from a second imaging examination of the patient conducted after the application the therapy to generate second lesion functional data for the first detected lesion;a trend analyzer which identifies a difference between the first and second lesion functional data. 2. The apparatus of claim 1 further including a lesion tracker which determines a correspondence between a lesion detected in image data from the first imaging examination and a lesion detected in image data from the second imaging examination. 3. The apparatus of claim 2 including a motion model, wherein the lesion tracker uses the motion model to determine the correspondence. 4. (canceled) 5. The apparatus of claim 1 further including a physiology model, wherein the lesion quantifier uses the physiology model to model an expected behavior of a tracer applied to the patient in connection with the first functional medical imaging examination. 6. (canceled) 7. The apparatus of claim 1 wherein the first imaging examination is conducted according to a first protocol and the second imaging is conducted according to a second protocol and the lesion quantifier applies a correction which reduces an effect of a variation between the first and second protocols. 8. The apparatus of claim 1 wherein the lesion quantifier applies a patient specific morphological correction or specific physiological correction. 9. (canceled) 10. The apparatus of claim 1 further including a disease model, wherein the trend analyzer uses the disease model to model a response of the first detected lesion to the applied treatment. 11. The apparatus of claim 1 further including a PET/CT scanner. 12. The apparatus of claim 1 wherein the lesion is a tumor and applied therapy includes at least one of an external radiotherapy, a chemotherapy, a brachytherapy, an ablation therapy, and a molecular therapy. 13. The apparatus of claim 1 wherein the lesion functional data includes a hypoxia, cell proliferation, or standard uptake value. 14. The apparatus of claim 1 wherein the lesion quantifier uses third functional image data generated from a third functional imaging examination of the patient conducted after the application of a second therapy to the patient to generate third lesion functional data for the first detected lesion. 15. The apparatus of claim 1 wherein the lesion quantifier uses the first functional image data to generate first lesion functional data for a second detected lesion, the lesion quantifi... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/05
US20100280332 | A1 | 2010-11-04 | Methods and systems for monitoring bioactive agent use | 2008-07-16 | Methods, computer program products, and systems are described that include accepting at least one indication of a bioactive agent use by an individual and/or assigning an artificial sensory experience to monitor at least one side effect of the bioactive agent on the individual. | 1-74. (canceled) 75. A system, comprising: an acceptor module;an assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual. 76. The system of claim 75, wherein the acceptor module comprises: a bioactive agent identification acceptor module. 77. The system of claim 75, wherein the acceptor module comprises: a substance abuse indication acceptor module. 78. The system of claim 75, wherein the acceptor module comprises: a nutraceutical identification acceptor module. 79. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: a sensate experience assigner module. 80. The system of claim 79, wherein the sensate experience assigner module comprises: a stimulus assigner module. 81. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: an automated medical device receiver module. 82. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: an assigner module configured to assign a virtual world, a modification to a virtual world, a computer game, a modification to a computer game, a website, a modification to a website, an online course, or a modification to an online course. 83. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: a monitorer module. 84. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: a neurophysiological measurement monitorer module. 85. The system of claim 84, wherein the neurophysiological measurement monitorer module comprises: a physiologic activity measurer module. 86. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: a brain activity surrogate marker measurer module. 87. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: a near real time measurer module. 88. The system of claim 75, wherein the assigner module configured to assign an artificial sensory experience to monitor at least one side effect of a bioactive agent on an individual comprises: a test function output measurer module. 89. The system of claim 75, wherein the assigner module configured to assign an artificial sensor... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00
US20100312070 | A1 | 2010-12-09 | Artificial Intelligence and Device for Diagnosis, Screening, Prevention and Treatment of Materno-Fetal Conditions | 2004-12-02 | The present invention relates to a time-oriented artificial intelligence system to handle any diagnostic screening or treatment of complications or risks throughout pregnancy. A user can insert a problem or query relating to clinical case management during a pregnancy and receive case oriented output guiding the management of the case via at least one algorithm. | 1. An expert system for optimizing health during pregnancy comprising: at least one database of pregnancy related health data including data representing time oriented information about pregnancy health complications;at least one input for inputting diagnostic and screening data, including time oriented information about said diagnostic and screening data; andat least one indicator for reporting a decision as a function of the inputted diagnostic and screening data and said pregnancy related heath data. 2. The expert system of claim 1, the system comprising: a plurality of time oriented data menus, said data menus comprising categorically defined pregnancy related health conditions, said data menus being organized as a function of the pregnancy time period, said health complications being classified in said data menus. 3. The expert system of claim 2, wherein the categorically defined pregnancy conditions comprise: inherited genetic abnormalities;fetal structural anomalies without detectable abnormal genetic pattern;idiopathic fetal malformations and diseases;fetal disease from pathology of physiology of mother;teratogenetic or other exposures; orsporadic genetic mutations. 4. The expert system of claim 1 wherein said report of time oriented information comprises: data representing the earliest known time for detection of the health complication; ordata representing a course of action regarding the complication;data representing the likelihood of such complication; ordata representing the class of a complication. 5. The expert system of claim 1, wherein the inputted diagnostic and screening data includes data from: a diagnostic and screening tool;a diagnostic and screening test; orinformation garnered from a person reporting on a pregnancy; ora report on the patient. 6. The expert system of claim 5, wherein the diagnostic and screening tool comprises: an ultrasound pattern recognition device;a genetic testing device;a genetic counseling system;a device for biochemical testing; ora magnetic resonance device. 7. The expert system of claim 5, wherein the diagnostic and screening tests comprises: a genetic test;an ultrasound test; ora biochemical test. 8. The expert system of claim 5, wherein the information garnered comprises: information from a patient interview;information provided by someone other than patient; orinformation volunteered by a patient. 9. The expert system of claim 1, said system comprising an intelligent agent further comprising at least one algorithmic rule adapted to apply to data inputted into the intelligent agent, said rule designed to produce at least one decision about a pregnancy case. 10. The expert system of claim 9, wherein said decision comprises: scheduling at least one action to be taken with respect to the complication, said action including an action for screening for at least one said complication;treating said complication. 11. The expert system of claim 5, wherein the report on the patient comprises: a... | - | A61B005/00 G06N005/02
US07853309 | B2 | 2010-12-14 | X-ray CT apparatus and myocardial perfusion image generating system | 2005-03-10 | An X-ray CT apparatus comprises a blood-flow information acquisition unit, a correction value calculating unit and a blood-flow image generating unit. The blood-flow information acquisition unit obtains information of a relative blood flow rate in the myocardium of the subject based on the CT image. The correction value calculating unit obtains a correction value based on the CT image in a concentration transition period defined to be a period from immediately after start of a continuous injection of a contrast medium into the subject until the contrast medium reaching the myocardium increases and is be in a state where it can be considered that the contrast medium is saturated at a constant value. The blood-flow image generating unit generates a blood flow value image in the myocardium by correcting the information of the relative blood flow rate with the correction value. | 1. An X-ray CT apparatus, comprising: an X-ray tube to radiate an X-ray to a subject so as to scan the subject;an X-ray detector to detect the radiated X-ray;a reconstruction unit configured to reconstruct a CT image of the subject based on obtained projection data;a controlling unit configured to control an injector so that the injector continuously injects a contrast medium into the subject and to detect a state in which a concentration of the contrast medium in a myocardial of the subject can be considered to be constant;a blood-flow information obtaining unit configured to obtain information of a relative blood flow rate in the myocardium of the subject based on the CT image;a correction value calculating unit configured to obtain a correction value based on a CT image obtained in a concentration transition period, which is defined to be a period from immediately after a start of a continuous injection of the contrast medium into the subject until the contrast medium reaching the myocardium increases and is in a state in which the contrast medium is saturated at a constant value, as detected by the controlling unit, the correction value calculating unit further configured to obtain the correction value by calculating a rate of change of the concentration of the contrast medium, which is obtained based on multiple CT images obtained in the concentration transition period; anda blood-flow image generating unit configured to generate a blood flow value image representing a blood flow rate in the myocardium by correcting the information of the relative blood flow rate with the correction value. 2. An X-ray CT apparatus according to claim 1, wherein the blood-flow information acquisition unit is configured to obtain the information of the relative blood flow rate by performing subtraction between a value based on the CT value of the myocardium and a CT image in a state in which the concentration of the contrast medium in a myocardial portion can be considered to be constant. 3. An X-ray CT apparatus according to claim 1, further comprising means for collecting projection data regarding part of the myocardium in the concentration transition period, and for collecting projection data regarding other portions of the myocardium following the concentration transition period. 4. An X-ray CT apparatus according to claim 1, wherein the correction value calculating unit is configured to obtain the correction value based on the projection data in part of the myocardium, and the blood-flow image generating unit is configured to correct the information of the relative blood flow rate in other portions of the myocardium with the correction value. 5. An X-ray CT apparatus, comprising: an X-ray tube to radiate an X-ray to a subject so as to scan the subject;an X-ray detector to detect the radiated X-ray;a reconstruction unit configured to reconstruct a CT image of the subject based on obtained projection data;a controlling unit configured to determine a cons... | Toshiba Medical Systems Corporation The Johns Hopkins University | A61B005/05
US20100312148 | A1 | 2010-12-09 | POSTURE EVALUATING APPARATUS | 2009-03-09 | A posture evaluating apparatus according to the invention includes: a grasp portion that is grasped with user's hands; a pedestal portion on which the user rides with both feet; inclination detecting means for detecting inclination of the grasp portion; barycentric position detecting means for detecting a barycentric position of a load acting on the pedestal portion; posture evaluating means for evaluating posture of the user based on the inclination of the grasp portion and the barycentric position of the load acting on the pedestal portion; and output means for outputting evaluation result of the posture evaluating means. | 1. A posture evaluating apparatus comprising: a grasp portion that is grasped with user's hands;a pedestal portion on which the user rides with both feet;inclination detecting means for detecting inclination of the grasp portion;barycentric position detecting means for detecting a barycentric position of a load acting on the pedestal portion;posture evaluating means for evaluating posture of the user based on the inclination of the grasp portion and the barycentric position of the load acting on the pedestal portion; andoutput means for outputting evaluation result of the posture evaluating means. 2. The posture evaluating apparatus according to claim 1, wherein the barycentric position detecting means is a load sensor that is provided in the pedestal portion. 3. The posture evaluating apparatus according to claim 2, wherein the load sensor also acts as weight measuring means for measuring a weight of the user. 4. The posture evaluating apparatus as in claim 1, wherein the inclination detecting means is an inclination sensor that is provided in the grasp portion. 5. The posture evaluating apparatus as in claim 1, wherein the posture evaluating means evaluates strain of a body based on the inclination of the grasp portion and the barycentric position of the load acting on the pedestal portion. 6. The posture evaluating apparatus as in claim 1, wherein the output means outputs advice to have good posture according to the evaluation result of the posture evaluating means. 7. The posture evaluating apparatus as in claim 1, further comprising body composition estimating means for measuring impedance of the user to estimate a body composition from the impedance. 8. The posture evaluating apparatus as in claim 2, wherein the inclination detecting means is an inclination sensor that is provided in the grasp portion. 9. The posture evaluating apparatus as in claim 3, wherein the inclination detecting means is an inclination sensor that is provided in the grasp portion. 10. The posture evaluating apparatus as in claim 2, wherein the posture evaluating means evaluates strain of a body based on the inclination of the grasp portion and the barycentric position of the load acting on the pedestal portion. 11. The posture evaluating apparatus as in claim 3, wherein the posture evaluating means evaluates strain of a body based on the inclination of the grasp portion and the barycentric position of the load acting on the pedestal portion. 12. The posture evaluating apparatus as in claim 2, wherein the output means outputs advice to have good posture according to the evaluation result of the posture evaluating means. 13. The posture evaluating apparatus as in claim 3, wherein the output means outputs advice to have good posture according to the evaluation result of the posture evaluating means. 14. The posture evaluating apparatus as in claim 2, further comprising body composition estimating means for measuring impedance of the user to estimate a body ... | Omron Healthcare Co., Ltd. | A61B005/103
US07844323 | B2 | 2010-11-30 | Means and method for the detection of cardiac events | 2007-06-22 | Disclosed is a system for the detection of cardiac events that includes an implanted device called a cardiosaver, a physician's programmer and an external alarm system. The system is designed to provide early detection of cardiac events such as acute myocardial infarction or exercise induced myocardial ischemia caused by an increased heart rate or exertion. The system can also alert the patient with a less urgent alarm if a heart arrhythmia is detected. Using different algorithms, the cardiosaver can detect a change in the patient's electrogram that is indicative of a cardiac event within five minutes after it occurs and then automatically warn the patient that the event is occurring. To provide this warning, the system includes an internal alarm sub-system (internal alarm means) within the cardiosaver and/or an external alarm system (external alarm means) which are activated after the ST segment of the electrogram exceeds a preset threshold. | 1. A system for identifying an ischemic event in a human patient including: at least two electrodes implanted in the patient for obtaining an electrical signal from the patient's heart; an implanted device including: (a) analog-to-digital converter circuitry for digitizing the electrical signal to produce a multiplicity of segments, each having a duration of least 1 second, each of the periods between successive segments being at least as long as the segment duration;(b) memory means designed to store a first segment at a first predetermined time;(c) memory means designed to store a recently collected segment at a second time that is later than the first predetermined time;(d) processor means coupled to said memory means for comparing the ST segment voltage of at least one beat of the recently collected segment at the second time with the average ST segment voltage of at least two beats of the first segment stored in the memory means at the first predetermined time; and,(e) means for identifying that the ischemic event has occurred by applying a test that is based on whether the difference between the ST segment voltage of the at least one beat of the recently collected segment and the average ST segment voltage of the at least two beats of the first segment exceeds a detection threshold, wherein the detection threshold is a preset percentage of the average signal amplitude of the at least two beats of the first segment. 2. The system of claim 1 where the average signal amplitude of the at least two beats of the first segment is the R wave peak voltage minus the average PQ segment voltage averaged over the at least two beats of the first segment. 3. The system of claim 1 where the average signal amplitude of the at least two beats of the first segment is the average peak-to-peak, R-to-S, signal amplitude of the at least two beats of the first segment. 4. The system of claim 1 where the average signal amplitude of the at least two beats of the first segment is the difference between the average PQ segment voltage and the peak S wave voltage averaged over the at least two beats of the first segment. 5. The system of claim 1 where the average signal amplitude of the at least two beats of the first segment is the larger of the R wave to PQ segment voltage difference or the PQ segment to S wave voltage difference. 6. The system of claim 1 wherein the first segment is a baseline segment. 7. The system of claim 1 wherein the test is based on whether the difference between the ST segment voltage of at least one beat of within a plurality of recently collected segments and the average ST segment voltage of the at least two beats of the first segment exceeds the detection threshold. 8. The system of claim 1 wherein the processor is configured to compute the difference between the ST segment voltage of a plurality of beats within the current electrogram segment and the average ST segment voltage of the at least two beats of the first segment, thereby... | Angel Medical Systems, Inc. | A61B005/00
US07854166 | B2 | 2010-12-21 | Method for evaluating an instrument operating force | 2008-06-24 | Different loads are applied to the skeletal muscle of at least one of an upper arm, an upper limb girdle, and a lower limb of the test subject to cause isometric muscle contraction under each load and measure the muscle activity of the skeletal muscle. Then, an upper limit of the muscle activity in a region where a magnitude of the load and an intensity of the muscle activity have a linear relationship is determined as the limit in an appropriate muscle activity range that suits the test subject. Subsequently, the muscle activity of the skeletal muscle of a test subject operating an instrument is measured, and the test subject's force for operating the instrument is evaluated based on a frequency distribution where the frequency of the muscle activity of the skeletal muscle being exercised to operate the instrument is distributed with respect to the obtained limit. | 1. An instrument operating force evaluation method for evaluating an instrument operating force required to operate an instrument when a test subject operates the instrument, the evaluation being made specifically according to a test subject, the method comprising: a step of measuring, when applying different loads onto a skeletal muscle of at least one of an upper limb, an upper limb girdle, and a lower limb of the test subject to cause isometric muscle contraction, a first muscle activity of the skeletal muscle observed under each of the different loads,a step of determining an upper limit of the first muscle activity in a region where a magnitude of each of the different loads at a time of the isometric muscle contraction and an intensity of the first muscle activity have a linear relationship as a limit of an appropriate muscle activity range that suits the test subject,a step of measuring a second muscle activity of the skeletal muscle of the test subject operating the instrument, anda step of comparing a frequency distribution of the second muscle activity of the skeletal muscle being exercised for operation of the instrument with the limit and evaluating the instrument operating force required to operate the instrument according to the test subject. 2. The instrument operating force evaluation method according to claim 1, wherein, when determining the limit, a position at which a rate of change in gradient representing a change in magnitude of the first muscle activity with respect to a change in each of the different loads exceeds a given value is obtained to determine a value of the first muscle activity at the position as the limit. 3. The instrument operating force evaluation method according to claim 1, wherein, when determining the limit, in a characteristic curve defined by the magnitude of each of the different loads and the intensity of the first muscle activity, a first region is approximated by one straight line, and a second region where each of the different loads is greater than in the first region is approximated by one curve to obtain a connection point where the approximated straight line and the approximated curve meet and determine a value of the first muscle activity at the connection point as the limit. 4. The instrument operating force evaluation method according to claim 1, wherein a frequency in a region of the frequency distribution where the second muscle activity exceeds the limit and a frequency in a region of the frequency distribution where the second muscle activity is lower than the limit are compared with a given value to evaluate the instrument operating force. 5. The instrument operating force evaluation method according to claim 1, wherein the instrument operating force is evaluated using a ratio of a frequency in a region of the frequency distribution where the second muscle activity exceeds the limit to a frequency in a region of the frequency distribution where the second muscle activity is lower ... | The Yokohama Rubber Co., Ltd. | A61B005/22
US07853302 | B2 | 2010-12-14 | Catheter with contractable mapping assembly | 2009-08-21 | The invention is directed to a mapping catheter useful for mapping tubular regions in and around the heart. The catheter comprises an elongated tubular catheter body. A mapping assembly is provided at the distal end of the catheter body. The mapping assembly comprises a tubular structure comprising a preformed generally circular main region generally transverse and distal to the catheter body and having an outer circumference. The tubular structure comprises a non-conductive cover over at least the main region of the mapping assembly. A plurality of electrodes are carried by the generally circular main region of the mapping assembly. A control handle is mounted at the proximal end of the catheter body. A contraction wire extends through the catheter body and non-conductive cover of the mapping assembly for contracting the generally circular main region of the mapping assembly. The contraction wire has a distal end anchored in the non-conductive cover and a proximal end anchored to a mechanism in the control handle that facilitates longitudinal movement of the contraction wire relative to the catheter body. The portion of the contraction wire extending through the non-conductive cover is positioned on the side of the generally circular main region closer to the center of the generally circular region. | 1. A mapping catheter comprising: an elongated tubular proximal shaft having an outer wall and proximal and distal ends;a distal shaft having proximal and distal ends, the proximal end of the distal shaft being attached to the distal end of the proximal shaft;a mapping assembly at the distal end of the distal shaft, the mapping assembly comprising: a tubular structure comprising a pre-formed generally circular main region generally transverse and distal to the proximal shaft and having an outer circumference, wherein the tubular structure comprises a non-conductive cover over at least the main region of the mapping assembly, anda plurality of electrodes carried by the generally circular main region of the mapping assembly;a control handle mounted at the proximal end of the proximal shaft; anda contraction wire extending through the proximal shaft and non-conductive cover of the mapping assembly for contracting the generally circular main region of the mapping assembly, the contraction wire having a distal end anchored in the non-conductive cover and a proximal end anchored to a mechanism in the control handle that facilitates longitudinal movement of the contraction wire relative to the catheter body, wherein the portion of the contraction wire extending through the non-conductive cover extends through a plastic tube and is positioned on the side of the generally circular main region closer to the center of the generally circular region. 2. A catheter according to claim 1, wherein the plastic tube comprises polyimide. 3. A catheter according to claim 1, wherein the plastic tube comprises a polyimide layer and a braided metal layer over the polyimide layer. 4. A catheter according to claim 1, wherein the mapping assembly further comprises a pre-formed support member having at least a portion extending through the non-conductive cover. 5. A catheter according to claim 4, wherein the pre-formed support member comprises nitinol. 6. A catheter according to claim 4, wherein the portion of the contraction wire and the portion of the pre-formed support member extending through the non-conductive cover extend through the plastic tube. 7. A catheter according to claim 6, wherein the plastic tube comprises polyimide. 8. A catheter according to claim 6, wherein the plastic tube comprises a polyimide layer and a braided metal layer over the polyimide layer. 9. A catheter according to claim 6, wherein the contraction wire and the pre-formed support member each have a distal end anchored to a metal tube. 10. A catheter according to claim 6, wherein the portion of the contraction wire extending through the non-conductive cover is positioned closer than the support member to the center of the generally circular region. 11. A catheter according to claim 4, wherein the portion of the contraction wire extending through the non-conductive cover is positioned closer than the support member to the center of the generally circular region. 12. A catheter accor... | Biosense Webster, Inc. | A61B005/04
US07841992 | B2 | 2010-11-30 | Tissue penetration device | 2005-12-22 | In one embodiment of the present invention a lancet is provided that has a sharpened distal tip, a shaft portion and a proximal drive head. The proximal drive head has a transverse dimension that is substantially larger than a transverse dimension of the shaft portion adjacent to the proximal drive head. | 1. A lancet module comprising a lancet channel formed within a module body portion;a lancet configured to be mechanically registered and coupled to a controllable force lancet driver;a sample chamber coupled to the lancet channel;wherein the lancet has a sharpened distal tip, a shaft portion and a drive head, the shaft portion being slidably disposed and free to slide within the lancet channel;wherein the lancet channel has a proximal end, a distal end, a series of lancet bearing guides configured to fit closely with a shaft of the lancet and confine the lancet to substantially axial movement, the lancet bearing guides includes a proximal lancet bearing guide and a distal lancet guide;a proximal fracturable seal positioned between a proximal lancet bearing guide and the shaft portion of the lancet that seals the sample; anda distal fracturable seal disposed between the lancet and the distal lancet guide to seal a distal end of the lancet channel and a sample input port. 2. The lancet module of claim 1, wherein the sharpened distal tip of the lancet is withdrawn into the lancet channel when the lancet is in a retracted position and the sharpened distal tip extending beyond a distal end of the lancet module when the lancet is in an extended position. 3. The lancet module of claim 1, wherein the module body portion is comprised of a polymer. 4. The lancet module of claim 1, wherein the module body portion has a length of about 0.5 to about 2 inch and a major transverse dimension of about 0.1 to about 0.5 inch. 5. The lancet module of claim 1, wherein the module body portion is comprised of PMMA. 6. The lancet module of claim 1, wherein the module body portion is substantially rectangular in shape and comprises a recess configured to be engaged by a ratchet drive mechanism. 7. The lancet module of claim 1, wherein the module body portion further comprises a lancet drive head slot disposed between a first protective strut and a second protective strut of the module body portion. 8. The lancet module of claim 7, wherein the first and second protective struts comprise elongated protective struts extending substantially parallel to the lancet. | Pelikan Technologies, Inc. | A61B005/00
US07848799 | B2 | 2010-12-07 | Drill device and method for forming microconduits | 2005-08-17 | The present invention relates to methods and devices for formation of microconduits in tissue, particularly using an impedance sensing drill to form microconduits. One embodiment of the invention is an impedance sensing drill comprising a drilling assembly, a control module, mechanically connected to the drilling assembly for controlling the depth of drilling by the drilling assembly; and a sensor, electrically connected to the drilling assembly and control module for detecting a change in an electrical impedance of a material being drilled. Another embodiment is a method of forming a microconduit in a material, which comprises the steps of drilling into the material, monitoring an electrical impedance of the material, and stopping the drilling into the material when a change in the impedance is detected, thereby forming microconduit. | 1. An impedance sensing drill comprising: a drilling assembly including a motor and a cutter which is adapted to be moved into and out of a tissue being drilled for the formation of microconduits;a control module mechanically connected to the drilling assembly for controlling the depth of drilling by the drilling assembly; anda sensor, electrically connected to the drilling assembly and the control module, for detecting a change in an electrical impedance of the tissue being drilled;wherein, upon detection of a change in the electrical impedance, the sensor sends a signal to the control module of the drill to automatically reverse the forward motion of the cutter to withdraw the cutter from the tissue, and thereby control the depth of the drilling. 2. The impedance sensing drill of claim 1, wherein the electrical impedance being measured is electrical resistance. 3. A method of forming a microconduit in mammalian tissue, which comprises the steps of drilling into the tissue with an impedance sensing drill according to claim 1, monitoring an electrical impedance of the tissue, and stopping the drilling into the tissue and automatically reversing the forward motion of the cutter with the control module when a change in the electrical impedance is detected, thereby forming a microconduit in the tissue. 4. The method of claim 3, wherein the mammalian tissue is selected from the group consisting of tissue, stratum corneum, epidermis, dermis, cartilage, nails, and bone. 5. The method of claim 4, wherein the mammalian tissue is nail tissue. 6. A method of forming a microconduit in mammalian tissue, which comprises the steps of drilling into the tissue with a drilling assembly including a motor and a cutter, monitoring electrical impedance of the tissue with a sensor, and stopping the drilling into the tissue and automatically reversing the forward motion of the cutter with a control module when a change in the electrical impedance is detected by the sensor, thereby forming a microconduit in the tissue. 7. The method of claim 6, wherein the mammalian tissue is selected from the group consisting of tissue, stratum corneum, epidermis, dermis, cartilage, nails, and bone. 8. The method of claim 7, wherein the mammalian tissue is nail tissue. 9. The method of claim 6, wherein the electrical impedance being measured is electrical resistance. 10. An impedance sensing drill comprising: a drilling assembly including a cutter;a control module mechanically connected to the drilling assembly for controlling the depth of drilling by the cutter;a sensor, electrically connected to the drilling assembly and control module for detecting a change in an electrical impedance of mammalian tissue being drilled; anda mount allowing the cutter to move into or out of mammalian tissue being drilled;wherein upon detection of a change in the electrical impedance in the mammalian tissue, the sensor sends a signal to the control module which automatically withdraws the cutte... | Path Scientific, LLC Galderma Pharma S.A. | A61B005/05
US07848787 | B2 | 2010-12-07 | Relative impedance measurement | 2005-07-08 | A method for calibrating impedance includes coupling at least first, second, and third electrodes at respective locations to a surface of a body of a subject. A first current passing through the body between the first and second body-surface electrodes is measured, and a second current passing through the body between the first and third body-surface electrodes is measured. From the first and second currents, a contact factor is derived that is indicative of the impedance between at least one of the body-surface electrodes and the surface of the body. Also described are methods for sensing the position of a probe and for detecting tissue contact based on a relation between currents from the probe to body-surface electrodes. | 1. A method for position sensing, comprising: inserting a probe comprising a probe electrode into a body of a subject;measuring a first current passing through the body between the probe electrode and a first body-surface electrode coupled to a surface of the body;measuring a second current passing through the body between the probe electrode and a second body-surface electrode coupled to the surface of the body;calculating a relation between the first and second currents, the relation being indicative of a value of relative impedance between the probe electrode and each of the body-surface electrodes, and being based on a quotient of the first current and a sum of the first and second currents; andtracking movement of the probe within the body responsively to the relation. 2. The method according to claim 1, wherein calculating the relation comprises: determining at least two values of relative impedance between the probe electrode and the first and second body-surface electrodes, andsolving a set of linear equations having parameters comprising the at least two values of relative impedance. 3. The method according to claim 1, wherein tracking the movement of the probe comprises determining that a change in the first and second currents that does not change the relation between the first and second currents is indicative of contact of the probe electrode with tissues of varying impedance within the body, and not due to the movement of the probe. 4. Apparatus for position sensing, comprising: a probe, comprising a probe electrode, adapted to be inserted into a body of a subject;a first body-surface electrode and a second body-surface electrode, each adapted to be coupled at respective locations to a surface of the body; anda control unit, adapted to measure a first current passing through the body between the probe electrode and the first body-surface electrode, to measure a second current passing through the body between the probe electrode and the second body-surface electrode, to calculate a relation between the first and second currents, and to track movement of the probe within the body responsively to the relation,wherein the relation is indicative of a value of relative impedance between the probe electrode and each of the body-surface electrodes, and is based on a quotient of the first current and a sum of the first and second currents. 5. The apparatus according to claim 4, wherein the control unit is adapted to derive at least two values of relative impedance and to track movement of the probe by solving a set of linear equations having parameters comprising the at least two values of the relative impedance. 6. The apparatus according to claim 4, wherein the control unit is adapted to determine that a change in the first and second currents that does not change the relation between the first and second currents is indicative of contact of the probe electrode with tissues of varying impedance within the body, and not due to the move... | Biosense Webster, Inc. | A61B005/04 A61B005/05 A61B005/06 A61B019/00
US08781194 | B2 | 2014-07-15 | Aneurysm detection | 2010-04-16 | An approach to automatically detecting, classifying and/or highlighting abnormal structures such as brain aneurysms is based on three-dimensional studies of the brain vessels. The approach is applicable to effectively all currently available modalities of acquisition of the cerebral vessels, including magnetic resonance angiography (MRA), computed tomography angiography (CTA), and conventional catheter-based three-dimensional rotational angiography (3DRA). | 1. A method for analysis of a vascular system comprising: accepting data characterizing a surface structure of a vascular system of a subject, the data identifying locations on the surface of vessels of the vascular system based on a segmented three-dimensional image of a portion of the subject's body including the vascular system;at each of the identified locations on the surface of the vessels, determining, using at least one processor, a local descriptor of vessel shape based on a computation of a Writhe Number that is based on a local neighborhood of the vessel surface corresponding to the identified location, based on quantitative relationships between each location of a plurality of locations in the local neighborhood and the identified location; andusing the determined local descriptors of vessel shape relative to a threshold to determine diagnostic information for presentation. 2. The method of claim 1 wherein determining the local descriptor of vessel shape includes determining the local neighborhood to include the identified location and extending along the direction of blood flow of the vessel. 3. The method of claim 1 wherein the extent of the local neighborhood in the direction of blood flow of the vessel includes at least a length of vessel greater than a diameter of the vessel in a vicinity of the identified location. 4. The method of claim 1 wherein the local descriptor of vessel shape at the identified location comprises an accumulation, for each of the plurality of locations in the local neighborhood, of a pairwise geometric relationship between the identified location and the location in the neighborhood. 5. The method of claim 1 wherein determining the local descriptor of vessel shape comprises: determining a set of neighboring locations on the surface of the vessel as the plurality of locations in the local neighborhood;for each neighboring location, determining a quantitative relationship between the identified location and the neighboring location; andcombining the determined quantitative relationships to form the local descriptor of vessel shape. 6. The method of claim 5 wherein combining the determined quantitative relationships comprises summing the determined quantitative relationships. 7. The method of claim 5 wherein determining a set of neighboring locations depends on an estimated radius of the vessel at the identified location. 8. The method of claim 5 wherein determining the quantitative relationship between the identified location and the neighboring location comprises combining directions of a normal to the vessel surface at each of the identified location and the neighboring location. 9. The method of claim 8 wherein determining the quantitative relationship between the identified location and the neighboring location further comprises combining the directions of the normal to the vessel surface and a direction between the identified location and the neighboring location. 10. The method of claim 1 wher... | Tufts Medical Center, Inc. | A61B001/002 A61B001/005
US08565852 | B2 | 2013-10-22 | Auricle-installed device and bio-signal measurement apparatus | 2010-06-18 | Disclosed herein is an auricle-installed device including, a reference electrode to be placed on the ear lobe, a detection electrode to be placed on a surface exposed to bones surrounding the auricle of the ear, and a support body for supporting the reference and detection electrodes. | 1. An auricle-installed device comprising: a support body for supporting a reference electrode and a detection electrode, the support body comprising:a main section formed to fit a base of an auricle and provided with the reference electrode on one end of said main section; anda flexible wire-shaped member for connecting said main section to said detection electrode; andwherein the detection electrode has a structure that can be held in a hollow between the crus of helix of said auricle and the superior crux of antihelix of said auricle, and the reference electrode has a structure including:a substantially circular electrode plate configured to contact an ear lobe,a clip configured to connect the electrode plate to the one end of the support body, anda link section configured to enable the electrode plate to pivot about an end of the clip such that a user can secure the reference electrode to the ear lobe by pivoting the electrode plate about the link section toward the one end of the support body. 2. A bio-signal measurement apparatus comprising an auricle-installed device including: a support body for supporting a reference electrode and a detection electrode, the support body comprising: a main section formed to fit a base of an auricle and provided with the reference electrode on one end of said main section;a flexible wire-shaped member for connecting said main section to said detection electrode; anda differential amplifier configured to amplify a difference in electric potential between the reference electrode and the detection electrode as a bio-signal, andwherein the detection electrode has a structure that can be held in a hollow between the crus of helix of said auricle and the superior crux of antihelix of said auricle, andthe reference electrode has a structure including: a substantially circular electrode plate configured to contact an ear lobe,a clip configured to connect the electrode plate to the one end of the support body, anda link section configured to enable the electrode plate to pivot about an end of the clip such that a user can secure the reference electrode to the ear lobe by pivoting the electrode plate about the link section toward the one end of the support body. | Sony Corporation | A61B005/0478 A61B005/0492 A61B005/0496
US08391964 | B2 | 2013-03-05 | Detecting electrical conduction abnormalities in a heart | 2009-05-11 | Techniques are described for detecting conduction abnormalities in a heart of a patient. In particular, an IMD may be configured to obtain electrical signals corresponding to cardiac activity of the heart of the patient and periodically analyze a most recent electrical signal of the obtained electrical signals to detect an electrical conduction abnormality of the heart. The IMD adjusts a frequency at which the most recent electrical signal is analyzed based on at least one physiological parameter of the patient. For example, the IMD may increase the frequency at which the most recent electrical signal is analyzed when a heart rate parameter has significantly changed and the number of detected premature ventricular contractions (PVCs) is greater than or equal to a threshold number. In this manner, the most recent electrical signal is analyzed at a higher frequency in situations in which conduction abnormalities are more likely. | 1. An implantable medical device (IMD) comprising: a sensing module configured to acquire cardiac electrical signals corresponding to cardiac activity of a heart of a patient via at least one sensor; anda processor configured to obtain the cardiac electrical signals corresponding to cardiac activity of the heart of the patient, periodically analyze, at a first frequency, a most recent cardiac electrical signal of the obtained cardiac electrical signals to monitor for an electrical conduction abnormality of the heart, compute at least a heart rate parameter of the patient and a number of premature ventricular contractions (PVCs) using the acquired cardiac electrical signals, compare the heart rate parameter of the patient to a heart rate parameter threshold, compare the number of PVCs to a PVC threshold, and increase the frequency at which the most recent electrical signal is analyzed to a second frequency when the heart rate parameter of the patient is greater than the heart rate parameter threshold and the number of PVCs detected is greater than the PVC threshold, wherein the first frequency and the second frequency are greater than zero. 2. The IMD of claim 1, wherein the processor decreases the frequency at which the most recent electrical signal is analyzed when the heart rate parameter falls below a second heart rate parameter threshold. 3. The IMD of claim 1, wherein the heart rate parameter comprises one of a heart rate and a relative change in the heart rate. 4. The IMD of claim 1, wherein the processor is configured to obtain a template electrogram and determine whether one or more characteristics of the most recent electrical signal have significantly changed relative to the template electrical signal. 5. The IMD of claim 4, wherein the processor analyzes one of an ST segment of the electrical signal, a T-wave segment of the electrical signal and a QRS segment of the electrical signal. 6. The IMD of claim 1, wherein the processor is configured to obtain an EGM corresponding with a most recent beat, obtain a template EGM corresponding to a beat during a normal sinus, and periodically analyze the most recent cardiac electrical signal by comparing the EGM corresponding with a most recent beat to the template EGM using a wavelet transform analysis. 7. The IMD of claim 1, wherein the electrical conduction abnormality is associated with at least one of a cardiac ischemia, a cardiac infarction, a branch bundle block, or fascicular block. 8. An implantable medical device comprising: a sensing module configured to acquire cardiac electrical signals corresponding to cardiac activity of a heart of a patient via at least one sensor; anda processor configured to obtain the cardiac electrical signals corresponding to cardiac activity of the heart of the patient, periodically analyze, at a first frequency, a most recent cardiac electrical signal of the obtained cardiac electrical signals to monitor for an electrical conduction abnormality of th... | Medtronic, Inc. | A61B005/04
US08364225 | B2 | 2013-01-29 | Estimating transform values using signal estimates | 2009-05-20 | According to embodiments, estimated values for a signal transform may be generated using estimated values for the signal. Signal parameters may then be determined based on the estimated signal transform. A first portion of a signal may be obtained. A second portion of the signal may be estimated. The second portion of the signal may correspond to a portion of the that is unknown, that is not yet available and/or that is obscured by noise and/or artifacts. A transform (e.g., a continuous wavelet transform) of both of the signal portions may be performed. One or more parameters corresponding to the signal may then be determined from transformed signal. | 1. A method for estimating transform values using a signal estimate comprising: obtaining a first portion of a biological signal from a sensor;estimating a second portion of the signal based at least in part on the first portion of the signal;performing a transform of the signal based at least in part on the first portion of the signal and the second portion of the signal; anddetermining a biological parameter corresponding to the signal based at least in part on the transformed signal. 2. The method of claim 1 wherein the biological signal comprises a photoplethysmograph (PPG) signal and wherein the determined parameter comprises blood oxygen saturation level, respiration rate, respiration effort metric, pulse rate, and/or blood pressure. 3. The method of claim 1 wherein the transform comprises a continuous wavelet transform. 4. The method of claim 3 wherein estimating the second portion of the signal further comprises: selecting a range of scales for the continuous wavelet transform; anddetermining a length for the second portion of the signal based at least in part on the selected range of scales. 5. The method of claim 4 wherein selecting the range of scales for the continuous wavelet transform comprises locating one or more features within a continuous wavelet transform of the signal. 6. The method of claim 1 wherein estimating the second portion of the signal comprises: determining an amplitude value of the first portion of the signal;estimating the determined amplitude value as a constant value for the second signal. 7. The method of claim 1 wherein estimating the second portion of the signal comprises: fitting a function to the first portion of the signal; andestimating the second signal using the function. 8. The method of claim 1 wherein estimating the second portion of the signal comprises: determining that the first portion of the signal is a known signal type;determining one or more characteristic values of the first portion PPG signal; andestimating the second signal by generating a signal portion of the known signal type having at least one of the determined characteristic values. 9. The method of claim 1 further comprising determining an estimation confidence value for the second portion of the signal. 10. The method of claim 1 further comprising updating the transform of the signal in response to obtaining signal values corresponding to the second portion of the signal. 11. The method of claim 1 wherein the second signal portion is generally adjacent to an end of the first signal portion. 12. The method of claim 1 wherein the second signal portion is generally within the first signal portion. 13. A system for processing a signal to estimate transform values using a signal estimate, the system comprising: a sensor to receive data indicative of a biological signal;a processor coupled to the sensor, wherein the processor is capable of: obtaining a first portion of the received signal;estimating a second portion of the si... | Nellcor Puritan Bennett Ireland | A61B005/00 A61B005/02 A61B005/08 G01N033/48 G01N033/50
US08396532 | B2 | 2013-03-12 | MRI-guided devices and MRI-guided interventional systems that can track and generate dynamic visualizations of the devices in near real time | 2010-06-16 | An MRI-guided medical device includes an elongated sheath, an elongated dilator, and an elongated needle. The sheath has a distal end, an opposite proximal end, and a central lumen extending between the proximal and distal ends. The sheath comprises MRI-compatible material and includes a tracking member located adjacent to the sheath distal end that is visible in an MRI image. The dilator comprises MRI-compatible material and is movably disposed within the sheath lumen. A distal end of the dilator is configured to extend outwardly from the sheath distal end and dilator includes at least one tracking member that is visible in an MRI image. The needle is movably disposed within the dilator lumen and is movable between stored and operative positions relative to the dilator. An RF shield may be coaxially disposed within the elongated sheath so as to surround a portion of the sheath central lumen. | 1. An MRI-compatible medical device, comprising: an elongated sheath having a distal end, an opposite proximal end, and a central lumen extending between the proximal and distal ends, wherein the sheath includes at least one tracking member adjacent the sheath distal end;a dilator movably disposed within the sheath lumen and having a distal end, an opposite proximal end, and a lumen extending between the dilator proximal and distal ends, wherein the dilator includes at least one tracking member adjacent the dilator distal end, wherein the at least one tracking member includes a cable configured to electrically connect the at least one tracking member to an MRI scanner, and wherein the dilator distal end is configured to extend outwardly from the sheath distal end; anda needle movably disposed within the dilator lumen and having a distal end, an opposite proximal end, and a central lumen extending between the needle proximal and distal ends, wherein the needle comprises a main body portion of an MRI-compatible material and a tip portion, wherein the tip portion is located at the needle distal end wherein the needle is movable between retracted and extended positions, wherein the needle distal end is within the dilator lumen when in the retracted position and wherein the needle distal end extends outwardly from the dilator distal end when in the extended position, wherein the needle tip portion includes a tracking member that is closely spaced-apart from the at least one tracking member on the dilator distal end when the needle is in the extended position. 2. The device of claim 1, wherein a portion of the sheath adjacent the sheath distal end has a generally curved configuration. 3. The device of claim 1, wherein the at least one tracking member adjacent the sheath distal end comprises a coating of MRI visible material on the sheath. 4. The device of claim 1, wherein the at least one tracking member adjacent the sheath distal end comprises at least one RF coil. 5. The device of claim 1, wherein the at least one tracking member adjacent the dilator distal end comprises at least one RF coil. 6. The device of claim 5, wherein the at least one RF coil adjacent the dilator distal end is embedded within a wall of the dilator. 7. The device of claim 1, wherein the at least one tracking member adjacent the dilator distal end comprises a pair of RF coils in adjacent spaced-apart relationship. 8. The device of claim 1, wherein the needle tip portion comprises multiple sections of conductive and non-conductive material alternately connected together. 9. The device of claim 1, wherein the needle tip portion comprises a plurality of sections of different diameters arranged in a telescopic configuration. 10. The device of claim 1, wherein the needle tip portion comprises a coating of material. 11. The device of claim 1, wherein a portion of the needle adjacent the needle distal end has a generally curved configuration. 12. The device of claim 1, wher... | MRI Interventions, Inc. | A61B005/05
US08391962 | B2 | 2013-03-05 | Cardiac function circadian variation analysis system and method | 2009-05-13 | Systems and Methods for stratifying relative risks of adverse cardiac events by processing a duration of electrocardiograph recordings generally recorded by a Holter type of device. The duration of electrocardiograph recordings are processed to resolve RR interval related data, QT interval related data, and are fitted to formulas to at least partially establish fitting related measures. The fitting formulas incorporate circadian related periodic factors, and can further incorporate additional processing including utilizing Lissajous analysis techniques, among others. | 1. A method of stratifying risks of an adverse cardiac event comprising the steps of: ascertaining, from a first duration of electrocardiographic records, RR and QT interval data;determining mean RR and QT interval data, denoted by RRMEAN and QTMEAN, respectively;fitting, with a calculator, at least a portion of at least one of the RR and QT interval data to at least one of first and second formulas that include fitting related measures and define RRFIT and QTFIT, respectively, said first formula being, RRfit=RRmean+Cr cos [(2π/24)(t+φR)],and said second formula being, QTfit=QTmean+Cq cos [(2π/24)(t+φQ)],wherein t represents a temporal measure and said fitting at least partially establishes the fitting related measures Cr, Cq, φR, and φQ, said Cr and Cq denoting RR and QT circadian amplitude coefficients, respectively, and said φR and φQ denoting RR and QT acrophases, respectively;establishing at least a first correlation between at least a first cardiac adversity relative risk related attribute and at least one of the fitting related measures; andstratifying relative risk of at least a first cardiac adversity in response to the at least a first correlation, said stratifying involving at least a first stratifying mathematical expression, said first stratifying mathematical expression involving at least two of the fitting related measures. 2. A method of stratifying risks of an adverse cardiac event according to claim 1, wherein said at least two fitting related measures are selected from a group consisting of said Cr, Cq, φR, φQ, RRFIT, and QTFIT. 3. A method of stratifying risks of an adverse cardiac event according to claim 1, wherein the first stratifying mathematical expression determines a Lissajous slope ΔLJ from a third formula, ΔLJ≡(Cq/Cr) cos [(2π/24)(φR−φQ)]. 4. A method of stratifying risks of an adverse cardiac event according to claim 1, wherein the first stratifying mathematical expression determines a phase difference φ from a fourth formula, φ=φR−φQ. 5. A method of stratifying risks of an adverse cardiac event according to claim 1, said fitting involving varying at least one of said Cr and Cq. 6. A method of stratifying risks of an adverse cardiac event according to claim 1, said fitting at least partially involving regression analysis. 7. A method of stratifying risks of an adverse cardiac event according to claim 1, further comprising the step of plotting said RRfit in relation to said QTfit to produce a RR-QT Lissajous curve representation of at least a portion of the first duration of electrocardiograph records and resolving at least a first Lissajous curve representation feature, denoted by FLjC, of the plotted representation, wherein said resolving the FLjC entails said at least two fitting related measures. 8. A method of stratifying risks of an adverse cardiac event according to claim 1, further comprising the step of distinguishing at least a first treatment in response to said relative risk stratifying, said first... | - | A61B005/04
US08473028 | B2 | 2013-06-25 | K-space sample density compensation for magnetic resonance image reconstruction | 2008-11-21 | The present invention provides a method for producing density-compensated MR images having improved signal-to-noise ratio and with reduced computational burden. The method includes sampling MR data using a rotationally symmetric acquisition trajectory, generating a cross-correlation matrix, and applying a discreet Fourier transform (DFT) to the acquired MR data and the cross correlation matrix. The method further includes employing the transformed MR data and transformed cross-correlation matrix in a linear system to produce a set of density-compensated transformed MR data, which may transformed into density compensated MR data by application of an inverse DFT. The density compensated MR data may be reconstructed by a variety of techniques to produce density compensated images. | 1. A method for producing density-compensated magnetic resonance (MR) images with a magnetic resonance imaging (MRI) system, the steps comprising; a) acquiring MR data using a rotationally symmetric acquisition trajectory;b) generating a cross-correlation matrix from the acquired MR data;c) applying a Fourier transform (FT) to the acquired MR data and the cross-correlation matrix produced in step b);d) computing a set of density-compensated, transformed MR data from the transformed cross-correlation matrix and transformed MR data produced in step c);e) applying an inverse FT to the density-compensated, transformed MR data to produce a set of density-compensated MR data; andf) performing image reconstruction on the density-compensated MR data to produce a density-compensated MR image. 2. The method as recited in claim 1 wherein the rotationally symmetric acquisition trajectory of step a) is at least one of a periodically rotated overlapping parallel lines with enhanced reconstruction (PROPELLER), BLADE, radial, concentric rings, and spiral trajectory. 3. The method as recited in claim 1 wherein the cross-correlation matrix of step b) has a Toeplitz form. 4. The method as recited in claim 1 wherein step c) further includes applying a discrete Fourier transform (DFT) along a rotationally symmetric dimension and wherein step e) further includes applying an inverse DFT along a rotationally symmetric dimension. 5. The method as recited in claim 1 wherein step d) further includes employing a linear system to compute a set of density-compensated, transformed MR data from the transformed cross-correlation matrix and the transformed MR data. 6. The method as recited in claim 5 further including employing a scaling factor to produce a set of density-compensated, transformed MR data having reduced amplitudes. 7. The method as recited in claim 6 wherein the employment of a scaling factor allows the production of density compensated MR images having an improved signal-to-noise ratio and reduced error. 8. The method as recited in claim 5 wherein the linear system includes a plurality of independent operations configured to be solved in parallel. 9. The method as recited in claim 1 wherein step f) further includes regridding the density-compensated MR data to a Cartesian grid and using a Fourier transform technique to reconstruct the regridded, density-compensated MR data to a density-compensated MR image. 10. The method as recited in claim 1 wherein step f) further includes using a non-uniform Fourier transform technique to reconstruct the density-compensated MR data to a density-compensated MR image. 11. A method for weighting image data to account for non-uniform sampling density and producing a weighted image therefrom, the method comprising the steps of: a) acquiring a set of non-uniform image data in a first image space having a non-uniform sampling density;b) generating elements of a matrix having a Toeplitz form from the acquired set of image d... | The Brigham and Women's Hospital Children's Medical Center Corporation | A61B005/055 G06K009/00
US09163273 | B2 | 2015-10-20 | Biosensors and methods for making and using them | 2010-05-20 | Embodiments of the invention provide analyte sensors having optimized permselective membranes and methods for making and using such sensors. Embodiments of the invention also provide analyte sensors such as those having porous matrices coated with an analyte sensing composition and methods for making and using such sensors. Illustrative embodiments include electrochemical glucose sensors having glucose oxidase coatings. | 1. A composition comprising a porous matrix comprising a lattice of spherical particles having a surface coated with an immobilized enzyme, wherein: the porous matrix comprises a micro-porous metal that functions as an electrode in an electrochemical sensor, wherein the electrode consumes hydrogen peroxide in the electrochemical sensor;the porous matrix is at least 10, 100, or 1000 microns thick; andthe composition is implantable within a mammal. 2. The composition of claim 1, wherein the enzyme is glucose oxidase, glucose dehydrogenase, lactate oxidase, hexokinase or lactate dehydrogenase. 3. The composition of claim 1, wherein the porous matrix has a surface area that is at least 2, 4, 6, 8, 10, 12, 14, 16 or 18 times the surface area of a non-porous matrix of same dimensions. 4. The composition of claim 1, wherein the enzyme coated on the surface is in a coating layer having a thickness selected from the group consisting of less than 1, 0.5, 0.25 and 0.1 microns. 5. The composition of claim 1, wherein the porosity range of the porous matrix is 40%-99%. 6. An analyte sensor apparatus for implantation within a mammal, the analyte sensor apparatus comprising a porous matrix comprising a lattice of spherical particles having a surface coated with an immobilized enzyme, wherein: the porous matrix comprises a micro-porous metal;the porous matrix is at least 1 micron in thickness;the immobilized enzyme is coated with an analyte modulating layer comprising a siloxane;and the porous matrix functions as a working electrode in the analyte sensor apparatus, andwherein the electrode consumes hydrogen peroxide in the analyte sensor. 7. The analyte sensor apparatus of claim 6, wherein: the immobilized enzyme is disposed within an analyte sensing layer disposed on the working electrode, such that the analyte sensing layer detectably alters the electrical current at the working electrode in the presence of an analyte. 8. The analyte sensor apparatus of claim 7, wherein: the analyte modulating layer is disposed on the analyte sensing layer and, modulates the diffusion of oxygen therethrough. 9. The analyte sensor apparatus of claim 8, further comprising an adhesion promoting layer disposed on the analyte sensing layer, wherein the adhesion promoting layer promotes the adhesion between the analyte sensing layer and an analyte modulating layer disposed on the analyte sensing layer. 10. The analyte sensor apparatus of claim 8, further comprising a protein layer disposed between the analyte sensing layer and the analyte modulating layer. 11. The analyte sensor apparatus of claim 8, further comprising an interference rejection layer disposed between the surface of the working electrode and the analyte sensing layer. 12. The analyte sensor apparatus of claim 8, further comprising a cover layer disposed on at least a portion of the analyte modulating layer, wherein the cover layer further includes an aperture that exposes at least a portion of the analyte m... | Medtronic MiniMed, Inc. | A61B005/05 C12Q001/00
US08613892 | B2 | 2013-12-24 | Analyte meter with a moveable head and methods of using the same | 2009-06-30 | The present disclosure provides in vitro analyte meters that include a meter portion that is moveable relative to at least one other meter portion. Embodiments include moveable meters that are integrated with in vivo analyte systems. Also provided are methods, systems and kits. | 1. An in vitro glucose meter, comprising: a housing including a user interface portion, the housing further including an in vitro glucose test strip contacting portion pivotally mounted to the user interface portion of the housing so that the in vitro glucose test strip contacting portion is selectively pivotable relative to the user interface portion, wherein the in vitro glucose test strip contacting portion pivots through a range of about 180 degrees, and further comprising a lock to substantially fix the in vitro glucose test strip contacting portion in a position throughout the range. 2. The meter of claim 1, wherein the in vitro glucose test strip contacting portion is pivotable towards a reporting module of the user interface portion. 3. The meter of claim 2, wherein the reporting module comprises a display. 4. The meter of claim 1, wherein the in vitro glucose test strip contacting portion is pivotable away from a reporting module of the user interface portion. 5. The meter of claim 4, wherein the reporting module comprises a display. 6. The meter of claim 1, wherein the in vitro glucose test strip contacting portion is pivotable to provide an angle of about 90 degrees relative to the user interface portion. 7. The meter of claim 1, wherein the in vitro glucose test strip contacting portion comprises an in vitro glucose test strip port that is pivotally mounted to the user interface portion of the housing. 8. The meter of claim 1, further including an in vivo glucose monitoring system integrated in the housing. 9. A method, comprising: selectively moving an in vitro glucose test strip contacting portion of an in vitro glucose meter relative to a user interface portion of the meter;contacting an in vitro glucose test strip to the in vitro glucose test strip contacting portion;applying a biological sample to the in vitro glucose test strip; andobtaining information about glucose from the user interface portion of the meter; wherein the in vitro glucose meter includes a housing, the housing further including the in vitro glucose test strip contacting portion pivotally mounted to the user interface portion so that the in vitro glucose test strip contacting portion is selectively pivotable relative to the user interface portion, wherein the in vitro glucose test strip contacting portion pivots through a range of about 180 degrees, and further including a lock to substantially fix the in vitro glucose test strip contacting portion in a position throughout the range. 10. The method of claim 9, comprising moving the in vitro glucose test strip contacting portion towards a reporting module of the user interface portion. 11. The method of claim 9, comprising moving the in vitro glucose test strip contacting portion away from a reporting module of the user interface portion. 12. The method of claim 9, comprising moving the in vitro glucose test strip contacting portion to provide an angle of about 90 degrees relative to the user interface... | Abbott Diabetes Care Inc. | G01N035/00
US20100280360 | A1 | 2010-11-04 | TWO-PART LOCAL SHOULDER COIL, AND MAGNETIC RESONANCE TOMOGRAPHY APPARATUS AND METHOD EMPLOYING SAME | 2010-04-30 | Improved patient comfort is provided by a local coil arrangement designed to support the production of an MRT image of a shoulder, wherein the housing of the local coil arrangement is composed of two separable parts. | 1. A magnetic resonance tomography local shoulder coil arrangement, comprising: a local shoulder coil configured to at least receive magnetic resonance signals from a shoulder of a patient, said local shoulder coil having a shape and dimensions conforming said local shoulder coil to a shape of a human shoulder;a housing containing said local shoulder coil; andsaid housing and said local shoulder coil contained therein being comprised of two separable parts, and said housing comprising a detachable electrical and mechanical connection configured to detachably connect said two parts. 2. A local shoulder coil arrangement as claimed in claim 1 wherein said housing is configured to allow a patient having a shoulder located in said local shoulder coil to lie on a patient table, and wherein a first of said two parts is configured to be located closer to said patient table and wherein a second of said two parts is configured to be located farther from said patient table. 3. A local shoulder coil arrangement as claimed in claim 1 wherein said first of said parts comprises a detachable connection to said patient table. 4. A local shoulder coil arrangement as claimed in claim 2 wherein said table has a base plate thereon, and wherein said first of said two parts is connectable to said base plate. 5. A local shoulder coil arrangement as claimed in claim 1 wherein at least one of said two separable parts is configured to be movable toward said shoulder. 6. A local shoulder coil arrangement as claimed in claim 1 wherein said two separable parts are configured to be moved toward each other in order to connect said two parts with each other. 7. A local shoulder coil arrangement as claimed in claim 1 wherein a first of said two separable parts is a part from among a set of parts having respectively different sizes and dimensions to respectively conform two differently sized shoulders, with each part in said set of parts being configured for connection to a second of said two parts. 8. A local shoulder coil arrangement as claimed in claim 1 wherein said two parts have a plane of symmetry therebetween allowing use of said two separable parts both for the left shoulder and the right shoulder of a patient. 9. A local shoulder coil arrangement as claimed in claim 1 comprising a first electrical connection at a first of said two separable parts configured to electrically connect a portion of said local shoulder coil in said first of said two separable parts to a magnetic resonance imaging system, and a second electrical connection at a second of said two separable parts configured to electrically connect a portion of said local shoulder coil contained in said second of said two separable parts to said magnetic resonance imaging system. 10. A magnetic resonance tomography apparatus comprising: a magnetic resonance data acquisition unit configured to receive a patient therein;a local shoulder coil arrangement in said data acquisition unit; andsaid local shoulder... | - | A61B005/055
US20100292586 | A1 | 2010-11-18 | Wireless automatic ankle-brachial index (AABI) measurement system | 2009-05-13 | A central processor controlled system and test procedure for independently, contemporaneously and plethysmographically monitoring arterial blood pressure at a subject's arms and legs. Micro-controlled occluding and sensing cuffs containing sensors are inflated and deflated at each limb of a supine patient. Sensed AC and DC pressure data is wirelessly linked to the central processor where the DC sensor data is sampled to derive median filtered, fitted and derivative waveforms that are iteratively processed and scored to determine a table of sample indices indicative of lowest pressure point. Second scoring and fittings about the lowest pressure point at the DC sensor data and original occluding cuff pressure data identify each limb's systolic pressure. The derived systolic limb pressure values are then processed to determine right and left ABI values. | 1. Apparatus for determining a blood pressure of a subject comprising: a) a pressure monitor adapted to be coupled to a limb of a test subject, wherein the monitor comprises 1) first and second inflatable cuffs and first and second pressure sensors respectively coupled to said first and second cuffs, 2) a compressor, 3) a deflation valve coupled to said first and second cuffs, 4) a first processor and wherein said first processor i) controls said compressor to inflate said first cuff to engage said first sensor to the limb and until blood flow through the limb is occluded and, ii) controls said compressor to inflate said second cuff to engage said second sensor to the limb at a location distal to said first cuff, iii) controls said deflation valve to deflate said first cuff in a plurality of steps, iv) stores data from said first sensor defining the pressure of said first cuff at occlusion and at each of said steps, and v) stores blood flow pressure data from said second sensor until termination of the deflation of the first cuff;b) a wireless communication link;c) a second processor coupled to said first processor via said communication link to receive data collected from said first and second sensors; andd) wherein the second processor processes the data received from said first and second sensors by i) sampling the occluding pressure data from said first sensor and the blood flow pressure data from said second sensor in a filtering operation to obtain first and second tables of indexed sample data and whereby the sample data exhibits reduced artifact and noise variances, ii) sampling the blood flow pressure sample data from table two in a derivative function to obtain a third table of sampled derivative data including at least one value identifying at least one slope transition of a waveform defined by the sample data of table two from declining to inclining, iii) arithmetically scoring seriatim groupings of the filtered sample data of table two relative to a plurality of conditions of the waveform of the filtered sample values and relative to the derivative values of said table three to determine an arithmetic score defining a low point sample value where the at least one slope transition of the filtered blood flow pressure sample data changes from declining to inclining, iv) correlating the low point sample value of table two relative to the unfiltered pressure data measured by the sensor one to determine a systolic pressure when blood flow returns to the monitored limb. 2. Apparatus as set forth in claim 1 wherein said first and second sensors are integrated into said first and second cuffs and wherein the occluded first cuff is deflated in a linear fashion with a constant pressure drop at each deflation step. 3. Apparatus as set forth in claim 1 wherein during inflation of said first and second cuffs the compressor is gradually slowed as an occlusion pressure for said first cuff and a sensor retention pressure for said second cuff is nea... | - | A61B005/02
US20100305435 | A1 | 2010-12-02 | Bone Marking System and Method | 2009-05-27 | An apparatus for marking bone of a patient includes a marker, a tracking system and a display. The marker includes pigment and is configured to impart the pigment on the bone. The tracking system is in communication with the marker. The tracking system is configured to determine a location of the marker relative to the bone and to match at least one anatomical feature of the bone of the patient to an image of the bone. The display is in communication with the tracking system and adapted to show an image of the location of the marker relative to the image of the bone while the marker imparts the pigment on the bone. | 1. An apparatus for marking bone of a patient, comprising: a marker comprising pigment and being configured to impart the pigment on the bone;a tracking system in communication with the marker, configured to determine a location of the marker relative to the bone, and configured to match at least one anatomical feature of the bone of the patient to an image of the bone; anda display in communication with the tracking system and adapted to show an image of the location of the marker relative to the image of the bone while the marker imparts the pigment on the bone. 2. The apparatus of claim 1 further comprising a reference sensor in communication with the tracking system and positionable relative to the bone of the patient, the reference sensor configured to provide for the tracking system a frame of reference fixed to the bone. 3. The apparatus of claim 1 wherein the tracking system comprises and is in separate communication with: a transmitter disposed in a known position relative to the bone of the patient; anda receiver attachable to the marker, the receiver sensing a signal from the transmitter to provide location information of the marker so that the tracking system can determine the location of the marker relative to the bone. 4. The apparatus of claim 1 wherein the tracking system comprises and is in separate communication with: a transmitter attachable to the marker; anda receiver disposed in a known position relative to the bone of the patient, the receiver sensing a signal from the transmitter to provide location information of the marker so that the tracking system can determine the location of the marker relative to the bone. 5. The apparatus of claim 1 wherein the at least one anatomical feature includes at least one tooth of the patient. 6. The apparatus of claim 1 wherein the at least one anatomical feature includes a feature of a surface of the bone. 7. The apparatus of claim 1 wherein the at least one anatomical feature includes three anatomical features of the bone. 8. The apparatus of claim 1 wherein the tracking system is adapted to match the at least one anatomical feature of the bone of the patient to the image of the bone without exogenous fiducial markers. 9. The apparatus of claim 1 wherein the tracking system determines location from location information including position and orientation of the marker relative to the bone. 10. The apparatus of claim 1 wherein the marker is a pencil. 11. A method of marking a bone of a patient, comprising: matching at least one anatomical feature of the bone of the patient to an image of the bone;determining a location of a marker comprising pigment relative to the bone;showing an image of the location of the marker relative to the image of the bone on a display;using the display to facilitate positioning the marker relative to a marking locating on the bone; andusing the marker to impart the pigment to the marking locating on the bone. 12. The method of claim 11 further compr... | - | A61B005/05
US20100298691 | A1 | 2010-11-25 | Method of examining dynamic cardiac electromagnetic activity | 2010-07-27 | A method of examining cardiac electromagnetic activity over a heart for diagnosing the cardiac functions of the heart is disclosed. The method may include constructing a phase diagram of electromagnetic signals over a heart by collecting sets of time-dependent magnetic signals, determining the zeroth and the first derivations of each set of the magnetic signals at a given time, and categorizing the zeroth and the first derivations of the magnetic signals in either of the four phases: (+, +), (−, −), (+, −), (−, +). The method may also include monitoring a wave propagation of the magnetic signals. | 1. A method of examining cardiac electromagnetic activity by monitoring a propagation of a wave of magnetic signals of a heart. 2. The method of claim 1, wherein the magnetic signals are either two-dimensionally or three-dimensionally distributed over the heart. 3. The method of claim 1, wherein the magnetic signals exhibit features of a P-wave, a Q-wave, a R-wave, a S-wave and a T-wave corresponding to an electrocardiography, and the propagation of the P-wave, the Q-wave, the R-wave, the S-wave, the T-wave or a combination thereof is monitored. | - | A61B005/05
US20100292557 | A1 | 2010-11-18 | METHOD AND DEVICE FOR SUBSTANCE MEASUREMENT | 2008-03-19 | Embodiments of the present disclosure present systems, devices and methods for ISF glucose monitoring that more accurately reflects the blood glucose levels by introducing a treatment element allowing stable and accurate prediction of blood glucose levels based measured glucose levels from interstitial fluid. | 1-55. (canceled) 56. An apparatus for measuring an interstitial fluid (“ISF”) in a tissue, comprising: a sensor arranged to be in fluid communication with ISF during use of the apparatus for measuring a level of a substance;a treatment element inserted to or otherwise in contact with tissue in which said sensor is arranged and/or tissue adjacent thereto; anda processor capable of communicating with said sensor and said treatment element, wherein said processor induces activity of said treatment element in relation to said sensor measurements of said substance. 57. The apparatus according to claim 56, wherein said substance comprises glucose. 58. The apparatus according to claim 56, wherein said energy source supplies energy selected from the group consisting of: radiation, mechanical vibrations, pressurization, suction, massaging, acoustic stimulation, magnetic field, electrical stimulation and RF (radiofrequency) energy, light, topical application of an additional substance, infusion of an additional substance and any combination of the foregoing. 59. The apparatus according to claim 58, wherein said treatment element comprises a heating element. 60. The apparatus according to claim 57, further comprising an insulin pump for infusion of insulin in response to measurement of a level of glucose. 61. The apparatus according to claim 60, wherein said insulin is infused in the vicinity of the sensor. 62. The apparatus according to claim 61, wherein said treatment element comprises a heating element, such that it heats the infused and measured tissue vicinity without heating the insulin above a limiting temperature that might damage the insulin. 63. The apparatus according to claim 56, further comprising a catheter inserted to the tissue for receiving the ISF. 64. The apparatus according to claim 63, wherein said sensor is disposed within said catheter. 65. The apparatus according to claim 63, further comprising a fluid pump or microdialysis pump for pumping the ISF to said sensor. 66. The apparatus according to claim 56, further comprising a display in communication with said processor for displaying at least one of a treatment level and said level of said substance. 67. The apparatus according to claim 66, wherein said treatment comprises application of heat and said treatment level is temperature. 68. The apparatus according to claim 67, wherein said heat is produced by a heating element embedded into the tissue. 69. The apparatus according to claim 56, wherein at least one element is disposable and at least one element is reusable. 70. The apparatus according to claim 56, wherein said processor is configured to analyze data of one or more previous glucose measurements in at least one of the ISF and blood level glucose measurements or additional logged information; anddetermine whether to apply the treatment accordingly. 71. The apparatus according to any of claims 56, wherein said treatment element is configured to perform at least... | - | A61B005/00
US20100310471 | A1 | 2010-12-09 | ANTIOXIDANT AND PARAMAGNETIC HEPARIN-NITROXIDE DERIVATIVES | 2008-08-22 | The present invention relates to novel heparin-nitroxide derivatives comprising heparin and at least two and more nitroxides/polynitroxide radicals that are covalently coupled to heparin by derivatisation of glycosaminoglycan carboxyl or amino groups. The heparin-nitroxide derivatives are useful as therapeutic or diagnostic agents. This invention further concerns novel methods for the production of the heparin-nitroxide agents, and methods of their uses for specifically targeting and labelling of biological vessels. The inventions also suggest the uses of the heparin-nitroxide derivatives for treatment of oxidative stress-mediated diseases. Furthermore, the heparin-nitroxide derivatives according to the present invention are in particular useful for electron paramagnetic resonance imaging (EPRI), for magnetic resonance imaging (MRI), and for preservation of biological transplants. | 1. Heparin-nitroxide derivative, comprising heparin and at least two and more nitroxides or polynitroxides that are covalently coupled to heparin by derivatisation of glycosaminoglycan carboxyl groups and/or glycosaminoglycan amino groups of heparin, the heparin-nitroxide derivative comprising the following general structure:wherein R is a nitroxide/nitroxyl radical and L is an optional linker, wherein x,y>0. 2. The heparin-nitroxide derivative of claim 1, wherein the linker L comprises natural amino acids or short peptides residues, amide bond —C(O)NR′— or diamide bonds separated by hydrocarbon chain —C(O)NR′—(CH2)m—C(O)NR′— (CH2)n— or —C(O)NR′—(CH2)m—NR′C(O)—(CH2)n— or by any other suitable linker, R′ is a hydrogen or an alkyl substituent, wherein m, n≦0. 3. The heparin-nitroxide derivative of claim 2, wherein n is an integer from 0 to 2, and m is an integer from 1 to 6. 4. The heparin-nitroxide derivative of any of the preceding claims, wherein the nitroxide is a 5- or 6-atom N-heterocycle selected from the group consisting of piperidine, tetrahydropyridine, pyrroline, pyrrolidine, imidazoline, imidazolidine or oxazolidine. 5. The heparin-nitroxide derivative of any of the preceding claims, wherein more than 50% of the possible carboxyl and/or amino group binding sites at the glycoaminoglycan backbone of heparin are occupied by nitroxides/nitroxyl radicals. 6. The heparin-nitroxide derivative of any of the preceding claims, wherein the nitroxide is TEMPO (2,2,6,6-tetramethylpiperidine-1-oxyl). 7. The heparin-nitroxide derivative of any of the preceding claims, wherein more than 70% of the disaccharides of heparin are modified by TEMPO (2,2,6,6-tetramethylpiperidine-1-oxyl). 8. The heparin-nitroxide derivative of any of the preceding claims, wherein the agent comprises one of the following structures (I)-(IV):wherein R═R1-11 may be one of the following residues:R1=2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl,R2=3-amino-2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl,R3=4-alkyloxycarbonyl-2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl,R4=4-hydroxyimino-2,2,6,6-tetramethyl-1-oxylpiperidin-3-yl,R5=2,2,5,5-tetramethyl-1-oxylpyrrolidin-3-yl,R6=2,2,6,6-tetramethyl-1-oxylpiperidin-4-diyl,R7=2,2,5,5-tetramethyl-1-oxylpyrrolidin-3-diyl,R8=2,2,6,6-tetramethyl-1-oxyl-1,2,5,6-tetrahydropyridin-4-yl,R9=4-acetylamino-2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl,R10°=2,2,5,5-tetramethyl-1-oxylpyrrolin-3-yl,R11=2,2,5,5-tetramethyl-4-bromo-1-oxylpyrrolin-3-yl. 9. The heparin-nitroxide derivative of any of the preceding claims, wherein the agent comprises one of the following structures (I)-(III):wherein R═R1-11 may be one of the following residues:R1=2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl,R2=3-amino-2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl,R3=4-alkyloxycarbonyl-2,2,6,6-tetramethyl-1-oxylpiperidin-4-yl,R4=4-hydroxyimino-2,2,6,6-tetramethyl-1-oxylpiperidin-3-yl,R5=2,2,5,5-tetramethyl-1-oxylpyrrolidin-3-yl,R6=2,2,6,6-tetramethyl-1-oxylpiperidin-4-diyl,R7=2,2,5,5-tetramethyl-... | Johannes Gutenberg-Universität Mainz | A61K049/12 C08B037/10 A61K031/727 A61P009/00 A61P025/28 A61P029/00 A61P035/00 A61P003/10 A61P009/12 A61P009/10 A61P019/02 A61B005/055
US20100280350 | A1 | 2010-11-04 | Chinese medicine tele-diagnostics and triage system | 2009-05-02 | The present invention is tele-diagnostics and patient triage method and system for the practices of traditional Chinese medicine (TCM) that acquire patient health condition info through 4 diagnostics process, provide pathogenesis analysis (disease cause, progression, projection), and treatment and prescription options remotely. This method and system established medical business modules with standard syntax through the implementation of medical decision support tools such as relational database, lookup tables, calculators, decision trees, manifestation reference charts, case comparison and statistics, and visualization in disease, pattern and syndrome identification, differentiation and determination. It essentially transform complex Chinese medicine practices into a networked, expert supported system with a knowledge base that deliver the needed information to assist a TCM physician to make medical decisions. Used in medical call centers and large general hospitals, this method and system help a plural of physicians quickly screen and diagnose a large number of patients and triage patients to different level of health care resources according to disease, condition and care needed for the patient. This system is to provide the method for TCM practitioner to treat patients of common disease with low risks that does not warrant a face to face diagnostics. This system also integrates with patient location based information to establish a quick referral system that connects call centers with clinics and hospitals nearby to provide face to face physician service and maintain continuity of care. A network of this system can alleviate the increasing pressure of resource shortage in rural and remote areas, particularly in developing countries. | 1. A method and a computer software that a Chinese medicine physician can interact with a patient through a plural of medical sensors and a networked computer software program, comprising of: incorporating weather, environmental, family data as part of medical info;gathering patient health condition data through 4 diagnostic methods, namely Auscultation and Olfaction, Inquiry, Pulse-taking and Palpation, together with medical sensors;analyzing cause of disease, progression, and possible outcomes with 8 principles and 6 dialectical methods;determining the severity and risk of patient and whether it is necessary to have a face to face consultation with a physician;if not needed, providing treatment options and prescription recommendations;using Chinese medical body charts and sample cases as references;a database structure of patient electronic health record established around Chinese medicine with both English and Chinese as primary language;if needed, determining location data with communication device and linking the system to clinic or hospital for care continuity and pharmacy or labs for special services;using medical logic modules and tools such as lookup tables, calculators, decision trees, statistics analysis, and visualization tools;determining subhealth conditions and provide a disease management plan. 2. The Observation and Inspection process and software in claims 1 that the physician can retrieve real time patient image for imaging sensors and gather 4 categories (Shen, Se, Xing, Tai) health condition info, comprising of: a plural of observation and inspection targets or areas such as whole body, body parts, and their exterior properties including shape, color etc;a plural of measurement methods and units that determine normal and abnormal and different level of severity;a link to possible disease cause and nature in different database;a reference chart and relation established between an internal organs and body parts to external parts so that an ill internal part may be diagnostic;with references charts and uses cases in database as references to physician;allowing physician with a few clicks to mark and record comments during the process. 3. The Auscultation and Olfaction process and software module in claims 1 that the physician can collect acoustic and odor data from the patients acoustic and odor sensors, or if necessary through face to face or on site inspection, that comprising of: a plural of acoustic and odor sensors;having a list of targets and items such voice and body odor as in a database that the physician will take sample of;having links to possible disease causes and properties and level of disorders; 4. A Inquiry process in claims 1 that a physician can survey the patients through questions and answers over the phone and other communication media such as web or email, that comprising: a list of questions that is related to patient demographic info, life styles, family medical history, personal medical history and p... | - | A61B005/05 A61B005/00 G06Q050/00 G06F017/30
US07833177 | B2 | 2010-11-16 | Breast feeding quantification | 2006-08-29 | Presented is a method for quantifying breastfeeding between a mother and a baby, the method including measuring a physiological volume indicative of stomach fullness volume for the baby; setting a signal threshold value of the physiological volume to correspond to a stomach level that is less than or equal to the stomach fullness volume; obtaining an objective measurement of the physiological volume indicative of a level of fullness of the baby's stomach; and providing an indication to the mother when the objective measurement equals or exceeds the signal threshold value. | 1. A method for quantifying breastfeeding between a mother and a baby, the method comprising: measuring a physiological volume indicative of stomach fullness volume for the baby;setting a signal threshold value of the physiological volume to correspond to a stomach level that is less than or equal to the stomach fullness volume;obtaining an objective measurement of the physiological volume indicative of a level of fullness of the baby's stomach; andproviding an indication to the mother when the objective measurement equals or exceeds the signal threshold value. 2. The method of claim 1, wherein the physiological volume is the fullness level of the baby's stomach. 3. The method of claim 1, wherein the mother has a breast, and wherein the physiological volume is the reduction in fullness level of the mother's breast. 4. The method of claim 1, further comprising providing instructions to the mother to discontinue breastfeeding when the indication is provided to the mother. 5. The method of claim 1, further comprising providing instructions to the mother with respect to correlating the indication with the actions and/or demeanor of the baby. 6. The method of claim 1, wherein measuring includes using a stomach monitor to measure the stomach fullness volume. 7. The method of claim 6, wherein the stomach monitor includes a strain gauge. 8. The method of claim 6, wherein the stomach monitor includes an ultrasonic transducer. 9. The method of claim 1, further comprising using a monitor to obtain the objective measurement. 10. The method of claim 9, wherein the monitor includes a strain gauge. 11. The method of claim 9, wherein the monitor includes an ultrasonic transducer. 12. The method of claim 9, wherein the monitor includes a durometer-measuring device. 13. The method of claim 9, wherein the monitor includes a size-measuring device. 14. The method of claim 13, wherein the size-measuring device measures a linear dimension. 15. The method of claim 13, wherein the size-measuring device measures a surface area. 16. The method of claim 13, wherein the size-measuring device measures a volume. 17. The method of claim 1, further comprising informing the mother that an appropriate nutritional training opportunity is occurring based on the indication. 18. The method of claim 1, wherein the physiological volume indicative of stomach fullness volume is a lowest observed physiological volume indicative of stomach fullness volume. 19. The method of claim 1, wherein the physiological volume indicative of stomach fullness volume is a highest observed physiological volume indicative of stomach fullness volume. 20. The method of claim 1, wherein the physiological volume indicative of stomach fullness volume is an average of observed physiological volumes indicative of stomach fullness volume. 21. The method of claim 1, wherein obtaining an objective measurement is performed concurrently with breastfeeding. | Kimberly-Clark Worldwide, Inc. | A61B005/103
US07840419 | B1 | 2010-11-23 | System and method for reducing the placebo effect in controlled clinical trials | 2009-08-21 | A method and system for performing a clinical trial having a reduced placebo effect is disclosed. The method includes randomizing study participants into three or more treatment groups and performing a first phase of testing on the groups. In a typical embodiment, the first phase of testing includes administering an active treatment to a first group, and administering a placebo to a second group and to a third group. Responders and non-responders are determined for each group. A second phase of testing is then performed. The second phase of testing includes administering the placebo to non-responders in the first group, administering the active treatment to non-responders in the second group, and administering the placebo to non-responders in the third group. The data from the first phase of testing and from the second phase of testing is pooled and analyzed to determine response rates to active treatment and placebo. | 1. A method of performing a clinical trial comprising: performing a first phase of testing comprising: randomizing participants into a first plurality of treatment groups including a first active treatment group and a first placebo group;administering the first active treatment group with an active treatment;administering the first placebo group with a placebo;analyzing the participants for response;sorting the participants, based upon the analyzing the participants for response step, by creating a first non-responders to active treatment group, a first responders to active treatment group, a first non-responders to placebo group, and a first responders to placebo group;creating a non-responder group comprising a combination of one or more members from the first non-responders to active treatment group and one or more members from the first non-responders to placebo group;performing a second phase of testing comprising: randomizing one or more members of the non-responder group into a second active treatment group and a second placebo group;administering the second active treatment group with the active treatment;administering the second placebo group with the placebo;analyzing the randomized one or more members of the non-responder group for response;analyzing, using a processor, the efficacy of the active treatment by; analyzing a response rate of the first phase of testing based on a number of participants in the first responders to active treatment group and a number of participants in the first responders to placebo group;analyzing a response rate of the second phase of testing based on a number of participants in the second active treatment group, determined based on said analyzing the randomized one or more members of the non-responder group for response step, and a number of participants in the second placebo group, also determined based on said analyzing the randomized one or more members of the non-responder group for response step;applying a first weighting factor to the response rate of the first phase of testing;applying a second weighting factor to the response rate of the second phase of testing;wherein the summation of the first and second weighting factor equals a constant; anddetermining the efficacy of active treatment based on a combination of the result of the step of applying a first weighting factor to the response rate of the first phase of testing and the result of the step of applying a second weighting factor to the response rate of the second phase of testing. 2. A method as in claim 1, wherein analyzing a response rate of the first phase of testing based on a number of participants in the first responders to active treatment group and a number of participants in the first responders to placebo group further comprises: analyzing a number of participants in the first non-responders to active treatment group and a number of participants in the first non-responders to placebo group. 3. A method as in claim 1, wherein th... | The General Hospital Corporation | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07837621 | B2 | 2010-11-23 | Computer-aided bone distraction | 2003-08-07 | Systems and methods of generating updated surgical plans are described herein. In one embodiment, a method of generating an updated surgical plan can include generating a three dimensional (3D) model of a bone, generating a surgical plan based on the 3D model, in which the surgical plan can include: locations on the bone upon which to dispose a fixator and settings of struts of the fixator, disposing the fixator on the bone based on the surgical plan, and, based on data associated with the placement of the fixator disposed on the bone, generating an updated surgical plan including updated settings for the struts. | 1. Utilizing a computer system comprising at least one input device, at least one output device, at least one storage device, and at least one suitably-programmed computer, a method of generating and carrying out an updated surgical plan for application to an anatomical bone, the method comprising: (a) generating in at least one suitably-programmed computer a three dimensional (3D) model of the anatomical bone,wherein the 3D model is based on:(1) one or more two dimensional (2D) images of the anatomical bone, the said 2D images being input into the computer system by means of at least one input device, and(2) a 3D bone template, chosen from a set of 3D bone templates stored in at least one storage device, to which a free form deformation is applied,wherein the 3D model is generated in the at least one suitably-programmed computer by successively comparing at least one 2D anatomical bone image with at least one view of the 3D bone template, computing an associated error, and optimizing 3D bone template free form deformation parameters to minimize said error,(b) based on the 3D model, generating in at least one suitably-programmed computer a surgical plan including: (1) one or more locations on the anatomical bone upon which to dispose one or more fixators, and(2) one or more settings of one or more struts of the one or more fixators,(c) outputting the surgical plan by means of at least one output device;(d) based on the surgical plan, disposing the one or more fixators on the anatomical bone,(e) based on data associated with the placement of the one or more fixators disposed on the anatomical bone, the said data being input into the computer system by means of at least one input device, generating in at least one suitably-programmed computer an updated surgical plan including updated settings for one or more of the one or more struts;(f) outputting the updated surgical plan by means of at least one output device; and(g) based on the updated surgical plan, adjusting at least one setting of a strut of at least one fixator. 2. The method of claim 1, wherein the one or more locations include one or more locations for one or more of an osteotomy and a coricotomy. 3. The method of claim 1, wherein the one or more locations on the bone upon which to dispose the one or more fixators include one or more locations on the bone upon which to dispose one or more pins for receiving the one or more fixators. 4. The method of claim 1, wherein the one or more settings of the one or more struts include one or more periodic adjustments of the one or more struts. 5. The method of claim 1, wherein the surgical plan further includes one or more surgical tools for disposing the one or more fixators on the bone. 6. The method of claim 1, wherein the data associated with the placement of the one or more fixators disposed on the bone include one or more images of the one or more fixators disposed on the bone. 7. The method of claim 6, wherein the one or more images i... | Carnegie Mellon University | A61B005/00 G06G007/48
US07835928 | B2 | 2010-11-16 | Connecting consumers with service providers | 2009-08-26 | A request is received from a consumer of services to consult with a service provider having a service provider profile that satisfies at least some attributes in a set of attributes that define a suitable service provider; an available service provider satisfying at least some of the attributes in the set of attributes is identified; and a communication channel is provided to establish a communication between the consumer of services and the identified service provider. | 1. A computer-implemented method comprises: receiving a request from a consumer of services to consult with a service provider having a service provider profile that satisfies at least some attributes in a set of attributes that define a suitable service provider;identifying, by one or more computers, one or more available service providers satisfying at least some of the attributes in the set of attributes;generating, by one or more computers, a graphical user interface, with the graphical user interface when displayed on a display device, rendering: a visual representation corresponding to at least one of the one or more identified service providers; anda link through which the consumer selects one or more of the identified service providers to cause a communication channel to be established between the consumer of services and the selected service provider; andreceiving data indicating the consumer's selection of an available service provider. 2. The method of claim 1 wherein identifying one or more available service providers comprises: monitoring availability of the one or more identified service providers. 3. The method of claim 1 further comprising: activating the communication channel in response to selection of the link. 4. The method of claim 1 wherein the graphical user interface comprises: a first portion providing a listing of service providers having a service provider profile that satisfies at least some attributes in the set of attributes that define the suitable service provider;a second portion providing information associated with the selected service provider;a third portion providing a current status of the selected service provider; anda fourth portion providing one or more links through which the consumer activates the communication channel with the selected service provider. 5. The method of claim 1 wherein the selected service provider is of a first type of service provider, and the method further comprises: identifying, by one or more computers, a second available service provider of a second type of service provider satisfying at least some of the attributes in the set of attributes; andgenerating, by one or more computers, a second graphical user interface, with the second graphical user interface when displayed on a display device, rendering: a visual representation corresponding to at least one or more of the second identified service providers; anda link through which the consumer selects one or more of the second identified service providers to cause a second communication channel to be established between the consumer of services and the second selected service provider. 6. The method of claim 5 wherein if the communication channel with the first selected service provider is established before the second communication channel, the method further comprises: providing information describing an outcome of a communication between the consumer of services and the first selected service provider to the second selec... | American Well Corporation | G06Q010/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07844325 | B2 | 2010-11-30 | Bioelectricity impedance measuring device, a malnutrition measurement system, a malnutrition measurement method | 2008-06-04 | A bioelectricity impedance measuring device includes a contact surface capable of being placed on a part of the body, a gripper which is formed to be easily grasped by a single hand, and electrodes on the contact surface. In addition, the bioelectricity impedance measuring device further includes a malnutrition measuring device which measures malnutrition. An operation and control device measures Phase angle theta and/or an RcXc ratio, and measures a nutrient state. The malnutrition measuring device measures the state of malnutrition from a value of Phase angle theta. A display shows the results of measurement with a bar classified by display of a color. In addition, the operation and control device further measures muscular volume. The operation and control device, by being capable of being provided data or regression, can also measure muscular volume for a child aged 5 or younger or an adult aged 60 or older. | 1. A bioelectricity impedance measuring device, which comprises: a housing with a contact surface capable of being placed on a part of the body;a handheld gripper configured to be easily grasped;electrodes facing the same direction from the contact surface of the housing;an operation and control device configured to measure a state of malnutrition from a phase angle theta and/or an RcXc ratio eta by a malnutrition measurement program; anda display which shows the measured state of malnutrition,wherein the operation and control device is configured to store values of Phase angle theta and/or an RcXc ratio eta in a data table as threshold values of Phase angle theta and/or values of RcXc ratio eta to measure a state of malnutrition from Phase angle theta and/or an RcXc ratio eta,wherein the values of Phase angle theta and/or RcXc ratio eta are equivalent to values of an upper arm circumference corresponding to several states of malnutrition determined by the Mid-Upper Arm Circumference (MUAC) method, andwherein the operation and control device thereby is configured to measure a state of malnutrition by comparing values of Phase angle theta and/or RcXc ratio eta calculated from measurement with the threshold values of Phase angle theta and/or RcXc ratio eta stored in the data table. 2. The bioelectricity impedance measuring device according to claim 1, wherein the gripper forms a concave surface, andthe electrodes facing from the contact surface are located at both ends of the concave surface of the gripper. 3. The bioelectricity impedance measuring device according to claim 1, wherein the display shows a color-coded indication of a judged result of the state of malnutrition with a bar classified by color, wherein the display is configured to provide a numerical value of the state of malnutrition along with the color-coded indication of the state of malnutrition, andwherein the display is configured to include a further notification to warn of serious malnutrition and thus a danger of death when an acute state of malnutrition has been measured. 4. The bioelectricity impedance measuring device according to claim 1, wherein the operation and control device is configured to further measure muscular volume by a muscular volume measuring program. 5. The bioelectricity impedance measuring device according to claim 4, wherein the operation and control device is configured to measure muscular volume for a child aged 5 or younger or an adult aged 60 or older, the operation and control device being provided with at least one of data from the data table or regression analysis. 6. The bioelectricity impedance measuring device according to claim 4, wherein the operation and control device is configured to measure a state of malnutrition and/or measure a muscular volume by individual parameters with sex, height and age. 7. The bioelectricity impedance measuring device according to claim 1, wherein the operation and control device is configured to measure th... | Tanita Corporation | A61B005/06
US20100298739 | A1 | 2010-11-25 | Multifunctional Medical Access Device | 2010-08-02 | A medical access device is provided including a syringe needle and a safety shield supportable on the syringe needle. The safety shield includes a pair of spaced legs each having a distal segment and a proximal segment, each of the proximal segments having a proximal end fixed with respect to the syringe needle and a distal end hingedly attached to the distal segment at a hinge member, each proximal segment including a camming surface. The safety shield also includes a nose member connected to a distal end of each distal segment of the pair of spaced legs. The safety shield also includes a trigger being movable to move the camming member into engagement with the camming surfaces of the proximal segments to effect movement of the legs between a first position and a second position. | 1. A medical access device, comprising: a syringe needle having a proximal end and a distal end configured for percutaneous use; anda safety shield supportable on the syringe needle, the safety shield including: a pair of spaced legs each having a distal segment and a proximal segment, each of the proximal segments having a proximal end fixed with respect to the syringe needle and a distal end hingedly attached to the distal segment at a hinge member, each proximal segment including a camming surface;a nose member connected to a distal end of each distal segment of the pair of spaced legs, wherein the nose member is translatably disposed on the syringe needle, wherein the nose member defines a blunt tip; anda trigger having a camming member and being movable to move the camming member into engagement with the camming surfaces of the proximal segments to effect movement of the legs between: a first position in which the nose member and the distal ends of the distal segments shield at least a distal end of the syringe needle; anda second position in which the distal end of the syringe needle is shielded by the nose member and capable of being at least partially exposed from a distal end of the nose member when the nose member is moved in a proximal direction relative to the syringe needle. 2. The medical access device according to claim 1, wherein in the first position, the distal and proximal segments are substantially linearly aligned with the hinge member of each leg positioned adjacent the needle such that when a force, acting in a substantially linear proximal direction, is applied to a distal end of the nose member of the safety shield, the legs are retained in the first position. 3. The medical access device according to claim 1, wherein the camming member of the trigger and the camming surfaces of the proximal segments are positioned and configured such that movement of the camming member of the trigger into engagement with the camming surfaces of the proximal segments splays the hinge member of each of the legs outwardly to retract the nose member in a proximal direction relative to the syringe needle. 4. The medical access device according to claim 1, wherein each of the legs defines a channel dimensioned to receive the syringe needle. 5. The medical access device according to claim 1, further comprising a needle hub configured for selective connection with a complementary feature of a fluid source, wherein the trigger is pivotally secured to the needle hub. 6. The medical access device according to claim 5, wherein the proximal end of the syringe needle is pivotably connected to the needle hub, and is axially movable with respect to the needle hub. 7. The medical access device according to claim 6, wherein the needle hub includes a locking feature configured to maintain the syringe needle in a retracted position. 8. The medical access device according to claim 1, further comprising a biasing member operatively associated with a le... | TYCO HEALTHCARE GROUP LP | A61M005/32 A61B005/153
US20100324432 | A1 | 2010-12-23 | METHOD AND DEVICE TO DETECT EATING, TO CONTROL ARTIFICIAL GASTRIC STIMULATION | 2007-02-21 | In a method and device for detecting the intake of food in a subject at least one parameter related to the blood flow and/or perfusion of a blood vessel and/or an organ in the digestive system of a patient is monitored by a sensor attached to, or in, a blood vessel or organ of the digestive system. The value of each monitored parameter is analyzed and may be used to control the activity of a gastric stimulator. | 1.-9. (canceled) 10. A device to detect intake of food by a subject, comprising: a sensor that senses a parameter indicative of blood flow to the digestive system of a subject, and that emits a sensor output signal representing said parameter; anda sensor carrier to which said sensor is mounted, said sensor carrier being configured for intracorporeal placement in the subject at a location allowing said sensor to sense said parameter indicative of blood flow to the digestive system of the subject. 11. A device as claimed in claim 10 wherein said sensor measures perfusion of an organ of the digestive system of the subject, and wherein said sensor carrier is configured to place said sensor on or adjacent to the digestive system of the subject. 12. A device as claimed in claim 10 wherein said sensor measures oxygen saturation in said blood flow, and wherein said sensor carrier is configured to place said sensor on or adjacent to the digestive system of the subject. 13. A device as claimed in claim 10 wherein said sensor is selected from the group consisting of photoplethysmographs, resistance sensors, impedance sensors, strain gauges, expandable cuffs, ultrasound detectors, magnetic flow detectors, vortex counters, elbow flow meters, thermistor flow meters and calorimetric flow meters. 14. A device as claimed in claim 10 wherein said sensor carrier is configured for placement at a location selected from the group consisting of attached to the stomach of the subject, in a blood vessel supplying blood to an organ of the digestive system of the subject, and in a blood vessel that accepts blood from an organ of the digestive system of the subject. 15. A gastric stimulation system comprising: a sensor that senses a parameter indicative of blood flow to the digestive system of a subject, and that emits a sensor output signal representing said parameter;a sensor carrier to which said sensor is mounted, said sensor carrier being configured for intracorporeal placement in the subject at a location allowing said sensor to sense said parameter indicative of blood flow to the digestive system of the subject; anda gastric stimulator configured for intracorporeal implantation in the body of the subject, said gastric stimulator being configured to interact with said digestive system of the subject to artificially apply an electrical stimulation thereto, said gastric stimulator being in communication with said sensor and being configured to administer said electrical stimulation dependent on said sensor output signal. 16. A gastric stimulation system as claimed in claim 15 wherein said gastric stimulator is configured to perform a function selected from the group consisting of initiating said electrical stimulation, modifying said electrical stimulation, and stopping said electrical stimulation, dependent on said sensor output signal. 17. A method for detecting intake of food in a subject, comprising the steps of: intracorporeally implanting a sensing device... | - | A61B005/026
US20100331731 | A1 | 2010-12-30 | BLOOD AND INTERSTITIAL FLUID SAMPLING DEVICE | 2010-07-29 | A device and method for lancing a patient, virtually simultaneously producing and collecting a small fluid sample from a body. The device comprises a blood collection system including a lancing needle, drive mechanism, kneading or vibration mechanism, optional suction system, and sample ejection mechanism. The device is preferably sized to be hand-held in one hand and operable with one hand. The device can optionally contain integral testing or analysis component for receiving the sample and providing testing or analysis indication or readout for the user. In one configuration, the lancing needle may be withdrawn very slightly from the point of maximum penetration to create an opening in which blood can pool before being suctioned through to device. | 1. An apparatus for obtaining a sample of body fluid, comprising: a skin piercing member; anda drive mechanism connected to the skin piercing member;wherein the drive mechanism is adapted to extend the skin piercing member to a point of maximum penetration beneath a skin surface and withdraw the skin piercing member from the point of maximum penetration while maintaining the skin piercing member beneath the skin surface to create an opening in which body fluid can pool. 2. The apparatus of claim 1, further comprising: a double stop system, wherein a first stop defines the point of maximum penetration and a second stop defines the point where the skin piercing member creates the opening in which body fluid can pool. 3. The apparatus of claim 1, wherein the skin piercing member includes a fluid passage, the apparatus further comprising: a suction mechanism adapted to create suction to draw body fluid through the fluid passage. 4. The apparatus of claim 1, further comprising: a sample collection chamber connected to the drive mechanism, wherein the drive mechanism is reusable and the sample collection chamber is disposable. 5. The apparatus of claim 1, wherein the skin piercing member is disposable and the drive mechanism is reusable. 6. The apparatus of claim 1, wherein the point of maximum penetration is at a depth of 0.3 to 0.6 millimeters below the skin surface. 7. The apparatus of claim 1, further comprising: a penetration depth adjustment mechanism adapted to adjust the depth to which the skin piercing member penetrates beneath the skin surface. 8. The apparatus of claim 1, further comprising: a needle guard connected to the drive mechanism, the needle guard adapted to contact the skin surface and inhibit the skin piercing member from contacting the skin while the skin piercing member is in a retracted position. 9. A method, comprising: extending a skin piercing member of a fluid sampling device to a point of maximum penetration beneath a skin surface with a drive mechanism; andwithdrawing the skin piercing member from the point of maximum penetration with the drive mechanism while maintaining the skin piercing member beneath the skin surface and forming an opening between the skin piercing member and the skin in which body fluid can pool. 10. The method of claim 9, further comprising: moving fluid from the opening between the skin piercing member and the skin through a passage in the skin piercing member. 11. The method of claim 9, further comprising: forming a vacuum; anddrawing fluid from the opening between the skin piercing member and the skin using the vacuum. 12. The method of claim 9, further comprising: moving fluid from the opening between the skin piercing member and the skin to a sample collection chamber connected to the drive mechanism. 13. The method of claim 9, wherein said extending results in the skin piercing member reaching a maximum depth of 0.3 to 0.6 millimeters below the skin surface. 14. The method of clai... | - | A61B005/151
US20100331656 | A1 | 2010-12-30 | SYSTEMS AND METHODS FOR PROCESSING, TRANSMITTING AND DISPLAYING SENSOR DATA | 2010-09-10 | Systems and methods for continuous measurement of an analyte in a host are provided. The system generally includes a continuous analyte sensor configured to continuously measure a concentration of analyte in a host and a sensor electronics module physically connected to the continuous analyte sensor during sensor use, wherein the sensor electronics module is further configured to directly wirelessly communicate displayable sensor information to a plurality of different types of display devices. | 1. A method for customizing displayable sensor information that is transmitted to display devices, the method comprising: measuring a concentration of an analyte in a host using a continuous analyte sensor;processing a data stream associated with the analyte concentration measured by the continuous analyte sensor to generate sensor data based on the data stream;wirelessly transmitting the sensor data to a key fob device; anddisplaying a first portion of the sensor data on a user interface of the key fob device. 2. The method of claim 1, further comprising authenticating a wireless connection between the continuous analyte sensor and the key fob device. 3. The method of claim 1, wherein wirelessly transmitting the sensor data to a key fob device comprises wirelessly transmitting the sensor data to a key fob device using a standardized wireless protocol. 4. The method of claim 3, wherein the standardized wireless protocol is a Bluetooth protocol. 5. The method of claim 1, further comprising: wirelessly transmitting sensor data to a secondary display device; anddisplaying a second portion of the sensor data on a user interface of the secondary display device. 6. The method of claim 5, further comprising authenticating a wireless connection between the key fob device and the secondary display device. 7. The method of claim 5, wherein the secondary display device comprises a phone. 8. The method of claim 5, wherein wirelessly transmitting sensor data to a secondary display device comprises wirelessly transmitting sensor data to a secondary display device using a standardized wireless protocol. 9. The method of claim 8, wherein the standardized wireless protocol is a Bluetooth protocol. 10. The method of claim 1, wherein the key fob device is sized to fit on a key chain. 11. The method of claim 1, further comprising storing sensor data on a data storage device of the key fob device. 12. The method of claim 1, wherein processing the data stream is performed on a sensor electronics module physically connected to the continuous analyte sensor during operation of the continuous analyte sensor, and wherein sensor data comprising displayable sensor information is generated. 13. The method of claim 12, wherein wirelessly transmitting sensor data to the key fob device comprises wirelessly transmitting sensor data comprising the displayable sensor information to the key fob. 14. The method of claim 1, further comprising processing sensor data on the key fob device to provide displayable sensor information. 15. The method of claim 14, wherein displaying the first portion of the sensor data on the user interface of the key fob device comprises displaying a first portion of the displayable sensor information. 16. The method of claim 1, further comprising receiving user input on the key fob device. 17. The method of claim 16, wherein the user interface of the key fob device comprises one or more buttons configured to accept the user input. 18. The ... | DexCom, Inc. | A61B005/00 G08B023/00
US07853308 | B2 | 2010-12-14 | System and method for patient positioning for radiotherapy in the presence of respiratory motion | 2005-02-11 | A system and method for positioning a patient for radiotherapy is provided. The method comprises: acquiring a first x-ray image sequence of a target inside the patient at a first angle; acquiring first respiratory signals of the patient while acquiring the first x-ray image sequence; acquiring a second x-ray image sequence of the target at a second angle; acquiring second respiratory signals of the patient while acquiring the second x-ray image sequence; synchronizing the first and second x-ray image sequences with the first and second respiratory signals to form synchronized first and second x-ray image sequences; identifying the target in the synchronized first and second x-ray image sequences; and determining three-dimensional (3D) positions of the target through time in the synchronized first and second x-ray image sequences. | 1. A method for positioning a patient for radiotherapy, comprising: acquiring a first x-ray image sequence of a target inside the patient by projecting a radiation beam from a linear accelerator positioned at a first angle;acquiring first respiratory signals of the patient while acquiring the first x-ray image sequence;acquiring a second x-ray image sequence of the target by projecting a radiation beam from the linear accelerator positioned at a second angle;acquiring second respiratory signals of the patient while acquiring the second x-ray image sequence, wherein the second x-ray image sequence and second respiratory signals are acquired after the first x-ray image sequence and first respiratory signals;retrospectively synchronizing the first and second respiratory signals with each other to form synchronized first and second x-ray image sequences;identifying the target in the synchronized first and second x-ray image sequences; anddetermining three-dimensional (3D) positions of the target through time in the synchronized first and second x-ray image sequences. 2. The method of claim 1, wherein the first and second x-ray image sequences are two-dimensional (2D) image sequences. 3. The method of claim 1, wherein the first angle is 0 degrees and the second angle is 90 degrees. 4. The method of claim 1, wherein the step of synchronizing the first and second respiratory signals, comprises: resolving a time delay between the first and second respiratory signals. 5. The method of claim 1, wherein the synchronized first and second x-ray image sequences represent the first and second x-ray image sequences if they were acquired at the same time. 6. The method of claim 1, wherein the step of identifying the target in the synchronized first and second x-ray image sequences, comprises: selecting the target from one image of the synchronized first and second x-ray image sequences; andtracking the target in the remaining images of the synchronized first and second x-ray image sequences. 7. The method of claim 1, wherein the step of determining the 3D position of the target through time, comprises: extracting the target from the synchronized first and second x-ray image sequences; andperforming a triangulation using the first and second angles. 8. The method of claim 1, further comprising: matching the 3D positions of the target from the synchronized first and second x-ray image sequences with four-dimensional (4D) positions of the target acquired from a computed tomography (CT) scan of the patient. 9. The method of claim 8, further comprising: generating a graphical user interface (GUI) having a first window for displaying the 3D positions of the target that match the 4D positions of the target. 10. The method of claim 9, wherein the GUI includes a second window for displaying a metric representing a difference between the 3D positions of the target and the 4D positions of the target. 11. The method of claim 8, further comprising: generating a firs... | Siemens Medical Solutions USA, Inc. | A61B005/05
US20100280334 | A1 | 2010-11-04 | PATIENT STATE DETECTION BASED ON SUPPORT VECTOR MACHINE BASED ALGORITHM | 2010-01-26 | A patient state is detected with at least one classification boundary generated by a supervised machine learning technique, such as a support vector machine. In some examples, the patient state detection is used to at least one of control the delivery of therapy to a patient, to generate a patient notification, to initiate data recording, or to evaluate a patient condition. In addition, an evaluation metric can be determined based on a feature vector, which is determined based on characteristics of a patient parameter signal, and the classification boundary. Example evaluation metrics can be based on a distance between at least one feature vector and the classification boundary and/or a trajectory of a plurality of feature vectors relative to the classification boundary over time. | 1. A method comprising: generating a signal based on a sensed parameter of a patient;determining a plurality of feature vectors over time based on the signal;applying a support vector machine based algorithm to classify a patient state based on the plurality of feature vectors, wherein the support vector machine algorithm based algorithm defines a classification boundary in a feature space;determining a trajectory of the feature vectors within the feature space relative to the classification boundary; andgenerating an indication based on the trajectory of the feature vectors within the feature space. 2. The method of claim 1, wherein generating the indication comprises determining an evaluation metric for the patient state based on the trajectory of the feature vectors within the feature space. 3. The method of claim 2, wherein determining the evaluation metric comprises determining the evaluation metric based on a number of feature values approaching the classification boundary within a predetermined range of time. 4. The method of claim 2, wherein determining the evaluation metric comprises: determining a distance between at least one of the feature values and the classification boundary; anddetermining the evaluation metric based on the distance. 5. The method of claim 4, wherein the distance comprises at least one of a mean, median or lowest distance of distances between at least two of the feature values and the classification boundary. 6. The method of claim 4, wherein the distance comprises the distance between the classification boundary and the feature value determined based on a most recent segment of the signal. 7. The method of claim 2, wherein determining the evaluation metric comprises determining a number of consecutive feature vectors that define a trajectory towards the classification boundary. 8. The method of claim 2, wherein determining the evaluation metric comprises determining a number of feature vectors within the trajectory that are less than a threshold distance away from the classification boundary. 9. The method of claim 1, further comprising controlling delivery of therapy to the patient based on the indication. 10. The method of claim 9, wherein controlling delivery of therapy to the patient comprises at least one of deactivating, activating or adjusting therapy delivery to the patient if the feature vectors define the trajectory toward the classification boundary over time. 11. The method of claim 10, wherein at least one of deactivating, activating or adjusting therapy delivery to the patient if the feature vectors define the trajectory toward the classification boundary over time comprises at least one of deactivating, activating or adjusting therapy delivery to the patient if a threshold number of feature vectors for consecutive segments of the signal define the trajectory toward the classification boundary over time. 12. The method of claim 9, wherein controlling delivery of therapy to the patient bas... | Medtronic, Inc. | A61B005/00 A61N001/02
US08460215 | B2 | 2013-06-11 | Systems and methods for predicting potentially difficult intubation of a subject | 2010-06-08 | Systems and methods for predicting potential difficult intubation of a subject acquire, using a facial structure analysis system, facial structure data of the subject. The systems and methods also extract, using facial structure analysis software of the facial structure analysis system, at least one facial structure variable value from the facial structure data and obtain a predicted potential difficult intubation value using an airway classification model and the at least one facial structure variable value. The systems and methods generate a report based on the predicted potential difficult intubation value. | 1. A non-transitory computer-readable storage medium having facial structure analysis software code stored therein that, when executed by a processor, causes the processor to perform analysis for predicting potential difficult intubation of a subject including the steps of: a) acquiring facial structure data of the subject;b) extracting at least one facial structure variable value from the facial structure data;c) obtaining a predicted potential-difficult-intubation value by inputting the at least one facial structure variable value into an airway classification model; andd) generating a report based on the predicted potential difficult intubation value;wherein step b) includes extracting at least three facial structure variable values from the facial structure data, and step c) includes obtaining the predicted potential difficult intubation value by inputting the at least three facial structure variable values into the airway classification model; andwherein the at least three facial structure variable values include a face-brow-nose-chin ratio, at least one jaw-neck slope high and low value, and at least one nose-tilt down and up value. 2. The computer-readable storage medium of claim 1, further comprising the step of measuring a thyromental distance of the subject, and wherein step c) includes obtaining the predicted potential difficult intubation value by inputting the at least three facial structure variable values and the thyromental distance into the airway classification model. 3. The computer-readable storage medium of claim 1, wherein the at least one facial structure variable value is an eigenface-transformed variable value. 4. The computer-readable storage medium of claim 1, wherein the predicted potential difficult intubation value is a predicted likelihood of the subject being difficult to intubate. 5. A non-transitory computer-readabe storage medium having facial structure analysis software code stored therein that, when executed by a processor, causes the processor to perform analysis for predicting potential difficult intubation of a subject including the steps of: a) acquiring facial structure data of the subject;b) extracting at least one facial structure variable value from the facial structure datac) obtaining a predicted potential-difficult-intubation value by inputting the at least one facial structure variable value into an airway classification model; andd) generating a report based on the predicted potential difficult intubation value;wherein the airway classification model is a quadratic logit model defined by the equation:L⁡(z)=11+ⅇ-zwhere L(z) is the predicted potential difficult intubation value, and:z=β0-(∑i=1n⁢ςi⁢(xi-αi)22⁢σi2)where n is the number of facial structure variable values extracted from the facial structure data in step b), x1 is the at least one facial structure variable value, αi represents an apex of a quadratic curve of the quadratic logit model, σi represents a steepness of sides of the quadrat... | The Brigham and Women's Hospital, Inc. | A61B005/103 A61B005/107
US08267883 | B2 | 2012-09-18 | Photocatalytic implant having a sensor | 2010-06-17 | An implant comprises a photocatalytic layer on at least one surface and is adapted to act as a sensor. In some embodiments, the photocatalytic layer is a semiconductor oxide that is doped. According to some embodiments, the implant comprises a wave guide. According to some embodiments the implant comprises a light port. According to some embodiments, the implant comprises a reflective material on a surface of the waveguide. According to some embodiments the implant comprises a composite material comprising a first material that has a transmissivity when exposed to a predetermined wavelength of light and a second material that has photocatalytic activity when exposed to the predetermined wavelength of light. According to some embodiments the implant comprises a light source adapted to irradiate the photocatalytic surface. | 1. A medical implant comprising: a) photocatalytic surface, andb) a light source adapted to irradiate the photocatalytic surface with light in an amount sufficient to activate photocatalysis thereon,wherein the implant is adapted to act as a sensor. 2. The implant of claim 1 wherein the sensor comprises the photocatalytic surface. 3. The implant of claim 1 wherein the sensor comprises the light source. 4. The implant of claim 3 wherein the sensor comprises a detection system. | DePuy Spine, Inc. | A61M005/00 A61N001/30 A61F002/00 A62B007/08
US08324975 | B2 | 2012-12-04 | Propagation tuned oscillator for orthopedic parameter measurement | 2010-06-29 | A measurement system for capturing a transit time, phase, or frequency of energy waves propagating through a propagation medium (702) is disclosed. The measurement system comprises two different closed-loop feedback paths. The first path includes a transducer driver (726), a transducer (704), a propagation structure (702), a transducer (706), and a zero-crossing receiver (740). The series and parallel resonance of the transducer (704) does not overlap the series and parallel resonance of the transducer (706). A second path includes a transducer driver (1126), a transducer (1104), a propagation medium (1102), a reflecting surface (1106), and an edge-detect receiver (1140). Each positive closed-loop path maintains the emission, propagation, and detection of energy waves in the propagation medium (702, 1102). In either path, a propagation tuned oscillator maintains positive closed-loop feedback of the system that sustains detection, emission, and propagation of energy waves or pulses in a medium. | 1. A propagation tuned oscillator (PTO) to maintain positive closed-loop feedback of energy waves in an energy propagating structure where energy waves propagate through a medium configured to couple to the muscular-skeletal system such that detection of a propagated energy wave initiates an energy wave emission into the medium, where the medium is configured to measure a parameter of the muscular-skeletal system, and where the parameter affects the medium. 2. The PTO of claim 1 where the propagation tuned oscillator tunes a resonant frequency of the energy waves in accordance with physical changes in the energy propagating structure. 3. The PTO of claim 2 where an integer number of energy waves propagates in the medium when positive closed-loop feedback is applied. 4. The PTO of claim 2 where the PTO measures one of transit time, frequency, or phase of a propagated energy wave. 5. The PTO of claim 4 where a parameter to be measured is applied to the medium and where changes in the parameter affect one of transit time, frequency, or phase of a propagated energy wave. 6. The PTO of claim 4 where a positive closed-loop feedback path of the PTO comprises a transducer driver configured to drive at least one transducer for the emission of energy waves into the medium, the energy propagating structure, and a zero-crossing receiver configured to detect the propagated energy wave. 7. The PTO of claim 4 where the positive closed-loop feedback path of the PTO comprises a transducer driver configured to drive at least one transducer for the emission of energy waves into the medium, the energy propagating structure, and a edge-detect receiver configured to detect the propagated energy wave. 8. The PTO of claim 4 where the energy propagating structure comprises: a first transducer;the medium where the first transducer is coupled to the medium at a first location; anda second transducer coupled to a second location of the medium. 9. The PTO of claim 8 where a series and parallel resonance of the first transducer is configured not to overlap a series and parallel resonance of the second transducer. 10. The PTO of claim 4 where the energy propagating structure comprises: a transducer;the medium where the transducer is coupled to the medium at a first location; anda reflecting surface at a second location of the medium. 11. A wireless sensing assembly comprising: a sensor comprising: a first transducer;a medium where the first transducer is coupled to the medium at a first location; anda second transducer where the second transducer is coupled to a second location of the medium and where a series and parallel resonance of the first transducer does not overlap a series and parallel resonance of the second transducer;a propagation tuned oscillator operatively coupled to the sensor to maintain positive closed-loop feedback of energy waves in an energy propagating structure where energy waves propagate through a medium such that detection of a propagated en... |  | G01R027/04 A61B005/103
US08579833 | B2 | 2013-11-12 | Thrombosis and osteoporosis prophylaxis | 2010-06-29 | Use of a device to reduce risk of thrombosis and osteoporosis, which uses a pressure transducer to measure the forefoot force generated by a person with their forefoot, and generates a signal so that the person can recognize whether their forefoot force has reached a minimum value. | 1. A method for reducing risk of thrombosis and osteoporosis, comprising: providing a pressure sensor (1) set to measure forefoot force produced by a person via their forefoot (5) and generate a signal that informs the person whether the forefoot force has reached a predetermined minimum value sufficient for reducing risk of thrombosis or osteoporosis, andhaving said person press against said sensor (1) with the forefoot (5) and generate the signal indicating to said person that the forefoot force has reached the predetermined minimum value sufficient for reducing risk of thrombosis or osteoporosis,wherein the signal is given upon reaching a predetermined cumulative forefoot force. 2. The method according to claim 1, wherein the forefoot force is produced via the ball of the foot and the toes. 3. The method according to claim 1, wherein the pressure sensor (1) is arranged relative to the person that the force generated by the plantarflexion of the foot and transmitted by the forefoot (5) is measured. 4. The method according to claim 1, wherein the forefoot force is generated by contraction of the calf muscles. 5. A method for reducing risk of thrombosis and osteoporosis, comprising: providing a pressure sensor (1) set to measure forefoot force produced by a person via their forefoot (5) and generate a signal that informs the person whether the forefoot force has reached a predetermined minimum value sufficient for reducing risk of thrombosis or osteoporosis, andhaving said person press against said sensor (1) with the forefoot (5) and generate the signal indicating to said person that the forefoot force has reached the predetermined minimum value sufficient for reducing risk of thrombosis or osteoporosis,wherein the pressure sensor (1) is arranged relative to the person so that the force generated by the plantarflexion of the foot and transmitted by the ball of the foot and the toes is measured,wherein said forefoot force corresponds to 30-70% of body weight of the person, andwherein the forefoot force is generated by contraction of the gastrocnemius muscle, the soleus muscle and the toe-flexing muscles. 6. A method for reducing risk of thrombosis and osteoporosis, comprising: providing a pressure sensor (1) set to measure forefoot force produced by a person via their forefoot (5) and generate a signal that informs the person whether the forefoot force has reached a predetermined minimum value sufficient for reducing risk of thrombosis or osteoporosis, andhaving said person press against said sensor (1) with the forefoot (5) and generate the signal indicating to said person that the forefoot force has reached the predetermined minimum value sufficient for reducing risk of thrombosis or osteoporosis,wherein the signal is produced upon reaching a given forefoot force which is programmed based on the weight of the person. 7. The method according to claim 6, wherein said forefoot force corresponds to at least 10% of body weight of the person. ... | - | A61B005/00
US08172752 | B2 | 2012-05-08 | Automatic wireless PAN/LAN switching | 2010-06-28 | A communication system is used by a portable patient monitoring device for connecting to other devices, including a docking station, via one or more networks. The communications system includes an adaptive communication interface which, when the patient monitoring device is docked, establishes communication with a first network, and receives from a first docking station a first unique identifier that identifies the first docking station, allowing the patient monitoring device to communicate patient parameters to the first destination. While the portable patient monitoring device is undocked, the adaptive communication interface is inhibited from establishing communication with a network providing an identifier to the portable patient monitor that does not match the first unique identifier. Upon docking to a second docking station, the patient monitoring device establishes communication with a second network using a new unique identifier received from the second docking station via a second exclusive communication link. | 1. A communication system configured to connect a portable patient monitoring device to a plurality of networks via separate communication links, comprising: a first patient area network including a first docking station and a first destination;a second patient area network, different from the first patient area network, including a second docking station and a second destination; anda portable patient monitoring device including an adaptive communication interface configured to: establish communication with the first patient area network when the portable patient monitoring device is docked to the first docking station, receive a first unique identifier from the first docking station via a first exclusive communication link between the first docking station and the portable patient monitoring device, and transmit patient parameters to the first destination via the first exclusive communication link, andmaintain communication with the first patient area network when the portable patient monitoring device is undocked from the first docking station and the first exclusive communication link is inoperative, by establishing a first wireless communication link between the portable patient monitoring device and the first patient area network and obtaining the first unique identifier from the first patient area network via the first wireless link, and transmit patient parameters via the first wireless communication link to the first destination,when the portable patient monitoring device is undocked and is in an area where the first and second patient area networks overlap and are within range of the patient monitoring device, the adaptive communication interface is configured to maintain connection with the first patient area network, and is actively inhibited from establishing communication with the second network when the second network provides an identifier that does not match the first unique identifier, andwhen the portable patient monitoring device is docked to the second docking station, the adaptive communication interface is configured to establish a second exclusive communication with the second patient area network, receive a second unique identifier from the second docking station via the second exclusive communication link between the second docking station and the portable patient monitoring device, and transmit patient parameters to the second destination. 2. The communication system of claim 1, wherein the first destination is different from the second destination. 3. The communication system of claim 1, wherein the first destination is identical to the second destination. 4. The communication system of claim 1, wherein each of the first and second unique identifiers identifies messages to and from at least one of the first and second network. 5. The communication system of claim 1, wherein each of the first and second unique identifier identifies a geographic location of at least one of the first and second docking station. 6. Th... | Draeger Medical Systems, Inc. | A61B005/00 H04M001/66
US08160671 | B2 | 2012-04-17 | Calibration techniques for a continuous analyte sensor | 2010-09-01 | Disclosed herein are systems and methods for calibrating a continuous analyte sensor, such as a continuous glucose sensor. One such system utilizes one or more electrodes to measure an additional analyte. Such measurements may provide a baseline or sensitivity measurement for use in calibrating the sensor. Furthermore, baseline and/or sensitivity measurements may be used to trigger events such as digital filtering of data or suspending display of data. | 1. A method for continuous glucose sensing, the method comprising: generating a glucose signal associated with a glucose concentration in a host using a continuous glucose sensor;detecting oxygen deprivation; anddetermining, using a processor module, a stability of the continuous glucose sensor based on the oxygen deprivation detection. 2. The method of claim 1, wherein detecting oxygen deprivation comprises examining a signal for oxygen deprivation-induced artifacts. 3. The method of claim 1, wherein the continuous glucose sensor comprises an electrode, and wherein detecting oxygen deprivation comprises switching, cycling or pulsing a voltage applied to the electrode and evaluating, using a processor module, a response to the applied voltage. 4. The method of claim 1, wherein determining a stability further comprises determining a sensitivity change over a time period. 5. The method of claim 4, wherein the sensitivity change is not based on a reference glucose measurement from a reference glucose monitor. 6. The method of claim 1, wherein the glucose signal is generated in vivo. 7. The method of claim 1, further comprising controlling a display of the glucose signal responsive to the stability determination. 8. The method of claim 1, further comprising suspending a display of the glucose signal responsive to the stability determination. 9. A system for continuous glucose sensing, the system comprising: a continuous glucose sensor configured to generate a glucose signal associated with a glucose concentration in a host; andelectronic circuitry configured to determine a stability of the continuous glucose sensor by detecting oxygen deprivation. 10. The system of claim 9, wherein the electronic circuitry is configured to detect oxygen deprivation by examining a signal for oxygen deprivation-induced artifacts. 11. The system of claim 9, wherein the continuous glucose sensor comprises an electrode, and wherein the electronic circuitry is configured to detect oxygen deprivation by switching, cycling or pulsing a voltage applied to the electrode and to evaluate a response to the applied voltage. 12. The system of claim 9, wherein the electronic circuitry is further configured to determine a sensitivity change over a time period. 13. The system of claim 12, wherein the electronic circuitry is configured to determine the sensitivity change without a reference glucose measurement from a reference glucose monitor. 14. The system of claim 9, wherein the electronic circuitry is configured to control a display of the glucose signal responsive to the stability determination. 15. The system of claim 9, wherein the electronic circuitry is configured to suspend a display of the glucose signal responsive to the stability determination. 16. A system for continuous glucose sensing, the system comprising: means for generating a glucose signal associated with a glucose concentration in a host using a continuous glucose sensor;means for detecting oxygen ... | DexCom, Inc. | A61B005/05
US09492103 | B2 | 2016-11-15 | Automatic guidewire maneuvering system and method | 2010-09-02 | System for guiding a catheter through a lumen system of a body of a patient, to a predetermined location within the lumen system, the system including a medical positioning system, a moving mechanism coupled with the catheter, and a controller coupled with the medical positioning system and with the moving mechanism, the medical positioning system including at least one position detector, the position detector being firmly attached to a distal portion of the catheter, the medical positioning system determining the position of the position detector, the controller controlling the operation of the moving mechanism to move the catheter to the predetermined location, according to the position and according to a topological representation of at least a portion of the lumen system. | 1. A system for guiding a catheter to a predetermined location within a lumen system of a body of a patient, comprising: a medical positioning system including a detector located at and attached to a distal portion of said catheter and configured to produce a signal respective of a three-dimensional position of said detector, said medical positioning system being configured to determine the position of said detector in response to said signal as said catheter travels through said lumen system;a moving mechanism coupled with said catheter; andan electronic controller coupled with said medical positioning system and with said moving mechanism configured to control the operation of said moving mechanism to automatically move said catheter according to the determined position of said detector along a preplanned path between an origin and said predetermined location relative to said lumen system, said controller being coupled to a storage unit wherein a topological representation of said lumen system in stored in said storage unit, said controller being configured to determine said preplanned path from said topological representation of said lumen system, said topological representation of said lumen system comprises a mapping of said lumen system;said controller being further configured to determine a modified path different from said preplanned path when said controller determines that said catheter has encountered an obstruction in said lumen system, said modified path being configured to overcome the obstruction;wherein said controller is further configured to determine an orientation of said catheter distal portion at a given point in said lumen system and to determine a computed slope of said preplanned path at said given point wherein said controller is configured to detect said obstruction when said orientation is different from said slope. 2. The system of claim 1 wherein when said controller detects said obstruction, said controller is further configured to control said moving mechanism to move said catheter in a corrective movement. 3. The system of claim 2 wherein said corrective movement includes retreating said catheter backward within said lumen system, performing one of twisting and bending of said distal portion of said catheter, and advancing said catheter along said modified path to thereby overcome said obstruction in said preplanned path. 4. The system of claim 1 wherein said topological representation of said lumen system comprises one of a two-dimensional and a three-dimensional mapping of said lumen system. 5. The system of claim 1 further comprising: an imaging system for detecting an image of a part of said lumen system having said preplanned path along which said catheter is to be guided; anda superimposition processor coupled with said medical positioning system and with said imaging system and a display coupled with said superimposition processor, said superimposition processor superimposing a representation of said ca... | MediGuide Ltd. | A61F011/00 A61B005/06 A61B001/00 A61M025/01 A61M025/09 A61B006/03 A61B006/00 A61B017/00
US07840250 | B2 | 2010-11-23 | Method for neural current imaging | 2002-08-16 | A method for neural current imaging. Electromagnetic fields produced unknown distribution of unknown sources in the body is sensed at a plurality of remote locations. The inverse problem is solved to produce a first fuzzy “image” of the sources. In addition, a standard imaging method, such as MRI, is used to independently image the body, to obtain a second fuzzy image of the sources. Additional independently obtained fuzzy images may also be provided. All or a selected subset of the images are pooled as components to form an enhanced image with greater resolution or clarity than the component images. | 1. A method for imaging electrical current flowing in a body, the current defining a field outside the body, the method comprising: obtaining field data, by use of at least one of (a) EEG and (b) MEG, representative of the field at a selected time;obtaining a solution to the inverse problem posed by said field data;producing, based on said solution, a first image showing the current at a first resolution;scanning the body at the same said time as part of an MRI procedure;producing by said MRI procedure a second image showing the current at a second resolution; andcombining said first and second images so as to produce an enhanced image of the current at a third resolution that is greater than the first and second resolution. 2. The method of claim 1, wherein said step of combining comprises employing Bayesian inference. 3. The method of claim 2, wherein the current is neural current. 4. The method of claim 1, wherein the current is neural current. | Electrical Geodesics, Inc. | A61B005/00
US07860542 | B2 | 2010-12-28 | Analyte detection system with reduced sample volume | 2005-12-21 | In certain embodiments, a method samples a body fluid of a patient. The method includes providing a fluid handling system having one or more fluid passageways. The method further includes infusing an infusion fluid by way of a fluid handling system into a patient through one or more fluid passageways. The method further includes obtaining a sample of body fluid by way of a fluid handling system from a patient through one or more fluid passageways. The obtained sample is no more than 400 microliters in volume. The method further includes analyzing at least an analyzed portion of the obtained sample by way of an analyte detection system operatively associated with the fluid handling system to determine a concentration of at least one analyte. | 1. A body fluid handling system configured to reduce sample dilution, the system comprising: a fluid handling network configured to provide access to a body fluid in a patient, the fluid handling network comprising a patient end, a junction in fluid communication with the patient end, a first fluid passageway connected to the junction, and a second fluid passageway connected to the junction;a body fluid detection sensor configured to detect the presence of substantially undiluted body fluid at a location in the fluid handling network and generate a signal;a pump system coupled to said fluid handling network, said pump system having an infusion mode in which said pump system is operable to infuse an infusion fluid toward said patient via said fluid handling network;one or more air injectors configured to position an air bubble at the entrance to the second fluid passageway adjacent the junction between the first and second passageways;said pump system being configured to: draw at least a portion of a sample of said body fluid from said patient, through the patient end, through the junction, and into the first fluid passageway, wherein the body fluid detection sensor identifies substantially undiluted body fluid at the location in the fluid handling network and generates a signal to facilitate the positioning of substantially undiluted body fluid at the junction between the first and second passageways,draw at least a portion of the sample through the junction and into the second fluid passageway after the air bubble is positioned such that the leading portion of the sample that enters the second fluid passageway immediately following the air bubble is substantially undiluted blood; andmaintain the air bubble immediately forward of the undiluted leading portion of the sample as the sample is transported through the second fluid passageway such that the air bubble provides a barrier to reduce dilution of the leading portion of the sample that enters the second fluid passageway by insulating the leading portion of the sample from dilution that may result from contact with other fluids; andan analyte detection system accessible via said fluid handling network and configured to measure an analyte in an analyzed a portion of said drawn sample of body fluid. 2. A method of sampling a body fluid of a patient and reducing sample dilution, the method comprising: providing a fluid handling system configured to provide access to the body fluid of the patient, the fluid handling system comprising a patient end, a junction in fluid communication with the patient end, a first fluid passageway connected to the junction, and a second fluid passageway connected to the junction;with the fluid handling system, infusing an infusion fluid into the patient through the patient end;positioning air at the entrance to the second fluid passageway adjacent the junction between the first and second passageways;with said fluid handling system, drawing at least a portion of a s... | OptiScan Biomedical Corporation | A61B005/1455
US07846121 | B2 | 2010-12-07 | Device for extracorporeal irradiation of a liquid containing bilirubin, and method therefor | 2004-05-28 | The invention concerns a device for the extracorporeal irradiation of a patient's (P) bodily fluid containing bilirubin, comprising a first (10) and a second (12) line which can be connected to the patient (P), an impermeable irradiation unit for bodily fluids (14) connected therewith and located between the first (10) and the second (12) lines, and at least one adjustable feed unit for bodily fluids (16) located in the first (10) and/or the second (12) line, wherein by means of at least one feed unit for bodily fluids (16) an adjustable flow of bodily fluid through the lines (10, 12) and the irradiation unit for bodily fluids (14) is achievable, in which the first line (10) is designed to continually circulate the bodily fluid drawn from the patient (P) and route it to the irradiation unit for bodily fluids (14), and the second line (12) is designed to continually circulate the irradiated bodily fluid to the patient (P), and in which the irradiation unit for bodily fluids (14) contains an radiation source for the emission of electromagnetic radiation with a wavelength greater than 430 nm, preferably in a spectrum from 450 to 530 nm. The invention furthermore concerns a corresponding method. | 1. Device for extracorporeal irradiation of a bodily fluid from a patient past the onset of puberty, said bodily fluid containing bilirubin, comprising a first and a second line which can be connected to the patient, an impermeable irradiation unit for bodily fluids connected therewith and located between the first and the second lines, and at least one adjustable feed unit for bodily fluids located in the first and/or the second line, wherein a feed unit for bodily fluids is configured to provide an adjustable flow of bodily fluid through the lines and the irradiation unit for bodily fluids in which the first line is configured to continually circulate the bodily fluid drawn from the patient and route it to the irradiation unit for bodily fluids, and the second line is configured to continually circulate the irradiated bodily fluid to the patient, and in which the irradiation unit for bodily fluids contains a radiation source for the emission of electromagnetic radiation with a wavelength greater than 430 nm and an irradiation intensity of greater than 4 μW/(cm2 nm) as applied directly to said bodily fluids in said irradiation unit, the device further comprising a plasma filter or cell separator with a cell-containing side and a cell-free side, wherein a fluid feed inlet to the cell-containing side is connected to a third patient-connectable line, a fluid discharge outlet from the cell-containing side is connected to a fourth patient-connectable line, a fluid discharge outlet from the cell-free side is connected to the first line, and a fluid feed inlet to the cell-free side is connected to the second line, said device configured to irradiate free and bound bilirubin obtained from said bodily fluid, and configured for use with a patient past the onset of puberty. 2. Device according to claim 1, in which a photoisomer-removal unit for the removal of water-soluble photoisomers of bilirubin is located in the second line. 3. Device according to claim 1, in which the irradiation unit for bodily fluids contains a partially translucent flow reservoir. 4. Device according to claim 3, in which the partially translucent flow reservoir is a flow cuvette. 5. Device according to claim 1, in which the photoisomer-removal unit is built as an adsorber, hemofilter and/or a dialyzer. 6. Device according to claim 1, in which the feed unit for bodily fluids is a pump. 7. Device according to claim 1, in which the bodily fluid is blood or blood plasma. 8. A system comprising a device according to claim 1 in combination with a bioartificial liver system. 9. The device according to claim 1, wherein said electromagnetic radiation has a wavelength in the spectrum from 450 to 530 nm. 10. Method for the reduction of bilirubin levels in a patient, comprising: drawing a bodily fluid containing bilirubin from a patient not treatable by phototherapy directed to the skin,processing said bodily fluid to produce a cell-free bodily fluid comprising free and bound bilirubi... | Fresenious Medical Care Deutschland GmbH | A61M037/00 A61N001/30 A61M031/00 A61M001/00 A61B005/00
US20100278745 | A1 | 2010-11-04 | COMPOUNDS FOR FLUORESCENCE IMAGING | 2007-12-21 | Methods and compositions involving enzyme-activatable fluorophore polymer imaging agents for photodetection of specific tissues and/or human diseases and disorders are provided. In certain embodiments, the imaging agent comprises a hydrophilic polymer backbone (e.g., poly(L)lysine), a hydrophobic fluorophore, and a hydrophilic solubilizing agent. The solubilizing agent may comprise a quarternary ammonium group (e.g., a 1-methylnicotinic group) to enhance self-quenching of the fluorophore. Various methods for the generation and purification of imaging agents are also provided, including methods involving solid phase probe extraction or ion exchange purification. | 1. An imaging probe comprising a biocompatible hydrophilic polymer backbone conjugated to a plurality of hydrophobic fluorophores and a plurality of hydrophilic solubilizing moieties; wherein each solubilizing moiety is independently eitherwherein R is independently hydrogen, a heteroatom-substituted or heteroatom-unsubstituted aryl(C1-C10), or aralkyl(C2-C10);further wherein R1, R2, R3, R4 and R5 are each independently hydrogen, hydroxy, amino, cyano, halo, nitro, mercapto, or a heteroatom-substituted or heteroatom-unsubstituted alkyl(C1-C10), aryl(C1-C10), aralkyl(C2-C10), acyl(C1-C10), alkoxy(C1-C10), aryloxy(C1-C10), aralkoxy(C2-C10), acyloxy(C1-C10), alkylamino(C1-C10), arylamino(C1-C10), aralkylamino(C2-C10), amido(C1-C10), or the carbonyl group that conjugates to the hydrophilic polymer backbone, provided that one of R1, R2, R3, R4 and R5 is said carbonyl group that conjugates to the hydrophilic polymer backbone;further wherein R6, R7, R8, R9 and R10 are each independently hydrogen, hydroxy, amino, cyano, halo, nitro, mercapto, or a heteroatom-substituted or heteroatom-unsubstituted alkyl(C1-C10), aryl(C1-C10), aralkyl(C2-C10), acyl(C1-C10), alkoxy(C1-C10), aryloxy(C1-C10), aralkoxy(C2-C10), acyloxy(C1-C10), alkylamino(C1-C10), arylamino(C1-C10), aralkylamino(C2-C10), or amido(C1-C10);further wherein R11, R12, and R13, are each independently a heteroatom-substituted or heteroatom-unsubstituted aryl(C1-C10), aryl(C1-C10), or aralkyl(C2-C10); andfurther wherein at least some of said fluorophores are in fluorescence-quenching or self-quenching permissive positions which are separable by enzymatic cleavage. 2. The imaging probe of claim 1, wherein R is alkyl(C1-6) or n=1-6. 3. The imaging probe of claim 1, wherein the solubilizing moieties comprisewherein R is an alkyl(C1-4). 4. The imaging probe of claim 3, wherein R is —CH3. 5. The imaging probe of claim 1, wherein the fluorophores comprise a near-infrared fluorophore, 1-pyrenebutyric acid (pyrene), NIR-667, NIR-641 N-succinimidyl ester, 11-((5-dimethylaminonaphhalene-1-sulfonyl)amino)undecanoic acid, cis-parinaric acid, BODIPY® 581/591, BODIPY® FL (505/511), BODIPY® 530/550, BODIPY (C1-BODIPY 500/510 C12, C4-BODIPY 500/510 C9, C8-BODIPY 500/510 C5, a distyryl-boradiazaindacene, Dansyl (336/517), Marina Blue®, M12652 Marina Blue®, pacific blue™, NBD (463/536), fluorescein (496/519), Oregon green® 488 (501/526), tetrametylrhodamine (540/566), lissamine rhodamine (560/581), lissamine rhodamine B sulfonyl chloride, texas red® (582/601), dansyl undecanoic acid, 6-(N-(7-nitrobenz-2-oxa-1,3-diazol-4-yl)amino)hexanoic acid, (4-fluoro-7-nitrobenz-2-oxa-1,3-diazole), (1-dodecyl,1′-(4-butyric acid NHS)-3,3,3′,3′-tetramethylindocarbocyanine halide), phthalocyanines (N-Ethyl-N-[5-(N-succininmidyloxycarbonyl)pentyl]-3,3,3,3-tetramethyl-2), 2-indocarbocyanine chloride, 11-Chloro-1,1′-dipropyl-3,3,3′,3′-tetramethyl-10,12-trimethyleneindatricarbocyanine iodide, NIR-664-N-succinimidyl ester, 2-[(E)-2-(2-... | - | A61B005/00 C08G069/48 C08G063/91 G01N033/53 C12Q001/34
US20100280346 | A1 | 2010-11-04 | Analyte Monitoring Device and Methods of Use | 2010-07-15 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for determining a stability of an analyte sensor, the method comprising: exposing an analyte sensor to a biological sample, wherein the analyte sensor comprises a first working electrode that measures an analyte component and a baseline component and a second working electrode that measures a baseline component substantially without an analyte component;receiving sensor data from the analyte sensor, wherein the sensor data comprises first sensor data associated with the first working electrode and second sensor data associated with the second working electrode; anddetermining a stability of the analyte sensor based at least in part on the second sensor data. 3. The method of claim 2, wherein the second sensor data comprises a plurality of time spaced data points, and wherein the step of determining a stability of the analyte sensor comprises evaluating a change in amplitude of two or more of the plurality of time spaced data points. 4. The method of claim 3, wherein the step of exposing the analyte sensor to a biological sample comprises intermittently exposing the analyte sensor to a biological sample and to a solution. 5. The method of claim 4, wherein the step of evaluating comprises comparing a plurality of time spaced points measured when the analyte sensor is exposed to the biological sample. 6. The method of claim 4, wherein the step of evaluating comprises comparing a first point measured when the analyte sensor is exposed to the biological sample to a second point measured when the analyte sensor is exposed to the solution. 7. The method of claim 4, wherein the step of evaluating comprises evaluating transient information of the second sensor data during a step change from exposure of the sensor to the solution to exposure of the sensor to the biological sample. 8. The method of claim 3, further comprising determining a predetermined level of stability when a change in an amplitude of the second sensor data meets one or more criteria. 9. The method of claim 8, further comprising controlling a display associated with the analyte sensor system in response to the determination of the predetermined level of stability. 10. The method of claim 9, wherein the step of controlling a display comprises providing at least one of a numeric estimated analyte value, a directional trend of analyte concentration, and a graphical representation of a plurality of estimated analyte values. 11. The method of claim 9, wherein the step of controlling a display comprises requesting reference analyte data. 12. The method of claim 2, wherein the step of determining a stability of the analyte sensor is performed during a predetermined time period. 13. The method of claim 2, wherein the step of determining a stability of the analyte sensor is performed after a predetermined time period. 14. The method of claim 2, further comprising repeating the step of determining a stability, wherein the step is conducted after a system re-s... | Abbott Diabetes Care Inc. | A61B005/05
US20100331730 | A1 | 2010-12-30 | CIRCUIT BOARD FOR BODY FLUID COLLECTION, METHOD FOR PRODUCING THE SAME, METHOD FOR USING THE SAME, AND BIOSENSOR | 2009-01-21 | A circuit board for body fluid collection includes a plurality of measurement units including a puncture needle and an electrode for making contact with a body fluid collected by puncturing with the puncture needle, the plurality of measurement units being arranged in parallel in a predetermined direction, and a support portion extending along the parallel arrangement direction and supporting the plurality of measurement units, wherein the support portion is rolled so that the plurality of measurement units can be disposed radially. | 1. A circuit board for body fluid collection comprising: a plurality of measurement units including a puncture needle and an electrode for making contact with a body fluid collected by puncturing with the puncture needle, the plurality of measurement units being arranged in parallel in a predetermined direction, anda support portion extending along the parallel arrangement direction and supporting the plurality of measurement units,wherein the support portion can be rolled so that the plurality of measurement units are arranged radially. 2. The circuit board for body fluid collection according to claim 1, wherein the plurality of measurement units extend outwardly with respect to the support portion when the plurality of measurement units are arranged radially. 3. The circuit board for body fluid collection according to claim 1, wherein the plurality of measurement units extend inwardly with respect to the support portion when the plurality of measurement units are arranged radially. 4. The circuit board for body fluid collection according to claim 1, further comprising a reinforcing portion that is disposed in a spaced apart relationship to the support portion, and connects, between the plurality of measurement units, the plurality of measurement units that are adjacent to each other. 5. The circuit board for body fluid collection according to claim 1, further comprising engage portions at one end portion and at the other end portion of the support portion in the parallel arrangement direction for retaining the rolling of the support portion. 6. A circuit board for body fluid collection comprising: a plurality of measurement units including a puncture needle and an electrode for making contact with a body fluid collected by puncturing with the puncture needle, the plurality of measurement units being arranged in parallel in a predetermined direction, anda support portion extending along the parallel arrangement direction and supporting the plurality of measurement units,wherein the plurality of measurement units are arranged radially by bending a boundary between the plurality of measurement units and the support portion, and rolling the support portion. 7. The circuit board for body fluid collection according to claim 6, wherein the plurality of measurement units extend outwardly with respect to the support portion. 8. The circuit board for body fluid collection according to claim 6, wherein the plurality of measurement units extend inwardly with respect to the support portion. 9. A biosensor comprising: a circuit board for body fluid collection comprising: a plurality of measurement units including a puncture needle and an electrode for making contact with a body fluid collected by puncturing with the puncture needle, the plurality of measurement units being arranged in parallel in a predetermined direction, anda support portion extending along the parallel arrangement direction and supporting the plurality of measurement units,wherein... | NITTO DENKO CORPORATION | A61B005/151
US20100286553 | A1 | 2010-11-11 | Devices and Methods for Screening of Vagal Nerve Stimulation | 2009-05-08 | The present disclosure provides systems and methods for screening vagus nerve stimulation to determine the potential efficacy of permanent stimulation systems. In one aspect, the system includes temporary electrode assemblies adapted for temporary placement on or in the body adjacent the vagus nerve. In another aspect, a method is provided to place a stimulating electrode adjacent the posterior of the carotid sheath. | 1. A method of screening vagus nerve stimulation to determine efficacy, the method comprising: placing at least one temporary electrode adjacent the vagus nerve;energizing the electrode to stimulate at least a portion of the vagus nerve;monitoring patient response to stimulation of the vagus nerve;determining whether vagus nerve stimulation had a beneficial effect on the patient condition; andremoving the temporary electrode. 2. The method of claim 1, wherein said placing includes positioning the temporary electrode on the skin of a patient. 3. The method of claim 2, wherein said securing includes adhering at least a portion of the electrode to the skin of a patient. 4. The method of claim 1, wherein said placing includes making initial contact between the patient and the temporary electrode and advancing the temporary electrode toward the vagus nerve. 5. The method of claim 4, wherein the initial contact is on the skin of the patient with a first distance between the temporary electrode and the vagus nerve; and the advancing includes pushing the temporary electrode through the skin in a direction toward the vagus nerve to a second position with a second distance between the temporary electrode and the vagus nerve, the second distance being less than the first distance. 6. The method of claim 5, wherein said securing includes applying a compression member to maintain the electrode in the second position. 7. The method of claim 1, wherein said placing includes forming an opening in the skin of the patient and positioning the temporary electrode beneath the skin of the patient. 8. The method of claim 7, wherein said positioning includes: performing a blunt needle stick adjacent the carotid sheath of the patient;imaging a portion of the electrode to monitor the electrode position within the patient;advancing the temporary electrode along the carotid sheath substantially parallel to the vagus nerve; andsecuring the temporary electrode. 9. The method of claim 8, wherein said advancing includes moving the temporary electrode along the exterior of the carotid sheath. 10. The method of claim 8, wherein said advancing includes moving the temporary electrode within the interior of the carotid sheath. 11. The method of claim 8, wherein said advancing includes providing a blunt catheter, advancing the blunt catheter along the carotid sheath, the electrode associated with the blunt catheter to position the electrode substantially parallel to the vagus nerve. 12. The method of claim 1, wherein said placing includes positioning the electrode adjacent the posterior of a carotid sheath. 13. The method of claim 7, wherein said positioning includes: creating an incision above the clavicle;forming a pocket parallel to and outside of the carotid sheath of the patient;inserting a temporary lead having at least two electrodes disposed thereon; andpositioning the temporary lead electrodes to engage at least a portion of the carotid sheath facing the vagus ne... | - | A61B005/05
US20100286501 | A1 | 2010-11-11 | METHOD AND APPARATUS FOR MONITORING AN OBJECT OF INTEREST IN A MEDICAL SYSTEM | 2008-09-16 | The invention relates to a method and apparatus for monitoring an object of interest (100). According to the invention, a cushion (104) is inflated such that the inflated cushion extends inwards to envelop the object of interest (100), whereupon the pressure exerted on the object of interest (1600) by the inflated cushion (104) is measured and a signal is output to stop inflating when the pressure reaches a predefined threshold. After that, the volume of the inflated cushion is measured and the size of the object of interest (100) is derived from the volume of the inflated cushion (104). In this way, the invention can keep the object of interest still during monitoring. Furthermore, the size of the object of interest can be used for imaging reconstruction to improve the quality of reconstructed images. | 1. An apparatus for monitoring an object of interest (100) comprising: an elastic cushion (104);a pump (102) arranged for inflating the cushion such that the inflated cushion extends in an inward direction so as to wrap itself around the object of interest (100);a pressure sensor (103) arranged for detecting the pressure exerted on the object of interest (100) by the inflated cushion (104) and outputting a signal to the pump (102) to stop inflating when the pressure reaches a predefined threshold;a measuring means (108) arranged for measuring the volume of the inflated cushion (104); anda processor (105) arranged for deriving the size of the object of interest (100) from the volume of the inflated cushion (104). 2. An apparatus as claimed in claim 1, further comprising a frame (101), with the cushion (104) being attached inside said frame (104). 3. An apparatus as claimed in claim 2, wherein the frame (101) further comprises a magnetic emitting means (109) arranged for emitting electromagnetic signals; a magnetic detection means (110) arranged for detecting magnetic induction signals; and an image reconstruction means (106) arranged for reconstructing images based on the detected signals and the derived size of the object of interest (100). 4. An apparatus as claimed in claim 3, wherein the processor is further arranged for outputting a signal based on the size of the object of interest (100) for the purpose of adjusting the magnetic emitting means (109). 5. An apparatus as claimed in claim 2, wherein the apparatus further comprises ultrasound transducers (201) for measuring ultrasound signals. 6. An apparatus as claimed in claim 5, wherein the apparatus further comprises an additional cushion (117) attached to inside of the cushion (104), and the ultrasound transducers (201) are attached to the additional cushion (117). 7. An apparatus as claimed in claim 2, wherein the frame (101) comprises at least two sections that can be disassembled or opened to allow the object of interest (100) to be moved into or out of the frame (101). 8. An apparatus as claimed in claim 7, wherein the apparatus comprises a flexible pipe (302) having two open ends for two respective neighboring sections, which open ends are inserted into said two respective neighboring sections of the cushion such that the sections of the cushion (104) are at a uniform pressure during inflation. 9. An apparatus as claimed in claim 1, wherein the apparatus comprises an elastic tissue (107) attached to the inside of the cushion, which tissue (107) is replaceable for individual use. 10. An apparatus as claimed in claim 1, wherein the substance pumped into the cushion (104) by the pump (102) for inflation is air or a liquid. 11. A method of monitoring an object of interest (100), comprising steps of: inflating an elastic cushion such that the inflated cushion extends in an inward direction so as to wrap itself around the object of interest (100);measuring the pressure exerted on t... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/05 A61B008/00
US07826896 | B2 | 2010-11-02 | Method and apparatus for impedance signal localizations from implanted devices | 2006-11-02 | Method and apparatus for monitoring a plurality of physiological factors contributing to physiological conditions of a heart, that determines a first impedance, corresponding to the plurality of physiological factors, across a plurality of vectors, and a second impedance, corresponding to the plurality of physiological factors, across the plurality of vectors subsequent to determining the first impedance. A relative change in impedance corresponding to the plurality of vectors is determined in response to the first impedance and the second impedance, first minimally contributing physiological factors of the plurality of physiological factors associated with a first physiological factor of the plurality of physiological factors are determined, and relative change in tissue resistivity corresponding to the first physiological factor in response to physiological factors of the plurality of physiological factors other than the first minimal contributing physiological factors and the relative change in impedance are determined. | 1. A method for monitoring a plurality of physiological factors contributing to physiological conditions of a patient in an implantable medical device, wherein the plurality of physiological factors are selected from the group consisting of lung resistivity, blood resistivity, heart muscle resistivity, skeletal muscle resistivity, heart volume and lung volume, and wherein one or more of the plurality of physiological factors contribute to the impedance of a vector, the method comprising: determining a first impedance across each of a plurality of vectors;determining a second impedance across each of the plurality of vectors subsequent to determining the first impedance;for each of the plurality of vectors, determining a difference value between the first impedance for the vector and the second impedance for the vector;determining, based on an identification of a first physiological factor of the plurality of physiological factors and first minimally contributing physiological factors of the plurality of physiological factors associated with the first physiological factor, a relative change in tissue resistivity corresponding to the first physiological factor based on at least some of the determined difference values and physiological factors of the plurality of physiological factors other than the first minimally contributing physiological factors; and,evaluating the relative change in tissue resistivity corresponding to the first physiological factor to determine whether the first physiological factor is an indicator of a medical condition,wherein, for each of the first minimally contributing physiological factors, a sensitivity of at least two of the plurality of vectors to the first physiological factor is substantially similar. 2. The method of claim 1, wherein the first physiological factor is blood resistivity. 3. The method of claim 1, wherein the plurality of vectors comprises a first vector and a second vector, the first vector including a first stimulation path and a first sense path extending between a first electrode, positioned within a ventricle of the heart, and an uninsulated portion of a housing of the device, and the second vector including a second stimulation path, extending between a second electrode, positioned within a ventricle of the heart, and the uninsulated portion of the housing, and a second sense path extending between the first electrode and the uninsulated portion of the housing. 4. The method of claim 3, further comprising: based on an identification of a second physiological factor of the plurality of physiological factors and second minimally contributing physiological factors of the plurality of physiological factors associated with the second physiological factor, determining a relative change in tissue resistivity corresponding to the second physiological factor based on physiological factors of the plurality of physiological factors other than the second minimally contributing physiological factors, at l... | Medtronic, Inc. | A61B005/05
US07856260 | B1 | 2010-12-21 | Implantable cardiac patch for measuring physiologic information | 2006-09-06 | An implantable cardiac patch is configured to be joined to a surface of a heart and includes a body portion having an inner surface configured to be joined to the surface of the heart. The body portion also has a plurality of pores extending into the body portion from the inner surface. A plurality of electrodes are attached to the body portion such that the electrodes are positioned proximate the surface of the heart, and a lead is electrically connected to the electrodes at a connecting portion of the lead. The lead has a proximal end configured for joining to an implantable device. | 1. An implantable cardiac patch configured to be joined to a surface of a heart, the device comprising: a body portion having an inner surface configured to be joined to the surface of the heart, the body portion having a plurality of pores extending into the body portion from the inner surface;a plurality of electrodes attached to the body portion; anda lead electrically connected to the electrodes at a connecting portion of the lead, the lead having a proximal end configured for joining to an implantable device;wherein the body portion includes a backing layer comprising a flexible printed circuit board having traces thereon that extend to the electrodes. 2. The cardiac patch of claim 1, wherein the pores extend completely through the body portion. 3. The cardiac patch of claim 1, wherein the pores extend completely through an inner layer of the body portion and the pores are covered by an outer layer of the body portion. 4. The cardiac patch of claim 1, wherein the pores are circular. 5. The cardiac patch of claim 1, wherein the pores are sized and shaped to allow ingrowth of the heart tissue at least partially into the pores to substantially increase the adherence of the body portion to the surface of the heart. 6. The cardiac patch of claim 1, wherein the body portion comprises a material selected from a group consisting of at least one of polyester, polypropylene, polytetrafluoroethylene, polyurethane, silicone, and silicone rubber. 7. The cardiac patch of claim 1, wherein the body portion further includes an attachment layer, and the pores are included in the attachment layer. 8. The cardiac patch of claim 1, wherein the electrodes include at least one of a monitoring electrode, a sensing electrode and a pacing electrode. 9. The cardiac patch of claim 1, wherein a distal end of the lead includes the connecting portion and is attached to the body portion. 10. The cardiac patch of claim 1, wherein a distal end of the lead includes a lead electrode. 11. An implantable cardiac patch configured to be joined to a surface of a heart, the device comprising: a body portion having an inner surface configured to be joined to the surface of the heart, the body portion having a plurality of pores extending into the body portion from the inner surface;a plurality of electrodes attached to the body portion; anda lead electrically connected to the electrodes at a connecting portion of the lead, the lead having a proximal end configured for joining to an implantable device;wherein the body portion includes a backing layer comprising multiple printed circuit boards interconnected by a mesh layer comprising the pores, each printed circuit board having traces thereon that extend to at least one of the electrodes. 12. The cardiac patch of claim 11 further comprising a plurality of processors, at least one processor on each of the printed circuit boards and connected to at least one of the electrodes. 13. An implantable cardiac patch configured to b... | Pacesetter, Inc. | A61B005/042 A61N001/05
US20100331715 | A1 | 2010-12-30 | SYSTEMS AND METHODS FOR DETECTING EFFORT EVENTS | 2009-06-30 | A method and system for detecting effort events is disclosed. Effort may be determined through feature analysis of the signal as transformed by a continuous wavelet transform, which may be compared against a reference effort measure to trigger an effort event flag that signals the onset and/or severity of an effort event. For example, a respiratory effort measure may be determined based at least in part on a wavelet transform of a photoplethysmograph (PP G) signal and features of the transformed signal. A respiratory reference effort measure may be based at least in part on past values of the respiratory effort measure, and a threshold test may be used to trigger an effort event flag, which may indicate a marked change in respiratory effort exerted by a patient. | 1. A method for monitoring respiratory effort, comprising: receiving an electronic photoplethysmograph signal;transforming the photoplethysmograph signal into a transformed signal based at least in part on a continuous wavelet transform;deriving, using processing equipment, a respiratory effort measure based at least in part on the transformed signal; andgenerating an electronic event flag representative of a respiratory effort event based at least in part on a comparison between the respiratory effort measure and at least one reference respiratory effort measure. 2. The method of claim 1, wherein the at least one reference respiratory effort measure is based at least in part on past instances of the respiratory effort measure within a time window. 3. The method of claim 2, wherein the at least one reference respiratory effort measure is based at least in part on a smoothed gradient of the past instances of the respiratory effort measure within the time window. 4. The method of claim 2, wherein the at least one reference respiratory effort measure is based at least in part on an area under the past instances of the respiratory effort measure within the time window. 5. The method of claim 2, wherein the at least one reference respiratory effort measure is based at least in part on a mean value of the past instances of the respiratory effort measure within the time window. 6. The method of claim 1, wherein generating an electronic event flag comprises: applying a threshold test to the derived effort measure to generate a result, wherein the threshold test is based at least in part on the at least one reference effort measure; andgenerating an electronic event flag based at least in part on the result. 7. The method of claim 6, wherein applying the threshold test comprises comparing the derived effort measure against multiple thresholds. 8. The method of claim 7, wherein the multiple thresholds are based at least in part on the variability of the respiratory effort measure. 9. The method of claim 6, wherein the threshold test is based at least in part on at least one of current and past occurrences of the electronic event flag. 10. The method of claim 6, wherein the threshold test is based at least in part on a physiological signal that is not the respiratory effort measure. 11. A system for monitoring respiratory effort, comprising: at least one memory device;an indicator device, capable of indicating a respiratory effort event in response to an event flag; anda processor, communicably coupled to the at least one memory device and the indicator device and capable of receiving an electronic photoplethysmograph signal, the processor being capable of: calculating a transformed signal based at least in part on at least on the electronic photoplethysmograph signal and a continuous wavelet transform;deriving a respiratory effort measure based at least in part on the transformed signal; andgenerating an event flag representative of a respiratory... | Nellcor Puritan Bennett Ireland | A61B005/08 A61B005/113
US20100292609 | A1 | 2010-11-18 | Method for manufacturing a carrier tape comprising lancets and tape magazines for storing lancets | 2009-11-16 | A method for manufacturing a carrier tape comprising lancets and corresponding products in which a plurality of disposable lancets having a lancet tip for piercing a user's skin are provided on a carrier tape that is or can be wound onto a spool and used for obtaining a sample of body fluid. The lancets can each be arranged in a package and the lancet packages formed in this manner can be applied individually to the carrier tape and fastened thereon. | 1. A method for manufacturing a carrier tape, adapted to be wound onto a spool, the carrier tape having a plurality of disposable lancets provided thereon, the lancets each having a lancet tip to pierce a user's skin and obtain a sample of body fluid, comprising: providing sterilized lancets and arranging each lancet in a package so that the lancets are thereby sterilely enclosed at least in the area of their tips before they are applied to the carrier tape; andsubsequently individually applying the lancet packages formed in this manner to the carrier tape and attaching them thereon. 2. The method of claim 1 wherein the packages are formed from a base foil and a cover foil that is joined thereto. 3. The method of claim 2 wherein the base foil and cover foil are brought into contact with one another and joined while enclosing the lancets. 4. The method of claim 2 wherein the base foil and the cover foil are joined together by a hot-melt adhesive activated by a hot stamp. 5. The method of claim 1 wherein a plurality of the lancet packages are manufactured in a connected manner as a package tape and are separated before application to the carrier tape. 6. The method of claim 5 wherein the package tape is formed by bringing together and joining a base foil strip and a cover foil strip. 7. The method of claim 5 wherein the package tape is manufactured by folding down a foil strip and forming lancet pockets in the folded foil strip. 8. The method of claim 1 wherein the lancet packages are at the same time cut and the border is sealed while enclosing the lancets by using a hot cut. 9. The method of claim 1 wherein the lancets are permanently joined to the packages at a proximal end section of the lancets. 10. The method of claim 1 wherein at least the tips of the lancets are protected against damage in the packages. 11. The method of claim 1 wherein the lancet tips are arranged in pockets formed by the packages. 12. The method of claim 1 wherein the lancet packages are formed by a base foil and a sealing compound applied thereon which encloses at least the lancet tips. 13. The method of claim 1 wherein at least the tips of the lancets in the lancet packages are sterilized before they are applied to the carrier tape. 14. The method of claim 1 wherein the lancet packages are each fastened on the carrier tape at a joining surface that faces away from the lancet. 15. The method of claim 1 wherein the lancet packages are applied to the carrier tape in the form of labels or stickers. 16. The method of claim 1 wherein the lancet packages are provided with pieces of adhesive tape to glue them onto the carrier tape. 17. The method of claim 1 wherein the lancet packages are held on the carrier tape by mechanical connecting means or by welding. 18. The method of claim 1 wherein test elements to which the sample of body fluid can be applied are applied to the carrier tape. 19. The method of claim 18 wherein the test elements are applied in alterna... | - | A61B005/151 B23P011/00 B65B035/30
US20100280404 | A1 | 2010-11-04 | ELECTRICAL IMPEDANCE TOMOGRAPHY EQUIPMENT | 2008-09-25 | There is described an electrical impedance tomography equipment, intended for providing images representing the inside of a portion of the body of a patient by means of electrical impedance measurements taken between electrodes positioned on the skin, comprising a plurality of acquisition units (15, 27) provided with circuits for processing low level analog signals picked up by the cables (13, 29) that interconnect the said equipment with the electrodes (14) applied to the patient, the said acquisition units being substantially symmetrically arranged to provide as a result a substantially equal irradiation of the heat generated in these units, easing the calibration of the analog channels and common mode rejection. The equipment further comprises at least one plane constituted by a printed circuit board whereto the said units are electrically connected, the said plane comprising complementary elements and devices, such as the electrical connection terminals for the said cables. | 1. Electrical impedance tomography equipment comprising a plurality of acquisition units (15, 27), as well as complementary elements and devices, the said acquisition units being provided with circuits for processing of low level analog signals picked up by a plurality of cables (13, 29) that interconnect the said equipment with the electrodes (14) applied to the patient, characterized in that the said acquisition units are provided substantially symmetrically in relation to at least one axis of symmetry (X, Y, Z). 2. Equipment, as claimed in claim 1, characterized in that the said at least one axis of symmetry is a vertical axis (Z). 3. Equipment, as claimed in claim 1, characterized in that the said complementary elements and devices are comprised in at least one plane (11, 25, 26) constituted by a printed circuit board, to which are electrically connected, in symmetrical fashion, the said acquisition units (15, 27). 4. Equipment, as claimed in claim 3, characterized in that the said acquisition units comprise two regions provided side by side, one of such regions comprising the circuits for analog processing of signals and the other comprising the circuits for digital processing of signals. 5. Equipment, as claimed in claim 3, characterized in that the said acquisition units (27) comprise a lower region (27a) containing the circuits for analog processing of signals and an upper region (27b) containing the circuits for digital processing of signals. 6. Equipment, as claimed in claim 4 or 5, characterized in that each acquisition unit (27) comprises at least one printed circuit board. 7. Equipment, as claimed in claim 6, characterized in that the said complementary elements and devices are comprised in a first plane (11, 25) and in a second plane (12, 26) provided parallel to one another. 8. Equipment, as claimed in claim 7, characterized in that the said first plane (11, 25) comprises devices that are substantially provided for processing of analog signals and the said second plane (11, 26) comprises devices that are substantially provided for processing of digital signals. 9. Equipment, as claimed in claim 1, characterized in that the said at least one axis of symmetry is a horizontal axis (X, Y). 10. Equipment, as claimed in claim 9, characterized by comprising two axes of symmetry (X, Y) defining a horizontal plane. 11. Equipment, as claimed in claim 1, characterized in that the said first (11, 25) and second (12, 26) planes are provided parallel in relation to the horizontal plane formed by the said two horizontal axes (X, Y). 12. Equipment, as claimed in claim 8, characterized in that the said first plane (11, 25) is electrically connected to the side (27c) adjacent to the region (27a) containing the said analog processing circuits, in the said acquisition units (27), and the said second plane (12, 26) is electrically connected to the side (27d) adjacent to the region (27b) containing the said digital processing circuits, in the... | Dixtal Biomedica Industria E Comercio Ltda. | A61B005/053
US07860732 | B2 | 2010-12-28 | Medicare pharmacy calculator II | 2007-10-09 | A method and apparatus for providing a Medicare Part D Pharmacy Benefits Calculator. A user inputs data for prescriptions to be purchased by a prospective insured for a plan period. The prescription data includes a quantity, a days supply, and a number of fills. The data for each prescription the user enters is used to calculate anticipated out-of-pocket costs for the prescriptions for one or more pharmacy benefits plans available to the prospective insured. The cost data is presented to the user in a way that illustrates how out-of-pocket costs will vary over the course of a plan period. The user may view costs for alternative drugs and alternatives sources of drugs. The user may vary and expand the presentation of the cost data to better understand the available plan benefits. The invention helps prospective insureds decide among alternative Medicare Part D pharmacy benefits plans offered by one or more insurance companies. | 1. A computerized method for calculating and displaying anticipated out-of-pocket pharmacy costs to a computer user under at least one Medicare Part D pharmacy benefit plan, comprising: a) receiving at a server from said computer user prescription data for a plurality of drugs, said prescription data comprising a drug identifier, a number of doses, a number of days supply, and a number of fills during a plan period;b) determining at said server for each of said plurality of drugs an indicator of whether said drug is a maintenance drug or non-maintenance drug;c) defining at said server a schedule for purchasing said drugs by assigning a drug fill for each fill of said plurality of drugs to a month in which said drug fill will be purchased during said plan period wherein said assignments are made according to: i. said number of doses for said drug;ii. said number of days supply for said drug;iii. said number of fills for said drug; andiv. said indicator of whether said drug is a maintenance drug or a non-maintenance drug;d) calculating at said server an anticipated out-of-pocket cost for each drug fill according to said assignments of each fill of said plurality of drugs to a month in which each drug fill will be purchased during said plan period wherein calculating an anticipated out-of-pocket cost comprises: i. determining a cost for said drug fill under said at least one Medicare Part D pharmacy benefit plan;ii. determining if at least one Medicare Part D stage is applicable to said drug fill;iii. determining out-of-pocket costs for said drug fill according to each Medicare Part D stage applicable to said drug fill as determined in step ii and said cost for said drug fill under said at least one Medicare Part D pharmacy benefit plan;iv. storing data for said anticipated out-of-pocket cost for each drug fill according to said assignments of each fill of said plurality of drugs to a month in which each drug fill will be purchased during said plan period;v. repeating steps i-v for each drug fill; ande) presenting to said computer user anticipated out-of-pocket costs for said plan period according to said schedule by: i. presenting for each applicable month of said plan period an anticipated out-of-pocket cost for each drug fill assigned to said month wherein said anticipated out-of-pocket cost is determined according to said assignments of each drug fill for each of said plurality of drugs to a month in which said drug fill will be purchased during said plan period; andii. presenting in association with each applicable month an indicator of whether more than one Medicare Part D stage is applicable to said month. 2. The computerized method of claim 1 further comprising prompting said computer user to change said schedule defining an order in which to purchase said plurality of drugs. 3. The computerized method of claim 2 further comprising: f) receiving from said computer user said prescription data for said plurality of drugs according to a new s... | Humana Inc. | A61B005/00 G06F019/00
US07835922 | B2 | 2010-11-16 | Diagnostic system and method | 2005-06-29 | A diagnostic system and method. In one embodiment, the invention is a diagnostic system comprising: a form comprising one or more diagnostic questions and one or more possible answers to each diagnostic question; a first overlay comprising one or more windows and one or more diagnostic fields, wherein when the first overlay is positioned over the form, the answers to the diagnostic questions that are indicative of the user having a first medical condition are visible through the windows of the first overlay; and wherein the one or more diagnostic fields contain criteria for diagnosing the user with the first medical condition based on the user's selection of answers on the form that are visible through the windows of the first overlay. In other embodiments, the invention is a computer-based version of the diagnostic system and a diagnostic method. | 1. A diagnostic system comprising: a form comprising one or more diagnostic questions and one or more possible answers to each diagnostic question;a first overlay comprising one or more windows and one or more diagnostic fields, wherein when the first overlay is positioned over the form, the answers to the diagnostic questions that are indicative of the user having a first medical condition are visible through the windows of the first overlay;wherein the one or more diagnostic fields contain criteria for diagnosing the user with the first medical condition based on the user's selection of answers on the form that are visible through the windows of the first overlay; anda second overlay having one or more windows and one or more diagnostic fields, wherein when the second overlay is positioned over the form, the answers indicative of the user having a second medical condition are visible through the windows of the second overlay;wherein the diagnostic fields of the second overlay contain criteria for diagnosing the user with the second medical condition based on the user's selection of answers that are visible through the windows of the second overlay,wherein the first overlay and the second overlay are connected along at least one side to form a file,wherein the form contains first and second sides, the form containing diagnostic questions and answer fields on both the first and second sides of the form; andwherein when the form is inserted into the file, the answers on the first side of the form that are indicative of the user having the first medical condition are visible through the windows of the first overlay, and the answers on the second side of the form that are indicative of the user having the second medical condition are visible through the windows of the second overlay. 2. The diagnostic system of claim 1 further: wherein the form comprises a plurality of diagnostic questions and corresponding answer fields;wherein the first overlay comprises a plurality of windows; andwherein answers indicative of the user having the first medical condition are visible through the windows of the first overlay when the first overlay is positioned over the form. 3. The diagnostic system of claim 2 wherein the diagnostic questions pertain to frequency of headaches, length of headaches, onset of aura symptoms, location of pain, type of pain, severity of pain, physical activities that increase or decrease headaches, or presence of additional headache-related symptoms, or any combination thereof. 4. The diagnostic system of claim 3 wherein the windows on the first overlay are arranged so that when the first overlay is positioned over the form, the answers to the diagnostic questions that are indicative of the user having a migraine are visible through the windows. 5. The diagnostic system of claim 3 wherein the windows on the first overlay are arranged so that when the first overlay is positioned over the form, the answers to the diagnostic questions th... | AstraZeneca AB | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07859261 | B2 | 2010-12-28 | Magnetic resonance method and apparatus for automatically determining objects that attenuate penetrating radiation | 2008-03-21 | In a method and an apparatus for automatic determination of objects that attenuate high energy/penetrating radiation by magnetic resonance, the magnetic resonance apparatus scans and prepares MR images, and the MR images contain information about the T2 relaxation time constant. Subsequently, penetrating radiation-attenuating objects are determined in the MR images by means of the T2 relaxation time constant. | 1. A method for therapeutically treating tissue in a living subject with high-energy radiation comprising the steps of: designating a target volume in a living subject containing tissue to be irradiated with high-energy radiation, said high-energy radiation having an energy associated therewith that produces a therapeutically beneficial modification of said tissue;prior to irradiating said target volume with said high-energy radiation, obtaining a magnetic resonance image, with a magnetic resonance apparatus, of a region of the subject containing said target volume by exposing the subject in the magnetic resonance apparatus to a magnetic resonance imaging sequence that causes the magnetic resonance image to contain information that is representative of the T2 relaxation time constant of all objects contained in said region;in a processor, automatically identifying a degree of attenuation of the high-energy radiation in the said region represented in said obtained magnetic resonance image by analyzing the T2 relaxation time constant information;designating a propagation path of said high-energy radiation proceeding in the subject to said target volume; andselectively orienting said propagation path dependent on the identified degree of attenuation of the high-energy radiation in the propagation path. 2. A method as claimed in claim 1 comprising identifying said degree of attenuation in said propagation path by identifying objects in said propagation path represented by the obtained magnetic resonance image exhibiting a T2 relaxation time constant below a predetermined threshold value. 3. A method as claimed in claim 2 comprising identifying additional objects in said obtained magnetic resonance image exhibiting a T2 relaxation time constant above said predetermined threshold value, and then subsequently suppressing said additional objects in said obtained magnetic resonance image. 4. A method as claimed in claim 3 comprising suppressing said additional objects in said obtained magnetic resonance image by subtracting said additional objects from said obtained magnetic resonance image. 5. A method as claimed in claim 2 comprising employing a value in a time range between 1 ms and 5 ms as said predetermined threshold value. 6. A method as claimed in claim 2 comprising additionally identifying, using said T2 relaxation time constant information, signals in said obtained magnetic resonance image originating from a cladding of a local coil in the magnetic resonance apparatus used to generate the obtained magnetic resonance image and, from said signals, identifying said local coil in said obtained magnetic resonance image as a hiqh-enerqv radiation-attenuating object. 7. A method as claimed in claim 1 comprising generating said obtained magnetic resonance image by exciting nuclear spins in the subject by exposing the subject to a radio-frequency pulse, and after a predetermined time interval, beginning to acquire magnetic resonance data signals fro... | Siemens Aktiengesellschaft | G01V003/00 A61B005/055
US07828742 | B2 | 2010-11-09 | Method and system of monitoring respiratory signal by radio | 2007-07-19 | A method and system is provided for monitoring a respiratory signal by radio. The method includes the steps of: converting a change in electric resistance, which is caused by a change in abdominal circumference measured through a rubber waistband that is made of conductive rubber and is mounted on a lower garment of a testee during respiration, into a voltage signal, performing A/D conversion on the voltage signal, and transmitting the converted digital signal to a short distance by radio using a wireless communication protocol for ZigBee; and receiving the respiratory signal transmitted by radio, transmitting it to a computer unit by wire through an RS-232 port that is a serial communication port, and enabling a tester to monitor the respiratory signal through a screen. | 1. A method of monitoring a respiratory signal by radio, the method comprising: attaching a single rubber waistband that is made of a conductive rubber to a lower garment of a testee so that the waistband completely surrounds the testee's body;a respiratory signal measuring and transmitting step of converting a change in electric resistance, which is caused by a change in abdominal circumference measured through the rubber waistband during respiration, into a voltage signal, performing analog/digital (A/D) conversion on the voltage signal, and transmitting the converted digital signal to a short distance by radio using a wireless communication protocol for ZigBee; anda respiratory signal receiving and monitoring step of receiving the respiratory signal transmitted by radio, transmitting it to a computer unit by wire through an RS-232 port that is a serial communication port, and enabling a tester to monitor the respiratory signal through a screen. 2. The method as claimed in claim 1, wherein the rubber waistband is produced by mixing conductive particles into rubber so as to have the electric resistance from tens of Ω to several kΩ. 3. The method as claimed in claim 1, wherein the rubber waistband in the respiratory signal measuring and transmitting step is operated by one of four resistors constituting a Wheatstone bridge circuit, and undergoes the conversion of the electric resistance change thereof into the voltage signal. 4. The method as claimed in claim 1, wherein the voltage signal converted by the electric resistance change of the rubber waistband corresponds to the respiratory signal that is amplified through a differential amplifier circuit and minimizes high-band noise through a low pass filter circuit. 5. The method as claimed in claim 1, wherein, prior to the respiratory signal receiving and monitoring step, a calibration procedure of performing a CO2 inhalation test on each testee once, analyzing correlation between an abdominal respiratory signal and a tidal volume of the testee to calculate a regression line, and substituting the abdominal respiratory signal, measured in the respiratory signal measuring and transmitting step, into the regression line to measure the lung volume. | Chungbuk National University Industry Academic Cooperation Foundation | A61B005/08
US07860547 | B2 | 2010-12-28 | Visceral fat measuring apparatus, visceral fat measuring method, program, and recording medium | 2004-02-17 | A subcutaneous fat measuring unit measures the subcutaneous fat thickness by pressing a predetermined face against the living body surface. A pressing pressure measuring unit measures the pressure applied from the subcutaneous fat measuring unit to the living body surface, and a database holds information on the relationship among the values of a plurality of subcutaneous fat thicknesses in the condition where no pressure is applied to the living body surface, the pressure applied to the living body surface with respect to each of the plurality of subcutaneous fat thicknesses and the subcutaneous fat thickness in a condition where the pressure is applied. A calculating unit calculates the subcutaneous fat thickness in a condition where no pressure is applied to the living body surface, by use of the database based on the subcutaneous fat thickness measured by the subcutaneous fat measuring unit and the pressure applied to the living body surface which pressure is measured by the pressing pressure measuring unit. And a visceral fat calculating unit calculates the amount of visceral fat from the calculated subcutaneous fat thickness. | 1. A visceral fat measuring apparatus comprising: a subcutaneous fat measuring unit having a contact face shaped to come into contact with a living body surface, and configured to measure a subcutaneous fat thickness with said contact face being pressed against said living body surface;a pressing pressure measuring unit configured to monitor a pressure at which said contact face is pressed against the living body surface; anda visceral fat calculating unit configured to calculate an amount of visceral fat by use of an Expression 2 based on the subcutaneous fat thickness measured by said pressing pressure measuring unit, and the pressure monitored by said pressing pressure measuring unit: S=D×((T−Be−CF)/(A−1)e−CF+1)+E (Expression 2)where S is the amount of visceral fat, F is the monitored pressure, T is the measured subcutaneous fat thickness, A, B, C, D and E are predetermined constants, and e is a base of a natural logarithm,wherein said subcutaneous fat measuring unit is an optical subcutaneous fat measuring apparatus. 2. A visceral fat measuring apparatus comprising: a subcutaneous fat measuring unit having a contact face shaped to come into contact with a living body surface, and configured to measure a subcutaneous fat thickness with said contact face being pressed against said living body surface;a pressing pressure measuring unit configured to monitor a pressure at which said contact face is pressed against the living body surface; anda visceral fat calculating unit configured to calculate an amount of visceral fat by use of an Expression 2 based on the subcutaneous fat thickness measured by said pressing pressure measuring unit, and the pressure monitored by said pressing pressure measuring unit: S=D×((T−Be−CF)/(A−1)e−CF+1)+E (Expression 2)where S is the amount of visceral fat, F is the monitored pressure, T is the measured subcutaneous fat thickness, A, B, C, D and E are predetermined constants, and e is a base of a natural logarithm,wherein said subcutaneous fat measuring unit comprises:an ultrasonic element configured to measure the subcutaneous fat thickness by use of a reflected wave;a domical lid to form an enclosed space between said ultrasonic element and said living body surface; anda soft body of being filled in the space. 3. A visceral fat measuring apparatus comprising: a subcutaneous fat measuring unit having a contact face shaped to come into contact with a living body surface, and configured to measure a subcutaneous fat thickness with said contact face being pressed against said living body surface;a pressing pressure measuring unit configured to monitor a pressure at which said contact face is pressed against the living body surface; anda visceral fat calculating unit having an abdominal girth inputting unit configured to input an abdominal girth, said visceral fat calculating unit being configured to calculate an amount of visceral fat by use of an Expression 5 based on the measured subcutaneous fat thickness, the mo... | Panasonic Electric Works Co., Ltd. | A61B005/05
US07833173 | B2 | 2010-11-16 | Skin incision instrument and method for incising skin with the same | 2010-02-03 | Provided are skin incision instrument to efficiently incise minimal portions and a method for incising skin with the skin incision instrument. The skin incision instrument according to the present invention comprises a holder, a needle, a needle drive unit, two or more skin expander sets, and a reader, wherein the skin expander sets comprise a first skin expander and a second skin expander; the first skin expander and the second skin expander are capable of expanding skin at both sides of the linear incision portion away from the linear incision portion and expanding the linear incision portion in a direction to expand the linear incision portion; and, minimal portions are efficiently incised by selecting a single skin expander set from two or more of the skin expander set based on the direction of the linear incision portion, such that an angle is adjusted to 45 degrees or more and 90 degrees or less wherein the angle is formed by the direction of the linear incision portion read by the reader and the direction to expand the skin by the single skin expander set. | 1. A skin incision instrument comprising: a holder; a needle; a needle drive unit; n skin expander sets, wherein n is a natural number of two or more; and a reader; wherein: the needle and the skin expander are mounted at an end of the holder, the needle drive unit is mounted in the holder,the needle drive unit is configured to form a linear incision portion in skin that is in contact with an end of the holder by moving the needle into contact with the skin,each of the skin expander sets comprises a first skin expander and a second skin expander,the first skin expander and the second skin expander are positioned symmetrically around an axis defined by the needle,the first skin expander and the second skin expander are capable of expanding skin at both sides of the linear incision portion away from the linear incision portion and expanding the linear incision portion in a direction to expand the linear incision portion,the reader reads a direction of the linear incision portion, and m skin expander set(s), wherein m is a natural number of one or more, and n is greater than m, is/are selected by the instrument from n skin expander sets based on the direction of the linear incision portion read by the reader, such that an angle is adjusted between 45 degrees and 90 degrees wherein the angle is formed by the direction of the linear incision portion read by the reader and the direction to expand the skin by the m skin expander set(s). 2. The instrument according to claim 1, wherein said skin expander sets operate in association with said reader. 3. The instrument according to claim 1, wherein said reader comprises an imaging unit, a light source, and a calculating unit. 4. The instrument according to claim 1, wherein said reader comprises a mechanism which operates in association with said skin expander sets to read the direction of the linear incision portion. 5. The instrument according to claim 1, wherein said holder comprises a mark indicating a direction of said blade surface. 6. A blood test kit comprising the instrument according to claim 1. | Panasonic Corporation | A61B005/15
US07860731 | B2 | 2010-12-28 | Monitoring and feedback wireless medical system and method | 2005-12-20 | A system and method are provided for obtaining medical data from a user, analyzing it and providing a feedback message. Existing home medical devices are connected to a data translation device which transmits the data to a cellular phone. The cellular phone transmits the data to a server where it is analyzed and a feedback message is transmitted back to the phone. | 1. A system for monitoring a user's health and for modifying behavior related thereto, comprising: at least one medical device for detecting a specific user physical data, said at least one medical device having connected thereto an interface device capable of transmitting said specific user physical data to a wireless cellular telephone;at least one wireless cellular telephone programmed to receive said specific user physical data and to transmit said data to an analysis system having a database; andan analysis system having a database storing at least one type of user physical data acquired over time for a specific user, said analysis system having a predetermined ordered list of a plurality of message parts and programmed to analyze specific user physical data received from said wireless cellular telephone, and in relation to previously acquired and stored user physical data, a medical circumstance of the user, a category of the user and recent messages sent to the user, as user parameter settings, for assembling a specific message for the user from said plurality of message parts based on one or more medical circumstances determined from the physical data received, category of user, recent messages sent to the user, and user parameter settings, either automatically or through clinician interaction, and for transmitting a message, assembled from specified message parts based on the medical circumstances determined, category of user, recent messages sent to the user, and user parameter settings to the wireless cellular telephone concerning the user's medical circumstances resulting from analysis conducted. 2. The system of claim 1, wherein said at least one medical device comprises a plurality of medical devices. 3. The system of claim 1, wherein said interface device and said wireless telephone are Bluetooth™ function enabled for conducting communication therebetween. 4. The system of claim 1, wherein said at least one medical device comprises an IrDA heart monitor, an IrDA pedometer, a serial digital precision scale and a self-calibrating serial blood pressure cuff for conducting obesity management, and each of said medical devices having its own interface device. 5. The system of claim 1, wherein said at least one medical device comprises a continuous glucose monitoring device for conducting diabetes management. 6. The system of claim 1, wherein said at least one medical device comprises an ECG/EKG monitor, a pulse oximetry monitor and a weight scale for conducting cardiovascular management. 7. The system of claim 1, wherein said analysis system comprises a server connected to a computer for allowing a clinician to review a user's data for creating a feedback message for the user. 8. The system of claim 7, wherein said server is also a web server which hosts a site allowing a user to securely view their data over the internet. 9. The system of claim 1, wherein said interface device comprises a wireless interface, a USB interface, a s... | Confidant Hawaii, LLC | A61B005/00 G06F019/00
US07831295 | B2 | 2010-11-09 | Localization device cross check | 2003-06-05 | A method and apparatus for cross checking points palpated with a localization device. Points corresponding to the palpated points are obtained from a preoperative image of an object such as an X-Ray of a tibia bone. During a surgical procedure, the accuracy of palpated points are determined based on the palpated points and corresponding points obtained from the preoperative image. If a palpated point is inaccurate, an indicator is generated to indicate the inaccuracy of the palpated point. | 1. A method for cross checking data points obtained with a localization device during a surgical procedure, said method comprising the steps of: palpating at least three unaligned points on an object during the surgical procedure using a surgical tool tracked by a localization device, the three unaligned palpated points corresponding to three points obtained from a preoperative image of said object;determining the accuracy with which said localization device determines a location of one of said at least three unaligned palpated points by developing a coordinate transformation based on said at least three unaligned palpated points and said three corresponding points obtained from said preoperative image; palpating a check point during the surgical procedure using said surgical tool, the palpated check point corresponding to a fourth point obtained from the preoperative image; transforming said fourth corresponding point obtained from the preoperative image using said coordinate transformation; and comparing said palpated check point to said transformed corresponding point to determine if it is accurate; andgenerating an indicator if said palpated check point is inaccurate, said palpated check point being inaccurate if said palpated check point differs from said transformed corresponding point by more than a predefined amount. 2. The method of claim 1, wherein said determining step further comprises: comparing a feature associated with said at least three palpated points to a corresponding feature associated with said three corresponding points obtained from said preoperative image of said object. 3. The method of claim 2, wherein said one of said at least three palpated points is determined to be inaccurate if said feature differs from said corresponding feature by more than a predefined amount. 4. The method of claim 2, wherein said comparing step comprises at least the step of: comparing the distance between two of the at least three palpated points and the distance between corresponding points obtained from said preoperative image. 5. The method of claim 4, wherein said step of comparing the distance comprises at least the step of: scaling the distance between said corresponding points obtained from said preoperative image to match the scale of the distance between said two of the at least three palpated points. 6. The method of claim 2, wherein said comparing step comprises at least the step of: comparing a first angle to a second angle, said first angle formed between a first line extending between two of said at least three palpated points and a second line extending between two of said at least three palpated points, at least one point within said first line not being in said second line, said second angle formed between a third line extending between points on said preoperative image corresponding to said two of said at least three palpated points defining said first line and a fourth line extending between points on said preoperative... | Aesculap AG & Co. KG | A61B005/05
US08708921 | B2 | 2014-04-29 | Method and device for detecting and assessing reactive hyperemia using segmental plethysmography | 2010-06-02 | A method for measuring reactive hyperemia in a subject is disclosed. The method includes performing a first segmental cuff plethysmography to generate a baseline arterial compliance curve and/or a baseline pressure-area (P-A) curve, performing a second segmental cuff plethysmography to generate a hyperemic arterial compliance curve and/or a hyperemic P-A curve, and calculating an area between the baseline and the hyperemic curves. The size of the area can be used as an indication of endothelial dysfunction (ED) and ED-related diseases. | 1. A method for measuring reactive hyperemia in a subject, comprising: performing a first segmental cuff plethysmography;generating a baseline arterial compliance curve and/or a baseline pressure-area (P-A) curve on a portion of the body of the subject, wherein the cuff pressure is increased to a first peak cuff pressure and immediately reduced from the first peak cuff pressure;performing a second segmental cuff plethysmography;generating a hyperemic arterial compliance curve and/or a hyperemic P-A curve, wherein the cuff pressure is increased to a second peak level, maintained at the second peak cuff pressure for a predetermined period of time, and then reduced from the second peak cuff pressure;calculating the difference between the baseline arterial compliance curve and the hyperemic arterial compliance curve as an area between the arterial compliance curves, and/or the difference between the baseline P-A curve and the hyperemic P-A curve as an area between the P-A curves; andmeasuring a level of reactive hyperemia based on the area between the arterial compliance curves and/or the area between the P-A curves,wherein a first cuff compliance curve is generated during the first segmental cuff plethysmography and a second cuff compliance curve is generated during the second segmental cuff plethysmography, wherein said first and second cuff compliance curves are generated using a metering pump to measure the volume change in a cuff, and a pressure sensor to measure pressure change for each known volume change, and wherein the metering pump is a low-frequency pump operating at a frequency that is not higher than the arterial cycle frequency. 2. The method of claim 1, further comprising: determining a level of endothelial dysfunction (ED) in the subject based on the result of the measuring step. 3. The method of claim 1, further comprising: determining a risk of an ED-related disease based on the result of the measuring step. 4. The method of claim 3, wherein the ED-related disease is selected from the group consisting of hypertension, atherosclerosis, cardiovascular disease, stroke, atrial fibrillation, congestive heart failure, peripheral vascular disease, septic shock, hypercholesterolemia, type I and II diabetes, erectile dysfunction, rheumatic arthritis, HIV, liver diseases and pre-eclampsia. 5. The method of claim 1, wherein the first peak cuff pressure is in the range of 150-180 mmHg. 6. The method of claim 1, wherein the second peak cuff pressure is a pressure that is above the systolic pressure of the subject. 7. The method of claim 6, wherein the second peak cuff pressure is in the range of 150-180 mmHg. 8. The method of claim 1, wherein the predetermined period of time is about 2-10 minutes. 9. The method of claim 8, wherein the predetermined period of time is about 5 minutes. 10. The method of claim 1, wherein the systolic blood pressure of the subject is determined during the first segmental cuff plethysmography. 11. A method ... | Cordex Systems, LLC | A61B005/022 A61B005/0295
US08473042 | B2 | 2013-06-25 | Tachyarrhythmia detection using dynamic duration | 2010-06-04 | An apparatus comprises an implantable cardiac signal sensing circuit configured to produce a sensed cardiac signal representative of cardiac activity of a subject and a controller communicatively coupled to the cardiac signal sensing circuit. The controller includes a sensing module configured to detect a cardiac depolarization using the sensed cardiac signal and an arrhythmia detection module. The arrhythmia detection module is configured to detect a depolarization rate or interval that satisfies a first tachyarrhythmia detection rate zone threshold, adjust a specified detection duration threshold for the tachyarrhythmia detection zone according to the detected depolarization rate, declare that the detected depolarization rate is an episode of tachyarrhythmia when the detected depolarization rate is sustained for the adjusted detection duration threshold, and provide an indication of the tachyarrhythmia to a user or process. | 1. An apparatus comprising: an implantable cardiac signal sensing circuit configured to produce a sensed cardiac signal representative of cardiac activity of a subject;a controller communicatively coupled to the cardiac signal sensing circuit, wherein the controller includes: a sensing module configured to detect a cardiac depolarization using the sensed cardiac signal; andan arrhythmia detection module configured to: detect a detection episode in which a depolarization rate or interval satisfies a first tachyarrhythmia detection rate zone threshold;adjust, for the same detection episode, a specified detection duration threshold for the tachyarrhythmia detection zone according to as a function of the detected depolarization rate or interval, wherein the detection duration threshold is adjustable to a plurality of values for depolarization rates or intervals within the same tachyarrhythmia detection rate zone, and wherein the detection duration threshold is adjustable according to a different function for a different tachyarrhythmia detection zone;generate an indication of detection of an episode of tachyarrhythmia when the detected depolarization rate is sustained for a period of time that meets or exceeds the adjusted detection duration threshold; andprovide the indication of the detected episode of tachyarrhythmia to a user or process; anda therapy circuit communicatively coupled to the controller and configured to provide an anti-tachyarrhythmia therapy,wherein the arrhythmia detection module is configured to determine, following expiration of the adjusted detection duration threshold, whether the declared tachyarrhythmia episode is sustained during a specified sustained rate duration (SRD), andwherein the controller is configured to initiate a delivery of an anti-tachyarrhythmia therapy in response to an expiration of the SRD during which the declared tachyarrhythmia episode is sustained. 2. The apparatus of claim 1, including: a memory circuit included with or communicatively coupled to the controller and configured to store detection duration values using depolarization rate or depolarization interval as an index; andwherein the arrhythmia detection module is configured to determine the detection duration threshold using a detection duration value determined using the index. 3. The apparatus of claim 1, wherein the arrhythmia detection module is configured to calculate the adjusted detection duration threshold as a function of the detected depolarization rate and a base detection duration value, wherein the base detection duration value is specified according to the tachyarrhythmia detection rate zone threshold. 4. The apparatus of claim 1, wherein the arrhythmia detection module is configured to: determine that a detected depolarization rate or interval falls within one of a plurality of specified tachyarrhythmia detection rate zones; andadjust the detection duration threshold for the tachyarrhythmia rate zone according to the detected ... | Cardiac Pacemakers, Inc. | A61B005/0464
US08636661 | B2 | 2014-01-28 | Embeddable modules for measuring blood glucose levels | 2009-06-04 | An embeddable module for measuring a blood glucose level may be embedded in an external host and may comprise a host interface, a microcontroller, an ASIC, a temperature measuring element, and a measurement strip port. The host interface may permit electrical communication between the host and the microcontroller. The microcontroller and ASIC may communicate with each other and may work together to measure the blood glucose level of a blood sample. The measurement strip port may receive a removable measurement strip having the blood sample. The ASIC may communicate with the measurement strip port such that the ASIC may receive signals from the measurement strip port related to the blood glucose level of the blood sample. The temperature measuring element may be in electrical communication with the microcontroller and may measure the temperature of the module and transmit signals to the microcontroller related to the temperature of the module. | 1. A blood glucose measurement module embeddable in an external host for measuring a blood glucose level of a blood sample provided on a disposable measurement strip, the module consisting of: a host interface;a microcontroller;an application-specific integrated circuit (ASIC);a temperature measuring element;a measurement strip port,an external code key separate from the measurement strip and configured to contain calibration information related to the measurement strip,a code key port configured to receive the external code key wherein the code key port resides completely within a housing of the host; anda non-volatile memory;wherein:the host interface is operable to permit electrical communication between the host and the microcontroller such that the module is operable to transmit a result of a blood glucose level measurement to the host and the module is operable to send to the host an event interrupt signal to indicate to the host that an event has occurred on the module and to send to the host a module ready signal to indicate to the host that the module is ready to receive a message from the host, wherein the event comprises detecting insertion of the measurement strip into the measurement strip port;the microcontroller is in electrical communication with the ASIC wherein the microcontroller and the ASIC are operable to measure the blood glucose level of the blood sample,the measurement strip port is operable to receive the disposable measurement strip having the blood sample,the ASIC is in electrical communication with the measurement strip port such that the ASIC is operable to receive signals from the measurement strip port related to the blood glucose level of the blood sample provided on the measurement strip,the temperature measuring element is in electrical communication with the microcontroller and is operable to measure the temperature of the module and transmit signals to the microcontroller related to the temperature of the module; andwherein the non-volatile memory is operable to store data relating to the operation of the module or calibration data for the measurement strip, wherein the non-volatile memory is in electrical communication with the microcontroller such that the microcontroller is operable to read data from and write data to the memory. 2. The blood glucose measurement module of claim 1 wherein: the code key is operable to be removably connected to the module; andthe microcontroller is in electrical communication with the code key such that the microcontroller is operable to read the calibration data stored on the code key when the code key is connected to the module. 3. The blood glucose measurement module of claim 1 wherein the module is operable to request calibration data for the measurement strip from the host. 4. The blood glucose measurement module of claim 1 wherein the host is in electrical communication with the code key having calibration data for the measurement strip and is operable to communicate ... | Roche Diagnostics Operations, Inc. | A61B005/00 G06F013/12
US07950163 | B2 | 2011-05-31 | Foot type and plantar pressure measuring device | 2009-12-30 | A foot type and plantar pressure measuring device is provided. The device includes a color pad and a foggy film on the color pad. The device further includes a mesh pad on the foggy film and/or a resilient pad under the color pad, the resilient pad having protrusions. When a patient stands still on the device for a few seconds and then steps off, the patient's foot type and plantar pressure is shown by color pattern displayed on surface areas of the foggy film. The color pattern corresponds to the plantar pressure exerted on the device, and tones of the color pattern represent the plantar pressure exerted on the surface areas. | 1. A foot type and plantar pressure measuring device, comprising: a soft, flexible color pad made of a sticky material; anda semi-transparent film on the color pad;wherein surface areas of the semi-transparent film display colors in response to a pressure exerted on the semi-transparent film and the color pad, andwherein tones of the colors are associated with the pressure exerted on the surface areas. 2. The device of claim 1, further comprising a meshed pad having a plurality of holes, the meshed pad being placed on the semi-transparent film. 3. The device of claim 2, wherein at least one of the holes are of a circular shape, a circle-like shape, or a polygonal shape. 4. The device of claim 1, wherein the semi-transparent film comprises a foggy film. 5. The device of claim 1, wherein the color pad is of a dark color. 6. The device of claim 5, wherein the dark color comprises any color from gray to black. 7. The device of claim 1, further comprising a resilient pad formed under the color pad. 8. The device of claim 7, wherein the resilient pad includes a plurality of protrusions for exerting a reactive force on the color pad. 9. A device for measuring plantar pressure, comprising: a resilient pad including a plurality of protrusions;a color pad on the resilient pad; anda semi-transparent film on the color pad, wherein surface areas of the semi-transparent film display colors in response to a pressure exerted on the semi-transparent film and the color pad. 10. The device of claim 9, wherein the protrusions contact the color pad. 11. The device of claim 9, wherein tones of the colors are associated with the pressure exerted on the surface areas. 12. The device of claim 9, further comprising a meshed pad on the semi-transparent film, the meshed pad having a plurality of holes. 13. A device for measuring plantar pressure, comprising: a resilient pad including a plurality of protrusions;a color pad on the resilient pad, the color pad being made of a sticky material;a light transmitting film on the color pad; anda meshed pad on the light transmitting film, the meshed pad having a plurality of holes. 14. The device of claim 13, wherein at least one of the holes are of a circular shape, a circle-like shape, or a polygonal shape. 15. The device of claim 13, wherein the protrusions contact the color pad. | - | A61B005/103 G01B021/20
US08005540 | B2 | 2011-08-23 | Monitoring of chronobiological rhythms for disease and drug management using one or more implantable device | 2010-07-14 | The health state of a subject is automatically evaluated or predicted using at least one implantable device. In varying examples, the health state is determined by sensing or receiving information about at least one physiological process having a circadian rhythm whose presence, absence, or baseline change is associated with impending disease, and comparing such rhythm to baseline circadian rhythm prediction criteria. Other chronobiological rhythms beside circadian may also be used. The baseline prediction criteria may be derived using one or more past physiological process observation of the subject or population of subjects in a non-disease health state. The prediction processing may be performed by the at least one implantable device or by an external device in communication with the implantable device. Systems and methods for invoking a therapy in response to the health state, such as to prevent or minimize the consequences of predicted impending heart failure, are also discussed. | 1. A method comprising: sensing or receiving at an implantable device, information about at least one physiological process having a chronobiological rhythm whose presence, absence, or change is statistically associated with a disease;comparing the chronobiological rhythm of the at least one physiological process to one or more chronobiological rhythm prediction criteria; andat least one of predicting, detecting, or identifying an occurrence of disease using the comparison. 2. The method of claim 1, wherein predicting the occurrence of disease includes predicting an occurrence of impending disease occurring during a specified prediction time period. 3. The method of claim 1, wherein sensing or receiving the information about the at least one physiological process includes sensing or receiving at least one of body temperature, heart rate, heart rate variability, respiration rate, respiration rate variability, minute ventilation, tidal volume, activity, blood pressure, posture, sleep pattern, thoracic impedance, or at least one heart sound. 4. The method of claim 1, further comprising providing an associated collection time to the chronobiological rhythm of the at least one physiological process. 5. The method of claim 1, further comprising sensing or receiving information about at least one arrhythmia incidence; and wherein predicting the occurrence of disease includes using a time of day of the arrhythmia incidence. 6. A method comprising: sensing or receiving at an implantable device, information about at least one physiological process having a chronobiological rhythm whose presence, absence, or change is statistically associated with a disease;comparing the chronobiological rhythm of the at least one physiological process to one or more chronobiological rhythm prediction criteria; andat least one of predicting, detecting, or identifying an occurrence of disease using the comparison, wherein predicting the occurrence of disease is computed using one or more stored weighting factor, each weighting factor corresponding to a chronobiological rhythm of a different one of the at least one physiological process. 7. The method of claim 6, further comprising adjusting or initiating a therapy using the predicted, detected, or identified occurrence of disease. 8. The method of claim 7, wherein adjusting or initiating the therapy includes determining a drug delivery time using the chronobiological rhythm of the at least one physiological process. 9. The method of claim 7, wherein adjusting or initiating the therapy includes recovering the chronobiological rhythm of the at least one physiological process using one or both of drug delivery or neurostimulation. 10. The method of claim 7, further comprising monitoring the efficacy of the therapy using a post-therapy chronobiological rhythm of the at least one physiological process. 11. A method comprising: sensing or receiving at an implantable device, information about at least one physiological pr... | Cardiac Pacemakers, Inc. | A61N001/02
US08162846 | B2 | 2012-04-24 | Head-stabilized, nystagmus-based repositioning apparatus, system and methodology | 2010-06-29 | Structure and methodology involving mountable and head-wearable frame structure which is positionally stabilized, during use, relative a human subject's head, and which carries a selection of positionally anchored data sensors, and stimuli deliverers, that are relevant to the diagnosis and treatment of vestibular disorders. Special configurations are provided for two types of stimulators, one for sound application and air-pressure modification, and the other for the introduction of fluids to the ear. Stabilization enables tight and accurate correlation of data which is quickly analyzable by a connected, properly algorithmed computer, which can also be used for feedback control in a designed “expert” system. The invention enables, among other things, practical and significant differentiation between physiological and pathological nystagmus. | 1. A method of enabling assessment and/or treatment of vestibular-function in a human subject, comprising: providing a goggle-like frame structure configured for securing to a subject's head in a non-relative-motion condition therewith, wherein the frame structure includes an eye-enclosing, ambient light-excluding housing, and further includes one or more tensive bands extending from the housing and configured to securely grip a portion of the subject's head;providing one or more image-capture devices coupled in a non-relative-motion condition with the housing, wherein the one or more image-capture devices is/are configured to obtain, within a darkened environment of the housing, imagery of a visually-observable oculo-vestibular response to a vestibular response-evoking stimulus administered to the subject;providing circuitry operably coupled with the one or more image-capture devices, the circuitry being configured to convert the imagery to computer-readable vestibular response data; andproviding a computing device specially configured with correlation instructions which, when executed by the computing device, correlate the vestibular response data with subject spatial orientation-indicating data and to display to a user a viewable representation of the correlated data via a display device operably coupled with the computing structure. 2. The method of claim 1, further comprising: one or more subject spatial-orientation sensing devices configured to produce data indicating a spatial orientation of one or more of the subject's semi-circular canals, wherein the spatial orientation temporally-corresponds to the administered stimulus. 3. The method of claim 2, wherein the one or more subject spatial-orientation sensing devices are selected from the group consisting of (a) a linear accelerometer, (b) an angular accelerometer, (c) an inclinometer, (d) a geographical direction indicating device and (e) a subject spatial-orientation repositioning device. 4. The method of claim 1, further comprising: providing one or more vestibular-response stimulating devices likewise coupled with the frame structure in a condition of immobility relative thereto. 5. The method of claim 4, wherein the one or more vestibular-response stimulating devices are selected from the group consisting of (a) a sound deliverer, (b) an air pressure modifier configured to engage with the subject's ear and affect an air pressure within the subject's auditory canal, (c) a fluid-flow structure configured to engage with the subject's ear and dispense a fluid into the subject's auditory canal, (d) a light-emitting structure, (e) a visual image-presenting structure, (f) a galvanic stimulus structure, (g) a caloric stimulus structure, and (h) a vibration-generating structure. 6. The method of claim 1, wherein the viewable representation includes one or more selected from the group consisting of (a) a graphical representation of a spatial orientation of one or more of the subject's semi... | Epley Research LLC | A61B013/00 A61B005/00 A61B005/103 A61B005/117
US09402544 | B2 | 2016-08-02 | Analyte sensor and apparatus for insertion of the sensor | 2010-02-01 | An apparatus for insertion of a medical device in the skin of a subject is provided. | 1. A method of operating an inserter comprising a base, a spring, an introducer mechanism, and a shuttle that is coupled with the introducer mechanism, the shuttle having a surface in contact with the spring, the method comprising: positioning an end of the base of the inserter on a skin surface of a subject, the inserter containing an analyte monitoring assembly including a power source and electronics coupled to an analyte sensor and operable to communicate analyte data;moving the shuttle, the shuttle surface in contact with the spring, the analyte monitoring assembly, and the introducer mechanism from a first position within the inserter to a second position where at least a portion of the analyte sensor of the analyte monitoring assembly is brought into contact with fluid in the subject; andautomatically retracting the shuttle and the introducer mechanism away from the skin surface and into the inserter with a force applied by the spring against the shuttle surface, while the end of the base is maintained on the skin surface. 2. The method of claim 1, further comprising, prior to positioning the end of the base, disengaging an end cap from a bottom end of the base to expose an interior space of the base. 3. The method of claim 1, wherein automatically retracting the shuttle and the introducer mechanism comprises automatically retracting the shuttle and the introducer mechanism upon the introducer mechanism reaching a predetermined position. 4. The method of claim 3, further comprising disengaging the introducer mechanism from a retention member upon reaching the predetermined position. 5. The method of claim 1, further comprising preventing the displacement of the analyte monitoring assembly within the inserter between the first position and the second position when an end cap is engaged with a bottom end of the housing. 6. The method of claim 1, wherein the analyte sensor is a glucose sensor. 7. The method of claim 1, further comprising maintaining the power source of the analyte monitoring assembly in an inactive state when the analyte monitoring assembly is positioned at or near one of the first position or the second position. 8. The method of claim 1, further comprising engaging the analyte monitoring assembly with at least a portion of the introducer mechanism within the inserter during movement from the first position to the second position. 9. The assembly of claim 1, wherein the electronics include an analyte sensor electronic circuit that is activated when the analyte monitoring assembly is at or near the second position. 10. The method of claim 1, further comprising positioning a portion of the analyte monitoring assembly on the skin surface at the second position. 11. The method of claim 10, further comprising activating the analyte monitoring assembly when the portion of the analyte monitoring assembly is positioned on the skin surface. 12. The method of claim 1, further comprising positioning at least a portion of the... | ABBOTT DIABETES CARE INC. | A61B005/00 A61B005/0205 G06K007/10 A61B005/145 A61B005/1486 A61M005/158 A61M005/172 A61M005/32
US20100292554 | A1 | 2010-11-18 | Analyte Monitoring Device and Methods of Use | 2010-07-29 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor has, for example, a substrate, a recessed channel formed in the substrate, and conductive material disposed in the recessed channel to form a working electrode. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte. The analyte monitor may also be part of a drug delivery system to alter the level of the analyte based on the data obtained using the sensor. | 1. (canceled) 2. A method for classifying a level of noise in a signal obtained from a continuous glucose sensor, the method comprising: receiving a signal from a continuous glucose sensor; andclassifying a level of noise on the signal. 3. The method of claim 2, wherein the step of classifying comprises applying a low pass filter to the signal to determine a signal strength. 4. The method of claim 3, wherein the step of classifying comprises defining one or more noise thresholds for classification of the level of noise on the signal, wherein the one or more noise thresholds are based at least in part on a percentage of the signal strength. 5. The method of claim 4, wherein the step of classifying comprises applying one or more low pass filters to the noise signal to obtain one or more noise indicators and comparing the noise indicators with the one or more noise thresholds. 6. The method of claim 2, further comprising determining a noise signal from the sensor signal, and wherein the step of classifying comprises applying one or more filters to the noise signal to obtain one or more noise indicators and comparing the noise indicators with one or more noise thresholds. 7. The method of claim 2, wherein the step of classifying comprises performing spectral analysis to determine at least one of a signal strength and a noise indicator. 8. The method of claim 2, wherein the step of classifying a level of noise comprises using hysteresis. 9. The method of claim 2, further comprising controlling an output based at least in part on the level of noise. 10. The method of claim 2, further comprising controlling a display based at least in part on the level of noise. 11. The method of claim 10, wherein the step of controlling a display comprises controlling the display of raw and/or filtered data based at least in part on the level of noise. 12. The method of claim 10, wherein the step of controlling a display comprises displaying rate of change information based at least in part on the level of noise. 13. The method of claim 2, further comprising a step of controlling at least one alarm indicative of at least one of hypoglycemia, hyperglycemia, predicted hypoglycemia, and predicted hyperglycemia based at least in part on the level of noise. 14. The method of claim 2, further comprising a step of controlling medicament delivery and/or medicament therapy instructions based at least in part on the level of noise. 15. The method of claim 2, further comprising a step of diagnosing a sensor condition based at least in part on the level of noise. 16. The method of claim 2, further comprising a step of suspending display of sensor data based at least in part on the level of noise. 17. The method of claim 2, further comprising a step of shutting down a sensor session based at least in part on the level of noise. 18. The method of claim 2, further comprising a step of displaying the level of noise on the user interface. 19. The method of claim 2, fu... | Abbott Diabetes Care Inc. | A61B005/05
US20100331637 | A1 | 2010-12-30 | Method and System for Measuring a Composition in a Blood Fluid | 2009-04-13 | A system (10) and method for measuring a composition in the blood fluid is disclosed. The system (10) comprises a non-invasive measuring unit (12) for measuring the composition; and at least one neural network (16) for processing a plurality of measurements taken by the non-invasive measuring unit (12) to determine an overall measurement of the composition in the blood fluid. A further aspect of the invention discloses a computer-readable medium for performing the above method. | 1. A system for measuring a composition of a blood fluid comprising: at least one neural-network for processing a plurality of measurements taken by a non-invasive measuring unit to determine an overall measurement of the composition in the blood fluid, wherein a linear equation associated with each output node of the at least one neural network is determined from a controlled source prior to training the at least one neural network. 2. A system for measuring a composition of a blood fluid comprising: a non-invasive measuring unit for measuring the composition; andat least one neural network for processing a plurality of measurements taken by the non-invasive measuring unit to determine an overall measurement of the composition in the blood fluid, wherein a linear equation associated with each output node of the at least one neural network is determined from a controlled source prior to training the at least one neural network. 3. A method of measuring a composition in a blood fluid comprising: Obtaining a plurality of measurements from a non-invasive measuring unit, andprocessing the plurality of measurements by at least one neural network to determine an overall measurement of the composition in the blood fluid, wherein a linear equation associated with each output node of the at least one neural network is determined from a controlled source prior to training the at least one neural network. 4. A system or method for measuring a composition in the blood fluid according to any of claim 1 to 3, wherein the linear equation associated with each output node is obtained based on a line of best fit. 5. A system or method for measuring a composition in the blood fluid according to any of claim 1 to 4, where the linear equation associated with each hidden node is determined through automated processes. 6. A system or method for measuring a composition in the blood fluid according to any one of claims 1 to 5, where the at least one neural network implements a back propagation algorithm. 7. A system or method for measuring a composition in the blood fluid according to claim 5, where the number of nodes in an input layer of the at least one neural network matches the number of measurements in the plurality of measurements taken by the non-invasive measuring unit. 8. A system or method for measuring a composition in the blood fluid according to any one of claim 5 to 7, where the at least one neural network comprises at a hidden layer of at least four nodes. 9. A system or method for measuring a composition in the blood fluid according to any one of claims 6 to 8, where the output value for the hidden node is a summation of weighted measurements. 10. A system or method for measuring a composition in the blood fluid according to any one of claims 6 to 9, where the output value for the output node is a summation of weighted normalized hidden node output values. 11. A system or method for measuring a composition in the blood fluid according to any on... | GLUCOSTATS SYSTEM PTE. LTD. | A61B005/1455
US20100286524 | A1 | 2010-11-11 | SYSTEM AND METHOD FOR ORGANIZING, RECORDING AND DISPLAYING IMAGES IN ULTRASOUND IMAGING SYSTEMS | 2007-07-06 | An ultrasound imaging system includes a video processor receiving coherent echo signals from a beamformer. The video processor converts the coherent echo signals to image data, which are applied to a streaming digital video processor. The system also includes an ECG monitor and processor that obtains an ECG signal and uses the signal to provide heart cycle data indicative of the beginning and end of each of the plurality of heart cycles. The streaming digital video processor converts the image data into streaming digital data, which are divided into the heart cycle sections as indicated by the heart cycle data. The system also includes a digital video disk recorder that records a digital video disk containing the digital video data divided into chapters containing respective heart cycle sections. The recorder may also record chapters of ECG data derived from the ECG signal along with the digital video data. | 1.-11. (canceled) 12. A method of organizing and storing digital data corresponding to an ultrasound image of an anatomical structure of a patient, the method comprising: obtaining an ECG waveform from the patient;using the ECG waveform to identify each of a plurality of heart cycles of the patient;dividing the digital data into sections each of which corresponds to a respective ultrasound image obtained during a respective one of the identified heart cycles; andrecording the digital data on a digital video disk, the digital data being recorded on the digital video disk in chapters corresponding to the respective sections into which the digital data are divided. 13. The method of claim 12, further comprising adjusting the beginning and ending points of each of the plurality of heart cycles with respect to the ECG waveform. 14. The method of claim 12 wherein the digital data to which the ultrasound image corresponds comprises a two-dimensional image. 15. The method of claim 12 wherein the digital data to which the ultrasound image corresponds comprises a three-dimensional volumetric image. 16. A method of displaying an ultrasound image, comprising: providing a digital video disk storing digital data corresponding to a plurality of ultrasound images, and an ECG waveform indicative of a plurality of heart cycles occurring during a period of time that the digital data corresponding to the plurality of ultrasound images were obtained, the digital data being divided into a plurality of chapters each of which contains the stored digital data obtained during a respective one of the heart cycles;reading the stored digital data from the digital video disk; andusing one of the chapters of digital data read from the digital video disk to display one of the ultrasound images and the ECG waveform corresponding to one of the read chapters of digital data. 17. The method of claim 16, further comprising displaying an index containing respective index entries for at least some of the chapters of digital data stored on the digital video disk, the index including an index entry for at least the chapter of digital data corresponding to the displayed ultrasound image and ECG waveform. 18. The method of claim 17, further comprising: selecting an index entry for one of the chapters in the displayed index;reading the digital data in the chapter indicated by the selected index entry; andusing the read digital data in the chapter indicated by the selected index entry to display the corresponding ultrasound image and ECG waveform. 19. The method of claim 16, further comprising displaying an image cursor adjacent the ECG waveform indicative of the time in a heart cycle where the digital image data corresponding to the displayed ultrasound image were obtained. 20. The method of claim 16, further comprising displaying cursors adjacent the ECG waveform indicative of the beginning and end of each heart cycle corresponding to the sections into which the digital image dat... | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61B008/14 A61B005/04
US20100310511 | A1 | 2010-12-09 | DETERMINING DISEASE OR IMMUNE SYSTEM FEATURES BASED ON ANALYSIS OF IMMUNE SYSTEM CHARACTERISTICS | 2010-03-05 | Methods for analyzing one or more elements, markers, ligands or other characteristics of the immune system of a body (e.g., of a human or other animal), and based on the analysis, determine a location of disease, identify immune system failure, and/or determine treatments based on disease location or immune system failure. All or a part of the immune system process may be analyzed to identify specific features of the immune system, e.g., which may indicate that the disease will evade the immune system and produce a negative outcome. Therapy may be employed to correct immune system failure rather than addressing the disease directly. | 1. A method for determining disease location in a body, comprising: analyzing one or more elements, markers, ligands, or other characteristics of an immune system of a body; anddetermining a location in the body at which a disease is present based on a result of analyzing the one or more characteristics of the immune system. 2. The method of claim 1, wherein the step of analyzing comprises: analyzing one or more of cells, cytokines, antibodies, soluble factors, or gradients related to the immune system. 3. The method of claim 2, wherein the step of analyzing comprises: analyzing cells using one or more of the following: assaying a surface of the cells,extracting cell contents,imaging the cells,assaying factors secreted by the cells,tracking a migratory behavior of the cells, orassaying one or more binding characteristics of the cells. 4. The method of claim 3, wherein the step of analyzing cells comprises: assaying a surface of the cells by assaying one or more of T-cell receptors, MHC-antigen complexes, integrins, selectin ligands, chemokine receptors, B-cell receptors, KIR receptors, cytokine receptors, or NKG2 receptors. 5. The method of claim 3, wherein the step of analyzing cells comprises: assaying factors secreted by the cells, wherein the factors include cytokines that recruit other cells. 6. The method of claim 3, wherein the step of analyzing cells comprises: assaying factors secreted by the cells, wherein the factors include factors that imprint tissue-specific homing properties onto other cells of the immune system. 7. The method of claim 1, wherein the step of analyzing comprises: analyzing characteristics of the immune system including body temperature or other physical characteristics of the body with a spatial resolution of about 500 micrometers or greater. 8. A method for detecting failure of one or more phases of an immune system, comprising: analyzing one or more characteristics of the immune system of a body during a response of the immune system to a cancer in the body; anddetermining failure of at least one phase of the immune systein in responding to the cancer based on a result of analyzing the one or more characteristics of the immune system. 9. The method of claim 8, wherein the step of analyzing comprises: assessing a function of the immune system at one or more phases of its sequential process; andwherein immune system failure is determined based on a result of assessing the function of the immune system at one or more phases of its sequential process in responding to the cancer. 10. The method of claim 9, wherein the step of analyzing comprises: assessing the function of the immune system at multiple phases of its sequential process. 11. The method of claim 9, wherein phases of the sequential process include activation of the innate immune system and activation of the adaptive immune system. 12. The method of claim 9, wherein phases of the sequential process include acute phase response, behavior of macroph... | Immuneering Corporation | A61K038/19 G01N033/53 G01N033/50 A61K035/74 A61K038/21 A61P037/04 A61B005/01
US20100305411 | A1 | 2010-12-02 | Control of operating characteristics of devices relevant to distance of visual fixation using input from respiratory system and/or from eyelid function | 2007-08-14 | Determination of distance of visual fixation using input from respiratory system and/or from eyelid function, for the purpose of controlling applications including the focus of image capture and viewing devices.;The invention relates to an apparatus and method for using respiratory and/or eyelid function data as a control system for applications which would benefit from a determination of a person's distance of visual fixation, such as to control the focusing of image capture and viewing devices. | 1-43. (canceled) 44. A method of controlling an operating characteristic of a device wherein that operating characteristic is related to a person's distance of visual fixation, characterised in that the method includes detecting selected physiological data of the person, which data is representative of changes in the distance of visual fixation of the person controlling the operating characteristic in response to the physiological data of the person. 45. A method according to claim 44 wherein the physiological data includes at least one of respiratory function data and eyelid function data. 46. A method according to claim 44 where the determination of changes in distance of visual fixation is used to control the focus of image viewing and image capture devices. 47. A method according to claim 44 where the determination of changes in distance of visual fixation is used to control the focus of image capture for a camera. 48. A method according to claim 44 where the determination of changes in distance of visual fixation is used to control the focus of at least one of: spectacles such as liquid crystal bifocals;binoculars;telescopes;microscopes;night vision goggles. 49. An apparatus for controlling an operating characteristic of a device wherein that operating characteristic is relevant to a person's distance of visual fixation, characterised in that the apparatus includes a computation unit which is operable to control the operating characteristic in response to an input from a sensor which collects physiological data of the person, to determine changes in the distance of visual fixation of the person. 50. An apparatus according to claim 49 wherein the physiological data includes at least one of the person's respiratory function data and eyelid function data. 51. An apparatus according to claim 49 which includes a sensor directly measuring pressure within the respiratory system such as a pressure sensor held in the mouth, held between the lips or implanted into a paranasal sinus chamber. 52. An apparatus according to claim 49 which includes at least one of: a sensor which measures pressure within the respiratory system by detecting changes in the sound of the user's breathing or output of the vocal system;a sensor that monitors heart rate in order to compare this input to the changing respiratory and/or eyelid function parameters;a motion sensor such as used in pedometers in order to compare this input to the changing respiratory and/or eyelid function parameters; andan electroencephalogram sensor or nerve sensor to detect electrical activity of the brain or nervous system controlling respiratory and/or eyelid functions. 53. An apparatus according to claim 49 which includes at least one plus lens in between the eye of the user and the subject being viewed to exaggerate the sensitivity of the user's respiratory and eyelid function response to shifts in attention from near to more distant objects. 54. An apparatus according to claim 49 whi... | VP UK IP LIMITED | A61B005/00
US20100311482 | A1 | 2010-12-09 | Electro-Biometric Methods and Apparatus | 2010-05-29 | A method and apparatus for tracking changes in an electro-cardiologic signal using a signal processing device having a sensor. The method comprises communicatively coupling the sensor of the signal processing device with a subject; reading into a memory an electro-cardiologic signal of the subject via the sensor and the signal processing device to create an enrolled electro-cardiologic signal; tracking, via the signal processing device, possible variations in the electro-cardiologic signal of the subject over a time period to modify the enrolled electro-cardiologic signal; and adjusting the enrolled electro-cardiologic signal when changes in the electro-cardiologic signal occur. | 1. A method for tracking changes in an electro-cardiologic signal using a signal processing device having a sensor, the method comprising: communicatively coupling the sensor of the signal processing device with a subject;reading into a memory an electro-cardiologic signal of the subject via the sensor and the signal processing device to create an enrolled electro-cardiologic signal;tracking, via the signal processing device, possible variations in the electro-cardiologic signal of the subject over a time period to modify the enrolled electro-cardiologic signal; andadjusting the enrolled electro-cardiologic signal when changes in the electro-cardiologic signal occur. 2. The method of claim 1, wherein the signal processing device is worn on a subject's body. 3. The method of claim 1, wherein the signal processing device is worn on a wrist of the subject. 4. The method of claim 3, wherein the signal processing device is located within a watch. 5. The method of claim 1, further comprising creating an individualized biometric signature of the subject from the electro-cardiologic signal of the subject. 6. The method of claim 5, further comprising monitoring a variation in the individualized biometric signature of the subject, wherein the biometric signature varies based on relative physiological states of the subject. 7. The method of claim 6, further comprising measuring and examining changes in the biometric signature of the subject. 8. The method of claim 7, wherein the measuring and examining are performed using an individualized historical database. 9. The method of claim 7, wherein the changes in the biometric signature relate to a physiological state of the subject. 10. The method of claim 7, wherein the changes in the biometric signature relate to a physiological mood of the subject. 11. The method of claim 7, wherein the changes in the biometric signature relate to a pulse rate of the subject 12. The method of claim 7, wherein the changes in the biometric signature relate to physiological stress of the subject. 13. A method for monitoring a subject, the method comprising: tracking changes in an electro-cardiologic signal using a signal processing device having a sensor; andsimultaneously performing biometric signature identification. 14. The method of claim 13, wherein performing biometric signature identification is initiated subsequently in time to commencement of tracking changes in the electro-cardiologic signal. 15. The method of claim 13, wherein tracking changes in the electro-cardiologic signal is initiated subsequently in time to commencement of performing biometric signature identification. 16. The method of claim 13, wherein tracking changes in an electro-cardiologic signal comprises: communicatively coupling the sensor of the signal processing device with the subject;reading into a memory an electro-cardiologic signal of the subject via the sensor and the signal processing device to create an enrolled electro-cardio... | - | A63F013/00 G06F007/04 A61B005/04
US07848790 | B2 | 2010-12-07 | System and method of imaging using a variable speed for thorax imaging | 2002-02-28 | A system and method of medical imaging using a variable speed patient positioning table are provided. The patient positioning table is configured to operate at a plurality of table speeds during acquisition of data from a selected region, such as the thorax region baying predefined cardiac and non-cardiac regions. In a non-cardiac region, the table is controlled to move at one speed and when the cardiac region is detected, the table is controlled to move at another speed, preferably faster than in the cardiac region to speed data acquisition and eliminate motion artifacts. | 1. A cardiac imaging method comprising the steps of: positioning a patient on a variable speed table of an imaging device;determining, along a common axis, a starting position and an ending position of each of a cardiac scanning region and at least one non-cardiac scanning region to be imaged in a single scan;moving the variable speed table at a primary velocity during imaging of the cardiac scanning region;moving the variable speed table at a secondary velocity, different from the primary velocity, during imaging of the at least one non-cardiac scanning region, wherein the secondary velocity is at a speed greater than the primary velocity;acquiring data during imaging of the cardiac scanning region and the at least one non-cardiac scanning region in the single scan; andreconstructing an image based on the data acquired in the single scan at differing table speeds. 2. The method of claim 1 further comprising the step of transitioning the variable table speed incrementally during changes in the speed of the variable table speed. 3. The method of claim 1 wherein the primary velocity is an EKG gating speed of the patient. 4. The method of claim 3 wherein the EKG gating speed is determined according to a diastole phase of a patient's cardiac cycle. 5. The method of claim 1 further comprising the step of weighting the acquired data according to a distance between a position where each data value is acquired and an image reconstruction location. 6. The method of claim 5 wherein the weighting of the acquired data is reduced as the distance between the position where each data is acquired and the image reconstruction location increases. 7. The method of claim 1 wherein the imaging device includes a CT scanner. 8. A computed tomography system comprising: a high frequency electromagnetic energy projection source to project high frequency energy towards an object;a detector to receive high frequency electromagnetic energy attenuated by the object;a plurality of electrical interconnects configured to transmit detector outputs to a data processing system;a patient positioning table having variable speed capability and configured to operate at a plurality of table speeds during acquisition of data from a selected region, the selected region including a starting position and an ending position of a cardiac scanning region input by a technician and including a non-cardiac scanning region; anda computer programmed to: calculate an image reconstruction point that is centrally positioned about the starting position and the ending position;control the patient positioning table during a single scan to operate at a first table speed in the cardiac scanning region and at a second table speed, different from the first table speed, in the non-cardiac region; andproduce a visual display based upon the detector outputs transmitted to the data processing system about the image reconstruction point. 9. The system of claim 8 wherein the selected region is a thorax reg... | General Electric Company | A61B005/05
US07860733 | B2 | 2010-12-28 | Optimization of timing for data collection and analysis in advanced patient management system | 2010-06-08 | Advanced patient management systems include a unit for collecting data from a device associated with a patient, and a host in communication with the unit, the host identifying a time for the device to update data on the device. The system can identify a time for the device to update data associated with the device by developing a histogram of the availability of the device for interrogation. The system can also optimize processing loads for the system by developing a histogram of the time at which the caregiver is most likely to access the system to review the collected and processed data. Methods for optimizing device data update and processing times are also included. | 1. A machine readable storage medium comprising instructions, which when executed by a machine, cause the machine to: maintain a history of when a caregiver accesses a patient management system;identify a time period during which patient data associated with an ambulatory patient device is forecasted to be accessed by the caregiver based on the history;schedule data analysis of the patient data associated with the ambulatory patient device to provide a scheduled data analysis time, the scheduled data analysis time based on the time period during which the patient data associated with the ambulatory patient device is forecasted to be accessed; andanalyze the patient data at the scheduled data analysis time by: identifying when the patient data from the ambulatory patient device is available to the patient management system; andscheduling the analysis time between when the patient data is available and the time period during which data associated with the patient is forecasted to be accessed by the caregiver, wherein the analysis time is scheduled at least in part to balance computational load at the patient management system. 2. The machine readable medium of claim 1, wherein instructions to identify the time period during which patient data associated with the ambulatory patient device is forecasted to be accessed by the caregiver further comprise instructions to create a histogram of times during which the caregiver accesses the patient data on the patient management system to determine the time period during which the patient data associated with the ambulatory patient device is forecasted to be accessed by the caregiver. 3. The machine readable medium of claim 2, wherein instructions to create a histogram further comprise instructions to: aggregate a plurality of instances of the time period to provide an aggregated value; andcalculate the histogram using the aggregated value. 4. The machine readable medium of claim 3, wherein the plurality of instances of the time period comprises a plurality of consecutive days. 5. The machine readable medium of claim 3, wherein the plurality of instances of the time period comprises a plurality of recurring weekend days. 6. The machine readable medium of claim 3, wherein the plurality of instances of the time period comprises a plurality of recurring days of a week. 7. The machine readable medium of claim 1, wherein the instructions further comprise instructions, which when executed by a machine, cause the machine to recurrently recalculate the time period. 8. The machine readable medium of claim 1, wherein the instructions to identify the time period further comprise instructions to access a time-of-day manually set by the caregiver. 9. A method for gathering and processing patient data associated with an ambulatory patient device of a patient by a patient management system, comprising: identifying, at a computer system, a time period during which the ambulatory patient device is forecasted to be a... | Cardiac Pacemakers, Inc. | G06Q050/00 A61B005/02
US07853306 | B2 | 2010-12-14 | Navigation of remotely actuable medical device using control variable and length | 2005-09-27 | A method of navigating a medical device includes determining the location of a medical device at a point in an operating region in a subject's body, the medical device being responsive to at least one control variable to assume a desired configuration includes storing information representative of the at least one control variable being applied to the medical device at the point, and more preferably storing information representative of the at least one control variable and the device length. | 1. A method of navigating a magnetically navigable medical device to return to a point to which the medical device had previously been magnetically navigated, the method comprising storing information representative of the magnetic field applied to the medical device when the device was at a selected point, and using the stored information representative of the magnetic field to apply the same magnetic field to the medical device to return the medical device to the selected point. 2. The method according to claim 1 further comprising the step of storing information representative of the free length of the medical device when the device was at the selected point, and using the stored information representative of the free length to change the free length of the medical device to return the medical device to the selected point. 3. The method according to claim 2 wherein the information representative of the magnetic field applied to the magnetically navigable device includes the magnetic field direction. 4. The method according to claim 2 wherein the magnetic field is applied with at least one external source magnet, and wherein information representative of the magnetic field applied to the magnetically navigable device includes the position of the at least one external source magnet. 5. The method according to claim 2 wherein the magnetic field is applied with at least one external source magnet, and wherein information representative of the magnetic field applied to the magnetically navigable device includes the orientation of the at least one external source magnet. 6. The method according to claim 2 wherein the magnetic field is applied with at least one external source magnet, and wherein information representative of the magnetic field applied to the magnetically navigable device includes the position and orientation of the at least one magnet. 7. The method according to claim 2 wherein the magnetic field is applied with at least one electromagnet, and wherein information representative of the magnetic field applied to the magnetically navigable device includes the current supplied to the at least one electromagnet. 8. A method of navigating a magnetically navigable medical device substantially along a previously traversed path through a volume in a subject's body, the method comprising: applying a magnetic field to the magnetically navigable medical device to navigate the medical device along a path through the volume in a subject's body;storing information representative of the magnetic field applied to the medical device at each of a plurality of points along the path;subsequently navigating the magnetically navigable medical device substantially along the previously traversed path by successively applying to the magnetically navigable medical device, a magnetic field corresponding to the magnetic field previously applied to the medical device at the same point in the path. 9. The method according to claim 8 further comprising the... | Stereotaxis, Inc. | A61B005/05
US07844318 | B2 | 2010-11-30 | Patient support device, such as a patient bed, table or chair for use with magnetic resonance imaging apparatuses | 2005-03-25 | The invention relates to a patient support device, for use with Magnetic Resonance imaging apparatuses, comprising a support surface (7) adapted to accommodate at least a part, particularly the whole of the patient body. According to the invention, this device comprises at least one, receptacles (307) for holding and/or removably coupling a coil (13′) adapted to receive signals from an anatomic region particularly from the spinal region, which receptacle (307) is arranged over the surface (7) of the patient support device in areas corresponding to the position of the anatomic region to be imaged. The invention also relates to a Magnetic Resonance imaging apparatus comprising a magnet structure (2) which delimits a cavity to receive at least a part of the patient body, in which an imaging volume is created. In accordance with the invention, the apparatus comprises a patient support device as defined in the invention and as described above. The invention further relates to a Magnetic Resonance imaging method providing the use of the apparatus and the patient support device according to the invention. | 1. A patient support device for use with Magnetic Resonance Imaging apparatuses, the device comprising: a patient support surface having such a size as to accommodate at least a part of the patient body;at least one receptacle for housing and/or removably coupling a coil adapted to receive signals from anatomic regions upon excitation thereof by the Magnetic Resonance Imaging apparatus, said at least one receptacle being arranged in an area of the patient support surface corresponding to a position of an anatomic region to be imaged, wherein the receptacle for housing and/or removably coupling the coil is formed by a niche or an enclosure with an open upper side extending flush with the patient support surface and wherein the coil is completely embedded inside the receptacle without protruding above the patient support surface;a mattress provided on a top face of the support surface, the mattress having one or more holes therein;a base of each coil has an axial extension for introduction into the respective receptacle through one of the holes in the mattress;and one or more closing members for closing any of the holes in the mattress through a coil axial extension is not extending. 2. A device as claimed in claim 1, wherein the receptacle for housing/coupling the receiver coil is accessible from the patient support surface side. 3. A device as claimed in claim 1, wherein the receptacle for housing/coupling the receiver coil includes means for removable mechanical connection of the receiver coil. 4. A device according to claim 1, wherein each receptacle provides hidden housings for said coils that are accessible from outside the device and beneath the patient support surface, whereby the coils are hiddenly received beneath the patient support surface. 5. A device according to claim 1, wherein the receiving coil is adapted for imaging of the spine of the patient. 6. A device as claimed in claim 1, further comprising a magnet structure of a Magnetic Resonance imaging apparatus, wherein the patient support device is configured to be moved into a predetermined imaging volume within a cavity defined by said magnet structure, said imaging volume being a fixed reference for limiting the displacement of the support device or at least the support surface in one or two transverse, orthogonal directions, and wherein motor-driven displacement means and/or displacement stroke limiting means and/or means for detecting the position of at least the support surface and/or means for detecting the stroke in one or both displacement directions are controlled in such a manner that the displacement stroke limiters and/or the motor-driven displacement means are actuated so that the support surface is automatically stopped when the receiver coil, whose presence is detected in one of the coupling receptacles, is coincident with or at least partly contained in the imaging volume. 7. A device as claimed in claim 6, further comprising a footrest situated at the patient... | Esaote S.p.A. | A61B005/05
US07833964 | B2 | 2010-11-16 | Methods of reducing visceral fat by increasing levels of insulin-like growth factor-I (IGF-I) | 2004-08-20 | The present invention provides methods and compositions for reducing visceral fat by administering to the subject a therapeutically effective amount of a compound that increases the bioactive serum levels of insulin-like growth factor-I (IGF-I) in the subject, thereby ameliorating negative effects of visceral obesity. The invention is useful in the treatment, prevention, or amelioration of one or more symptoms of visceral obesity or IGF-I deficiency related condition, including, for example, cardiovascular disease and the metabolic syndrome. | 1. A method for reducing visceral fat in a subject having an excess of visceral fat, the method comprising: a) diagnosing the subject as having an excess of visceral fat; andb) administering to the subject a compound that increases the bioactive serum levels of insulin-like growth factor-I (IGF-I) in an amount effective to reduce visceral fat in said subject, wherein the compound is not growth hormone. 2. The method of claim 1, wherein either (i) the subject is an adult human male and is type II diabetic and has a waist size of at least 40 inches before treatment with such compound or (ii) the subject is an adult human female and is type II diabetic and has a waist size of at least 35 inches before treatment with such compound. 3. The method of claim 2, wherein the reduction in visceral fat of the subject is assessed by determining the ratio of visceral fat to subcutaneous fat. 4. The method of claim 2, wherein the reduction in visceral fat of the subject is assessed by a decrease in a ratio of waist measurement to hip measurement of the subject. 5. The method of claim 2, wherein the reduction in visceral fat of the subject is assessed by computer tomography (CT) scan. 6. The method of claim 2, wherein the compound is IGF-I. 7. The method of claim 6, wherein the IGF-I is complexed with insulin-like growth factor binding protein-3 (IGFBP-3). 8. The method of claim 6, wherein the amount of IGF-I is greater than or equal to 40 μg/kg/day. 9. The method of claim 6, wherein said IGF-I is administered subcutaneously. 10. The method of claim 2, wherein the compound is administered with an effective amount of a growth-promoting agent. 11. The method of claim 10, wherein the growth-promoting agent is an insulin-like growth factor binding protein (IGFBP). 12. The method of claim 11, wherein the IGFBP is IGFBP-3. 13. The method of claim 10, wherein the compound is IGF-I. 14. The method of claim 1, wherein the subject is an adult human that presents, before treatment with such compound, a body composition characterized by a ratio of visceral fat to subcutaneous fat that is higher than the ratio of visceral fat to subcutaneous fat that characterizes the average body composition of normal adult humans of the same body mass index (BMI) and gender as the subject. 15. The method of claim 1, wherein the subject is an adult subject that presents, before treatment with such compound, a condition characterized by type II diabetes and three or more symptoms selected from the group consisting of: (1) a waist size that is greater than 40 inches if the subject is a male and greater than 35 inches if the subject is a female, (2) a fasting blood glucose level greater than or equal to 110 mg/dL, (3) a fasting triglyceride level of greater than or equal to 150 mg/dL, (4) a high-density lipoprotein (HDL) level of less than 40 mg/dL if the subject is a male and less than 50 mg/dL if the subject is a female, and (5) a blood pressure level of greater than 130/85. 1... | Tercica, Inc. | A61K038/00 A61B005/55
US07837648 | B2 | 2010-11-23 | Medicament dispensing system | 2002-11-21 | A medicament dispensing system for treatment of a disease is disclosed, the system comprising: (i) detecting means for detecting one or more biochemical marker(s) indicative of the disease; (ii) electronic processing means for processing information obtainable from the detecting means; and (iii) medicament dispensing means for dispensing a quantity of medicament determinable by the electronic processing means. The medicament dispensing system is particularly useful for the treatment of respiratory disorders, such as asthma, or for other disorders which may be treated by inhalation, either through the nose or mouth. | 1. A hand-held inhalation device for treatment of a disease by inhalation of medicament through the mouth or nose of a patient comprising: (i) a medicament dispenser for dispensing a quantity of medicament,(ii) a detector for detecting one or more biochemical marker(s) in an exhaled breath of the patient indicative of said disease, the detector comprising a sampler into which the patient is able to exhale, and;(iii) an electronic processor for processing information obtainable from the detector;wherein said medicament dispenser dispenses a quantity of medicament determinable by the electronic processor, and; said medicament dispenser, detector and electronic processor are integrated components of the handheld inhalation device. 2. An inhalation device according to claim 1, wherein the detector comprises a sensor. 3. An inhalation device according to claim 2, wherein the sensor is a non-specific bio- or chemosensor. 4. An inhalation device according to claim 3, wherein the detector further comprises a pre-separation device. 5. An inhalation device according to claim 4, wherein the pre-separation device comprises a miniaturised gas chromatograph device, a capillary electrophoresis chip or other electrophoresis device, or a lateral flow membrane. 6. An inhalation device according to claim 2, wherein the sensor comprises a biosensor. 7. An inhalation device according to claim 6, wherein said biosensor comprises a sensing layer and a transducer layer. 8. An inhalation device according to claim 7, wherein said sensing layer comprises one or more materials selected from the group consisting of enzymes, receptors, micro-organisms, antibodies, antigens, nucleic acids, cells, peptides, proteins and biochemicals. 9. An inhalation device according to claim 7 wherein the sensing layer comprises one or more artificial or synthetic materials analogous to enzymes, receptors, micro-organism, antibodies, antigens, nucleic acids, cells, peptides, proteins and biochemicals. 10. An inhalation device according to claim 7, wherein the sensing layer comprises a renewable sensing surface. 11. An inhalation device according to claim 7, wherein the sensing layer comprises a non-renewable sensing surface. 12. An inhalation device according to claim 11, wherein the sensing layer comprises a disposable component incorporating the non-renewable sensing surface. 13. An inhalation device according to claim 9, wherein said transducer layer comprises an optical, electrical, semi-conductor, thermal or acoustic mechanism. 14. An inhalation device according to claim 1, wherein the detector comprises spectroscopic detecting means for direct spectroscopic detection of one or more biochemical marker(s). 15. An inhalation device according to claim 14, wherein the spectroscopic detector is suitable for use in infra red and Raman spectroscopic detection methods. 16. An inhalation device according to claim 1, wherein said device comprises more than one detector for detecting ... | Glaxo Group Limited | A61M031/00 A61M011/00 A61B005/08
US07825815 | B2 | 2010-11-02 | Apparatus, systems, and methods for gathering and processing biometric and biomechanical data | 2010-01-29 | Apparatus, systems, and methods are provided for measuring and analyzing movements of a body and for communicating information related to such body movements over a network. In certain embodiments, a system gathers biometric and biomechanical data relating to positions, orientations, and movements of various body parts of a user performed during sports activities, physical rehabilitation, or military or law enforcement activities. The biometric and biomechanical data can be communicated to a local and/or remote interface, which uses digital performance assessment tools to provide a performance evaluation to the user. The performance evaluation may include a graphical representation (e.g., a video), statistical information, and/or a comparison to another user and/or instructor. In some embodiments, the biometric and biomechanical data is communicated wirelessly to one or more devices including a processor, display, and/or data storage medium for further analysis, archiving, and data mining. In some embodiments, the device includes a cellular telephone. | 1. An apparatus for gathering, transmitting, and displaying and/or storing body movement data, the apparatus comprising: a first wearable sensor comprising at least two of an accelerometer, a gyroscope, and a magnetometer, the first sensor configured to gather and transmit position and movement data for a first location on a wearer's body, the first sensor configurable to change a first sampling rate at which the data is gathered and/or transmitted;a second wearable sensor comprising at least two of an accelerometer, a gyroscope, and a magnetometer, the second sensor configured to gather and transmit position and movement data for a second location on a wearer's body, the second sensor having a second sampling rate at which the data is gathered and/or transmitted;a third wearable sensor comprising at least two of an accelerometer, a gyroscope, and a magnetometer, the third sensor configured to gather and transmit position and movement data for a third location on a wearer's body, the third sensor having a third sampling rate at which the data is gathered and/or transmitted;a first processor configured to receive wireless transmission of position and movement data from each of the three wearable sensors, determine the relative movement of the three sensors, optionally display representations of that movement, and provide that data to a database,wherein the first processor is further configured to allow independent selection of the first and second sampling rates and allow viewing of a three-dimensional animation of body movement that resulted in the data, the animation created by the processor interpolating between the first, second, and third locations on the user's body to show an avatar;wherein the database is configured to: store the data in raw and/or processed form for later analysis and comparison;allow access to users to compare stored data gathered at one time or from one user to data gathered at another time or from another user; andallow access to professionals to further analyze the data and provide analysis to the user. 2. The apparatus of claim 1, wherein the first processor is further configured to access the database and display animations of a user's body movement compared with animations of another person's body movement. 3. The apparatus of claim 1, wherein the database is further configured to store historical data and the first processor is further configured to show trends or patterns in the user's movements. 4. The apparatus of claim 1, wherein the database is further configured to store body movement data that has been normalized and averaged over a selected group to allow a benchmark for comparison to the user's body movement data. 5. The apparatus of claim 1, wherein the professionals are doctors, coaches, or physical therapists. 6. The apparatus of claim 1, wherein the first processor is selected from the group consisting of a laptop computer, smart phone, PDA, or golf interface box. 7. A motion sensing system for ... | Applied Technology Holdings, Inc. | G08B023/00 A63B069/36 A61B005/22
US08134367 | B2 | 2012-03-13 | Irradiating coil and magnetic resonance imaging apparatus using the same | 2008-04-04 | An irradiating coil includes a first ring conductor located on a plane, a second ring conductor having a smaller diameter than a diameter of the first ring conductor and located substantially on the plane of the first ring conductor, a plurality of first line conductors connected the first ring conductor and to the second ring conductor, and an arrangement for making uniform a radio-frequency-pulse that is transmitted by the irradiating coil. The pulse uniforming arrangement includes a diverger which diverges a current flow of the first line conductors. | 1. An irradiating coil comprising: a first ring conductor located on a plane;a second ring conductor having a smaller diameter than a diameter of the first ring conductor and located substantially on the plane of the first ring conductor;a plurality of first line conductors connected the first ring conductor and to the second ring conductor; andmeans for making uniform a radio-frequency-pulse that is transmitted by the irradiating coil;wherein the pulse uniforming means includes diverging means for diverging a current flow of the first line conductors. 2. An irradiating coil according to claim 1, wherein the diverging means includes at least one branch from the first line conductor to the second ring conductor. 3. An irradiating coil according to claim 2, wherein the branch is provided on both sides of each first line conductor. 4. An irradiating coil according to claim 2, wherein the branch is for bypassing a current flow between the first ring conductor and a connecting point of the first line conductor with the branch. 5. An irradiating coil according to claim 2, wherein the branch has a plurality of connecting points with the first line conductor. 6. An irradiating coil according to claim 2, wherein the branch has a second line conductor connected to the second ring conductor. 7. An irradiating coil according to claim 6, wherein each current flowing on the second line conductor and the first line conductor between the second ring conductor and a connecting point of the first line conductor with the branch is substantially a same current. 8. An irradiating coil according to claim 2, wherein two adjacent branches are disposed between two adjacent first line conductors and are electrically connected to each other via at least one resonant capacitor. 9. An irradiating coil according to claim 8, wherein the electrical connections between the branches and the resonant capacitors build up a circuit in a circumferential direction between the first ring conductor and the second ring conductor. 10. An MRI apparatus comprising: a static magnetic field generating means for applying a static magnetic field to an object to be examined;a gradient magnetic field coil for applying to the object a slicing gradient magnetic field, a readout gradient magnetic field, and an encoding gradient magnetic field;an irradiating coil for applying to the object an irradiation pulse for generating magnetic resonance, including a first ring conductor located on a plane, a second ring conductor having a smaller diameter than a diameter of the first ring conductor and located substantially on the plane of the first ring conductor, and a plurality of first line conductors connecting the first ring conductor and the second ring conductor;a receiving coil for detecting a magnetic resonance signal; andimage reconstructing means for obtaining an image of the object using the signal detected by the receiving coil;wherein the irradiating coil includes means for making the i... | Hitachi Medical Corporation | G01V003/00
US08783260 | B2 | 2014-07-22 | Apparatuses and methods for mandibular protrusion | 2010-06-04 | Embodiments of methods and mandibular protruders and positioning devices, comprising, for example: an upper mounting bracket having an upper dental appliance; a lower mounting bracket having a lower dental appliance; and a rail system configured to couple the upper mounting bracket to the lower mounting bracket. Some embodiments comprise a drive motor and/or an initial position adjustment mechanism. | 1. A mandibular protruder comprising, an upper mounting bracket having an upper dental appliance and a portion that extends from the upper dental appliance, the upper dental appliance being configured for use with dental impression material;a lower mounting bracket having a lower dental appliance and a portion that extends from the lower dental appliance, the lower dental appliance being configured for use with dental impression material;wherein at least a portion of each extending portion is configured to extend outside the mouth of a subject when the upper and lower dental appliances are positioned in the mouth of the subject;a rail system coupling the extending portions such that relative motion of the extending portions is substantially prevented in a lateral direction, wherein the rail system comprises an engaged portion of the upper and lower extending portions; andwherein the upper mounting bracket or the lower mounting bracket is configured to be coupled to a drive-motor connector of a drive motor, the drive-motor connector having a longitudinal axis that is substantially parallel to the direction of actuation of the drive motor, such that the longitudinal axis of the upper mounting bracket or the lower mounting bracket is substantially parallel to the longitudinal axis of the drive-motor connector. 2. A mandibular protruder comprising, an upper mounting bracket having an upper dental appliance and a portion that extends from the upper dental appliance, the upper dental appliance being configured for use with dental impression material;a lower mounting bracket having a lower dental appliance and a portion that extends from the lower dental appliance, the lower dental appliance being configured for use with dental impression material;wherein at least a portion of each extending portion is configured to extend outside the mouth of a subject when the upper and lower dental appliances are positioned in the mouth of the subject;a rail system coupling the extending portions such that relative motion of the extending portions is substantially prevented in a lateral direction, wherein the rail system comprises an engaged portion of the upper and lower extending portions; andwherein the rail system comprises a portion of one of the lower mounting bracket and the upper mounting bracket that is configured to wrap around a portion of the other of the lower mounting bracket and the upper mounting bracket. 3. The mandibular protruder of claim 2, wherein the portion of one of the upper mounting bracket and the lower mounting bracket that is configured to wrap around the other of the upper mounting bracket and the lower mounting bracket is a portion of the lower mounting bracket that is configured to wrap around a portion of the upper mounting bracket. 4. The mandibular protruder of claim 3, wherein the portion of the lower mounting bracket that is configured to wrap around the portion of the upper mounting bracket is coupled in fixed relation to the ... | ZST Holdings Inc. | A61F005/56
US08940298 | B2 | 2015-01-27 | High affinity anti-prostate stem cell antigen (PSCA) antibodies for cancer targeting and detection | 2008-09-04 | The present invention provides novel high affinity antibodies and fragments thereof that bind to the cancer antigen PSCA. The antibodies of the present invention may be used for cancer diagnosis, prognosis, treatment, visualization, and the like. The present invention also provides methods for the detection, visualization, and treatment of various cancers expressing PSCA. | 1. An antigen binding protein construct that binds to human prostate stem cell antigen (PSCA) comprising an antigen binding domain comprising a heavy chain variable domain comprising the complementarity determining regions (CDRs) of the heavy chain variable domain of SEQ ID NO: 7 and a light chain variable domain comprising the CDRs of the light chain variable domain of SEQ ID NO: 7. 2. The antigen binding protein construct of claim 1, wherein said antigen binding domain comprises a light chain variable domain comprising the light chain variable domain of SEQ ID NO: 7. 3. The antigen binding protein construct of claim 1, wherein said antigen binding domain comprises a heavy chain variable domain comprising the heavy chain variable domain of SEQ ID NO: 7. 4. The antigen binding protein construct of claim 1, wherein the construct is a minibody. 5. The antigen binding protein construct of claim 4, wherein said minibody comprises an amino acid sequence that is at least 95% identical to SEQ ID NO: 12. 6. The antigen binding protein construct of claim 5, wherein said minibody comprises the amino acid sequence of SEQ ID NO: 12. 7. The antigen binding protein construct of claim 1, wherein the construct is a diabody. 8. The antigen binding protein construct of claim 1, wherein the construct is a humanized antibody. 9. The antigen binding protein construct of claim 1, wherein the construct is an affinity matured antibody. 10. The antigen binding protein construct of claim 1, wherein the construct is conjugated to a therapeutic agent. 11. The antigen binding protein construct of claim 10, wherein the therapeutic agent is a cytotoxic agent. 12. The antigen binding protein construct of claim 11, wherein the cytotoxic agent is selected from a group consisting of ricin, ricin A-chain, doxorubicin, daunotubicin, paclitaxel, ethiduim bromide, mitomycin, etoposide, tenoposide, vincristine, vinblastine, colchicine, dihydroxy anthracin dione, actinomycin D, diphteria toxin, Pseudomonas, exotoxin (PE) A, PE40, abrin, arbrin A chain, modeccin A chain, alpha-sarcin, gelonin mitogellin, retstrictocin, phenomycin, enomycin, curicin, crotin, calicheamicin, sapaonaria officinalis inhibitor, maytansinoids, and glucocorticoidricin. 13. The antigen binding protein construct of claim 10, wherein the therapeutic agent is a radioactive isotope. 14. The antigen binding protein construct of claim 13, wherein the radioactive isotope is 212Bi, 131I, 111In, 90Y or 186Re. 15. The antigen binding protein construct of claim 1, wherein the construct is linked to an anti-cancer pro-drug activating enzyme capable of converting a pro-drug to its active form. 16. The antigen binding protein construct of claim 1, which is labeled with a detectable marker. 17. The antigen binding protein construct of claim 16, wherein the detectable marker is a radioisotope, a fluorescent compound, a bioluminescent compound, a chemiluminescent compound, a metal chelator or an enzyme. 18. The a... | The Regents of the University of California | A61K039/395 C07K016/30
US08412299 | B2 | 2013-04-02 | Biological information measurement device capable of accurately conducting measurement, that can safely be used | 2009-12-23 | A biological information measurement device is constituted of an upper cover and a lower cover and a main body portion. Then, the upper cover and the lower cover are assembled such that they partially overlap with each other. A fingertip insertion portion is formed of a holding member, on which a substrate and a line of the main body portion are placed. By covering the whole main body by assembling the upper cover and the lower cover, the substrate and the line of the main body portion arranged on a side surface of the main body in the inside cannot externally be seen. | 1. A biological information measurement device provided with a light emission portion for emitting light to an inserted fingertip and a light reception portion for receiving transmitted light, comprising: a first holding portion;a second holding portion forming a fingertip insertion recess together with said first holding portion, wherein a fingertip is to be inserted into the recess, the second holding portion is opposed to the first holding portion with said fingertip insertion recess lying therebetween, and the second holding portion being capable of relative displacement with respect to the first holding portion in a direction of opening and closing of a fingertip insertion port of said fingertip insertion recess; anda restriction member for restricting relative displacement of said first and second holding portions to avoid any exposed gap between the first holding portion and the second holding portion, except for the fingertip insertion port of said fingertip insertion recess, forming around the fingertip insertion port of said fingertip insertion recess when said fingertip insertion portion is opened. 2. The biological information measurement device according to claim 1, wherein said first holding portion is constituted of a first holding member and a first housing provided to cover said first holding member,said second holding portion is constituted of a second holding member and a second housing provided to cover said second holding member,an electronic component implementing said light emission portion is provided in at least one of said first and second holding members, andan electronic component implementing said light reception portion is provided in at least one of said first and second holding members. 3. The biological information measurement device according to claim 1, wherein each of said first and second holding portions has side portions arranged on respective sides of said fingertip insertion portion,the side portions of said first holding portion overlap with the side portions of said second holding portion when said fingertip insertion portion is closed, andsaid restriction member restricts relative displacement of said first and second holding portions at a position where such a state that the side portions of said first holding portion and the side portions of said second holding portion overlap with each other when said fingertip insertion portion is opened is maintained. 4. The biological information measurement device according to claim 1, wherein said first holding portion includes a wall portion extending across said fingertip insertion recess. 5. A biological information measurement device provided with a light emission portion for emitting light to an inserted fingertip and a light reception portion for receiving transmitted light, comprising: a first holding portion;a second holding portion forming with said first holding portion a fingertip insertion recess to receive a fingertip, the second holding portion... | Nintendo Co., Ltd. | A61B005/1455
US08509869 | B2 | 2013-08-13 | Method and apparatus for detecting and analyzing variations in a physiologic parameter | 2009-05-15 | The present disclosure is generally directed to identifying and/or analyzing high resolution variations in a measured physiologic parameter, such as blood oxygen saturation (SpO2) measured using pulse oximetry. Present embodiments may include a system including a sensor comprising an emitter capable of emitting light at different wavelengths into a tissue bed, and a detector capable of detecting the light from the emitter after dispersion and/or reflection by the tissue bed. Further, the system may include a pulse oximeter capable of receiving signals from the sensor that are indicative of characteristics of the light detected by the detector, and utilizing the signals to estimate blood oxygen saturation values over time at a high resolution to facilitate detection of variations in the blood oxygen saturation values that are smaller in magnitude than an accuracy, display precision, and/or calibration of the blood oxygen saturation values. | 1. A pulse oximetry system, comprising: a sensor comprising: an emitter configured to emit light at different wavelengths into a tissue bed; and a detector configured to detect the light from the emitter after dispersion and/or reflection by the tissue bed; anda pulse oximeter configured to: receive signals from the sensor that are indicative of characteristics of the light detected by the detector;utilize the signals to estimate blood oxygen saturation values over time at a high resolution;detect variations in the blood oxygen saturation values that are smaller in magnitude than an accuracy, display precision, and/or calibration of the blood oxygen saturation values; andperform control and/or pattern detection functions based on the variations. 2. The pulse oximetry system of claim 1, wherein the pulse oximeter is configured to estimate the blood oxygen saturation values with a precision of 0.1%, 0.01% or 0.001%. 3. The pulse oximetry system of claim 1, wherein the pulse oximeter is configured to report changes in a value of partial pressure of oxygen in blood based on the detection of variations in the blood oxygen saturation values. 4. The pulse oximetry system of claim 1, wherein the emitter is configured to emit light having a wavelength in a bounded range of 600-660 nm. 5. The pulse oximetry system of claim 4, wherein the emitter comprises at least two light emitting diodes, wherein one of the at least two light emitting diodes is configured to emit light in the bounded range of 600-620 nm. 6. A method of performing pulse oximetry, comprising: emitting light from an emitter component of a pulse oximeter sensor at first and second wavelengths useful for measuring oxygen saturation;detecting the light with a detector component of the pulse oximeter sensor;transmitting signals indicative of characteristics of the detected light to a pulse oximeter;estimating blood oxygen saturation values over time at a high resolution based on the signals with the pulse oximeter;detecting variations in the blood oxygen saturation values that are smaller in magnitude than an accuracy, display precision, and/or calibration of the blood oxygen saturation values; anddetecting a patient condition and/or controlling a physiologic parameter based on the variations. 7. The method of claim 6, wherein the first wavelength is between 650-670 nm. 8. The method of claim 6, wherein the first wavelength is within a range of 600-620 nm and the second wavelength is within a range of 800-940 nm. 9. The method of claim 6, comprising controlling a partial pressure of oxygen in a patient's blood based on the variations in the blood oxygen saturation values. 10. The method of claim 9, wherein controlling the partial pressure of oxygen comprises controlling delivery of oxygen to the patient based on the variations in the blood oxygen saturation values. 11. The method of claim 6, comprising identifying a pattern indicative of sleep apnea based on the variations in the bloo... | Covidien LP | A61B005/00
US08452612 | B2 | 2013-05-28 | Method for evaluating and prognosticating the daily emotive behavior states and psychophysiological activity of a person according to the measures of night hypersympathicotonia syndrome | 2008-12-26 | The invention relates to the field of identifying the results of medical measurements. The use thereof ensures the technical result in the form of a possibility for estimating confidently and prognosticating reliably the emotional-behavioral states and psychophysiological functioning of a human by means of data obtained as a result of studying the cardiac rhythm thereof. This result is achieved owing to the method including step of: a) performing an intraday monitoring of the patient's ECG; b) storing the data of this monitoring in the computer memory; c) carrying out a computer processing of the data stored in the step b) for constructing, in accordance with those data, an intervalogram of the patient's cardiac rhythm variability per 24-hour period; d) determining, in accordance with the intervalogram constructed in the step c), those spans corresponding to the nighttime and having at least the predetermined length, where the cardiac rhythm variability is reduced in comparison with the average daily characteristics at the intervalogram spans having the same length and corresponding to the daytime, herewith said reduction of the cardiac rhythm variability shows that the given patient has a night hypersympathicotonia syndrome; e) comparing the differences determined in the step d) with the values from the preformed correspondence set stored in the computer memory, thus defining a degree of manifestation of the predefined indices of the night hypersympathicotonia syndrome; and f) performing, on the basis of the comparison made in the step e), predictive estimations of the human day emotional-behavioral states and psychophysiological functioning. | 1. A method for predicting human emotional states and psychophysiological functioning by measuring indices of night hypersympathicotonia syndrome, the method comprising steps of: a) monitoring a patient's ECG using an ECG device for a period of at least approximately 24 hours and recording the patient's ECG data;b) entering and storing the patient's ECG data in computer memory;c) constructing an intervalogram using a computer processor and using the patient's ECG data, and storing the intervalogram in computer memory;d) selecting using a computer a night time span in the intervalogram corresponding to nighttime, and selecting using a computer a day time span in the intervalogram corresponding to daytime, wherein both time spans have an equal duration, the durations being not less than 30 minutes;e) comparing cardiac rhythm variability of the day time span and of the night time span and determining a patient cardiac variability profile comprising that comparison;f) providing a comparison set in computer memory, the comparison set comprising physiological data indicative of a presence or absence of night hypersympathicotonia syndrome, including daytime and nighttime cardiac rhythm variability data, for a plurality of comparison individuals, some of whom have night hypersympathicotonia syndrome and some of whom do not have night hypersympathicotonia syndrome, the comparison set also including psychological characteristics for the comparison individuals determined using psychological testing; andg) comparing the patient cardiac variability profile and the comparison set, and predicting the patient's psychological characteristics based on the physiological and psychological characteristics of the comparison individuals contained in the comparison set. 2. The method of claim 1, wherein the comparison set is created ahead of time by a process comprising monitoring and recording the comparison individuals' ECG for a period of at least approximately 24 hours, and performing psychological testing of each comparison individual after their ECG monitoring ends. 3. The method of claim 2, further comprising creating an intervalogram for each comparison individual. 4. The method of claim 3, the process for creating the comparison set further comprising selecting a night time span in each comparison individual intervalogram corresponding to nighttime, and selecting a day time span in each comparison individual intervalogram corresponding to daytime, wherein both time spans have an equal duration, the duration being not less than 30 minutes; for each comparison individual intervalogram, comparing the cardiac rhythm variability of the day time span and the cardiac rhythm variability night time span to determine a cardiac variability ratio;correlating the cardiac variability ratios for the plurality of comparison individuals with the psychological characteristics of the plurality of comparison individuals and including said correlations in the comparison set. 5.... | - | G06Q010/00 G06Q050/00
US08690929 | B2 | 2014-04-08 | Orthopedic screw for measuring a parameter of the muscularskeletal system | 2010-06-29 | A dual-mode closed-loop measurement system (100) for capturing a transit time, phase, or frequency of energy waves propagating through a medium (122) is disclosed. A first module comprises an inductor drive circuit (102), an inductor (104), a transducer (106), and a filter (110). A second module housed in a screw (335) comprises an inductor (114) and a transducer (116). The screw (335) is bio-compatible and allows an accurate delivery of the circuit into the muscular-skeletal system. The inductor can be attached and interconnected on a flexible substrate (331) that fits into a cavity in the screw (335). The first and second modules are operatively coupled together. The first module provides energy to power the second module. The second module emits an energy wave into the medium that propagates to the first module. The transit time of energy waves is measured and correlated to the parameter by known relationship. | 1. An orthopedic screw having threads thereon configured to couple to the muscular-skeletal system comprising: a screw including a cavity; anda passive circuit within the cavity where the passive circuit is configured to emit an energy wave to the screw and where the screw is configured to couple the energy wave to a medium;an inductor; and a transducer operatively coupled to the inductor; andfurther including a flexible substrate where the inductor and transducer are mounted to the substrate and where the substrate includes interconnect for coupling the inductor to the transducer. 2. The orthopedic screw of claim 1 where a die of the inductor and transducer is stacked. 3. The orthopedic screw of claim 1 where a flexible substrate is altered from a planar shape to fit in the cavity of the screw. 4. The orthopedic screw of claim 1 further including a module configured to be in closed-loop feedback with the inductor and the transducer of the orthopedic screw. 5. The orthopedic screw of claim 4 where the module is configured to be coupled in a first path to the inductor of the orthopedic screw electromagnetically. 6. The orthopedic screw of claim 5 where the module is configured to be coupled in a second path to the transducer of the orthopedic screw ultrasonically. 7. The orthopedic screw of claim 6 where the closed-loop feedback between the module and the orthopedic screw forms a propagation tuned oscillator for a continued emission, propagation, and detection of energy waves in the medium and where an energy wave transit time, frequency, or phase is monitored over a predetermined distance through the medium. 8. The orthopedic screw of claim 1 where the orthopedic screw is configured to be used to measure bone density of a muscular-skeletal system. 9. An orthopedic screw comprising: a screw having a cavity therein; anda passive circuit comprising an inductor operatively coupled to a transducer within the cavity where the transducer is configured to emit an energy wave to the screw and where the screw is configured to couple the energy wave to a medium;where the transducer is an ultrasonic transducer; andfurther including a flexible substrate where the inductor and transducer are mounted to the substrate and where the substrate includes interconnect for coupling the inductor to the transducer. 10. The orthopedic screw of claim 7 where the transducer is coupled to the screw and where the screw is configured to radiate an ultrasonic signal. 11. The orthopedic screw of claim 10 where the flexible substrate is altered from a planar shape to fit in the cavity of the screw. 12. The orthopedic screw of claim 11 where a die of the inductor and transducer is stacked. 13. The orthopedic screw of claim 9 further including a module configured in closed-loop feedback with the inductor and the transducer of the orthopedic screw. 14. The orthopedic screw of claim 13 where the module is configured to couple in a first path to the inductor of the ortho... | Orthosensor Inc. | A61B017/86 A61N001/00 A61B005/05
US09283339 | B2 | 2016-03-15 | Life support and monitoring apparatus with malfunction correction guidance | 2010-05-14 | A life support and monitoring apparatus with malfunction correction guidance is provided. The life support and monitoring apparatus of the present disclosure identifies the root cause or potential cause of a fault/failure and then prompts an operator to take appropriate steps to assure the continuance of life support and critical physiologic monitoring. When multiple faults/failures exist, the apparatus automatically prioritizes them based on risk to the patient and prompts the operator to do the most appropriate intervention to assure patient safety. | 1. A ventilation apparatus comprising: at least one sensor for sensing at least one physiological parameter of a patient;a ventilator module for providing mechanical ventilation and supplying an adjustable level of inspired oxygen to the patient;a controller for controlling the ventilator module to achieve the adjustable level of inspired oxygen based on the sensed at least one physiological parameter of the patient;a user interface for communicating events to an operator; andat least one memory for storing a plurality of event codes and data associated with each event code,wherein the controller determines if an event has occurred based at least in part on the at least one sensed physiological parameter of the patient, wherein if the controller has determined an event has occurred, the controller determines the event code for the event that was determined to have occurred and retrieves the associated data from the memory based on the determined event code and causes the user interface to generate an indication of the associated data,wherein the associated data of each event comprises at least a description and mitigation instructions,wherein the mitigation instructions are presented as a list to resolve the event,wherein, if multiple events occur, the controller is configured to prioritize the events based on risk to the patient and indicate the multiple events on the user interface from the event with the highest priority to the event with the lowest priority and to update the user interface as new events occur such that the event with the highest priority is indicated first, andwherein the controller is further configured to lock out indication of at least one lower priority event such that the lower priority event cannot be presented until the highest priority event is rectified. 2. The ventilation apparatus as in claim 1, wherein the associated data of each event further comprises at least an alarm name. 3. The ventilation apparatus as in claim 1, wherein the associated data of each event further comprises if not resolved instructions and the if not resolved instructions comprise indicators to distinguish the if not resolved instructions from the list of mitigation instructions. 4. The ventilation apparatus as in claim 3, wherein the if not resolved instructions are presented after an end of the list of mitigation instructions. 5. The ventilation apparatus as in claim 1, wherein the mitigation instructions are presented in the form of a video or audible message. 6. The ventilation apparatus as in claim 1, further comprising a communication interface for communicating with an external computing device, wherein the external computing device stores additional associated data. 7. The ventilation apparatus as in claim 1, wherein the sensed at least one physiologic parameter includes oxyhemoglobin saturation (SpO2), end-tidal CO2 (ETCO2), heart rate, blood pressure and temperature. 8. An event correction guidance method for a life support a... | ZOLL MEDICAL CORPORATION | A61M016/00 A61B005/00 A61B005/08 A61M016/10
US10376213 | B2 | 2019-08-13 | System, method and apparatus for sensor insertion | 2009-06-30 | Embodiments provide a sensor insertion tool (SIT) that provides a motive force for insertion of an analyte sensor into/through skin. A SIT may be releasably locked to one or more components of a sensor insertion system, such that components of the sensor insertion system remain securely coupled during sensor insertion. A SIT may include a release member that unlocks or uncouples the SIT and the other components after sensor insertion. In various embodiments, a SIT may be a component of a sensor insertion system configured for assembly by an end user, a health care professional, and/or a caretaker prior to sensor insertion, and may act in cooperation with other sensor insertion system components. Additional components and methods of assembly and use are also provided herein. | 1. A sensor insertion tool for inserting a sensor through intact skin without a cannula or trocar, comprising: a channel guide assembly, comprising a channel guide, anda channel guide cover, wherein the channel guide and the channel guide cover together form a channel guide assembly channel;a motive force assembly to store and release energy, comprising a motive force member,a hammer channel separate from and non-contiguous with the channel guide assembly channel, and having a curved portion,a hammer extending through the channel guide cover into the channel guide assembly channel, anda hammer pin coupled to and extending transverse from the hammer, the hammer pin configured to extend from the hammer into the hammer channel and move along the curved portion of the hammer channel to guide the hammer to strike a sensor subassembly positioned within the channel guide assembly channel;a cocking assembly coupled to the motive force assembly to load energy into the motive force assembly; anda trigger assembly having a trigger member coupled to the motive force assembly to release stored energy in response to actuation of the trigger member. 2. The sensor insertion tool of claim 1, wherein the sensor insertion tool is disposable. 3. The sensor insertion tool of claim 1, wherein the channel guide assembly is permanently coupled to the motive force assembly. 4. A sensor insertion tool for inserting a sensor through intact skin without a cannula or trocar, comprising: a motive force assembly comprising a motive force member,a curved track,a hammer, anda hammer pin coupled to and extending transverse from the hammer into the curved track, the hammer pin configured to move along a curved portion of the curved track to guide the hammer to strike a sensor subassembly within a channel guide assembly channel separate from and non-contiguous with the curved track, the channel guide assembly comprising a channel guide, and a channel guide cover that together form the channel guide assembly channel;a cocking assembly coupled to the motive force assembly to load energy into the motive force assembly;a trigger assembly coupled to the motive force assembly to release the energy in response to actuation of the trigger assembly; anda retention element that couples the sensor subassembly to the channel guide assembly. 5. The sensor insertion tool of claim 4, wherein the motive force member comprises one or more springs. 6. The sensor insertion tool of claim 4, further comprising a release member to engage a component of the channel guide assembly to retain the sensor subassembly, the actuation of the release member adapted to cause an uncoupling of the sensor subassembly from the channel guide assembly. 7. The sensor insertion tool of claim 4, further comprising an elongated housing. 8. The sensor insertion tool of claim 4, further including the channel guide assembly. 9. The sensor insertion tool of claim 8, wherein the channel guide assembly is permanently coup... | WAVEFORM TECHNOLOGIES, INC. | A61B005/05 A61B005/00 A61B005/145 A61B017/34
US08260403 | B2 | 2012-09-04 | Photoacoustic imaging apparatus and photoacoustic imaging method | 2010-06-16 | A photoacoustic wave detector detects a photoacoustic wave generated inside a specimen by light irradiated thereto. A signal processing device: forms first volume data from a first signal, the first signal being the detection signal acquired from the detector or a signal obtained by adjusting an amplitude of the detection signal; forms second volume data from a second signal, the second signal being a signal obtained by changing a phase of the first signal; forms third volume data from the first and second volume data; and generates and outputs image data representing information on an interior of the specimen from the third volume data. | 1. An in vivo photoacoustic imaging apparatus comprising: a detector configured to detect a photoacoustic wave generated inside a specimen by irradiating the specimen with light; anda signal processing device that images and outputs information on an interior of the specimen based on a detection signal acquired from said detector,wherein said signal processing device forms first volume data from a first signal, the first signal being the detection signal itself or a signal obtained by adjusting an amplitude of the detection signal,generates a second signal by changing a phase of the first signal,forms second volume data from the second signal,forms third volume data from the first volume data and the second volume data, andgenerates output image data from the third volume data. 2. The photoacoustic imaging apparatus according to claim 1, wherein said signal processing device generates the second signal by delaying a phase of a positive frequency component contained in the first signal by 90 degrees and advancing a phase of a negative frequency component contained in the first signal by 90 degrees. 3. The photoacoustic imaging apparatus according to claim 1, wherein said signal processing device generates the first signal by multiplying the detection signal by a coefficient proportional to a period of time from the irradiation of the specimen with light until the detection of the photoacoustic wave. 4. The photoacoustic imaging apparatus according to claim 1, wherein the third volume data has a value proportional to a square root of a sum of squares of the first volume data and the second volume data. 5. The photoacoustic imaging apparatus according to claim 1, wherein the third volume data has a value proportional to a sum of absolute values of the first volume data and the second volume data. 6. An in vivo photoacoustic imaging method used in a photoacoustic imaging apparatus, the method comprising the steps of: detecting, by means of a detector, a photoacoustic wave which is generated inside a specimen by light irradiated thereto;forming first volume data from a first signal, the first signal being a detection signal acquired from the detector or a signal obtained by adjusting an amplitude of the detection signal;generating a second signal by changing a phase of the first signal,forming second volume data from the second signal;forming third volume data from the first volume data and the second volume data; andgenerating and outputting, from the third volume data, image data which represents information on an interior of the specimen. 7. A non-transitory computer readable storage medium storing, in executable form, a program for causing a computer to perform an in vivo method comprising the steps of: acquiring a detection signal of a photoacoustic wave which is generated inside a specimen by light irradiated thereto;forming first volume data from a first signal, the first signal being the detection signal acquired or a signal obtained by ... | Canon Kabushiki Kaisha | A61B006/00
US08611987 | B2 | 2013-12-17 | Method and device for assisting in determination of the suitability of a patient for a scan of the patient's heart using an X-ray computer tomograph and method and X-ray computer tomograph for scanning the heart of a patient | 2010-06-18 | A method is disclosed for assisting in determination of the suitability of a patient for a scan of the heart of the patient using an X-ray computer tomograph. In at least one embodiment of the method, a) an electrocardiogram of the patient is recorded; b) the electrocardiogram is evaluated by predicting the occurrence time of at least the immediately following R wave on the basis of at least four immediately consecutive R waves of the electrocardiogram which were measured last, and comparing this with the actual measured occurrence time of the next R wave; and c) wherein the quality of the prediction is visualized qualitatively. Further, in at least one embodiment, a device is disclosed including an ECG instrument and a computation device for carrying out the method. At least one embodiment of the invention furthermore relates to a method and a device for scanning the heart of a patient on the basis of a prediction of R waves. | 1. A method for assisting in determination of the suitability of a patient for a scan of a heart of the patient in a resting phase of only one cardiac cycle of the patient's heart using an X-ray computer tomograph, the method comprising: recording an electrocardiogram of the patient;evaluating, by a processor, the electrocardiogram by predicting an occurrence time of at least an immediately following R wave based upon at least four immediately consecutive R waves of the electrocardiogram which were previously measured;comparing, by the processor, the occurrence time with an actual measured occurrence time of the next R wave to obtain an accuracy of the prediction;determining, by the processor, a quality of the prediction by comparing the accuracy with a threshold; anddisplaying, by the processor, the quality of the prediction. 2. The method as claimed in claim 1, wherein the recording of the electrocardiogram is performed while the patient stops breathing, only a part of the electrocardiogram which is recorded while the patient has stopped breathing for at least one of two to five cardiac cycles and three to four seconds is used in the evaluating. 3. The method as claimed in claim 2, wherein in the evaluating, the occurrence times of the immediately following fifth and sixth waves are predicted on a basis of a first, a second, a third and a fourth immediately consecutive measured R waves of the electrocardiogram, and compared with the actual measured occurrence times of the immediately following fifth and sixth waves. 4. The method as claimed in claim 1, wherein in the evaluating, the occurrence times of immediately following fifth and sixth R waves are predicted on a basis of a first, a second, a third and a fourth immediately consecutive measured R waves of the electrocardiogram, and compared with the actual measured occurrence times of the immediately following fifth and sixth waves. 5. The method as claimed in claim 1, wherein the evaluating is carried out repeatedly in order to determine the quality of the prediction. 6. The method as claimed in claim 3, wherein the evaluating is carried out in an overlapping fashion in relation to the R waves, an advance by one R wave taking place each time the evaluating is carried out. 7. The method as claimed in claim 1, wherein a prediction of the occurrence time of an R wave is classed as adequate if a difference between the predicted time and the actual measured occurrence time of the R wave is less than 5%. 8. The method as claimed in claim 7, wherein the quality of the prediction is visualized green overall if the relevant predictions have predominantly been classed as adequate, and wherein the quality of the predictions is visualized red overall if the relevant predictions have predominantly been classed as too inaccurate. 9. A method for recording 2D X-ray projections of a heart of a patient from different projection directions in a resting phase of only one cardiac cycle of the patient us... | Siemens Aktiengesellschaft | A61B005/05
US08634629 | B2 | 2014-01-21 | Estimating risk of future bone fracture utilizing three-dimensional bone density model | 2006-11-10 | A novel approach for analyzing a patient's body part of interest to assess bone strength and/or risk of future fracture includes obtaining a priori information regarding the body part of interest, performing X-ray absorptiometric scans of the patient's body part of interest and collecting X-ray absorptiometry data from the scans, constructing a three-dimensional model of the patient's body part of interest, by utilizing the a priori information along with the X-ray absorptiometric data, and performing measurements of various geometric parameters on the three-dimensional model for determining geometric and structural properties. | 1. A computer-implemented method for analyzing a patient's body part of interest to assess bone strength and/or risk of future fracture, said method comprising: obtaining a priori information regarding the body part of interest, said a priori information comprising a three-dimensional (3D) generic model of a comparable body part;performing X-ray absorptiometric scans of the patient's body part of interest and collecting X-ray absorptiometry data from the X-ray absorptiometric scans, said data being derived using x-ray imaging detectors included in a computer-driven data acquisition system, said data representing a plurality of two-dimensional (2D) images of the patient's body part, each of said 2D images of the patient's body part being at a respective orientation of an imaging x-ray beam relative to the patient's body part;constructing a three-dimensional model of the patient's body part of interest, by utilizing the a priori information along with the X-ray absorptiometric data, said constructing comprising operating computerized equipment to carry out the steps of: (i) obtaining two-dimensional (2D) projection views of the 3D generic model that correspond in orientation to the 2D images of the patient's body part,(ii) comparing each of the 3D projection views of the generic model with the corresponding 2D image of the patient's body part to calculate a cost function,(iii) modifying the 3D model according to said cost function, and(iv) repeating steps (i) through (iii), each time using in step (i) the 3D model as modified in the previous iteration of step (iii), until a selected threshold has been reached;performing measurements of various geometric parameters on the three-dimensional model as modified through said constructing, and determining at least one of geometric and structural properties of the patient's body part; andestimating force of a hypothetical fall by the patient;wherein said structural properties include bone strength, and said bone strength is considered in combination with one or more other parameters to estimate said force of the fall. 2. The method of claim 1, wherein said a priori information is obtained from an anatomical atlas or statistical atlas of the body part of interest. 3. The method of claim 1, further comprising modifying the 3D model according to said cost function in a manner that reduces a value of the cost function. 4. The method of claim 1, wherein said a priori information is obtained from an anatomical atlas or a statistical atlas of the body part of interest, said anatomical or statistical atlas utilizes eigenvectors to represent anatomical variations, and the cost function is minimized by applying a post-and-scale technique. 5. The method of claim 1, wherein said at least one of the geometric and structural properties include at least one of cross sectional moment of inertia, section modulus and buckling ratio. 6. The method of claim 1, wherein said other parameters include one or more of age, hei... | Hologic, Inc. | G06K009/00
US08536470 | B2 | 2013-09-17 | Body composition monitor that corrects a body composition evaluation by comparing it with a user's change in weight reading to cancel errors caused by dryness of hands and way of gripping | 2009-03-16 | A body composition monitor of the present invention includes means for measuring a body composition, means for storing a measured body composition value, and body composition comparison means for comparing the measured body composition value and a stored body composition value, in which the body composition monitor further includes means for inputting a body weight value, means for storing an inputted body weight, body weight comparison means for comparing the inputted body weight and stored body weight, body composition change determination means for determining a change degree of the body composition value with using a comparison result of the body composition comparison means and a comparison result of the body weight comparison means, and display means for displaying a determination result determined by the body composition change determination means. | 1. A body composition monitor, comprising: means for measuring a body composition;means for storing a measured body composition value; andbody composition comparison means for comparing the measured body composition value and a stored body composition value, wherein the body composition comparison means determines a body composition evaluation level by comparing a difference between the measured body composition value and the stored composition value with a first predetermined threshold value,the body composition monitor further comprising:means for inputting a body weight value;means for storing an inputted body weight; andbody weight comparison means for comparing the inputted body weight and the stored body weight, wherein the body weight comparison means determines a body weight evaluation level by comparing a difference between the inputted body weight and the stored body weight with a second predetermined threshold value,the body composition monitor further comprising:body composition change determination means for determining a corrected body composition evaluation level by comparing the body composition evaluation level with the body weight evaluation level; anddisplay means for displaying the corrected body composition evaluation level. 2. The body composition monitor according to claim 1, wherein the means for inputting a body weight value comprises means for measuring the body weight. 3. The body composition monitor according to claim 1, further comprising: means for inputting at least one of age, gender and height as personal identification information of a measured person, whereinthe body composition change determination means determines the corrected body composition evaluation level by using the personal identification information. 4. The body composition monitor according to claim 3, wherein the body composition change determination means is provided with a plurality of different determination criteria, and selects one of the determination criteria to be used for the measured person based on the personal identification information. 5. The body composition monitor according to claim 1, wherein the body composition comparison means compares a plurality of different body composition values, and the body composition change determination means determines a corrected body composition evaluation level for each of the plurality of different body composition values. 6. The body composition monitor according to claim 5, wherein the display means displays side by side the corrected body composition evaluation level for each of the plurality of different body composition values determined by the body composition change determination means. 7. The body composition monitor according to claim 5, wherein the display means displays the corrected body composition evaluation level for each of the plurality of different body composition values determined by the body composition change determination means so that the corrected body composition e... | Omron Healthcare Co., Ltd. | A61B005/05 A61B005/053 G01G019/40
US09026223 | B2 | 2015-05-05 | Therapy system including multiple posture sensors | 2009-04-30 | Posture-responsive therapy is delivered by the medical system based on posture state input from only one of multiple posture sensors at any given time. An example implantable medical system includes a first posture sensor and a second sensor. A processor controls therapy delivery to the patient based on at least one of a patient posture state or a patient activity level determined based on input from only one of the first or second posture sensors. In some examples, one of multiple posture sensors of an implantable posture-responsive medical system is used to automatically reorient another posture sensor (of the system), which has become disoriented. The disoriented posture sensor may be automatically reoriented for one or more posture states at a time. | 1. An implantable medical system comprising: a first multiple-axis posture sensor;a second multiple-axis posture sensor; anda processor that controls therapy delivery to a patient based on a patient posture state determined based on input selected from only one of the first or second posture sensors. 2. The system of claim 1, wherein the processor toggles between activating the first posture sensor and deactivating the second posture sensor, and deactivating the first posture sensor and activating the second posture sensor based on an initial patient posture state detected by one or both of the first or the second posture sensors. 3. The system of claim 1, wherein the processor selects the one of the first or second posture sensors to determine the patient posture state based on an initial patient posture state detected by one or both of the first or the second posture sensors. 4. The system of claim 1, wherein the first and the second posture sensors are each associated with one or more of a plurality of posture states for the patient. 5. The system of claim 4, wherein the processor selectively determines the patient posture state based on input from one of the first or second posture sensors based on one or more associations between the one of the first or second posture sensors and the posture state of the patient. 6. The system of claim 4, wherein the first posture sensor is associated with a sitting posture state, and wherein the processor controls therapy delivery to the patient based on the sitting posture state determined based on input selected from only the first sensor. 7. The system of claim 1, wherein the processor selectively determines the patient posture state based on input from the one of the first or the second posture sensors based on a determined status of the other of the first and the second posture sensors. 8. The system of claim 7, wherein the processor selectively determines the patient posture state based on input from the one of the first or the second posture sensors when the other of the first and the second posture sensors fails. 9. The system of claim 7, wherein the processor selectively determines the patient posture state based on input from the one of the first or the second posture sensors when the other of the first or the second posture sensors changes orientation. 10. The system of claim 9, wherein the processor selectively determines the patient posture state based on input from the one of the first or the second posture sensors when the processor determines that the other of the first or the second posture sensors has changed orientation relative to an initial orientation. 11. The system of claim 1, wherein at least one of the first and the second posture sensors comprises one or more of a multiple-axis accelerometer, multiple single-axis accelerometers, a gyroscope, a piezoelectric crystal or a pressure transducer. 12. The system of claim 1 further comprising: an implantable electrical stimulat... | Medtronic, Inc. | A61N001/00 A61M005/142 A61N001/36 A61B005/11 A61B005/00 A61N001/372
US08577443 | B2 | 2013-11-05 | Magnetic resonance imaging apparatus and image processing apparatus | 2010-08-17 | A magnetic resonance imaging apparatus includes a first data acquisition unit, a second data acquisition unit and an image data generating unit. The first data acquisition unit acquires first data from a slice to be a target after a first delay time from a reference of a first heart rate synchronized with an electrocardiogram. The second data acquisition unit acquires second data from the slice after a second delay time from a reference of a second heart rate which is different from the first heart rate. The image data generating unit generates image data with image reconstruction processing using the first data and the second data. | 1. A magnetic resonance imaging apparatus comprising: a preparation scan unit configured to acquire signal values of a blood flow while changing delay times from an R wave in synchronization with an electrocardiogram;a graph generating unit configured to generate a graph which shows a relation between the signal values of the blood flow acquired by the preparation scan unit and the delay times;a display unit configured to display the generated graph;a determining unit configured to determine cardiac phases between systole and diastole in which blood flow signal intensity varies by referring to the displayed graph;a three-dimensional image data acquisition unit configured to acquire three or more echo data sets respectively corresponding to the determined cardiac phases and to generate three or more three-dimensional image data sets respectively corresponding to the determined cardiac phases by image reconstruction processing;a subtraction processing unit configured to generate two or more three-dimensional time-series blood flow image data sets by performing subtraction processing between a reference image data set selected from the generated three-dimensional image data sets and each of the generated three-dimensional image data sets other than the reference image data set or between two of the generated three-dimensional image data sets at near cardiac phases;a projection unit configured to generate two or more projection image time-series data sets by performing projection processing to the three-dimensional blood flow age data sets; andan interface unit configured to set an image processing condition with an operation of an input device through a window displayed on a monitor so as to make at least one of the subtraction processing and the projection processing automatically performed after acquisition of the echo data sets. 2. The magnetic resonance imaging apparatus of claim 1, wherein: the determining unit determines the cardiac phase range and the number of the cardiac phases used for acquiring the three or more echo data sets, andthe three-dimensional image data acquisition unit acquires the echo data sets based on the determined cardiac phase range and the number of the cardiac phases. 3. The magnetic resonance imaging apparatus of claim 2, wherein the cardiac phase range for acquisition by the three-dimensional image data acquisition unit is determined to be narrower than that used in the preparation scan. 4. A magnetic resonance imaging apparatus comprising: a preparation scan unit configured to acquire signal values of a blood flow while changing delay times from an R wave in synchronization with an electrocardiogram;a graph generating unit configured to generate a graph which shows a relation between the signal values of the blood flow acquired by the preparation scan unit and the delay times;a display unit configured to display the generated graph;a determining unit configured to determine cardiac phases between systole and dias... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | A61B005/05
US08400316 | B2 | 2013-03-19 | Electronic thermometer | 2009-02-17 | An electronic thermometer has a probe having a temperature measuring unit at its end, a temperature sensor arranged in the temperature measuring unit of the probe, a contact sensor arranged in a position shifter from the temperature measuring unit of the probe toward a thermometer body, a determining unit for determining whether a state of contact between the temperature measuring unit and a measurement target portion of a user is good or not, based on outputs of both the temperature sensor and the contact sensor, and a notifying unit for providing a notification according to a result of the determination by the determining unit. | 1. An electronic thermometer comprising: a probe having a temperature measuring unit at the an end of said probe;a first sensor arranged in said temperature measuring unit of said probe for measuring a temperature;a second sensor arranged in a position between said temperature measuring unit of said probe and a thermometer body for sensing a contact with a human body;a determining unit for determining whether a state of contact between said temperature measuring unit and a measurement target portion of a user is good or not, based on both outputs of said first and second sensors; anda notifying unit for performing notification according to a result of determination of said determining unit. 2. The electronic thermometer according to claim 1, wherein said first sensor serves also as a sensor for measuring a body temperature of said user. 3. The electronic thermometer according to claim 1, wherein said notifying unit can notify at least that the contact state is good, during measurement of the body temperature. 4. The electronic thermometer according to claim 1, wherein when the contact state is bad, said determining unit determines a level of a failure in the contact state based on the outputs of said first and second sensors, and said notifying unit performs the notification according to said level of the failure. 5. The electronic thermometer according to claim 1, wherein when the contact state is bad, said determining unit estimates a cause of a failure in the contact state based on the outputs of said first and second sensors, and said notifying unit performs the notification according to said cause of the failure. 6. The electronic thermometer according to claim 5, wherein said notifying unit provides guidance on solution to said cause of the failure. 7. The electronic thermometer according to claim 1, wherein said notifying unit includes an LED, and performs the notification utilizing a state of turn-on/off of said LED. 8. The electronic thermometer according to claim 7, wherein said notifying unit performs the notification by turning on said LED in different colors depending on the good contact state and the bad contact state, respectively. 9. The electronic thermometer according to claim 8, wherein said notifying unit includes a buzzer for performing the notification by turning on said buzzer when the contact state is bad. 10. The electronic thermometer according to claim 1, wherein the second sensor includes an electrostatic capacity sensor. 11. The electronic thermometer according to claim 1, wherein the determining unit determines whether a state of contact between said temperature measuring unit and a measurement target portion of a user is good or not, based on the output with regard to the temperature of said first sensor and the output with regard to the contact state of said second sensor. | Omron Healthcare Co., Ltd. | G08B017/00
US08290572 | B2 | 2012-10-16 | Method and system for navigating a catheter probe in the presence of field-influencing objects | 2010-09-13 | A system for and method of determining and compensating for the effect of a field influencing object on a field sensor, preferably a coil, that is within a navigational domain. The navigational domain contains navigational magnetic energy and disturbing magnetic energy, and the field influencing object produces the disturbing magnetic energy in response to the navigational magnetic energy. The correction system includes a first transmitter for projecting into the navigational domain field energy in a first waveform sufficient to induce a first signal value in the sensing coil. The system also includes a second transmitter for projecting into the navigational domain field energy in a second waveform sufficient to induce a second signal value in the sensing coil. The system further includes a signal processor for receiving the first signal value and for receiving the second signal value to determine the effect of the electrically conductive object on the field sensor. In other embodiments, the correction system may utilize a second sensing coil to receive and process an alternate aspect of the field energy from a position that is different from the first field sensor, or the system may utilize a storage device that contains information relating to the effect of the field influencing device within the navigational domain. | 1. A navigation system to compensate for distortion of an electromagnetic (EM) field caused by a tool in the EM field, comprising: a plurality of coils, wherein each of the plurality of coils is configured to generate an EM field;a tool positionable relative to a patient to assist in performing a procedure on the patient, wherein the tool includes a ferromagnetic material that distorts the EM field generated by each of the plurality of coils to form a distorted field;a tool coil positioned adjacent to at least a portion of the tool, wherein the tool coil is configured to sense the EM field generated by each of the plurality of coils and the distorted field and further configured to generate a tool coil signal based on the sensed EM field generated by each of the plurality of coils and the distorted field;a sensor coil configured to sense the EM field generated by each of the plurality of coils and the distorted field and further configured to generate a sensor coil signal based on the sensed EM field generated by each of the plurality of coils and the distorted field; anda signal processor configured to receive the tool coil signal and the sensor coil signaland further configured to determine a proper orientation and position of the sensor coil by adjusting the sensor coil signal with the tool coil signal to compensate for the distorted field. 2. The navigation system of claim 1, wherein the plurality of coils includes: three unidirectional coil sets, each of the unidirectional coil sets being driven by a unit capable of driving the unidirectional coil sets at a first waveform at a first frequency, andsix delta coil sets, each of the delta coil sets being driven by a drive unit capable of driving the delta coil sets at the first waveform at the first frequency, such that the three unidirectional coil sets and the six delta coil sets generate the EM field at said first frequency. 3. The navigation system of claim 2, wherein the tool includes at least one of a drill and a microscope. 4. The navigation system of claim 2, wherein the tool forms the distorted field by non-linearly distorting the EM field generated by each of the plurality of coils. 5. The system of claim 4, wherein the EM field generated by each of the plurality of coils or the distorted field is generated in a region and the sensor coil is navigated in a navigation domain defined by at least a portion of the EM field generated by each of the plurality of coils or the distorted field; wherein the tool is positioned in a peripheral domain outside of the navigation domain. 6. The navigation system of claim 4, wherein the EM field generated by each of the plurality of coils or the distorted field is generated in a region and the sensor coil is navigated in a navigation domain defined by at least a portion of the EM field generated by each of the plurality of coils or the distorted field; wherein the tool is positioned in the navigation domain in which the instrument coil is operable... | Medtronic Navigation, Inc. | A61B005/05
US08151801 | B2 | 2012-04-10 | Septum port locator system and method for an implantable therapeutic substance delivery device | 2010-08-31 | An implantable medical device system includes an implantable device and an external locator device for percutaneously locating detecting port opening of the implantable device. The implantable device includes a port chamber forming the port opening, a septum sealing the port chamber relative to an exterior of the device, and a coil positioned at a known location relative to the port opening. The locator device includes a controller, at least one X-loop electrically coupled to the controller and oriented along a first major axis, and at least one Y-loop electrically coupled to the controller and oriented along a second major axis differing from the first major axis. The system is configured such that when the locator is spatially proximate the coil, an induced voltage in at least the Y-loop(s) is read by the controller to indicate a location of the coil relative to the locator device. | 1. A locator device for externally detecting a coil carried by a substance delivery device implanted within a patient, the locator device comprising: a controller; anda locator grid including: a plurality of interleaved X-loops electrically coupled to the controller, each of the X-loops defining a major axis that is oriented parallel to a first major axis,at least one Y-loop electrically coupled to the controller and oriented along a second major axis differing from the first major axis;wherein the locator device is configured such that when the locator device is spatially proximate the coil, energization of the coil induces a voltage in the at least one Y-loop that is read by the controller, and further wherein the controller is adapted to determine a location of the coil relative to the locator grid based upon the read voltage. 2. The locator device of claim 1, wherein the plurality of X-loops includes a central X-loop oriented along a first central axis and the at least one Y-loop comprises a central Y-loop oriented along a second central axis, and further wherein the locator device defines a needle opening extending through a housing of the locator device at a location where the first central axis intersects the second central axis. 3. The locator device of claim 1, wherein the second central axis is orthogonal to the first central axis. 4. The locator device claim 1, wherein each of the X-loop is staggered relative to, and overlaps, a neighboring X-loop. 5. The locator device of claim 1, wherein the locator grid includes a plurality of Y-loops, each Y-loop including a major axis that is parallel to the second major axis such that each Y-loop is staggered relative to a neighboring Y-loop. 6. The locator device of claim 5, wherein the plurality of Y-loops are interleaved. 7. The locator device of claim 6, wherein the plurality of X-loops are electrically decoupled from the plurality of Y-loops. 8. The locator device of claim 1, wherein at least one of the plurality of X-loops comprises a rectangular loop defining a first side opposite a second side and opposing ends, the first and second sides being parallel to the first major axis and the opposing ends being parallel to a minor axis of the at least X-loop, the minor axis being orthogonal to the first major axis. 9. The locator device of claim 1, wherein the locator device further includes a power source. 10. The locator device of claim 9, wherein the locator device is adapted such that when the locator grid is positioned proximate the coil and a current is applied to one of the plurality of X-loops, the coil is energized and a voltage is induced in the at least one Y-loop. 11. The locator device of claim 10, wherein at least one Y-loop includes a plurality of Y-loops, and further wherein the controller is adapted to iteratively sample and compare signals generated at each of the Y-loops in at least partially determining a location of the coil relative to the locator grid. 12. The l... | Medtronic, Inc. | A61B019/00
US08326397 | B2 | 2012-12-04 | Mountable unit for biological signal measurement and biological signal measuring method | 2010-06-16 | Disclosed herein is a mountable unit for biological signal measurement, including: three electrodes; and a support body configured to support the three electrodes at positions corresponding to a positional relationship among a predetermined position on a forehead on one side of a face bounded by a midline, a temple position on the one side of the face, and a predetermined position on a zygomatic body on the one side of the face. | 1. A mountable unit for biological signal measurement, comprising: three electrodes; anda support body configured to support the three electrodes at positions corresponding to a positional relationship among a predetermined position on a forehead on one side of a face bounded by a midline, a temple position on the one side of the face, and a predetermined position on a zygomatic body on the one side of the face, whereinthe support body is provided with a fixture for fixing the support body by sandwiching head hair. 2. The mountable unit for biological signal measurement according to claim 1, wherein the predetermined position on the forehead is a position above the temple position, and the predetermined position on the zygomatic body is a position beside a tail of an eye. 3. The mountable unit for biological signal measurement according to claim 1, wherein the support body has a plate shape, andat least part of each of the three electrodes is exposed in one surface of the support body having the plate shape, and suction cups or adhesive materials surrounding the exposed electrode parts as concave parts are provided. 4. The mountable unit for biological signal measurement according to claim 1, wherein the support body is provided with a circuit that measures potential difference between the electrode supported at the position corresponding to the predetermined position on the forehead and the electrode supported at the position corresponding to the temple position, and potential difference between the electrode supported at the position corresponding to the predetermined position on the zygomatic body and the electrode supported at the position corresponding to the temple position. | Sony Corporation | A61B005/0478 A61B005/0492
US08394020 | B2 | 2013-03-12 | Skin laxity measurement system | 2010-05-07 | A skin laxity measurement system which provides a skin laxity measurement tool and methods of measuring skin laxity. | 1. A skin laxity measurement tool, comprising: a) a first linear measurement guide;b) a first arm fixedly coupled to said first linear measurement guide;c) a second arm coupled to said first linear measurement guide, said first arm and said second arm disposed to provide opposed inner surfaces in moveable relation, said first arm and said second arm corresponding provide a first arm terminal and a second arm terminal disposed a distance from said first linear measurement guide; andd) a cross member slidely coupled to both said first arm and said second arm in substantially perpendicular relation to said opposed inner surfaces to allow positioning of said cross member at a location between said first arm end and said first arm terminal and said second arm end and said second arm terminal. 2. The skin laxity measurement tool described in claim 1, wherein said first arm slidely couples with said first linear measurement guide and said second arm fixedly couples to said first linear measurement guide to provide said opposed inner surfaces disposed in moveable relation. 3. The skin laxity measurement tool described in claim 1, wherein both of said first arm and said second arm slidely couple to said first linear measurement guide to provide said opposed inner surfaces disposed in moveable relation. 4. The skin laxity measurement tool described in claim 1, further comprising a first graduated scale coupled to said first linear measurement guide. 5. The skin laxity measurement tool described in claim 4, further comprising a second graduated scale coupled to said first arm or said second arm. 6. The skin laxity measurement tool described in claim 5, wherein said first graduated scale coupled to said first linear measurement guide comprises a plurality of measurement indicia in alignable relation to movement of said first arm to provide a distance value which corresponds to a distance between said opposed inners surfaces of said first arm and said second arm. 7. The skin laxity measurement tool described in claim 6, wherein said second graduated scale coupled to said first arm or said second arm comprises a plurality of measurement indicia in alignable relation to movement of said cross member to provide a height value corresponding to a height of said cross member from said first arm terminal and said second arm terminal. 8. The skin laxity measurement tool described in claim 5, wherein said first graduated scale coupled to said first liner measurement guide comprises a plurality of measurement indicia translatable by operation of a distance calculator to display said distance value. 9. The skin laxity measurement tool described in claim 8, wherein said second graduated scale coupled to said first or second arm comprises a plurality of measurement indicia translatable by operation of a height calculator to display said height value. 10. The skin laxity measurement tool described in claim 9, wherein said distance calculator comprises a digital di... | American Network of Lipolysis, LLC | A61B005/00
US08452360 | B2 | 2013-05-28 | System and method for non-invasive photothermal radiometric measurement | 2010-05-18 | There is provided a glucose monitoring method and apparatus based on the principle of Wavelength-Modulated Differential Laser Photothermal Radiometry (WM-DPTR). Two intensity modulated laser beams operating in tandem at specific mid-infrared (IR) wavelengths and current-modulated synchronously by two electrical waveforms 180 degrees out-of-phase, are used to interrogate the tissue surface. The laser wavelengths are selected to absorb in the mid infrared range (8.5-10.5 μm) where the glucose spectrum exhibits a discrete absorption band. The differential thermal-wave signal generated by the tissue sample through modulated absorption between two specific wavelengths within the band (for example, the peak at 9.6 and the nearest baseline at 10.5 μm) lead to minute changes in sample temperature and to non-equilibrium blackbody radiation emission. This modulated emission is measured with a broadband infrared detector. The detector is coupled to a lock-in amplifier for signal demodulation. Any glucose concentration increases will be registered as differential photothermal signals above the fully suppressed signal baseline due to increased absorption at the probed peak or near-peak of the band at 9.6 μm at the selected wavelength modulation frequency. The emphasis is on the ability to monitor blood glucose levels in diabetic patients in a non-invasive, non-contacting manner with differential signal generation methods for real-time baseline corrections, a crucial feature toward precise and universal calibration (independent of person-to-person contact, skin, temperature or IR-emission variations) in order to offer accurate absolute glucose concentration readings. | 1. A method of detecting an analyte within a material, said method comprising the steps of: a) directing and overlapping first and second intensity modulated optical beams onto a surface of said material;wherein an absorption spectrum of said material comprises a spectral region substantially free of spectral features in the absence of said analyte, and wherein the presence of said analyte produces a spectral peak within said spectral region;wherein said first and second optical beams are provided with substantially equal intensity, and wherein said first and second optical beams are modulated at a substantially equal frequency with a phase difference of approximately 180 degrees; andwherein wavelengths of said first and second optical beams are selected to lie within said spectral region so that the presence of said analyte produces increased absorption of said first optical beam relative to absorption of said second optical beam;b) obtaining a signal by detecting photothermal emission radiated in response to differential absorption of said first and second optical beams, wherein said photothermal emission is detected with a phase-sensitive detection apparatus provided with a reference signal related to a phase of said modulated optical beams; andc) correlating said signal with an amount of said analyte in said material. 2. The method according to claim 1 wherein said spectral region substantially free of spectral features comprises a substantially flat spectral region. 3. The method according to claim 1 wherein said spectral region is substantially free of spectral features from other species that may be present in said material. 4. The method according to claim 1 wherein said analyte is glucose. 5. The method according to claim 1 wherein said material comprises a biological tissue. 6. The method according to claim 5 wherein said tissue comprises skin. 7. The method according to claim 5 wherein said spectral region lies between approximately 8.5 μm and 10.5 μm. 8. The method according to claim 7 wherein said first optical beam has a wavelength of approximately 9.6 μm and said second optical beam has a wavelength of approximately 10.5 μm. 9. The method according to claim 1 wherein said step of detecting photothermal emission comprises the steps of: collecting photothermal emission from said surface;spectrally filtering said collected photothermal emission to remove scattered optical radiation from said first and second beams; anddetecting said photothermal emission with said phase-sensitive detection apparatus. 10. The method according to claim 9 wherein said phase-sensitive detection apparatus comprises a thermal detector and a lock-in amplifier. 11. The method according to claim 9 wherein said step of collecting said photothermal emission is performed using one of solid angle and reflectivity optimized curved mirrors and a mid-infrared collecting fiber optic system. 12. The method according to claim 1 wherein said frequency is selec... | - | A61B005/1455
US08353852 | B2 | 2013-01-15 | Ophthalmic sizing devices | 2009-01-21 | An ophthalmic sizing device includes a sizing tube or set of such sizing tubes for use in determining a correct size of a bypass tube required in lacrimal surgery, each sizing tube having a flanged end at one end of a longitudinal length of tubing, the remaining end of the tubing distal to the flanged end having a series of ruled markings longitudinally along and towards the flange end. When used, a set of sizing tubes includes tubes having at least two relatively different absolute longitudinal lengths of tubing; and/or diameter flanges. The device can be in the form of a dilator having a sizing section. The use of such sizing devices allows the length of an ophthalmic bypass tube to be determined that leaves the tip of the bypass tube, when in situ, to protrude an optimally desired distance beyond the outer (lateral) nasal wall and into the nasal cavity of each particular patient, while remaining clear of the midline septum. Such a proud projection of bypass tubing helps prevent the bypass tube from being blocked after its permanent insertion. | 1. An ophthalmic sizing be for use in determining a correct size of a bypass tube required in tear-duct surgery, said ophthalmic sizing tube consisting of a cylindrical tubular element having a longitudinal length, the tabular element defining a lumen, and the tubular element defining a flanged end at one end of the tubular element, the remaining length of the tubular element distal to the flanged end having a series of ruled markings, the series extending longitudinally along the remaining length and towards the flanged end, and the markings being visible by virtue of at least one of: a contrast between color of the markings and color of the remainder of the element over which the markings are located, andthe markings being raised or depressed with respect to intervening sections of the tubular element;such that in use the ophthalmic sizing tube may be inserted into a drainage passage made during tear-duct surgery, the drainage passage extending from an eye to a nasal passage, with the flanged end locating at the inner canthus of the eye, the markings within the nasal passage being visible. 2. A set of ophthalmic sizing tubes for use in determining a correct size of a bypass tube required in tear-duct surgery comprising: a plurality of sizing tubes as claimed in claim 1, wherein the plurality of sizing tubes comprises sizing tubes of at least two relatively different absolute longitudinal lengths. 3. A set of sizing tubes as claimed in claim 2, wherein the plurality of sizing tubes comprises sizing tubes having at least two relatively different diameters of flanges at their flanged ends. 4. A set of ophthalmic sizing tubes for use in determining a correct size of a bypass tube required in tear-duct surgery comprising: a plurality of sizing tubes as claimed in claim 1, wherein the plurality of sizing tubes comprises sizing tubes having at least two relatively different diameters of flanges at their flanged ends. 5. A sizing tube as claimed in claim 1, manufactured of a material that may be cut and trimmed to a desired length. 6. A sizing tube as claimed in claim 1, manufactured of a plastic or silicone material. 7. An ophthalmic dilator for use in determining a correct size of a bypass tube required in tear-duct surgery, said ophthalmic dilator defining a lumen and defining a cylindrical tubular element having a longitudinal length, and a shoulder to a portion having a larger diameter at one end of the cylindrical tubular element, the end of the tubular element distal to the shouldered end having a series of ruled markings, the series extending longitudinally along the tubular element and towards the shouldered end, and the markings being visible by virtue of at least one of: a contrast between color of the markings and color of the remainder of the element over which the markings are located, andthe markings being raised or depressed with respect to intervening sections of the tubular element;such that in use the ophthalmic dilator may be ... | LJT Projects Ltd. | A61B005/00
US08475370 | B2 | 2013-07-02 | Method for measuring patient motion, activity level, and posture along with PTT-based blood pressure | 2009-05-20 | The invention provides a system and method for measuring vital signs (e.g. SYS, DIA, SpO2, heart rate, and respiratory rate) and motion (e.g. activity level, posture, degree of motion, and arm height) from a patient. The system features: (i) first and second sensors configured to independently generate time-dependent waveforms indicative of one or more contractile properties of the patient's heart; and (ii) at least three motion-detecting sensors positioned on the forearm, upper arm, and a body location other than the forearm or upper arm of the patient. Each motion-detecting sensor generates at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which it is affixed. A processing component, typically worn on the patient's body and featuring a microprocessor, receives the time-dependent waveforms generated by the different sensors and processes them to determine: (i) a pulse transit time calculated using a time difference between features in two separate time-dependent waveforms, (ii) a blood pressure value calculated from the time difference, and (iii) a motion parameter calculated from at least one motion waveform. | 1. A method for monitoring a patient, comprising the following steps: (a) detecting a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a first sensor comprising a first detector configured to be worn on the patient's body;(b) detecting a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a second sensor comprising a second detector configured to be worn on the patient's body;(c) detecting sets of time-dependent motion waveforms with at least two motion-detecting sensors positioned on different locations on the patient's body, wherein each motion-detecting sensor detects a set of motion waveforms indicative of motion of the location on the patient's body to which the motion-detecting sensor is affixed;(d) processing the first and second time-dependent physiological waveforms to determine at least one vital sign from the patient;(e) analyzing at least a portion of each set of time-dependent motion waveforms, or a mathematical derivative thereof, to determine a motion parameter;(f) processing the motion parameter to determine a probability that the patient is undergoing a specific activity state; and(g) estimating the patient's activity state based on the probability,wherein step (e) further comprises analyzing a time-dependent variation of least one motion waveform, and calculating a mathematical derivative of at least one motion waveform to determine the motion parameter. 2. A method for monitoring a patient, comprising the following steps: (a) detecting a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a first sensor comprising a first detector configured to be worn on the patient's body;(b) detecting a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a second sensor comprising a second detector configured to be worn on the patient's body;(c) detecting sets of time-dependent motion waveforms with at least two motion-detecting sensors positioned on different locations on the patient's body, wherein each motion-detecting sensor detects a set of motion waveforms indicative of motion of the location on the patient's body to which the motion-detecting sensor is affixed;(d) processing the first and second time-dependent physiological waveforms to determine at least one vital sign from the patient;(e) analyzing at least a portion of each set of time-dependent motion waveforms, or a mathematical derivative thereof, to determine a motion parameter;(f) processing the motion parameter to determine a probability that the patient is undergoing a specific activity state; and(g) estimating the patient's activity state based on the probability,wherein step (e) further comprises analyzing a time-dependent variation of least one motion waveform, and comparing at least a portion o... | Sotera Wireless, Inc. | A61B005/00
US08622902 | B2 | 2014-01-07 | Controlling an alarm in a medical instrument | 2009-01-19 | The invention concerns a method for controlling an alarm in a medical instrument or system, the medical instrument or system detecting at least one physiological parameter of the patient. The present value of the physiological parameter is consecutively detected and an alarm delay is determined as a function of at least one detected value of the physiological parameter wherein the function yields a shorter alarm delay for increasing values of the deviation from a normal value and a longer alarm delay for decreasing values of the deviation from the normal value. Further, the duration the value of the physiological parameter exceeds or under-runs at least one predefined threshold for the physiological parameter defining an upper or lower limit for a normal range of the physiological parameter, respectively, is measured and the alarm is generated when the duration the determined value of the physiological parameter exceeds or under-runs the predefined threshold exceeds the alarm delay. This method allows for the avoidance of nuisance alarms while still indicating severe conditions of the monitored patient reliably. | 1. A method for controlling an alarm unit in a medical instrument or system, the medical instrument or system detecting at least one physiological parameter of the patient, the method comprising: with the medical instrument or system, consecutively detecting a present value of the physiological parameter;with a determination unit, after detecting the present value of the physiological parameter, determining an alarm delay as a function of at least one detected value of the physiological parameter wherein the function yields a shorter alarm delay if the degree by which two consecutively determined values of the physiological parameter deviate from a normal value increases, and wherein the function yields a longer alarm delay if the degree by which two consecutively determined values of the physiological parameter deviate from the normal value decreases;measuring a time the physiological parameter has exceeded or under-run at least one predefined threshold for the physiological parameter defining an upper or lower limit for a normal range of the physiological parameter, respectively;starting a limit violation counter running after a first of multiple directly consecutive events of detected present values of the physiological parameter that exceed or under-run the threshold, respectively, and wherein the value of the alarm delay is continuously updated according to the function of at least the detected value of the physiological parameter;determining the value of the limit violation counter when a detected present value of the physiological parameter has returned into a normal range, a reduction value is calculated as a function of the detected present value of the limit violation counter and/or the degree by which the detected present value of the physiological parameter exceeds or under-runs the predefined threshold, and the reduction value is decremented in time; andcontrolling the alarm unit to generate an alarm when the time of the physiological parameter exceeds or under-runs the predefined threshold, respectively, exceeds the alarm delay. 2. The method according to claim 1, wherein the most recently detected value of the physiological parameter is used for determining the alarm delay. 3. The method according to claim 1, wherein the determination unit considers the degree by which the detected value exceeds or under-runs the predefined threshold, respectively, and wherein the function yields a longer alarm delay for a lesser degree of exceeding or under-running the predefined threshold, respectively, and vice versa. 4. The method according to claim 1, wherein a maximum alarm delay and/or a minimum alarm delay are defined. 5. The method according to claim 1, wherein the reduction value is equal to the value of the limit violation counter when the detected present value of the physiological parameter has returned into the normal range. 6. The method according to claim 1, wherein a subsequently alarm delay is reduced by the present reduction... | Koninklijke Philips N.V. | A61B005/00 G06F019/00
US08626276 | B2 | 2014-01-07 | Cardiac activation sequence monitoring for ischemia detection | 2010-07-30 | Cardiac monitoring and/or stimulation methods and systems that provide one or more of monitoring, diagnosing, defibrillation, and pacing. Cardiac signal separation is employed to detect, monitor, track and/or trend ischemia using cardiac activation sequence information. Ischemia detection may involve sensing composite cardiac signals using implantable electrodes, and performing a signal separation that produces one or more cardiac activation signal vectors associated with one or more cardiac activation sequences. A change in the signal vector may be detected using subsequent separations. The change may be an elevation or depression of the ST segment of a cardiac cycle or other change indicative of myocardial ischemia, myocardial infarction, or other pathological change. The change may be used to predict, quantify, and/or qualify an event such as an arrhythmia, a myocardial infarction, or other pathologic change. Information associated with the vectors may be stored and used to track the vectors. | 1. A system, comprising: a housing configured for implantation in a patient;a plurality of implantable electrodes configured for sensing a plurality of composite signals associated with an unknown number of signal sources;a memory provided in the housing and configured to store a baseline angular orientation of a cardiac signal vector representative of the patient's baseline ischemia status; anda processor provided in the housing and coupled to the electrodes and the memory, the processor configured to: open a window associated with an ST-T segment within which the plurality of composite signals are sensed for each of a plurality of cardiac cycles;perform recurring source separation operations by which the plurality of composite signals sensed within the window are separated according to their respective sources to produce windowed cardiac signal vectors indicating a dominant angular orientation of repolarization wavefronts over the plurality of cardiac cycles; anddetect a deviation in the dominant angular orientation of the windowed cardiac signal vectors relative to the baseline angular orientation as an indication of a change in the patient's ischemia status relative to the patient's baseline ischemia status. 2. The system of claim 1, wherein the processor is configured to center the window on the ST-T segment of each of the plurality of cardiac cycles. 3. The system of claim 1, wherein the processor is configured to determine a change in one or more ST-T segment features of the plurality of composite signals sensed within the window indicative of the change in the patient's ischemia status. 4. The system of claim 1, wherein the processor is configured to determine a change in one or more ST-T segment features of the plurality of composite signals sensed within the window indicative of an elevation or a depression of the ST-T segment. 5. The system of claim 1, wherein: the memory is configured to store a baseline magnitude of a cardiac signal vector representative of the patient's baseline ischemia status; andthe processor is configured to detect a deviation in angular orientation and a deviation in magnitude of the windowed cardiac signal vectors relative to the baseline angular orientation and baseline magnitude, respectively, as an indication of the change in the patient's ischemia status relative to the patient's baseline ischemia status. 6. The system of claim 1, wherein the processor is configured to predict myocardial infarction based on the detected deviation in angular orientation. 7. The system of claim 1, wherein the processor is configured to implement a confirmation procedure to confirm that candidate windowed cardiac signal vectors comprise a cardiac signal. 8. The system of claim 1, wherein the processor is configured to: implement a first confirmation procedure that attempts to confirm that the windowed cardiac signal vectors comprise a cardiac signal; andif the first confirmation procedure is unsuccessful, implement a s... | Cardiac Pacemakers, Inc. | A61B005/04
US09125600 | B2 | 2015-09-08 | Medical device with incorporated disinfecting wipe and method of using same | 2010-06-21 | A medical device including a barrel having an open first end and capable of supporting a needle at a second end of the barrel is provided and includes a lid releasably adhered to and overlying the open first end of the barrel in a fluid tight manner; and a pad secured to a surface of the lid such that the pad is disposed within a cavity of the barrel when the lid is adhered to the barrel, wherein the pad comprises a disinfecting solution. | 1. A medical needle device comprising: a barrel having an open first end and a second end having an opening configured to support a needle, the barrel defining a cavity in fluid communication with the opening, the cavity being configured to receive a blood collection tube;a lid releasably connected to and overlying the open first end of the barrel, an axial surface of the lid facing the cavity of the barrel; anda pad secured to the axial surface of the lid and disposed within the cavity of the barrel when the lid is connected to the barrel, the lid being positioned such that the pad extends into and is exposed to the cavity of the barrel when the lid is connected to the barrel, wherein the pad comprises a disinfecting solution. 2. The medical needle device according to claim 1, wherein the disinfecting solution is one of alcohol, isopropyl alcohol and benzalkonium chloride. 3. The medical needle device according to claim 1, wherein the lid is adhered to the barrel in at least one of a fluid tight and vapor tight manner. 4. The medical needle device according to claim 2, wherein the lid includes a tab extending therefrom. 5. The medical needle device according to claim 2, wherein the lid is fabricated from at least one of a fluid impermeable material and a vapor impermeable material. 6. The medical needle device according to claim 1, wherein the medical needle device is a blood collection device. 7. The medical needle device according to claim 1, wherein the medical needle device is a syringe. 8. A medical needle device including a barrel having an open first end and capable of supporting a needle at a second end of the barrel, the medical needle device comprising: a lid releasably connected to the barrel, the lid including: a first layer releasably adhered to and overlying the open first end of the barrel; anda second layer releasably adhered to the first layer, wherein the first layer and the second layer define at least one of a fluid and vapor tight pocket therebetween, wherein the second layer of the lid is adhered to the first layer of the lid such that upon application of force, the second layer is separated from the first layer before the lid is separated from the barrel; anda pad disposed within the pocket of the lid, wherein the pad comprises a disinfecting solution. 9. The medical needle device according to claim 8, wherein the disinfecting solution is one of alcohol, isopropyl alcohol and benzalkonium chloride. 10. The medical needle device according to claim 8, wherein the lid includes a tab extending from and integrally connected with the second layer. 11. The medical needle device according to claim 8, wherein the second layer is integrally connected to the first layer along a fold line. 12. The medical needle device according to claim 8, wherein the lid includes an additional pad secured to the first layer at a location within the barrel. 13. The medical needle device according to claim 8, wherein the pad is secured to ... |  | A61B017/32 A61B017/14 A61B005/15 A61B005/153 A61M005/00 A61M005/31
US08512255 | B2 | 2013-08-20 | Device and method for the geometric determination of electrical dipole densities on the cardiac wall | 2009-01-16 | Disclosed are devices, a systems, and methods for determining the dipole densities on heart walls. In particular, a triangularization of the heart wall is performed in which the dipole density of each of multiple regions correlate to the potential measured at various locations within the associated chamber of the heart. To create a database of dipole densities, mapping information recorded by multiple electrodes located on one or more catheters and anatomical information is used. In addition skin electrodes may be implemented. | 1. A device for creating a database of dipole densities d(y) at the surface of one or more cardiac chambers of a patient, said device comprising: a first receiver configured to receive mapping information from multiple electrodes located on one or more catheters, the mapping information received when the multiple electrodes are placed in the one or more cardiac chambers;a second receiver configured to receive anatomical information from at least one imaging instrument configured to produce a geometrical depiction of the one or more cardiac chambers; anda dipole density module configured to generate the database of dipole densities d(y), wherein the dipole density module determines a dipole density for individual triangle shaped projections onto the cardiac chamber wall, where each triangle projection at a location y contributes {acute over (ω)}(x,y) times the dipole density d(y) to a potential V(x) at a point x, wherein {acute over (ω)}(x,y) is the solid angle for that triangle projection, and where: a) x represents a series of locations within one or more cardiac chambers; andb) V(x) is a measured potential at point x, said measured potential recorded by the multiple electrodes. 2. The device of claim 1 wherein said triangle projections are sized such that the dipole density for each triangle projection is substantially constant. 3. The device of claim 1 wherein the dipole density is determined for at least 1000 triangle shaped projections. 4. The device of claim 1 wherein the dipole density is determined by a number of triangle shaped projections, said number determined by the size of a cardiac chamber. 5. The device of claim 1 wherein the multiple electrodes are included in a single catheter. 6. The device of claim 1 wherein the multiple electrodes are included in two or more catheters. 7. The device of claim 1 wherein the imaging instrument is selected from a group consisting of: a computed tomography (CT) instrument; a magnetic resonance imaging (MRI) instrument; an ultrasound instrument; a multiple electrode mapping catheter and mapping system; and combinations thereof. 8. The device of claim 1 wherein the imaging instrument comprises a standard anatomical geometry which is uploaded to the dipole density module. 9. The device of claim 1 wherein the dipole density module includes a mathematical processing element that comprises one or more of: a computer; an electronic module; a computer program stored in a memory and executable by a processor; a microcontroller; a microprocessor; and combinations thereof. 10. The device of claim 1 wherein the dipole density module implements a progressive algorithm configured to improve at least one of a spatial resolution and a time resolution of the database of dipole densities d(y). 11. The device of claim 1 wherein the dipole density module uses a linear system of equations to determine the database of dipole densities d(y). 12. The device of claim 1 wherein the dipole density module determin... | Christoph Scharf | A61B005/04
US08512256 | B2 | 2013-08-20 | Method of locating the tip of a central venous catheter | 2010-09-09 | A method of locating a tip of a central venous catheter (“CVC”) having a distal and proximal pair of electrodes disposed within the superior vena cava, right atrium, and/or right ventricle. The method includes obtaining a distal and proximal electrical signal from the distal and proximal pair of electrodes and using those signals to generate a distal and proximal P wave, respectively. A deflection value is determined for each of the P waves. A ratio of the deflection values is then used to determine a location of the tip of the CVC. Optionally, the CVC may include a reference pair of electrodes disposed within the superior vena cava from which a reference deflection value may be obtained. A ratio of one of the other deflection values to the reference deflection value may be used to determine the location of the tip of the CVC. | 1. A method of locating a tip of a central venous catheter having a portion disposed within the superior vena cava, the catheter comprising a distal pair of electrodes and a proximal pair of electrodes spaced longitudinally from the distal pair of electrodes, the method comprising: obtaining a distal electrical signal from the distal pair of electrodes;generating a first P wave from the distal signal;obtaining a proximal electrical signal from the proximal pair of electrodes;generating a second P wave from the proximal signal;comparing the first and second P waves;obtaining a reference electrical signal in the superior vena cava;generating a reference P wave from the reference electrical signal;comparing the reference P wave with at least one of the first and second P waves; anddetermining a location of the tip of the central venous catheter based on the comparison of the first and second P waves and the comparison of the reference P wave with at least one of the first and second P waves. 2. The method according to claim 1, wherein the first P wave has a first shape, the second P wave has a second shape, and the reference P wave has a reference shape, comparing the first and second P waves comprising comparing the first and second shapes, and comparing the reference P wave with at least one of the first and second P waves comprising comparing the reference shape with at least one of the first and second shapes. 3. The method according to claim 2, wherein the first shape has a first height and the second shape has a second height, determining the location of the tip of the central venous catheter based on the comparison of the first and second P waves and the comparison of the reference P wave with at least one of the first and second P waves comprising concluding the tip of the central venous catheter is located in the right atrium when the first height is at least twice the second height. 4. The method according to claim 2, wherein the first shape has a first height and the second shape has a second height, determining the location of the tip of the central venous catheter based on the comparison of the first and second P waves and the comparison of the reference P wave with at least one of the first and second P waves comprising concluding the tip of the central venous catheter is located in the right ventricle when the first height is at most half the second height. 5. The method according to claim 2, wherein the first shape has a first height, the second shape has a second height, and the reference shape has a reference height, determining the location of the tip of the central venous catheter based on the comparison of the first and second P waves and the comparison of the reference P wave with at least one of the first and second P waves comprising: concluding the tip of the central venous catheter is located in the superior vena cava when the first height is approximately equal to both the second height and the reference height; andcon... | Bard Access Systems, Inc. | A61B005/02 A61B005/04 A61B005/05
US08376968 | B2 | 2013-02-19 | Method and system for quantifying an intention of movement of a user | 2009-05-15 | A method for quantifying an intention of movement of a user, the method comprising: extracting a mechanomyogram (MMG) signal in response to surface vibration of a muscle of the user; and averaging the MMG signal by calculating the MMG root mean squared (RMS) amplitude or frequency variance in a predetermined time window; where an increasing RMS amplitude of the MMG signal corresponds to increasing muscle voluntary contraction or a decreasing frequency variance of the MMG signal corresponds to increasing muscle voluntary contraction. | 1. A method for determining an intention of movement of a user, the method comprising: extracting a mechanomyogram (MMG) signal from one or more vibration sensors in response to surface vibration of a muscle of the user;averaging the MMG signal by calculating the MMG root mean squared (RMS) amplitude or frequency variance in a predetermined time window;determining a change in the RMS amplitude or frequency variance of the MMG signal over time, wherein an increasing RMS amplitude or a decreasing frequency variance of the MMG signal is used to identify a voluntary contraction of the muscle; anddetermining an intention of the user to contract the muscle based on the determined change in the RMS amplitude or frequency variance of the MMG signal. 2. The method according to claim 1, wherein the frequency variance of the MMG signal is the scale or degree of frequency being spread out in a power density spectrum (PDS). 3. The method according to claim 1, wherein both RMS MMG amplitude and the frequency variance of the MMG signal are used to estimate a level of muscle contraction. 4. The method according to claim 1, wherein the muscle contraction level is normalized to a percentage of maximal voluntary contraction (MVC). 5. The method according to claim 4, further comprising instructing a robotic system according to the estimated percentage of MVC. 6. The method according to claim 1, wherein the RMS amplitude is normalized as a percentage of a span between a maximal and a minimum value. 7. The method according to claim 1, wherein the frequency variance of the MMG signal is normalized as a percentage of a span between a maximal and a minimum value. 8. The method according to claim 1, wherein the predetermined time window is 1 second with 100 ms step forward. 9. A system for determining an intention of movement of a user, the system comprising: a measurement unit to extract a mechanomyogram (MMG) signal using one or more vibration sensors in response to surface vibration of a muscle of the user; anda signal processing unit to average the MMG signal by calculating the MMG root mean squared (RMS) amplitude or frequency variance in a predetermined time window, to determine a change in the RMS amplitude or frequency variance of the MMG signal over time, wherein an increasing RMS amplitude or a decreasing frequency variance of the MMG signal is used to identify a voluntary contraction of the muscle, and to determine an intention of the user to contract the muscle based on the determined change in the RMS amplitude or frequency variance of the MMG signal. 10. The system according to claim 9, further comprising another signal processing unit to pre-process a raw signal corresponding to the surface vibration of a muscle of the user. 11. The system according to claim 9, further comprising a modeling unit to translate the RMS amplitude signal and the frequency variance of the MMG signal into a set of instructions for recording or actuating a robot or an act... | The Hong Kong Polytechnic University | A61B005/00
US08992430 | B2 | 2015-03-31 | Breath sampling system | 2007-10-25 | There is provided herein a breath sampling device comprising a controller adapted to initiate a signal to trigger a valve functionally associated with a breath sampling tube upon said controller receiving a signal indicating a blockage in the breath sampling tube. The valve is adapted to facilitate an increase of pressure within the breath sampling tube between the blockage and a pump functionally associated with the breath sampling tube. There is further provided herein an airway adapter comprising a restrictor located in proximity to a connection point between the airway adapter and a breath sampling tube, wherein the restrictor may be adapted to reduce the flow rate of liquids entering the breath sampling tube. There is further provided herein an airway adapter comprising one or more openings located near a connection point between said airway adapter and a breath sampling tube, wherein said one or more openings are adapted to allow inflow of air to reduce the flow rate of liquids entering the breath sampling tube. | 1. A breath sampling device comprising: a breath sampling tube;a pump configured to create a negative pressure in the breath sampling tube;a pressure sensor configured to monitor pressure in said breath sampling tube, wherein pressure reaching a predetermined threshold is indicative of a blockage;a valve associated with said breath sampling tube; anda controller configured to deactivate said pump when the predetermined pressure threshold is reached and, while said pump is deactivated, to initiate said valve to toggle back and forth between a NORMAL mode of operation, and an AUTO-ZERO mode of operation, said AUTO-ZERO mode of operation allowing air to enter said breath sampling tube between the blockage and the pump, thereby increasing pressure in said breath sampling tube so as to prevent additional liquids from entering said breath sampling tube. 2. The device according to claim 1, wherein the predetermined pressure threshold is a predetermined pressure value within said breath sampling tube. 3. The device according to claim 1, wherein the predetermined pressure threshold is a predetermined pressure change rate within said breath sampling tube. 4. The device according to claim 1, wherein said valve is a solenoid valve. 5. The device according to claim 1, further comprising a restrictor located in proximity to a connection point between an airway adapter and said breath sampling tube, wherein said restrictor is configured to reduce a flow rate of liquids entering said breath sampling tube. 6. The device according to claim 5, wherein said restrictor comprises a restricting hole. 7. The device according to claim 1, further comprising one or more openings located near a connection point between an airway adapter and said breath sampling tube, wherein said one or more openings are configured to allow inflow of air to reduce a flow rate of liquids entering said breath sampling tube. | Oridion Medical 1987 Ltd. | A61B005/02 A61B005/087 A61B005/083
US08956293 | B2 | 2015-02-17 | Graphical ‘mapping system’ for continuously monitoring a patient's vital signs, motion, and location | 2009-05-20 | The invention provides a body-worn monitor that measures a patient's vital signs (e.g. blood pressure, SpO2, heart rate, respiratory rate, and temperature) while simultaneously characterizing their activity state (e.g. resting, walking, convulsing, falling). The body-worn monitor processes this information to minimize corruption of the vital signs by motion-related artifacts. A software framework generates alarms/alerts based on threshold values that are either preset or determined in real time. The framework additionally includes a series of ‘heuristic’ rules that take the patient's activity state and motion into account, and process the vital signs accordingly. These rules, for example, indicate that a walking patient is likely breathing and has a regular heart rate, even if their motion-corrupted vital signs suggest otherwise. | 1. A system for continuously monitoring a plurality of patients, the system comprising: a plurality of patient monitors, each patient monitor configured to be worn on the body of one of the plurality of patients, each patient monitor comprising: a first sensor comprising a first detector configured to detect a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;a second sensor comprising a second detector configured to detect a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart;at least two motion-detecting sensors positioned on locations selected from a forearm, upper arm, and a body location other than the forearm or upper arm of the patient, each of the motion-detecting sensors generating at least one motion waveform indicative of motion of a body location onto which the motion-detecting sensor is affixed;a location sensor comprising a wireless component and a location processing component, the location processing component configured to receive a signal from the wireless component and process this signal to determine a physical location of the patient to which the location sensor is attached;a processing component configured to receive the time-dependent physiological and motion waveforms and determine therefrom: (i) at least one vital sign calculated using at least one of the first and second time-dependent physiological waveforms; (ii) motion parameters corresponding to the patient's posture and activity state calculated using at least one motion waveform generated by each of the motion-detecting sensors; and (iii) an alarm parameter calculated from the at least one of the vital sign and motion parameters; and,a wireless transceiver configured to wirelessly transmit the vital sign, motion parameter, location of the patient, and alarm parameter through a wireless system; and,a remote computer system comprising a display and an interface to the wireless system, the interface configured to receive from the wireless system a vital sign, motion parameter, location of the patient, and motion-dependent alarm parameter corresponding to each of the plurality of patients, andthe display configured to render a first user interface that, for each patient in the plurality of patients, displays (i) at least one vital sign, (ii) a graphical icon simultaneously indicating patient posture and activity state, (iii) a location of the patient, and (iv) motion-dependent alarm parameter, and to render a second user interface that displays a map corresponding to a monitored area of a hospital, the map depicting the current location of each patient in the plurality of patients within the monitored area of the hospital. 2. The system of claim 1, wherein the display is configured to render the motion-dependent alarm parameter as a graphical icon for each of the plurality of patients comprising a patient monitor. 3. The system of claim 1... | Sotera Wireless, Inc. | A61B005/0205 A61B005/103 G08B019/00 A61B005/00 A61B005/021 A61B005/11 G08B021/04 A61B005/01 A61B005/024 A61B005/0245 A61B005/04 A61B005/0404 A61B005/0452 A61B005/0456 A61B005/08 A61B005/145 G06F019/00
US08461997 | B2 | 2013-06-11 | Transceiver unit in a measurement system | 2010-04-30 | A measurement system, comprises a sensor wire and a transceiver unit. The sensor wire comprises an insertable portion configured to be inserted in a blood vessel of a patient's body and a sensor disposed within the insertable portion at a distal end of the sensor wire. The sensor is configured to measure a variable inside the patient. The transceiver unit comprises: a housing adapted to be connected to a proximal end of the sensor wire; and a first communication module within the housing adapted to wirelessly communicate by a communication signal with an external second communication module in order to transfer information to the external second communication module. The communication signal, including sensor values, is generated by the transceiver unit and transferred as a data stream. The housing is configured to remain external to the patient's body. | 1. A pressure measurement system, comprising: a pressure sensor wire comprising an insertable portion configured to be inserted in a blood vessel of a patient's body and a pressure sensor disposed within the insertable portion at a distal end of the pressure sensor wire, wherein the pressure sensor is configured to measure pressure inside the patient; anda transceiver unit, comprising: a housing adapted to be connected to a proximal end of the pressure sensor wire; anda first communication module within the housing adapted to wirelessly communicate by a communication signal with an external second communication module in order to transfer pressure information to the external second communication module; andwherein the communication signal, including pressure sensor values, is generated by the transceiver unit and transferred as a data stream, andwherein the housing is configured to remain external to the patient's body. 2. The pressure measurement system according to claim 1, wherein said communication signal is a radio frequency signal. 3. The pressure measurement system according to claim 2, wherein radio frequency communication is performed by Bluetooth. 4. The pressure measurement system according to claim 2, wherein said radio frequency signal has a frequency of 2.4 GHz. 5. The pressure measurement system according to claim 2, wherein radio frequency communication is performed by a frequency hopping technique. 6. The pressure measurement system according to claim 1, wherein said data stream is in a form of data packets. 7. The pressure measurement system according to claim 1, wherein the first communication module is configured to perform bi-directional communication with the second communication module. 8. The pressure measurement system according to claim 1, wherein said communication signal is one of an infrared signal, an ultrasound signal, and a light signal. 9. The pressure measurement system according to claim 1, wherein the transceiver unit further comprises an activator configured to activate and initiate the transceiver unit. 10. The pressure measurement system according to claim 9, wherein the transceiver unit further comprises a female connector with electrical contact surfaces, and wherein said activator activates and initiates the transceiver unit when electrical connections are established between the electrical contact surfaces of the female connector and electrical contact surfaces of a male connector at the proximal end of the pressure sensor wire. 11. The pressure measurement system according to claim 1, wherein the transceiver unit further comprises a guide to guide the pressure sensor wire during insertion into the patient. 12. The pressure measurement system according to claim 1, wherein the transceiver unit further comprises an energizer to energize the transceiver unit and circuitry of the connected pressure sensor wire. 13. The pressure measurement system according to claim 1, further comprising a communi... | St. Jude Medical Systems AB | G08B023/00 A61B005/00 A61N001/00
US08562951 | B2 | 2013-10-22 | Methods for classifying and treating physiologic brain imbalances using quantitative EEG | 2010-05-25 | Neurophysiologic information such as quantitative electroencephalography (QEEG) is used in a method for classifying, diagnosing, and treating physiologic brain imbalances. Neurophysiologic information is also used to guide sample selection in clinical tests for psychopharmacologic drug candidates. Finally, neurophysiologic information is used for remotely assessing and treating patients with physiologic brain imbalances. | 1. A method for classifying quantified electroencephalographic data, comprising: a) providing; i) a first data set collected from an untreated patient wherein said first data set comprises quantified electroencephalographic data and behavioral indicia scores; andii) a second data set collected from a plurality of asymptomatic individuals wherein said second data set comprises quantified electroencephalographic data and individual behavioral indicia scores;b) establishing pattern differences by comparing said first data set and said second data set;c) classifying said untreated patient quantified electroencephalographic data as defining a responsivity profile based upon said pattern differences. 2. The method of claim 1, wherein said method further comprises extracting at least one untreated patient multivariable from said first data set. 3. The method of claim 1, wherein said method further comprises extracting a plurality of asymptomatic individual multivariables from said second data set. 4. The method of claim 1, wherein said pattern differences comprise a plurality of Z scores comparing said untreated patient quantified electroencephalographic data set with said asymptomatic individual quantified electroencephalographic data. 5. The method of claim 1, wherein said pattern differences comprise a plurality of Z scores comparing said untreated patient behavioral indicia scores and said asymptomatic patient behavioral indicia scores. 6. The method of claim 1, wherein said untreated patient behavioral indicia scores are selected from the group consisting of clinical global improvement scores, Beck Depression scores, and Hamilton-D scores. 7. The method of claim 1, wherein said asymptomatic individual behavioral indicia scores are selected from the group consisting of clinical global improvement scores, Beck Depression scores, and Hamilton-D scores. 8. The method of claim 2, wherein said at least one untreated patient multivariable is derived from an output measurement selected from a group consisting of absolute power, relative power, frequency, intrahemispheric coherence, interhemispheric coherence, intrahemispheric asymmetry, and interhemispheric asymmetry, and ratios or combinations thereof 9. The method of claim 3, wherein said plurality of asymptomatic individual multivariables are derived from output measurements selected from a group consisting of absolute power, relative power, frequency, intrahemispheric coherence, interhemispheric coherence, intrahemispheric asymmetry, and interhemi spheric asymmetry, and ratios or combinations thereof. 10. The method of claim 1, wherein said responsivity profile comprises changes in said untreated patient quantified electroencephalographic data selected from the group consisting of absolute power average, frontal midline progression index, ratio of frontal/posterior alpha indicies, and average midline theta/beta indicies. 11. The method of claim 1, wherein said responsivity profile comprise ch... | CNS Response | A61K049/00
US08654986 | B2 | 2014-02-18 | System for remotely obtaining audiometric measurements and adjusting hearing aids via the internet | 2008-10-24 | System and method for remotely obtaining audiometric measurements and adjusting hearing aids via the Internet, using telematic means such as video conferencing and e-mail to establish a remote connection between the patient who may be at home, in a medical center or in a pharmacy and a hearing aid specialist at an audiometry laboratory or clinic, whereby the specialist remotely provides the patient with instructions for carrying out an audiometry test and, depending on the results, subsequently transmits adjustment signals via the Internet to the patient's hearing aid which is connected to a computer to which the patient has access. | 1. A method for remotely obtaining audiometric measurements and adjusting hearing aids between a patient and a hearing aid specialist, comprising the following steps: a computer from said hearing aid specialist's medical office is connected to the patient's computer via a telematic network,a web camera from the hearing aid specialist sends data and images with instructions to the patient's computer screen,a series of incrementally increasing and decreasing sequences in pure tone intensity, for a specific number of frequencies, are sent from the hearing aid specialist's computer to the patient's computer, and which are reproduced through loud speakers connected to the patient's computer, or through a hearing aid connected to the patient's computer,the patient sends, at least, affirmative signals confirming that they can hear, from their computer to the hearing aid specialist's computer,the hearing aid specialist calculates a minimum hearing threshold of the patient for each one of the pure tone emission frequencies according to the affirmative signals confirming that they could hear them,the hearing aid specialist's computer sends, via the telematic network, an adjustment signal of the patient's hearing aid generated according to the patient's said minimum hearing thresholds and which are received in the patient's computerthe hearing aid connected to the patient's computer is adjusted according to said adjustment signal. 2. A method for remotely obtaining audiometric measurements and adjusting hearing aids between a patient and a hearing aid specialist, according to the previous claim, additionally comprising the following steps: the patient receives audio signals to check whether their hearing aid is correctly reprogrammed to their necessities by means of an anechoic chamber,the patient sends information through telematic means to the hearing aid specialist with the measurement data in the anechoic chamber,the hearing aid specialist sends a new signal to adjust the patient's hearing aid via the telematic network, the signal being calculated according to said measurement data in the anechoic chamber, the hearing aid being adjusted accurately. 3. A method for remotely obtaining audiometric measurements and adjusting hearing aids between a patient and a hearing aid specialist, according to claim 1, whereas the hearing aid specialist complements the audiometry by sending the patient a series of standard digitalized and tested words. 4. A method for remotely obtaining audiometric measurements and adjusting hearing aids between a patient and a hearing aid specialist, according to claim 2, whereas the hearing aid specialist complements the audiometry by sending the patient a series of standard digitalized and tested words. 5. A method for remotely obtaining audiometric measurements and adjusting hearing aids between a patient and a hearing aid specialist, according to claim 1, whereas the hearing aid specialist's computer registers the historical re... | - | H04R029/00
US09022954 | B2 | 2015-05-05 | Device for measuring real-time pressure exerted by a support surface on a perineal artery | 2010-05-12 | Provided are methods and devices for measuring real-time pressures on the perineum of a user positioned on a surface. In an embodiment the surface is a bicycle seat and the real-time measurement is a pressure map of the pressure distribution during a bicycle ride. The methods relate to connecting a pressure sensor to the perineum region of the user, positioning the user on the surface, wherein the pressure sensor is positioned between the perineum and the surface, and obtaining output from the pressure sensor while the user is positioned on the surface, such as during a bicycle ride. The devices relate to a plurality of pressure sensors capable of being connected to the perineum region and a microcontroller containing a drive circuit and chip for analog-to-digital conversion and storage. One use of the methods and devices are to assess risk of erectile dysfunction for bicycle riders. | 1. A method of measuring the percentage of the arterial occlusion pressure exerted by a surface on one or more arteries of the perineum of a human male user positioned on the surface, said method comprising the steps of: connecting a pressure sensor to the perineum region of the human male user, wherein the pressure sensor is not affixed to the surface;positioning the human male user on the surface, wherein the pressure sensor is positioned between one or more arteries of the perineum and the surface;obtaining output from the pressure sensor while the human male user is positioned on the surface;comparing the output to the arterial occlusion pressure of the one or more arteries of the perineum; anddetermining the output's percentage of the arterial occlusion pressure. 2. The method of claim 1, wherein the surface is formed from a bicycle seat and the method further comprises the step of: evaluating the risk of erectile dysfunction by the bicycle seat on the user from the obtained output. 3. The method of claim 2, wherein a plurality of pressure sensors are connected to the perineum region of the user to obtain a pressure map of the perineum region. 4. The method of claim 3, wherein the pressure map corresponds to the pressure of one or more of the following arteries of the perineum region selected from the group consisting of: left proximal artery; left distal artery; right proximal artery, right distal artery, and internal pudendal artery. 5. The method of claim 3, wherein the pressure map is generated by six individual pressure sensors. 6. The method of claim 3, wherein the output is a time course of the pressure map. 7. The method of claim 6, wherein the time course is for a period that is greater than or equal to two hours. 8. The method of claim 2, wherein the obtained output is from non-stationary bicycle riding for a time period. 9. The method of claim 2, wherein the evaluation step further comprises comparing the obtained output to a user-selected pressure level. 10. The method of claim 9, wherein the risk of erectile dysfunction is identified as high for a bicycle seat that exerts a maximum pressure on the perineum that is greater than or equal to the user-selected pressure level for a time period that is greater than or equal to a user-selected time period, wherein the user-selected pressure level is related to a perineal artery occlusion pressure. 11. The method of claim 10, further comprising the step of determining perineal artery occlusion pressure for the user by: exerting a pressure on a perineal artery of the user;identifying the pressure as the perineal artery occlusion pressure for the user; andidentifying the output of the pressure sensor for the perineal artery occlusion pressure to obtain an occlusion pressure sensor value. 12. The method of claim 11 wherein the evaluating step comprises comparing a maximum output of the pressure sensor against the occlusion pressure sensor value; andidentifying the risk of erecti... | The Board of Trustees of the University of Illinois | A61B005/02 A61B005/117 A61B005/103 A61B005/00
US08147403 | B2 | 2012-04-03 | Capsule propulsion device and propulsion method | 2010-06-01 | A capsule propulsion device has: at least one acoustic streaming generator for generating acoustic streaming as a flow of the a medium in a space; a control unit for driving and controlling the acoustic streaming generator; and a distance obtaining unit for obtaining an object distance between the acoustic streaming generator and an object existing in the space. The control unit drives and controls the acoustic streaming generator on the basis of the distance information obtained by the distance obtaining unit. | 1. A capsule propulsion device which is introduced in a space in which a medium for transmitting a sound wave exists, the capsule propulsion device comprising: at least one acoustic streaming generator that generates acoustic streaming as a flow of the medium in the space;a control unit that drives and controls the at least one acoustic streaming generator;a distance obtaining unit that obtains an object distance between one of the at least one acoustic streaming generator and an object existing in the space; andan input receiving unit that receives an input in a travel direction of the capsule propulsion device in the space in accordance with an operation instruction input by an operation instruction input unit that is separated from the capsule propulsion device,wherein the control unit drives and controls the at least one acoustic streaming generator on the basis of the distance information obtained by the distance obtaining unit, andthe control unit specifies one or more acoustic streaming generator to be driven for travel in the travel direction in accordance with an operation signal transmitted corresponding to the operation instruction input by the operation instruction input unit and, when an object distance obtained by the distance obtaining unit is less than a predetermined threshold that is defined in advance to be a distance required for generating the acoustic streaming between the specified one or more acoustic streaming generator and the object, the control unit stops driving of the specified specified one or more acoustic streaming generator. 2. The capsule propulsion device according to claim 1, wherein the control unit specifies at least one acoustic streaming generator for travel in a direction opposite to the travel direction and drives it for predetermined time. 3. The capsule propulsion device according to claim 1, wherein the at least one acoustic streaming generator includes at least one first acoustic streaming generator that gives a propulsion force to the capsule propulsion device by generating acoustic streaming and at least one second acoustic streaming generator that gives a rotational force to the capsule propulsion device by generating acoustic streaming, andthe control unit drives the second acoustic streaming generator, assuring the object distance obtained by the distance obtaining unit, which is equal to or larger than the predetermined threshold, to turn the capsule propulsion device, and controls so that the object distance between the one of the at least one first acoustic streaming generator and the object becomes the predetermined threshold or larger. 4. The capsule propulsion device according to claim 1, wherein the control unit specifies at least one first acoustic streaming generator for travel in the travel direction and, when an object distance between the first acoustic streaming generator and the object, obtained by the distance obtaining unit is less than the predetermined threshold, the control... | Olympus Medical Systems Corp. | A61B005/00
US08666466 | B2 | 2014-03-04 | Device and method for monitoring of absolute oxygen saturation and tissue hemoglobin concentration | 2010-06-10 | A method and medical device for detecting signals that detects emitted light scattered by a volume of tissue delivered along a first pathway and a second pathway different from the first pathway, detects emitted light scattered by a volume of tissue delivered along a third pathway and a fourth pathway different from the third pathway, determines a first uniformity corresponding to the emitted light detected along the first pathway and the second pathway, determines a second uniformity corresponding to the emitted light detected along third pathway and the fourth pathway, determines a total uniformity in response to the determined first uniformity and the determined second uniformity, and alters sensing by the device in response to the determined total uniformity. | 1. A method of detecting signals in a medical device, comprising: detecting emitted light scattered by a volume of tissue delivered along a first pathway and a second pathway different from the first pathway;detecting emitted light scattered by a volume of tissue delivered along a third pathway and a fourth pathway different from the third pathway;determining a first uniformity corresponding to the emitted light detected along the first pathway and the second pathway;determining a second uniformity corresponding to the emitted light detected along third pathway and the fourth pathway;determining a total uniformity in response to the determined first uniformity and the determined second uniformity; andaltering sensing by the device in response to the determined total uniformity. 2. The method of claim 1, further comprising comparing the first uniformity and the second uniformity and altering sensing by the device in response to the comparing. 3. The method of claim 1, further comprising: determining whether a first difference between the emitted light detected along the first pathway and the emitted light detected along the second pathway is greater than a predetermined threshold;determining whether a second difference between the emitted light detected along the third pathway and the emitted light detected along the fourth pathway is greater than a predetermined threshold; anddetermining the total uniformity in response to the determined first difference and the determined second difference. 4. The method of claim 3, wherein determining the total uniformity comprises: determining whether the first difference and the second difference are both less than the predetermined threshold; anddetermining if both one of the first difference and the second difference is less than the threshold and the other of the first difference and the second difference is not less than the predetermined threshold. 5. The method of claim 4, wherein altering sensing by the device comprises: changing the functionality of the device in response to both one of the first difference and the second difference being less than the threshold and the other of the first difference and the second difference not being less than the predetermined threshold; anddisabling sensing by the device in response to determining both the first difference and the second difference not being less than the predetermined threshold. 6. The method of claim 4, further comprising detecting a change in uniformity in response to the determined total uniformity, and wherein altering sensing by the device comprises generating an alert in response to the detected change. 7. The method of claim 1, wherein determining a first uniformity comprises: detecting emitted light scattered by the volume of tissue delivered along the first pathway to generate corresponding first detected light intensity output signals;detecting emitted light scattered by the volume of tissue delivered along the second pathway to ge... | Medtronic, Inc. | A61B005/00
US08833133 | B2 | 2014-09-16 | Apparatus for evaluation of a device for measuring lung diffusion capacity | 2010-06-01 | A test apparatus for evaluating a measuring device for measuring a diffusion capacity of a person's lungs by measuring a change of concentration of a gas species in a mixture, when the mixture is inhaled and subsequently exhaled by a person, wherein the test apparatus includes (a) a port for connecting to the measuring device; (b) a compartment coupled to the port; (c) a circulation circuit configured to circulate gas from the compartment back to the compartment; and (d) a gas species removing element configured to selectively reduce a partial pressure of the gas species at a point in the circulation circuit. | 1. A test apparatus for evaluating a measuring device for measuring a diffusion capacity of a person's lungs by measuring a change of concentration of a gas species in a mixture, when the mixture is inhaled and subsequently exhaled by a person, the test apparatus comprising: a port for connecting to the measuring device;a compartment coupled to the port;a circulation circuit configured to circulate gas from the compartment back to the compartment; anda gas species removing element configured to selectively reduce a partial pressure of the gas species at a point in the circulation circuit,further comprising a catalytic converter that removes the gas species incorporated in the circulation circuit. 2. The test apparatus according to claim 1, further comprising a flow regulator configured to regulate a flow of the circulating gas to a set flow rate. 3. The test apparatus according to claim 1, further comprising a mixer configured to mix the mixture in the compartment. 4. The test apparatus according to claim 1, wherein the gas species removing element is configured to reduce the partial pressure substantially to zero. 5. The test apparatus according to claim 1, wherein the gas species is carbon monoxide. 6. The test apparatus according to claim 1, wherein the compartment has an adjustable volume. 7. A test apparatus for evaluating a measuring device for measuring a diffusion capacity of a person's lungs by measuring a change of concentration of a gas species in a mixture, when the mixture is inhaled and subsequently exhaled by a person, the test apparatus comprising: a port for connecting to the measuring device;a compartment coupled to the port;a circulation circuit configured to circulate gas from the compartment back to the compartment; anda gas species removing element configured to selectively reduce a partial pressure of the gas species at a point in the circulation circuit, further comprising a drying material in the circulation circuit preceding the gas species removing element. 8. A test apparatus for evaluating a measuring device for measuring a diffusion capacity of a person's lungs by measuring a change of concentration of a gas species in a mixture, when the mixture is inhaled and subsequently exhaled by a person, the test apparatus comprising: a port for connecting to the measuring device;a compartment coupled to the port;a circulation circuit configured to circulate gas from the compartment back to the compartment; anda gas species removing element configured to selectively reduce a partial pressure of the gas species at a point in the circulation circuit,further comprising a variable volume chamber in the circulation circuit, coupled between the gas species removing element and an output coupled to the compartment and valves coupled between the gas species removing element and the compartment and between the variable volume chamber and the compartment. 9. The test apparatus according to claim 8, further comprising a control c... | Erasmus University Medical Center Rotterdam | G01N021/00
US08894718 | B2 | 2014-11-25 | System for remote management in ambient intelligence environments using electromyographic signals | 2008-11-26 | The system is comprised of an interface (1) consisting of one or more electrodes capturing electromyographic (biolectric muscle activity) signals, an earphone/microphone and a radiofrequency transmitter/receiver; a nexus module (2), with means for the encoding and interpretation of the electromyographic signals which taking into account both their duration as well as their amplitude and time lapse interval, wirelessly connected to the transmitter/receiver of the interface (1); and a controller (3), consisting of means for controlling and managing the devices comprising the ambient intelligence environment of a domotic, inmotic, military or other types, the input of which is connected to the output of the nexus module (2).;Additionally, the communication facilitated by the earphone/microphone is completed with glasses incorporating a device capable of projecting visual messages on the inner side thereof. | 1. A system for remote management in ambient intelligence environments comprising: an interface (1) further comprising: an earphone or a microphone,at least one electromyographic signal-capturing electrode, wherein the electromyographic signal is generated by voluntary movement of a muscle or a muscle group,and a means for calibration, encoding and interpretation of a plurality of electromyographic signals, determining a Myogenic Language based on three parameters, the first two parameters being an amplitude (A) and a duration (D) of an electromyographic signal, the amplitude is categorized into a (AS) variable or a (AL) variable, wherein the (AL) variable represents an electromyographic signal having a larger amplitude than an amplitude of a designated electromyographic signal, and the duration is categorized into a (DS) variable or a (DL) variable, wherein the (DL) variable represents an electromyographic signal having a longer duration than the designated electromyographic signal, such that a combination of these variables is used to generate at least one myogenic four letter code including “AS/DS”, “AS/DL”, “AL/DS” and “AL/DL”,the third parameter being defined as a time lapse interval between a generation of one myogenic four letter code and a next myogenic four letter code;a nexus module (2), an input of which is connected to the interface (1), and an output of which is a wireless transmitter-receiver;a controller (3) comprising means for controlling and managing a plurality of devices comprising an ambient intelligence environment, an input of the controller is wirelessly connected to the output of the nexus module (2). 2. The system for remote management in ambient intelligence environments in accordance with claim 1, wherein the interface (1) further comprises a micro-projector device capable of projecting visual messages including texts, icons or graphics, onto glasses connected to the system. 3. The system for remote management in ambient intelligence environments in accordance with claim 1, wherein the controller for controlling and managing the plurality of devices comprising the ambient intelligence environment is further comprised of PLCs, actuators and telecommunications systems. | - | A61F002/70 A61B005/04 A61B005/0488 A61B005/00 G06F003/01 A61B005/11
US08295430 | B2 | 2012-10-23 | Image guided radiation therapy | 2010-06-02 | Radiation therapy of a lesion in a part of a patient is carried out by maintaining the patient on a patient support table in fixed position using an immobilization device while the table is rotated between a magnetic resonance imaging system, a CT imaging system for generating a 360 degree scanned image of the patient at the location of the lesion and the radiation therapy system for generating a beam of radiation for treatment of the lesion and for scanning the beam 360 degrees around the lesion. The MR image locates the lesion and the CT system is used to calculate the treatment. Registration between the MR image, the CT image and the radiation treatment is provided by the fixed position of the part of the patient on the table which is held fixed using an immobilization system suitable for the part concerned which may include a molded head mask where the head is involved. | 1. A system for use in radiation therapy of a lesion in a part of a patient comprising: a patient support table for receiving and supporting the patient during planning for and effecting of the therapy of the lesion;an immobilization device for locating the part of the patient on the table so as to remain fixed relative to the table during the planning and therapy;a magnetic resonance imaging system for obtaining MR images of the lesion of the patient including a magnet, at least one RF coil and an imaging control system;a CT imaging system for generating a 360 degree scanned image of the patient at the location of the lesion;a radiation therapy system for generating a beam of radiation for treatment of the Lesion and for scanning the beam 360 degrees around the lesion;the magnetic resonance imaging system, CT imaging system and the radiation therapy system being arranged at angularly spaced positions around the patient support tablethe magnetic resonance imaging system being arranged so as to provide relative movement between the table and the system to enable the imaging of the patient while the patient remains in fixed position relative to the table;the CT imaging system being arranged so as to provide relative movement between the table and the system to enable the imaging of the patient while the patient remains in fixed position relative to the table;the radiation therapy system being arranged so as to provide relative movement between the table and the system to enable the radiation treatment of the patient while the patient remains in fixed position relative to the table;the CT imaging system and the radiation therapy system being arranged so that the image and the treatment being arranged such that they lie in a common plane around the lesion in the patient;the table and immobilization device being MR compatible, radiation translucent, and non-conductive so as to allow the imaging and treatment of the patient while stationary on the table;and a treatment control system arranged to locate the lesion using the MR image and to plan the radiation treatment to vary the intensity of radiation around the plane using the CT image. 2. The system according to claim 1 wherein the immobilization device includes a head neck thermoplastic mask which is molded to match the patient anatomy. 3. The system according to claim 1 wherein the immobilization device includes an immobilization board arranged to be placed in fixed position on the table. 4. The system according to claim 3 wherein where the RF coil includes a base part which is incorporated into the immobilization board which is capable of accepting the head neck thermoplastic mask. 5. The system according to claim 1 wherein the table includes a perforated section located underneath the location of the lesion. 6. The system according to claim 5 wherein the perforated section is located at a perforated head extension board. 7. The system according to claim 6 wherein head extension board is for... | Imris Inc. | A61B006/00 A61N005/10
US08463345 | B2 | 2013-06-11 | Device and method for monitoring of absolute oxygen saturation and total hemoglobin concentration | 2010-06-10 | A medical device for monitoring of oxygen saturation includes an optical sensor adapted for positioning adjacent to a tissue volume. The optical sensor has a light emitting portion capable of emitting light at a plurality of wavelengths and a light detecting portion capable of generating an electrical output signal corresponding to light incident on the detecting portion. A control module coupled to the optical sensor controls the light emitted by the light emitting portion. A monitoring module receives the output signal from the light detecting portion and computes a volume-independent measure of oxygen saturation in the volume of tissue using the output signal. | 1. A medical device for monitoring oxygen saturation, comprising: a sealed housing comprising a first opening and a second opening;a light emitting portion comprising a first lens sealed in the first opening of the housing, the light emitting portion emitting light at a plurality of wavelengths including at least four spaced apart wavelengths through the lens into a volume of a patient's body tissue adjacent to the lens;a light detecting portion comprising a second lens sealed in the second opening of the housing, the light detecting portion generating an electrical output signal corresponding to light incident on the light detecting portion, the incident light comprising light emitted by the emitting portion, scattered by the body tissue volume, and passing through the second lens;a control module coupled to the light emitting portion to provide control signals controlling the light emitted by the light emitting portion; anda monitoring module coupled to the light detecting portion to receive the output signal and to compute a volume-independent measure of oxygen saturation in the volume of tissue in response to the output signal, wherein each of the light emitting portion and the light detecting portion comprise a light source for emitting light and a light detector for detecting light and the functionality of the light emitting portion and the light detecting portion as one of emitting and detecting can be selectively switched. 2. The device of claim 1, wherein the light emitting portion comprises at least four light sources each emitting light at separate spaced apart wavelengths. 3. The device of claim 2, wherein the light emitting portion comprises a light diffusing material to diffuse the separate spaced apart wavelengths. 4. The device of claim 1, wherein the separate spaced apart wavelengths comprises: a first wavelength spacing between a first wavelength and a second wavelength; anda second wavelength spacing between the second wavelength and a third wavelength, wherein the first and the second wavelength spacings are approximately equal. 5. The device of claim 2, wherein the separate spaced apart wavelengths comprises: a first wavelength spacing between a first wavelength and a second wavelength; anda second wavelength spacing between the second wavelength and a third wavelength, wherein the first and the second wavelength spacings are unequal. 6. The device of claim 1, wherein the light emitting portion comprises a reference photodetector providing a reference signal corresponding to an intensity of emitted light. 7. The device of claim 6 wherein the control module receives the reference signal, the control module configured to adjust a drive signal applied to the light emitting portion in response to the reference signal. 8. The device of claim 6 wherein the monitoring module receives the reference signal and computes the volume-independent measure of oxygen saturation using the reference signal. 9. The device of claim 1, whe... | Medtronic, Inc. | A61B005/0215
US08750979 | B2 | 2014-06-10 | Method and implantable medical device (IMD) for monitoring permeability status of cell membranes | 2008-01-29 | An implantable medical device comprises a signal generator for generating a current signal having a frequency in a frequency window slightly less than the β-dispersion frequency of a tissue and applying the signal over the tissue. A signal measurer measures the resulting voltage signal and an impedance parameter is calculated from the applied and measured signal by a parameter determiner. A status monitor monitors the permeability status of cell membranes in the tissue based on this impedance parameter. | 1. An implantable medical device comprising: at least two implantable electrodes configured to communicate electric signals with a tissue;a memory configured to store a predetermined value representative of a β-dispersion frequency associated with cell membranes in said tissue;a signal generator configured to generate a first signal at a frequency in a first frequency window as a function of the stored predetermined value representative of the β-dispersion frequency, the frequency window extending from the stored predetermine value representative of the β-dispersion frequency to a frequency corresponding to about 90% of the stored value representative of said β-dispersion frequency, said first signal being communicated to said tissue via said at least two implantable electrodes;a signal measurer configured to measure a first resulting signal collected by said at least two implantable electrodes;a parameter determiner configured to determine, based on said first signal and said first resulting signal, a first impedance parameter representative of an impedance of said tissue; anda status monitor configured to monitor a permeability status of said cell membranes in said tissue based on said first impedance parameter. 2. The device according to claim 1, wherein said signal generator generates said first signal at a frequency in a first frequency window extending from a frequency corresponding to about 99% of said β-dispersion frequency to said frequency corresponding to about 90% of said β-dispersion frequency. 3. The device according to claim 1, wherein said status monitor monitors said permeability status based on a comparison of said first impedance parameter and a reference impedance parameter. 4. The device according to claim 3, wherein said reference impedance parameter is a previously determined impedance parameter representative of an impedance of said tissue at a previous time instance. 5. The device according to claim 3, wherein said status monitor comprises: a calculator configured to calculate a condition parameter based on said first impedance parameter and said reference impedance parameter; anda detector configured to detect a condition resulting in an increase in cell permeability if said condition parameter is not present within a pre-defined parameter interval. 6. The device according to claim 1, wherein said signal generator is further configured to generate a second signal at a frequency in a second frequency window extending from said β-dispersion frequency to a frequency corresponding to about 110% of said β-dispersion frequency, said second signal being communicated to said tissue via said at least two implantable electrodes, said signal measurer is further configured to measure a second resulting signal collected by said at least two electrodes, said parameter determiner is further configured to determine, based on said second signal and said second resulting signal, a second impedance parameter representative of an imped... | St. Jude Medical, AB | A61B005/053
US08414488 | B2 | 2013-04-09 | Medical system, apparatus and method | 2007-11-13 | According to some embodiments, there is provided a medical monitoring device that includes one or more sensors, wherein a sensor is adapted to sense a parameter of a patient and computator adapted to receive an output of at least one sensor and to compute a condition-index-value directly related to a condition of the patient. | 1. A medical monitoring device comprising: two or more sensors adapted to sense at least two parameters of a patient, wherein said two or more sensors comprise at least a capnograph and a pulse oximeter; anda computing unit adapted to continuously receive at least two output parameters at least from said capnograph and said pulse oximeter and to continuously compute single unit-less index-values based on integration of said at least two output parameters,wherein the single unit-less index-value is indicative of a respiratory status of the patient; such that a change in the single unit-less index-value is indicative of a change in the patient respiratory status. 2. The medical monitoring device of claim 1, wherein said at least two parameters of a patient comprises a respiratory related parameter, pulmonary related parameter, cardiac related parameter or any combination thereof. 3. The medical monitoring device of claim 1, wherein said at least two parameters of a patient comprises respiration rate, EtCO2, SpO2, heart rate, or any combination thereof. 4. The medical monitoring device of claim 1, wherein said two or more sensors further comprise a heart rate monitor, blood pressure monitor, ECG, EEG, or any combination thereof. 5. The medical monitoring device of claim 1, wherein said single unit-less index-value is in the range of 1 to 10. 6. The medical monitoring device of claim 1, wherein an increase in said unit-less index-value is indicative of an improvement in a patient's status. 7. The medical monitoring device of claim 1, wherein a decrease in said unit-less index-value is indicative of a deterioration of a patient's status. 8. The medical monitoring device of claim 1, wherein said computing unit is adapted to compute said single unit-less index-value according to an average of an output of at least one sensor. 9. The medical monitoring device of claim 1, wherein said computing unit is adapted to compute said single unit-less index-value by mathematical operation of a medical significance level of each of the at least two output parameters. 10. The medical monitoring device of claim 9, wherein said medical significance level of each of the at least two output parameters is determined by correlating the output value of each of said at least two output parameters with that of a normal level. 11. The medical monitoring device of claim 9, wherein said mathematical operation comprises multiplication. 12. The medical monitoring device of claim 1, wherein said device is further adapted to compute a trend of said single unit-less index-value. 13. The medical monitoring device of claim 1, wherein said device is further adapted to compute a reliability index of said singe unit-less index-value. 14. The medical monitoring device of claim 1, further adapted to provide a pause frequency parameter. 15. The medical monitoring device of claim 1, wherein said device is further adapted to provide medical recommendation. 16. The medical monit... | Oridion Medical 1987 Ltd. | A61B005/00
US08323208 | B2 | 2012-12-04 | Neurologic monitoring system and method | 2010-05-06 | A method, system, and products that allow for accurate localization, virtual depiction, and testing of nerve(s) or neural structures that course through or around a proposed surgical site are described. The method and system employs a variety of types of neurodiagnostic tests that incorporate sensory, motor, and mixed nerve conduction studies in an orthodromic and antidromic fashion as well as utilizing triggered electromyography. This method, system, and products allow for a much more specific and sensitive method to assess the baseline function and continued integrity of nerves and neural structures prior to, during, and after a surgical procedure. | 1. A system for monitoring neural function of a nerve, comprising: a first electrode configured to be suitable for nerve stimulus delivery and nerve response recordation;a first slave box connected to the first electrode, the slave box further comprising electrical stimulus circuitry and amplification circuitry to either deliver a stimulus to the first electrode via the stimulus circuitry or to receive a response from the first electrode via the amplification circuitry;a first switch having a first position to connect the first electrode to the electrical stimulus circuitry and disconnect the first electrode from the amplification circuitry to allow delivery of a nerve stimulus at the first electrode, and a second position to connect the first electrode to the amplification circuitry and disconnect the first electrode from the electrical stimulus circuitry to allow recording a nerve response at the first electrode; anda second electrode attached to the slave box, wherein the switch in the second position connects the second electrode to the electrical stimulus circuitry and disconnects the second electrode from the amplification circuitry, and the switch in the first position connects the second electrode to the amplification circuitry and disconnects the second electrode from the electrical circuitry. 2. The system of claim 1, wherein the switch is a hard wired switch or a software switch. 3. The system of claim 1, further comprising a neuromonitoring machine that controls the connection of the first electrode to the stimulus circuitry and to the amplification circuitry. 4. The system of claim 1, comprising a hard switch that controls the connection of the first electrode to the stimulus circuitry and the amplification circuitry. 5. The system of claim 1, further comprising a ground electrode connected to the amplification circuitry. 6. The system of claim 1, wherein in the first switch position the first electrode is connected to an anode in the stimulus circuitry, and the switch in the second position connects the first electrode to an active contact in the amplification circuitry. 7. The system of claim 1, wherein the amplification circuitry comprises differential amplification circuitry. 8. The system of claim 1, wherein the system further comprises a neuromonitoring machine configured to automatically connect the first electrode to the electrical stimulus circuitry, and automatically connect the first electrode to the amplification circuitry. 9. The system of claim 8, further comprising a second electrode configured to be suitable for nerve stimulus delivery and nerve response recordation; a second slave box connected to the second electrode, the second slave box further comprising second electrical stimulus circuitry and second amplification circuitry to either deliver a stimulus to the second electrode via the second stimulus circuitry or to receive a response from the second electrode via the second amplification circuitry;and a ... |  | A61B005/04
US07985387 | B2 | 2011-07-26 | Protease inhibitor sample collection system | 2009-12-02 | A collection container and a method for collecting a biological sample, particularly whole blood, includes at least one stabilizing agent in an amount effective to stabilize and inhibit protein degradation and/or fragmentation. The stabilizing agent is able to stabilize proteases in the biological sample, particularly at the point of collection, by inhibiting protein degradation and/or fragmentation in the sample when the sample is stored. The stabilizing agent includes one or more protease inhibitors. | 1. A sterile and evacuated tube for collecting and stabilizing whole blood or a component thereof, comprising a first end and a second end and at least one interior wall defining a reservoir portion for receiving the blood, wherein said reservoir comprises at least one antioxidant and a protein stabilization agent comprising at least two lyophilized protease inhibitors, wherein a first inhibitor is an inhibitor of a serine protease and a second inhibitor is an inhibitor of another class of protease; anda closure pierceable by a needle for supplying blood to said reservoir. 2. The tube of claim 1, wherein said interior wall comprises plastic or glass. 3. The tube of claim 1, wherein said reservoir comprises more than two lyophilized protease inhibitors. 4. The tube of claim 1, wherein said lyophilized protease inhibitors comprise at least two inhibitors of a serine protease and at least one inhibitor of a cysteine protease. 5. The tube of claim 4, wherein said inhibitor of a serine protease is selected from the group consisting of AEBSF, aprotinin, antithrombin, elastatinal, leupeptin, APMSF, DFP and trypsin, and mixtures thereof. 6. The tube of claim 4, wherein said inhibitor of a serine protease comprises AEBSF, aprotinin and leupeptin. 7. The tube of claim 4, wherein said inhibitor of a cysteine protease is E-64. 8. The tube of claim 1, wherein said lyophilized protease inhibitors comprise at least one inhibitor of a serine protease, at least one inhibitor of a cysteine protease, and an inhibitor of a cysteine/serine protease. 9. The tube of claim 8, wherein said inhibitor of a cysteine/serine protease is selected from the group consisting of leupeptin, TPCK, TLCK-HCl (L-1-chloro-3-[4-tosylamidol-7-amino-2-heptanone-HCl), and antipain-HCl, and mixtures thereof. 10. The tube of claim 1, wherein said lyophilized protease inhibitors comprise at least one inhibitor of a serine protease, at least one inhibitor of a cysteine protease, and at least one of an inhibitor of a cysteine/serine protease, an inhibitor of an aspartic protease, an inhibitor of a metalloprotease, and an inhibitor of an exopeptidase. 11. The tube of claim 10, wherein said inhibitor of an aspartic protease is selected from the group consisting of calpain, N-acetyl-leu-leu-norleucinal, N-acetyl-leu-leu-methioninal, pepstatin and VdLPFFVdL, and mixtures thereof. 12. The tube of claim 10, wherein said inhibitor of a metalloprotease is selected from the group consisting of EDTA, 1,10-phenanthroline, phosphoramodon and mixtures thereof. 13. The tube of claim 10, wherein said inhibitor of an exopeptidase is selected from the group consisting of amastatin, bestatin, diprotin A and diprotin B, and mixtures thereof. 14. The tube of claim 1, wherein said at least two lyophilized protease inhibitors comprise AEBSF in a concentration of less than 1 mM at whole blood volume of the collected blood or component thereof. 15. The tube of claim 1, further comprising an anticoagulant.... | Becton, Dickinson and Company | B01L003/00
US08475388 | B2 | 2013-07-02 | Detection of disordered breathing | 2010-07-30 | Devices and methods for detecting disordered breathing involve determining that the patient is asleep and sensing one or more signals associated with disordered breathing indicative of sleep-disordered breathing while the patient is asleep. Sleep-disordered breathing is detected using the sensed signals associated with disordered breathing. The sensed signals associated with disordered breathing may also be used to acquire a respiration pattern of one or more respiration cycles. Characteristics of the respiration pattern are determined. The respiration pattern is classified as a disordered breathing episode based on the characteristics of the respiration pattern. One or more processes involved in the detection of disordered breathing are performed using an implantable device. | 1. A method of operating a medical device, comprising: sensing respiration of a patient and generating a respiration signal;determining respiration signal characteristics from the respiration signal;comparing the respiration characteristics to respiration threshold values;sensing non-respiration related artifacts and generating an artifact signal;sensing two or more sleep related signals;comparing at least one of the sensed sleep related signals to a first sleep threshold, the first sleep threshold adjusted by another of the one or more sensed sleep related signals;determining a sleep status of the patient based at least in part on the sleep threshold comparison;detecting disordered breathing including discriminating between apnea, hypopnea, and episodes of mixed apnea and hypopnea, using a processor, based on the sleep status and comparing the respiration signal characteristics to the respiration threshold values;reducing erroneous detection of the disordered breathing based on the artifact signal; andstoring historical data related to the disordered breathing in memory of the medical device. 2. The method of claim 1, wherein the sleep related signals comprise an accelerometer signal and the respiration signal. 3. The method of claim 1, wherein detecting the disordered breathing comprises detecting Chine-Stokes respiration. 4. The method of claim 1, wherein the non-respiration related artifacts comprise one or both of motion and posture related artifacts. 5. The method of claim 1, wherein the first sleep threshold is a patient activity threshold and the first sleep threshold is adjusted using a minute ventilation signal. 6. The method of claim 1, wherein storing the historical data related to the disordered breathing comprises storing one or more of a duration, severity, and frequency of the disordered breathing episodes. 7. The method of claim 1, wherein storing the historical data related to the disordered breathing comprises storing respiratory signals of the disordered breathing episode. 8. A method of operating a medical device, comprising: sensing respiration of a patient and generating a respiration signal;determining tidal volumes and breath intervals from the respiration signal;comparing the tidal volumes and the breath intervals to respiration indices, including comparing the tidal volumes to a hypopnea index and comparing the breath intervals to an apnea index;sensing non-respiration related artifacts and generating an artifact signal;sensing one or more sleep related signals;comparing at least one sleep related signal to a first sleep threshold;determining a sleep status of the patient based at least in part on comparison of the sleep related signal to the first sleep threshold;detecting disordered breathing based on the sleep status and on comparing the tidal volumes and the breath intervals to the hypopnea and apnea indices; andreducing erroneous detection of the disordered breathing based on the artifact signal, wherein det... | Cardiac Pacemakers, Inc. | A61B005/08
US08596277 | B2 | 2013-12-03 | Tracheal tube with lumen for tracheal pressure measurement and technique for using the same | 2009-06-18 | According to various embodiments, a tracheal tube may include a pressure monitoring lumen configured to sample the gases in the tracheal space. The pressure monitoring lumen may be in fluid communication with a pressure transducer that provides pressure measurements. An opening of the lumen may be covered with a porous membrane that allow gas to diffuse freely in and out of the lumen, but that prevents secretions from clogging the lumen. | 1. A system comprising: a tracheal tube comprising a distal end and a proximal end, wherein the tube is capable of transferring a fluid to a patient's lungs;an inflatable cuff associated with the tracheal tube, the inflatable cuff comprising a distal shoulder and a proximal shoulder;a lumen disposed in a wall of the tracheal tube, wherein the lumen comprises an opening in the wall of the tracheal tube between the distal shoulder of the cuff and the distal end of the tracheal tube; anda porous membrane covering the opening, wherein the porous membrane is substantially permeable to the fluid; anda monitor comprising: a memory circuit storing data representative of a relationship between a secretion buildup on the porous membrane and a tracheal pressure value, wherein the data comprises a correction factor representative of an estimated secretion buildup on the porous membrane; anda processor configured to adjust the tracheal pressure value to account for the secretion buildup based on the correction factor. 2. The system of claim 1, wherein the memory circuit stores data representative of a relationship between a pressure in the lumen and a the tracheal pressure value. 3. The system of claim 1, wherein the memory circuit is associated with a connection or cable on the tracheal tube. 4. The system of claim 1, comprising a pressure transducer in fluid communication with the lumen. 5. The system of claim 1, wherein the porous membrane comprises a hydrophobic material. 6. The system of claim 5, wherein the hydrophobic material comprises a material with a water contact angle greater than about 70 degrees. 7. The system of claim 1, wherein the porous membrane comprises a material with a break strength greater than about 60 cm H20. 8. The system of claim 1, wherein the porous membrane comprises a material with a pore size greater than about 0.4 micrometers. 9. The system of claim 1, wherein the porous membrane comprises one or more of polyvinylidene fluoride, poly(ethyleneterephthalate), or polytetrafluoroethylene. 10. The system of claim 1, wherein the porous membrane encircles at least a portion of the tracheal tube. 11. A system for determining trachea pressure comprising: a tracheal tube comprising a distal end and a proximal end and configured to transfer a fluid to a patient's lungs; a lumen disposed in a wall of the tracheal tube, wherein the lumen comprises an opening in the wall of the tracheal tube adjacent to the distal end; anda porous membrane covering the opening, wherein the porous membrane is substantially permeable to the fluid;a pressure transducer in fluid communication with the lumen; anda monitor comprising: a memory circuit storing instructions to determine a corrected trachea pressure based upon a signal from the pressure value transducer and a correction factor representative of an estimated secretion buildup on the porous membrane; anda processor configured to execute the instructions. 12. The system of claim 11, compri... | Covidien LP | A61M016/00
US08620615 | B2 | 2013-12-31 | Apparatuses and methods for evaluating a person for a sleep system | 2008-11-14 | A method for evaluating a person for a sleep system, the method including: while the person is not positioned on an evaluating member, adjusting a pressure of a comfort layer inflatable member disposed within a comfort layer of the evaluating member to an initial comfort value; positioning the person on the evaluating member in a first position; while the person is positioned on the evaluating member in the first position, measuring a pressure of the comfort layer inflatable member as a first measured comfort value; calculating a difference between the first measured comfort value and the initial comfort value as comfort pressure 1; calculating a first optimal pressure level for the comfort layer inflatable member using comfort pressure 1; and recommending a sleep support member for the person using the calculated first optimal pressure level for the comfort layer inflatable member and using data measuring quality of sleep. | 1. A method for evaluating a person for a sleep system, the method comprising: while the person is not positioned on an evaluating member, adjusting a pressure of a comfort layer inflatable member disposed within a comfort layer of the evaluating member to an initial comfort value;positioning the person on the evaluating member in a first position;while the person is positioned on the evaluating member in the first position, measuring a pressure of the comfort layer inflatable member as a first measured comfort value;calculating a difference between the first measured comfort value and the initial comfort value as ΔCOMFORT PRESSURE 1;calculating, using at least one processor, a first optimal pressure level for the comfort layer inflatable member using ΔCOMFORT PRESSURE 1; andrecommending, using the at least one processor, a sleep support member for the person using both the calculated first optimal pressure level for the comfort layer inflatable member and data measuring quality of sleep. 2. The method of claim 1, further comprising adjusting the comfort layer inflatable member to the calculated first optimal pressure level for the comfort layer inflatable member. 3. The method of claim 1, further comprising: while the person is not positioned on the evaluating member, adjusting a pressure of a support layer inflatable member disposed within a support layer of the evaluating member to an initial support value;while the person is positioned on the evaluating member in the first position, measuring a pressure of the support layer inflatable member as a first measured support value;calculating a difference between the first measured support value and the initial support value as ΔSUPPORT PRESSURE 1;calculating a first optimal pressure level for the support layer inflatable member using ΔSUPPORT PRESSURE 1; andrecommending the sleep support member for the person using the calculated first optimal pressure level for the support layer inflatable member. 4. The method of claim 3, further comprising adjusting the support layer inflatable member to the calculated first optimal pressure level for the support layer inflatable member. 5. The method of claim 1, further comprising: while the person is not positioned on the evaluating member, measuring a pressure at a first region of a foundation layer of the evaluating member as an initial foundation value;while the person is positioned on the evaluating member in the first position, measuring a pressure at the first region of the foundation layer as a first measured foundation value;calculating a difference between the first measured foundation value and the initial foundation value as ΔFOUNDATION PRESSURE 1;calculating a first optimal pressure level for the first region of the foundation layer using ΔFOUNDATION PRESSURE 1; andrecommending the sleep support member for the person using the calculated first optimal pressure level for the foundation layer. 6. The method of claim 3, wherein the support layer... | Kingsdown, Inc. | G01L017/00
US08689377 | B2 | 2014-04-08 | Radar-equipped patient bed for a medical imaging apparatus, and operating method therefor | 2010-05-14 | A patient bed for an imaging medical apparatus has a patient support plate that has at least one radar antenna to obtain physiological and/or geometric data from a patient the patient support plate. In a method for the operation of such a patient bed having at least one radar antenna in an imaging medical apparatus, the at least one radar antenna is operated to obtain data from the patient on the patient support plate. | 1. A patient bed for a medical imaging apparatus, comprising: a patient support plate configured to receive a patient thereon;at least one radar antenna mounted at said patient support plate and configured to operate at an operating frequency to radiate a primary signal to said patient on said patient support plate and to receive a secondary signal reflected from the patient on the patient support plate;said at least one radar antenna comprising a metal surface or metal layer having a thickness corresponding to a skin depth of the metal defined at said operating frequency; anda control and evaluation unit configured to operate said at least one radar antenna to radiate said primary signal and to receive said secondary signal, and configured to generate, from said secondary signal, at least one of physiological data and geometric data from the patient on the patient support plate. 2. A patient bed as claimed in claim 1 wherein said at least one radar antenna is a patch antenna. 3. A patient bed as claimed in claim 2 wherein said patch antenna has a layer thickness in a range between 2 μm and 20 μm. 4. A patient bed as claimed in claim 1 comprising a radar antenna array mounted at said patient support plate. 5. A patient bed as claimed in claim 1 comprising at least two radar antennas mounted at said patient support plate that are located adjacent to each other. 6. A patient bed as claimed in claim 1 wherein said patient support plate comprises a flat patient support plate component, selected from the group consisting of a film and a substrate layer, at which said at least one radar antenna is mounted. 7. A patient bed as claimed in claim 6 wherein said patient support plate component is comprised of circuit board material. 8. A patient bed as claimed in claim 6 wherein said patient support plate component comprises at least one electrically conductive coating, in addition to said at least one radar antenna. 9. A patient bed as claimed in claim 8 wherein said electrically conductive coating is a continuous coating on a side of said patient support plate component facing away from the patient. 10. A patient bed as claimed in claim 8 wherein said electrically conductive coating is a metallic coating. 11. A patient bed as claimed in claim 8 wherein said electrically conductive coating has a thickness in a range between 2 μm and 20 μm. 12. A patient bed as claimed in claim 6 wherein said patient support plate component has a thickness in a range between 0.5 mm and 2.5 mm. 13. A patient bed as claimed in claim 6 wherein said patient support plate component is integrated into said patient support plate. 14. A patient bed as claimed in claim 6 wherein said patient support plate component is applied at a side of said patient support plate facing toward the patient. 15. A patient bed as claimed in claim 1 comprising a base on which said patient support plate is mounted, and wherein said control and evaluation unit is located in said base. 16.... | Siemens Aktiengesellschaft | A47B013/00 A61B005/05
US08388542 | B2 | 2013-03-05 | System for cardiac pathology detection and characterization | 2010-04-28 | A system for heart performance characterization and abnormality detection includes an interface for receiving an electrical signal comprising a pressure indicative waveform indicating a heart blood pressure of a patient over a heart beat cycle. A timing detector determines multiple different time periods in at least one heart cycle from the pressure indicative waveform. A patient monitor monitors the multiple different time periods and in response to detection of a variation in at least one of the multiple different time periods exceeding a predetermined threshold or range, generates an alert message associated with the variation. | 1. A system for heart performance characterization and abnormality detection, comprising: an interface for receiving an electrical signal comprising a pressure indicative waveform indicating a heart blood pressure of a patient over a heart beat cycle;a timing detector for determining a plurality of different time periods in at least one heart cycle of the pressure indicative waveform, said plurality of different time periods comprising a time interval T1 substantially from first to second successive peaks in the pressure indicative waveform, and a further time interval T4 from a pressure peak to a successive end of Systolic point; anda patient monitor for monitoring said further time interval T4 and in response to detection of a variation in said further time interval exceeding a predetermined threshold or range, generating an alert message associated with said variation. 2. A system according to claim 1, wherein said plurality of different time periods include at least one of (a) a time interval T2 substantially comprising a duration of a Systolic period and(b) a time interval T3 substantially comprising a duration of a Diastolic period andsaid patient monitor derives and monitors a plurality of different ratios using said plurality of different time periods comprising at least one of, andT1T2+T3⁢⁢and(a)T4T2.(b) 3. A system according to claim 2, wherein said patient monitor substantially continuously monitors said plurality of different ratios for at least a 24 hour period. 4. A system according to claim 2, wherein said patient monitor uses predetermined mapping information, associating ranges of said different ratios with corresponding medical conditions, in comparing a ratio with said ranges and generates an alert message indicating a potential medical condition. 5. A system according to claim 4, wherein said predetermined mapping information associates ranges of said different ratios with particular patient demographic characteristics and with corresponding medical conditions and said patient monitor uses patient demographic data including at least one of, age, weight, gender and height in comparing a ratio with said ranges and generating an alert message indicating a potential medical condition. 6. A system according to claim 2, wherein said timing detector determines a time interval T5 substantially from a pressure minimum point to a successive end of Diastolic point andsaid plurality of different ratios includes at least one of,T2-T4T2,(i)T5T3⁢⁢and(ii)T3-T5T3.(iii) 7. A system according to claim 1, wherein said patient monitor substantially continuously monitors said plurality of different time periods for at least a 24 hour period. 8. A system according to claim 1, wherein said electrical signal comprises at least one of, (a) digital data representing said pressure indicative waveform and (b) an analog pressure indicative waveform andsaid electrical signal is acquired at a particular anatomical location. 9. A system according to c... | Siemens Medical Solutions USA, Inc. | A61B005/02
US07985188 | B2 | 2011-07-26 | Vessel, coating, inspection and processing apparatus | 2010-05-12 | Methods for processing a vessel, for example to provide a gas barrier or lubricity, are disclosed. First and second PECVD or other vessel processing stations or devices and a vessel holder comprising a vessel port are provided. An opening of the vessel can be seated on the vessel port. The interior surface of the seated vessel can be processed via the vessel port by the first and second processing stations or devices. Vessel barrier, lubricity and hydrophobic coatings and coated vessels, for example syringes and medical sample collection tubes are disclosed. A vessel processing system and vessel inspection apparatus and methods are also disclosed. | 1. A syringe or syringe parts comprising: a barrel and a piston, the barrel having an inner surface and the piston having a side surface engaging the barrel inner surface, wherein at least one of the inner surface and the side surface is a substrate coated with a lubricity layer configured to provide a lower piston sliding force or breakout force than the uncoated substrate the lubricity layer having one of the following atomic ratios, measured by X-ray photoelectron spectroscopy (XPS), SiwOxCy or SiwNxCy, where w is 1, x in this formula is from about 0.5 to 2.4, and y is from about 0.6 to about 3;the lubricity layer having a thickness by transmission electron microscopy (TEM) between 10 and 500 nm;the lubricity layer deposited by plasma enhanced chemical vapor deposition (PECVD) under conditions effective to form a coating from a precursor selected from a linear siloxane, a monocyclic siloxane, a polycyclic siloxane, a polysilsesquioxane, a linear silazane, a monocyclic silazane, a polycyclic silazane, a polysilsesquiazane, a silatrane, a silquasilatrane, a silproatrane, an azasilatrane, an azasilquasiatrane, an azasilproatrane, or a combination of any two or more of these precursors. 2. The syringe or syringe parts of claim 1, in which the precursor comprises hexamethyldisiloxane, octamethyltrisiloxane, decamethyltetrasiloxane, hexamethylcyclotrisiloxane, octamethylcyclotetrasiloxane, decamethylcyclopentasiloxane, dodecamethylcyclohexasiloxane, a poly(methylsilsesquioxane) according to the T8 cube formula:in which each R is methyl, a poly(Methyl-Hydridosilsesquioxane) according to the T8 cube formula in which 90% of the R groups are methyl and 10% are hydrogen atoms, methyltrimethoxysilane, or a combination of any two or more of these. 3. The syringe or syringe parts of claim 1, in which the precursor comprises octamethylcyclotetrasiloxane (OMCTS). 4. The syringe or syringe parts of claim 1, in which the precursor consists essentially of octamethylcyclotetrasiloxane (OMCTS). 5. The syringe or syringe parts of claim 1, in which the lubricity layer has a thickness of 10-200 nm. 6. The syringe or syringe parts of claim 1, in which the lubricity layer has a thickness of 20 to 100 nm. 7. The syringe or syringe parts of claim 1, in which the substrate comprises glass. 8. The syringe or syringe parts of claim 1, in which the substrate comprises a polymer. 9. The syringe or syringe parts of claim 1, in which the substrate comprises a polycarbonate polymer. 10. The syringe or syringe parts of claim 1, in which the substrate comprises an olefin polymer. 11. The syringe or syringe parts of claim 1, in which the substrate comprises a cyclic olefin copolymer. 12. The syringe or syringe parts of claim 1, in which the substrate comprises a polypropylene polymer. 13. The syringe or syringe parts of claim 1, in which the substrate comprises a polyester polymer. 14. The syringe or syringe parts of claim 1, in which the substrate comprises a polyethy... | CV Holdings LLC | A61B005/00 B65D081/00 C08F002/46 C08J007/18 H05H001/00 H05H001/24 C23C014/14 C23C014/16 B05D003/14 H01L021/469
US08239011 | B2 | 2012-08-07 | Atrial arrhythmia detection and discrimination based on intracardiac impedance | 2010-04-09 | An apparatus comprises an implantable impedance sensing circuit configured to sense an atrial impedance signal when coupled to a plurality of implantable electrodes, and an impedance signal analyzer circuit configured to detect a sudden change in a characteristic of the sensed atrial impedance signal that indicates atrial tachyarrhythmia. The impedance signal analyzer circuit classifies the atrial tachyarrhythmia indication as ST when the detected sudden change satisfies an ST threshold value of the characteristic, classifies the atrial tachyarrhythmia indication as AT when the detected sudden change satisfies an AT threshold value of the characteristic that is different from the ST threshold value, classifies the atrial tachyarrhythmia indication as AF when the detected sudden change satisfies an AF threshold value of the characteristic that is different from the ST and AT amplitude threshold values, and provides a classification of the tachyarrhythmia to a user or process. | 1. An apparatus comprising: an implantable impedance sensing circuit configured to sense an atrial impedance signal of a subject and adapted to be coupled to a plurality of implantable electrodes; andan impedance signal analyzer circuit, communicatively coupled to the implantable impedance sensing circuit, configured to: detect a sudden change, in a characteristic of the sensed atrial impedance signal, that indicates atrial tachyarrhythmia;classify the atrial tachyarrhythmia indication as sinus tachycardia (ST) when the detected sudden change satisfies an ST threshold value of the characteristic;classify the atrial tachyarrhythmia indication as atrial tachycardia (AT) when the detected sudden change satisfies an AT threshold value of the characteristic that is different from the ST threshold value;classify the atrial tachyarrhythmia indication as atrial fibrillation (AF) when the detected sudden change satisfies an AF threshold value of the characteristic that is different from the ST and AT amplitude threshold values; andprovide a classification of the tachyarrhythmia to a user or process. 2. The apparatus of claim 1, wherein the impedance signal analyzer circuit is configured to detect a sudden decrease in an amplitude of the sensed atrial impedance signal that is less than an amplitude threshold value, and classify the indication as ST, AT, or AF according to the detected decrease. 3. The apparatus of claim 2, wherein the impedance signal analyzer circuit is configured to detect at least one of: a sudden decrease in a calculated central tendency of the amplitude of the sensed atrial impedance signal;a sudden decrease in a baseline-to-peak amplitude of the atrial impedance signal; anda sudden decrease in a peak-to-peak amplitude of the atrial impedance signal. 4. The apparatus of claim 1, wherein the impedance signal analyzer circuit is configured to detect a sudden change in an atrial depolarization rate or interval in the sensed atrial impedance signal that satisfies a specified rate or interval threshold value. 5. The apparatus of claim 4, including: an implantable cardiac signal sensing circuit, communicatively coupled to the impedance signal analyzer circuit, configured to sense a cardiac signal associated with a cardiac action potential signal of the subject when coupled to a plurality of implantable electrodes, andwherein the impedance signal analyzer circuit is configured to detect the sudden change in the atrial depolarization rate or interval using both of the sensed atrial impedance signal and a sensed atrial action potential signal. 6. The apparatus of claim 1, wherein the impedance signal analyzer circuit is configured to detect a sudden increase in variation of atrial depolarization rate or interval measured using the sensed atrial impedance signal. 7. The apparatus of claim 1, wherein the impedance signal analyzer circuit is configured to detect a sudden decrease in a measure of regularity of the sensed atrial impedance sig... | Cardiac Pacemakers, Inc. | A61B005/02
US08282829 | B2 | 2012-10-09 | System and method for automated data collection of twenty-four hour ultrafiltration and other patient parameters using wired or wireless technology | 2009-05-20 | A dialysis system includes: a dialysis machine for performing a first dialysis session on a patient at a first time of a day; a remote exchange device for performing a second dialysis session on the patient at a second time of the day, the remote exchange device configured to record an amount of ultrafiltration (“UF”) that is removed from the patient over the second dialysis session; and a communication link between the dialysis machine and the remote exchange device, wherein the dialysis machine is configured to receive the recorded amount UF removed from the remote exchange device via the communication link and determine a total amount of UF removed from the patient over the first and second sessions. | 1. A dialysis system comprising: a dialysis machine for performing a first dialysis session on a patient at a first time of a day at a first location;a remote dialysis fluid exchange device for performing a second dialysis session on the patient at a second time of the day at a second location remote from the dialysis machine, the remote exchange device configured to record an amount of ultrafiltration (“UF”) that is removed from the patient over the second dialysis session; anda communication link between the dialysis machine and the remote exchange device,wherein the dialysis machine is configured to receive the recorded amount of UF removed from the remote dialysis fluid exchange device at the first location via the communication link and determine a total amount of UF removed from the patient over the first and second sessions. 2. The dialysis system of claim 1, wherein the communication link is wired or wireless. 3. The dialysis system of claim 1, wherein the dialysis machine is configured to record an amount of UF that is removed from the patient over the first dialysis session and add that amount to the amount of UF removed over the second dialysis session to determine the total amount of UF removed. 4. The dialysis system of claim 1, wherein the first dialysis session is a nighttime treatment and the second dialysis session is a day exchange. 5. The dialysis system of claim 1, wherein the communication link includes an intermediate data transfer unit, which is configured to receive the recorded amount of UF removed from the remote dialysis fluid exchange device. 6. The dialysis system of claim 5, wherein the intermediate data transfer unit and the remote dialysis fluid exchange device communicate (i) wirelessly or (ii) over a wire. 7. The dialysis system of claim 5, wherein at least one of: (i) the dialysis machine and the intermediate data transfer unit each include a memory storage member port, the memory storage member enabling the dialysis machine to receive the recorded amount of UF from the remote dialysis fluid exchange device; (ii) the dialysis machine and the intermediate data transfer unit are configured to communicate wirelessly so that the dialysis machine can receive the recorded amount of UF from the remote dialysis fluid exchange device; and (iii) the dialysis machine and the intermediate data transfer unit are configured to communicate via a wire so that the dialysis machine can receive the recorded amount of UF from the remote dialysis fluid exchange device. 8. The dialysis system of claim 5, wherein the intermediate data transfer unit is further configured to receive, for communication to the dialysis machine, at least one of: (i) blood pressure data, (ii) bioimpedance data, (iii) patient weight data, (iv) patient glucose level data, and (v) patient temperature data. 9. The dialysis system of claim 8, wherein at least one of: (i) each of the different data is sent from a different apparatus corresponding to the p... | Baxter International Inc. Baxter Healthcare S.A. | B01D061/34 B01D061/32
US08369936 | B2 | 2013-02-05 | Wearable apparatus for measuring heart-related parameters and deriving human status parameters from sensed physiological and contextual parameters | 2010-07-20 | A monitor device and associated methodology are disclosed which provide a self contained, relatively small and continuously wearable package for the monitoring of heart related parameters, including ECG. The detection of heart related parameters is predicated on the location of inequipotential signals located within regions of the human body conventionally defined as equivalent for the purpose of detection of heart related electrical activity, such as on single limbs. Amplification, filtering and processing methods and apparatus are described in conjunction with analytical tools for beat detection and display. | 1. An apparatus to determine physiological status of an individual, comprising: a. a wearable sensor device adapted to be worn in an equivalence region of said individual's body;b. at least two electrodes attached to said wearable sensor device, said electrodes adapted to be mounted within said equivalence region and to detect a heart-related signal; andc. a processor in electronic communication with said at least two electrodes and programmed to determine a heart-related status parameter from said heart-related signal, and to determine said individual's physiological status based on said heart-related signal. 2. The apparatus of claim 1, further comprising at least one sensor mounted to said wearable sensor device, said at least one sensor selected from the group consisting of a heat flux sensor, a galvanic skin response sensor, a skin temperature sensor, and an accelerometer. 3. The apparatus of claim 2, wherein said processor is further programmed to determine said individual's physiological status from data detected by said at least one sensor. 4. The apparatus of claim 1, wherein said individual's physiological status is at least one of fatigue, heat distress, dehydration, shock, and peripheral vascular shutdown. 5. The apparatus of claim 2, wherein said individual's physiological status is at least one of fatigue, heat distress, dehydration, hypothermia, and peripheral vascular shutdown. 6. The apparatus of claim 1, wherein said device is adapted to be in electronic communication with separate wearable sensor devices. 7. The apparatus of claim 6, wherein said processor is further programmed to determine said device's distance from at least one of said separate wearable sensor devices. 8. The apparatus of claim 6, wherein said devices communicates said physiological status to at least one of said separate wearable device. 9. The apparatus of claim 1, wherein said wearable sensor device is adapted to be worn on a left upper arm of said individual. 10. The apparatus of claim 2, wherein said processor is further programmed to determine said individual's contextual status from data detected by said at least one sensor. | BodyMedia, Inc. | A61B005/04
US08583204 | B2 | 2013-11-12 | Polymer membranes for continuous analyte sensors | 2010-03-05 | Devices and methods are described for providing continuous measurement of an analyte concentration. In some embodiments, the device has a sensing mechanism and a sensing membrane that includes at least one surface-active group-containing polymer and that is located over the sensing mechanism. The sensing membrane may have a bioprotective layer configured to substantially block the effect and/or influence of non-constant noise-causing species. | 1. A device for continuous measurement of an analyte concentration, the device comprising: a sensor configured to continuously measure a signal associated with an analyte concentration in a host; anda membrane located over the sensor, wherein the membrane comprises a polyurethane and a hydrophilic portion;wherein the device is configured to provide, at analyte concentrations of from about 40 mg/dL to about 400 mg/dL, a level of accuracy corresponding to a mean absolute relative difference of no more than 8% over a sensor session of at least about 3 days, wherein one or more reference measurements associated with calculation of the mean absolute relative difference are determined by analysis of blood. 2. The device of claim 1, wherein the sensor session is at least about 5 days. 3. The device of claim 1, wherein the sensor session is at least about 6 days. 4. The device of claim 1, wherein the sensor session is at least about 7 days. 5. The device of claim 1, wherein the sensor session is at least about 10 days. 6. The device of claim 1, wherein the membrane comprises a copolymer comprising a silicone segment. 7. The device of claim 1, wherein the mean absolute relative difference is no more than 7% over the sensor session. 8. The device of claim 1, wherein the membrane comprises an enzyme configured to react with the analyte. 9. The device of claim 1, wherein the membrane comprises a copolymer comprising a fluorocarbon segment. 10. The device of claim 1, wherein the membrane comprises a polycarbonate segment. 11. The sensor of claim 1, wherein the sensor is configured for transcutaneous implantation into a host. 12. The sensor of claim 1, wherein the sensor comprises a first portion and second portion, wherein the first portion is adapted for insertion under the host's skin and the second portion is adapted to reside ex vivo. 13. The sensor of claim 1, wherein the sensor is configured to reside in subcutaneous tissue. 14. The sensor of claim 1, wherein the sensor is configured to have a mean absolute difference of no more than 7 mg/dL at an analyte concentration range of from 40 mg/dL to 80 mg/dL. 15. The sensor of claim 1, wherein the host is diabetic. 16. The sensor of claim 1, wherein the sensor is configured to exhibit no more than an equivalent glucose signal of no more than about 30 mg/dL while immersed in a solution comprising about 3 mg/dL acetaminophen. 17. The sensor of claim 1, wherein the sensor is configured to exhibit no more than an equivalent glucose signal of no more than about 17 mg/dL while immersed in a solution comprising about 6 mg/dL ascorbic acid. 18. The sensor of claim 1, wherein the sensor is configured to exhibit no more than an equivalent glucose signal of no more than about 25 mg/dL while immersed in a solution comprising about 6 mg/dL uric acid. 19. A system for continuous measurement of an analyte concentration, the system comprising: a sensor comprising: a sensing region configured to continuousl... | DexCom, Inc. | A61B005/05 A61B005/00
US08463355 | B2 | 2013-06-11 | Apparatus to enable display of real-time graphical or numeric information within an MR image | 2009-06-09 | In a system or method to enable display of real-time graphical or numeric information within an MR image, at least one display segment having an MR visible substance therein is placed at least partially inside an imaging volume of the MR image. | 1. A system to enable display of real-time alphabetic and/or numeric changeable message information within an MR image of a volume of interest comprising at least a portion of a person, animal, or object being imaged, comprising: an imaging volume which defines a region in a magnetic field of an MR machine in which imaging planes are located for constructing the MR image and within which is located said volume of interest which said imaging planes intersect;at least one tube or frame display segment having an MR visible substance therein, said display segment being placed at least partially inside said imaging volume to provide said changeable message information during MR imaging; andsaid display segment having at least one respective coil on each of opposite sides of the display segment to create a Maxwell pair across the display segment inside the imaging volume. 2. The system of claim 1 wherein the MR visible substance is adapted to be resonant at or near a frequency of a nucleus being imaged. 3. The system of claim 1 wherein said display segment comprises a flat tube having said at least one respective coil on each opposite surface of the display segment. 4. A system to enable display of real-time graphical or numeric information within an MR image of a volume of interest comprising at least a portion of a person, animal, or object being imaged, comprising: an imaging volume which defines a region in a magnetic field of an MR machine in which imaging planes are located for constructing the MR image and within which is located said volume of interest which said imaging planes intersect;at least one tube or frame display segment having an MR visible substance therein, said display segment being placed at least partially inside said imaging volume;said display segment comprises a flat tube having at least one respective coil on each opposite surface of the display segment to create a Maxwell pair across the display segment inside the imaging volume; anda plurality of coils are placed on each of said opposite surfaces of the display segment so as to create a plurality of said Maxwell pairs across the display segment inside the imaging volume. 5. The system of claim 4 wherein when energized, the Maxwell pairs are adapted to create a plurality of large local field gradients along a length of the display segment preventing a coherent MR signal from the display segment during the imaging, the energized state representing an “OFF” state for the display segment, and an “ON” state being provided when the coils are de-energized. 6. The system of claim 3 wherein a longitudinal axis of the flat tube runs perpendicular to an imaging plane of the imaging volume. 7. A system to enable display of real-time graphical or numeric information within an MR image of a volume of interest comprising at least a portion of a person, animal, or object being imaged, comprising: an imaging volume which defines a region in a magnetic field of an MR machine in which i... | Siemens Medical Solutions USA, Inc. | A61B005/05
US08478538 | B2 | 2013-07-02 | Selection of signal regions for parameter extraction | 2009-05-07 | According to embodiments, techniques for extracting a signal parameter from a selected region of a generally repetitive signal are disclosed. A pulse oximetry system including a sensor or probe may be used to obtain an original photoplethysmograph (PPG) signal from a subject. A filter transformation may be applied to the original PPG signal to produce a baseline PPG signal. The baseline PPG signal may contain artifacts and/or noise, and a region of the baseline PPG signal suitable for extracting the signal parameter may be selected. A suitable region of the baseline PPG signal may be selected by applying one or more thresholds to the baseline PPG signal, where the values of the thresholds may be set based on derivative values, amplitude-based percentiles, and/or local minima and maxima of the baseline PPG signal. A portion of the original PPG signal corresponding to the selected region may be processed, and the signal parameter may be extracted from the processed region. In an embodiment, the signal parameter may correspond to the respiration rate of a patient. | 1. A method of extracting a signal parameter from a generally repetitive signal, comprising: obtaining, using a specifically programmed processor, an original photoplethysmograph (PPG) signal from a sensor;applying, using the specifically programmed processor, a filter transformation to the original PPG signal to substantially remove a pulse component from the original PPG signal and produce a baseline PPG signal;determining, using the specifically programmed processor, one or more regions of the baseline PPG signal by applying a threshold to the baseline PPG signal, wherein the values of the threshold depend, at least in part, on a set of signal amplitude values of the baseline PPG signal;selecting, using the specifically programmed processor, a region from the one or more regions, and processing a portion of the original PPG signal corresponding to the selected region; andextracting, using the specifically programmed processor, the signal parameter from the processed portion of the original PPG signal. 2. The method of claim 1, wherein the determining one or more regions of the baseline PPG signal further comprises: computing, processor, a set of derivative values of the baseline PPG signal;setting, using the specifically programmed processor, an upper threshold value and a lower threshold value based on the computed set of derivative values; anddetermining, using the specifically programmed processor, one or more regions of the baseline PPG signal by selecting regions of the baseline PPG signal having derivative values located within the lower threshold value and the upper threshold value. 3. The method of claim 1, wherein the determining one or more regions of the baseline PPG signal further comprises: computing, processor, a set of amplitude-based percentile values of the baseline PPG signal;setting, using the specifically programmed processor, an upper threshold value and a lower threshold value based on the computed set of amplitude-based percentile values; anddetermining, using the specifically programmed processor, one or more regions of the baseline PPG signal by selecting regions of the baseline PPG signal having amplitude-based percentile values located within the lower threshold value and the upper threshold value. 4. The method of claim 1, wherein the determining one or more regions of the baseline PPG signal further comprises: computing, using the specifically programmed processor, a set of local minima and local maxima of the baseline PPG signal;processing, using the specifically programmed processor, the set of local minima and local maxima to obtain a set of processed local minima and local maxima values;setting, using the specifically programmed processor, an upper threshold value and a lower threshold value based on the obtained set of processed local minima and local maxima values; anddetermining, using the specifically programmed processor, one or more regions of the baseline PPG signal by selecting regions of the baselin... | Nellcor Puritan Bennett Ireland | G06F007/00
US08478381 | B2 | 2013-07-02 | MRI guidewire | 2008-12-31 | A guidewire (100) for use with interventional magnetic resonance imaging has a guidewire body (102) having a distal end and a proximal end and reserving a space therein, a dipole antenna (108) disposed in the space reserved within the guidewire body, the dipole antenna being adapted to be electrically connected to a signal processing system through a first signal channel (110) through the proximal end of the guidewire body, and a loop antenna (112) disposed in the space reserved within the guidewire body toward the distal end of the guidewire body, the loop antenna (112) being adapted to be electrically connected to the signal processing system through a second signal channel (114) through the proximal end of the guidewire body. The dipole antenna and the loop antenna are each constructed to receive magnetic resonance imaging signals independently of each other and to transmit received signals through the first and second signal channels, respectively, to be received by the signal processing system. An interventional magnetic resonance imaging system includes an active guidewire. | 1. An interventional magnetic resonance imaging system, comprising: a magnetic resonance imaging scanner;an active guidewire disposed proximate said magnetic resonance imaging scanner, said active guidewire comprising a dipole antenna and a loop antenna positioned between a distal end and a proximal end of the active guidewire, wherein the loop antenna is positioned at the distal end of the active guidewire so as to visualize a tip of the guidewire and the dipole antenna is positioned between the proximal end and the loop antenna so as to visualize a shaft of the active guidewire; anda signal processing system in communication with said active guidewire to receive a first signal from said dipole antenna through a first signal channel and to receive a second signal through a second signal channel from said loop antenna. 2. An interventional magnetic resonance imaging system according to claim 1, wherein said signal processing system is constructed to receive signal information indicative of an extended portion of a shaft portion of said active guidewire through said first signal channel during operation of said magnetic resonance imaging scanner, andwherein said signal processing system is constructed to receive signal information indicative of a distal end of said active guidewire through said second signal channel during operation of said magnetic resonance imaging scanner. 3. An interventional magnetic resonance imaging system according to claim 2, further comprising a display system that is in communication with said signal processing system, wherein said display system is adapted to display a real-time image corresponding to said active guidewire during a surgical procedure. 4. An interventional magnetic resonance imaging system according to claim 2, wherein said active guidewire comprises: a guidewire body having a distal end and a proximal end and reserving a space therein;a dipole antenna disposed in said space reserved within said guidewire body, said dipole antenna being adapted to be electrically connected to said signal processing system through said first signal channel through said proximal end of said guidewire body; anda loop antenna disposed in said space reserved within said guidewire body toward said distal end of said guidewire body, said loop antenna being adapted to be electrically connected to said signal processing system through said second signal channel through said proximal end of said guidewire body,wherein said dipole antenna and said loop antenna are each constructed to receive magnetic resonance imaging signals independently of each other and to transmit received signals through said first and second signal channels, respectively, to be received by said signal processing system. 5. An interventional magnetic resonance imaging system according to claim 4, wherein said dipole antenna comprises: an outer hypotube disposed within said space reserved within said guidewire body;an inner hypotube disposed within a lume... | The United States of America, as Represented by the Secretary, Department of Health and Human Services | A61B005/055
US08172776 | B2 | 2012-05-08 | Systems and methods for detecting labor conditions via electromagnetic field disturbances | 2009-06-09 | A system may detect a labor condition in a laboring patient. The system may include an electromagnetic field generating unit, an electromagnetic field sensing unit, and a processing unit. The electromagnetic field generating unit may be operable to create an electromagnetic field about the laboring patient. The electromagnetic field sensing unit may be operable to detect a disturbance in the electromagnetic field. The processing unit may be operable to process the disturbance in the electromagnetic field to extract a labor condition. | 1. A system for detecting a labor condition in a laboring patient, the system comprising: an electromagnetic field generating unit operable to create an electromagnetic field about the laboring patient, the generated electromagnetic field being of relative specificity and sensitivity for patients of varying sizes throughout any stage of labor;an electromagnetic field sensing unit operable to detect a disturbance in the generated electromagnetic field due to a labor condition muscular contraction in an area through which the generated electromagnetic field propagates; anda processing unit operable to process the disturbance in the generated electromagnetic field due to the labor condition. 2. The system of claim 1, wherein the electromagnetic field generating unit comprises: one or more electromagnets; anda field generator operable to cause the electromagnets to produce an electromagnetic field. 3. The system of claim 2, wherein the field generator is operable to provide a current to the electromagnets. 4. The system of claim 2, wherein the field generator comprises a direct current source and a pulse width modulator. 5. The system of claim 2, wherein the electromagnets are positioned in one or more of the following: a labor bed, a labor mat, and a bed rail. 6. The system of claim 2, wherein the electromagnetic field generating unit further comprises an actuator operable to move or rotate at least one electromagnet. 7. The system of claim 6, wherein the processing unit is further operable to control the actuator in response to information provided by the electromagnetic field sensing unit. 8. The system of claim 1, wherein the electromagnetic field sensing unit comprises one or more electromagnetic sensors. 9. The system of claim 1, wherein the electromagnetic field sensing unit comprises: a belt device adapted for positioning about an abdomen of the laboring patient; andone or more electromagnetic sensors positioned within the belt device. 10. The system of claim 1, wherein the electromagnetic field sensing unit comprises: a catheter adapted for positioning in a bladder of the laboring patient; andone or more electromagnetic sensors associated with the catheter. 11. The system of claim 1, wherein the electromagnetic field sensing unit comprises: a probe adapted for positioning in a vagina or a uterus of the laboring patient; andone or more electromagnetic sensors associated with the probe. 12. The system of claim 1, wherein the processing unit is further operable to: receive electromagnetic field information from the electromagnetic field generating unit; andreceive disturbance information from the electromagnetic field sensing unit. 13. The system of claim 12, wherein the processing unit is further operable to process the electromagnetic field information and the disturbance information to identify a cause of the disturbance in the electromagnetic field. 14. The system of claim 13, wherein the processing unit is further operable to ... |  | A61B005/103
US08146422 | B2 | 2012-04-03 | High precision sensing for parameter measurement of the muscular-skeletal system | 2010-06-29 | A measurement system for capturing a transit time, phase, or frequency of energy waves propagating through a propagation medium is disclosed. The measurement system comprises a compressible waveguide (403), ultrasonic transducers (405, 406), and circuitry to sustain energy wave propagation in the waveguide (403). The circuitry includes a propagation tuned oscillator (404), a digital counter (409), a pulse generator (410), a phase detector (414), a counter (420), a digital timer (422), and a data register (424). The measurement system employs a continuous mode (CM), pulse mode, or pulse-echo mode of operation to evaluate propagation characteristics of continuous ultrasonic waves in the waveguide by way of closed-loop feedback to determine levels of applied forces on the waveguide. | 1. A force measurement system for the muscular-skeletal system comprising: a surface for coupling to the muscular-skeletal system;a compressible waveguide coupled to the surface;at least one ultrasonic transducer coupled to the compressible waveguide; andcircuitry to sustain continuous wave propagation in the waveguide during a measurement sequence where an integer number of energy waves is maintained in the waveguide such that a change in length of the waveguide results in a change in transit time, frequency, or phase of the energy waves. 2. The system of claim 1 where closed-loop feedback is applied to the circuitry during the measurement sequence. 3. The system of claim 2 where the circuitry includes a detection circuit in the closed-loop feedback path to detect propagated energy waves. 4. The system of claim 3 where the detection circuit is coupled to the at least one ultrasonic transducer and where the detection circuit triggers the at least one ultrasonic transducer to emit an energy wave into the waveguide upon detecting propagated energy waves. 5. The system of claim 1 further including a force sensor comprising: a first transducer coupled to a first location of the waveguide to emit energy waves into the waveguide; anda second transducer coupled to a second location of the waveguide to generate a signal corresponding to propagated energy waves. 6. The system of claim 1 where the measurement sequence comprises a predetermined number of energy waves propagating through the waveguide and an elapsed time of the measurement sequence. 7. The system of claim 6 further including a known relationship between waveguide length and a force applied thereto. 8. A force measurement system for the muscular-skeletal system comprising: a surface for coupling to the muscular-skeletal system;a compressible waveguide coupled to the surface;at least one ultrasonic transducer coupled to the compressible waveguide; andcircuitry to sustain pulsed wave propagation in the waveguide during a measurement sequence where an integer number of pulsed energy waves is maintained in the waveguide such that a change in length of the waveguide results in a change in transit time, frequency, or phase of the pulsed energy waves. 9. The system of claim 8 where closed-loop feedback is applied to the circuitry during the measurement sequence. 10. The system of claim 9 where the circuitry includes a detection circuit in the closed-loop feedback path to detect propagated pulsed energy waves, where the detection circuit is coupled to the at least one ultrasonic transducer, and where the detection circuit triggers the at least one ultrasonic transducer to emit a pulsed energy wave into the waveguide upon detecting a zero-crossing of propagated pulsed energy waves. 11. The system of claim 8 further including a force sensor comprising: a first transducer coupled to a first location of the waveguide to emit pulsed energy waves into the waveguide; anda second transducer coupled... |  | A61B005/22
US08827924 | B2 | 2014-09-09 | Diagnostic methods and systems based on urine analysis | 2010-02-23 | A diagnostic method and apparatus for detecting at least one change in a urinary parameter indicative of a body malfunction, the method comprising at least semi-continuously monitoring in real time at least one of a sodium level, an oxygen level, a potassium level, and combinations thereof in the urine of a catheterized patient; whereby at least one parameter is monitored so as to detect one or more changes in the at least one parameter to reflect at least one of a fluid state, an electrolyte balance, a kidney state, a kidney perfusion and an organ perfusion in the patient, indicative of the body malfunction in the patient, in which the monitoring is preferably performed through electrodes that are arranged perpendicularly to the flow of urine through a patient's catheter system. | 1. A monitoring apparatus for at least semi-continuous real-time in-line monitoring of urine from a patient, the patient having a catheter inserted therein, the catheter having a distal portion extending outside of the patient, the apparatus comprising: a catheter tube fluidly connected to said distal portion; an electrode system fluidly connected to said catheter tube, said electrode system comprising a receiving connector for fluidly connecting to said catheter tube, and a chamber fluidly connected to said receiving connector, said chamber comprising at least one glass electrode arranged perpendicularly to a flow of urine through said chamber, wherein said at least one electrode operates without reconditioning, wherein said catheter tube comprises a negative pressure tube for providing continuous flow of urine through said chamber, wherein said flow of urine flows directly from the patient through the catheter, through said catheter tube, into said receiving connector and into said chamber for in-line analysis without interruption. 2. The apparatus of claim 1, comprising a plurality of electrodes arranged perpendicularly to the flow of urine in said chamber. 3. The apparatus of claim 2, wherein said electrodes are arranged to measure one or more of sodium, potassium, oxygen, pH or a combination thereof. 4. The apparatus of claim 3, wherein said electrode system further comprises a measurement reader for determining a measurement of said one or more of sodium, potassium, oxygen, pH or a combination thereof. 5. The apparatus of claim 4, wherein said measurement reader further comprises a computational processor for processing said measurement. 6. The apparatus of claim 5, wherein said computation processor also processes said measurement according to a calibration reading. 7. The apparatus of claim 6, wherein said electrode system further comprises a color detector. 8. The apparatus of claim 7, wherein said electrode system further comprises a turbidity detector. 9. The apparatus of claim 8, wherein said electrode system further comprises a particle size detector. 10. A diagnostic method for detecting at least one change in a urinary parameter indicative of a body malfunction, the method comprising at least semi-continuously monitoring in real time at least one of a sodium level, an oxygen level, a potassium level, and combinations thereof in the urine of a catheterized patient; whereby at least one parameter is monitored so as to detect one or more changes in the at least one parameter to reflect at least one of a fluid state, an electrolyte balance, a kidney state, a kidney perfusion and an organ perfusion in the patient, indicative of the body malfunction in the patient, wherein the monitoring comprises arranging electrodes perpendicularly to and in the flow of urine through a patient's catheter system, in line to said system, wherein said electrodes comprise a glass electrode; receiving urine from the patient to said electrodes; and ... | FlowSense Ltd. | A61B005/00
US08696589 | B2 | 2014-04-15 | Method and system for heart failure status evaluation based on a disordered breathing index | 2010-07-30 | An evaluation of heart failure status is provided based on a disordered breathing index. Patient respiration is sensed and a respiration signal is generated. Disordered breathing episodes are detected based on the respiration signal. A disordered breathing index is determined based on the disordered breathing episodes. The disordered breathing index is trended and used to evaluate heart failure status. The disordered breathing index may be combined with additional information and/or may take into account patient activity, posture, sleep stage, or other patient information. | 1. A system for providing a patient assessment, the system comprising: a respiration sensor configured to generate a signal responsive to patient respiration; anda disordered breathing detector coupled to the respiration sensor and configured to detect disordered breathing episodes based on the respiration signal and discriminate between central disordered breathing episodes and obstructive disordered breathing episodes based on speed of transition to hyperventilation during the disordered breathing episodes; andan index processor configured to determine a ratio of the central disordered breathing episodes and the obstructive disordered breathing episodes and to determine a disordered breathing index using the ratio. 2. The system of claim 1, further comprising: a trend processor configured to trend the disordered breathing index; anda diagnostics unit configured to evaluate heart failure status based on the trend. 3. The system of claim 1, further comprising a rapid eye movement (REM) detector, wherein the index processor is configured to correlate the disordered breathing episodes with REM and non-REM sleep stages. 4. The system of claim 1, further comprising a posture sensor configure to provide an output indicative of tilt angle during the disordered breathing episodes, wherein the index processor is configured to determine the disordered breathing index including weighing the disordered breathing episodes based on tilt angle. 5. The system of claim 1, further comprising an activity sensor configured to generate a patient activity signal, wherein the index processor is configured to determine the disordered breathing index using the activity signal including applying different weights to disordered breathing episodes that occur during periods of low activity and disordered breathing episodes that occur during periods of high activity. 6. The system of claim 1, wherein the disordered breathing detector is configured to discriminate between the central disordered breathing episodes and the obstructive disordered breathing episodes based on hyperventilation portions of the respiration signal during the disordered breathing episodes. 7. The system of claim 1, further comprising a pressure sensor configured to generate a pressure signal, wherein the disordered breathing detector is configured to further discriminate between the central disordered breathing episodes and the obstructive disordered breathing episodes based on the pressure signal. 8. The system of claim 7, wherein the pressure sensor is configured to sense intrathoracic pressure, intracardiac pressure or intravascular pressure. | Cardiac Pacemakers, Inc. | A61B005/08
US08579815 | B2 | 2013-11-12 | Device, method, and graphical user interface for searching, filtering and displaying event-aligned information sets related to diabetes | 2009-06-10 | A device, method, and graphical user interface for displaying diabetes related information sets aligned by the occurrence of a particular event are disclosed. After receiving a request for an event-aligned display of diabetes related information based on an event, a search for a plurality of tagged occurrences of the event and retrieval of an information set for each of the plurality of tagged occurrences is provided. Each information set may include diabetes related information chronologically related to the tagged occurrence. The retrieved information sets is displayed on a display, wherein the retrieved information sets are shown on the display aligned by their tagged occurrences such that all diabetes related information is positioned relative the event based on its chronological relationship with the event. | 1. A method for displaying diabetes related information comprising: tagging at least manually with user defined identifying information in a diary database occurrences of various events, which repeat at least sporadically, to correlate the tagged occurrences of the various events with diary records such that a plurality of tagged occurrences are defined;receiving by a processor, via an input terminal, a request for an event-aligned display of diabetes related information based on an event from the various events, the request being at least an exact word or categorical title;searching, via the processor, the plurality of tagged occurrences in the dairy database for tagged occurrences of the event matching the request;retrieving, via the processor, an information set for each of the plurality of tagged occurrences from the diary database that match the request, wherein each information set comprises diabetes related information chronologically related to the tagged occurrence; anddisplaying the retrieved information sets on a display, wherein the retrieved information sets are shown by the processor on the display aligned by their tagged occurrences such that all diabetes related information are positioned relative the event based on their chronological relationship with the event. 2. The method of claim 1 wherein tagged occurrences comprise exercise, illness, stress, menstrual state, meal, ketosis, and/or therapeutic changes. 3. The method of claim 1, further comprising tagging occurrences automatically based on the diabetes related information or external conditions. 4. The method of claim 3 wherein the diabetes related information comprises glucose levels and events of hypoglycemia or hyperglycemia are automatically tagged when a glucose level drops below or rises above a threshold value. 5. The method of claim 1 wherein each information set may be selectively removed from the display. 6. The method of claim 1 wherein searching for a plurality of tagged occurrences comprises searching backwards in time from a most recent tagged occurrence. 7. The method of claim 6 wherein the plurality of tagged occurrences comprises between four and seven tagged occurrences. 8. The method of claim 1 wherein the diabetes related information comprises glucose levels. 9. The method of claim 8 wherein the glucose levels are obtained via a continuous blood glucose meter. 10. The method of claim 1 wherein displaying the information sets is performed on a touch screen display providing a graphical user interface which also receives the request. 11. The method of claim 10 wherein the display may be scrolled forward or backwards in time. 12. The method of claim 10 wherein each of the plurality of information sets may selectively be displayed through engaging the touch screen display. 13. The method of claim 1 further comprising tagging the tagged occurrences with additional categorical information such that a variety of search requests would find the same ev... | Roche Diagnostics Operations, Inc. | A61B005/00
US08391939 | B2 | 2013-03-05 | Metering glucose level in pulsing blood | 2007-07-18 | Method for the continuous measurement of the glucose concentration in blood undergoing pulsational flow, with the steps:—determination of a value for the glucose concentration for a first measurement cycle, and—repetition of the determination of this value in subsequent measurement cycles, where there is multiple detection, within each measurement cycle, of the transmittance and/or scattering power of the blood for at least two incident NIR wavelengths, calculation of an indicator value depending on the blood glucose concentration, and ascertaining the blood glucose concentration by comparing the indicator value with a previously determined calibration table, determination of the blood temperature during the detection of the transmittance and/or scattering power,—continuous measurement of the pulse duration of the pulsational blood flow, where the duration of the measurement cycle is arranged to keep in step as integral multiple of the pulse duration, where the first of the at least two NIR wavelengths is selected from the wavelength range 1560-1630 nm, and the second of the at least two NIR wavelengths is selected from the wavelength range 790-815 nm, and the ratio of the transmittance and/or scattering power of the at least two wavelengths is calculated, this ratio serving in relation to the blood temperature as indicator value for reading off the blood glucose concentration from the calibration table. | 1. A method for the continuous measurement of glucose level in blood undergoing pulsing flow, the method comprising the steps of: determining a value for the glucose level in succeeding and nonoverlapping time windows of the same length of a first measurement cycle;repeating the determination of this value in subsequent measurement cycles with the same number of nonoverlapping time windows of the same length such that there is multiple detection within each measurement cycle of the transmission or scattering power of the blood for at least two incident NIR wavelengths, calculation of an indicator value dependent on the blood-glucose level and ascertaining the blood-glucose level by comparing the indicator value with a previously determined calibration table;determining ensemble averages for all the time windows of each measurement cycle;determining the blood temperature during the detection of the transmission and/or scattering power,continuously measuring the pulse duration of the pulsing blood flow, the duration of the measurement cycle being set to keep in step as integral multiple of the pulse duration,selecting the first of the two NIR wavelengths from the wavelength range 1560-1630 nm, and the second of the NIR wavelengths from the wavelength range 790-815 nm, andusing the ratio of the transmission or scattering power of the two wavelengths relative to the blood temperature as an indicator value for reading off the blood-glucose level from the calibration table. 2. The method according to claim 1 wherein the at least two NIR wavelengths are irradiated in an amplitude-modulated manner with modulation frequencies above 1 MHZ. 3. The method according to claim 1 wherein each time window comprises less than 100 milliseconds. 4. The method according to claim 1 wherein the number of the measured values for transmission and/or scattering power of the blood in each time window is at least 10,000. 5. The method according to claim 1 wherein the average measured values in the determined time window are subjected to a Fourier transformation in the frequency range. 6. The method according to claim 5, further comprising the step of:evaluating the Fourier components of the average measured values for modulation frequency of the amplitude modulation of the irradiated NIR wavelengths as a gauge for the transmission or scattering power of the blood. 7. The method according to claim 6 wherein an integral in the frequency range over the Fourier components of the average measured values at an interval around the modulation frequency of the amplitude modulation of the irradiated NIR wavelengths is evaluated as a gauge of the transmission or scattering power of the blood. 8. An apparatus for carrying out the method according to claim 1 wherein the blood is taken from a living organism in a circulation, measured and returned, the apparatus further comprising: a measuring cell embodied for the continuous inflow and outflow of blood with a trans-illumination z... | Nirlus Engineering AG | A61B005/00 G01N033/48
US08702617 | B2 | 2014-04-22 | Simultaneous ambulatory pulse oximetry and pH monitoring for the diagnosis of GERD-related respiratory disease | 2010-04-27 | Respiratory diseases related to gastroesophageal reflux disease in a patient are diagnosed by detecting esophageal or pharyngeal impedance and pH in the patient, detecting oxygen saturation in the patient, and diagnosing the respiratory diseases from the esophageal or pharyngeal impedance, the pH, and the oxygen saturation. | 1. A method for diagnosing respiratory diseases related to gastroesophageal reflux disease in a patient, the method comprising: (a) providing a catheter having sensors for detecting esophageal or pharyngeal impedance, pH, and oxygen saturation in the patient;(b) inserting the catheter into an esophagus or pharynx of the patient;(c) detecting the esophageal or pharyngeal impedance and the pH in the patient, using the sensors;(d) detecting the oxygen saturation in the patient, using the sensors;(e) importing the esophageal or pharyngeal impedance and the pH detected in step (c) and the oxygen saturation detected in step (d) into a computing device; and(f) processing the esophageal or pharyngeal impedance and the pH detected in step (c) and the oxygen saturation detected in step (d) in the computing device to determine a time of a reflux event and a time of an oxygen desaturation event, to determine a reflux-desaturation association if the oxygen desaturation event occurs within a given time period of the reflux event, and to diagnose the respiratory diseases from the reflux-desaturation association. 2. The method of claim 1, wherein the sensors comprise a plurality of impedance sensors, and wherein step (c) comprises detecting the impedance using the plurality of impedance sensors. 3. The method of claim 2, wherein step (f) comprises determining that the patient is experiencing the reflux event if either the pH is below a threshold or the reflux event is detected in two of the plurality of impedance sensors. 4. The method of claim 1, wherein step (f) comprises detecting an abnormal multichannel intraluminal impedance-pH study if step (f) results in a determination that: (i) a DeMeester score exceeds a threshold;(ii) a number of acidic reflux episodes exceeding a predetermined number are detected; or(iii) an alkaline reflux episode is detected. 5. The method of claim 1, wherein step (f) comprises: (i) determining SpO2; and(ii) detecting the oxygen desaturation event if either (A) the SpO2 is below a first threshold or (B) the SpO2 drops by more than a second threshold. 6. The method of claim 1, wherein: the sensors comprise at least one pH sensor, a plurality of impedance sensors, and at least one oxygen sensor;step (c) is performed using the at least one pH sensor and the plurality of impedance sensors; andstep (d) is performed using the at least one oxygen sensor. 7. The method of claim 6, wherein the at least one pH sensor is an antimony type pH sensor. 8. A system for diagnosing respiratory diseases related to gastroesophageal reflux disease in a patient, the system comprising: a catheter configured to be insertable into an esophagus or pharynx of the patient;a sensor, disposed on the catheter, for detecting esophageal or pharyngeal impedance in the patient;a sensor, disposed on the catheter, for detecting pH in the patient;a sensor, disposed on the catheter, for detecting oxygen saturation in the patient; anda computing device, in commun... | University of Rochester | A61B005/14551 A61B005/0205 A61B005/021 A61B005/085
US07991452 | B2 | 2011-08-02 | Contrast enhanced MRA with highly constrained backprojection reconstruction using phase contrast composite image | 2007-12-28 | Mask projection views are obtained prior to the arrival of a contrast agent during a dynamic contrast enhanced MRA study. After the arrival of the contrast agent, a set of undersampled contrast enhanced projection views are obtained for each of a plurality of time frames. Corresponding mask projection views are subtracted from the contrast enhanced projection views to provide sparse contrast enhanced projection view sets. A phase contrast scan of a region of interest is performed prior to or after the arrival of the contrast agent. The phase contrast image is used as a composite image in a HYPR reconstruction of the sparse projection view sets to produce first pass contrast enhanced images. Iterative HYPR reconstructions can also be performed to remove venous information from the reconstructed images. | 1. A method for producing a contrast enhanced image of a subject positioned in the field of view (FOV) of a magnetic resonance imaging (MRI) system, the steps comprising: a) acquiring with the MRI system a set of phase contrast views of the subject positioned in the FOV;b) reconstructing a phase contrast composite image from the phase contrast views which indicates a motion value at each composite image pixel of the subject positioned in the FOV;c) acquiring with the MRI system a set of mask views;d) acquiring with the MRI system a time frame data set comprised of an undersampled set of contrast enhanced views of the subject positioned in the FOV after the arrival of the contrast agent;e) repeating step d) to acquire a plurality of time frame data sets during the in flow of contrast agent into the FOV;f) subtracting mask views from the contrast enhanced views to obtain sparse contrast enhanced views for each time frame; andg) performing a highly constrained projection (HYPR) reconstruction to produce a contrast enhanced image of the subject for each time frame using both the sparse contrast enhanced views for the time frame and the phase contrast composite image. 2. The method of claim 1 in which the views acquired in steps c) and d) are projection views and wherein step g) includes: g)i) backprojecting each sparse contrast enhanced projection view in the set into the FOV and weighting the value backprojected into each image pixel by the value of the corresponding pixel in the phase contrast composite image; andg)ii) summing the backprojected values for each image pixel. 3. The method as recited in claim 2 in which step e) is performed prior to performing the HYPR reconstruction in step g). 4. The method as recited in claim 1 which includes: i) producing an updated composite image for each time frame using the reconstructed contrast enhanced image for that time frame; andj) performing another HYPR reconstruction to produce another contrast enhanced image of the subject for each time frame using the sparse contrast enhanced projection views for the time frame and the updated composite image for that time frame. 5. The method as recited in claim 4 in which step c) is performed by acquiring a set of mask views of the subject and step f) is performed by subtracting from each contrast enhanced view in said one set a mask view acquired at a same view angle. 6. The method as recited in claim 5 which includes repeating steps d) and e) to reconstruct a plurality of CEMRA images and subtracting one of the CEMRA images from the remaining CEMRA images to suppress background signal therein. 7. The method as recited in claim 5 which further includes: h) reconstructing a CEMRA composite image using a plurality of sets of contrast enhanced views acquired in step d);i) performing a HYPR reconstruction using the set of sparse contrast enhanced views and information in the CEMRA composite image to produce the information in the set of sparse contrast enhanced... | Wisconsin Alumni Research Foundation | A61B005/055
US08260391 | B2 | 2012-09-04 | Medical sensor for reducing motion artifacts and technique for using the same | 2010-07-14 | A sensor may be adapted to reduce motion artifacts by damping the effects of outside forces and sensor motion. A sensor is provided with a motion damping structure adapted to reduce the effect of motion of a sensor emitter and/or detector. Further, a method of damping outside forces and sensor motion is also provided. | 1. A sensor comprising: a bandage-type sensor body;an emitter and a detector disposed on the sensor body and adapted to be positioned proximate to a tissue of a patient; anda motion damping structure disposed on an exterior surface of the sensor body, such that the motion damping structure is visible and faces away from the tissue when the sensor is applied to the patient, wherein the motion damping structure is capable of damping a force experienced by the sensor body, such that the motion damping structure moves at least partially independently of the emitter and the detector when an external force is applied to the motion damping structure and wherein the motion damping structure comprises a chamber, the chamber comprising a fluid layer and a retaining layer, such that the fluid layer is between the retaining layer and the sensor body and comprises a viscoelastic material. 2. The sensor, as set forth in claim 1, wherein the sensor comprises at least one of a pulse oximetry sensor or a sensor for measuring a water fraction. 3. The sensor, as set forth in claim 1, wherein the emitter comprises at least one light emitting diode, and wherein the detector comprises at least one photodetector. 4. The sensor, as set forth in claim 1, wherein the motion damping structure is only disposed about a region correlating to a tip of a digit of the patient. 5. The sensor, as set forth in claim 1, wherein the motion damping structure is disposed between the emitter and the detector. 6. The sensor, as set forth in claim 1, wherein the fluid comprises a gel. 7. A pulse oximetry system comprising: a pulse oximetry monitor; anda pulse oximetry sensor adapted to be operatively coupled to the monitor, the sensor comprising: a bandage-type sensor body;an emitter and a detector disposed on the sensor body and adapted to be positioned proximate to a tissue of a patient; anda motion damping structure disposed on an exterior surface of the sensor body, such that the motion damping structure is visible and faces away from the tissue when the sensor is applied to the patient, wherein the motion damping structure is capable of damping a force experienced by the sensor body, and wherein the motion damping structure comprises a chamber, the chamber comprising a fluid layer and a retaining layer, such that the fluid layer is between the retaining layer and the sensor body and comprises a viscoelastic material. 8. The system, as set forth in claim 7, wherein the emitter comprises at least one light emitting diode, and wherein the detector comprises at least one photodetector. 9. The system, as set forth in claim 7, wherein the motion damping structure substantially covers the exterior surface of the sensor body. 10. A sensor comprising: a conformable sensor body;an emitter and a detector disposed on an interior surface of the sensor body, wherein the interior surface is adapted to be proximate to a patient's tissue when the sensor is applied; anda motion damping struct... | Nellcor Puritan Bennett LLC | A61B005/1455
US08172760 | B2 | 2012-05-08 | Medical device encapsulated within bonded dies | 2009-06-18 | An implantable medical system includes a first die substrate with a first outer surface. The system also includes a second die substrate with a second outer surface. Furthermore, the system includes a medical device with a first portion that is mounted to the first die substrate and a second portion that is mounted to the second die substrate. The first and second die substrates are fixed to each other and substantially hermetically sealed to each other. Also, the medical device is substantially encapsulated between the first and second die substrates. The first portion is electrically connected to the second portion. Moreover, the first and second outer surfaces of the first and second die substrates are directly exposed to a biological material. | 1. An implantable medical system comprising: a first die substrate with a first outer surface;a second die substrate with a second outer surface; anda medical device with a first portion that is mounted to the first die substrate and a second portion that is mounted to the second die substrate, wherein the first portion comprises at least one first conductive element terminating at a mating surface and the second portion comprises at least one second conductive element terminating at a mating surface, the first and second die substrates being fixed to each other and substantially hermetically sealed to each other, the medical device substantially encapsulated between the first and second die substrates, the at least one first conductive element of the first portion being electrically connected to the at least one second conductive element of the second portion when the first and second die substrates substantially hermetically seal and substantially encapsulate the medical device, the first and second outer surfaces of the first and second die substrates being directly exposed to a biological material. 2. The implantable medical system of claim 1, wherein the first and second die substrates cooperate to define a cavity that is enclosed and substantially hermetically sealed between the first and second die substrates. 3. The implantable medical system of claim 2, wherein the first die substrate includes a first mating surface with an inner recess, wherein the second die substrate includes a second mating surface, and wherein the first and second mating surfaces are fixed to each other to substantially hermetically seal the inner recess and to define the cavity. 4. The implantable medical system of claim 2, wherein the first portion of the medical device includes a first sensor electrode that is disposed on an interior surface of the cavity on the first die substrate, and wherein the second portion of the medical device includes a second sensor electrode that is disposed on an interior surface of the cavity on the second die substrate, the first and second sensor electrodes cooperating with each other to detect a characteristic of the biological material. 5. The implantable medical system of claim 4, wherein at least one of the first and second die substrates includes a diaphragm that is deflectable by the biological material to move at least one of the first and second sensor electrodes relative to the other. 6. The implantable medical system of claim 5, wherein the medical device is a pressure sensor that detects a pressure of the biological material based on deflection of the diaphragm by the biological material. 7. The implantable medical system of claim 1, wherein the first outer surface includes an outer recess, and further comprising an outer electrode disposed within the outer recess, the outer electrode in electrical communication with the medical device. 8. The implantable medical system of claim 1, wherein the first and second die... | Medtronic, Inc. | A61N001/00
US08777863 | B2 | 2014-07-15 | Implantable medical device with internal piezoelectric energy harvesting | 2010-05-06 | Methods, systems, and apparatus for powering and/or recharging medical devices implanted within the body are described. An illustrative power generation module disposable within the interior space of an implantable medical device includes a module body that defines an interior cavity as well as a flexible diaphragm that spans the interior cavity. The flexible diaphragm includes a first electrical conductor, a piezoelectric layer disposed adjacent to the first electrical conductor, and a second electrical conductor disposed adjacent to the piezoelectric layer. The piezoelectric layer is configured to displace within the interior cavity and generate a voltage differential between the first electrical conductor and the second electrical conductor. | 1. A power generation module disposed within a first interior cavity of an implantable medical device, the implantable medical device comprising a housing defining the first interior cavity, wherein a portion of the housing includes a first flexible diaphragm configured to displace into the first interior cavity in response to blood pressure pulses impinging on the first diaphragm, the power generation module comprising: a module body, disposed within the first interior cavity, wherein the module body encloses a second interior cavity; anda second flexible diaphragm spanning the second interior cavity, the second flexible diaphragm including: a first electrical conductor;a piezoelectric layer disposed adjacent to the first conductor; anda second electrical conductor disposed adjacent to the piezoelectric layer;wherein the piezoelectric layer is configured to displace into the second interior cavity and generate a voltage differential between the first electrical conductor and the second electrical conductor in response to a change in pressure within the first interior cavity, wherein the change in pressure within the first interior cavity is caused by the blood pressure pulses impinging on the first diaphragm, wherein the first interior cavity is filled with air. 2. The power generation module of claim 1, further comprising a power conversion circuit that converts the voltage differential between the first and second electrical conductors into an operating current that can be used to power one or more components within the implantable medical device. 3. The power generation module of claim 1, wherein the second flexible diaphragm further comprises: an insulating layer disposed adjacent to the second electrical conductor;a third electrical conductor disposed adjacent to the insulating layer;a second piezoelectric layer disposed adjacent to the third electrical conductor; anda fourth electrical conductor disposed adjacent to the second piezoelectric layer. 4. The power generation module of claim 3, wherein the second flexible diaphragm has a neutral bending plane, and the insulating layer is positioned along the neutral bending plane. 5. The power generation module of claim 1, wherein the second interior cavity is at reduced pressure relative to a pressure exterior to the second interior cavity. 6. The power generation module of claim 1, further comprising a lower electrical conductor coupled to a lower surface of the second interior cavity. 7. The power generation module of claim 6, further comprising a controller configured to selectively electrically disconnect the lower electrical conductor and the first electrical conductor. 8. The power generation module of claim 7, wherein when the lower electrical conductor is electrically disconnected, movement of the second flexible diaphragm creates a voltage differential between the first electrical conductor and the second electrical conductor. 9. The power generation module of claim 7, wherein... | Cardiac Pacemakers, Inc. | A61B005/02
US09149564 | B2 | 2015-10-06 | Articles comprising large-surface-area bio-compatible materials and methods for making and using them | 2007-06-22 | The present invention provides articles of manufacture comprising biocompatible nanostructures comprising significantly increased surface area for, e.g., organ, tissue and/or cell growth, e.g., for bone, tooth, kidney or liver growth, and uses thereof, e.g., for in vitro testing of drugs, chemicals or toxins, or as in vivo implants, including their use in making and using artificial tissues and organs, and related, diagnostic, screening, research and development and therapeutic uses, e.g., as drug delivery devices. The present invention provides biocompatible nanostructures with significantly increased surface area, such as with nanotube and nanopore array on the surface of metallic, ceramic, or polymer materials for enhanced cell and bone growth, for in vitro and in vivo testing, cleansing reaction, implants and therapeutics. The present invention provides optically transparent or translucent cell-culturing substrates. The present invention provides biocompatible and cell-growth-enhancing culture substrates comprising elastically compliant protruding nanostructure substrates coated with Ti, TiO2 or related metal and metal oxide films. | 1. A product of manufacture comprising: (a) a biocompatible surface;(b) a plurality of macro or microscale structures fixed or loosely placed on the biocompatible surface, wherein the plurality of macro or microscale structures comprise:(i) a plurality of hairy-shaped or gauze shaped, mesh or mesh screen, wire, woven wire, spring, ribbon, powder, or flake shaped macro or microscale structures,wherein the macro or a microscale structure comprises a Ti, Zr, Hf, Nb, Ta, Mo and/or W metal material, a Ti, Zr, Hf, Nb, Ta, Mo and/or W alloy, a Ti, Zr, Hf, Nb, Ta, Mo and/or W oxide or nitride, and/or stainless steel or ceramic;(ii) a plurality of macro or microscale loose, short-fibers, or flake shaped structures, or macro or microscale particles, or macro or microscale spring-like fibers, made of a material comprising Ti, Zr, Hf, Nb, Ta, Mo and/or W metal material, a Ti, Zr, Hf, Nb, Ta, Mo and/or W alloy, a Ti, Zr, Hf, Nb, Ta, Mo and/or W oxide or nitride, and/or stainless steel or ceramic;(iii) a plurality of macro or microscale micro-wires, micro-fibers or micro-ribbons;(iv) a plurality of macro or microscale spring-like fibers or mesh screen shapes comprising a Ti or metal material, an alloy and/or a ceramic;or(v) any combination of the structures of (i) to (vi); and(c) a plurality of nanostructures on the surfaces of the plurality of macro or microscale structures, wherein the nanostructures comprise: (i) a biocompatible vertically aligned nanotube array structure;(ii) a lock-in nanostructure comprising a plurality of nanopores or nanotubes,wherein the nanopore or nanotube entrance has a smaller diameter or size than the rest (the interior) of the nanopore or nanotube to exhibit a re-entrant configuration;(iii) a dual structured biomaterial comprising (A) micro- or macro-pores, wherein the micro or macro pores has an average diameter, or equivalent diameter if the pores are not circular, in the range of between about 0.5-1,000 μm, or between about 1-100 μm; and,(B) a surface area covered with nanotubes;(iv) a plurality of enlarged diameter nanopores and/or nanotubes, wherein the nanopores and/or nanotubes comprise at least 150 nm; or(v) any combination of (i) to (iv)wherein the nanotube structures of any of (c)(i) to (c)(iv) have a thickness of at least about 0.5, or between about 1 and 50 nm, or between about 1 and 100 nm, or between about 10 nm to 500 nm,wherein at least a portion of, or substantially all of, or all of, the surface of the macro or microscale structures of (b) are covered or coated by one or a plurality of nanostructures structures of (c). 2. The product of manufacture of claim 1: wherein a macro or microscale members and/or a nanotube array structure, or a structure of any of (b) or (c) comprises a Ti, Zr, Hf, Nb, Ta, Mo and/or W metal; a Ti, Zr, Hf, Nb, Ta, Mo and/or W alloy; and/or, a Ti, Zr, Hf, Nb, Ta, Mo and/or W oxide or nitride, and/or stainless steel or ceramic. 3. The product of manufacture of claim 1, wherein the plura... | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA | A61K049/00 A01N063/00 A61F002/00 A61L027/54 A61L027/30 A61L027/38
US09220574 | B2 | 2015-12-29 | Tissue marker | 2008-09-24 | The invention relates to a tissue marker (1, 1′) for marking a lesion (16) in the body tissue (6). The invention is characterized in that the tissue marker (1, 1′) comprises a transponder (2). The transponder (2) can be activated by electromagnetic radiation (8) in such a manner that the transponder (2) emits a localization signal (9) in the form of electromagnetic radiation by means of which the position of the tissue marker (1, 1′) in the body tissue (6) can be detected. The invention further relates to a system for determining the position of such a tissue marker (1, 1′) and to the use of an RFID tag as a tissue marker (1, 1′) for marking a lesion (16) in the body tissue. According to an advantageous embodiment of the invention, the tissue marker (1, 1′) is biodegradable. | 1. Tissue marker for marking a lesion in body tissue, wherein the tissue marker comprises a transponder, whereby the transponder can be activated via electromagnetic radiation, specifically in such a manner that the transponder emits a localization signal to determine the position of the tissue marker in the body tissue as electromagnetic radiation, wherein the transponder is arranged to emit the localization signal in a coherent manner, with a defined and constant phasing, in order to assign a phase value to a position in space, and wherein the tissue marker is biodegradable and has a biodegradable sheath that encloses the transponder, whereby the biodegradation time can be pre-determined by the thickness and/or the composition of the sheath. 2. Tissue marker according to claim 1, wherein the transponder has an integrated electronic circuit and an antenna connected with the latter, for reception and transmission of electromagnetic radiation. 3. Tissue marker according to claim 2, wherein the transponder is configured as a passive transponder, whereby the power supply of the circuit takes place via the induction current generated in the antenna during reception of electromagnetic radiation. 4. Tissue marker according to claim 1, wherein the tissue marker has a sensor element connected with the transponder, whereby the transponder emits a sensor signal of the sensor element as electromagnetic radiation. 5. Tissue marker according to claim 1, wherein the transponder is a radio frequency identification (RFID) tag. 6. Tissue marker according to claim 5, wherein medical data with regard to the lesion to be marked are stored or can be stored in an electronic data memory of the RFID tag. 7. System for determining the position of a tissue marker according to claim 1 in body tissue, the system comprising the tissue marker according to claim 1 and further comprising a transmitter for emitting electromagnetic radiation for activation of the transponder of the tissue marker, a receiver for receiving the localization signal emitted by the transponder, and having a processor for evaluating the localization signal, wherein the processor is set up for determining the position of the tissue marker by determining the phasing of the electromagnetic radiation of the localization signal at the location of the reception unit, in order to assign a phase value to a position in space. 8. System according to claim 7, further comprising an aiming device for determining the position and/or orientation of a medical instrument relative to the position of the tissue marker, the aiming device comprising at least one of an aiming device transmitter affixed to the medical instrument and an aiming device receiver affixed to the medical instrument. 9. System according to claim 8, wherein the aiming device has a display that indicates the guidance direction or the orientation of the medical instrument to the person guiding the medical instrument, so that the medical instrume... | amedo smart tracking solutions GmbH | A61B005/05 A61B019/00 A61B005/00 A61B017/00
US09126049 | B2 | 2015-09-08 | Shock reduction using absolute calibrated tissue oxygen saturation and total hemoglobin volume fraction | 2010-06-10 | An implantable medical device for detecting and treating an arrhythmia includes an optical sensor adapted for positioning adjacent to a blood-perfused tissue volume. In one embodiment for controlling arrhythmia therapies delivered by the device, the optical sensor is controlled to emit light in response to detecting an arrhythmia, detect light scattered by the volume of blood perfused tissue including measuring an optical sensor output signal corresponding to the intensity of scattered light for at least four spaced-apart wavelengths, and compute a volume-independent measure of tissue oxygen saturation from the detected light. The hemodynamic status of the arrhythmia is detected in response to the measure of tissue oxygen saturation. | 1. A method, comprising: sensing cardiac depolarization signals;detecting an arrhythmia in response to the depolarization signals;controlling an optical sensor to emit light in response to the detected arrhythmia;detecting light scattered by a volume of blood perfused tissue wherein detecting light comprises measuring an optical sensor output signal corresponding to light attenuation of at least four spaced apart light wavelengths;computing a measure of tissue oxygenation from the optical sensor output signal, wherein computing the measure of tissue oxygenation comprises: computing a light attenuation for each of the at least four wavelengths;computing a second derivative of the light attenuation with respect to two different wavelengths;computing a measure of tissue oxygen saturation independent of total hemoglobin as a ratio of the second derivatives at the two different wavelengths;computing a measure of total hemoglobin volume fraction using the second derivative of a selected one of the two different wavelengths;detecting a hemodynamic status of the arrhythmia in response to the measure of tissue oxygenation, wherein the measure of tissue oxygenation is computed responsive to at least one of the measure of tissue oxygen saturation independent of total hemoglobin and the measure of total hemoglobin volume fraction;computing the measure of tissue oxygenation as a weighted combination of the measure of oxygen saturation and the measure of total hemoglobin volume fraction;determining if one of the measure of oxygen saturation and the measure of total hemoglobin volume fraction is out of a respective acceptable range; andadjusting a weighting factor used to compute the weighted combination in response to one of the measures of oxygen saturation and total hemoglobin volume fraction being out of a respective acceptable range. 2. The method of claim 1, wherein detecting a hemodynamic status comprises: detecting a hemodynamically unstable arrhythmia in response to a decreasing measure of tissue oxygen saturation;enabling a shock therapy for treating the arrhythmia in response to detecting the hemodynamically unstable arrhythmia; andwithholding delivering of a shock therapy for treating the arrhythmia in response to not detecting the hemodynamic status as being unstable. 3. The method of claim 1, wherein the measure of total hemoglobin volume fraction is computed using the measure of tissue oxygen saturation. 4. The method of claim 1, further comprising: determining a baseline measurement of tissue oxygenation prior to detecting the arrhythmia;computing an onset measurement of tissue oxygenation corresponding to a time at which the arrhythmia is detected in response to the depolarization signals;comparing the baseline measurement and the onset measurement; andenabling a shock therapy in response to the onset measurement being less than the baseline measurement. 5. The method of claim 1, further comprising: computing an onset measurement of tissue ... | Medtronic, Inc. | A61B005/00 A61N001/362 A61B005/02 A61B005/145 A61B005/1459 A61N001/365 A61N001/39 A61B005/0215 A61B005/042 A61B005/11
US08647271 | B2 | 2014-02-11 | Hydrogel implant for sensing metabolites in body tissue | 2008-05-23 | An implant (110) for detecting at least one analyte (126) in a body fluid, in particular an eye fluid, is proposed. The implant (110) is designed to be implanted in a tissue layer and/or a chamber of an eye of a patient, the implant (110) having a hydrogel matrix (110) with at least one hydrogel (114). The implant (110) also has sensor particles (116) dispersed in the hydrogel matrix (110), the sensor particles (116) having at least one sensor matrix (120) with a sensor matrix material (122) and at least one sensor material (124). | 1. An implant for detecting at least one analyte in a body fluid, the implant being designed to be implanted in a body tissue of a patient, the implant comprising a hydrogel matrix with at least one hydrogel and sensor particles dispersed homogeneously in the hydrogel matrix, wherein the sensor particles comprise at least one sensor matrix comprising a sensor matrix material and at least one sensor material, the implant further comprising reference components also dispersed homogeneously in the hydrogel matrix, wherein the reference components are at least substantially analyte-invariant. 2. The implant as claimed in claim 1, the sensor material comprising at least one of the following materials: concanavalin A; glucose oxidase; glucose dehydrogenase; hexokinase; glucose/galactose-binding protein; a protein and/or a fragment functionally equivalent to a protein; a mutant of hexokinase; a mutant of glucose/galactose-binding protein; and a borate ester derivative. 3. The implant as claimed in claim 1, the sensor particles comprising a shell that is at least partially permeable to the analyte. 4. The implant as claimed in claim 1, the sensor material comprising at least one of the following materials: a molecule and/or a group which, in the presence of an analyte, changes at least one chemical or physical property;a molecule and/or a group where the analyte, on approaching and/or reacting with the molecule and/or the group, changes at least one chemical or physical property; anda molecule and/or a group where the analyte undergoes a chemical or physical bonding to the molecule and/or the group, which bonding can be detected by a change in at least one chemical or physical property. 5. The implant as claimed in claim 1, the sensor matrix material. comprising at least one of the following materials; an alginate; a polysaccharide; a polysaccharide derivative; a synthetic polymer; a copolymer crosslinked covalently or via a hydrogen bond; an ionically crosslinked polymer or copolymer; a polyacrylamide; a copolymer containing an acrylic acid unit; an acrylic acid derivative; and a methacrylic acid derivative. 6. The implant as claimed in claim 1, wherein the reference component comprises at least one luminescent component, wherein the luminescence properties of the luminescent component are at least substantially analyte-invariant. 7. The implant as claimed in claim 1, wherein the implant comprises reference particles implanted in the hydrogel matrix, the reference particles comprising the at least one reference component, the reference component comprising, at least one reference matrix material, and wherein the reference component is bonded physically and/or chemically on the surface and/or the interior of the reference matrix material. 8. A method for detecting at least one analyte in a body fluid, comprising implanting an implant according to claim 1 in a body tissue and detecting the at least one analyte using qualitative or quantitative dete... | Eyesense AG | A61B005/00
US08706193 | B2 | 2014-04-22 | Catheter with obliquely-oriented coils | 2009-06-22 | A medical probe includes an elongate body having a longitudinal axis, and a plurality of magnetic field transducers, which are contained within the body and have respective transducer axes that are oriented obliquely with respect to the longitudinal axis. | 1. A medical probe, comprising: an elongate body having a longitudinal axis and a central lumen about the longitudinal axis for a passage of solution, substances or objects therethrough; anda plurality of magnetic field transducers, which are contained within the body, each magnetic field transducer having a respective magnetic field transducer axis that is oriented obliquely with respect to the longitudinal axis, wherein each magnetic field transducer is positioned along the respective magnetic field transducer axis so as not to interfere with the longitudinal axis and preserving an unobstructed lumen for the central lumen for the passage of solution, substances or objects therethrough and whereby the lumen does not pass through any of the magnetic field transducers. 2. The probe according to claim 1, wherein the elongate body comprises an insertion tube for insertion into an organ of a patient. 3. The probe according to claim 1, wherein the plurality of transducers comprises three transducers. 4. The probe according to claim 3, wherein the transducers are substantially mutually orthogonal. 5. The probe according to claim 1, wherein the transducers are positioned within the elongate body so that none of the transducers intersects the longitudinal axis. 6. The probe according to claim 1, wherein the transducers are positioned at a distal end of the elongate body, and wherein the probe comprises leads connecting the transducers to a connector located at a proximal end of the elongate body. 7. The probe according to claim 1, and comprising a cylinder-shaped fixture for holding the transducers within the elongate body, wherein the fixture has a central opening that is parallel with the longitudinal axis. 8. The probe according to claim 1, wherein the field transducers comprise coils. 9. A medical position tracking system, comprising: a medical probe, comprising: an elongate body having a longitudinal axis and a central lumen about the longitudinal axis for a passage of solution, substances or objects therethrough; anda plurality of magnetic field transducers, which are contained within the body, each magnetic field transducer having a respective magnetic field transducer axis that is oriented obliquely with respect to the longitudinal axis, wherein each magnetic field transducer is positioned along the respective magnetic field transducer axis so as not to interfere with the longitudinal axis and preserving an unobstructed lumen for the central lumen for the passage of solution, substances or objects therethrough and whereby the lumen does not pass through any of the magnetic field transducers; anda position measurement subsystem, which is arranged to exchange one or more magnetic fields with the magnetic field transducers in the probe, and to compute a position of the probe responsively to the exchanged fields. 10. The system according to claim 9, wherein the elongate body comprises an insertion tube for insertion into an organ of a patien... | Biosense Webster, Inc. | A61B005/00
US09107587 | B2 | 2015-08-18 | Device and process for calculating new indices of arterial stiffness, and/or for stroke volume monitoring | 2009-02-12 | A process for calculating an indice of arterial stiffness, including the step of extracting pulse wave analysis data from a recorded pressure of an artery, the recorded pressure is recorded as a function of time, the indice being calculated as a function of the extracted data, and the extracted pulse wave analysis data including at least one time interval. | 1. A process for calculating an indice of arterial stiffness of a human or another animal, comprising: recording an artery pressure as a function of time from the human or another animal;providing a processor configured to perform the following steps:extracting pulse wave analysis data from the recorded pressure of the artery of the human or another animal,calculating the indice of arterial stiffness of the human or another animal as a function of the extracted pulse wave analysis data comprising:at least one extracted time interval including an extracted systolic time, andan extracted diastolic pressure and an extracted pressure at an inflection point during systole, wherein the inflection point is not an inflexion point at the beginning of systole or at the end of systole,the indice of arterial stiffness of the human or another animal being thus calculated as a function of the extracted systolic time, the extracted diastolic pressure and the extracted pressure at said inflection point during systole; andmonitoring, in time, the calculated indice of arterial stiffness of the human or another animal in one of a medical unit, a surgical intensive care unit or an anesthesia unit. 2. The process according to claim 1, wherein the at least one time interval comprises a time to the pressure at an inflection point during systole, the indice being calculated as a function of the ratio: ST/Δt, where Δt is the time to the pressure at the inflection point during systole and ST is the systolic time. 3. The process according to claim 1, wherein the at least one time interval comprises a period of the artery, the indice being calculated as a function of the ratio: ST/T, where T is the period and ST is the systolic time. 4. The process according to claim 1, wherein the indice of arterial stiffness is a characteristic impedance calculated according to the following equation: Zc=[(Pi−DAP)*ST]/(2SV),where Zc is the characteristic impedance, DAP is the diastolic pressure, Pi is the pressure at the inflection point during systole, SV is a recorded stroke volume, and ST is the systolic time. 5. The process according to claim 2, wherein the indice of arterial stiffness is a total arterial compliance calculated according to the following equation: C=(SV*Δt)/[(Pi−DAP)*ST], where C is the total arterial compliance, DAP is the diastolic pressure, Pi is the pressure at the inflection point during systole, SV is a recorded stroke volume, ST is the systolic time, and Δt is the time to the pressure at the inflection point during systole. 6. The process according to claim 2, wherein the indice of arterial stiffness is a total arterial stiffness calculated according to the following equation: 1/C=[(Pi−DAP)*ST]/(SV*Δt),where 1/C is the total arterial stiffness, DAP is the diastolic pressure, Pi is the pressure at the inflection point during systole, SV is a recorded stroke volume, ST is the systolic time, and Δt is the time to the pressure at the inflection point during sys... | Universite Paris SUD Assistance Publique-Hopitaux de Paris Universite Paris 13 Ecole Normale Superieure de Technoiques Avancees | A61B005/02 A61B005/0215 A61B005/021 A61B005/029 A61B005/1455
US08430822 | B2 | 2013-04-30 | Blood pressure measuring apparatus and method of controlling the same | 2010-08-11 | A high-accuracy blood pressure value deriving method that is hardly influenced by individual differences of a person to be measured is provided. A blood pressure measuring apparatus includes a cuff member comprising a compression air bladder, a sub air bladder, and a pulse wave detection air bladder, pressure control means for pressurizing or depressurizing each air bladder, a pressure sensor which senses the internal pressure of each air bladder, pulse wave signal extracting means for extracting time-series data of a pulse wave signal superposed on a cuff internal pressure, in the process during which the pressure control means pressurizes or depressurizes each air bladder, and blood pressure value deriving means for deriving a systolic blood pressure value and/or a diastolic blood pressure value based on a change in feature amount of the pulse wave signal and a cuff internal pressure at a point of time of the change. | 1. A blood pressure measuring apparatus comprising: a cuff member comprising a compression air bladder which is adapted to be formed on a side to be brought into contact with a blood pressure measurement portion and which is adapted to press a whole blood pressure measurement portion, a sub air bladder which is adapted to be formed on a side of said compression air bladder, which is brought into contact with the blood pressure measurement portion, and which is adapted to press a heart side of a blood vessel in the blood pressure measurement portion, and a pulse wave detection air bladder which is adapted to be formed on the side of said compression air bladder, which is brought into contact with the blood pressure measurement portion, and which is adapted to detect a pulse wave on a slightly downstream side of a central portion of the blood vessel in the blood pressure measurement portion;pressure control means for pressurizing or depressurizing each air bladder of said cuff member;a pressure sensor which senses an internal pressure of each air bladder of said cuff member;pulse wave signal extracting means for extracting time-series data of a pulse wave signal superposed on a cuff internal pressure sensed by said pressure sensor, in the process during which said pressure control means pressurizes or depressurizes each air bladder of said cuff member; andblood pressure value deriving means for deriving a systolic blood pressure value and/or a diastolic blood pressure value based on a change in feature amount of the pulse wave signal and a cuff internal pressure at a point of time of the change,wherein said blood pressure value deriving means detects, for each of a plurality of one-period pulse wave signals contained in the pulse wave signal time-series data, a maximum gradient point of the one-period pulse wave in a period between a peak point and a bottom point appearing prior to the peak point, and derives the systolic blood pressure value and/or the diastolic blood pressure value based on a difference between a pulse wave amplitude value at a bottom point appearing prior to the detected maximum gradient point and a value, which is obtained at a time of the bottom point, of a tangent passing through the detected maximum gradient point. 2. A blood pressure measuring apparatus comprising: a cuff member comprising a compression air bladder which is adapted to be formed on a side to be brought into contact with a blood pressure measurement portion and which is adapted to press a whole blood pressure measurement portion, a sub air bladder which is adapted to be formed on a side of said compression air bladder, which is brought into contact with the blood pressure measurement portion, and which is adapted to press a heart side of a blood vessel in the blood pressure measurement portion, and a pulse wave detection air bladder which is adapted to be formed on the side of said compression air bladder, which is brought into contact with the blood pressur... | Terumo Kabushiki Kaisha | A61B005/02
US08372351 | B2 | 2013-02-12 | Glucose monitoring system with wireless communications | 2010-05-28 | A glucose monitor having only a test strip reader and wireless transmission to a remote hand held calculation processor. All glucose data processing occurs in a remote hand held calculator processor with the glucose monitor providing only sensing and wireless transfer of data. In an embodiment, the glucose monitor includes a button for use in wirelessly pairing the monitor to a smart phone for data communication with a non-proprietary communication protocol, and a light source to communicate the progress of the pairing. In a further embodiment, the glucose monitor includes a switch coupled to the strip reader such that power is applied to the monitor when a test strip is inserted and power is removed with the withdrawal of the test strip. In another embodiment, inserting the test strip and activating the switch also places the glucose monitor into the pairing search mode to communication with a host remote processor. | 1. A diabetes management system comprising: a portable glucose monitor including: a glucose sensor that senses a glucose level and provides glucose sensor data representative of the level of glucose sensed;a glucose monitor processor configured to receive the glucose sensor data, process the received glucose sensor data, and provide such processed glucose data for transmission;a wireless communication component including a transmitter configured to receive the processed glucose data and wirelessly transmit the processed glucose data by a non-proprietary protocol;a test strip switch located and configured to be activated by insertion of a test strip into the portable glucose monitor, the test strip switch providing an activation signal once activated;wherein the glucose monitor processor is connected to the test strip switch and is programmed to become active upon receiving the activation signal at which time the glucose monitor processor receives the glucose senor data and provides the processed glucose data for transmission, and controls the wireless communication component to transmit the processed glucose data;a portable housing in which are located only the glucose sensor, the glucose data processor, and the communication component and including no direct user interface other than the glucose sensor; anda remote host calculation device including: a remote host communication module configured to wirelessly pair with the portable glucose monitor using the non-proprietary protocol so that data communications between the two may occur;a display;a remote host processor; anda nonvolatile memory in which is stored a host glucose data processing program;wherein the remote host processor is connected to the remote host communication module; the display, and the remote host nonvolatile memory to access the memory to upload into the processor and run the host glucose data processing program which programs the remote host processor to:control the remote host communication module to pair with the portable glucose monitor and receive processed sensor data;analyze glucose data received from the portable glucose monitor;display a determined glucose level; andprovide a calculated action on the display in response to the determined glucose level. 2. The diabetes management system of claim 1, wherein the calculated action displayed by the host program includes a recommended dose of medication. 3. The diabetes management system of claim 1, wherein the glucose processing program further programs the remote host processor to time and date stamp the glucose data received. 4. The diabetes management system of claim 1, wherein the remote host calculation device comprises a smart phone having the host glucose data processing program as a downloadable application program. 5. The diabetes management system of claim 1, further comprising an input element configured to allow a user to input a manual coding for a test strip, wherein the host glucose data processing pro... | Abbott Diabetes Care Inc. | G01N033/00
US08005532 | B2 | 2011-08-23 | Electrocardiograph with extended lead function, and extended lead electrocardiogram deriving method | 2010-06-29 | There are provided an electrocardiograph with an extended lead function and an extended lead ECG deriving method capable of easily deriving an ECG signal of an extended lead ECG by an arithmetic operation, based on ECG signals of a standard 12-lead ECG measured by a potential detector 10. An ECG memory 12 stores the ECG signals measured as the standard 12-lead ECG by the potential detector 10. An extended lead ECG calculator 16 calculates extended lead ECGs V7-V9 from the ECG signals stored in the ECG memory 12, using coefficients α representing a relationship among leads. The extended lead ECGs V7-V9 calculated are displayed through an extended lead ECG waveform outputting device 18 on a display monitor 20. | 1. An electrocardiograph with an extended lead function comprising: storage means for storing lead vectors L1T, LIIT, L1T-L6T, and extended lead vectors L7T-L9T, L3RT-L6RT;a potential detector configured for measuring at least two chest lead potentials and at least two limb lead potentials out of all of the lead potentials of a standard 12-lead electrocardiogram (ECG); andextended lead potential calculating means for calculating extended lead potentials V7-V9, V3R-V5R, based on the lead potentials of the standard 12-lead ECG measured by the potential detector;wherein said extended lead potential calculating means applies the lead potentials measured by said potential detector and the measured lead vectors stored in said lead vector storage means to the following equations (1, 2) to calculate heart vectors with components hx, hy and hz, and calculates said extended lead potentials based on said heart vectors and said extended lead vectors stored in said lead vector storage means:(LITLIITL1TL2TL3TL4TL5TL6T)⁢(hxhyhz)=(VIVIIV1V2V3V4V5V6)(1)wherein (T indicates transposed vector)L=(lxlylz).(2) 2. The electrocardiograph with the extended lead function according to claim 1, further comprising: electrodes configured to be mounted on a body surface of a living body in order to obtain lead waveforms of the standard 12-lead ECG;the potential detector being arranged for measuring ECG signals of the standard 12-lead ECG from the respective electrodes;an ECG signal memory in said storage means for storing each of the ECG signals of the standard 12-lead ECG measured by the potential detector;standard 12-lead ECG waveform processing means for performing waveform processing of the standard 12-lead ECG with input of the ECG signals obtained via the ECG signal memory;the extended lead potential calculating means calculates the extended lead potentials with input of the ECG signals obtained via the ECG signal memory;extended lead ECG waveform outputting means for performing waveform processing of the extended lead ECG with input of the ECG signal calculated by the extended lead ECG signal calculating means; anda display monitor for simultaneously displaying on a screen, the ECG waveform outputs from the standard 12-lead ECG waveform outputting means and from the extended lead ECG waveform outputting means. 3. An extended lead ECG deriving method comprising the steps of: a first step of storing lead vectors L1T, LIIT, L1T-L6T, and extended lead vectors L7T-L9T, L3RT-L6RT in a signal memory;a second step of measuring at least two chest lead potentials and at least two limb lead potentials out of all of the lead potentials of a standard 12-lead ECG with a potential detector; anda third step of calculating, using a calculating device, extended lead potentials V7-V9, V3R-V5R, based on the lead potentials of the standard 12-lead ECG measured by the potential detector;wherein said third step applies the lead potentials and the measured lead vectors stored in said first s... | Daming Wei | A61B005/0402
US08174395 | B2 | 2012-05-08 | Transceiver unit in a measurement system | 2010-04-30 | A measurement system may comprise a sensor wire and a transceiver unit. The sensor wire may comprise an insertable portion configured to be inserted in a blood vessel of a patient's body and a sensor disposed within the insertable portion at a distal end of the sensor wire. The sensor is configured to measure a parameter when inserted inside the patient. The transceiver unit may comprise: a housing adapted to be connected to a proximal end of the sensor wire; and a first communication module within the housing adapted to wirelessly communicate by a communication signal with an external second communication module in order to transfer information to the external second communication module. | 1. A measurement system, comprising: a sensor wire comprising an insertable portion configured to be inserted in a blood vessel of a patient's body and a sensor disposed within the insertable portion at a distal end of the sensor wire, wherein the sensor is configured to measure a parameter when inserted inside the patient; anda transceiver unit, comprising: a housing adapted to be connected to a proximal end of the sensor wire; anda first communication module within the housing adapted to wirelessly communicate by a communication signal with an external second communication module in order to transfer information to the external second communication module,wherein the communication signal, including sensor values, is generated by the transceiver unit and transferred as a data stream,wherein the housing is configured to remain external to the patient's body. 2. The measurement system according to claim 1, wherein said communication signal is a radio frequency signal. 3. The measurement system according to claim 2, wherein radio frequency communication is performed by Bluetooth. 4. The measurement system according to claim 2, wherein said radio frequency signal has a frequency of 2.4 GHz. 5. The measurement system according to claim 2, wherein radio frequency communication is performed by a frequency hopping technique. 6. The measurement system according to claim 1, wherein said data stream is in a form of data packets. 7. The measurement system according to claim 1, wherein the first communication module is configured to perform bi-directional communication with the second communication module. 8. The measurement system according to claim 1, wherein said communication signal is one of an infrared signal, an ultrasound signal, and a light signal. 9. The measurement system according to claim 1, wherein the transceiver unit further comprises an activator configured to activate and initialize the transceiver unit. 10. The measurement system according to claim 9, wherein the transceiver unit further comprises a female connector with electrical contact surfaces, and wherein said activator activates and initializes the transceiver unit when electrical connections are established between the electrical contact surfaces of the female connector and electrical contact surfaces of a male connector at the proximal end of the sensor wire. 11. The measurement system according to claim 10, wherein said female connector comprises an insulating hollow housing containing a predetermined number of hollow contact surfaces to make contact with the electrical contact surfaces of the male connector. 12. The measurement system according to claim 1, wherein the transceiver unit further comprises a guide to guide the sensor wire during insertion into the patient. 13. The measurement system according to claim 1, wherein a sensor signal adapting circuitry is disposed within the housing and is adapted to filter, process, and format a signal received from the senso... | St. Jude Medical Systems AB | G08B023/00
US08133184 | B2 | 2012-03-13 | Calibration of in vivo blood pressure sensors | 2010-07-08 | A method for performing an in-vivo calibration of a blood pressure sensor that is associated with a balloon of an in-vivo balloon system, the sensor and balloon being associated such that the sensor is in-vivo when the balloon is in-vivo. The balloon is inflated so that a gas pressure in the balloon system is indicative of a patient's blood pressure. The patient's blood pressure is monitored through two channels, the gas pressure and the sensor. The blood pressure measurements obtained by monitoring the gas pressure are used as reference, or “true,” blood pressure measurements to determine a mathematical relationship between blood pressure measurements obtained through the sensor and the reference blood pressure measurements. In this manner, future blood pressure measurements obtained through the sensor can be modified according to the mathematical relationship to generate calibrated blood pressure measurements. | 1. A method for performing an in-vivo calibration of a blood pressure sensor that is associated with a balloon of an in-vivo balloon system, the sensor and balloon being associated such that the sensor is in-vivo when the balloon is in-vivo, the method comprising the steps of: inflating or deflating the balloon so that a gas pressure in the balloon system is indicative of a patient's blood pressure;monitoring the patient's blood pressure by observing the gas pressure through a gas pressure sensor while simultaneously monitoring the patient's blood pressure through the blood pressure sensor and recording a pressure signal generated by a gas pressure sensor and a pressure signal generated by the blood pressure sensor; andusing blood pressure readings obtained by monitoring the gas pressure as reference blood pressure measurements and determining a mathematical relationship between blood pressure measurements obtained through the sensor and the reference blood pressure measurements such that future blood pressure measurements obtained through the sensor can be modified according to the mathematical relationship to generate calibrated blood pressure measurements,wherein the step of using includes performing a correlation function between a window of the balloon-based data and the corresponding window of the blood pressure sensor, and performing a time-alignment of the pressure signal generated by the gas pressure sensor and the pressure signal generated by the blood pressure sensor to a point in time corresponding to a peak of the correlation function. 2. A method for performing an in-vivo calibration of a blood pressure sensor that is associated with a balloon of an in-vivo balloon system, the sensor and balloon being associated such that the sensor is in-vivo when the balloon is in-vivo, the method comprising the steps of: inflating the balloon so that a gas pressure in the balloon system is indicative of a patient's blood pressure;monitoring the patient's blood pressure by observing the gas pressure through a gas pressure sensor while simultaneously monitoring the patient's blood pressure through the blood pressure sensor and recording a pressure signal generated by a gas pressure sensor and a pressure signal generated by the blood pressure sensor from the time where a membrane in the in-vivo balloon system is toggled to inflate the balloon; andusing blood pressure readings obtained by monitoring the gas pressure as reference blood pressure measurements and determining a mathematical relationship between blood pressure measurements obtained through the sensor and the reference blood pressure measurements such that future blood pressure measurements obtained through the sensor can be modified according to the mathematical relationship to generate calibrated blood pressure measurements,wherein the step of using includes filtering the pressure signal generated by the blood pressure sensor to generate a filtered signal and using a correlation or cros... | Datascope Investment Corp. | A61B005/02
US08200321 | B2 | 2012-06-12 | Method for measuring patient posture and vital signs | 2009-05-20 | The invention provides a system and method for measuring vital signs (e.g. SYS, DIA, SpO2, heart rate, and respiratory rate) and motion (e.g. activity level, posture, degree of motion, and arm height) from a patient. The system features: (i) first and second sensors configured to independently generate time-dependent waveforms indicative of one or more contractile properties of the patient's heart; and (ii) at least three motion-detecting sensors positioned on the forearm, upper arm, and a body location other than the forearm or upper arm of the patient. Each motion-detecting sensor generates at least one time-dependent motion waveform indicative of motion of the location on the patient's body to which it is affixed. A processing component, typically worn on the patient's body and featuring a microprocessor, receives the time-dependent waveforms generated by the different sensors and processes them to determine: (i) a pulse transit time calculated using a time difference between features in two separate time-dependent waveforms, (ii) a blood pressure value calculated from the time difference, and (iii) a motion parameter calculated from at least one motion waveform. | 1. A method for monitoring vital signs and posture of a patient, the method comprising the following steps: (a) detecting a first time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a first sensor comprising a first detector configured to be worn on the patient's body; (b) detecting a second time-dependent physiological waveform indicative of one or more contractile properties of the patient's heart with a second sensor comprising a second detector configured to be worn on the patient's body; (c) detecting a set of time-dependent motion waveforms with at least one motion-detecting sensor positioned on the patient's torso, wherein the one motion-detecting sensor detects a set of motion waveforms indicative of motion of the patient's torso; (d) processing the first and second time-dependent physiological waveforms to determine at least one vital sign from the patient; (e) analyzing at least a portion of the set of time-dependent motion waveforms, or a mathematical derivative thereof, to determine a vector corresponding to motion of the patient's torso; (f) comparing the vector to a coordinate space representative of how the motion-detecting sensor is oriented on the patient to determine a posture parameter; and (g) comparing the posture parameter to a threshold value to estimate the patient's posture. 2. The method of claim 1, wherein step (e) further comprises analyzing three time-dependent motion waveforms, each corresponding to a unique axis, to determine the vector. 3. The method of claim 2, wherein step (f) further comprises comparing the vector to a coordinate space defined by three positional vectors, each indicating an orientation of the motion-detecting sensor on the patient. 4. The method of claim 3, wherein a first positional vector corresponds to a vertical axis in the coordinate space, a second positional vector corresponds to a horizontal axis in the coordinate space, and the third positional vector corresponds to a normal axis extending normal to the patient's chest. 5. The method of claim 3, wherein the posture parameter is an angle. 6. The method of claim 5, wherein step (f) further comprises comparing the vector to the coordinate space to determine the angle between the vector and at least one of the three positional vectors. 7. The method of claim 6, wherein step (f) further comprises comparing the vector to the coordinate space to determine the angle between the vector and a vector corresponding to a vertical axis, or a property of the angle between the vector and the vector corresponding to the vertical axis. 8. The method of claim 7, wherein step (g) further comprises comparing the angle to a threshold value that is substantially equivalent to 45 degrees, and wherein the patient's posture is estimated to be upright if the angle is less than the threshold value. 9. The method of claim 7, wherein step (g) further comprises comparing the angle to a threshold v... | Sotera Wireless, Inc. | A61B005/00
US08670811 | B2 | 2014-03-11 | Pulse oximetry system for adjusting medical ventilation | 2010-06-25 | A physiological monitoring system can include a physiological monitor having one or more processors that can derive oxygen saturation values from a patient. The oxygen saturation values can correspond to values of oxygen saturation in blood at a tissue site of the patient. The physiological monitor can output an indication of amplitude of the differences per respiratory cycle in the oxygen saturation values. | 1. A noninvasive method of analyzing oxygen saturation of a patient, the method comprising: deriving oxygen saturation values from a patient, the oxygen saturation values reflecting oxygen saturation of blood at a tissue site over a respiration cycle of the patient;calculating a difference between a maximum and a minimum of the derived oxygen saturation values over the respiration cycle;outputting an indication of the difference between the maximum and the minimum of the derived oxygen saturation values; andoutputting an indication of a risk of overextension of alveoli of the patient. 2. The method of claim 1, wherein said outputting the indication of the difference comprises outputting a graph of the difference between the maximum and the minimum of the derived oxygen saturation values. 3. The method of claim 1, wherein said outputting the indication of the difference comprises outputting a percentage value of variability. 4. The method of claim 1, further comprising determining whether the difference between the maximum and minimum of the derived oxygen saturation values is above a threshold value. 5. The method of claim 4, wherein said outputting the indication of the difference comprises generating an alarm in response to a variability being above the threshold value. 6. The method of claim 1, further comprising providing a control signal to a ventilator responsive to the difference between the maximum and minimum of the derived oxygen saturation values. 7. The method of claim 6, wherein the control signal to the ventilator increases or decreases an FiO2 amount delivered by the ventilator to the patient. 8. The method of claim 7, wherein the FiO2 amount is increased or decreased to keep the patient's oxygen saturation values on a steep part of a SpO2 versus FiO2 curve of the patient. 9. A physiological monitoring system, the system comprising: a physiological monitor comprising one or more processors configured to: derive oxygen saturation values from a patient, the oxygen saturation values corresponding to values of oxygen saturation in blood at a tissue site over a respiration cycle of the patient;calculate a difference between a maximum and a minimum of the derived oxygen saturation values over the respiration cycle;output an indication of the difference between the maximum and the minimum of the derived oxygen saturation values; andoutput an indication of a risk of overextension of alveoli of the patient. 10. The system of claim 9, wherein the physiological monitor is further configured to output the indication of the difference by at least outputting a graph of the difference between the maximum and the minimum of the derived oxygen saturation values. 11. The system of claim 9, wherein the physiological monitor is further configured to output the indication of the difference by at least outputting a percentage value representing the difference. 12. The system of claim 9, wherein the physiological monitor is further configured ... | Masimo Corporation | A61B005/00
US08398557 | B2 | 2013-03-19 | Spread-spectrum method for determining physiological parameters | 2007-02-06 | An apparatus including a receiver and an extractor is provided for determining a physiological parameter of a living being. The receiver is adapted to receive iterating optical signals, the optical signal including sequences, and a sequence including at least two bright periods during which a transmit light source adopts an on state and at least one dark period during which no transmit light source adopts an on state, and the at least two bright periods being irregularly arranged within a sequence. In addition, the apparatus includes an extractor. The extractor extracts information on the physiological parameter from a receive signal, the extractor being adapted to extract, on the basis of the information on the arrangement of the bright periods within the sequence, a value of a magnitude indicating a vital parameter. | 1. An apparatus for determining a physiological parameter of a living being, comprising: a receiver adapted to receive iterating optical signals, an optical signal comprising sequences, and a sequence comprising at least two bright periods during which a transmit light source adopts an on state and at least two dark periods during which no transmit light source adopts an on state, and the at least two bright periods being irregularly arranged within a sequence, andan extractor for extracting information on the physiological parameter from the receive signal, the extractor being adapted either to receive, to store or to generate binary code words comprising a length and indicating an arrangement of the bright and dark periods, and to link a binary code word with the signals in a block-by-block manner so as to obtain a digital signal of a bright transmit channel, the bright transmit channel being determined by the binary code word, and to link two other binary code words with the signals so as to obtain two digital signals of two dark transmit channels which are determined by the two other binary code words, and the extractor being adapted to extract, on the basis of the digital signal of the bright transmit channel and of the two digital signals of the dark transmit channels, a value of a magnitude indicating the physiological parameter, the extractor being implemented to form a differential signal from the two dark transmit channels, and to reduce an aliasing interference in the digital signal of the bright transmit channel by means of a spectral analysis of the differential signal. 2. The apparatus as claimed in claim 1, wherein the sequence of the optical signal is based on a clock in accordance with which the bright and dark periods occur, and the extractor comprises: a clock circuit implemented to receive the clock from a controller, to extract the clock from the receive signal, or to generate the clock by itself using a clock generator; andan analog/digital converter implemented to receive the clock signal from the clock circuit and to convert the receive signal to a digital receive signal in dependence on the clock. 3. The apparatus as claimed in claim 2, wherein the extractor is implemented to operate at a clock higher than 800 Hz. 4. The apparatus as claimed in claim 1, wherein the extractor is implemented to form a scalar product between the binary code word and a block of digital receive signal values of a length of the binary code word, to weight the result, to link it with other results and to associate an overall result with a transmit channel. 5. The apparatus as claimed in claim 1, wherein the extractor is implemented to subtract a values of a dark transmit channel from those of the bright transmit channel. 6. The apparatus as claimed in claim 1, wherein the sequence of the optical signal comprises at least two further bright periods originating from a second transmit light source, the extractor being implemented to extract a ... | Fraunhofer-Gesellschaft zur Foerderung der Angewandten Forschung E.V. | A61B005/02
US08849368 | B2 | 2014-09-30 | Implantable electronic devices for detecting hypoglycaemia using EEG signals | 2009-01-02 | Apparatus for capturing EEG signals comprising an implantable signal capture unit (34) connected to a proximal end of at least one implantable electrode (32) is implanted by passing a tubular guide sheath (24,26) along a subcutaneous path from an implantation site, passing the distal end of a said electrode through said guide sheath to bring said signal capture unit to said implantation site, and withdrawing said guide sheath from the electrode while progressively axially splitting the guide sheath to enable its removal. The signal capture unit may have a housing having at least two electrode contacts on its for capturing EEG signals in operative connection with circuitry within said housing for receiving said EEG signals and for wireless transmission thereof for reception at a location exterior to the body of the patient. | 1. Apparatus for use in preventing hypoglycaemia comprising an EEG signal based detector which detects the onset of hypoglycaemia on the basis of EEG signals, operatively connected to a generator for generating an alarm signal perceivable by a user of said apparatus upon the detection of the onset of hypoglycaemia, and further operatively connected to a pump for the administration of glucose or glucagon to a patient upon detection of the onset of hypoglycaemia, said pump being operatively connected to a reservoir containing glucose or glucagon for administration, said operative connection between the detector and the pump comprising a processor configured to initiate administration of said glucose or glucagon by said pump following a predetermined time lapse from the generation of a said alarm signal by said alarm signal generator, and comprising an input for a wireless cancellation signal by said user such that input of said wireless cancellation signal by the user prevents or halts actuation of said pump by said processor. 2. Apparatus as claimed in claim 1, wherein the EEG based detector comprises an implantable signal capture unit operatively connected to at least two EEG sensor electrodes adapted for subcutaneous implantation. | Hypo-Safe A/S | A61B005/04 A61B017/34 A61B005/0478 A61N001/00 A61M025/06
US08523773 | B2 | 2013-09-03 | Long term analyte sensor array | 2009-07-15 | A long term analyte sensor for measuring at least one analyte in the body of a user and which includes a housing, a plurality of analyte contacting sensor elements and at least one structure for relaying information away from the sensor. This plurality of analyte contacting sensor elements are typically disposed in an array. The analyte sensor further includes at least one sensor protection membrane that is controllable in a manner such that sensor elements may be activated (e.g. exposed to the external environment) at different times so as to extend the useful life of the sensor. In illustrative analyte sensors, the analyte is glucose. | 1. A method of sensing an analyte within the body of a mammal, the method comprising implanting an analyte sensor in to the mammal, the analyte sensor comprising: a plurality of analyte sensor elements adapted to contact and sense analyte, the analyte sensor elements comprising:a base layer;at least three working electrodes disposed on the base layer;a glucose oxidase layer disposed upon the at least three working electrodes, wherein the least three working electrodes are coated with the glucose oxidase layer so as to allow the analyte sensing device to sense glucose; an analyte modulating layer disposed on the glucose oxidase layer, wherein the analyte modulating layer comprises a hydrogel composition that includes a compound selected for its ability to crosslink to a siloxane composition;an adhesion promoting layer disposed between the glucose oxidase layer and the analyte modulating layer that functions to promote adhesion between the glucose oxidase layer and the analyte modulating layer, wherein the adhesion promoting layer comprises a siloxane composition crosslinked to the compound in the analyte modulating layer and;at least one analyte sensor membrane disposed upon at least one sealable well in a manner that reversibly prevents an analyte from contacting the at least one of the plurality of analyte sensors, wherein the permeability of the analyte sensor membrane can be controlled to allow an analyte to contact at least one of the plurality of analyte sensor elements; andat least one structure operatively coupled to the analyte sensing device for relaying information away from the analyte sensing device; andsensing an analyte within the body of a mammal. 2. The method of claim 1, further comprising controlling the analyte sensor membrane so that a second analyte sensor element in the plurality of analyte sensor elements contacts analyte after a first analyte sensor element in the plurality of analyte sensor elements exhibits a decrease in the ability to sense analyte due to biofouling, so that the useful life of the analyte sensing device is extended. 3. The method of claim 1, further comprising controlling the analyte sensor membrane so that a second analyte sensor element in the plurality of analyte sensor elements is adapted to contact analyte after a first analyte sensor element in the plurality of analyte sensor elements exhibits a decrease in the ability to sense analyte due to loss of activity of an analyte sensing enzyme disposed in the first analyte sensor element, so that the useful life of the analyte sensing device is extended. 4. The method of claim 1, wherein the analyte sensor membrane comprises a rupturable metallic membrane. 5. The method of claim 4, wherein the analyte sensor membrane hermetically seals the analyte sensor element. 6. The method of claim 1, wherein the analyte sensor membrane comprises a biodegradable polymer. 7. The method of claim 6, wherein the biodegradable polymer is comprised of materials sele... | Medtronic MiniMed, Inc. | A61B005/00
US09031640 | B2 | 2015-05-12 | Laser doppler blood flow measuring method and device | 2008-12-19 | [Subject] To provide laser Doppler blood flow measuring method and device which achieve multi-dimensional measurement efficiently at a high degree of accuracy over a wide range with a simple optical system and device.;[Solving Means] Laser light from a semiconductor laser 12 is split and formed into sheet lights Ls using a cylindrical lens 22, and the sheet lights Ls are crossed with each other at a predetermined position. A lens system 30 configured to form an image of scattered lights into a linear shape at a linear irradiation site Lx where the sheet lights Ls cross with each other is provided. An optical fiber array 32 having a plurality of optical fibers 34 is provided at an image-forming position of the lens system 30. Avalanche photodiodes 42 configured to convert the scattered lights which are shifted in frequency by the Doppler effect caused by the blood flow into electric signals for the each optical fiber 34 are provided. The blood flow velocity in a tissue of a biological body at the linear irradiation site Lx is calculated from the scattered lights transformed from the laser light by being shifted in frequency by the Doppler effect and is scanned for the each optical fiber 34, whereby the blood flow velocity in the blood vessel in a predetermined area in the biological body is calculated. | 1. A laser Doppler blood flow measuring method comprising: splitting infrared laser light emitted from a laser light source into split laser lights, transforming the split laser lights into thin sheet lights, and crossing the sheet lights with each other at a linear irradiation site at a position in a biological body, wherein the sheet lights are scattered into scattered lights at the linear irradiation site, and wherein among the scattered lights, scattered lights which are scattered by blood cells in a blood flow in a blood vessel in the biological body are shifted in frequency due to the Doppler effect, and scattered lights which are scattered by stationary tissue are not shifted in frequency due to the Doppler effect;receiving, by an optical system, the scattered lights from the linear irradiation site;emitting the scattered lights from the optical system such that the scattered lights are in a linear shape and are emitted onto a plurality of light-receiving elements which are arranged in a linear shape;performing photoelectric conversion of the scattered lights by a plurality of photoelectric conversion elements which correspond to the plurality of light-receiving elements to obtain electric signals corresponding to the scattered lights;processing the electric signals obtained by the photoelectric conversion, to obtain frequencies of the scattered lights;obtaining positional information about the blood vessel in the biological body from positions of ones of the plurality of light-receiving elements which receive the scattered lights which have been shifted in frequency by the Doppler effect; andcalculating a blood flow velocity in the blood vessel using Doppler frequency shifts of the scattered lights which were scattered by the blood cells in the blood flow;wherein the method further comprises scanning the sheet lights to move the linear irradiation site in two directions orthogonal to a longitudinal direction in which the linear irradiation site extends and along a direction parallel to the longitudinal direction, and obtaining information about the blood flow in the blood vessel in three dimensions using the frequencies of the scattered lights at a plurality of positions to which the linear irradiation site is moved by the scanning; andwherein the obtaining of the information about the blood flow in the blood vessel in three dimensions includes obtaining the positional information about the blood vessel and obtaining the blood flow velocity in the blood vessel at the plurality of positions to which the linear irradiation site is moved by the scanning. 2. The method according to claim 1, wherein the optical system is an image forming optical system which is positioned on a same side as a light-collecting optical system which applies the sheet lights to the linear irradiation site. 3. The method according to claim 1, wherein the optical system is an image forming optical system which is positioned on an opposite side of the tissue of the ... | Institute of National Colleges of Technology, Japan National Univeristy Corporation of Toyama | A61B005/02 A61B005/0285 A61B005/00 A61B005/026