## Sheet1
公告(開)號 | KindCode | 公告(開)日 | 名稱 | 申請日 | 摘要 | 申請範圍 | 專利權人 | IPC
US20100036264 | A1 | 2010-02-11 | DEVICE FOR COLLECTING PHYSIOLOGICAL INFORMATION OF AN ANIMAL, AND CORRESPONDING METHOD | 2007-07-11 | The device comprises:—a tight test enclosure (14) for receiving the animal (16),—a means (20, 21) for thermally regulating the test enclosure (14),—a means (27) for renewing air in the test enclosure (14) at a controlled flow rate,—a pressure sensor (41) for measuring pressure difference between the air in the test enclosure and a reference,—a means (43) for deducing a volume of air inhaled or expired by the animal during one inhalation or expiration from a pressure measurement of the pressure sensor (41), The test enclosure (14) comprises a floor having at least three dry electrodes, insulated from each other, each electrode being intended for collecting an electrical signal resulting from the cardiac activity of the animal (16) when the animal is in contact with the electrode. The device comprises a means (51) for determining an electrocardiogram of the animal (16) by using three electrical signals collected by the at least three electrodes. | 1. Device for collecting physiological information of an animal, in particular a young rodent such as a young mouse, the device comprising: a test tight enclosures for receiving the animal,a means for thermally regulating the test enclosures,a means for renewing air in the test enclosure at a controlled flow rate,a pressure sensor for measuring pressure difference between the air in the test enclosure and a reference,a means for deducing a volume of air inhaled or expired by the animal during one inhalation or expiration from a pressure measurement of the pressure sensor,the device being characterized in that: the test enclosures comprises a floor having at least three dry electrodes, insulated from each other, each electrode being intended for collecting an electrical signal resulting from the cardiac activity of the animal when the animal is in contact with the electrode,the device comprises a means for determining an electrocardiogram of the animal by using three electrical signals collected by the at least three electrodes. 2. Device according to claim 1, characterized in that the device comprises: a reference enclosure of identical dimension to the test enclosure, the reference enclosure being intended to remain empty,a means for thermally regulating the reference enclosure, in the same way as for the test enclosure,a means for renewing air in the reference enclosure at a controlled flow rate, in the same way as for the test enclosure,wherein the reference of the pressure sensor is the air of the reference enclosure. 3. Device according to claim 2, characterized in that the means for deducing the volume of air inhaled or exhaled by the animal is configured for using the Drorbaugh and Fenn equation. 4. Device according to claim 1, characterized in that it comprises a means which is connected to at least one electrode so as to send an electrical stimuli to the animal through the electrode. 5. Device according to claim 1, characterized in that the means for thermally regulating the test enclosure comprises a thermostatic bath in which the test enclosure is immerged. 6. Device according to claim 1, characterized in that the test enclosure comprises a window transparent to infrared radiation. 7. Device according to claim 6, characterized in that it comprises a infrared temperature sensor placed outside the test enclosure and adapted to be orientated toward the animal in the test enclosure through the window. 8. Device according to claim 1, characterized the device comprises an ultrasonic microphone disposed inside the test enclosure. 9. Device according to claim 1, characterized in that the test enclosure delimits a volume from 40 to 90 mliters. 10. Method for collecting physiological information of an animal, particularly a small rodent such as a small mouse, the method comprising the steps of: placing the animal in a closed test tight enclosure for receiving the animal, the test enclosure being thermally regulated and the air of the en... | - | A61B005/0205 A61B005/091
US20100010372 | A1 | 2010-01-14 | POROUS MULTIPLE SAMPLE SLEEVE AND BLOOD DRAWING DEVICE FOR FLASH DETECTION | 2009-05-11 | A device for drawing fluid from a lumen, and particularly blood from a blood vessel, is disclosed. The device may provide indication of the entry of an intravenous cannula into the lumen. The device may include a central body having an outer wall and an inner fluid passage. The device may include a luer-type adapter to permit the attachment of sample vial guide tube, or an I.V. infusion set of various lengths (“butterfly needle”) and or any luer-type fitting attached to an existing device. The outer wall of the central body of the flexible sleeve may be transparent or translucent to permit the detection of fluid within the inner passage. A front cannula may extend from one end of the central body and a rear cannula may extend from the other end of the central body. Both the front and rear cannulae may communicate with the inner fluid passage. An air-permeable blood-impermeable flexible sleeve may surround at least a tip portion of the rear cannula. The air-permeable blood-impermeable flexible sleeve acts as a venting member. The venting member permits the passage of air through it, and is substantially impermeable to liquids, such as blood. The flexible sleeve may be comprised of two types of material, a first material which is air-impermeable and a second material which is air-permeable blood-impermeable. | 1. A device for drawing fluid from a lumen, comprising: a central body having an outer wall and an inner fluid passage;a front cannula communicating with the inner fluid passage;a rear cannula communicating with the inner fluid passage; anda flexible sleeve made of air-impermeable material and a portion made of air-permeable blood-impermeable material, said sleeve surrounding at least a tip portion of the rear cannula. 2. (canceled) 3. (canceled) 4. (canceled) 5. (canceled) 6. (canceled) 7. (canceled) 8. The device of claim 1 wherein the air-permeable blood-impermeable portion is substantially porous for gas constituents less than about 5 microns in size, and substantially non-porous for liquid constituents about 5 microns and greater in size. 9. The device of claim 1 wherein the front cannula is incorporated into a butterfly needle. 10. (canceled) 11. (canceled) 12. A method of drawing blood from a lumen comprising the steps of: providing a blood drawing device including a rear cannula surrounded by an air-impermeable flexible sleeve having a portion made of air-permeable blood-impermeable material;inserting the blood drawing device into a lumen;receiving blood into the blood drawing device; andventing air through the air-permeable blood-impermeable portion of the flexible sleeve responsive to the receipt of blood in the device. 13. A flexible sleeve configured to receive a cannula of a blood drawing device located outside of a patient body, said sleeve comprising: an open end, a closed end made of air-impermeable material, and a portion made of air-permeable blood-impermeable material. 14. The sleeve of claim 13 wherein the air-permeable blood-impermeable portion is comprised of flexible material integrated into the sleeve. 15. The sleeve of claim 13 wherein the air-permeable blood impermeable portion is substantially porous for gas constituents less than about 5 microns in size, and substantially non-porous for liquid constituents about 5 microns and greater in size. 16. (canceled) 17. (canceled) 18. (canceled) 19. (canceled) 20. A method of making an air-permeable blood-impermeable flexible sleeve, comprising the steps of: providing a hydrophobic matrix material;mixing the hydrophobic matrix material with a hydrophilic porous material;forming the mixture of hydrophobic matrix material and hydrophilic porous material into a shape adapted to make up a portion of the flexible sleeve;drying the mixture of hydrophobic matrix material and hydrophilic porous material sufficiently to render air-permeable blood-impermeable material; andforming a flexible sleeve having a first portion made of air-impermeable material and a second portion made of the air-permeable blood-impermeable material. 21. The method of claim 20 wherein the hydrophilic porous material is polyacrylamide gel. 22. (canceled) 23. (canceled) 24. (canceled) 25. The sleeve of claim 1, wherein at least a portion of the sleeve is transparent or translucent. 26. The... | CLEARVIEW PATIENT SAFETY TECHNOLOGIES, LLC | A61B005/15 B29C071/00
US20100004521 | A1 | 2010-01-07 | IMPLANTABLE VOLTAIC CELL | 2007-05-07 | The invention provides a voltaic cell for implantation into the body of a subject that oxidizes oxidizable biological material to generate current to do work. The voltaic cell comprises a biologically inert shell having an inner compartment containing a cathodic environment and an anodic environment on the outer surface of the shell that is in contact with bodily fluid. A connector connects the anodic environment to the cathodic environment and has a component that provides resistance between the anodic and cathodic environments. The shell contains at least one salt bridge disposed within the shell that permits passage of small ions between the inner compartment and the bodily fluid, thereby completing the circuit. The invention also provides devices such as a glucometer which continuously detects glucose and transmits a signal to an external device which provides an output, such as blood glucose level. Methods of using the same are also provided. | 1. A voltaic cell for use in the body of a subject comprising: a biologically inert shell defining an inner compartment, said inner compartment containing an ionic solution that mediates a reduction reaction, said shell having an inner surface and an outer surface, said inner surface being in contact with said ionic solution; a cathodic environment within said shell in contact with said ionic solution;an anodic environment attached to said outer surface of said shell, in contact with bodily fluid wherein said anode accepts electrons from an oxidation reaction with at least one oxidizable biological material in said bodily fluid;a connector connecting said anodic environment and said cathodic environment;a component that provides resistance disposed along said connector between said anode and said cathode; andat least one salt bridge disposed within said shell permitting passage of ions between said inner compartment and said bodily fluid. 2. The voltaic cell of claim 1 wherein said cathodic environment comprises at least one cathode. 3. The voltaic cell of claim 1 wherein said anodic environment comprises at least one anode. 4. The voltaic cell of claim 2 wherein said anodic environment comprises at least one anode. 5. The voltaic cell of claim 1 wherein said oxidizable biological material is selected from the group consisting of carbohydrates, cholesterols, fatty acids, amino acids, polypeptides, lipids, hormones and polynucleotides. 6. The voltaic cell of claim 5 wherein said carbohydrate is glucose. 7. The voltaic cell of claim 5 wherein said hormone is selected from the group consisting of estrogen, progesterone, testosterone, growth hormone and thyroid hormone. 8. The voltaic cell of claim S wherein said polypeptide is prostate-specific antigen or a cancer-specific polypeptide. 9. The voltaic cell of claim 1 wherein said ionic solution comprises diamminesilver complex [Ag(NH3)2]+, silver, copper, or iron salts 10. The voltaic cell of claim 1 wherein said voltaic cell comprises a plurality of cathodes. 11. The voltaic cell of claim 1 wherein said voltaic cell comprises a plurality of anodes. 12. The voltaic cell of claim 11 wherein said voltaic cell comprises a plurality of cathodes. 13. The voltaic cell of claim 1 wherein said shell comprises a plurality of salt bridges. 14. The voltaic cell-of claim 1 wherein said connector is a wire. 15. The voltaic cell of claim 1 wherein said cathode is a non-magnetic metal. 16. The voltaic cell of claim 2 wherein said cathode is selected from the group consisting of silver, aluminum, lead, magnesium, platinum, gold, tin oxide, titanium dioxide, tungsten, non-magnetic alloys, semiconductors, semimetals and organometallic complexes. 17. The voltaic cell of claim 1 wherein said shell is formed from an impermeable material. 18. The voltaic cell of claim 1 wherein said shell is formed from a semi-permeable material. 19. The voltaic cell of claim 1 further comprising an inert barrier having ... | - | A61B005/1473
US20100022912 | A1 | 2010-01-28 | METHOD AND APPARATUS TO DETERMINE A PARAMETER PERTAINING TO THE PROSTATE OF A PATIENT | 2009-07-24 | In a method to determine at least one parameter pertaining to a prostate of a patient, an integrated diagnostic support system implements the following steps. A first risk value for prostate illness is determined using a screening method. In the event that this first characteristic risk value exceeds a first limit value, at least one second characteristic risk value for prostate illness and its location are determined using a method supplying a medical image of the patient. An integrated report system outputs at least one of the characteristic risk values and/or the location. A corresponding device has a mini-laboratory analysis unit for the screening test, an intelligent decision support system for the first characteristic risk value, a first interface with an imaging system, an evaluation system for the medical image for the second characteristic risk value, and a report system with a user interface. | 1. A method for determining at least one parameter pertaining to the prostate of a patient, comprising the steps of: in a single integrated diagnostic support system, determining a first risk value for the presence of a prostate illness using a screening method by interaction of the patient with the single integrated diagnostic support system;when said first characteristic risk value exceeds a limit value, determining at least one second characteristic risk value for the presence of a prostate illness at a suspected location of the prostate illness by generating a medical image of the prostate of the patient by interaction of the integrated diagnostic support system with the patient; andfrom said single integrated diagnostic support system, automatically generating a written report describing said first or second characteristic risk values and, if applicable, the location of the prostate illness. 2. A method as claimed in claim 1 wherein said limit value is a first limit value, and comprising, when said second characteristic risk value exceeds a second limit value, obtaining a tissue sample from the prostate of the patient by invasive interaction of the single integrated diagnostic support system with the patient. 3. A method as claimed in claim 1 comprising implementing the screening method with a mini-laboratory analysis unit in said single integrated diagnostic support system. 4. A method as claimed in claim 1 comprising automatically determining said first characteristic risk value using an IT-based decision support system in said single integrated diagnostic support system. 5. A method as claimed in claim 1 comprising determining said second characteristic risk value using prostate-specific variations. 6. A method as claimed in claim 1 comprising determining said second characteristic risk value as a local risk distribution using an automated evaluation system in said single integrated diagnostic support system. 7. A method as claimed in claim 1 comprising implementing said invasive method using said medical image for guidance. 8. A method as claimed in claim 1 comprising automatically determining said first characteristic risk value and said second characteristic risk value using a knowledge management system in said single integrated diagnostic support system. 9. An apparatus for determining at least one parameter pertaining to the prostate of a patient, comprising: a mini-laboratory analysis unit that interacts with a patient to conduct a prostate screening test on the patient;an artificially intelligent design support system that automatically determines a first characteristic value for the presence of a prostate illness using said screening test;a first interface connected to an imaging system that interacts with the patient to produce a medical image of the prostate of the patient;an automated evaluation system that determines at least one second characteristic risk value for the presence of the prostate illness, and a suspecte... | - | A61B010/00 A61B005/00 G06Q050/00
US20100036209 | A1 | 2010-02-11 | Circulatory monitoring systems and methods | 2008-08-07 | Systems and methods are described for obtaining and acting upon information indicative of circulatory health and related phenomena in human beings or other subjects. | 1. A medical or veterinary monitoring system comprising: circuitry for obtaining a local symptom of vascular occlusion; andcircuitry for selecting a first notification mode partly based on the local symptom of vascular occlusion and partly based on an additional indication of hemodynamic instability. 2. The medical or veterinary system of claim 1, in which the circuitry for selecting a first notification mode partly based on the local symptom of vascular occlusion and partly based on an additional indication of hemodynamic instability comprises: circuitry for deciding to use the first notification mode partly based on an apparent failure of a second notification mode. 3. The medical or veterinary system of claim 1, in which the circuitry for selecting a first notification mode partly based on the local symptom of vascular occlusion and partly based on an additional indication of hemodynamic instability comprises: circuitry for obtaining another local symptom of vascular occlusion as the additional indication of hemodynamic instability. 4. The medical or veterinary system of claim 1, in which the circuitry for selecting a first notification mode partly based on the local symptom of vascular occlusion and partly based on an additional indication of hemodynamic instability comprises: circuitry for causing a selection of the first notification mode responsive to one or more indications of clot movement. 5. The medical or veterinary system of claim 1, in which the circuitry for selecting a first notification mode partly based on the local symptom of vascular occlusion and partly based on an additional indication of hemodynamic instability comprises: circuitry for causing a selection of a higher-profile feature of the first notification mode in lieu of a lower-profile feature of a second notification mode. 6. The medical or veterinary system of claim 1, in which the circuitry for selecting a first notification mode partly based on the local symptom of vascular occlusion and partly based on an additional indication of hemodynamic instability comprises: circuitry for causing a selection of the first notification mode responsive to a subject-dependent profile. 7. The medical or veterinary system of claim 1, further comprising: circuitry for selecting a second notification mode in response to the additional indication of hemodynamic instability including one or more of an abnormally high heart rate measurement or an abnormal blood pressure measurement. 8. The medical or veterinary system of claim 1, in which the circuitry for obtaining a local symptom of vascular occlusion comprises: circuitry for obtaining an abnormal pressure measurement as the local symptom of vascular occlusion. 9. The medical or veterinary system of claim 1, in which the circuitry for obtaining a local symptom of vascular occlusion comprises: circuitry for obtaining an indication of a subject's local discomfort as the local symptom of vascular occlusion. 10. The medical or ve... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/026
US20100004525 | A1 | 2010-01-07 | MEDICAL CATHETER WITH SEVERAL POLES OR ELECTRODES | 2009-07-02 | A medical catheter (1) that has, on its distal end (2), several poles (3) or electrodes that are arranged axially one next to the other and that are used for stimulating and/or detecting potentials from an organ of the human body. A coil made from as many parallel running, electrically conductive windings as there are poles (3) is provided. The windings are bare, so that the entire cross section of the coil and thus of the catheter (1) can be kept small, and insulation (4), for example, in the form of a plastic cord, extends between the windings that form the poles (3). In this way, the outer dimensions of the catheter (1) are relatively small and, in addition, its distal end (2) has good flexibility. | 1. Medical catheter (1) comprising several poles (3) or electrodes arranged axially one next to the other or one behind the other for stimulating and/or detecting potentials from an organ of the human body, a coil formed from several electrically conductive windings extending parallel to each other, wherein each of the windings leads to one of the poles (3) arranged at a distal end (2) of the catheter (1), the electrically conductive windings for forming the poles (3) are bare windings at the distal end of the catheter and insulation (4) is arranged between the bare windings at least in a region of the poles (3) and the insulation is formed by windings that are made from insulating material or by wound plastic cords and that extend between the bare windings. 2. Catheter according to claim 1, wherein for forming the poles (3) by the windings, at least a last turn of the winding at the distal end is left bare at least partially or across an entire periphery of the coil or the windings for forming the poles (3) are left bare across a fraction of a length or across an entire length of the coil. 3. Catheter according to claim 1, wherein the wound insulation (4) arranged between the electrically conductive windings has a cross-sectional shape that is round. 4. Catheter according to claim 1, wherein the windings are formed from at least one of wires, multi-strand wires, flat wires, or round wires and shaped into a multi-turn coil. 5. Catheter according to claim 1, wherein the windings of the multi-turn coil have matching dimensions and extend coaxially one next to the other. 6. Catheter according to claim 1, further comprising an outer, insulating coating (5) that extends up to the distal end (2) of the coil with the poles (3), and openings (7) are located in the outer, insulating coating in a region of the bare windings for forming small surface-area poles (3). 7. Catheter according to claim 1, wherein the windings are coated or covered in a direction toward a proximal end of the catheter (1) with insulation such that the poles (3) do not repeat on a meridian line of the coil in the direction toward the proximal end. 8. Catheter according to claim 6, wherein the coating (5) has only one of the openings (7) on the periphery in the region of the bare windings for forming each of the poles (3) and the openings (7) are arranged coaxially one next to the other. 9. Catheter according to claim 1, further comprising electrically conductive or metallic rings (8) arranged coaxially on an outside of the bare windings forming the poles (3) on the distal end (2), and insulation (9) is arranged between the conductive or metallic rings and the bare windings, and one of the windings forming one of the poles (3) contacts a respective one of the rings (8). 10. Catheter according to claim 1, wherein a number of the poles is increased via additional conductors (10) or wires guided in an inner lumen of the windings of the coil, past the distal end of the coil to a... | - | A61B005/04 A61N001/05
US20100042009 | A1 | 2010-02-18 | DEVICE FOR USING SAVED FREQUENCY INFORMATION | 2007-12-31 | The invention relates to a device for applying saved frequency information in the form of modulated magnetic fields and/or modulated light to the body, wherein the device comprises at least one display, one memory, one control electronic, one D/A-converter, and a coil (6) for the application of modulated magnetic fields and/or a laser (7) for the application of modulated light. The invention also relates to a method for tapping signals of the electric potential of the heart and the pulse value and applying saved frequency information by means of a device according to the invention. | 1. A device for applying saved frequency information in the form of modulated magnetic fields and/or modulated light to the body wherein the device comprises at least one display, one memory, one control electronic, one D/A-converter and a coil for applying the modulated magnetic fields and/or a laser for applying modulated light. 2. The device as described in claim 1 wherein the device is a mobile computer equipped with a coil or a laser and a corresponding D/A-converter. 3. The device as described in claim 1 wherein the coil creates a magnetic field with a fluctuating frequency with a mean value of 9 Hz. 4. The device as described in claim 1 wherein the magnetic field is oriented symmetric transverse to the longitudinal axis of the device. 5. The device as described in claim 1 wherein the magnetic field created by the coil has a magnetic field strength of 0.1 to 70 μT at a measuring distance of 0 m. 6. The device as described in claim 1 wherein the laser is a semiconductor laser-diode which creates light with a fluctuating red wave length with a mean value of 650 nm. 7. The device as described in claim 1 that wherein the laser has a power of less than 1 mW. 8. The device as described in claim 1 that wherein the device comprises a transceiver for transmitting or receiving the frequency information to be applied. 9. The device as described in claim 1 wherein the frequency information to be applied is available in form of a digital file format. 10. The device as described in claim 1 wherein the device comprises in addition at least two, preferably three sensors for measuring the electric potential of the heart and at least one sensor, preferably in form of a photo transistor, for measuring the pulse value. 11. The device as described in claim 10 wherein three metal sensors in form of electrodes and a photo sensor are each arranged at the edge-portions of the device for the application of one finger at each sensor. 12. The device as described in claim 10 wherein two sensors are located on the front side of the device and two sensors are located opposite the first two sensors on the backside of the device. 13. The device as described in claim 10 wherein the control electronic of the device processes the signals received from the sensors for the computation of the heartrate-variability-heart-coherence-analysis and shows the result on the display. 14. The device as described in claim 13 wherein the result computed by the control electronic is compared with a database, which is located in the memory of the device, and that on the display of the device a corresponding selection of applicable frequency information and/or additional further information can be shown, where the shown selection of applicable frequency information is downloadable from the server by means of a transceiver or is stored in the memory of the device. 15. The device as described in claim 10 wherein the sensors, the memory, the control electronic, the D/A-converter, th... | - | A61B005/0402 A61N002/04 A61N005/06
US20100049059 | A1 | 2010-02-25 | APPARATUS AND METHOD FOR MEASURING BLOOD PRESSURE | 2009-07-28 | Disclosed are blood pressure measuring apparatus and method, in which a blood pressure measuring posture of a person to be examined is calculated on the basis of signals that are measured by an inclination measuring unit, so as to guide the person to be examined to maintain a reference measuring posture. When it is confirmed that the person to be examined maintains the reference measuring posture, blood pressure is measured on the basis of a measured living body signal of the person to be examined. By this disclosure, since the blood pressure measuring posture of the person to be examined can be correctly guided by inclination sensors, the person to be examined can accurately measure the blood pressure using a noninvasive and a nonpressurized method. | 1. A blood pressure measuring apparatus that measures blood pressure of a person to be examined, comprising: an inclination measuring unit that includes inclination sensors, and measures an inclination of the blood pressure measuring apparatus from signals that are detected by the inclination sensors;a living body signal measuring unit that receives a living body signal from the person to be examined and measures an electrocardiogram and a pulse wave; anda central processing unit that calculates a blood pressure measuring posture of the person to be examined on the basis of the signals measured by the inclination measuring unit, determines whether the person to be examined maintains a predetermined reference measuring posture by comparing the calculated blood pressure measuring posture of the person to be examined and the predetermined reference measuring posture, and measures the blood pressure of the person to be examined on the basis of the living body signal that is measured by the living body signal measuring unit. 2. The blood pressure measuring apparatus of claim 1, wherein the inclination sensors are acceleration sensors. 3. The blood pressure measuring apparatus of claim 1, wherein the inclination sensors are three-axis gravity acceleration sensors. 4. The blood pressure measuring apparatus of claim 1, wherein the inclination measuring unit measures an elevation angle that is an angle between the surface of the earth and the blood pressure measuring apparatus and a rotation angle of the blood pressure measuring apparatus, on the basis of the signals that are detected by the inclination sensors. 5. The blood pressure measuring apparatus of claim 4, wherein the central processing unit calculates a height difference between the heart of the person to be examined and the blood pressure measuring apparatus on the basis of the elevation angle and height information of the person to be examined, and calculates the blood pressure measuring posture of the person to be examined on the basis of the calculated height difference between the heart of the person to be examined and the blood pressure measuring apparatus. 6. The blood pressure measuring apparatus of claim 4, wherein the central processing unit sets a reference rotation angle on the rotation angle, compares the reference rotation angle and the measured rotation angle, and calculates the blood pressure measuring posture of the person to be examined according to whether the measured rotation angle corresponds to the reference rotation angle. 7. The blood pressure measuring apparatus of claim 4, wherein the central processing unit outputs data that is used to guide the person to be examined to maintain the reference measuring posture, when it is determined that the person to be examined does not maintain the reference measuring posture on the basis of the elevation angle and the rotation angle that are measured by the inclination measuring unit. 8. The blood pressure measuring apparat... | - | A61B005/02 A61B005/0402
US07650182 | B2 | 2010-01-19 | Dual sensing for brady-tachy pacemaker/ICD | 2005-07-08 | A system detects events related to cardiac activity. The system comprises a primary cardiac signal sensing circuit, at least one secondary cardiac signal sensing circuit having a higher sensitivity than the primary sensing circuit, and a controller circuit coupled to the primary and secondary cardiac signal sensing circuits. The controller circuit determines a rate of depolarization using the primary sensing circuit and detects tachyarrhythmia using the rate. The controller circuit also detects tachyarrhythmia using the secondary sensing circuit and also deems the tachyarrhythmia valid if the controller circuit detects the tachyarrhythmia using both the primary and secondary sensing circuit. | 1. A system comprising: a primary cardiac signal sensing circuit operable to sense at least a first intrinsic cardiac signal, wherein the primary cardiac signal sensing circuit is configured to be coupled to a first and a second implantable electrode and has a first signal sensing sensitivity configured to detect bradycardia and a second higher sensitivity to simultaneously detect tachyarrhythmia;at least one secondary cardiac signal sensing circuit operable to sense at least a second intrinsic cardiac signal, wherein the secondary cardiac signal sensing circuit is configured to be coupled to at least a third implantable electrode different from the first and second electrode, has a higher signal sensing sensitivity than the primary sensing circuit, and is configured to detect ventricular tachycardia; anda controller circuit, coupled to the primary and secondary cardiac signal sensing circuits, the controller circuit operable to: detect tachyarrhythmia and bradycardia using the primary sensing circuit;detect tachyarrhythmia using the secondary sensing circuit; anddeem the tachyarrhythmia valid if the controller circuit detects the tachyarrhythmia using both the primary and secondary sensing circuits. 2. The system of claim 1, wherein the controller circuit is operable to change the first sensitivity according to depolarization events sensed by the secondary sensing circuit. 3. The system of claim 1, wherein the controller circuit is operable to use the secondary sensing circuit to detect the tachyarrhythmia by shifting the second cardiac signal in time and multiplying the second cardiac signal and the first cardiac signal. 4. The system of claim 3, wherein the controller circuit is further operable to shift a cardiac signal in time to maximize an absolute peak of a multiplied signal. 5. The system of claim 1, wherein the secondary sensing circuit includes at least one electrode operable to deliver high energy shock therapy. 6. The system of claim 1, wherein the primary sensing circuit includes an electrode operable to sense intrinsic cardiac signals associated with a first ventricle and the secondary sensing circuit includes an electrode operable to sense intrinsic cardiac signals associated with a second ventricle. 7. The system of claim 1, wherein the primary sensing circuit includes a tip electrode and a ring electrode of a cardiac lead and the secondary sensing circuit includes a can electrode and one of a ring electrode or a tip electrode of another cardiac lead. 8. The system of claim 1, wherein the primary sensing circuit includes a tip electrode and a ring electrode of a cardiac lead and the secondary sensing circuit includes a can electrode and a shock therapy coil electrode. 9. The system of claim 1, wherein the secondary sensing circuit includes electrodes adaptable for placement across a myocardium of a patient, and wherein the controller circuit is operable to measure mechanical action of the heart using impedance signals sen... | Cardiac Pacemakers, Inc. | A61B005/0464
US07665615 | B2 | 2010-02-23 | Composite article having hydrophilic properties and method of manufacture | 2005-09-30 | A composite article includes a base material and a porous membrane laminated with the base material. The porous membrane has hydrophobic properties and includes at least one of expanded polytetrafluoroethylene, woven polytetrafluoroethylene, and non woven polytetrafluoroethylene. A coating layer is formed on at least a portion of the porous membrane. The coating layer has hydrophilic properties and includes at least one of an organofunctional siloxane and a polyether urethane polymer. | 1. A composite article comprising: a base material;a porous membrane laminated with the base material, said porous membrane having hydrophobic properties and comprising at least one of expanded polytetrafluoroethylene, woven polytetrafluoroethylene, and non woven polytetrafluoroethylene; anda coating layer formed on at least a portion of said porous membrane, said coating layer having hydrophilic properties and comprising at least one of an organofunctional siloxane and a polyether urethane polymer, wherein said composite article retains water wetability and can be dried and subsequently flow liquid without pre-wetting, and comprises a water permeability of between about 5 and about 60 milliliters per minute per square centimeter at about 27 inches of mercury vacuum. 2. A composite article in accordance with claim 1 wherein said base material comprises a material selected from a group of materials comprising polypropylene, polyethylene, polyester, acrylic, polyphenelene sulfide, aramid, polyimide, glass, and/or polytetrafluoroethylene, said base material being one of woven and non woven. 3. A composite article in accordance with claim 1 wherein said porous membrane comprises a weight of between about 0.05 and about 1 ounce per square yard. 4. A composite article in accordance with claim 3 wherein said porous membrane comprises a weight of between about 0.1 and about 0.5 ounces per square yard. 5. A composite article in accordance with claim 1 wherein said coating layer is formed from a coating composition comprising between about 0.1 and about 10 wt % solids based on the total weight of said coating composition. 6. A composite article in accordance with claim 5 wherein said coating composition comprises between about 1.0 and about 3.0 wt % solids based on the total weight of said coating composition. 7. A composite article in accordance with claim 1 wherein said coating layer is formed from a coating composition comprising a solvent having a surface tension of less than about 31 dynes per centimeter. 8. A composite article in accordance with claim 1 wherein said coating layer is formed from a coating composition comprising at least one solvent selected from a group of solvents comprising water, isopropyl alcohol, ethanol, and acetone. 9. A composite article in accordance with claim 1 wherein said composite article comprises an air permeability of between about 0.01 and about 10 CFM per square foot as measured in accordance with ASTM D737. 10. A composite article in accordance with claim 9 wherein said composite article comprises an air permeability of between about 0.75 and about 2.25 CFM per square foot as measured in accordance with ASTM D737. 11. A composite article in accordance with claim 1 wherein said composite article remains hydrophilic after at least three wet then dry cycles. 12. A composite article in accordance with claim 1 wherein said composite article comprises a water permeability of between about 45 and about 55 milli... | General Electric Company | B01D029/00 B01D009/46 B01D071/26 A61B005/103 B01D061/00
US20100036471 | A1 | 2010-02-11 | Method of Diagnosing and Treating Lower Urinary Tract Symptoms | 2008-08-05 | A method of diagnosing and treating a patient having lower urinary tract symptoms caused by insufficient blood flow to the urinary bladder, the urethra, a nerve innervating the urinary bladder, or a nerve innervating the urethra due to atherosclerosis of a pelvic artery is disclosed. A method of diagnosing the patient's condition includes determining if a stenosis exists within a pelvic vessel. A method of treating the patient's condition may include placing a stent within the stenosed pelvic artery. | 1. A method of diagnosing lower urinary tract symptoms correctable by arterial stenting, the method comprising: identifying a stenosis in a pelvic artery that is restricting blood flow to an artery that supplies blood to one of the urinary bladder, the urethra, a nerve innervating the urinary bladder, and a nerve innervating the urethra. 2. The method of claim 1, wherein the step of identifying a stenosis in a pelvic artery includes performing a pelvic angiogram. 3. The method of claim 1, wherein the step of identifying a stenosis in a pelvic artery includes performing a diagnostic procedure selected from the group consisting of a CT scan, a MRI, CT-angiography, MR-angiography, doppler sonography, color duplex sonography, and nuclear imaging. 4. A method of treating lower urinary tract symptoms caused by atherosclerosis in a pelvic artery that supplies blood to one of the urinary bladder, the urethra, a nerve innervating the urinary bladder, and a nerve innervating the urethra, the method comprising: tracking a delivery catheter having a stent to a stenosis within the artery; andrestoring blood flow through the artery by expanding the stent within the stenosis. 5. The method of claim 4, wherein the pelvic artery is one of the internal iliac artery, external iliac artery, and common iliac artery. 6. The method of claim 4, wherein the pelvic artery is one of the superior vesical artery, the umbilical artery, the vaginal artery, the obturator artery, and the inferior gluteal artery. 7. The method of claim 4, wherein the pelvic artery is one of the internal pudendal artery, the inferior vesical artery, and the middle rectal artery. 8. The method of claim 4, wherein the pelvic artery is one of a prostatic artery, the artery to the urethral bulb, and the urethral artery. 9. The method of claim 4, wherein pelvic artery supplies blood to one of the vesical nerve plexus, the pudendal nerve, and the cavernous nerve plexuses. 10. The method of claim 4, wherein the stent used in the step of restoring blood flow through the artery is balloon expandable. 11. The method of claim 4, wherein the stent used in the step of restoring blood flow through the artery is self-expanding. 12. The method of claim 4, wherein the stent used in the step of restoring blood flow through the pelvic artery is a drug-eluting stent. 13. A method of treating lower urinary tract symptoms caused by atherosclerosis in a pelvic artery, the method comprising: tracking a revascularization tool to a stenosis within the pelvic artery; andperforming a revascularization procedure with the revascularization tool within the pelvic artery to substantially eliminate the stenosis and restore flow through the pelvic artery. 14. The method of claim 13, wherein the revascularization procedure includes rotoblading. 15. The method of claim 13, wherein the revascularization tool includes a vibrating portion for breaking-up the stenosis via mechanical vibration. 16. The method of claim 13, ... | Medtronic Vascular, Inc. | A61F002/84 A61B005/05
US20100023348 | A1 | 2010-01-28 | REMOTELY TAKING REAL-TIME PROGRAMMATIC ACTIONS RESPONSIVE TO HEALTH METRICS RECEIVED FROM WORN HEALTH MONITORING DEVICES | 2008-07-22 | Health metrics can be received from at least one person wearing a health monitoring device in an unencumbered fashion that permits free motion. The health monitoring device can wirelessly convey the heath metrics to a remotely located computing device. The received health metrics can be compared against at least one configurable, yet previously established threshold. A change in a situation proximate to the person can be inferred based upon comparison results. A programmatic event can be fired based upon the inferred change. At least one programmatic action can be automatically initiated responsive to the firing of the programmatic event, wherein the programmatic action initiates a response to the inferred change. | 1. A method for utilizing health metrics comprising: receiving health metrics from at least one person wearing a health monitoring device in an unencumbered fashion that permits free motion, wherein the health monitoring device wirelessly conveys the heath metrics to a remotely located computing device;comparing the received health metrics against at least one configurable, yet previously established threshold;inferring a change in a situation proximate to the at least one person based upon comparison results;firing a programmatic event based upon the inferred change; andinitiating at least one programmatic action responsive to the firing of the programmatic event, wherein the programmatic action initiates a response to the inferred change. 2. The method of claim 1, wherein the receiving and comparing of metrics, the inferring the change, firing the programmatic event, and the initiating at least one programmatic action occur in real-time. 3. The method of claim 1, wherein the at least one person wearing the health monitoring device comprises a plurality of people, said method further comprising: comparing the received health metrics of the plurality of people against the threshold to generate the comparison results from which the change is inferred. 4. The method of claim 1, wherein the inferred change indicates that a wearing of the health monitoring device is in need of help, wherein said programmatic action automatically conveys a notification message to a responder that indicates the need of the wearer, wherein the conveyance of the notification message occurs without the wearing taking an active action. 5. The method of claim 1, wherein the inferred change indicates a change in a user state, wherein user states able to be inferred comprise unconsciousness, death, panic, and normality. 6. The method of claim 1, wherein a plurality of different people wear the health monitoring device concurrently participate in a common endeavor, said programmatic action comprising a notification action, said method further comprising: executing the notification action to automatically notify each of the different people of the inferred change, wherein the different people are notified if the inferred change in at least one of real-time and near real time. 7. The method of claim 1, a plurality of agents each wearing the health monitoring device, wherein the inferred change in the situation indicates a potential breach of a security perimeter established by said agents. 8. The method of claim 1, a plurality of agents each wearing the health monitoring device when deployed to a potentially hostile situation, wherein the inferred change is configured to immediate provide notification to at least one designated person when any of the plurality of agents is rendered unconscious and is configured to immediately provide a notification to at least one designated person when any of the plurality of agents becomes dead, wherein the provided notification occurs ... | INTERNATIONAL BUSINESS MACHINES CORPORATION | G06Q050/00 A61B005/00
US20100036289 | A1 | 2010-02-11 | Apparatus and Methods for Assessment of Vestibulo-Ocular Reflexes | 2008-01-15 | A rotary disk securely rotates a patient for the purpose of inducing vestibulo-ocular reflex eye movements. The apparatus provides a rotating platform with a positioning post to which a patient can be secured with a safety vest. A headrest on the positioning post permits a patient's head to be positioned at an angle conducive to inducing reflex eye movements. A camera or other recording device is used to view and record eye movements before, during and/or after a patient is subjected to a period of rotation. | 1. A rotary disk device for rotating a patient comprising: a base;a platform rotatably connected to the base and having a single axis of rotation;a positioning post attached to and rotatable with the platform;a headrest attached to the positioning post; anda means for viewing and recording patient eye movements, adjustably attached to the rotary disk. 2. The device, according to claim 1, wherein the positioning post is generally vertical and attached to the platform at or near the axis of rotation. 3. The device, according to claim 1, wherein the headrest is tilted such that a patient's head positioned therein is at a downward forward-facing angle. 4. The device, according to claim 1, wherein the headrest is tilted at an angle such that a patient's head positioned therein is at a downward forward-facing angle of approximately 20° to approximately 40° from the horizontal. 5. The device, according to claim 1, wherein the means for viewing and recording is a video camera. 6. The device, according to claim 1, wherein the means for viewing and recording is adjustably attached to the positioning post. 7. The device, according to claim 1, further comprising a safety vest for securing a patient against the positioning post. 8. A rotary disk device for rotating a patient comprising: a base;a platform rotatably connected to the base and having a single axis of rotation;a motor operably connected to the platform that rotates the platform about the axis of rotation;a positioning post fixedly attached to and rotatable with the platform;a headrest fixedly attached to the positioning post; anda means for viewing and recording patient eye movements, adjustably attached to the rotary disk. 9. The device, according to claim 8, wherein the positioning post is attached to the platform at or near the axis of rotation. 10. The device, according to claim 8, wherein the headrest is tilted such that a patient's head positioned therein is at a downward forward-facing angle. 11. The device, according to claim 10, wherein the headrest is tilted at an angle such that a patient's head positioned therein is at a downward forward-facing angle of approximately 20° to approximately 40° from the horizontal. 12. The device, according to claim 8, wherein the means for viewing and recording is a video camera. 13. The device, according to claim 8, wherein the means for viewing and recording is adjustably attached to the positioning post. 14. The device, according to claim 8, further comprising a safety vest for securing a patient against the positioning post. 15. The device, according to claim 8, further comprising a remote device for controlling the motor. 16. A method for eliciting a vestibulo-ocular reflex in a patient utilizing a device comprising: a base;a platform rotatably connected to the base and having a single axis of rotation;a positioning post fixedly attached to and rotatable with the platform;a headrest fixedly attached to the positioning post;a safety ve... | UNIVERSITY OF FLORIDA RESEARCH FOUNDATION, INC. | A61B005/11
US20100041978 | A1 | 2010-02-18 | CONTROLLING TELEMETRY DURING MAGNETIC RESONANCE IMAGING | 2009-10-20 | The invention is directed techniques for coordinating telemetry of medical devices with magnetic resonance imaging (MRI) techniques. By coordinating telemetry of a medical device with the performance of MRI techniques with, the use of telemetry during MRI may be facilitated. In one example, information indicative of electromagnetic radiation bursts in MRI techniques can be communicated to the medical device prior to execution. In another example, the medical device may identify the electromagnetic radiation bursts, e.g., by measuring for the presence of such bursts. In either case, the medical device can adjust its telemetry to improve communication during MRI. | 1. A method of coordinating a medical device with magnetic resonance imaging (MRI), the method comprising: identifying information associated with one or more MRI electromagnetic radiation bursts; andadjusting telemetry of a medical device during the electromagnetic radiation bursts based on the information. 2. The method of claim 1, wherein identifying information comprises receiving the information from an MRI device. 3. The method of claim 1, wherein identifying information comprises identifying timing information indicative of a timing of one or more electromagnetic radiation bursts. 4. The method of claim 3, wherein the timing information includes a start time of at least one burst interval, a duration of the burst interval(s), and a time interval between successive burst intervals. 5. The method of claim 1, wherein adjusting telemetry comprises disabling one or more components of a telemetry unit during the electromagnetic radiation bursts. 6. The method of claim 1, wherein adjusting telemetry comprises increasing power of telemetry signals during the electromagnetic radiation bursts. 7. The method of claim 1, wherein adjusting telemetry comprises selecting a packet size for effective communication between burst intervals. 8. The method of claim 1, further comprising: identifying information associated with one or more MRI magnetic gradients; and adjusting telemetry of a medical device during application of the MRI magnetic gradients. 9. A medical device comprising: a telemetry unit to perform telemetry with another device; anda control unit to identify information associated with one or more magnetic resonance imaging (MRI) electromagnetic radiation bursts and adjust the telemetry during the MRI electromagnetic radiation bursts. 10. The medical device of claim 9, wherein the telemetry unit receives information from an MRI device indicative of the electromagnetic radiation bursts, and wherein the control unit adjusts telemetry based on the information. 11. The medical device of claim 9, wherein the information comprises timing information indicative of a timing of one or more electromagnetic radiation bursts. 12. The medical device of claim 11, wherein the timing information includes a start time of at least one burst interval, a duration of the burst interval(s), and a time interval between successive burst intervals. 13. The medical device of claim 9, wherein the control unit adjusts telemetry by disabling one or more components of the telemetry unit during the electromagnetic radiation bursts. 14. The medical device of claim 9, wherein the control unit adjusts telemetry by increasing power of telemetry signals during the electromagnetic radiation bursts. 15. The medical device of claim 9, wherein the control unit adjusts telemetry by selecting a packet size for effective communication between burst intervals. 16. The medical device of claim 9, wherein the medical device comprises one of a programmer, an implantable medical... | Medtronic, Inc. | A61B005/055 A61N001/08
US07654969 | B2 | 2010-02-02 | Integrated spot monitoring device with fluid sensor | 2008-02-05 | A bodily fluid sampling device is operable to breach the skin surface and allow bodily fluid to emerge from the breach location. The bodily fluid sampling device further evaluates the amount of bodily fluid emerged from breach location and determines whether the amount is sufficient or insufficient for a particular purpose, such as sampling and testing. The determination is accomplished automatically without moving the device. The user may also intervene to perform a variety of tasks following the determination of the amount of bodily fluid. | 1. A body fluid sampling device, comprising: a meter;a lancing device configured to create an incision in a surface of tissue, wherein the lancing device includes a lancet;a fluid detector configured to detect sufficiency of body fluid bled from the incision while on the surface of the tissue;an analysis device configured to analyze the body fluid once the fluid detector detects that the body fluid on the surface of the tissue is sufficient for analysis, wherein the analysis device includes a test strip;an integrated lancing test strip including the lancet attached to the test strip;the integrated lancing test strip being detachably coupled to the meter, the lancet and test strip of the integrated lancing test strip being configured to be removed together from the meter as a single unit;wherein the fluid detector includes one or more sensors disposed on the integrated lancing test strip to detect sufficiency of the body fluid bled from the incision while on the surface of the tissue;the meter including a housing in which the fluid detector is disposed, the housing defining a cavity with an opening where the incision is created; andwherein the fluid detector includes an ultrasonic detector configured to detect a change in resonance of the cavity in the housing as the body fluid from the incision fills the cavity. 2. The body fluid sampling device of claim 1, further comprising: a controller operatively coupled to the lancing device and a firing mechanism; andan output device operatively coupled to the controller, wherein the output device is configured to positively indicate insufficiency of the body fluid when the fluid detector detects the insufficiency of the body fluid. 3. The body fluid sampling device of claim 1, wherein: the test strip defines a capillary channel with a capillary channel opening for drawing the body fluid; andthe one or more sensors are disposed proximal the capillary channel opening for locating the capillary channel opening during drawing of the body fluid. 4. The body fluid sampling device of claim 1, wherein the analysis device includes a test strip with a capillary channel for drawing the body fluid and an actuator coupled to the test strip for positioning the capillary channel to collect the body fluid when the fluid detector detects a sufficient amount of the body fluid on the surface of the tissue. 5. The device of claim 1, further comprising: a controller;an expression device operatively coupled to the fluid detector through the controller; andthe controller being configured to activate the expression device to automatically express the body fluid from the incision upon detecting with the fluid detector that the body fluid on the surface of the tissue is insufficient for analysis. 6. The body fluid sampling device of claim 5, further comprising: the housing having a skin contacting portion around the opening that contacts the tissue during creation of the incision; andwherein the expression device is configure... | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00
US07657292 | B2 | 2010-02-02 | Method for evaluating extracellular water concentration in tissue | 2005-11-18 | The present disclosure provides systems, devices, and/or methods for assessing body fluid-related metrics and/or changes therein. The disclosure further provides systems, devices, and/or methods for correlating body fluid-related metrics in a particular tissue with the corresponding whole-body metric. The disclosure also provides, systems, devices, and/or methods for assessment of such metrics to facilitate diagnosis and/or therapeutic interventions related to maintaining and/or restoring body fluid balance. | 1. A method for assessing an extracellular water metric in a subject, said method comprising: using at least one sensor: emitting light of at least one wavelength toward a tissue site of interest, wherein a portion of the emitted light is reflected by the tissue site of interest;detecting at least one wavelength of the light reflected by the tissue site of interest from a position about one millimeter to about five millimeters from the tissue site of interest;emitting light of at least one wavelength toward a reference tissue site, wherein a portion of the emitted light is reflected by the reference tissue site;detecting at least one wavelength of the light reflected by the reference tissue site from a position about one millimeter to about five millimeters from the reference tissue site; andusing a processor: processing detected reflected light from the tissue site of interest and detected reflected light from the reference tissue site to compute the extracellular water metric, wherein the hydrostatic pressure of the tissue site of interest differs from the hydrostatic pressure of the reference tissue site. 2. A method according to claim 1, wherein the extracellular water metric comprises intravascular water fraction, fwIV, interstitial water fraction fwIS, extracellular water fraction, fwEC, or water balance index, Q or combinations thereof. 3. A method according to claim 1, wherein the extracellular water metric comprises changes in the intravascular water fraction, ΔfwIV, changes in the interstitial water fraction, ΔfWIS, changes in the extracellular water fraction, ΔfWEC, or changes in the water balance index, ΔQ or combinations thereof. 4. A method according to claim 1, wherein the extracellular water metric is a metric of at least a portion of the subject's whole body. 5. A method according to claim 4 further comprising comparing the extracellular water metric with a control and determining from that comparison whether the at least a portion of the subject's whole body is dehydrated or over-hydrated. 6. A method according to claim 5 further comprising administering a hydration therapy to a dehydrated subject. 7. A method according to claim 5 further comprising administering a dehydration therapy to an over-hydrated subject. 8. A method according to claim 1 further comprising exhibiting the extracellular water metric on a display. 9. A method according to claim 1, wherein detecting at least one wavelength of the light reflected by the tissue site of interest comprises detecting light comprising a wavelength of 950-1400 nm, 1500-1800 nm, or 2000-2300 nm or any combination thereof. 10. A method according to claim 9 wherein the light comprises a wavelength of 1180 nm, 1125 nm, 1168 nm, 1170 nm, 1185 nm, 1190 nm, 1230 nm, 1245 nm, 1250 nm, 1274 nm, 1275 nm, 1300 nm, 1330 nm, 1710 nm, 1730 nm, or 1740 nm, or any combination thereof. 11. A method according to claim 9, wherein detecting at least one wavelength of the light reflected by t... | Nellcor Puritan Bennett LLC | A61B005/145
US07665364 | B2 | 2010-02-23 | Method and apparatus for remote sensing of molecular species at nanoscale utilizing a reverse photoacoustic effect | 2006-12-14 | A method and apparatus for identifying a sample, involves illuminating the sample with light of varying wavelengths, transmitting an acoustic signal against the sample from one portion and receiving a resulting acoustic signal on another portion, detecting a change of phase in the acoustic signal corresponding to the light of varying wavelengths, and analyzing the change of phase in the acoustic signal for the varying wavelengths of illumination to identify the sample. The apparatus has a controlled source for illuminating the sample with light of varying wavelengths, a transmitter for transmitting an acoustic wave, a receiver for receiving the acoustic wave and converting the acoustic wave to an electronic signal, and an electronic circuit for detecting a change of phase in the acoustic wave corresponding to respective ones of the varying wavelengths and outputting the change of phase for the varying wavelengths to allow identification of the sample. The method and apparatus can be used to detect chemical composition or visual features. A transmission mode and a reflection mode of operation are disclosed. The method and apparatus can be applied at nanoscale to detect molecules in a biological sample. | 1. A method for identifying a sample, the method comprising: exposing the sample to electromagnetic radiation of varying wavelengths to cause absorption of electromagnetic radiation in the sample;during said exposure, transmitting an acoustic wave onto the sample from one position and receiving the acoustic wave at another position;detecting a change of phase in the acoustic wave induced by the absorption of the electromagnetic radiation of varying wavelengths; andanalyzing the change of phase for the electromagnetic radiation of varying wavelengths to identify the sample. 2. The method of claim 1, wherein the sample is identified as to chemical composition. 3. The method of claim 1, wherein the sample is identified as to a character of being black or white. 4. The method of claim 1, wherein the sample is identified as to a character of being absorbing or non-absorbing. 5. The method of claim 1, wherein portions of the sample are identified as to color. 6. The method of claim 1, wherein an acoustic transmitter and an acoustic receiver are opposite each other, but offset from the sample by a distance so as to form an angle of incidence to the sample and an angle of reflection from the sample, and wherein the acoustic wave is transmitted to the sample from the acoustic transmitter and then reflected in a reflection mode to the acoustic receiver. 7. The method of claim 6, wherein the angle of incidence is equal to the angle of reflection and is approximately eighty degrees. 8. The method of claim 1, wherein an acoustic transmitter and an acoustic receiver are directly opposite each other; wherein the sample is midway between the acoustic transmitter and the acoustic receiver, and wherein the acoustic wave is transmitted in a transmission mode through the sample from the acoustic transmitter to the acoustic receiver. 9. The method of claim 1, wherein illuminating the sample with light of varying wavelengths further comprises generating a beam of light and passing said beam of light through a plurality of filters before the light contacts the sample. 10. The method in claim 1, wherein the sample is a human body part, and wherein the acoustic signal passes to or is reflected to an acoustic detector from the human body part, wherein the human body part is radiated with electromagnetic waves of varying frequencies for producing the absorption spectra of physiologically relevant materials such as glucose. 11. The method in claim 1, wherein the method is carried out to sense points of the sample within nanoscale dimensions, and wherein individual molecules are identified. 12. An apparatus for identifying a sample, the apparatus comprising: a controllable source of electromagnetic radiation for exposing the sample to electromagnetic radiation of varying wavelengths to cause absorption of electromagnetic radiation in the samplea transmitter for transmitting an acoustic wave onto the sample from a first position;a receiver for receiving the acousti... | UT-Battelle, LLC | G01N009/24 A61B005/00
US07662107 | B2 | 2010-02-16 | Use of 1,1,1,2-Tetrafluoroethane for measuring lung function, especially for determining the functional residual capacity (FRC) of the lungs | 2006-08-14 | 1,1,1,2-Tetrafluoroethane (norflurane) is used as a tracer/trace gas for measuring the lung function because of its simple detectability, physiological harmlessness, environmental friendliness and ready availability. Especially advantageous is the possibility of the optical concentration measurement of 1,1,1,2-tetrafluoroethane for the measurement of the lung function and especially for the determination of the FRC during respiration during anesthesia, without cross sensitivity occurring during the measurement with respect to anesthetics. | 1. A method comprising: providing a gas including a breathing gas with 1,1,1,2-tetrafluoroethane as a trace gas;delivering said breathing gas with the 1,1,1,2-tetrafluoroethane trace gas to the lungs of a subject via one of an anesthesia apparatus and a respirator; andmeasuring lung function of the subject during inhalation with one of said anesthesia apparatus and said respirator based on the 1,1,1,2-tetrafluoroethane trace gas, wherein said step of measuring lung function includes determining the functional residual capacity (FRC) of the lungs. 2. A method according to claim 1, wherein FRC of the lungs is determined by washout. 3. A method according to claim 1, wherein the 1,1,1,2-tetrafluoroethane is a component of an inhalation composition used during respiration/anesthesia. 4. A method according to claim 3, wherein the inhalation composition includes at least one drug. 5. A method according to claim 1, wherein said gas includes said breathing gas with 1,1,2,3,3,3,-heptafluoropropane as another trace gas, wherein the quantity ratio of the two trace gases 1,1,1,2,3,3,3-heptafluoropropane:1,1,1,2-tetrafluoroethane ranges from 1:2 to 1:3 as the maximum. 6. A method according to claim 1, wherein concentration during respiration of the trace gas 1,1,1,2-tetrafluoroethane is about 1 vol. % to 5 vol. %, maximum of the inspired breathing gas volume present in the lungs. 7. A method according to claim 6, wherein concentration during respiration of the trace gas 1,1,1,2-tetrafluoroethane is 0.1 vol. % to 1 vol. % of the inspired breathing gas volume present in the lungs. 8. A method according to claim 1, wherein said gas includes said breathing gas with 1,1,2,3,3,3-heptafluoropropane as another trace gas, wherein concentration during respiration of the trace gas 1,1,1,2-tetrafluoroethane combined with 1,1,1,2,3,3,3-heptafluoropropane is about 0.1 vol. % to 1 vol. %, of the inspired breathing gas volume present in the lungs. 9. A method according to claim 1, wherein using the 1,1,1,2-tetrafluoroethane trace gas to measure lung function includes dispensing 1,1,1,2-tetrafluoroethane continuously per breath by means of a spray dispenser during respiration. 10. A method comprising: providing a gas including a breathing gas with 1,1,1,2-tetrafluoroethane as a trace gas and a predetermined concentration of 1,1,1,2,3,3,3-heptafluoropropane as another trace gas;delivering said breathing gas with the 1,1,1,2-tetrafluoroethane trace gas and the 1,1,1,2,3,3,3-heptafluoropropane trace gas to the lungs of a subject via one of an anesthesia apparatus and a respirator; andmeasuring lung function of the subject during inhalation with one of said anesthesia apparatus and said respirator based on the 1,1,1,2-tetrafluoroethane trace gas and the 1,1,1,2,3,3,3-heptafluoropropane trace gas, wherein a concentration during respiration of the trace gas 1,1,1,2-tetrafluoroethane combined with 1,1,1,2,3,3,3-heptafluoropropane trace gas is about 1 vol. % to 5 vol. %, maxi... | Dräger Medical AG & Co. KG | A61B005/08
US08150488 | B2 | 2012-04-03 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2009-10-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. In a glucose sensing system comprising a glucose sensor configured for producing an output representative of glucose concentration in a host and processing circuitry configured to process the output, a method for processing data from the output, the method comprising: generating at least one data stream from the glucose sensor output;detecting signal loss, wherein detecting signal loss comprises monitoring the data stream to identify a negative rate of change and determining that the negative rate of change exceeds a predetermined threshold;generating an estimated data stream or data stream portion to reduce or replace the signal loss; andoutputting to an external device or displaying the at least one estimated data stream or data stream portion. 2. The method of claim 1, comprising filtering the data stream. 3. The method of claim 1, wherein the data stream comprises filtered data. 4. The method of claim 1, wherein the data stream comprises raw data. 5. The method of claim 1, wherein the data stream comprises a calibrated data stream associated with glucose concentration values. 6. A glucose sensor system for monitoring a glucose concentration in a host, the system comprising: a continuous glucose sensor configured to produce an output;a signal processing module configured to generate a data stream based on the output;a signal loss detector configured to detect signal loss by monitoring the data stream to identify a negative rate of change and determining that the negative rate of change exceeds a predetermined threshold; a signal estimator configured to generate an estimated data stream or data stream portion to reduce or replace the signal loss; anda user interface configured to display information representative of the estimated data stream or data stream portion responsive to detecting signal loss. 7. The system of claim 6, wherein the data stream comprises filtered data. 8. The system of claim 6, wherein the data stream comprises raw data. 9. The system of claim 6, wherein the signal estimator is housed in at least one of a sensor electronics housing or a receiver electronics housing. 10. The method of claim 6, wherein the data stream comprises a calibrated data stream associated with glucose concentration values. 11. In a glucose sensing system comprising a glucose sensor configured for producing an output representative of glucose concentration in a host and processing circuitry configured to process the output, a method for processing data from the output, the method comprising: generating at least one data stream from the glucose sensor output;detecting signal loss, wherein detecting signal loss comprises evaluating the at least one data stream for a high frequency content that exceeds a predetermined threshold;generating an estimated data stream or data stream portion to reduce or replace the signal loss; andoutputting to an external device or displaying the at least one estimated data stream or estimated data stream porti... | DexCom, Inc. | A61B005/05 A61B005/00
US08386017 | B2 | 2013-02-26 | Method for implementing an imaging examination method | 2009-07-22 | A method is disclosed for implementing an imaging examination method. In at least one embodiment, the method includes creating an overview data record containing angiography data relating to a patient; simulating flow conditions in vessels of the patient; determining a strain on the vascular walls with the aid of the simulation; identifying vascular regions, in which the strain on the vascular wall exceeds a threshold value and implementing an imaging examination in at least one of the identified vascular regions. | 1. A method for implementing an imaging examination method comprising: creating an overview data record with angiography data relating to a patient;simulating flow conditions in vessels of the patient;determining a strain on vascular walls with the aid of the simulation;identifying vascular regions, in which the determined strain on the vascular wall exceeds a threshold value by comparing the determined strain with the threshold value, an image resolution of the overview data record being lower than an image resolution of the identified vascular regions; andimplementing an imaging examination in at least one of the identified vascular regions, wherein the overview data record is created by a first MR examination, with a blood circulation of the patient being enforced by a contrast agent containing iron oxide, the identified vascular regions are examined by a second MR examination, with the blood circulation of the patient still being enforced by the contrast agent containing iron oxide and the contrast agent containing iron oxide is embodied such that the contrast agent containing iron oxide adheres to ligands of vulnerable plaque. 2. The method as claimed in claim 1, wherein a plurality of identified vascular regions are examined, with the plurality of the vascular regions being examined in an order of their exceeding of the threshold value. 3. The method as claimed in claim 2, wherein a measurement period is defined for examinations of the identified vascular regions and a number of examinations of vascular regions which is implementable during the measurement period is selected automatically. 4. The method as claimed in claim 2, wherein the identified vascular regions are shown in a graphical overview on a display medium, with a numerical value exceeding of the threshold value of the respective region being highlighted in color. 5. The method as claimed in claim 1, wherein a measurement period is defined for examinations of the identified vascular regions and a number of examinations of vascular regions which is implementable during the measurement period is selected automatically. 6. The method as claimed in claim 1, wherein the identified vascular regions are shown in a graphical overview on a display medium, with a numerical value exceeding of the threshold value of the respective region being highlighted in color. 7. The method as claimed in claim 1, wherein the contrast agent which contains iron oxide contains VSOP. 8. The method as claimed in claim 1, wherein the simulation of the flow conditions includes: segmenting the angiography data and identifying vessels; anddetermining flow speed and vascular wall pressure by way of numerical simulation. 9. The method as claimed in claim 1, wherein the determination of the strain on the vascular walls includes: identifying flow type changes and flow speed changes. 10. A non-transitory computer readable medium including program segments for, when executed on a computer device, causing the... | Siemens Aktiengesellschaft | A61B005/055 A61B005/02
US09549687 | B2 | 2017-01-24 | Patient transport apparatus with integrated subsystems for use in magnetic resonance imaging | 2008-08-08 | A patient transport apparatus having a table configured to support a patient, a base attached to the table, a docking system attached to the base, the docking system configured to couple to a mating docking system of an MR imaging system, and a plurality of bays formed in the base, with each bay configured to receive a patient care module therein. The patient transport apparatus further includes a control system configured to be electrically coupled to each patient care module received within the plurality of bays and configured to centrally control each patient care module. | 1. A patient transport apparatus comprising: a table configured to support a patient;a base attached to the table;a docking system attached to the base, the docking system configured to couple to a mating docking system of an MR imaging system;a plurality of bays formed in the base, each bay configured to receive a patient care module therein and magnetically shield the patient care module from a magnetic field generated by the MR imaging system; anda control system configured to be electrically coupled to each patient care module received within the plurality of bays and configured to centrally control each patient care module;wherein the table is detachable from the base and configured to be transferred into a bore of an MRI system magnet to enable an acquisition of MR images of the patient, such that the patient care modules received in the plurality of bays in the base are kept outside of the bore of the MRI system during the acquisition of MR images. 2. The patient transport apparatus of claim 1 wherein the plurality of bays comprises: a first bay configured to receive an anesthesia module, the anesthesia module configured to combine an anesthetic agent with a carrier gas;a second bay configured to receive a ventilator module, the ventilator module configured to deliver gases to the patient via a breathing tube; anda third bay configured to receive a patient monitoring module, the patient monitoring module configured to collect medical data from the patient and transmit the collected medical data to the control system. 3. The patient transport apparatus of claim 2 wherein the plurality of bays further comprises a fourth bay configured to receive one of an intravenous (IV) pump module configured to provide fluids intravenously to the patient and an injector module configured to provide an MRI contrast agent to the patient. 4. The patient transport apparatus of claim 1 further comprising a video display module configured to display patient data and patient care module systems data. 5. The patient transport apparatus of claim 4 wherein the control system comprises a user input device integrated with one of the video display module and the base. 6. The patient transport apparatus of claim 1 further comprising a compartment configured to store a container for life support gases. 7. The patient transport apparatus of claim 1 further comprising a compartment configured to house one of a battery and a fuel cell configured to provide power to a patient care module. 8. The patient transport apparatus of claim 1 further comprising a connection port configured to receive one of electrical power, patient gases, and electronic communications from a source external to the patient transport apparatus. 9. The patient transport apparatus of claim 1 wherein the base is configured to permit adjustment to a height of the table. 10. A patient imaging system comprising: a magnetic resonance imaging (MRI) system having a plurality of gradient coils positio... | GENERAL ELECTRIC COMPANY | A61B005/05 A61B005/055 A61G013/06 A61G013/10
US20100016679 | A1 | 2010-01-21 | Apparatus and Method for the Identification of a Subject Memory Trace and for the Intention Prevision | 2009-07-16 | An apparatus and method for the identification of a subject's memory trace and for intention prediction use a procedure for acquiring information by means of administering to the subject a series of tests; the apparatus records the answers and the reaction times of the subject and it elaborates the results in order to give, with an accuracy of 92%, the memory trace or the intention of the subject. The apparatus and method give the possibility to avoid or identify maligners and fakers. | 1. Apparatus for the measurement of a subject's memory trace and for the intention prediction, comprising a unit of elaboration comprising a software that implements a series of test,visual means and/or audio-visual means to administer the test to the subject, characterized in that the apparatus comprisesat least a first and a second key for the input, in the unit of elaboration, of the responses to the testmeans for measuring the reaction time of the responsesmeans for storing the subjects responses and the relative reaction time and in thatthe software of the apparatus comprises a plurality of tests divided in different phases, the tests being based on sentences describing concepts or actions to be classified, pressing respectively the first and second key, whereina first phase comprises the classification, by the subject, of innocent and guilty sentences, in the guilty sentences the subject declaring to have committed an action, while in the innocent sentences the subject declaring not to have committed an action, the classification comprising the pressing of the first and second key answering the sentences subdued to the subject, the apparatus storing the response and the relative reaction time of the first phase,a second phase comprises the classification of true and false sentences, the true sentences relating to facts absolutely true and the false sentences relating to facts absolutely false, the apparatus storing the response and the relative reaction time of the second phase,a third phase comprises the classification of true and guilty sentences with the same key and of false and innocent sentences with another key, the apparatus storing the response and the relative reaction time of the third phase,a fourth phase comprises the classification of guilty and innocent sentences, using the keys of the first phase in a reversed manner, the apparatus storing the response and the relative reaction time of the fourth phase,a fifth phase comprising the combinations of the sentences of the third phase, the classification of guilty and false sentences with the first key and of innocent and true sentences with the second key, the apparatus storing the responses and the relative reaction time of the fifth phase. 2. Apparatus according to claim 1, wherein the unit of elaboration comprises a software which eliminates the reaction time with latencies inferior to 150 ms and superior to 10000 ms, calculates the standard deviation, using latencies of the third and fifth phases, calculates the mean latencies of the correct classifications in the third and in the fifth phases, substitutes the reaction times of the wrong trials with the mean of the corresponded phase added of a time of 600 ms, calculates the mean of the reaction times of the critical phases (corrected for the errors), calculates the difference of the reaction times of the fifth and third phases, divides the result for the standard deviation, gets an index called IAT-d through which is possib... | Universita' Degli Studi di Padova | A61B005/16
US20100010318 | A1 | 2010-01-14 | Identifying Products Containing a Food Item That Cause a Food Sensitivity | 2008-07-11 | In one aspect a mobile health monitor that indicates that a food product is suitable for a sufferer of a food sensitivity is provided. The mobile health monitor includes a reader that reads the barcode of the food product, a database having information about a plurality of food products to determine the suitability of the respective food product, and an indicator that provides an indication when a food product is suitable for the suffer. The database is searched base on the read barcode to determine if the food product associated with the barcode is suitable. The indicator indicates that the food product is suitable when the food product is determined suitable based on the searched database. A method indicating that a food product is suitable for a sufferer of a food sensitivity is also provided. | 1. A mobile health monitor system that indicates that a food product is suitable for a sufferer of a food sensitivity; comprising: a mobile health monitor comprising a reader configured to read a barcode of the food product, a mobile health monitor database having information about a plurality of food products to determine the suitability of the food product, the database being searchable based on information obtained from reading the barcode, and an indicator configured to provide an indication to identify food product suitability; anda remote processing unit configured for releasable connection to the mobile health monitor, the remote processing unit having a food suitability database comprised of food product information: andwherein the mobile health monitor is configured to communicatively couple to the remote processing unit such that the mobile health monitor database is updated to include at least a portion of the food product information of the food suitability database. 2. The mobile health monitor system according to claim 1, wherein the indication is visual. 3. The mobile health monitor system according to claim 1, wherein the indicator is comprised of at least one LED. 4. The mobile health monitor system according to claim 1, wherein the indication is audible. 5. The mobile health monitor system according to claim 1, wherein only information for suitable food products is stored in the mobile health monitor database. 6. The mobile health monitor system according to claim 1 wherein the remote processin unit is a computer or a host computer center. 7. The mobile health monitor system according to claim 6, wherein only information for suitable food products is transferred to the mobile health monitor when the mobile health monitor database is updated. 8. A method for determining when a food product is suitable for a sufferer of a food sensitivity, comprising: entering food sensitivity information into a mobile health monitor;communicatively coupling a remote processing unit to the mobile health monitor to update a database of the mobile health monitor to include food product status information for food products based on food sensitivity information entered into the mobile health monitor:reading a food product via a mobile health monitor;determining food product information for the food product with the mobile health monitor;searching the database of the mobile health monitor to determine the suitability of the food product:indicating the suitability of the food product by an indication emitted or displayed by the mobile health monitor; andwherein the food product is indicated as being unsuitable when information for the food product is not found during the searching of the database. 9. The method according to claim 8, wherein the database is searched by product code. 10. The method according to claim 8, wherein the remote processing unit is a host computer center or a computer. 11. The method according to claim 8, wherein only f... | SIEMENS ENTERPRISE COMMUNICATIONS GMBH & CO. KG | A61B005/00 G06K007/10 G09B019/00
US20100007357 | A1 | 2010-01-14 | Electrical Impedance Tomography Method and Device | 2007-09-26 | Electrical impedance tomography method comprising: an electrical measurement step during which pre-determined electrical conditions are imposed on the surface of a medium to be imaged, while generating a mechanical disturbance at predefined points of the medium by locally modifying the impedance of the medium and an electrical parameter is measured at several points on the surface of the medium; and a calculation step during which the electrical impedance is determined at several points in the internal volume of the medium, taking into account the measurements carried out during the disturbance, as a function of a law for modification of the electrical impedance by this disturbance. | 1. An electrical impedance tomography method for imaging a medium having a certain internal volume delimited by an external surface, this method comprising: at least one electrical measurement step during which:predetermined electrical conditions are imposed on the surface of the medium by an electrical measurement apparatus controlled by a central unit, said electrical measurement apparatus having electrodes on the surface of the medium;and at least one electrical parameter is measured by said electrical measurement apparatus at several points on the surface of the medium while generating a mechanical disturbance at predefined points of the medium by a mechanical disturbance generating apparatus controlled by said central unit, thus locally modifying the impedance of the medium,and at least one calculation step during which at least one parameter, linked to the electrical impedance, is determined at several points in the internal volume of the medium, wherein during the calculation step, said parameter linked to electrical impedance is determined taking into account the measurements carried out during said disturbance, using a predetermined law for modification of the electrical impedance by said disturbance. 2. The method according to claim 1, in which the electric conditions imposed include at least one current imposed at least one point on the surface of the medium, and said measured electrical parameter is an electrical potential. 3. The method according to claim 1, in which said parameter linked to the electrical impedance is the conductivity. 4. The method according to claim 1, in which the mechanical disturbance is a wave focussed on at least one point of the medium. 5. The method according to claim 4, in which the wave is an acoustic wave. 6. The method according to claim 5, in which the acoustic wave is an ultrasonic wave. 7. The method according to claim 6, in which the wave corresponds to an amplitude-modulated signal at a modulation frequency suitable for generating an ultrasound radiation force resulting in a local displacement of the medium. 8. The method according to claim 4, in which the wave is an elastic wave. 9. The method according to, claim 4, in which the wave corresponds to an encoded signal. 10. The method according to claim 4, in which, during the calculation step, the following equation is solved:γ(z)=Ek(z)A∇uk(z)·∇uk(z)(8)for any point z of the medium to be imaged,where: k is an index denoting a set of at least one electric current jik at the surface of the medium, i being an index denoting each current of this set,Ek(z) is an energy corresponding to the disturbance produced by the wave during the application of the set of electric currents jik at the surface of the medium,uk(z) is the electrical potential at the point z of the medium,and A is a matrix representative of the shape of a focal spot produced by the wave about the point on which it is focussed. 11. The method according to claim 10, in which t... | Centre National De La Recherche | G01R027/28 G01N027/02 A61B005/053 G06F017/10
US20100016709 | A1 | 2010-01-21 | SYSTEM AND METHOD OF RECORDING AND DISPLAYING IN CONTEXT OF AN IMAGE A LOCATION OF AT LEAST ONE POINT-OF-INTEREST IN A BODY DURING AN INTRA-BODY MEDICAL PROCEDURE | 2009-07-14 | A method of displaying at least one point-of-interest of a body during an intra-body medical procedure. The method is effected by (a) establishing a location of the body; (b) establishing a location of an imaging instrument being for imaging at least a portion of the body; (c) defining at least one projection plane being in relation to a projection plane of the imaging instrument; (d) acquiring at least one point-of-interest of the body; and (e) projecting said at least one point-of-interest on said at least one projection plane; such that, in course of the procedure, the locations of the body and the imaging instrument are known, thereby the at least one point-of-interest is projectable on the at least one projection plane even in cases whereby a relative location of the body and the imaging instrument are changed. | 1. (canceled) 2. A method for providing navigational assistance for directing a tool at an angle to a surface of an intra-body cavity comprising: providing a catheter with a location sensor associated with a distal tip thereof;projecting a location and direction of said tip on an image plane related to an imaging picture of said intra-body cavity on a monitor;calculating a direction in which a local tissue portion is oriented;displaying on said monitor a representation of said direction in which said local tissue portion is oriented;displaying an indication of an angle at which said catheter is oriented relative to said representation. 3. The method of claim 2 wherein said calculated direction in which a local tissue portion is oriented comprises an iso-height curve along said tissue portion, relative to a reference from of coordinates. 4. The method of claim 2 wherein said direction in which a local tissue portion is oriented is displayed as a tissue plane. 5. The method of claim 4 wherein displaying an indication of an angle at which said catheter is oriented relative to said representation comprises displaying an angle of said tip relative to two perpendicular planes, each of which is perpendicular to said tissue plane. 6. The method of claim 2 wherein displaying an indication of an angle at which said catheter is oriented relative to said representation comprises displaying in a first view, a viewing angle of an imaging instrument and in a second view a perpendicular direction to said viewing angle. 7. The method of claim 2 wherein calculating a direction in which a local tissue portion is oriented comprises: recording positions of at least three points on said tissue portion;defining said direction in which a local tissue portion is oriented as being normal to a plane defined by said at least three points. 8. The method of claim 7 wherein recording positions of at least three points on said tissue portion comprises dragging said catheter tip along a portion of said tissue. 9. The method of claim 7 wherein recording positions of at least three points on said tissue portion comprises using an ultrasound probe equipped with a six degree of freedom locating system an a 3D modeling algorithm. 10. The method of claim 2 wherein calculating a direction in which a local tissue portion is oriented comprises using said catheter tip to draw a line directly on said tissue portion while recording positions of said tip. | - | A61B005/05
US07654967 | B2 | 2010-02-02 | Gas column pressure monitoring device | 2007-03-07 | The present invention provides a device for efficiently coupling a pressure monitoring device to a pressure transducer. In addition, the present invention reduces or eliminates the likelihood air will be unintentionally injected into the pressure monitoring device while connecting the pressure monitoring device to a pressure transducer. In one embodiment, the present invention includes a pressure monitoring coupler comprising a sealing member and a device body. The sealing member comprises a sealing flange connected to an engagement member. The engagement member includes a first and second O-ring positioned thereon. The device body comprises a pressurizing lumen and an intersecting monitoring lumen in communication with a receiving aperture. During use, a pressure monitoring catheter is inserted into a patient's tissue. The pressure monitoring catheter and a transducer are connected to the device body of the present invention. Thereafter, the sealing member is inserted into the device body while the pressure within the pressure monitoring catheter is maintained at or less than a state or equilibrium with the pressure applied by the tissue. | 1. A method of connecting a pressure monitoring catheter to a pressure transducer, comprising: providing a pressure monitoring connector having an outlet port open to the atmosphere;coupling a pressure monitoring catheter to said connector;coupling a transducer to said connector;adjusting a sealing member in said pressure monitoring connector after coupling said pressure monitoring catheter and said transducer to said connector so as to selectively vent said outlet port to the atmosphere in a first position of said sealing member and to both seal said outlet port and to displace a predetermined volume of air into a passage between said pressure monitoring device and said transducer in a second position of said sealing member, thereby limiting a pressure differential within said pressure monitoring catheter. 2. The method of claim 1, wherein said adjusting a sealing member further comprises moving a plurality of o-rings disposed on said sealing member. 3. The method of claim 1, wherein said coupling a transducer to said connector further comprises coupling said transducer for measuring pressure within a cranium. 4. The method of claim 1, wherein said adjusting a sealing member includes moving said sealing member relative to a body of said pressure monitoring connector. 5. The method of claim 4, wherein said moving said sealing member relative to a body of said pressure monitoring connector further comprises moving said sealing member in a passage of said body of said pressure monitoring connector. 6. The method of claim 5, wherein said outlet port open to the atmosphere is connected to said passage of said body. 7. The method of claim 1, wherein said providing a pressure monitoring connector having an outlet port open to the atmosphere further comprises providing said connector further including a first passage and a second passage. 8. The method of claim 7, wherein said coupling a pressure monitoring catheter to said connector further comprises connecting said first passage to said pressure monitor. 9. The method of claim 8, wherein said coupling a transducer to said connector further comprises connecting said transducer to said first passage. 10. The method of claim 9, wherein said providing said connector further including a first passage and a second passage further comprises providing said first passage intersected by said second passage. 11. The method of claim 10, wherein said providing said connector further comprises providing said second passage connected to said outlet port. | InnerSpace Medical, Inc. | A61B005/00
US07651471 | B2 | 2010-01-26 | Encapsulated medical device and method of examining, curing, and treating internal region of body cavity using encapsulated medical device | 2006-06-13 | An encapsulated medical device is passed through the lumen of a body cavity in order to examine, cure, or treat an internal region of the body cavity under the control of the extracorporeal device. A capsule body has a linkage hole bored in the extended portion thereof so that when the capsule body comes into close contact with the internal surface of a body cavity to block the lumen thereof, the linkage hole will realize a fluid passage which allows a fluid such as a gas or humor to flow into the forward and backward parts of the lumen. | 1. An encapsulated medical device comprising: a capsule body capable of passing through the lumen of a body cavity to permit examination or treatment; anda fluid passage, one of provided in and on the capsule body, for when the capsule body passes through the lumen of the body cavity, if the capsule body comes into close contact with the internal surface of the body cavity to block the lumen of the body cavity, linking forward and backward portions of the lumen of the body cavity;wherein the capsule body includes at least one medical-purpose examining means including at least one of an optical sensor, a PH sensor, a temperature sensor, a pressure sensor or an ultrasonic probe; andwherein the fluid passage is formed directly in the capsule body or in an armor member that is freely detachably attached to the periphery of the capsule body. 2. The encapsulated medical device according to claim 1, wherein an extended portion is provided of which cross section cut at right angle to the longitudinal axis is non-circular and enlarged in the direction of one axis of the capsule body of cross section that is substantially circular, and the fluid passage is provided to the extended portion. 3. The encapsulated medical device according to claim 1, wherein a plurality of fluid passages are formed in one direction with respect to the longitudinal axis of the capsule body. 4. The encapsulated medical device according to claim 1, wherein the capsule body includes at least one observing means. 5. The encapsulated medical device according to claim 1, wherein a plurality of fluid passages are formed substantially symmetrically to one another with respect to the longitudinal axis of the capsule body. 6. The encapsulated medical device according to claim 1, wherein the capsule body has a plurality of hard members linked by a flexible string-like member. 7. The encapsulated medical device according to claim 1, wherein the capsule body has a battery that supplies power and a switch to control power supply from the battery incorporated therein. 8. The encapsulated medical device according to claim 1, wherein the capsule body includes at least one opening means for a treating or examining means such as a medicine administration port or a humor infusion port. 9. The encapsulated medical device according to claim 1, wherein the capsule body includes at least one treating means such as an injection needle. 10. The encapsulated medical device according to claim 1, wherein the fluid passage is a spiral groove. 11. The encapsulated medical device according to claim 1, wherein the capsule body has at least a magnet incorporated therein. 12. The encapsulated medical device according to claim 1, wherein the capsule body has at least a magnet incorporated therein, and the fluid passage is a spiral groove. 13. The encapsulated medical device according to claim 1, wherein the fluid passage is a linkage hole formed between the capsule body and the armor member or in the a... | Olympus Corporation | A61B005/117 A61B005/103
US07647095 | B2 | 2010-01-12 | Method and apparatus for verifying a determined cardiac event in a medical device based on detected variation in hemodynamic status | 2006-04-28 | A method and apparatus for verifying a determined cardiac event in a medical device based on detected variation in hemodynamic status that includes a plurality of sensors sensing cardiac signals, and a physiologic sensor sensing physiologic signals to generate a plurality of variation index samples corresponding to the sensed signals. A microprocessor detects a cardiac event in response to the sensed cardiac signals, computes a variation index trend associated with a predetermined number of variation index samples of the plurality of variation index samples, determines a rate of change of the computed variation index trend, and confirms the determined cardiac event in response to the computed variation index trend and in response to the determined rate of change. | 1. A medical device comprising: a plurality of sensors configured to sense cardiac signals;a physiologic sensor sensing physiologic signals configured to generate a plurality of oxygen (O2) variation index samples corresponding to the sensed signals; anda microprocessor configured to: detect a cardiac event in response to the sensed cardiac signals;compute an O2 variation index trend associated with a predetermined number of O2 variation index samples of the plurality of O2 variation index samples in response to the detected cardiac event;determine, in response to the O2 variation index trend being less than a variation index trend threshold, whether a rate of change of the O2 variation index trend is within a predetermined cardiac event zone defined by a first zone threshold having a first rate of change corresponding to the O2 variation index trend becoming less than the variation index trend threshold in a first period of time and a second zone threshold having a second rate of change corresponding to the O2 variation index trend becoming less than the variation index trend threshold in a second period of time greater than the first period of time;and confirm the determined detected cardiac event in response to both the computed O2 variation index trend being less than the variation index trend threshold and the rate of change of the O2 variation index trend being within the predetermined cardiac event zone. 2. The device of claim 1, wherein the O2 variation index trend corresponds to a measure of change in tissue oxygenation and includes a start point corresponding to a first O2 variation index sample of the predetermined number of O2 variation index samples and an endpoint corresponding to a second O2 variation index sample of the predetermined number of O2 variation index samples. 3. The device of claim 1, wherein the microprocessor is configured to determine a deviation of the predetermined number of O2 variation index samples from the O2 variation index trend. 4. The device of claim 1, wherein the microprocessor is configured to determine whether the rate of change is one of greater than the first zone threshold and less than the second zone threshold. 5. The device of claim 1, wherein the microprocessor is configured to determine whether the rate of change is within one of a first range within the cardiac event zone corresponding to a ventricular tachycardia event and a second range within the cardiac event zone corresponding to a ventricular fibrillation event. 6. The device of claim 1, wherein the O2 variation index trend corresponds to relationships between proportions of red light intensities and infrared light intensities to corresponding baseline intensities. 7. The device of claim 1, wherein the O2 variation index trend corresponds toii0-i*i0*wherein is an intensity of red light, i* is an intensity of infrared light, i0 is an intensity of red light baseline, and i0* is an intensity of infrared light baseline. 8. A method of... | Medtronic, Inc. | A61B005/04
US07652468 | B2 | 2010-01-26 | Correction of measured values for a magnetic localization device | 2005-02-11 | The invention relates to a device and a method for correction of the position (x) of a field sensor (4) measured by means of a magnetic localization device. External field distortions, such as caused for example by the rotating components (1a, 1b) of a computer tomograph (1), are then determined with the help of reference sensor (3) placed at a known position. It is possible to deduce, for example, the current angle of rotation (Φ) of the computer tomograph (1) from the measurement signals of the reference sensor (3). Based on an empirically determined correction (δ(x, Φ)), the uncorrected determined positions (x) of the field sensor (4) can then be converted to corrected positions (x′) in relation to the field distortions. The field generator (2) and the reference sensor (3) are preferably fastened to the gantry in order to eliminate the dependency of the field distortions on an inclination of the gantry. | 1. A magnetic localization device, comprising: a) a field generator for generating a magnetic field;b) a field sensor for measuring the magnetic field;c) a reference sensor for measuring the magnetic field at a known reference position; andd) a control unit, which is arranged for determining a position of the field sensor relative to the field generator and thereby compensating for external field distortions by taking the reference sensor into consideration and correcting the determined position of the field sensor if external field distortions are present. 2. A localization device as claimed in claim 1, wherein the spatial position of the field generator is known. 3. A localization device as claimed in claim 1, wherein at least one of the field generator and the reference sensor is fastened to the gantry of a computer tomograph. 4. A localization device as claimed in claim 1, wherein the control unit contains a memory with a calibration function, which provides a correction shift for an uncorrected determined position of the field sensor based on measured signals of the reference sensor and the field sensor. 5. A localization device as claimed in claim 4, wherein the correction shift is further based on empirical measurements taken with a probe sensor at one or more support points in a volume of interest. 6. A localization device as claimed in claim 4, wherein the measured signals of the reference sensor are used to determine a correction parameter describing an angle of rotation of a computer tomograph situated in the vicinity of the localization device. 7. A method for position measurement with a magnetic localization device, comprising the steps of: a) collecting the signals of at least one of a field sensor and a field generator;b) collecting the signals of a magnetic reference sensor, which is placed at a known spatial position relative to the field generator or to the field sensor; andc) determining a position of the field sensor relative to the field generator, where external field distortions are compensated for by taking the signals of the reference sensor into consideration and correcting the determined position of the field sensor if external field distortions are present. 8. A method as claimed in claim 7, wherein a correction function is determined, which indicates a correction shift for an uncorrected determined position of the field sensor in dependence on the signal of the reference sensor and the uncorrected determined position of the field sensor. 9. A method as claimed in claim 8, wherein the correction shift further depends on empirical measurements taken with a probe sensor at one or more support points in a volume of interest. 10. A method as claimed in claim 8, wherein the signal of the reference sensor is used to determine a correction parameter describing an angle of rotation of a computer tomograph situated in the vicinity of the localization device. 11. A method as claimed in claim 7, wherein a parameter is d... | Koninklijke Philips Electronics N.V. | G01B007/0004 G01B007/14 A61B005/05
US07668734 | B2 | 2010-02-23 | Internet medical information system (IMED) | 2006-04-10 | A system for creating, maintaining and selectively accessing medical histories. The system allows a patient and other selected parties to build a medical history, but importantly places the authority to grant access to the medical history solely in the hands of the patient. When the patient's medical history file is created, the patient also pre-authorizes selected parties, including government agencies such as FEMA to have access to at least part of the patient's history. Each selected party is associated with an access module that dictates the level of access the party has to the medical history. Primary care physicians are associated with a full access module, which allows full access to the patient's medical history. Dentists and pharmacies are associated with a limited access module, which limits the party's access to only pertinent portions of the patient's medical history. The present system also includes a fingerprint scanner and/or retina scanner that can be used to identify unconscious patients and patients without an I.D. card. | 1. A system for creating, maintaining and accessing medical histories and, in emergencies, for providing pre-authorized access to the medical histories, wherein patients pre-authorize selected parties to access the histories and wherein access to the histories is limited for many of the selected parties, the system comprising: at least one server, which has stored therein: at least one database, the at least one database, storing multiple patient files, wherein each patient file is associated with a patient and contains patient information, wherein the patient information defines a medical history of the associated patient, the patient information being stored in multiple fields within each patient file;the at least one server also having stored therein multiple access modules, wherein each access module is associated with a selected party and at least one patient file, further wherein each access module comprises multiple data files, and at least one of the data files is an access file that contains limiting information that dictates an amount of the patient's medical history the associated selected party may access;a network, wherein the at least one sever is connected to the network;a computer, wherein the computer is associated with an access module and is connected to the network, wherein the selected party associated with the access module is able to gain access to the at least one patient file in the server in accordance with the access file, and wherein the patients can use the computer to authorize access to their medical histories without the aid of an identification card or a smart card; and,an identification device that is connected to the computer, capable of identifying the patient upon physical interaction with the patient, even when the patient is unconscious, and wherein upon identification, the selected party is able to access the patient's medical history in the server in accordance with the associated access module. 2. The system of claim 1, wherein the identification device is a biometric reading device, including a fingerprint scanner and a retina scanner. 3. The system of claim 1, wherein the identification device is an electro-magnetic card reader capable of reading a card with an electro-magnetic strip, and wherein only connection related information is stored on the card that assists in connection to the server and access to the files and modules stored in the server. 4. The system of claim 1, wherein at least one of the multiple data files is an administration file that provides an interface for the patient and the selected party and assists in creating the patient files and associating at least one access module to each patient file. 5. The system of claim 4, wherein the administration file also provides an interface for the patient and the selected party to update the patient's medical history. 6. The system of claim 1, wherein at least some patient information and at least one access module are stored on the com... | - | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07657297 | B2 | 2010-02-02 | Implantable analyte sensor | 2004-05-03 | An implantable analyte sensor including a sensing region for measuring the analyte and a non-sensing region for immobilizing the sensor body in the host. The sensor is implanted in a precisely dimensioned pocket to stabilize the analyte sensor in vivo and enable measurement of the concentration of the analyte in the host before and after formation of a foreign body capsule around the sensor. The sensor further provides a transmitter for RF transmission through the sensor body, electronic circuitry, and a power source optimized for long-term use in the miniaturized sensor body. | 1. An electrochemical glucose sensor comprising a three-electrode system, the sensor comprising: an electrochemical cell comprising a working electrode, reference electrode, and counter electrode; anda potentiostat that controls the potential between the working and reference electrodes, wherein the potentiostat sets a low voltage limit at which the counter electrode of the glucose sensor reacts with reducible species, other than oxygen, when oxygen becomes limited and at which no substantial bubble formation occurs, wherein the potentiostat comprises an operational amplifier comprising: (i) a first input coupled to the reference electrode;(ii) a second input coupled to a reference voltage source; and(iii) an output coupled to the counter electrode, wherein the reference voltage source is set to a voltage value relative to battery ground such that the low voltage limit at the counter electrode causes the counter electrode to react with reducible species, other than oxygen, when oxygen becomes limited, and substantial bubble formation does not occur at the counter electrode. 2. The sensor of claim 1, wherein the reference voltage setting is between about +0.6V and +0.8V with respect to battery ground. 3. The sensor of claim 2, wherein the reference voltage setting is about +0.7V with respect to battery ground. 4. The sensor of claim 1, wherein the electrochemical cell is configured and arranged for transcutaneous implantation. 5. The sensor of claim 1, further comprising an RF module, wherein the RF module is operably connected to the potentiostat. 6. The sensor of claim 5, further comprising a receiver configured to receive RF transmissions from the RF module. 7. The sensor of claim 6, wherein the receiver comprises a display configured to display sensor data. 8. The sensor of claim 1, wherein the potentiostat comprises a second operational amplifier comprising: (i) a first input coupled to the working electrode;(ii) a second input coupled to a working voltage source; and(iii) an output coupled to the working electrode through a resistor, wherein the potential between the working and reference electrodes is determined by the potential of the working voltage source and reference voltage source. 9. An electrochemical glucose sensor comprising a three-electrode system, the sensor comprising: an electrochemical cell comprising a working electrode, reference electrode, and counter electrode; anda potentiostat that controls the potential between the working and reference electrodes, wherein the potentiostat comprises an operational amplifier comprising:(i) a first input coupled to the reference electrode;(ii) a second input coupled to a reference voltage source;(iii) an output coupled to the counter electrode;(iv) a negative power supply input coupled to a first voltage source; and(v) a positive power supply input coupled to a second voltage source, wherein the reference voltage source, first voltage source and second voltage source set a low v... | DexCom, Inc. | A61B005/05 A61B005/00 G01N027/26
US07668583 | B2 | 2010-02-23 | Method and apparatus for control and location of an instrument or appliance | 2003-06-10 | The invention relates to an apparatus for location of an instrument or appliance having at least one magnet which produces a magnetic moment at right angles to the appliance shaft and can be rotated independently of the instrument or appliance. This results in a location system which allows precise location at an accurate time, determination of the axis direction and control of an appliance which is being operated in a channel or medium. | 1. An apparatus for locating an instrument, comprising: an instrument having a body;at least one magnet disposed inside the body of the instrument, the at least one magnet being rotatable independently of a rotation of the body of the instrument and adapted to produce a moving magnetic field extending outside the body of the instrument to generate a magnetic moment which is perpendicular to an axis of the instrument;a drive for rotating the at least one magnet independent of the body of the instrument;variation means for varying the magnetic field generated outside the body of the instrument by the at least one magnet, thereby creating a variable magnetic field component which relates to a roll angle of the instrument; anda receiver outside the body of the instrument configured for measuring the variable magnetic field component to determine the roll angle of the instrument. 2. The apparatus as claimed in claim 1, wherein the receiver is configured to detect three time-dependent magnetic field components. 3. The apparatus as claimed in claim 1, further comprising an evaluation unit for determining from the variable magnetic field component at least one parameter selected from the group consisting of position, direction of the instrument axis, and roll angle of the instrument. 4. The apparatus as claimed in claim 1, wherein the drive is an electrical drive. 5. The apparatus as claimed in claim 1, wherein the drive is a hydraulic drive using liquid to drive the magnet. 6. The apparatus as claimed in claim 1, further comprising means for providing a reproducible deflection of the magnet from its rotation axis. 7. The apparatus as claimed in claim 1, further comprising a coupling which temporarily interrupts the rotation of the magnet. 8. The apparatus as claimed in claim 1, wherein the magnet is composed of magnet elements that move with respect to one another and whose elements are shifted by a driver at a specific roll angle. 9. The apparatus as claimed in claim 1, wherein the instrument has a member selected from the group consisting of drill, culling or impact apparatus, at least one needle, and at least one set of forceps. 10. The apparatus as claimed in claim 1, wherein the instrument has at least one opening for ejection of a liquid. 11. The apparatus as claimed in claim 1, wherein the instrument has an apparatus for production/emission of light beams, laser beams, radioactive beams, sound waves or ultrasound waves. 12. The apparatus as claimed in claim 1, wherein the instrument has an apparatus for recording optical images or ultrasound images. 13. The apparatus as claimed in claim 1, wherein the instrument has an apparatus for emission of electrical pulses or for recording electrical data. 14. The apparatus as claimed in claim 1, further comprising two or more transmitters and/or additional receivers for processing signals commensurate with a position of the instrument at different points. 15. The apparatus as claimed in claim... | Rayonex Schwingungstechnik GmbH | A61B005/05 E21B025/16 A61B019/00 A61B001/00
US08048000 | B2 | 2011-11-01 | Patient weight and ankle displacement correlation device | 2009-08-24 | A device for correlating trend data with respect to a patient's weight ankle displacement can identify conditions indicative of congestive heart failure. A weight scale or similar device coupled with imaging mechanism operable to measure ankle displacement collects a plurality of measurements over a period of time. Over time trend analysis of both the patient's weight and the ankle displacement measurements can be obtained and compared to identify whether over a particular sample period an increase in a patient's ankle displacement is or is not correlated with an increase in the patient's weight. When an increase in ankle displacement is identified as not correlating to a corresponding change in the patient's weight an alert can be issued of conditions indicative of congestive heart failure. | 1. An apparatus for correlating ankle displacement with weight, the apparatus comprising: a weight measuring device operable to measure a weight measurement of a user;an ankle displacement measurement device operable to measure an ankle displacement measurement of the user substantially simultaneously with operation of the weight measuring device;a memory for storing the weight measurement and the ankle displacement measurement; anda processor capable of executing a program of instructions wherein said program of instructions comprises a plurality of program codes for determining correlation between user ankle displacement measurements and weight measurements, wherein said program of instructions includes program code configured to calculate a weight measurement trend and an ankle displacement measurement trend,program code configured to determine a correlation factor between the weight measurement trend and the ankle displacement trend, andresponsive to the correlation factor being below a predetermined value, program code to determine whether the correlation factor being below the predetermined value is statistically significant. 2. The apparatus of claim 1 wherein the ankle displacement device uses ultrasound waves to measure ankle displacement. 3. The apparatus of claim 1 wherein the ankle displacement device uses infrared waves to measure ankle displacement. 4. The apparatus of claim 1 further comprising a transmitter operable to wirelessly transmit an alert to a healthcare professional responsive to the correlation factor being below the predetermined value and being statistically significant. 5. The apparatus of claim 1 further comprising a display operable to display an alert when the correlation factor is below the predetermined value and is statistically significant. 6. The apparatus of claim 1 wherein the processor is coupled to a network and wherein the processor is operable to transmit an alert via the network to a healthcare professional responsive to the correlation factor being below the predetermined value and being statistically significant. 7. The apparatus of claim 1 further comprising program code configured to, responsive to the correlation factor being below the predetermined value being statistically significant, create an alert. 8. The apparatus of claim 7 further comprising program code configured to automatically convey the alert to a healthcare professional. 9. The apparatus of claim 7 further comprising program code configured to inform the user of the alert. 10. A method for monitoring correlation between a patient's weight and ankle displacement, the method comprising: determining for the patient over a period of time, a plurality of weight measurements and a plurality of ankle displacement measurements wherein each weight measurement is associated an ankle displacement measurement forming a measurement pair and wherein each measurement pair is associated with a date regarding when over the period of time t... | STMicroelectronics, Inc. | A61B005/107
US08082025 | B2 | 2011-12-20 | ECG data acquisition device | 2008-08-14 | An apparatus for generating ECG recordings and a method for using the same are disclosed. The apparatus includes a handheld device and a controller. The handheld device has first, second, third, and fourth electrodes on an outer surface of the handheld device. The controller records signals from the first, second, third, and fourth electrodes and generates an ECG recording from the signals. The controller also provides placement information regarding locations on a user's body that are to be brought into contact with the first, second, third, and fourth electrodes while the controller records the signals. The handheld device can include a cellular telephone or a PDA. The controller stores an exemplary normal ECG recording, and compares a newly generated ECG recording with the stored ECG recording and generates an indication that indicates whether the newly generated ECG recording is consistent with the stored ECG recording. | 1. A method for generating an ECG recording of a patient comprising: selecting said ECG recording from a plurality of predetermined ECG recordings; providing a handheld device having first, second, third, and fourth electrodes on an outer surface of said handheld device; causing said patient to grip said handheld device such that said first and second electrodes contact a first hand of said patient such that said first and second electrodes contact different locations on said first hand, said third electrode being in contact with a predetermined location on said patient's other hand and said fourth electrode contacts a predetermined point on said patient's body chosen from the group consisting of said patient's lower abdomen, one of said patient's legs, and a precordial measurement point on said patient, said predetermined point depending on said selected ECG recording; measuring signals as a function of time between said first electrode and each of said second, third, and fourth electrodes while said patient simultaneously contacts said first, second, third, and fourth electrodes as described above; and converting said measured signals to said ECG recording. 2. The method of claim 1 wherein said first and second electrodes are in contact with the thumb and finger on said patient's first hand. 3. The method of claim 1 wherein said third electrode is contact with a finger on said patient's other hand. 4. The method of claim 1 wherein said fourth electrode is in contact with a predetermined point on one of said patient's legs. 5. The method of claim 1 wherein said fourth electrode is in contact with a Precordial measurement point on said patient's chest. 6. The method of claim 1 further comprising comparing said generated recording with an ECG recording stored in said handheld device and generating an indication that is receivable by said patient, said indication indicating whether said generated ECG recording is consistent with said stored ECG recording. 7. The method of claim 6 further comprising communicating said generated ECG recording to a location remote from said handheld device if said indication indicates that said generated ECG recording is not consistent with said stored ECG recording. | - | A61B005/0404
US08298153 | B2 | 2012-10-30 | System and method for the detection of acute myocardial infarction | 2008-07-09 | A system and method are provided for the detection of acute myocardial infarction (AMI) using a staged approach for accurate and rapid detection. Physiological signals in a patient's body are sensed and corresponding physiological parameters are derived in a staged approach in order to determine the probability that AMI is occurring in a patient in a first detection stage. If the computed probability from physiological signals indicates the possibility of AMI, then the patient is prompted, such as through a patient-wearable device, to answer specific AMI-related questions to assist in diagnosis of AMI in a second stage. AMI is detected when the computed probability in the second stage exceeds a predefined detection threshold. A patient or physician alert may then be generated, which may further include the transfer of data via a communication link or network, in response to an AMI detection signal. | 1. A method comprising: sensing at least one physiological signal in a first stage;deriving a corresponding physiological parameter from each sensed signal in the first stage;computing a probability of acute myocardial infarction from the physiological parameters in the first stage;determining that the computed probability of acute myocardial infarction has reached a sufficient threshold in the first stage;in response to the determination that the computed probability of acute myocardial infarction reached the sufficient threshold in the first stage, collecting additional information related to acute myocardial infarction in a second stage;computing a probability of acute myocardial infarction from the collected additional information in the second stage; andclassifying an event as acute myocardial infarction when the computed probability in the second stage exceeds a predefined detection threshold. 2. The method of claim 1, further comprising collecting additional information in the second stage by sensing at least one additional physiological signal and deriving a corresponding physiological parameter from each additional sensed signal. 3. The method of claim 1, further comprising collecting additional information in the second stage by collecting information from a patient relating to acute myocardial infarction-related symptoms being experienced by the patient. 4. The method of claim 3, further comprising querying the patient with acute myocardial infarction-related questions in response to the computed probability of acute myocardial infarction reaching the sufficient threshold in the first stage. 5. The method of claim 1, wherein the physiological signals comprise at least one of a cardiac electrical signal, a respiratory signal, a temperature signal and a heart sound signal. 6. The method of claim 1, further comprising generating an alert when the computed probability in the second stage exceeds a predefined detection threshold. 7. The method of claim 1, further comprising: performing the sensing, deriving and computing actions of the first stage in an implantable medical device implanted in a patient; andperforming the collecting, computing and detecting actions of the second stage in a device external to the patient and in communication with the implantable medical device. 8. A system comprising: at least one physiological sensor configured to sense a corresponding physiological signal;a controller configured to derive a corresponding physiological parameter from each sensed signal, to compute a probability of acute myocardial infarction from the physiological parameters in a first stage, and to determine that the computed probability of acute myocardial infarction has reached a sufficient threshold in the first stage;the controller further configured to, in response to determining that the computed probability of acute myocardial infarction has reached the sufficient threshold in the first stage, collect additional information re... | Medtronic, Inc. | A61B005/02
US08388529 | B2 | 2013-03-05 | Differential diagnosis of neuropsychiatric conditions | 2008-07-08 | A computer implemented method, apparatus, and computer program product for generating neuropsychiatric diagnoses. Quantitative information describing diagnostic characteristics associated with a patient is generated based on an analysis of a set of patient scans. The quantitative information comprises a set of indictors associated with regions of interest in the set of scans for the patient. The set of indicators of potential neuropsychiatric conditions is compared with a set of diagnostic signatures. A diagnostic signature comprises a set of indicators of a known neuropsychiatric condition. Matching signatures are identified. A matching signature is a diagnostic signature that corresponds to at least one indicator in the set of indicators to form a set of signatures. A diagnosis associated with each signature in the set of signatures is identified to form a set of potential diagnoses. The set of potential diagnoses is presented with links to relevant portion of the medical literature. | 1. A computer implemented method executing on a processor for generating neuropsychiatric diagnoses, the computer implemented method comprising: automatically determining a set of regions of interest in a set of patient scans stored on computer readable storage media based on comparing the set of patient scans with a set of baseline scans stored on computer readable storage media;generating quantitative information describing diagnostic characteristics associated with a patient based on an analysis of the set of patient scans, wherein the quantitative information comprises a set of indicators of potential neuropsychiatric conditions, the set of indicators including one or more indicators associated with each region in the set of regions of interest in the set of scans for the patient including a first indicator associated with a first region of interest in the set of regions of interest and a second indicator associated with a second region of interest in the set of regions of interest, wherein the first indicator is of a different form than the second indicator;comparing each of the set of indicators of potential neuropsychiatric conditions with a set of diagnostic signatures, wherein a diagnostic signature comprises a set of indicators describing diagnostic characteristics of a known neuropsychiatric condition;identifying a plurality of matching signatures, wherein a matching signature is a diagnostic signature describing a first diagnostic characteristic that corresponds to a second diagnostic characteristic described by at least one indicator in the set of indicators to form a set of signatures; andidentifying a plurality of diagnoses including a diagnosis associated with each signature in the set of signatures to form a set of potential diagnoses. 2. The computer implemented method of claim 1 further a identifying a weighting associated with each matching signature, wherein a matching signature having a greater weighting is identified as a diagnosis having a higher degree of certainty than a diagnosis associated with a matching signature having a lower weighting. 3. The computer implemented method of claim 1 wherein generating the quantitative information comprises: deriving the quantitative information from a correlation of the regions of interest with portions of interest in the medical literature. 4. The computer implemented method of claim 1 wherein generating the quantitative information comprises: analyzing a medical history for the patient with the regions of interest and portions of interest in the medical literature to form the quantitative information, wherein the medical history comprises information describing previously diagnosed medical conditions of the patient, previously prescribed medications, and currently prescribed medications. 5. The computer implemented method of claim 1 wherein generating the quantitative information comprises: analyzing clinical data with the regions of interest and portions of interest in the me... | International Business Machines Corporation | A61B005/00 G06K009/00
US08660631 | B2 | 2014-02-25 | Torsional support apparatus and method for craniocaudal rotation of animals | 2009-06-01 | A torsional support apparatus is disclosed for craniocaudal rotation of test animals to enable multiple-view imaging. The animal is supported in a U-shaped loop of optically transparent material and the loop is moved to apply torsion to the animal to rotate it about its craniocaudal axis. Methods of imaging also are disclosed that use the torsional support technique. | 1. Apparatus for imaging an animal having a craniocaudal axis, comprising: an elongated flexible film for supporting such an animal in a position with its craniocaudal axis transverse to an axis of elongation of the elongated flexible film;a forming mechanism for limning the elongated flexible film into an upwardly open, U-shaped loop, the loop being sized for receiving and engaging such an animal;a moving mechanism for moving the elongated flexible film in a direction of the axis of elongation while maintaining the U-shaped loop, whereby movement of the elongated flexible film applies torsion to such an animal so that it is rotated about its craniocaudal axis; andan imaging system for imaging such an animal at various angles of craniocaudal rotation. 2. Apparatus according to claim 1, wherein the elongated flexible film is made from an optically transparent, flexible film. 3. Apparatus according to claim 1, wherein the forming mechanism for forming comprises: a pair of horizontally spaced, parallel drive shafts attached to the elongated flexible film at locations spaced along the axis of elongation; anda guide system for guiding the elongated flexible film to form the U-shaped loop between the drive shafts. 4. Apparatus according to claim 3, wherein the elongated flexible film has first and second longitudinally extending edges and the guide system comprises a pair of guides, one on each of the edges for forming the U-shaped loop. 5. Apparatus according to claim 4, wherein the U-shaped loop is tangent to a focal plane of the imaging system. 6. Apparatus according to claim 4, wherein the moving mechanism comprises a drive for rotating the drive shafts. 7. Apparatus according to claim 4, wherein each of the guides comprises a central passage, further comprising an opening on an upper side of the central passage of one guide for receiving a tail of such an animal and an anterior expandable tubing segment extending through the central passage of the other guide for engaging a head of such an animal. 8. Apparatus according to claim 7, further comprising a delivery system for introducing anesthesia into the anterior expandable tubing segment to anesthetize such an animal. 9. Apparatus according to claim 3, further comprising a base plate, mounting hub on the base plate for supporting the drive shafts and a window through the base plate through which light or radiation may pass toward the means for imaging. 10. Apparatus according to claim 9, wherein the imaging system comprises an X-ray source above the U-shaped loop and a phosphor screen movable beneath the window to receive radiation passing through such an animal in the U-shaped loop. 11. Apparatus according to claim 1, wherein there are a first plurality of elongated flexible films, a corresponding second plurality of forming mechanisms for forming, and a corresponding third plurality of moving mechanisms for moving, whereby a corresponding fourth plurality of such animals may be imaged.... | Bruker Biospin Corporation | A61B005/05
US08706178 | B2 | 2014-04-22 | Method and device for determining oxygen saturation of hemoglobin, for determining hematocrit of blood, and/or for detecting macular degeneration | 2006-05-02 | A method and device for accurately determining oxygen saturation of hemoglobin by the measurement of the optical density of a sample, such as a blood vessel, in response to illumination by light having at least three wavelengths (λ1, λ2, λ3, . . . ) within a range of about 460 nm to about 523 nm. The hematocrit of a sample may be determined from optical density measurements at the three or more wavelengths in conjunction with a known path length. The device may be an intravenous or intra-arterial fiber optic catheter used to deliver the interrogating light signal to the blood and to detect the reflected signal. A method and device of determining the thickness of the retinal well using spectroscopic information are also disclosed. | 1. A device for determining the oxygen saturation of blood, the device comprising a light source for illuminating blood with at least three distinct wavelengths (λ) of light, each within the range of about 460 nm and 523 nm; anda processing element configured to determine optical density of the blood sample at each of the at least three wavelengths and to determine a parabolic function based on the correspondence of measured optical density to the respective wavelengths said processing element also configured to determine the minimum wavelength (λmin) that corresponds to an inflection point of the parabolic function, and said processing element further configured to correspond the determined minimum wavelength (λmin) to an oxygen saturation concentration of the blood. 2. A method for determining the oxygen saturation of blood, comprising the steps of: providing a device according to claim 1;illuminating blood with at least three distinct wavelengths (λ) of light, each within the range of about 460 nm and 523 nm;measuring optical density of the blood sample at each of the at least three wavelengths;determining a parabolic function based on the correspondence of measured optical density to the respective wavelengths;determining the minimum wavelength (λmin) that corresponds to an inflection point of the parabolic function; and,corresponding the determined minimum wavelength (λmin) to an oxygen saturation concentration of the blood. 3. The method of claim 2, wherein the blood is illuminated in vitro or in vivo. | UAB Research Foundation The University of Alabama in Huntsville | A61B005/1455
US20100018525 | A1 | 2010-01-28 | Apparatus and method for increasing, monitoring, measuring, and controlling perspiratory water and solid loss at a reduced ambient pressure | 2009-07-22 | A device for increasing, monitoring, and measuring perspiration water and solid loss at reduced ambient pressure, comprising a sealed chamber capable of maintaining less than atmospheric pressure for an extended period of time and a gasket-sealed door accessing the chamber. The chamber provides a controlled hypobaric environment for treatment of various medical conditions. Internal controls of ambient pressure inside, temperature, humidity, oxygen, CO2 and other environmental conditions within the chamber are included, as well as a mechanism for recovering perspiration water loss from an occupant of the chamber. Methods for determining the perspiration volume of water and solid loss experienced by an occupant in the chamber are also disclosed including a computer program for determining quantity of sweat produced by the occupant of the chamber on a continuous basis. An algorithm allowing for continuous calculations of sweat loss and fluid replacement requirements of the occupant of the chamber is disclosed. | 1. A device for increasing, monitoring, and measuring perspiration water and solid loss at reduced ambient pressure, comprising: a sealed chamber capable of maintaining less than atmospheric pressure for an extended period of time; a gasket-sealed door accessing said chamber; means for reducing ambient pressure inside said chamber; and means for recovering perspiration water loss from an occupant of said chamber. 2. The device of claim 1, further including means for recovering perspiration solid loss from an occupant of said chamber. 3. The device of claim 1, further including means for monitoring and controlling the humidity inside said chamber. 4. The device of claim 1, further including means for monitoring and controlling the ambient temperature inside said chamber. 5. The device of claim 1, further including means for measuring and controlling the quality and quantity of oxygen supplied to said chamber. 6. The device of claim 1, further including means for computerized control of bioinstrumentation and environmental control systems in said chamber. 7. The device of claim 1, further including oxygen gas mask means for providing oxygen under certain treatment conditions to said occupant of said chamber. 8. The device of claim 1, further including sweat collection means for collecting sweat from said occupant of said chamber. 9. The device of claim 8, wherein said sweat collection means comprises a layered garment. 10. A method for determining the perspiration volume of water and solid loss experienced by an occupant in said chamber, comprising the steps of: measuring the water and solid input; measuring the beginning weight of said occupant in said chamber; measuring the ending weight of said occupant of said chamber; measuring the metabolic activity experienced by said occupant of said chamber; measuring water and solid loss experienced by said occupant of said chamber; and calculating the perspiration volume water and solid loss. 11. The method of claim 10, wherein said method includes a computer program means for determining quantity of sweat produced by said occupant of said chamber on a continuous basis. 12. The method of claim 11, wherein said computer program utilizes an algorithm allowing for continuous calculations of sweat loss and fluid replacement requirements of said occupant of said chamber. | - | A61G010/02 A61B005/00 A61M016/00
US20100036265 | A1 | 2010-02-11 | APPARATUS AND METHOD FOR MEASURING BLOOD PRESSURE | 2009-05-20 | A method of measuring blood pressure includes measuring a characteristic of body fat of a measured body part, detecting a sphygmus wave and a pressure of a blood vessel of the measured body part, detecting a base blood pressure of the measured body part based on a result of the detecting the sphygmus wave and the pressure of the blood vessel of the measured body part, and detecting a continuous blood pressure using a blood pressure calibration. The blood pressure calibration uses the base blood pressure and a blood pressure calculation regression equation including the characteristic of the body fat. | 1. An apparatus for measuring blood pressure, the apparatus comprising: a pressure sensor which detects a sphygmus wave and a pressure of a blood vessel of a measured body part;a light source which emits light toward the measured body part;a light detector which detects reflective light reflected from the measured body part;a body fat measuring unit which calculates a characteristic of body fat of the measured body part based on an intensity of the reflective light;a base blood pressure detecting unit which detects a base blood pressure of the measured body part based on a characteristic of the sphygmus wave and the pressure of the blood vessel detected by the pressure sensor; anda continuous blood pressure detecting unit which detects a continuous blood pressure based on the base blood pressure and a blood pressure calculation regression equation including the characteristic of the body fat calculated by the body fat measuring unit. 2. The apparatus of claim 1, wherein the characteristic of the body fat is a thickness of the body fat. 3. The apparatus of claim 1, wherein the base blood pressure detecting unit detects the base blood pressure using an oscillometric method. 4. The apparatus of claim 1, wherein the continuous blood pressure detecting unit detects the continuous blood pressure using a tonometric method. 5. The apparatus of claim 1, further comprising a compression unit which compresses the blood vessel of the measured body part. 6. The apparatus of claim 1, wherein the measured body part is a wrist, andthe blood vessel of the measured body part is a radial artery of the wrist. 7. The apparatus of claim 6, further comprising a wrist band which receives the wrist. 8. The apparatus of claim 1, wherein the light source comprises a light emitting diode. 9. The apparatus of claim 1, further comprising a plurality of light sources, wherein light sources of the plurality of light sources are spaced apart from each other,the light detector is spaced apart from each of the light sources by gaps having different respective lengths,the light detector detects reflective lights corresponding to lights emitted from each of the light sources, andthe body fat measuring unit calculates the characteristic of the body fat of the measured body part based on an intensity of the reflective lights. 10. The apparatus of claim 1, wherein the lights emitted from the light sources have a same wavelength band. 11. The apparatus of claim 9, wherein, when a gap between the light detector and a first light source is greater than a gap between the light detector and a second light source, an amount of light emitted from the first light source is greater than an amount of light emitted from the second light source. 12. A method of measuring blood pressure, the method comprising: measuring a characteristic of body fat of a measured body part;detecting a sphygmus wave and a pressure of a blood vessel of the measured body part;detecting a base blood pressure ... | SAMSUNG ELECTRONICS CO., LTD. | A61B005/02
US20100004517 | A1 | 2010-01-07 | NON-INVASIVE METHOD AND APPARATUS FOR DETERMINING A PHYSIOLOGICAL PARAMETER | 2009-09-14 | The present invention relates to an apparatus and method for the non-invasive analysis of physiological attributes, such as heart rate, blood pressure, cardiac output, respiratory response, body composition, and blood chemistry analytes including glucose, lactate, hemoglobin, and oxygen saturation. Using a combination of multi-functioning disparate sensors, such as optical and electrical, improvements are made over existing physiological measurement devices and techniques. The special configuration of one or more multi-functional sensors is used to non-invasively measure multi-wavelength optical plus one or more of ECG, Bio-impedance, and RF-impedance spectroscopic data. This information is used to develop self-consistent, non-linear algorithm in order to derive the physiological attributes while compensating for various forms of interfering effects including motion artifacts, sensor attachment variability, device component variability, subject physical and physiology variability, and various interfering physiological attributes. | 1. A method of noninvasively measuring a global body level physiological parameter of a subject comprising: a) locating respective noninvasive sensors on the body of the subject, said sensors outputting separate signals representing respective measured values of at least two disparate measurable physical properties of the subject, the value of each of said disparate physical properties varying in dependence on said physiological parameter and being independently capable of giving a measurement of said body level physiological parameter;b) obtaining calibration data reflecting the effect of changes in said physiological parameter on the values of each of said at least two disparate measurable physical properties so as to permit the identification of the values of said body level physiological parameter corresponding to particular combinations of values of said disparate measurable physical properties;c) processing said signals representing said respective measured values of said disparate physical properties in a processor configured to transform said signals into a final value signal representing a quantitative measurement of said body level physiological parameter that is derived from a plurality of said signals and based on the particular combination of the measured values of said disparate properties represented by said signals, whereby said final value is based on the aggregate effect of changes in said physiological parameter on said at least two disparate physical properties predetermined in step b); andd) outputting said final value of said body level physiological parameter. 2. An apparatus for noninvasively measuring a global body level physiological parameter of a subject comprising: a) noninvasive sensors for location on the body of the subject, said sensors outputting separate signals representing respective measured values of at least two disparate measurable physical properties of the subject, the value of each of said disparate physical properties varying in dependence on said physiological parameter and being independently capable of giving a measurement of said body level physiological parameter;b) a memory storing calibration data reflecting the effect of changes in said physiological parameter on the values of each of said at least two disparate measurable physical properties so as to permit the identification of the values of said body level physiological parameter corresponding to particular combinations of values of said disparate measurable physical properties;c) a processor configured to processing said signals representing said respective measured values of said disparate physical properties so as to transform said signals into a final value signal representing a quantitative measurement of said body level physiological parameter that is derived from a plurality of said signals and based on the particular combination of the measured values of said disparate properties represented by said signals, whereby said final value... | BIOPEAK CORPORATION | A61B005/00
US20100022856 | A1 | 2010-01-28 | IMPLANTABLE OPTICAL HEMODYNAMIC SENSOR INCLUDING LIGHT TRANSMISSION MEMBER | 2008-07-28 | An implantable medical device (IMD) includes an optical hemodynamic sensor comprising at least one optical emitter and at least one detector. In some examples, the at least one optical emitter may be optically coupled to at least one light transmission member that extends from a housing of the IMD. In addition, in some examples, the at least one detector may be optically coupled to at least one light transmission member that extends from the housing of the IMD. In other examples, an optical emitter and/or detector of a hemodynamic sensor may be carried by an extension member that extends from a housing of the IMD. The elongated member may electrically couple the optical emitter and/or detector to a controller or other components within the IMD housing. | 1. An implantable medical system comprising: an implantable housing;a light transmission member extending from the housing;an optical emitter coupled to the housing; anda detector coupled to the housing, wherein the detector is configured to generate a signal indicative of an intensity of light emitted by the optical emitter and transmitted to the detector,wherein at least one of the optical emitter or the detector are optically coupled to the light transmission member. 2. The implantable medical system of claim 1, wherein the light transmission member comprises at least one of an optical fiber or an optical waveguide. 3. The implantable medical system of claim 1, wherein the optical emitter comprises at least one of a light emitting diode, a laser diode or a vertical cavity surface emitting laser. 4. The implantable medical system of claim 1, wherein the light transmission member is movable relative to the housing. 5. The implantable medical system of claim 1, wherein the light transmission is substantially rigid. 6. The implantable medical system of claim 1, wherein the light transmission member has a smaller cross-sectional size than the housing. 7. The implantable medical system of claim 1, wherein the optical emitter is coupled to the light transmission member. 8. The implantable medical system of claim 1, wherein the detector is coupled to the light transmission member. 9. The implantable medical system of claim 1, wherein the light transmission member comprises a first light transmission member, the system further comprising a second light transmission member, wherein the optical emitter is optically coupled to the first light transmission member and the detector is optically coupled to the second light transmission member. 10. The implantable medical system of claim 9, wherein the first and second light transmission members are substantially parallel. 11. The implantable medical system of claim 1, wherein the light transmission member comprises a first light transmission member, the system further comprising a second light transmission member, wherein the at least one of the optical emitter or the detector is optically coupled to the first and second light transmission members. 12. The implantable medical system of claim 11, further comprising a third light transmission member optically coupled to the at least one of the optical emitter or the detector, wherein the first light transmission member comprises a first opening for at least one of receiving or emitting light, the second light transmission member comprises a second opening for at least one of receiving or emitting light, and the third light transmission member comprises a third opening for at least one of receiving or emitting light, and wherein the first, second, and third openings define at least one of a two-dimensional array or a three-dimensional array. 13. The implantable medical system of claim 1, further comprising a light deflection member optically coupled ... | Medtronic, Inc. | A61B005/1455
US20100016675 | A1 | 2010-01-21 | Method of assessing a condition using sucking patterns | 2008-07-18 | A method of assessing a condition includes comparing a sucking pattern for a target infant with a sucking pattern associated with the condition. A method of monitoring progress of a treatment includes comparing the rhythmic sucking percentage for a target infant before treatment for a condition with the rhythmic sucking percentage for the infant after treatment of the condition to assess the effectiveness of the treatment. A method of monitoring progress of a treatment includes using sucking patterns to compare an infant's rate of progress at a given point in time to a population's rate of progress at the same point in time. | 1. A method of assessing a condition in a target infant comprising, acquiring a first sucking pattern associated with the condition;acquiring a second sucking pattern for the target infant;comparing the second sucking pattern with the first sucking pattern; anddetermining if the second sucking pattern is consistent with the first sucking pattern. 2. The method of claim 1 wherein the first sucking pattern is an aggregate pattern derived from a population of infants expressing the condition. 3. The method of claim 1 wherein the second sucking pattern is an aggregate pattern derived from the target infant on two or more occasions. 4. The method of claim 1 wherein at least one of the first and the second sucking pattern is acquired with a pacifier adapted to measure at least one of intensity, frequency, pauses, number of sucks per envelope, or envelope function of the sucking patterns. 5. The method of claim 4 wherein the second sucking pattern is acquired in the target infant's residence and the comparing step and the determining steps are performed by a computer adapted to receive data from the pacifier. 6. The method of claim 1 wherein the condition is colic or cognitive alertness. 7. The method of claim 1 wherein the determining step is a subjective evaluation performed by a trained technician. 8. The method of claim 1 wherein the determining step is an objective evaluation based, at least in part, on pattern analyzing software, artificial intelligence, or neural networks. 9. The method of claim 1 wherein the second sucking pattern is acquired in a clinic and the comparison is performed in a clinic. 10. The method of claim 1 wherein the second sucking pattern is acquired in a first location and the comparing step is performed at a second location different than the first location. 11. The method of claim 10 wherein the first location is the infant's residence and wherein the second location is a clinic. 12. A method of monitoring progress of a treatment of a target infant having a condition comprising, acquiring a first sucking pattern for the target infant over a first period of time;calculating a first rhythmic sucking pattern percentage for the first period of time;acquiring a second sucking pattern for the target infant over a second period of time after the treatment of the target infant for the condition;calculating a second rhythmic sucking pattern percentage for the second period of time;comparing the first rhythmic sucking pattern percentage to the second rhythmic sucking pattern percentage; andassessing the effectiveness of the treatment. 13. The method of claim 12 further comprising acquiring a third sucking pattern for the target infant over a third period of time after a second treatment of the target infant for the condition, wherein the second treatment is different than the first treatment,calculating a third rhythmic sucking pattern percentage for the third period of time;comparing the third rhythmic sucking pattern p... | - | A61B005/00
US20100030102 | A1 | 2010-02-04 | Active Delivery and Flow Redirections: Novel Devices and Method of Delivery of Materials to Patients | 2007-05-14 | A medical device and method for planning or performing a method for delivering material through tissue into a defined area of a patient may comprise: a material delivery element through which the material may flow out of a delivery end; and observing the migration, flow and persistence of material delivered and developing an plan or optimizing a plan for the delivery of material into the defined area. Novel catheter devices are provided to support these methods. | 1. A method for the provisioning and positioning of a flowable material into a region of a patient comprising: identifying a region of a patient to be viewed or treated;identifying a region wherein the region forms a potential volume between two opposed different and distinct surfaces;penetrating at least one of the two opposed different and distinct surfaces with a material delivery device; andproviding material from the material delivery device into the potential volume to create a volume containing at least delivered material. 2. The method of claim 1 wherein the volume is mass transfer stable in that less than 80% by volume of delivered material is removed from the created volume by natural biological activity in less than 5 minutes. 3. The method of claim 1 wherein the material delivery device forms an at least partial seal around a puncture formed by penetration of only one of the at least two opposed different and distinct surfaces. 4. The method of claim 3 wherein the seal is formed covering a surface area extending circumferentially away from and around a diameter of the material delivery device. 5. The method of claim 3 wherein the seal is formed on one or both sides of the puncture relative to the opposed surface penetrated. 6. The method of claim 4 wherein a structure forming the seal is a) inflated to assist in sealing the puncture, b) is flexible and distorts to apply pressure over the puncture, c) seal is formed by expansion of a component on the material delivery device around a region of a puncture created by the penetration or d) seal is formed by expansion of a component on the material delivery device around a region of a puncture created by the penetration, and the component is selected from a parasol or a balloon. 7. The method of claim 1 wherein the two opposed surfaces comprise the pia and the cortex. 8. A medical device for delivering material through tissue into a defined area of a patient that is between two adjacent and different opposed tissue surfaces, the area having a potential volume upon injection of a liquid comprising: a material delivery element through which the material may flow out of a delivery end;the delivery end having an opening that can be inserted through a surface of the tissue; anda sealing system proximal to the delivery end that can extend radially away from the material delivery element while under operator control through the material delivery device and move against the surface of the tissue and apply pressure to the surface of the tissue after the material delivery element has been inserted through the surface of the tissue to provide a seal over a hole or puncture through which the delivery device was inserted through the tissue and redirect fluid flow along an interface between the two different opposed surfaces and sustain a volume in the area. 9. The device of claim 8 wherein the material delivery device is selected from the group consisting of a catheter and needle. 10. The devi... | - | A61B005/03 A61M025/00 A61M025/10 G06F015/18 G06N005/04
US20100049026 | A1 | 2010-02-25 | FIXATION OF A MEDICAL IMPLANT TO THE EXTERIOR OF A BODY ORGAN | 2009-09-18 | In general, invention is directed to devices for use in surgical procedures in which a medical implant is affixed to or implanted within an exterior surface of a body organ. A system, for example, is described that includes a cannula and a delivery instrument disposed within the cannula to fix a medical implant to an exterior surface of an organ. The delivery instrument has a distal end including a cavity and a vacuum port to draw a portion of the exterior surface of the organ into the cavity. The medical implant is affixed to the portion of the exterior surface drawn into the cavity of the delivery instrument. | 1. A system comprising: an electrode comprising:a conductive material defining a cavity; and a vacuum inlet that receives vacuum pressure from a delivery instrument and applies the vacuum pressure to the cavity to draw tissue from an exterior surface of an organ of a patient into the cavity. 2. The system of claim 1, wherein the electrode further comprises an insulative backing coupled to the conductive material. 3. The system of claim 1, further comprising a wireless transceiver that communicates with an external device. 4. The system of claim 1, further comprising a pulse generator that delivers electrical stimulation to the conductive material. 5. The system of claim 1, further comprising a lead electrically coupled to the conductive material. 6. The system of claim 1, further comprising the delivery instrument, wherein the delivery instrument defines a chamber sized to hold the electrode, wherein when the electrode is held in the chamber, the cavity defined by the conductive material is exposed. 7. The system of claim 1, further comprising a pin that extends through the conductive material and the cavity to fix the electrode to the tissue drawn into the cavity. 8. The system of claim 1, wherein the electrode further comprises an eyelet that receives a securing structure to secure the electrode to the tissue. 9. The system of claim 1, further comprising: a cannula capable of being introduced into an abdomen of the patient via an incision in the abdomen; andwherein the delivery instrument is disposed within the cannula, the delivery instrument comprising a vacuum outlet through which the vacuum pressure is applied to the vacuum inlet of the electrode to draw a portion of the exterior surface of the organ into the cavity defined by the conductive material. 10. The system of claim 9, further comprising: a source of vacuum pressure coupled to the delivery instrument; anda conduit through the delivery instrument to apply the vacuum pressure from the source to the exterior surface of the organ. 11. The system of claim 9, further comprising a trocar disposed within the cannula, wherein the delivery instrument is disposed within the trocar. 12. The system of claim 1, further comprising: means for introducing a delivery instrument into the patient via an incision in an abdomen of the patient; andmeans for applying vacuum pressure to an exterior surface of an organ to draw at least a portion of the exterior surface into the cavity defined by the conductive material. 13. The system of claim 12, further comprising means for affixing the electrode to the portion of the exterior surface drawn into the cavity. 14. The system of claim 13, further wherein the means for affixing the electrode comprises means for anchoring the electrode to the portion of the exterior surface drawn into the cavity defined by the conductive material. | Medtronic, Inc. | A61B005/04
US20100049036 | A1 | 2010-02-25 | Medical Image Display Method and Program Thereof | 2007-03-06 | It has been difficult to display medical three-dimensional images taken by different imaging devices (imaging methods or modalities) individually in a common display direction. In the invention, a coordinate system 1 on the basis of landmarks on the bones and a coordinate system 2 on the basis of landmarks on the skins are set in a CT image, and a transformation coefficient (correction amount) Tc between the coordinate system 1 and the coordinate system 2 is calculated and stored. In a case where image data acquired by another imaging method is displayed, a coordinate system same as the coordinate system 1 in the CT image can be set in a case where the image data includes a bone display layer, and a coordinate system same as the coordinate system 2 in the CT image can be set in a case where the image displays a skin display layer. By correcting the coordinate system thus set using the transformation coefficient Tc, the coordinate system 1 can be corrected to the coordinate system 2 or the coordinate system 2 to the coordinate system 1. It is thus possible to display images according to the common coordinate system between different sets of image data. | 1. A display method of a medical image according to image data acquired by different imaging methods, comprising the steps of: displaying a first image according to first medical three-dimensional image data of a patient acquired by a first imaging method;setting a first reference coordinate system used to determine a direction of the first image on the basis of a plurality of first landmarks included in the first image;setting a second reference coordinate system used to determine the direction of the first image on the basis of a plurality of second landmarks included in the first image and having at least one landmark different from the first landmarks;computing a difference of directions between the first reference coordinate system and the second reference coordinate system and calculating a transformation coefficient between the both coordinate systems on the basis of the difference;displaying a second image according to second medical three-dimensional image data of the patient acquired by a second imaging method;setting a third reference coordinate system used to determine a direction of the second image on the basis of a plurality of third landmarks included in the second image;correcting the third reference coordinate system to a fourth reference coordinate system using the transformation coefficient that has been calculated; anddisplaying the second image in a direction according to the fourth reference coordinate system set by correction. 2. A display method of a medial image, comprising the steps of: preparing a transformation coefficient on the basis of a difference between a pre-calculated coordinate system in reference to landmarks on hard tissues represented by bones and a pre-calculated coordinate system in reference to landmarks on soft tissues represented by skins;constructing a first reference coordinate system used to determine a direction of an image on the basis of a plurality of landmarks on an image showing the soft tissues in a case where a medical image of a patient taken by a predetermined imaging method is an image that shows the soft tissues represented by the skins but does not show the hard tissues represented by the bones;correcting the first reference coordinate system to a second reference coordinate system using the transformation coefficient; anddisplaying an image of the patient in a direction according to the second reference coordinate system. 3. A display method of a medical image according to image data acquired by different imaging methods, comprising the steps of: computing a difference between two coordinate systems set on first medical three-dimensional image data acquired by a first imaging method and calculating a transformation coefficient between the both coordinate systems on the basis of the difference;setting a coordinate system common with either one of the two coordinate systems on second medical three-dimensional image data acquired by a second imaging method; and correcting the coordinat... | IMAGNOSIS INC. | A61B005/05 G06T015/00
US20100036281 | A1 | 2010-02-11 | Integrated Sensor and Lancet Device and Method for Collecting Body Fluid Using the Same | 2006-11-30 | The present invention relates to a sensor/lancet integrated device 1 including a lancet including a needle part 14 for puncturing skin, and a sensor including a reagent part 17. The lancet is integrally molded with the needle part 14. The device 1 preferably further includes a capillary 15 for supplying to the reagent part 17. Preferably, the needle part 14 is formed to a hollow form, and has the interior communicated to an introduction port. The needle part 14 may be formed projecting from a location different from the introduction port. | 1. A sensor/lancet integrated device comprising a lancet including a needle part for puncturing skin, and a sensor including a reagent part; wherein the lancet is integrally molded with the needle part. 2. The sensor/lancet integrated device according to claim 1, wherein the sensor further includes an electrode integrally incorporated in the lancet. 3. The sensor/lancet integrated device according to claim 1, wherein the sensor and the lancet are formed as separate bodies; andthe sensor is held in an internal space formed in the lancet to be integrated to each other. 4. The lancet/lancet integrated device according to claim 1, wherein the sensor has a mode in which a cover is joined with respect to a substrate;the lancet includes a joint for attaching to the sensor; andthe joint is fixed to the sensor so as to sandwich the substrate. 5. The lancet/lancet integrated device according to claim 4, wherein the joint is fixed to the sensor so as to further sandwich the cover. 6. The lancet/lancet integrated device according to claim 5, further comprising: a capillary for supplying body fluid to the reagent part; whereinthe needle part is projected from a location different from an introduction port of the body fluid in the capillary. 7. The sensor/lancet integrated device according to claim 1, wherein the reagent part includes a color producing agent and is formed in an internal space formed in the lancet; andat least one part of the lancet is formed transparent to measure a color of the color producing agent. 8. The sensor/lancet integrated device according to claim 1, wherein the needle part is formed hollow. 9. The lancet/lancet integrated device according to claim 8, further comprising: a capillary for supplying body fluid to the reagent part; whereinthe interior of the needle part is communicated to the interior of the capillary. 10. The sensor/lancet integrated device according to claim 8, wherein the interior of the needle part is filled with resin before use. 11. The lancet/lancet integrated device according to claim 9, further comprising: a capillary for supplying body fluid to the reagent part; whereinthe lancet includes a groove for defining the capillary. 12. A sensor/lancet integrated device comprising a lancet including a needle part for puncturing skin, and a sensor including an introduction port for introducing body fluid; wherein the needle part is formed projecting from a location different from the introduction port. 13. The sensor/lancet integrated device according to claim 12, wherein the needle part is integrally molded to a component of the sensor. 14. The sensor/lancet integrated device according to claim 12, further comprising: a cover to be joined to the sensor; whereinthe cover includes a groove for defining a capillary for introducing the body fluid. 15. The sensor/lancet integrated device according to claim 12, wherein the sensor further includes an electrode. 16. The sensor/lancet integrated device according... | ARKRAY, INC. | A61B005/151
US20100041984 | A1 | 2010-02-18 | Impedance sensing device and catheter system | 2008-08-12 | A sensing system including a collection catheter and a sensing device. The sensing system can be used in determining an optimal location of a collection catheter for the removal of a medium from coronary circulation. The sensing system can also be used in determining the optimal time for the collection of blood containing a medium. The sensing system can further be used to maintain a baseline flow rate through at least a portion of the coronary circulation system during a medical procedure. | 1. A method of locating a collection catheter within a coronary sinus, the method comprising the steps of: a) providing a sensing device at a distal end of the collection catheter, the sensing device including at least a first pair of electrodes;b) positioning a distal end of the collection catheter within the coronary sinus at a first position;c) monitoring the impedance measured by the first pair of electrodes to detect the presence of a medium flowing into the coronary sinus;d) determining by the impedance measurements from which particular vein the medium is flowing; ande) moving the collection catheter to a second position such that an opening at the distal end of the collection catheter is located downstream of the antegrade flow from the particular vein. 2. The method of claim 1, further including extending the sensing device through the opening located at the distal end of the collection catheter. 3. The method of claim 2, further including extending the sensing device from a first extension position to a second extension position relative to the collection catheter to determine by the impedance measurements from which particular vein the medium is flowing. 4. The method of claim 1, further including inflating a balloon when the collection catheter reaches the second position to occlude the coronary sinus during operation of the collection catheter. 5. The method of claim 4, further including removing the medium from the coronary sinus by operation of the collection catheter. 6. The method of claim 5, further including deploying a vessel support device when the collection catheter reaches the second position to maintain the patency of the coronary sinus during operation of the collection catheter. 7. The method of claim 1, wherein the sensing device further includes a second pair of electrodes, and wherein the step of monitoring the impedance includes monitoring the impedance measured by each of the first and second pairs of electrodes, the first pair of electrodes being located farther from the opening of the collection catheter than the second pair of electrodes. 8. The method of claim 7, wherein the step of determining from which particular vein the medium is flowing includes comparing a first impedance measured by the first pair of electrodes with a second impedance measured by the second pair of electrodes, and wherein an impedance measurement generally equal to that of blood indicates no medium is present at that corresponding electrode pair, and wherein an impedance measurement different from that of blood indicates medium is present at that corresponding electrode pair. 9. The method of claim 1, further including providing a vessel support device that maintains the patency of a vessel during removal of the medium, the sensing device being carried by the vessel support device. 10. The method of claim 1, wherein the medium is a contrast media. 11. An impedance sensing system, comprising: a) an elongated collection tube hav... | - | A61B005/055 A61M025/10 A61B005/05
US20100041964 | A1 | 2010-02-18 | Methods and systems for monitoring and modifying a combination treatment | 2008-09-30 | Methods, computer program products, and systems are described that include monitoring at least one health attribute of an individual during an artificial sensory experience, associating a characteristic of the artificial sensory experience with the at least one health attribute of the individual, and/or modifying at least one of a bioactive agent or the artificial sensory experience at least partly based on associating a characteristic of the artificial sensory experience with the at least one health attribute of the individual. | 1-78. (canceled) 79. A system, comprising: a monitorer module;an associator module; anda modifier module configured to modify at least one of a bioactive agent or an artificial sensory experience at least partly based on associating a characteristic of the artificial sensory experience with the at least one health attribute of an individual. 80. The system of claim 79, wherein the monitorer module comprises: a data receiver module. 81. The system of claim 79, wherein the monitorer module comprises: a physiological measurement monitorer module. 82. The system of claim 79, wherein the monitorer module comprises: a neurophysiological activity monitorer module. 83. The system of claim 82, wherein the neurophysiological activity monitorer module comprises: a physiologic activity measurer module. 84. The system of claim 79, wherein the monitorer module comprises: a recorder module. 85. The system of claim 79, wherein the monitorer module comprises: an observer module. 86. The system of claim 79, wherein the associator module comprises: a physiological response associator module. 87. The system of claim 86, wherein the physiological response associator module comprises: an artificial sensory experience characteristic associator module. 88. The system of claim 79, wherein the associator module comprises: a report acceptor module. 89. The system of claim 79, wherein the associator module comprises: an algorithm utilizer module. 90. The system of claim 79, wherein the associator module comprises: a comparer module. 91. The system of claim 79, wherein the modifier module configured to modify at least one of a bioactive agent or an artificial sensory experience at least partly based on associating a characteristic of the artificial sensory experience with the at least one health attribute of an individual comprises: an access modifier module. 92. The system of claim 91, wherein the access modifier module comprises: a restrictor module. 93. The system of claim 91, wherein the access modifier module comprises: a granter module. 94. The system of claim 79, wherein the modifier module configured to modify at least one of a bioactive agent or an artificial sensory experience at least partly based on associating a characteristic of the artificial sensory experience with the at least one health attribute of an individual comprises: a visual object modifier module. 95. The system of claim 94, wherein the visual object modifier module comprises: a color modifier module. 96. The system of claim 94, wherein the visual object modifier module comprises: a text modifier module. 97. The system of claim 79, wherein the modifier module configured to modify at least one of a bioactive agent or an artificial sensory experience at least partly based on associating a characteristic of the artificial sensory experience with the at least one health attribute of an individual comprises: an alterer module. 98. The system of claim 97, wherein the alterer module co... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/00 A61M021/00
US07654985 | B2 | 2010-02-02 | Controlled detachment of intra-luminal medical device | 2004-03-30 | An intra-luminal medical device includes a fixation mechanism to attach the medical device to tissue within a body lumen, and a detachment mechanism to permit selective detachment of the medical device from the tissue attachment site without the need for endoscopic or surgical intervention. An electromagnetic device may be provided to mechanically actuate the detachment mechanism. Alternatively, a fuse link may be electrically blown to detach the medical device. As a further alternative, a rapidly degradable bonding agent may be exposed to a degradation agent to detach the medical device from a bonding surface within the body lumen. The medical device may eliminate problems associated with uncertain and inconsistent detachment of intra-luminal medical devices. | 1. A medical device comprising: a device housing sized for introduction into and residence completely within the body lumen;a fixation mechanism to attach the device housing to a surface within the body lumen, wherein the fixation mechanism includes a cavity formed in the device housing and a shaft to capture luminal tissue within the cavity and a spring to bias the shaft toward the tissue;a controlled detachment mechanism mechanically actuating the fixation mechanism to selectively detach the device housing from the surface of the body lumen without endoscopic intervention, wherein the detachment mechanism includes a solenoid coil wound about the shaft and a circuit to energize the solenoid coil to drive the shaft against the spring bias and thereby release the luminal tissue; anda controller responsive to a control signal, wherein the controller activates the controlled detachment mechanism, and wherein the medical device remains completely within the body lumen until after the device is detached from the surface. 2. The medical device of claim 1, wherein the cavity includes a vacuum port for application of vacuum pressure to draw the tissue into the cavity. 3. The medical device of claim 1, wherein the circuit is responsive to a control signal to energize the solenoid coil. 4. The medical device of claim 1, further comprising a power source to power the detachment mechanism. 5. The medical device of claim 4, wherein the power source includes a battery. 6. The medical device of claim 4, wherein the power source includes an inductive coupling circuit to generate power from an inductive element external to the body lumen. 7. The medical device of claim 1, wherein the controller includes a telemetry circuit to receive the control signal as a telemetry signal from an external controller. 8. The medical device of claim 1, wherein the controller includes an inductive coupling circuit to sense the presence of an external inductive element as the control signal. 9. The medical device of claim 1, wherein the controller includes an inductive coupling circuit to generate power from an inductive element external to the body lumen and thereby drive the detachment mechanism with the generated power. 10. The medical device of claim 1, wherein the device housing is sized for introduction into the esophagus. 11. The medical device of claim 1, wherein the device housing is sized for passage through the gastrointestinal tract. 12. The medical device of claim 1, further comprising a sensor, mounted to the device housing, to sense at least one condition within the body lumen. 13. The medical device of claim 1, further comprising a sensor, mounted to the device housing, to sense at least one of pH, flow, temperature, and pressure within the body lumen. 14. The medical device of claim 1, further comprising: an electrical pulse generator, mounted within the device housing, to generate an electrical stimulation waveform;one or more electrodes electrically ... | Given Imaging Ltd. | A61B005/05 A61M005/32
US07643866 | B2 | 2010-01-05 | Method for producing a computed tomography display of tissue structures by applying a contrast medium | 2005-12-15 | A method is disclosed for producing a computed tomography display of tissue structures by applying a contrast medium. A contrast medium is applied to a patient for better visualization of the tissue structure to be examined. An X-ray scan is performed during the presence of the one contrast medium, and computed tomography 2- or 3-dimensional pictures are subsequently reconstructed from the X-ray scan data obtained. In the method, absorption data are measured for at least two different energy spectra, a computed tomography intermediate image is reconstructed per energy spectrum, and the distribution of the one contrast medium in the tissue is determined from the different energy-specific absorption behavior between tissue and contrast medium. | 1. A method of determining a distribution of a contrast medium in a tissue structure, an X-ray scan being performed during the presence of the contrast medium, comprising: measuring absorption data for at least two different energy spectra;reconstructing a computed tomography intermediate image per energy spectrum;determining the distribution of the contrast medium in the tissue structure from different energy-specific absorption behavior between the tissue structure and the contrast medium, anda calculating significancy statement referring to the distribution of the contrast medium from the at least two intermediate images,by forming the quotient of the intermediate images. 2. The method as claimed in claim 1, wherein display image values with a positive significancy statement referring to the presence of contrast medium obtain a predetermined image value in a CT image. 3. The method as claimed in claim 2, wherein the image value is at least one of a maximum and minimum gray scale value. 4. The method as claimed in claim 2, wherein the image value is a color value that is not present in a remaining image display. 5. The method as claimed in claim 1, wherein a dedicated radiation source is used per spectrum to determine the absorption data for the at least two different energy spectra. 6. The method as claimed in claim 1, wherein a dedicated anode material per spectrum is used in the radiation source to determine the absorption data for the at least two different energy spectra. 7. The method as claimed in claim 1, wherein a combination of anode material and contrast medium is used, the anode material having an intensity maximum in an bremsstrahlung spectrum, and the contrast medium having an absorption maximum in an energy range of the intensity maximum of the bremsstrahlung spectrum generated by the anode material. 8. The method as claimed in claim 1, wherein the measuring uses a spring focus. 9. The method as claimed in claim 1, wherein a dedicated detector is used per spectrum to determine the absorption data for the at least two different energy spectra. 10. The method as claimed in claim 1, wherein to determine the absorption data for the at least two different spectral energies, at least one detector that can resolve the energy spectrum of the detected radiation for at least two different energies is used. 11. A method for producing a computer tomography display of a tissue structure by applying a contrast medium, the method comprising: applying the contrast medium to a patient for better visualization of the tissue structure to be examined;performing an X-ray scan during the presence of the contrast medium;measuring absorption data for at least two different energy spectra;reconstructing a computed tomography intermediate image per energy spectrum;determining the distribution of the contrast medium in the tissue structure from different energy-specific absorption behavior between the tissue structure and the contrast medium;reco... | Siemens Aktiengesellschaft | A61B005/05
US07653558 | B2 | 2010-01-26 | Consolidation of consumer interactions within a medical brokerage system | 2008-04-07 | In general in one aspect, agenda items are generated based on a consumer's one or more interactions with a system. The agenda items are consolidated into an agenda, a graphical user interface is produced that when rendered on a display displays the agenda, and a request from the consumer to interact with one or more of the agenda items is received. | 1. A computer-implemented method, the method comprises: receiving, by one or more computers, a medical history for a consumer of medical services from one or more health data repositories and the consumer's one or more interactions with a medical brokerage system;applying a rules engine to the medical history for the consumer of medical services;generating, based on the application of the rules engine to the medical history, a list of one or more medical issues associated with the consumer of medical services;applying the rules engine to the list of one or more medical issues associated with the consumer of medical services;generating, based on the application of the rules engine to the list of one or more medical issues, a subset of the list of medical issues, the subset including medical issues that the provider of medical services is qualified to address with the consumer of medical services;producing, by one or more computers, a graphical user interface that when rendered on a display displays the subset of the list of medical issues;sending, by one or more computers, the subset of the list of the medical issues to one or more of the providers of medical services and the consumer of medical services; andestablishing a real-time communication channel between the consumer of medical services and the provider of medical services. 2. The method of claim 1 wherein generating a subset of the list of medical issues comprises: applying rules to the consumer's list of medical issues based on one or more attributes of the list of medical issues. 3. The method of claim 1 wherein the generation of the subset of the list of medical issues adds new medical issues to a prior list of medical issues. 4. The method of claim 2 wherein applying rules comprises joining new medical issues associated with the consumer with prior medical issues associated with the consumer. 5. The method of claim 2 wherein applying rules comprises: determining a date of removal from the list of medical issues for the one or more medical issues included on the list; andremoving, on the determined date of expiration, one or more medical issues from the list of medical issues. 6. The method of claim 1 wherein the medical history includes the consumer's prior interactions with the system. 7. The method of claim 1 wherein the consumer's one or more interactions with a medical brokerage system include the medical brokerage system's recommended future interactions with the system for the consumer. 8. The method of claim 1 wherein the consumer's one or more interactions with a medical brokerage system include one or more of the consumer's appointments. 9. The method of claim 1 wherein the consumer's one or more interactions with a medical brokerage system include one or more waiting room interactions. 10. The method of claim 1 wherein the consumer's one or more interactions with a medical brokerage system include conversations with one or more providers of services. 11. A computer... | American Well Inc. | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07647084 | B2 | 2010-01-12 | Medical sensor and technique for using the same | 2006-07-28 | A sensor assembly is provided that includes a skeletal frame comprising a loop structure. The sensor assembly includes at least one physiological sensor attached to the loop structure. A coating is provided over the skeletal frame and the at least one physiological sensor to form a sensor assembly. The sensor may be placed on a patient's finger, toe, and so forth to obtain pulse oximetry or other physiological measurements. Methods of manufacturing the sensor and sensor body and of applying the sensor are also provided. | 1. A sensor assembly, comprising: a skeletal frame comprising a loop structure;at least one physiological sensor attached to the loop structure; anda coating molded over the skeletal frame and the at least one physiological sensor to form a sensor assembly. 2. The sensor assembly of claim 1, wherein the at least one physiological sensor comprises at least one of a light emitting diode or a photodetector. 3. The sensor assembly of claim 1, wherein the sensor assembly comprises at least one of a pulse oximetry sensor, a sensor for measuring a water fraction, or a combination thereof. 4. The sensor assembly of claim 1, wherein the skeletal frame comprises a single molded part. 5. The sensor assembly of claim 1, wherein the frame comprises two or more parts joined to form the skeletal frame. 6. The sensor assembly of claim 1, wherein the coating comprises a thermoplastic elastomer. 7. The sensor assembly of claim 6, wherein the thermoplastic elastomer comprises at least one of a thermoplastic polyolefin, a thermoplastic vulcanizate alloy, thermoplastic polyurethane, silicone, or a combination thereof. 8. The sensor assembly of claim 1, wherein the coating comprises a conformable material. 9. The sensor assembly of claim 1, wherein no moving parts of the frame remain exposed from the coating. 10. The sensor assembly of claim 1, wherein the sensor assembly is fluid tight. 11. The sensor assembly of claim 1, wherein the coating is chemically bonded to the skeletal frame. 12. The sensor assembly of claim 1, comprising a pulse oximeter monitor configured to receive data from the sensor assembly. 13. The sensor assembly of claim 1, comprising a multi-parameter monitor configured to receive data from the sensor assembly. 14. The sensor assembly of claim 1, comprising at least one integrated circuit device. 15. The sensor assembly of claim 1, comprising a cable comprising one or more integrated circuits. 16. The sensor assembly of claim 1, comprising an adapter cable connected to one or more signal transmission structures of the sensor assembly. 17. The sensor assembly of claim 1, wherein the loop structure biases an emitter and a detector toward one another. 18. A method of applying a sensor, comprising: applying force to a sensor body comprising a loop structure portion of an overmolded skeletal frame, wherein the applied force squeezes together opposing sides of the loop structure such that the sensor body is distorted; andsituating the sensor body on a digit of a patient while the force is applied to the sensor body. 19. The method of claim 18, comprising: ceasing to apply force to the sensor body once the sensor body is situated on the digit. | Nellcor Puritan Bennett LLC | A61B005/145
US07654966 | B2 | 2010-02-02 | Method and apparatus for monitoring dynamic cardiovascular function using n-dimensional representatives of critical functions | 2002-10-11 | A method, system, apparatus and device for the monitoring, diagnosis and evaluation of the state of a dynamic pulmonary system is disclosed. This method and system provides the processing means for receiving sensed and/or simulated data, converting such data into a displayable object format and displaying such objects in a manner such that the interrelationships between the respective variables can be correlated and identified by a user. This invention provides for the rapid cognitive grasp of the overall state of a pulmonary critical function with respect to a dynamic system. | 1. A medical display method for using a data processor to provide an integrated representation of a patient's pulmonary system for diagnostic purposes and display to a user, comprising: displaying an inspired gas object on a display, wherein the inspired gas object is configured to depict a patient's inspired gas volume as measured by a first sensor configured to detect inspired gas;displaying a lung object on the display, the lung object being associated with the inspired gas object and configured to represent a patients' lung condition as sensed by a second sensor configured to detect a lung condition;displaying an expired gas object on the display, the expired gas object being configured to represent a patient's expiration volume as measured by third sensor configured to detect expired gas, and the expiration object being displayed in combination with the lung object and inspired gas object to exhibit interaction between objects to display pulmonary states of the patient; anddisplaying a normal reference frame on the display, the normal reference frame representing expected normal pulmonary values for inspired gas, lungs and expired gas. 2. A medical diagnostic display method as in claim 1, further comprising: displaying an airway object representing a constriction state of a patient's airway;displaying a ventilation object in communication with the airway object representing tidal volume of the patient's pulmonary system. 3. A medical diagnostic display method as in claim 2, wherein a ventilation object is configured in a symbolic lung shape. 4. A medical diagnostic display method as in claim 1, wherein the lung object represents alveoli or a site of gas interchange. 5. A medical diagnostic display method as in claim 1, further comprising displaying a lung reference frame having lung reference frame thickness mapped inversely to lung compliance, wherein a thicker lung reference frame represents increased lung stiffness. 6. A medical diagnostic display method as in claim 5, wherein the lung object is mapped to Positive End Expiratory Pressure (PEEP). 7. A medical diagnostic display method as in claim 5, wherein the lung object can expand in width past the lung reference frame to represent high internal lung pressure or volume. 8. A medical diagnostic display method as in claim 1, wherein changes on an x-axis of the airway object represent Positive Inspiratory Pressure (PIP). 9. A medical diagnostic display method as in claim 1, wherein the inspired gas object is configured to represent fractional inspired oxygen concentration (FIO2). 10. A medical diagnostic display method as in claim 1, wherein the expired gas object is configured to represent end tidal carbon dioxide concentration (ETCO2). 11. A medical diagnostic display method as in claim 1, wherein the normal reference frame and object within the reference frame are configured to show positive increases past normalization by expanding outside the reference frame and negative dec... | University of Utah Research Foundation | A61B005/08
US08321003 | B2 | 2012-11-27 | Methods and apparatuses for detection of myocardial ischemia upon exertion | 2009-07-23 | Various method embodiments of the present invention concern sensing patient-internal pressure measurements indicative of physiological exertion, identifying one or more steady state periods of physiological exertion based on the patient-internal pressure measurements, sensing extra-cardiac response data and cardiac response data corresponding to the one or more physiological exertion steady state periods, respectively comparing the extra-cardiac response data and the cardiac response data to extra-cardiac response information and cardiac response information associated with equivalent levels of physiological exertion intensity of the one or more steady state periods, and determining the likelihood that myocardial ischemia occurred during the one or more steady state periods based on the comparison of the extra-cardiac response data to the extra-cardiac response information and the cardiac response data to the cardiac response information. | 1. A system for characterizing myocardial ischemia, comprising: an implantable static exertion sensor configured to output an exertion signal responsive to patient static exertion, the exertion signal containing static exertion data;at least one implantable physiological parameter sensor configured to output a physiological signal responsive to a physiological parameter, the physiological signal containing physiological exertion-response data, the at least one implantable physiological parameter sensor comprising: an implantable extra-cardiac parameter sensor configured to output an extra-cardiac signal responsive to an extra-cardiac parameter, the extra-cardiac signal containing extra-cardiac response data; andan implantable cardiac parameter sensor configured to output a cardiac signal responsive to a cardiac parameter, the cardiac signal containing cardiac response data;memory; anda processor configured to execute program instructions stored in memory to cause the system to: identify steady state periods of physiological exertion based on the static exertion data and assign one of a plurality of exertion intensity levels to each of the steady state periods;organize the physiological exertion-response data according to the assigned exertion intensity levels of the identified steady state periods of physiological exertion during which the physiological exertion-response data was respectively output;compare the physiological exertion-response data with stored physiological response information associated with equivalent levels of physiological exertion intensity of the steady state periods to which the physiological exertion-response data is organized; anddetermine the likelihood that myocardial ischemia occurred during one or more of the periods based on the comparison of the physiological response information to the physiological exertion-response data, wherein: processor organization of the physiological response data comprises organizing the respective extra-cardiac response and cardiac response data according to the assigned exertion intensity levels of the identified steady state periods of physiological exertion during which the extra-cardiac response and cardiac response data was respectively output;processor comparison of the physiological exertion-response data with stored physiological response information comprises respectively comparing the extra-cardiac and cardiac response data with extra-cardiac and cardiac response information associated with equivalent levels of physiological exertion intensity of the steady state periods to which the extra-cardiac and cardiac response data is organized; anddetermining the likelihood that myocardial ischemia occurred during one or more of the periods is based on the comparison of the extra-cardiac and cardiac response data to the extra-cardiac and cardiac response information. 2. The system of claim 1, wherein: the likelihood that myocardial ischemia occurred during a period of the identified s... | Cardiac Pacemakers, Inc. | A61B005/04
US08057479 | B2 | 2011-11-15 | Surgical orientation system and method | 2009-07-13 | A system and method for detecting and measuring changes in angular position with respect to a reference plane is useful in surgical procedures for orienting various instruments, prosthesis, and implants with respect to anatomical landmarks. One embodiment of the device uses dual orientation devices of a type capable of measuring angular position changes from a reference position. One such device provides information as to changes in position of an anatomical landmark relative to a reference position. The second device provides information as to changes in position of a surgical instrument and/or prosthesis relative to the reference position. | 1. A method of positioning a medical prosthesis with respect to an anatomical feature of a patient, the method comprising: providing a surgical instrument adapted to support the medical prosthesis or an alignment guide adapted to contact the anatomical feature, and identifying a first position of the surgical instrument relative to the anatomical feature;coupling a first measuring device to the surgical instrument in the first position and zeroing the first measuring device, the first measuring device being adapted to measure first positional changes;attaching the first measuring device to the patient near the anatomical feature;coupling a second measuring device to the surgical instrument in the first position and zeroing the second measuring device, the second measuring device being adapted to measure second positional changes; anddetermining a relative position of the second measuring device to the anatomical feature, based on the first positional changes and the second positional changes. 2. The method of claim 1 wherein the second measuring device includes a digital display to provide an indication of the second positional changes. 3. The method of claim 1 wherein the first measuring device is operatively coupled to the second measuring device by a communication channel. 4. The method of claim 3 wherein the communication channel is a wireless infrared or radio frequency telemetry link. 5. The method of claim 3 wherein the second measuring device receives the first positional changes from the first measuring device via the communication channel. 6. The method of claim 5 wherein the second measuring device includes a digital display to provide an indication of the position of the second measuring device relative to the anatomical feature. 7. The method of claim 1 wherein the anatomical feature is an acetabulum and the medical prosthesis is an acetabular prosthetic socket. 8. The method of claim 7 wherein the step of identifying a first position of the surgical instrument relative to the anatomical feature includes coupling the alignment guide to the surgical instrument and contacting a distal end of the alignment guide to a rim of the acetabulum. 9. The method of claim 8 further comprising positioning the surgical instrument at a predetermined angular orientation relative to the acetabulum. 10. The method of claim 9 wherein the predetermined angular orientation is a predetermined anteversion angle. 11. The method of claim 1 further comprising identifying a reference position of the anatomical feature using the surgical instrument, and wherein the first positional changes are positional changes of the anatomical feature from the reference position. 12. The method of claim 1 wherein the first and second measuring devices are three-dimensional electronic orientation devices and wherein each of the first and second measuring devices includes at least one rate sensor that provides a rate signal indicative of a change in angular position ... | OrthAlign, Inc. | A61B017/58
US08049497 | B2 | 2011-11-01 | MRI RF encoding using multiple transmit coils | 2007-10-31 | A common method of RF encoding assumes that the B1 field generated by the RF coils is linear, which is likely not the case in many situations. It is therefore desirable to have a method of operating an MR system to reconstruct an image of a subject, wherein the method is capable of also handling arbitrary B1 fields used for RF encoding. Accordingly, such an MR system employing one or more RF coils is disclosed herein. The method comprises obtaining transmit sensitivities and weighting factors for individual RF coils. Each RF coil is activated based on its respective weighting factor to apply RF excitation to a subject under examination in the MR system. MR signals—such as free induction decays (FID) signals or echo signals—generated from the subject in response to the RF excitation are received and processed based on the transmit sensitivities to generate an MR image or spectrum representative of the subject. | 1. A magnetic resonance method employing one or more radio-frequency coils, the method comprising: obtaining transmit sensitivities for individual radio-frequency coils;obtaining weighting factors for individual radio-frequency coils, wherein obtaining the weighting factor for each radio-frequency coil comprises: defining a B1 gradient value for the one or more radio-frequency coils; andcalculating the weighting factor for each radio-frequency coil based on a superposition of the respective B1 gradient value and transmit sensitivity;applying radio-frequency excitation to a subject by activating each radio-frequency coil based on its respective weighting factor;receiving magnetic resonance signals from the subject generated in response to the applied radio-frequency excitation; andprocessing the received magnetic resonance signals based on the transmit sensitivities, to generate a magnetic resonance image or a spectrum representative of the subject. 2. The method of claim 1, wherein obtaining the transmit sensitivities includes: measuring the transmit sensitivity for each radio-frequency coil for the radio-frequency excitation in the subject. 3. The method of claim 1, wherein activating each radio-frequency coil based on its respective weighting factor includes: setting an amplitude and/or phase of the excitation signal applied by the radio-frequency coil. 4. The method of claim 1, wherein receiving magnetic resonance signals from the subject generated in response to the applied radio-frequency excitation includes: receiving a complex data point from the subject for each radio-frequency excitation. 5. The method of claim 1, wherein processing the received magnetic resonance signals based on the transmit sensitivities includes: multiplying the received magnetic resonance signals by the inverse of a matrix of the transmit sensitivities. 6. The method of claim 1, wherein processing the received magnetic resonance signals based on the measured transmit sensitivities includes: backprojecting the received magnetic resonance signals. 7. The method of claim 1, wherein the defined B1 gradient value is non-linear. 8. The method of claim 1, wherein the resonance signals are spatially encoded by the B1 gradient. 9. A non-transitory computer readable medium carrying software which controls one or more processors of a magnetic resonance system to perform the method of claim 1. 10. A magnetic resonance system comprising: one or more radio frequency coils;an RF transmitter;an RF receiver; andone or more processors programmed to perform the method of claim 1. 11. A magnetic resonance system comprising: one or more radio-frequency coils;a mapping unit which obtains transmit sensitivities for individual radio-frequency coils;a weighting unit which obtains weighting factors for individual radio-frequency coils, the weighting unit including: a gradient control unit which defines a B1 gradient value for each radio-frequency coil, anda computing unit which cal... | Koninklijke Philips Electronics N.V. | G01V003/00
US08571662 | B2 | 2013-10-29 | Transvascular nerve stimulation apparatus and methods | 2008-01-29 | Electrode structures for transvascular nerve stimulation combine electrodes with an electrically-insulating backing layer. The backing layer increases the electrical impedance of electrical paths through blood in a lumen of a blood vessel and consequently increases the flow of electrical current through surrounding tissues. The electrode structures may be applied to stimulate nerves such as the phrenic, vagus, trigeminal, obturator or other nerves. | 1. A method for regulating breathing of a person, the method comprising: implanting at least one of: a first electrode structure at a position along a lumen of the left subclavian vein that is proximate to the left phrenic nerve; a second electrode structure at a position along a lumen of the superior vena cava that is proximate to the right phrenic nerve; a third electrode structure at a position along a lumen of the left interior jugular vein that is proximate to the left phrenic nerve; and a fourth electrode structure at a position along a lumen of the right interior jugular vein that is proximate to the right phrenic nerve;stimulating at least one of the left- and right-phrenic nerves by applying stimulation signals to electrodes of the first and second electrode structures. 2. A method for regulating breathing of a person, the method comprising: implanting a first electrode structure at a position along a lumen of the left subclavian vein that is proximate to the left phrenic nerve;implanting a second electrode structure at a position along a lumen of the superior vena cava that is proximate to the right phrenic nerve;stimulating the left- and right-phrenic nerves by applying stimulation signals to electrodes of the first and second electrode structures. 3. A method according to claim 2 comprising inserting the first and second electrode structures through a cannulation site in the person's left subclavian vein. 4. A method according to claim 3 comprising inserting the first electrode structure and subsequently introducing the second electrode structure through a bore of the first electrode structure. 5. A method according to claim 2, wherein the electrode structures each comprise an electrode supported on an electrically-insulating backing sheet; and, a structure for holding the backing sheet against the inner wall of a blood vessel with the electrode in contact with the inner wall of the blood vessel and the method comprises providing a balloon within the electrode structure and inflating the balloon to urge the backing sheet against an inner wall of the lumen of the blood vessel. | Simon Fraser University | A61N001/18
US09050049 | B2 | 2015-06-09 | Urethra gauge and methods of manufacture, and operation thereof | 2009-06-11 | A urethra gauge for measuring an inner size of a urinary tract. The gauge permits a clinician to measure the size of the urethral lumen (proximal to an outer orifice of the urethra) before a catheterization is performed. Selecting a properly-sized catheter for insertion into a patient by measuring the size of the urethral lumen before performing a catheterization is believed to reduce: (1) catheter-associated-urinary-tract infections; (2) urethral injuries due to inserting a catheter that is too large for the urethra; and/or (3) leakage associated with catheters that are too small for a particular patient. | 1. A system for determining a measured internal size of a urinary tract lumen proximate to a urethral oval outer opening of a patient, prior to urinary catheterization of said patient, the system comprising: a gauge for insertion into the urethra, constructed of a single rod-shaped measuring device having a distal-end, configured for insertion into a urethra and which forms the point of initial insertion into the outer opening of the urethra;a proximal-end of progressively increasing cross-sectional circumference relative to the distal end, until an inner circumference of the urethra is coextensive with an inserted portion of the gauge; anda solid or semisolid oval cross-section, wherein length of a cross-section of the gauge is greater than width at a widest point thereof; andwherein an array of indicia are located on an outer-peripheral surface of the gauge, such that each indicia corresponds to a potential measured internal circumference, area, diameter, perimeter or boundary of the urethral lumen; andwherein the gauge is contoured for insertion into the urethra only until resistance is first encountered without injury to the patient; anda stain for application to an outer surface of the urethra and a portion of the gauge coextensive with the stained outer surface of the urethra when the gauge has been inserted into the urethra only until resistance is first encountered without injury to the patient. 2. The in system of claim 1, wherein the the gauge has a conical longitudinal shape. 3. The system of claim 1, wherein the gauge is formed of a flexible material. 4. The system of claim 1, wherein the gauge is made from material selected from the group consisting of: rubber, polymer, silicone, glass, wood, plastic, metal, cardboard, polypropylene, polyurethane, and one or more combinations thereof. 5. The system of claim 1, wherein the gauge is sterile. 6. The system of claim 1, wherein the gauge is packaged in a sterile container. 7. A method of measuring the internal size of a urinary tract lumen proximate to a urethral oval outer opening of a patient using a gauge constructed of a single rod-shaped measuring device having a distal-end, configured for insertion into a urethra and which forms the point of initial insertion into the outer opening of the urethra;a proximal-end of progressively increasing cross-sectional circumference relative to the distal end, until an inner circumference of the urethra is coextensive with an inserted portion of the gauge; anda solid or semisolid oval cross-section, wherein length of a cross-section of the gauge is greater than width at a widest point thereof; andwherein an array of indicia are located on an outer-peripheral surface of the gauge, such that each indicia corresponds to a potential measured internal circumference, area, diameter, perimeter or boundary of the urethral lumen; andwherein the gauge is contoured for insertion into the urethra only until resistance is first encountered without injury... | - | A61B005/103 A61B005/00 A61B005/107
US08684920 | B2 | 2014-04-01 | Classification of severity of medical condition by wavelet based multi-resolution analysis | 2008-07-07 | A system and method are provided for classifying the severity of a medical condition detected by a medical device using wavelet based multi-resolution analysis of biomedical signals. Biomedical signals are converted into a set of wavelet coefficients including scales of wavelet coefficients at different resolutions ranging from a finest scale to a coarsest scale. Features of the biomedical signal useful in diagnosing a biomedical condition are identified by determining whether corresponding identifying features appear in at least one of the plurality of scales of wavelet coefficients. The identifying features determined to appear in the various scales of wavelet coefficients are conditionally linked together and used to classify a degree of severity of a diagnosed biomedical condition based on the number of identifying features appearing in different scales of wavelet coefficients. In some embodiments, alerts can be generated based on the diagnosed degree of severity of the biomedical condition. | 1. A method comprising: converting biomedical signal data into a set of wavelet coefficients, the set of wavelet coefficients including a plurality of scales of wavelet coefficients extending from wavelet coefficients at a fine scale to wavelet coefficients at a coarse scale;identifying features of the biomedical signal data based on whether corresponding identifying features appear in at least one of the plurality of scales of wavelet coefficients;conditionally linking identified features in each of the plurality of scales of wavelet coefficients to determine a degree of severity of a biomedical condition based on the identified features of the biomedical signal data. 2. The method of claim 1, further comprising: determining whether an identified feature appears in wavelet coefficients at a first selected scale;selecting the identified featured appearing in wavelet coefficients at the first selected scale as a base point for the conditional linking. 3. The method of claim 2, further comprising determining whether an identified feature appears in an adjacent scale of wavelet coefficients from the base point scale within a first time period window surrounding the base point. 4. The method of claim 3, further comprising: determining whether the identified feature appearing in the adjacent scale of wavelet coefficients meets certain threshold criteria; andif the certain threshold criteria are met, identifying the identified feature as a linking point and conditionally linking the identified feature appearing in the adjacent scale of wavelet coefficients to the base point to represent a chain of conditional linking points. 5. The method of claim 4, further comprising repeatedly performing the following operations until a chain of conditional linking points has been completed without break between consecutive scales of wavelet coefficients: determining whether an identified feature appears in a next adjacent scale of wavelet coefficients from a linking point that was last added to the chain of conditional linking points in a time period window surrounding the last added linking point;if an identified feature is located in the next adjacent scale of wavelet coefficients, determining whether the identified feature appearing in the next adjacent scale of wavelet coefficients meets certain threshold criteria, anddetermining whether the identified feature appearing in the next adjacent scale of wavelet coefficients is located within an overall time period window for the entire chain of conditional linking points; andif the certain threshold criteria and overall time period window conditions are satisfied, identifying the identified feature as a linking point and conditionally linking the identified feature appearing in the next adjacent scale of wavelet coefficients to the chain of conditional linking points having identified features appearing in the other scales of wavelet coefficients. 6. The method of claim 5, wherein the chain of conditional link... | Medtronic, Inc. | A61B005/00
US08378831 | B2 | 2013-02-19 | System and method for coordinated health monitoring, emergency response, and medical record delivery | 2009-08-04 | A computer-based system for providing coordinated health monitoring, emergency response, and medical record delivery. The system can include computing devices configured to process emergency-related indicators and data. The system can also include monitoring devices linked to the computing devices. The monitoring devices can monitor a particular area for the emergency-related indicators and data, where the monitoring devices detect among speech, sounds, images and other detectable emergency-related indicators. The monitoring devices can also transmit the emergency-related indicators and data to the computing devices. Furthermore, the system can include a module linked to the monitoring devices and configured to execute on the computing devices. The module can analyze the transmitted emergency-related indicators and data to determine whether there is an emergency, communicate with a monitoring service to validate that an emergency exists, and provide access to patient records to authorized personnel, based upon whether an emergency exists. | 1. A computer-based method for providing coordinated health monitoring, emergency response, and medical record delivery, the method comprising the steps of: actively monitoring a particular area associated with a patient for emergency-related indicators and data by utilizing at least one monitoring device to detect the emergency-related indicators and data, the emergency related indicators and data comprising at least one of speech, sound, or images associated with the particular area;transmitting the emergency-related indicators and data to at least one computing device, wherein the at least one computing device is communicatively linked to the at least one monitoring device;analyzing the emergency-related indicators and data to determine whether there is a medical emergency regarding the patient, wherein the analysis is performed by the at least one computing device;communicating with a monitoring service to validate the emergency; andproviding access to patient records for the patient to authorized personnel if the medical emergency is validated. 2. The method of claim 1, wherein the at least one monitoring device comprises at least one among a microphone, a speaker, a smoke detector, a heat detector, a device enabling encryption of the emergency-related indicators and data, a camera, a video camera, an intercom, a baby monitor, and a motion sensor. 3. The method of claim 1, wherein the at least one monitoring device and at least one computing device are adapted to be portable so as to be carried by an individual and communicatively and securely link to at least one of monitoring services, offices, WiFi enabled facilities, and mobile services. 4. The method of claim 1, wherein the monitoring step further comprises generating additional indicators and data for the emergency-related indicators and data by utilizing at least one health-monitoring sensor associated with the patient to detect the additional indicators and data, the additional indicators and data comprising at least one among blood pressure, temperature, heart-rate, and other health-related patient metrics. 5. The method of claim 1, further comprising communicatively linking to emergency services and signaling an alarm to indicate that an emergency is occurring. 6. The method of claim 1, further comprising storing and forwarding a record of the emergency-related indicators and data to at least one of the monitoring service and an emergency service. 7. The method of claim 1, further comprising integrating drug dispensing devices, wherein the drug dispensing devices administer drugs so as to enable remote medical maintenance. 8. A computer-based system for providing coordinated health monitoring, emergency response, and medical record delivery, the system comprising: at least one computing device configured to process emergency-related indicators and data;at least one monitoring device communicatively linked to the at least one computing device, wherein the at least one monitor... | The Quantum Group, Inc. Noel J. Guillama | G08B023/00
US08357105 | B2 | 2013-01-22 | Anti-reflux mechanism for urine collection systems | 2008-08-07 | A urine collection system includes an anti-flux mechanism to reduce the amount of urine that may flow from within the collection bag back into the urine meter. The urine collection system includes a urine meter, a collection member defining at least a first opening for receiving urine from the urine meter and a shield positioned adjacent the opening of the collection member. The shield is affixed to the collection member at predetermined connection points along a periphery of the shield whereby urine flows between the connection points from the urine meter to the collection member. The shield may be affixed to an internal surface of the bag. The shield may be spot welded to the collection member at the connection points. The shield may include at least a first slot configured to permit the passage of fluid therethrough. | 1. A urine collection system, which comprises: a urine meter having a volume to accommodate urine therein;a collection member connected to the urine meter and defining at least a first opening for receiving urine from the urine meter; anda shield covering the at least first opening of the collection member and being permanently affixed to the collection member at predetermined connection points along a periphery of the shield such that urine must flow between the connection points to pass from the urine meter to the collection member. 2. The urine collection system according to claim 1 wherein the shield is affixed to an internal surface of the collection member. 3. The urine collection system according to claim 1 wherein the shield is spot welded to the collection member at the connection points. 4. The urine collection system according to claim 1, further including a flap disposed between the shield and the collection member. 5. The urine collection system according to claim 4, wherein the flap is configured to deflect towards the collection member as urine flows from the urine meter to the collection member. 6. The urine collection system according to claim 1, wherein the urine meter is configured to be tilted to effect urine drainage from the urine meter to the collection member. 7. A urine collection system, which comprises: a urine meter having a volume to accommodate urine therein, the urine meter defining a dump port;a collection member connected to the urine meter and defining an opening for receiving urine from the urine meter; anda guard covering the dump port of the urine meter, the guard having a plurality of passages formed through the guard which communicate the dump port with the opening in the collection member to allow passage of urine from the urine meter into the collection member, wherein the plurality of passages control the flow of fluid into the collection member and control the reflux of the fluid from the collection member into the urine meter. 8. The urine collection system according to claim 7 wherein at least some of the passages of the guard define different internal dimensions. 9. The urine collection system according to claim 8 wherein the internal dimensions of the passages of the guard increase from a lower end of the guard to an upper end of the guard. 10. The urine collection system according to claim 8 wherein the plurality of passages of the guard define a plurality of slots. | Covidien LP | A61B005/00
US09968784 | B2 | 2018-05-15 | Posture state redefinition based on posture data | 2009-04-30 | The disclosure is directed towards posture-responsive therapy. To avoid interruptions in effective therapy, an implantable medical device may include a posture state module that detects the posture state of the patient and automatically adjusts therapy parameter values according to the detected posture state. A system may include a memory that stores posture state definitions, a posture state module that records a plurality of postures of a patient over a period of time, and a processor that identifies a set of the plurality of postures that fall within a posture state, and redefines a boundary of the posture state based on where the postures fall within the posture state. | 1. A system comprising: a memory configured to store posture state definitions;a posture state module configured to sense a plurality of postures of a patient over a period of time; anda processor configured to: identify, based on one of the posture state definitions, a set of two or more of the plurality of postures that fall within a posture state;identify a portion of the posture state where a specified amount of the postures fall; andautomatically update the one of the posture state definitions to redefine at least one boundary of the posture state to change a size of the posture state and to correspond to the identified portion based on where the postures in the set of postures fall within the posture state and the identified portion. 2. The system of claim 1, wherein the specified amount comprises at least a specified percentage of postures of the set of the plurality of postures. 3. The system of claim 1, wherein the processor is configured to update the one of the posture state definitions to redefine the at least one boundary to change the size of the posture state by at least modifying a tolerance from a posture reference vector at a center of the posture state. 4. The system of claim 1, wherein the posture state module is configured to periodically sample the posture of the patient. 5. The system of claim 1, further comprising a medical device, wherein the medical device comprises at least one of the memory, posture state module, and processor. 6. The system of claim 5, wherein the medical device comprises an implantable medical device. 7. The system of claim 1, further comprising an external programmer, wherein the external programmer comprises at least one of the memory, posture state module, and processor. 8. The system of claim 7, wherein the external programmer comprises at least one of a patient programmer or a clinician programmer. 9. The system of claim 1, wherein the processor is configured to update the one of the posture state definitions to change the size of the posture state by shrinking the posture state. 10. The system of claim 1, wherein the processor is configured to update the one of the posture state definitions to change the size of the posture state by expanding the posture state. 11. The system of claim 2, wherein the specified percentage of postures comprises at least 80 percent of postures of the set of the plurality of postures. 12. The system of claim 3, wherein the tolerance comprises one of an angle or a cosine value. 13. A system comprising: means for sensing a plurality of postures of a patient over a period of time;means for identifying, based on one of a plurality of posture state definitions, a set of two or more of the plurality of postures that fall within a posture state;means for identifying a portion of the posture state where a specified amount of the postures fall; andmeans for automatically updating the one of the posture state definitions to redefine at least one boundary of the po... | Medtronic, Inc. | A61B005/103 A61B005/117 A61N001/36 A61B005/11 G06F019/00 A61N001/365 A61B005/00
US10080499 | B2 | 2018-09-25 | Implantable medical system including multiple sensing modules | 2008-07-30 | A medical system includes at least two sensing modules that each generate an optical signal that changes as a function of a physiological parameter of a patient. The sensing modules may be coupled to a common light source and a common receiver via an optically transmissive member. At least a first sensing module that is closest to the light source along a length of the optically transmissive member may include a waveguide to split the light emitted by the light source. A first portion of the light may be directed toward the first sensing module and a second portion of the light may be directed toward a second sensing module that is placed downstream of the first sensing module in a direction substantially along the direction of light flow through the optically transmissive member and away from the light source. | 1. An implantable medical system comprising: a light source configured to emit light;an implantable optically transmissive member comprising a waveguide comprising an optical cavity, a first segment and a second segment, the first segment optically coupled to the light source and configured to propagate the light emitted by the light source to a distal end of the first segment, the distal end of the first segment received at a first end of the waveguide so that the light emitted by the light source and propagated from the light source through the first segment toward the distal end of the first segment exits the distal end of the first segment and is received at the optical cavity, the second segment having a proximal end received at a second side of the waveguide, the first side of the waveguide separated from the second side of the waveguide by the optical cavity, wherein the first segment is physically separated from the second segment by the optical cavity and is optically coupled to the second segment through at least one optical path formed through the optical cavity:a light dividing member positioned within the optical cavity and configured to divide the light from the light source received at the optical cavity through the first segment into at least a first portion and a second portion, wherein the light dividing member comprises at least one surface within the optical cavity that occludes a first portion of the optical cavity and is configured to direct the first portion of the light, and wherein a second portion of the optical cavity is unobstructed by the light dividing member, the second portion of the optical cavity comprising the at least one optical path between the first segment and the second segment through the optical cavity for the second portion of light:a first sensing module optically coupled to the first and second segments of the optically transmissive member, the first sensing module configured to receive the first portion of light and generate a first optical signal; anda second sensing module optically coupled to the second segment of the optically transmissive member, the second sensing module configured to receive only the second portion of the light and generate a second optical signal,wherein the first sensing module and the second sensing module are configured to modulate one or more properties of the first and second portions of light as a function of one or more physiological parameters of a patient. 2. The implantable medical system of claim 1, wherein the light dividing member comprises a first light dividing member, the system further comprising a second light dividing member that divides the second portion of the light into at least a third portion and at least a fourth portion, wherein the second sensing module is configured to receive the third portion of the light. 3. The implantable medical system of claim 2, further comprising a third sensing module optically coupled to the optically transmissive mem... | Medtronic, Inc. | A61B005/1459 A61B005/01 A61B005/0215 A61B005/11 A61B005/00 A61B005/145 A61B005/0205 A61N001/365
US09585603 | B2 | 2017-03-07 | CPR assist device for measuring compression parameters during cardiopulmonary resuscitation | 2009-01-16 | A device for the determination of at least one compression parameter during administration of cardiopulmonary resuscitation (CPR) on a patient comprising: a field generator, a field detector, and a processor. Position information and the compression parameter are determined from the field detected by the field detector. One of the field generator and the field detector is a position sensor and the other is a reference sensor. | 1. A device for determination of at least one compression parameter during administration of cardiopulmonary resuscitation (CPR) on a patient comprising: a position sensor including a field generator adapted to generate a field;a reference sensor including a field detector adapted to detect the field generated by the field generator and generate a response signal;wherein the field generator comprises at least one field generating coil and the field detector comprises at least one field detecting coil, and the generated field is an electromagnetic field,wherein at least one of the field generator and the field detector comprises at least three orthogonally wound field generating coils or field detecting coils, respectively,a defibrillator unit having two defibrillator pads;a processor adapted to determine, from the response signal, position and/or orientation information for the field detector relative to the field generator, to determine the at least one compression parameter using the determined position and/or orientation information of the field detector, the at least one compression parameter including a compression angle, and to determine whether a rescuer is administering CPR at an effective or desired compression angle to the patient's chest, the processor being provided in the defibrillator unit;wherein the position sensor is attached to or embedded in one of the two defibrillator pads which is configured to be placed on the patient's chest, and the reference sensor is attached to or embedded in the other of two defibrillator pads configured to be placed at a different location on the patient's chest;wherein the position sensor is adapted to move in accordance with the patient's chest and the reference sensor is adapted to be stationary relative to the patient, andwherein the position sensor is configured to determine a position of the one of the two defibrillator pads having the position sensor relative to the other of the two defibrillator pads having the reference sensor in order to determine if the rescuer is administering CPR at the effective or desired angle to the patient's chest and/or instruct the rescuer in maintaining the effective or desired compression angle. 2. The device of claim 1 wherein at least one of the field generator and the field detector comprises more than three coils. 3. The device of claim 1 wherein the field generator comprises at least two field generating coils, and each field generating coil sequentially generates a field detectable by the at least one field detecting coil. 4. The device of claim 1 wherein the field generator comprises at least two field generating coils and at least two field detecting coils, and each field generating coil generates a field at a different frequency, the field detector having at least one bandpass filter for separating the fields at different frequencies. 5. The device of claim 1 wherein the device comprises a feedback component configured to provide feedback to a CPR a... | Physio-Control Canada Sales Ltd. | A61B005/08 A61B005/11 A61N001/362 A61H031/00 A61N001/39 G01B007/004 G09B023/28 A61B005/113 A61N001/04 A61B005/00
US20100016703 | A1 | 2010-01-21 | BIO-ELECTRODE POSSESSING A HYDROPHILIC SKIN-CONTACTING LAYER AND AN ELECTROLYTE SUBSTANCE | 2007-07-13 | A bio-electrode for conveying electrical signals to, or from a body is constructed with two components, a pre-laminated member and a re-usable electrode assembly. The pre-laminated member comprises a first substrate layer and a second electrolyte-containing layer. The substrate contacts the body and is constructed of a material that is biocompatible, hydrophilic and inherently electrically semi-conductive or conductive. The electrolyte-containing layer, which does not contact the body, is composed of an adhesive, electrolytic hydrogel. Small amounts of moisture and electrolytes from the hydrogel diffuse into the substrate thereby reducing and stabilizing the substrate's electrical resistance. The pre-laminated member can be used with existing re-usable electrode assemblies. | 1. An electrode for detecting physiologic signals from or delivering electrical signals to a body, comprising: a) a body contacting substrate layer;b) an electrolyte-containing layer; andc) a conducting member, said electrolyte-containing layer being in contact with said body contacting substrate layer and said conducting member;wherein said conducting member is adapted to deliver said signals between said body and an external source, through said electrolyte-containing layer and said body contacting substrate layer. 2. The electrode of claim 1 wherein said body contacting substrate layer is comprised of a material that is at least partially conductive. 3. The electrode of claim 1 wherein said body contacting substrate layer has the capacity to absorb electrolyte from said electrolyte-containing layer. 4. The electrode of claim 1 wherein said electrolyte-containing layer overlies at least a portion of said body contacting substrate layer. 5. The electrode of claim 1 wherein said body contacting substrate layer comprises a non-metallic, semi-conducting polymer. 6. The electrode of claim 5 wherein said body contacting substrate layer comprises an inherently dissipative polymer. 7. The electrode of claim 6 wherein said body contacting substrate layer further comprises a thermoplastic polyolefin elastomer. 8. The electrode of claim 1 wherein said electrolyte-containing layer is a biocompatible, aqueous hydrogel. 9. The electrode of claim 1 wherein said body contacting substrate layer is hydrophilic. 10. The electrode of claim 1 wherein said conducting member comprises a conducting plate and active electronic mean and wherein said electrolyte-containing layer is in contact with said conducting plate. 11. An electrode for detecting physiologic signals from or transmitting electrical signals to a body, comprising: a) a pre-laminated member comprising a body contacting layer and an electrolyte-containing layer; andb) a re-usable electrode assembly in contact with said electrolyte-containing layer;wherein said body contacting layer comprises a hydrophilic material that is at least partially conductive and which is capable of absorbing electrolyte from said electrolyte-containing layer, and wherein the electrolyte containing layer provides an electrical connection between said body contacting layer and said re-usable electrode assembly. 12. The electrode of claim 11 wherein said pre-laminated member is disposable. 13. An electrode for detecting physiologic signals from or delivering electrical signals to a body, comprising: a) a body contacting substrate layer having a first surface and a second surface; said second surface being in contact with said body; andb) a conducting member adapted to deliver said signals between said body and an external source; said conducting member being in contact with said first surface;wherein a portion of said first surface and said conducting member is in contact with an electrolyte-containing substance. 14. A... | CARDIAC BIO-SYSTEMS INC. | A61B005/04
US20100030092 | A1 | 2010-02-04 | Adaptive Hypoglycaemia Alert System and Method | 2007-11-14 | The onset or presence of a hypoglycaemic condition in a living being is estimated by obtaining ECG signals, which are processed together with secondary input, such as skin impedance, information about a condition of the living being, or blood glucose level, to estimate whether the living being is experiencing or approaching a hypoglycaemic condition. An output alert may be generated. Feedback indicative of a possible misestimate of the onset or presence of the hypoglycaemic condition is received to allow a system incorporating the invention to adapt and personalize to a specific living being. The feedback may influence one or more sensitivity or specificity settings. Settings of the system may be influenced by the location or rate of movement of the living being. | 1. A system for estimating the onset or presence of a hypoglycaemic condition in a living being, the system comprising a memory, a processor, and: at least one first sensor element for obtaining ECG signals representative of the living being's ECG;at least one secondary input element for receiving at least one secondary input parameter;an output generator for generating an alert signal;the processor being programmed to: process the ECG signals, the at least one secondary input parameter and at least one reference value stored in the a memory to estimate the onset or presence of the hypoglycaemic condition;cause the output generator to generate the alert signal in case the hypoglycaemic condition is occurring or approaching;receive a feedback input indicative of a possible misestimate of the onset or presence of the hypoglycaemic condition, adjust said reference value in response to the feedback input, and store the adjusted reference value in the memory. 2. A system according claim 1, wherein: the at least one reference value comprises at least one sensitivity setting;the feedback input comprises at least one sensitivity-influencing input;the processor is programmed to vary the sensitivity setting under influence of the sensitivity-influencing input. 3. A system according to claim 2, wherein the sensitivity-influencing input comprises at least one of: quantity of a past medication of the living being;time lapse since a past medication of the living being;the living being's blood glucose level;input describing the living being's past meal and/or drink intake;time;dose of last medication of the living being;a physical location of the living being;an input, upon which the level of the living being's physical activity can be determined;a manually entered input; anda user-specified sensitivity setting. 4. A system according to claim 3, wherein the sensitivity-influencing input comprises the physical location of the living being, and wherein the processor is programmed to vary the sensitivity setting in case said physical location is a vehicle. 5. A system according to claim 3, wherein the sensitivity-influencing input comprises the physical location of the living being, the system further comprising a location detector for detecting said physical location. 6. A system according to claim 5, wherein at least the processor and the output generator are incorporated in a portable device, and wherein the location detector is adapted to be secured to a specific physical location, the system further comprising a location interface causing the system to change the sensitivity when the location detector is within communication reach of the portable device. 7. A system according to claim 3, wherein the sensitivity-influencing input comprises the input, upon which the level of the living being's physical activity can be determined, and wherein the processor is programmed to set change the sensitivity in response to said input. 8. A system according to clai... | Novo Nordisk A/S | A61B005/0402
US20100030067 | A1 | 2010-02-04 | SYSTEM, METHOD, COMPUTER-READABLE MEDIUM AND USE FOR IMAGING OF TISSUE IN AN ANATOMICAL STRUCTURE | 2007-11-14 | A system for imaging of prostate cancer in a prostate in vivo is provided. The system utilizes Diffuse Optical Tomography (DOT) for creating a 3D image for the detection of suspicious prostate tissue. The DOT image may be used to guide the biopsy, thereby reducing the number of false negatives. A method, computer-readable medium and use are also provided. | 1. A system (10) for imaging of prostate cancer in a prostate in vivo, said system comprising: at least three units selected from: an electromagnetic radiation source (11), and a detector unit (12), forming a plurality of electromagnetic radiation paths,wherein said electromagnetic radiation source is configured to emit incident electromagnetic radiation on said prostate, and said detector unit is configured to receive said electromagnetic radiation, wherein said electromagnetic radiation has been scattered multiple times in said prostate, said system further comprising:an image reconstruction unit (13) for reconstructing a Diffuse Optical Tomography image dataset of said prostate based on said received scattered electromagnetic radiation by said at least one detector unit; anda discrimination unit (14) for discriminating between healthy and diseased tissue based on information in said image dataset. 2. The system according to claim 1, wherein said at least one electromagnetic radiation source and said at least one detector unit is located on either side of the cancer to be imaged. 3. The system according to claim 1, wherein at least one electromagnetic radiation source is comprised in a urethral unit, said urethral unit being suitable for insertion through urethra and in use positioned in the vicinity of the prostate gland. 4. The system according to claim 1, wherein at least one detector unit is comprised in a transrectal unit, said transrectal unit being suitable for rectal insertion through rectum and in use positioned in the vicinity of the prostate gland. 5. The system according to claim 1, wherein said image dataset is a 2D, 3D, or multi-dimensional image dataset. 6. The system according to claim 1 wherein the distance between each electromagnetic radiation source and each detector unit is 2 mm to 10 cm. 7. The system according to claim 1, further comprising an ultrasound unit for providing an ultrasound image dataset of said prostate. 8. The system according to claim 7, wherein said ultrasound unit is integrated in said transrectal unit, and in use configured to provide an ultrasound image dataset of said prostate. 9. The system according to claim 7, wherein said ultrasound image dataset is used to guide a biopsy of said tissue utilizing the information of said Diffuse Optical Tomography image dataset. 10. A method for imaging of prostate cancer in a prostate in vivo, said method comprising emitting incident electromagnetic radiation on said prostate,receiving said electromagnetic radiation, wherein said electromagnetic radiation has been scattered multiple times in said prostate, wherein said emitting incident electromagnetic radiation and said receiving said electromagnetic radiation forming plurality of electromagnetic radiation paths, said method further comprisingreconstructing a Diffuse Optical Tomography image dataset of said prostate based on said received scattered electromagnetic radiation, anddiscriminating between hea... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/05 A61B008/14
US07660625 | B2 | 2010-02-09 | Catheter with compactly terminated electronic component | 2005-05-12 | A cable assembly includes an elongated catheter defining a lumen. A cable within the lumen has at least first and second electrically independent conductors. A first electrical component is connected to the catheter and has a first lead and second lead. The first lead is connected to the first conductor at a first connection, and the second lead connected to the second conductor at a second connection. The first and second connections are spaced apart along the length of the catheter, and the component is positioned at an intermediate location between the first and second connections. The component may be a coil, with leads extending in opposite directions, and only one of the conductors may pass through the coil. The assembly may include multiple components spaced apart along the length of the catheter. | 1. A cable assembly comprising: an elongated catheter having an exterior and defining a lumen;a cable within the lumen and having at least first and second electrically independent conductors;a first electrical component connected to the catheter and having a first lead extending in a first direction toward a first end of the catheter and a second lead extending in a second direction, different from the first direction, toward a second end, different from the first end, of the catheter, the first component being a coil surrounding the exterior of the entire catheter;the first lead connected to the first conductor at a first connection;the second lead connected to the second conductor at a second connection;the first and second connections being spaced apart along the length of the catheter; andthe first component being positioned at an intermediate location between the first and second connections. 2. The cable assembly of claim 1 wherein the first conductor is a central conductor of a coaxial cable, and the second conductor is a shield surrounding the first conductor. 3. The cable assembly of claim 2 wherein the cable is stripped to provide: a first section comprising an exposed free end portion of the central conductor,an adjacent second section comprising the central conductor surrounded by dielectric,an adjacent third section comprising dielectric surrounded by exposed shield, andan adjacent fourth section comprising shield surrounded by an insulating sheath. 4. The cable assembly of claim 3 wherein the first component is aligned with the second section. 5. The cable assembly of claim 1 including a second cable connected to a second component and connected to the catheter, the second component being located at a position spaced apart from the first component along the length of the catheter. 6. The cable assembly of claim 5 wherein the catheter defines multiple lumens, and wherein the first and second cables are received in different lumens. 7. A cable assembly comprising: an elongated catheter having an exterior, defining a lumen and having a free end and an opposed connection end;a first electrical component connected to the catheter proximate the free end and having a first lead and second lead, the first component being a coil surrounding the exterior of the entire catheter;a first cable within the lumen, having at least first and second electrically independent conductors, and extending from the connection end of the catheter to the component;the first lead of the first component extending toward the free end of the catheter; andthe second lead of the first component extending toward the connection end of the catheter. 8. The cable assembly of claim 7 wherein the first lead is connected to the first conductor at a first connection, the second lead is connected to the second conductor at a second connection, the first and second connections are spaced apart along the length of the catheter, and the component is positioned at an int... | Tyco Electronics Corporation | A61B005/05
US20100010356 | A1 | 2010-01-14 | BLOOD PRESSURE MONITOR WITH REMOTE DISPLAY | 2008-07-09 | A blood pressure monitor having a wireless transmitter is disclosed. A pressure sensor measures blood pressure data from a target, and the blood pressure data is transmitted to a remote device through the wireless transmitter. In one embodiment, the remote device is a mobile phone or a PDA. An alternate blood pressure monitor having an electronic display unit displaying the instantaneous pressure inside a cuff is also disclosed. In an embodiment, the blood pressure monitor computes the blood pressure value of the target, and the user can choose the measuring mode of the blood pressure monitor. | 1. A blood pressure monitor comprising: a) a cuff;b) at least one pump disposed on a surface of said cuff;c) at least one valve disposed on a surface of said cuff;d) a pressure sensor fitted inside said cuff to measure pressure inside said cuff;e) a microcontroller unit connected to said at least one pump, said at least one valve and said pressure sensor;f) a wireless transmitter connected to said microcontroller unit; andg) a remote device physically separate from said cuff; said remote device comprising a wireless receiver;wherein said wireless transmitter transmits blood pressure data measured from said pressure sensor to said remote device. 2. The blood pressure monitor according to claim 1, wherein said remote device further comprises a display unit. 3. The blood pressure monitor according to claim 1, wherein said remote device is a mobile phone. 4. The blood pressure monitor according to claim 1, wherein said remote device is a personal digital assistant. 5. The blood pressure monitor according to claim 1, further comprising a display unit attached to said cuff. 6. The blood pressure monitor according to claim 1, wherein said remote device is configured to forward said blood pressure data to other remote devices. 7. The blood pressure monitor according to claim 1, wherein said remote device comprises a microprocessor to process said blood pressure data. 8. A blood pressure monitor comprising: a) a cuff;b) at least one pump disposed on a surface of said cuff;c) at least one valve disposed on a surface of said cuff;d) a pressure sensor fitted inside said cuff to measure pressure inside said cuff;e) a microcontroller unit connected to said at least one pump, said at least one valve and said pressure sensor;wherein an instantaneous pressure value is displayed on said electronic display unit such that a user can use a stethoscope for measuring blood pressure of a target by referring to said displayed pressure value. 9. The blood pressure monitor according to claim 8, wherein said blood pressure monitor does not have a systolic and diastolic sensor. 10. The blood pressure monitor according to claim 8, wherein said microcontroller unit computes blood pressure values of said target from said blood pressure data. 11. The blood pressure monitor according to claim 8, wherein said displayed pressure value is an instantaneous pressure value inside said cuff. 12. The blood pressure monitor according to claim 8, wherein said displayed pressure value is calculated by said microcontroller unit from at least one parameter. 13. The blood pressure monitor according to claim 12, wherein said at least one parameter is a high level, a low level and a time for said cuff to deflate from said high level to said low level. 14. The blood pressure monitor according to claim 8, further comprising a recording mechanism for said user to record said displayed value when said recording mechanism is activated. 15. The blood pressure monitor according to claim 1... | - | A61B005/0225
US20100010366 | A1 | 2010-01-14 | METHOD TO EVALUATE PSYCHOLOGICAL RESPONSES TO VISUAL OBJECTS | 2006-12-22 | A method of evaluating the response of a subject to visual features of a visual display, the method including the steps of: (a) presenting a visual display having particular visual features to the subject during a first period; (b) determining brain activity of the subject during the first period; (c) presenting reference display material to a subject during a second period; (d) determining reference brain activity of the subject during the second period; (e) tracking the gaze position of at least one of the eyes of the subject on the visual display during the first period; and (f) evaluating the response of the subject to particular visual features of the visual display by determining differences in brain activity determined between steps (b) and (d) when the gaze of the subject is directed at the particular features. | 1. A method of evaluating the response of a subject to visual features of a visual display, the method including the steps of: (a) presenting a visual display having particular visual features to the subject during a first period;(b) determining brain activity of the subject during the first period;(c) presenting reference display material to a subject during a second period;(d) determining reference brain activity of the subject during the second period;(e) tracking the gaze position of at least one of the eyes of the subject on the visual display during the first period; and(f) evaluating the response of the subject to particular visual features of the visual display by determining differences in brain activity determined between steps (b) and (d) when the gaze of the subject is directed at the particular features. 2. A method as claimed in claim 1 wherein the visual display is printed advertising material, text layout, product design, packaging website, the interior or exterior of a building. 3. A method as claimed in claim 2 wherein the visual display is displayed on a video screen. 4. A method as claimed in claim 3 including the step of selecting the visual features of the visual display and determining the areas where the selected visual features are located on the video screen and wherein step (e) determines when the gaze position of the subject falls upon respective area of the selected visual features on the video screen. 5. A method as claimed in claim 4 wherein the differences in brain activity determined in step (f) are averaged for each selected visual feature. 6. A method as claimed in claim 1 wherein steps (a) to (e) are presented to a plurality of subjects and step (f) includes the steps of averaging the differences in brain activities of the subjects. 7. A method as claimed in claim 1 wherein steps (b) and (d) are carried out by determining gamma or high frequency EEG or MEG activity. 8. A method as claimed in claim 1 wherein steps (b) and (d) are carried out by detecting EEG or MEG activity in the frequency range 8 to 13 Hz. 9. A method as claimed in claim 1 wherein steps (b) and (d) are carried out by assessment of the phase of steady state visually evoked potentials (SSVEP) in EEG signals obtained from the subject or subjects or by assessment of steady state visually evoked responses (SSVER) in MEG signals obtained from the subject or subjects. 10. A method as claimed in claim 1 wherein steps (a) and (c) include the steps of placing electrodes at scalp sites to obtain output EEG signals which enable assessment of: visual attention to detail of the visual features;emotional intensity associated with the visual features;long term memory encoding associated with the visual features;engagement with the visual features;attraction associated with the visual features;desirability associated with the visual features; and/orlikeability associated with the visual features. 11. A method as claimed in claim 10 including the step... | - | A61B005/0484 A61B005/00
US20100049088 | A1 | 2010-02-25 | Specimen container | 2008-08-25 | A specimen container includes an opening for receiving biological fluid specimens, and a body, wherein at least one of a body mid-section and a container base is substantially wider than the opening. The body mid-section may be curved and wider than the container base. Alternatively, the body mid-section may be flat and narrower than the container base. A detachable lid may be configured to cover the opening. In another embodiment, a specimen container includes a base member, a set of wall members fixedly connected to the base member, and a deflector cap configured to deflect fluid specimen material back into the specimen container. Advantages of the present invention include a specimen container that allows more complete and easier specimen collection and retention. These advantages may be achieved, at least in part, through better shape optimization of the specimen container. | 1. A specimen container, comprising: an opening for receiving biological fluid specimens; anda body, wherein at least one of a body mid-section and a container base is substantially wider than the opening. 2. The specimen container of claim 1, wherein the body mid-section is curved and wider than the container base. 3. The specimen container of claim 2, wherein the body mid-section is roughly spherical. 4. The specimen container of claim 1, wherein the body mid-section is flat and narrower than the container base. 5. The specimen container of claim 4, wherein the overall shape of the body is roughly conical. 6. The specimen container of claim 1, further comprising a detachable lid configured to cover the opening. 7. The specimen container of claim 1, wherein the specimen container is made of at least one of the group consisting of: a plastic;a polymer;a synthetic rubber;a composite;a metal; anda ceramic. 8. A specimen container, comprising: an opening for receiving biological fluid specimens; anda body, wherein at least a portion of a body mid-section is situated substantially below the opening and a container base is situated substantially below the body mid-section, when the specimen container is in a substantially upright orientation. 9. The specimen container of claim 8, wherein the overall shape of the body is roughly a tear-drop shape. 10. The specimen container of claim 8, further comprising a detachable lid configured to cover the opening. 11. The specimen container of claim 8, wherein the specimen container is made of at least one of the group consisting of: a plastic;a polymer;a synthetic rubber;a composite;a metal; anda ceramic. 12. A specimen container, comprising: a base member;a set of wall members fixedly connected to the base member; anda deflector cap, hingedly connected to a top of at least one wall member, the deflector cap configured to deflect fluid specimen material back into the specimen container. 13. The specimen container of claim 12, wherein the shape of the base member is selected from the group consisting of: a square;a rectangular;a triangle;a curved-end rectangle; anda hemisphere. 14. The specimen container of claim 12, wherein the deflector cap is hingedly connected to the top of the at least one wall member by at least one flexible connection to the at least one wall member. 15. The specimen container of claim 12, wherein the specimen container is made of at least one of the group consisting of: a plastic;a polymer;a synthetic rubber;a composite;a metal; anda ceramic. 16. The specimen container of claim 12, further comprising wall flaps, attached to at least one of the deflector cap and the wall members, the wall flaps configured to extend at least one wall member as the deflector cap is raised and to deflect the fluid specimen material back inside the specimen container. 17. The specimen container of claim 16, wherein each wall flap is fixedly connected to the deflector cap. | IP Filepoint, LLC, A Delaware LLC | A61B005/00
US20100041969 | A1 | 2010-02-18 | MEASURING DEVICE AND METHOD FOR OPTICALLY DETERMINING THE CONCENTRATION OF BLOOD SUGAR AND/OR LACTATE IN BIOLOGICAL SYSTEMS | 2008-03-16 | The invention relates to a measuring device for optically determining the concentration of blood sugar and/or lactate in biological systems, comprising at least one IR radiation source, that radiates IR light on a volume that is to examined, and at least one measuring detector that detects light coming from the volume that is to be examined in order to determine the concentration of blood sugar and/or lactate, also by laymen in a simple manner and anywhere. According to the invention, the IR light radiated on the volume that is to be examined is supplied, prior to entry into the volume, to a reference detector. | 1. Measurement device for optical determination of the concentration of blood sugar and/or lactate in biological systems, having at least one IR radiation source that radiates IR light onto a volume to be investigated, having at least one measurement detector that absorbs light proceeding from the volume to be investigated, and having at least one reference detector to which the IR light radiated onto the volume to be investigated is passed before entry into the volume, whereby the radiation source, the reference detector, and the measurement detector are connected with one another by means of a lock-in, and an optical measurement path is formed between the IR radiation source and the measurement detector, having a first measurement path characteristic, and an optical reference path is formed between the IR radiation source and the reference detector, having a second measurement path characteristic that deviates from the first measurement path characteristic, wherein the IR radiation source radiates IR light onto the volume to be investigated in at least two discrete wavelengths, i.e. in at least two discrete wavelength bands. 2. Measurement device according to claim 1, wherein a beam splitter is disposed between the volume to be investigated and the radiation source, in such a manner that part of the light that proceeds from the radiation source is directed at the reference detector, and wherein at least part of the light that proceeds from the volume is directed at the measurement detector. 3. Measurement device according to claim 1, wherein the reference detector is directed at scattered light of the IR radiation source. 4. Measurement device according to claim 1, wherein the measurement detector is oriented in linear manner with regard to the IR light radiated onto the volume. 5. Measurement device according to claim 1, wherein the measurement detector is oriented at an angle with regard to the IR light radiated onto the volume. 6. Measurement device according to claim 1, wherein the radiation source is a laser diode. 7. Measurement device according to claim 6, wherein the laser diode has an emission frequency between 1,000 nm and 2,000 nm and/or between 2,000 nm and 3,000 nm. 8. Measurement device according to claim 6, further comprising means for modulation of the laser diode over a bandwidth below 170 nm, preferably below 20 nm. 9. Measurement device according to claim 8, further comprising two laser diodes. 10. Measurement device according to claim 1, further comprising means for modulation of the wavelength of the IR radiation source at least within a discrete wavelength band. 11. Measurement device according to claim 10, wherein the modulation means are connected to interact with a lock-in. 12. Method for optical determination of the concentration of blood sugar and/or lactate in biological systems, in which IR light is radiated onto a volume to be investigated, along an optical measurement path, and a value relevant to the c... | - | A61B005/1455 G01N021/35
US20100049068 | A1 | 2010-02-25 | IN-VEHICLE ELECTROCARDIOGRAPH DEVICE AND VEHICLE | 2009-08-21 | An in-vehicle electrocardiograph device includes: a direct electrode that is used to detect an electric potential of a body of a vehicle occupant in a state in which the direct electrode contacts a skin of the occupant; a capacitive-coupled electrode that is used to detect the electric potential of the body of the occupant in a state in which the capacitive-coupled electrode does not contact the skin of the occupant; and an electrocardiographic waveform determination unit that determines an electrocardiographic waveform of the occupant based on an electric potential at the direct electrode and an electric potential at the capacitive-coupled electrode. | 1. An in-vehicle electrocardiograph device, comprising: a direct electrode that is used to detect an electric potential of a body of a vehicle occupant in a state in which the direct electrode contacts a skin of the occupant;a capacitive-coupled electrode that is used to detect the electric potential of the body of the occupant in a state in which the capacitive-coupled electrode does not contact the skin of the occupant; andan electrocardiographic waveform determination unit that determines an electrocardiographic waveform of the occupant based on an electric potential at the direct electrode and an electric potential at the capacitive-coupled electrode. 2. The in-vehicle electrocardiograph device according to claim 1, wherein the electrocardiographic waveform determination unit determines the electrocardiographic waveform of the occupant based on a difference between the electric potential at the direct electrode and the electric potential at the capacitive-coupled electrode. 3. The in-vehicle electrocardiograph device according to claim 1, wherein the direct electrode includes one electrode. 4. The in-vehicle electrocardiograph device according to claim 1, wherein: the electrocardiographic waveform determination unit includes an amplifier that amplifies a signal indicating the electric potential at the capacitive-coupled electrode; andthe electrocardiographic waveform determination unit determines the electrocardiographic waveform of the occupant based on the amplified signal indicating the electric potential at the capacitive-coupled electrode. 5. The in-vehicle electrocardiograph device according to claim 4, wherein the capacitive-coupled electrode is integrally formed with the amplifier. 6. The in-vehicle electrocardiograph device according to claim 1, wherein: the direct electrode is connected to a ground;the capacitive-coupled electrode includes a plurality of electrodes; andthe electrocardiographic waveform determination unit uses the electric potential at the direct electrode as a reference electric potential, and determines the electrocardiographic waveform of the occupant based on a difference in electric potential between the plurality of electrodes included in the capacitive-coupled electrode. 7. The in-vehicle electrocardiograph device according to claim 6, wherein each one of the plurality of electrodes is connected to the ground. 8. The in-vehicle electrocardiograph device according to claim 1, wherein: the direct electrode is directly connected to a ground;the capacitive-coupled electrode is connected to the ground via a resistance; andthe electrocardiographic waveform determination unit uses the electric potential at the direct electrode as a reference electric potential, and determines the electrocardiographic waveform of the occupant based on a difference between the reference electric potential and the electric potential at the capacitive-coupled electrode. 9. The in-vehicle electrocardiograph device according to cla... | - | A61B005/0408 H01R043/00
US07657445 | B1 | 2010-02-02 | Method and system for managing healthcare facility resources | 2002-07-05 | The present invention provides a method and system for managing resources of an enterprise, in particular a healthcare enterprise. One aspect of the present invention provides a method and system for remote access and exchange of information for managing enterprise resources using a distributed computing environment such as the Internet or World Wide Web. The system provides real-time management of healthcare facility resources including clinical hours provided per patient per day. | 1. A healthcare facility management system comprising: a data terminal at a healthcare facility located in a geographic location having regulatory required hours of care to be provided to a patient per day based upon the acuity level of the patient; anda management control unit having a processor, the processor configured to: receive patient census data indicating the total number of patients present in the facility and the acuity level of each patient;compute a census projection for a predetermined period of time based upon received patient census data;calculate total budgeted clinical hours for the facility from the patient census data and a predetermined number of hours of care to be provided to a patient per day based upon the acuity level of the patient;automatically determine total required clinical hours for the facility from the patient census data and the regulatory required number of hours of care to be provided to a patient per day based upon the acuity level of the patient;monitor clinical hours provided at the facility;compare clinical hours provided at the facility to total clinical hours budgeted to project a deficit or surplus between total actual clinical hours provided and total budgeted clinical hours;compute a burn rate based upon at least the total clinical hours budgeted and actual clinical hours provided, the burn rate indicating a rate of revenue loss or gain;compute projected profit or loss for the facility for the predetermined period of time based upon at least the burn rate and the census projection;compare clinical hours provided at the facility to total required clinical hours to project a deficit or surplus between total actual clinical hours provided and total required clinical hours; andprovide a staffing schedule adjustment recommendation based upon the projected profit or loss and a comparison of clinical hours provided and projected deficit or surplus such that: actual clinical hours provided will be no greater than the total clinical hours budgeted and no less than the total required clinical hours for each facility; and a computation based upon at least the adjusted staffing schedule and the census projection will indicate a profit for the facility. 2. A computer readable medium having program instructions tangibly stored thereon executable by a processor to perform a method for managing clinical hours in a regulated inpatient healthcare facility, the method comprising: receiving patient census data indicating the total number of patients present in a healthcare facility and the acuity level of each patient;determining the total required clinical hours mandated by regulations to be provided by the facility during a predetermined period of time based upon the patient census data;receiving total budgeted clinical hours for the facility for the predetermined period of time;monitoring in real-time actual clinical hours provided at the healthcare facility;calculating the projected actual clinical hours provided d... | Rise Above Technologies, LLC | A61B005/00 G06F019/00
US07657294 | B2 | 2010-02-02 | Compliant diaphragm medical sensor and technique for using the same | 2005-08-08 | A unitary sensor is provided that includes a frame upon which electrical and optical components may be disposed and a coating, such as an overmolded coating provided about the frame. The unitary sensor includes one or more diaphragm structures formed from thin regions of the coating material and disposed generally about at least one of the optical components. The diaphragm structure or structures allow the optical components to be separated, such as by application of an opposing force to the lateral sides of the sensor, such that the sensor may be placed on a patient. The sensor may thereby be placed on a patient's finger, toe, and so forth to obtain pulse oximetry or other physiological measurements. | 1. A sensor assembly, comprising: a frame;an emitter disposed on the frame;a detector disposed on the frame; anda coating molded over at least the frame to form at least one diaphragm structure. 2. The sensor assembly of claim 1, wherein the sensor assembly comprises at least one of a pulse oximetry sensor, a sensor for measuring a water fraction, or a combination thereof. 3. The sensor assembly of claim 1, comprising one or more signal transmission structures connected to at least one of the emitter or the detector. 4. The sensor assembly of claim 1, comprising a first diaphragm structure disposed generally about the emitter. 5. The sensor assembly of claim 1, comprising a first diaphragm structure disposed generally about the detector. 6. The sensor assembly of claim 1, comprising a first diaphragm structure disposed generally about the emitter and a second diaphragm structure disposed generally about the detector. 7. The sensor assembly of claim 1, wherein the frame comprises a single molded part. 8. The sensor assembly of claim 1, wherein the frame comprises two or more parts joined to form the frame. 9. The sensor assembly of claim 1, comprising two opposing lateral sides generally transverse to the at least one diaphragm structure. 10. The sensor assembly of claim 9, wherein the emitter and the detector are configured to be moved apart when an opposing force is applied to the two opposing lateral sides. 11. The sensor assembly of claim 1, wherein the coating comprises a thermoplastic elastomer. 12. The sensor assembly of claim 11, wherein the thermoplastic elastomer comprises at least one of a thermoplastic polyolefin, a thermoplastic vulcanizate alloy, thermoplastic polyurethane, silicone, or a combination thereof. 13. The sensor assembly of claim 1, wherein the coating comprises a conformable material. 14. The sensor assembly of claim 1, wherein the sensor assembly is fluid tight. 15. The sensor assembly of claim 1, wherein the coating is chemically bonded to the frame. 16. The sensor assembly of claim 1, comprising a pulse oximeter monitor configured to receive data from the sensor assembly. 17. The sensor assembly of claim 1, comprising a multi-parameter monitor configured to receive data from the sensor assembly. 18. The sensor assembly of claim 1, comprising at least one integrated circuit device. 19. The sensor assembly of claim 1, comprising a cable comprising one or more integrated circuits. 20. The sensor assembly of claim 1, comprising an adapter cable connected to one or more signal transmission structures of the sensor assembly. 21. The sensor assembly of claim 1, wherein each diaphragm structure comprises one or more regions of the coating configured to allow at least one of the emitter or the detector to be displaced upon placement of the sensor assembly on a patient. 22. The sensor assembly of claim 1, wherein the emitter comprises one or more light emitting diodes. 23. The sensor assembly of claim 1, w... | Nellcor Puritan Bennett LLC | A61B005/1455
US07662112 | B2 | 2010-02-16 | Medical measuring device and method | 2005-05-13 | A subsurface wound measuring device and method are provided for measuring a subsurface wound or cavity of a patient, the device including a flexible strip having first and second sections, the sections having first and second ends respectively with an intermediate bend being located between the first and second ends. The first end is adapted for being located internally to a patient, the second end being adapted for being located externally to the patient. The first and second sections have generally concave-shaped cross-sections with measuring indicia being located along at least one of the sections. The device also optionally includes a measuring strip and a receiver mounted on the second end for receiving the measuring strip, where the measuring strip and the flexible strip form complementary structure for measuring a subsurface wound or cavity. | 1. A wound measuring method for measuring at least one physical parameter of a patient with a subsurface wound having a cavity with a cavity opening and a sidewall, said method comprising: providing a measuring device with a flexible measuring strip having a first end, a second end and an exterior surface with measuring indicia;providing said measuring device with a flexible guide strip having a first end, a second end and an exterior surface;providing each said strip with a generally arcuate, concave, transverse, cross-sectional configuration;providing a collar or sleeve with a passage extending therethrough;slidably receiving said measuring strip in said collar or sleeve passage in a frictional, gripping relation;inserting the flexible guide strip at least partially into the cavity by the flexible guide strip second end, the second end extending towards the cavity opening, the first end extending towards the cavity sidewall;inserting the flexible measuring strip at least partially into the cavity in sliding contact on the guide strip with the measuring strip first end extending towards the cavity sidewall;locating the measuring device in association with the cavity sidewall;identifying a distance along the cavity sidewall with the measuring indicia;forming a bend in the flexible measuring strip at a variably repositionable bend location between said measuring strip first and second ends;relating the bend location to a dimension of the cavity;dividing said flexible measuring strip at said bend into first and second semi-rigid sections terminating at said first and second strip ends respectively;moving the collar or sleeve into position in proximity said bend to maintain the bend and to provide a visual reference for the bend location along the flexible measuring strip, and further to lock the position of the flexible measuring strip in place;identifying the distance with said first and second semi-rigid sections extending along respective portions of said cavity;recording the distance;providing a receptacle with a closed end at the measuring strip first end;providing the receptacle with an opening in spaced relation and oriented away from said measuring strip first end;obtaining a tissue or fluid sample in said receptacle from within said wound; andextracting said sample from the wound and analyzing said sample. 2. The wound measuring method according to claim 1, which includes the additional step of providing said receptacle with a radiused configuration. | KCI Licensing, Inc. | A61B005/103
US07662111 | B2 | 2010-02-16 | Vacuum assisted auto-lancing device | 2004-02-20 | Disclosed herein is a vacuum assisted auto-lancing device. The device includes a housing (200) having a body (210) which has a lever hole. An actuating lever (300) includes an actuating switch (310) seated on the lever hole and having first and second actuating steps, a switch cap (320), and a switch cover (330). The device also includes a holding unit (400) having a first stem (410), a second stem (430), and a stem cap (450). A trigger (500) has a first trigger unit activated by the first actuating step, and a second trigger unit activated by the second actuating step. A lancet holder (524) holding a lancet (522) is secured to an end of the first trigger unit. The device includes a blood collecting unit (600) having an adjusting screw, an adjusting slider, and an end cap, and a vacuum unit (700) having a plunger and a body cap. | 1. A vacuum assisted auto-lancing device, comprising: a housing having a body, the body being partitioned into a first chamber and a second chamber by a first partition plate, with a lever hole provided on an outer periphery of the first chamber at a position adjacent to a first end of the first chamber, and a locking protrusion vertically protruding from a position around the lever hole;an actuating lever, comprising: an actuating switch seated on the lever hole, and having at a lower portion thereof first and second actuating steps which pass through the lever hole;a switch cap covering the actuating switch; anda switch cover locking the actuating switch and the switch cap to a predetermined position of the body;a holding unit, comprising: a first stem coupled at a first end thereof to an inner periphery of the first chamber, a second end of the first stein being exposed to an outside of the first chamber;a second stem integrally coupled to the first end of the first stem and placed in the first chamber; anda stem cap to close a first end of the second stem;a trigger comprising: a first trigger unit disposed at a position in the first and second stems, and activated by a first actuating step, with a lancet holder holding a lancet and being secured to an end of the first trigger unit which extends to the first stem; anda second trigger unit rotatably mounted to cover both sides and a top of the second stem, and activated by a second actuating step;a blood collecting unit, comprising: an adjusting screw positioned at the first end of the first chamber and rotatably secured to the first stem;an adjusting slider fastened to the adjusting screw; andan end cap mounted to the first stem; anda vacuum unit, comprising: a plunger placed to reciprocate in the second chamber, and selectively passing through the first partition plate to engage with the second trigger unit; anda body cap to lock the plunger to the second chamber. 2. The vacuum assisted auto-lancing device according to claim 1, wherein the actuating switch is configured so that the second actuating step thereof is shorter in length than the first actuating step thereof, and a first vacuum hole extends from an upper portion of the actuating switch and passes between the first and second actuating steps, and the switch cap is made of a flexible material to seal a gap between the actuating switch and the outer periphery of the body, with a second vacuum hole formed on an upper surface of the switch cap to communicate with the first vacuum hole, and the switch cover has at a central portion thereof a through hole to allow the actuating switch surrounded by the switch cap to pass through the switch cover, with a plurality of locking recesses provided on a lower surface of the switch cover to engage with the locking protrusions. 3. The vacuum assisted auto-lancing device according to claim 1, wherein the first partition plate has on a central portion thereof a first spring mounting protrusion pro... | - | A61B005/00
US07668589 | B2 | 2010-02-23 | Integrated, plural-signal, heart-signal sensor system and related data-presentation methodology | 2005-10-25 | A system and a related methodology for gathering, during a selected time span, and from a common anatomical site, time-contemporaneous ECG-electrical and heart-sound signals including (1) processing such signals to effect (a) time-based, related ECG fiducials, and (b) systolic and diastolic heart-sound indicators, and (2) creating a reportable data stream which communicates such effected fiducials and indicators in a manner whereby time-based relationships between them, and non-time-based differentiation between systolic and diastolic heart-sound indicators, are made visually discernible. The methodology of the invention may also be implemented strictly for the gathering and processing of heart sounds. | 1. A patient-portable heart-condition monitoring and reporting system comprising an anatomy-contact sensor including an audio sensing element, and plural electrical-signal sensing elements, andsignal-processing circuit structure operatively connected to said elements, operable, during a selected time interval, (a) to collect therefrom audio and electrical-ECG signals detected by said elements under a circumstance with said sensor contacting a selected anatomical site on a person's anatomy, and (b) to create therefrom an output signal constructed to generate, on a common time scale, a time-based graphical, laddergram display keyed to selected ECG fiducials, and including time-specific audio-event heart-sound marks datumed to, and projecting linearly from opposite, upper and lower sides of, a generally horizontal neutral axis, said marks including both systolically-related heart-sound marks deployed on one of said opposite sides of said axis, and diastolically-related heart-sound marks deployed on the other, opposite side of said axis. 2. The system of claim 1, wherein said systolically-related and diastolically-related marks are characterized with excursion heights measured relative to said axis which are indicators of the intensities of the respective, associated heart sounds. 3. The system of claim 2, wherein said output signal is further constructed to generate, in said display, certain heart-sound marks having lateral widths whose sizes are related to frequency and/or intensity-content characteristics of the respective, associated heart sounds. 4. The system of claim 1, wherein said output signal is further constructed to generate, in said display, certain heart-sound marks having lateral widths whose sizes are related to frequency and/or intensity-content characteristics of the respective, associated heart sounds. 5. The system of claim 1, wherein said output signal is further constructed to generate additional heart-sound information in the forms of ECG and heart-sound-acoustic waveforms that are time aligned with said laddergram-display marks. 6. During a selected time span, gathering time-contemporaneous ECG-electrical and heart-sound-acoustic signals, processing such electrical and acoustic signals to create a display-effecting output signal, andemploying the created output signal, producing therefrom a time-based laddergram display (a) designed for reading in synchronized relation with at least one of an ECG waveform marked with selected fiducials, and an audio waveform picturing heart acoustic activity, and (b) possessing time-specific, audio-event, heart-sound marks datumed to, and projecting linearly from opposite, upper and lower sides of, a generally horizontal, neutral axis, such marks including both systolically-related marks deployed on one of the opposite axis sides, and diastolically-related marks deployed on the other axis side. 7. During a selected time span, gathering time-contemporaneous ECG-electrical and heart-sound... | Inovise Medical, Inc. | A61B005/044
US07658192 | B2 | 2010-02-09 | Method and device for treatment of obstructive sleep apnea | 2007-07-03 | This invention relates to devices and methods for maintaining upper airway patency during sleep for the treatment of sleeping disorders, upper airway resistance syndrome, snoring and obstructive sleep apnea. | 1. A method for maintaining a patient's tongue to prevent sleep related breathing disorders comprising the steps of: forming a lateral channel at the base of the patient's tongue;positioning a restrainer, said restrainer being elastic, and said restrainer being sized and dimensioned to fit into the lateral channel wherein the restrainer is capable of suspending the tongue away from the patient's throat during sleep; andsecuring the restrainer to the interior surface of the patient's mandible to suspend the tongue away from the patient's throat during sleep. 2. The method of claim 1 further comprising providing the restrainer in the form of a cylinder with a solid center section throughout the entire length of the cylinder. 3. The method of claim 2 further comprising providing the solid center section in the form of filaments and fibers. 4. The method of claim 1 wherein the lateral channel is formed by tongue tissue incision and dissection with a scalpel. 5. The method of claim 1 wherein the lateral channel is formed with an electrosurgical device. 6. The method of claim 1 wherein the lateral channel is formed by a needle passer that pulls the restrainer through the tongue tissue and tunnels through the tongue tissue. 7. The method of claim 1 further comprising providing the restrainer in the form of a planar sheet. 8. The method of claim 1 further comprising providing the restrainer in the form of at least one planar sheet with a solid portion positioned between two porous sections. 9. The method of claim 1 further comprising providing the restrainer in the form of at least two individual planar sections. 10. The method of claim 1 further comprising providing the restrainer in the form of two solid center end sections. 11. The method of claim 1 further comprising providing the restrainer in the form of a solid central section. 12. The method of claim 1, wherein the restrainer comprises an elastic material. 13. The method of claim 1 wherein the restrainer comprises pores that are sized and dimensioned to support growth of vascularized tissue within the restrainer. | Douglas C. Harrington | A61F005/56 A61F013/00 A61B005/117
US07652582 | B2 | 2010-01-26 | System and method for ergonomic tracking for individual physical exertion | 2007-11-20 | A system and method are provided for tracking a user's posture. The system and method include an environmental module for determining the user's physical environment, a biomechanical module for determining at least a user's posture in the user's physical environment, and an output module for outputting to the user an indication of at least the user's posture relative to at least a target correct posture. The physical environment can include a computer workstation environment, a manufacturing environment, a gaming environment, and/or a keypadding environment. | 1. A computer system for providing real-time feedback about a user's posture, comprising: an environmental module for determining a subject physical environment;a biomechanical module for determining at least a portion of a user's posture in the determined subject physical environment; andan output module for providing to the user a real-time indication of the determined user's posture relative to a target correct posture associated with the determined subject physical environment, the target correct posture being in accord with a standardized posture for the determined subject physical environment. 2. The system of claim 1, wherein the standardized posture is an American National Standards Institute (ANSI) standard. 3. The system of claim 1, wherein the standardized posture is a published standard. 4. The system of claim 1, wherein the standardized posture is an industry standard. 5. A computer system for providing real-time feedback about a user's posture, comprising: an environmental module for determining a subject physical environment of the user;a biomechanical module for determining at least a portion of a user's posture in the determined subject physical environment, the portion of a user's posture including a body portion controlled by at least one of small muscles and large muscles; andan output module for providing to the user a real-time indication of the determined user's posture relative to a target correct posture associated with the determined subject physical environment. 6. The computer system of claim 5, wherein the body portion controlled by at least one of small muscles and large muscles includes at least one of: a torso; a hand area; a neck; an arm area; and a leg area. 7. A computer system for providing real-time feedback about a user's posture, comprising: an environmental module for determining a subject physical environment;a biomechanical module for determining at least a portion of a user's posture in the determined subject physical environment; andan output module for providing to the user a real-time indication of the determined user's posture relative to a target correct posture associated with the determined subject physical environment, the user's posture being one of a seated posture, a standing posture, a reaching posture, and a manipulating posture. 8. The computer system of claim 7 wherein the biomechanical module further determines at least a portion of a user's posture when transitioning from a first posture to a subsequent posture; and wherein the output module further provides a real-time indication of the determined user's transition relative to a target transition from the first posture to the subsequent posture. 9. The computer system of claim 7 wherein the user's posture is static. 10. A method for providing real-time feedback about a user's posture, comprising computer implemented steps of: determining a subject physical environment;determining at least a portion of a user's posture based on ... | - | G08B023/00 A61B005/00
US07648467 | B2 | 2010-01-19 | Device for non-surgical correction of congenital inverted nipples and/or collection of nipple aspirate fluid | 2005-12-02 | The invention is an apparatus (10) for treating inverted nipples and/or the collection of NAF. The invention includes an assembly which creates an airtight seal upon attachment to at least an areola (16) of an inverted nipple (18) and by rotation of a part thereof draws a partial vacuum to apply an outward force to the nipple; and a syringe (26) including a curved longitudinal axis (28), the syringe being carried by the assembly and being in fluid communication with a chamber (24) of the assembly upon attachment of the assembly to at least an areola of an inverted nipple, the syringe including a plunger (30) which is sealed against an inner surface (34) of the syringe and is rotatable to move along the longitudinal axis relative to the chamber which moves the plunger within a syringe body to expel air from the syringe body away from the chamber upon rotation of the plunger in a first direction and to aspirate air from the chamber upon attachment of the apparatus to the breast surrounding the nipple upon rotation in a second direction opposite to the first direction to create a partial vacuum within the chamber to evert the nipple. | 1. An apparatus for treating inverted nipples and/or collecting nipple aspirate fluids comprising: a flexible attachment part including an opening, the attachment part being adapted for attachment to at least an areola of a nipple to create an airtight seal between the attachment part and the areola;a chamber extending away from the opening toward which the nipple may project upon application of a vacuum to the chamber while the attachment part is attached;a syringe plunger;a syringe body including a curved longitudinal axis which is curved around the chamber, the syringe body engaging an outer surface of the chamber, the syringe body including a first open end for receiving the syringe plunger which creates an airtight seal with an interior surface of the syringe body when received in the syringe body and a second end;a valve in fluid communication with the body, the syringe body being in fluid communication with the chamber, the valve opening to expel air away from the chamber upon movement of the syringe plunger within the syringe body toward the second end, movement of the syringe plunger within the syringe body away from the second end causing aspiration of air from the chamber to create a partial vacuum within the chamber when the attachment part is attached at least to the areola; anda cap rotatably attached to an outer surface of the chamber, the cap carrying the syringe plunger and rotation in a first direction while the syringe plunger is received in the syringe body causing air to be expelled from the syringe body through the valve and when an airtight seal exists between the attachment part and the breast surrounding the nipple air is rotation in a second direction moving the syringe plunger away from the second end to produce a partial vacuum within the chamber to apply an outward force to the nipple. 2. An apparatus in accordance with claim 1 comprising: an absorbent material mounted within the chamber which upon contact with the nipple absorbs any nipple aspirate fluid; anda biasing mechanism for biasing the absorbent material to contact the nipple which applies a force to the material to cause contact with the nipple while the apparatus is worn on the breast surrounding the nipple. 3. An apparatus in accordance with claim 2 wherein the biasing mechanism includes a spring attached to an inner end of the cap and to a material holder which removably receives the material, the spring causing the material holder to cause contact of the material with the nipple. 4. An apparatus in accordance with claim 1 comprising: means for removably and rotatably attaching the cap to a curved portion of an outer surface of the chamber;the fluid communication of the syringe body with the chamber being through a conduit connecting the syringe body with an interior of the chamber; anda racket mechanism for rotatably stopping the cap relative to the chamber so that the syringe plunger is settable in different rotational positions relative to the chamb... | - | A61B005/00 B65D081/00
US07647096 | B2 | 2010-01-12 | Methods and apparatus for calculating and presenting the probabilistic functional maps of the human brain | 2002-02-18 | The present invention proposes a method for calculating, presenting, and combining probabilistic functional maps (PFMs) of the human brain representing the probability of structures existing. The method comprises three major steps: reading of data containing the coordinates of contacts, calculating the PFMs, and presenting the PFMs. The data can be read from a file in text or binary format or from a database as local or remote client. The PFM calculation comprises the following steps: forming 3D models of contacts, normalizing the contact models, voxelizing the contact models, calculating an atlas function, and calculating the PFM. The PFM can be presented alone or along with anatomical atlases. Both 3D and 2D interfaces can be used for presentation. The proposed method also includes different ways of combining the contact data and/or existing PFMs from multiple sources. This mechanism is the basis of an internet portal for stereotactic and functional neurosurgery. | 1. A computer-implemented method for obtaining and presenting a functional distribution of brain structures, the method comprising: obtaining contact data specifying the locations, within the brains of a plurality of individuals, of active electrical contacts which detect a brain structure;generating an active contact model for each of the active contacts, wherein each active contact model represents a physical shape and/or electrical properties of one of the active contacts;transforming the generated active contact models into a common space for the individuals;calculating a probability function for locating the brain structure based on the obtained contact data and the transformed active contact models, said probability function indicating the space density of the active contacts; andpresenting a representation of the calculated probability function. 2. A method according to claim 1 wherein said transforming includes normalizing the generated active contact models to transform said generated active contact models into the common space. 3. A method according to claim 2 in which the normalization is performed using a homeomorphism. 4. A method according to claim 1 wherein said calculating includes deriving an atlas function which, at every point in a three-dimensional region, is the number of active contact normalized models which are sensitive to a structure at that point. 5. A method according to claim 4 which includes deriving an analytic representation of the atlas function. 6. A method according to claim 5 in which the analytic representation is a set of ellipsoidal functions, where each has a centre at a mean of the atlas function defined on some cluster and semi-axial lengths determined by standard deviations and correlations of the atlas function defined on the cluster. 7. A method according to claim 4 further including smoothing the atlas function. 8. A method according to claim 4 further including using the atlas function to derive a value for each voxel of an array of voxels. 9. A method according to claim 1 further including calculating a respective atlas function for each voxel of an array of voxels, wherein said atlas function represents the number of active contacts which are sensitive to a brain structure located in or intersecting with that voxel, and wherein said calculating the probability function for locating the brain structure is based on the atlas function. 10. A method according to claim 9 in which the atlas function for each respective voxel is compensated based on a numerical factor which indicates a proportion of the corresponding voxel which is within the models of the active contacts. 11. A method according to claim 9 in which the voxels are not cubic. 12. A method according claim 1 in which the contact data is obtained from a pre-generated data file. 13. A method according to claim 1 in which the obtained data is represented as a hierarchical tree structure with levels corresponding to (i) brain structur... | Kent Ridge Digital Labs | A61B005/0476 A61B005/05 G06K009/00
US08121667 | B2 | 2012-02-21 | Interfaced base unit and display system for an MRI magnet room | 2007-05-10 | A base unit operable with a wireless patient monitoring unit used to acquire physiological data during an MRI examination is constructed to have a detachable display unit. The detachable display unit wirelessly communicates with the base unit when in a roving mode. The detachable display unit may also have a magnetically-hardened power supply that does not saturate during an MRI examination. Such a magnetically-hardened power supply allows the display unit to operate when proximate or in the magnetic field generated by an MRI machine during an MRI examination. | 1. A base station for an MRI compatible patient monitor device, the base station positionable outside a bore of a magnet during an MRI examination, the base station comprising: a movable stanchion disposed in a shielded magnetic resonance room;a monitor unit which receives wireless transmissions of data from a patient sensor that acquires physiological signals from a patient in the bore of the magnet, the monitor being mounted on the stanchion; anda display unit which wirelessly communicates with the monitor unit and displays information associated with the data received from the patient sensor by the monitor unit, the display unit being movable independently of the monitor unit in the shielded magnetic resonance room, the display init including:a rechargeable battery that provides power to the display unit when the display unit is remote from the stanchion, the stanchion providing charging current to the rechargeable battery when the display unit is mounted to the stanchion, anda magnetically hardened power supply operative, without saturation, in high-flux magnetic fields up to 15,000 Gauss to manage the rechargeable battery during an MRI examination. 2. The base station of claim 1 wherein the monitor unit receives wireless transmissions of data from the patient sensor on radio waves on a first channel and transmits data to the display unit on radio waves on a second channel different from the first channel. 3. The base station of claim 1 wherein the display unit includes: an LCD display screen; andLEDs which backlight the LCD display screen. 4. The base station of claim 1 wherein the display unit includes: a backlight display screen and backlights which operate in the bore of the magnet and do not interfere with the MRI examination;a rechargeable battery which powers the backlight display; anda magnetically hardened power supply which recharges the rechargeable battery without saturation when positioned adjacent to the bore of the magnet. 5. The base station of claim 1 wherein the stanchion carries a first set of contacts and the display device carries a second set of contacts, the first and second sets of contacts communicating to transfer power from the stanchion to the display unit when the display unit is temporarily mounted to the stanchion. 6. The base station of claim 5 wherein the monitor unit automatically suspends wireless communications to the display unit when the display unit is mounted to the stanchion. 7. A base station for an MRI compatible patient monitor device, the base station positionable outside the bore of the magnet during an MRI examination and comprising: a monitor unit which receives wireless transmissions of data from a patient sensor that acquires physiological signals from a patient;a display unit positionable remotely from the monitor unit and in wireless communication with the monitor unit to display information associated with the data received from the patient sensor by the monitor unit;a stanchion, the ... | Koninklijke Philips Electronics N.V. | A61B005/05
US08301217 | B2 | 2012-10-30 | Multiple wavelength sensor emitters | 2009-09-28 | A physiological sensor is adapted to removably attach an emitter assembly and a detector assembly to a fingertip. The emitter assembly is adapted to transmit optical radiation having multiple wavelengths into fingertip tissue. The detector assembly is adapted to receive the optical radiation after attenuation by the fingertip tissue. The sensor has a first shell and a second shell hinged to the first shell. A spring is disposed between the shells and urges the shells together. | 1. A physiological sensor adapted to removably attach an emitter assembly and a detector assembly to a fingertip, the emitter assembly adapted to transmit optical radiation having a plurality of wavelengths into fingertip tissue, the detector assembly adapted to receive the optical radiation after attenuation by the fingertip tissue, the sensor comprising: a first shell including a proximal end and a distal end and an axis running from the proximal end to the distal end;a second shell hingedly coupled to the first shell at a first location proximate the proximal end of the first shell;a spring disposed between and urging together the shells, the spring including: a coil located proximate the first location,a first leg extending from the coil to a second location proximate the distal end of the first shell;an emitter assembly fixedly attached to one of the first shell or second shell; anda detector assembly fixedly attached to the other of the first shell and second shell. 2. The physiological sensor according to claim 1, further comprising an anchor configured to movably anchor the first leg of the spring to the first shell. 3. The physiological sensor according to claim 1, wherein the spring is configured so as to create a pivot point along a finger gripped between the first shell and the second shell that is substantially behind a fingertip. 4. The physiological sensor of claim 1, wherein the first leg and the second leg extend from a coiled portion of the spring and the first leg is longer than the second leg. 5. The physiological sensor of claim 4, wherein the first leg extends from outer portions of the coiled portion of the spring and the second leg extends from an inner portion of the coiled portion of the spring. 6. The physiological sensor of claim 1, wherein the first leg comprises a first portion that extend along a first side of the first shell and a second portion that extends along a second side of the first shell. 7. A physiological sensor comprising: a first shell including a proximal end and a distal end;a second shell at least partially hinged to the first shell at a first location proximate the proximal end of the first shell;a spring disposed between and urging together the shells, the spring including a first portion extending along at least a portion of a length of the first shell from a second location proximate the first location to a third location proximate the distal end of the first shell. 8. The physiological sensor of claim 7, further comprising a pad configured to position a tissue portion of a patient. 9. The physiological sensor according to claim 8 further comprising a contour defined along the pad and generally shaped to conform to a fingertip. 10. The physiological sensor according to claim 9 further comprising a tissue guide raised from, and defined along, the pad and adapted to increasingly conform to a fingertip inserted along the pad. 11. The physiological sensor according to claim 10 further comp... | Cercacor Laboratories, Inc. | A61B005/1455
US08079999 | B2 | 2011-12-20 | Fluid injector system | 2009-07-02 | Fluid injector systems used for a variety of imaging and injection procedures are disclosed. The systems may include separate modules or assemblies that may be located in different rooms of a hospital or imaging facility. Various single use and multiple use components may also be used with the modules or assemblies of the system. In addition, the injector system may include hydraulic and/or pneumatic fluid sources to power the system modules of the present invention. | 1. A method of injecting a first fluid into a patient during a magnetic resonance imaging procedure, said method comprising: providing communication between a cylinder assembly in a control room of an imaging facility and a console in said control room, said console comprising a user control device;providing communication between a hydraulic motor in a patient room of said imaging facility and a peristaltic pump in said patient room;providing communication between said cylinder assembly in said control room and said hydraulic motor in said patient room;driving said cylinder assembly with a motor of said console;providing a second fluid from said cylinder assembly under pressure;driving said hydraulic motor with said second fluid that is fed through said hydraulic motor to cause said hydraulic motor to generate rotary forces;driving said peristaltic pump with said rotary forces generated by said hydraulic motor; andinjecting said patient with said first fluid using said peristaltic pump. 2. The method of claim 1, wherein said injecting of said first fluid is caused by rotation of a rotor of said peristaltic pump which causes rollers of said rotor to compress a tubing, thereby drawing in said first fluid and transporting said first fluid through said tubing and into said patient. 3. A method of injecting a first fluid into a patient during a magnetic resonance imaging procedure, said method comprising: providing communication between a cylinder assembly in a control room of an imaging facility and a console in said control room, said console comprising a user control device;providing communication between a hydraulic motor in a patient room of said imaging facility and a gear cartridge in said patient room;providing communication between said cylinder assembly in said control room and said hydraulic motor in said patient room;driving said cylinder assembly with a motor of said console;providing a second fluid from said cylinder assembly under pressure;driving said hydraulic motor with said second fluid that is fed through said hydraulic motor to cause said hydraulic motor to generate rotary forces;driving said gear cartridge with said rotary forces generated by said hydraulic motor; andinjecting said patient with said first fluid using said gear cartridge. 4. The method of claim 3, wherein said injecting of said first fluid is caused by rotation of a rotor which causes gears in said gear cartridge to rotate and compress a tubing located between said gears, thereby drawing in said first fluid and transporting said first fluid through said tubing and into said patient. 5. A system to inject a first fluid into a patient during a magnetic resonance imaging procedure, said system comprising: a console in a control room of an imaging facility, said console comprising a user control device;a cylinder assembly in said control room of said imaging facility, said cylinder assembly being in communication with said console;a peristaltic pump in a patient r... | ACIST Medical Systems, Inc. | A61M031/00
US08252080 | B2 | 2012-08-28 | Liquid absorbing filter assembly and system using same | 2009-10-06 | A filter assembly for use in sidestream gas sampling assembly. The filter assembly of the present invention includes a hydrophilic. The filter further includes a Hydrophobic member positioned downstream of the hydrophilic liner. The hydrophobic disk acts as a second line of defense against moisture reaching the sensing mechanism member that is constructed from a porous material and is situated such that gas entering the filter assembly passes through at least a portion the hydrophilic member. In a further embodiment, the filter includes a hydrophobic member positioned downstream of the hydrophilic member. The hydrophobic member acts as a second line of defense against moisture reaching the sensing mechanism in the gas sampling assembly. | 1. A filter assembly for use in a sidestream gas sampling assembly comprising: a housing having an inlet, an outlet, and an inner wall defining a channel through the housing from the inlet to the outlet;a hydrophobic member disposed across the channel at a location proximate to the outlet; anda hydrophilic member having an inlet, an outlet, an outer wall, and a pore volume, wherein the hydrophilic member comprises a hydrophilic liner disposed on the inner wall of the housing defining the channel at a location proximate to the hydrophobic member such that fluid passing through the channel passes over at least a portion of the outer wall of the hydrophilic liner before passing through at least a portion of the hydrophobic member. 2. The filter assembly of claim 1, wherein the pore volume ranges from approximately 40% to approximately 90%. 3. The filter assembly of claim 1, wherein the hydrophilic member has a first edge that contacts a surface of the hydrophobic member. 4. The filter assembly of claim 1, wherein the hydrophilic member covers a majority of the inner wall extending between the inlet and the hydrophobic member. 5. The filter assembly of claim 1, wherein the outer wall of the hydrophilic member is substantially in contact with a majority of the inner wall extending between the inlet and the hydrophobic member. 6. The filter assembly of claim 1, further comprising a gas sampling line coupled to the inlet. 7. The filter assembly of claim 6, further comprising an exit conduit coupled to the outlet. 8. The filter assembly of claim 1, wherein the housing is cylindrical and wherein the hydrophobic member is disc-shaped. 9. A sidestream gas monitoring system comprising: (a) a sampling line having a first end and a second end, wherein the sampling line is adapted to carry a flow of gas from a patient circuit;(b) a filter assembly comprising: (1) a housing having an inlet adapted to be coupled to the second end of the sampling line, an outlet, and an inner wall defining a channel through the housing from the inlet to the outlet,(2) a hydrophobic member disposed across the channel at a location proximate to the outlet, and(3) a hydrophilic member having an inlet, an outlet, an outer wall, and a pore volume, wherein the hydrophilic member comprises a hydrophilic liner disposed on the inner wall of the housing defining the channel at a location proximate to the hydrophobic member such that fluid passing through the channel passes over at least a portion of the outer wall of the hydrophilic liner before passing through at least a portion of the hydrophobic member;(c) a sample cell operatively coupled to an outlet of the housing; and(d) a detecting system adapted to measure a constituent of gas contained within the sample cell. 10. The system of claim 9, wherein the hydrophilic member has a first edge that contacts a surface of the hydrophobic member. 11. The system of claim 9, wherein the hydrophilic member covers a majority of the inner ... | RIC Investments, LLC | B01D046/00
US08632469 | B2 | 2014-01-21 | Devices, systems, and methods for mapping organ profiles | 2008-01-23 | Devices, systems, and methods for the localization of body lumen junctions and other intraluminal structure are disclosed. Various embodiments permit clinicians to identify and locate lesions and/or anatomical structures within a lumen and accurately place leads and/or devices within a lumen, through determining the intralumen conductance and/or cross-sectional area at a plurality of locations within the body lumen. | 1. A system comprising: a device having a proximal end and a distal end, the distal end of the device configured for placement into a body lumen, the device comprising a first electrode located between the proximal and distal ends of the device; anda processor operably connected to the first electrode of the device, the processor operable to collect conductance data from the device to determine a profile of the body lumen based upon the conductance data, the conductance data collected at a plurality of locations within the body lumen and determined at each of the plurality of locations when the first electrode is immersed in a fluid within the body lumen;wherein the processor further comprises a computational program configured to detect a bifurcation in the body lumen based upon the conductance data obtained at the plurality of locations and changes in cross-sectional area identified using the conductance data, the conductance data including data obtained at the bifurcation using the device, and distinguishing the bifurcation from a lesion. 2. The system of claim 1, wherein the processor is further capable of calculating a cross-sectional area of the body lumen at each of the plurality of locations within the body lumen using the conductance data. 3. The system of claim 1, wherein the computational program is further configured to adjust the conductance data to eliminate a conductance spike due to the detection of a bifurcation. 4. The system of claim 1, wherein the device is selected from the group consisting of an impedance catheter and an impedance wire. 5. The system of claim 1, wherein, the body lumen comprises an aortic lumen. 6. The system of claim 5, wherein the lesion comprises an aneurysm. 7. The system of claim 1, wherein the body lumen comprises a lumen of a coronary artery. 8. The system of claim 1, wherein the device further comprises a third electrode and a fourth electrode, wherein the first and third electrodes comprise excitation electrodes and the second and fourth electrodes comprise detection electrodes; and the third and fourth electrodes are operably connected to the processor. 9. The system of claim 8, wherein the fluid comprises blood. 10. The system of claim 1, wherein the computational program comprises a smoothing algorithm. 11. The system of claim 10, wherein the computational program further comprises a threshold limit. 12. The system of claim 1, wherein the computational program comprises a gradient. 13. The system of claim 1, wherein the computational program comprises an interpolation. 14. The system of claim 1, wherein the computational program comprises a threshold limit. 15. A system comprising: a catheter having a proximal end and a distal end, the distal end of the catheter configured for placement into a body lumen, the catheter comprising a first electrode and a second electrode located between the proximal and distal ends of the catheter, the catheter further comprising a passageway configur... | 3DT Holdings, LLC | A61B005/02
US09198597 | B2 | 2015-12-01 | Leading-edge cancer treatment | 2009-05-22 | Kits and methods for the treatment of certain types of cancers, specifically various primary brain cancers. In some embodiments, treatment is directed toward the known areas of cancer cell infiltration and along pathways of likely migration ahead of established areas of cancer cell infiltration. In some embodiments, cancer cells are targeted where they have likely spread, but yet remain undetected. Some embodiments relate to a means of automatically directing radiological analysis along likely pathways of cancer cell migration to precisely determine the extent of detectable cancer spread. In some embodiments, treatments are directed to a predetermined distance along such pathways. | 1. A method for treating a cancerous tumor in a brain, the method comprising: identifying a bulk of the tumor or a nidus;predicting regions of likely cancer-cell migration based at least in part on patient population data corresponding to regions of likely or demonstrated cancer-cell migration in a patient population, wherein predicting the regions of likely cancer-cell migration comprises analyzing a location of the bulk of tumor or the nidus;imaging the brain to determine whether cancer cells have already migrated to the predicted regions;tracing a trajectory along a tissue region that is separated from the bulk of the tumor or the nidus by a sufficient distance to ensure that cancer cells have not spread beyond that distance, the trajectory being spaced apart from the bulk of the tumor or the nidus by a minimum separation;selectively treating likely healthy white matter pathways in the regions of likely cancer-cell migration using a stereotactic radiosurgery surgical apparatus to provide radiation treatment to tissue residing beyond the spreading cancer cells, wherein the white matter pathways selected for treatment are determined at least in part by the location of the bulk of tumor or the nidus. 2. A computer-implemented method of determining a target region for treating a cancerous tumor in a brain using one or more computer processors configured to execute the steps, the method comprising: distinguishing areas of non-cancerous tissue from areas of cancerous tissue;identifying a location of a bulk of the tumor or a nidus;predicting areas of likely cancer-cell migration lying at or beyond an outermost portion of the bulk of the tumor or the nidus and along a white matter pathway based at least in part on the location of the bulk of the tumor or the nidus and patient population data corresponding to regions of likely or demonstrated cancer-cell migration in a patient population;imaging the brain to determine whether cancer cells have already migrated to the predicted areas;tracing a trajectory along a tissue region that is separated from the outermost portion of the nidus by a sufficient distance to ensure that cancer cells have not spread beyond that distance, the trajectory being spaced apart from the bulk of the tumor or the nidus by a minimum separation; andtreating a likely healthy region of the white matter pathway beyond the spreading cancer cells using a stereotactic radiosurgery surgical apparatus to prevent migration of cancer cells. 3. The method of claim 2, wherein distinguishing areas of non-cancerous tissue from areas of cancerous tissue comprises using Fluid-Attenuated Inversion-Recovery. 4. The method of claim 2, wherein establishing areas of noncancerous tissue from areas of cancerous tissue comprises using magnetic resonance spectroscopy. 5. The method of claim 2, wherein distinguishing areas of non-cancerous tissue from areas of cancerous tissue comprises using Fluid-Attenuated Inversion-Recovery and establishing areas ... | Christopher Duma | A61B005/05 A61B005/055 G01R033/465 A61N005/10
US08473027 | B2 | 2013-06-25 | Process for draping breast MRI imaging coils | 2009-02-18 | A draping provides at least partial and preferably complete coverage of MRI coils, an MRI table and an MRI support system thereby protecting a patient and the equipment. Draping notch fold down flaps may be pushed into shaped apertures of an MRI apparatus by the patient's breast due to the design of the draping. This offers a barrier between the patient and the equipment pad providing the patient with warmth, comfort and protection from infectious diseases such as Methicillin-resistant staphylococcus aureus (MRSA). The draping reduces the clean up and improves the procedure turnaround time for the technologist. | 1. A method for reducing contact between a patient and a portion of a Magnetic Resonance Imaging (MRI) apparatus, the MRI apparatus including a patient interface surface and an aperture extending through the patient interface surface, the method comprising: positioning a barrier over at least a portion of the patient interface surface, the barrier comprising an opening;arranging a periphery of the barrier to encompass a periphery of the patient interface surface;securing the barrier to the MRI apparatus;locating the opening over the aperture of the MRI apparatus;inserting a patient breast through the barrier opening and into the aperture of the MRI apparatus; andmanipulating a plurality of flaps of the barrier surrounding the opening to be within the aperture of the MRI apparatus and adjacent a sidewall and/or coil of the MRI apparatus, wherein the plurality of flaps are interposed between a) the patient breast, and b) the sidewall and/or coil of the MRI apparatus, and wherein at least a portion of the manipulating step occurs via the inserting step. 2. The method of claim 1, wherein the locating step further comprises arranging a periphery of the opening to be encompassed by a periphery of the aperture. 3. The method of claim 1, wherein the locating step further comprises: aligning the barrier opening and the aperture of the MRI apparatus. 4. The method of claim 3, wherein the aligning step further comprises arranging a center axis of each of the opening and the aperture of the MRI apparatus to be collinear with each other. 5. The method of claim 1, wherein the securing step further comprises securing the barrier to the patient interface surface. 6. The method of claim 5, wherein the securing step further comprises securing the barrier to a ledge associated with the patient interface surface. 7. The method of claim 1, wherein the securing step further comprises: manipulating at least one securing member interconnected to the barrier to detachably connect the barrier to the periphery of the MRI apparatus to restrict relative movement between the barrier and the MRI apparatus. 8. The method of claim 7, wherein the at least one securing member is comprises at least one of: elastic, hook and loop, string, tape, snaps, ribbons, sticky glue, buttons, ties and tab hooks. 9. A method for reducing contact between a patient and a portion of a Magnetic Resonance Imaging (MRI) apparatus, wherein the MRI apparatus includes a patient interface surface and an aperture extending through the patient interface surface, wherein the aperture is at least partially surrounded by a sidewall and/or coil of the MRI apparatus, and wherein the method comprises: positioning a barrier over at least a portion of the patient interface surface, the barrier comprising an opening inside a periphery thereof, wherein the positioning step comprises arranging a periphery of the barrier to encompass a periphery of the MRI apparatus;locating the opening over the aperture of th... | QSUM Biopsy Disposables LLC | A61B005/055
US08560057 | B2 | 2013-10-15 | State machine interface system | 2007-05-28 | State machine interface system, heart state machine analyzer and/or stimulator, comprising state machine algorithms and a graphical user interface, adapted to receive signals from a sensor device that are related to physiological activities of the heart and/or the circulatory system of a living being and the state machine algorithms determine states of heart cycles based upon the signals. The determined heart cycle states are graphically presented at the graphical user interface such that the temporal relation between the different states are illustrated. The graphical user interface may be circular diagrams or bar graphs including parts representing the temporal relation between the different states. | 1. A method for graphical representation of time-related heart cycle state diagrams based on measured heart functions and heart pumping models, the method comprising: measuring one or more heart functions of a patient, the measuring selected from one or more of ultrasound imaging, electrocardiogram, magnetic resonance imaging, spinning computer aided tomography, strain rate imaging, x-ray imaging, gamma radiation imaging, pulse plethysmography, oximetry, tissue velocity imaging, pulse measuring, flow measuring, pressure measuring, or volume measuring;analyzing the measured heart functions using a dynamic displacement pumping model of a heart using one or more processing devices, the pumping model determining heart cycle states by simulating heart pumping, and by identifying durations of heart cycles, from the measured heart functions; anddisplaying a graphical model of the heart cycle states on a graphical user interface such that a temporal relation between different heart cycle states is depicted in a single state diagram. 2. The method for graphical representation of heart cycle state diagrams based on measured heart functions and heart pumping models according to claim 1, further comprising measuring velocity changes of a heart and using the measured velocity changes as triggering points correlating to hydraulic or mechanical heart activity. 3. The method for graphical representation of heart cycle state diagrams based on measured heart functions and heart pumping models according to claim 1, wherein the different heart cycle states are graphically presented as one or several overlapping diagrams presenting different activities of the heart at one or several locations arranged as heart cycle state diagrams in which heart cycle states are represented as segments of a circle having lengths corresponding to a duration of heart cycle state. 4. The method for graphical representation of heart cycle state diagrams based on measured heart functions and heart pumping models according to claim 3, wherein the different activities of the heart are represented by three or more rings in which a first ring represents a state diagram from a left ventricle, a second ring represents a state diagram from a right ventricle and a third ring positioned between the first and second rings represents a state diagram of a ventricular septum. 5. The method for graphical representation of heart cycle state diagrams based on measured heart functions and heart pumping models according to claim 3, wherein the heart cycle states are graphically presented in one or more bar graphs divided into parts, each part representing different states and wherein a size of each part correlates to a duration of a respective heart cycle state. 6. The method for graphical representation of heart cycle state diagrams based on measured heart functions and heart pumping models according to claim 5, wherein a first bar graph represents states of a right heart half and a second bar graph ... | Gripping Heart AB | A61B005/044
US08911379 | B2 | 2014-12-16 | Method and system for non-invasive measurement of cardiac parameters | 2007-09-04 | A method and system are presented for use in assessment of at least one cardiac parameter of an individual. An electrodes arrangement is applied to an individual's body, for applying an electrical field to the body and providing an electrical output indicative of a systolic impedance change and of a velocity of said change during a cardiac cycle. Also provided is additional data indicative of at least of the following conditions of the individual: a value of total peripheral resistance (TPR), a value of cardiac index (CI), and existence of the AHF condition. Data corresponding to these condition is analyzed to determine whether the TPR satisfies a first predetermined condition and/or the CI satisfies a predetermined second condition and/or whether the acute heart failure (AHF) condition is identified, to thereby use the data indicative of the measured electrical output and selectively calculate said at least one cardiac parameter based on either the systolic impedance change data or on said data of the velocity of the impedance change. | 1. A computer implemented method for use in assessment of at least one cardiac parameter of an individual, the method comprising: receiving, by a processor, measured data provided by a measuring unit and being indicative of a systolic impedance change, ΔR, and of a velocity of said change, δR/δt, of an individual during a cardiac cycle;providing input data indicative of at least one condition of the individual, said at least one condition being indicative of whether the individual is sustained by an acute heart failure (AHF);analyzing, by a processor, said input data indicative of said at least one condition of the individual and utilizing, by the processor, the measured data for calculating the at least one cardiac parameter of the individual, wherein said utilizing of the measured data comprises:upon identifying that said input data is indicative of that the individual is sustained by AHF, using the data of the velocity of impedance change in the calculation, andupon identifying that said input data is indicative of that the individual is not sustained by AHF, either using the data of systolic impedance change in the calculation or analyzing additional input data indicative of at least one other predetermined condition of the individual in order to determine whether to usethe systolic impedance change data or the velocity of said change, said at least one other condition comprising at least one of a total peripheral resistance (TPR) value and a cardiac index (CI) value; andgenerating an output indicative of the calculation result. 2. The method of claim 1, wherein said at least one cardiac parameter is selected from the following: stroke volume, cardiac output, cardiac index, stroke index, heart rate. 3. The method of claim 1, wherein said TPR data is obtained by operating an electrodes arrangement, located on the individual's body and adapted to apply an electrical field to the body and provide an electrical output, to obtain first measured data of the systolic impedance change; providing second measured data indicative of at least an individual's blood pressure, utilizing the first and second measured data to determine the individual's TPR. 4. The method of claim 1, wherein said CI is determined as CI=(ΔR/R)/BSA, where BSA is the individual's body surface area, and R is measured basal resistance. 5. The method according to claim 1, wherein said at least one other predetermined condition is satisfied when the measured TPR value is higher than a value ranging from 1800 dynes/sec−5 to 1900 dynsec−5. 6. The method of claim 1, wherein said at least one other predetermined condition is satisfied when the cardiac index is above 2.5. 7. The method of claim 1, wherein the calculation of said at least one cardiac parameter based solely on the systolic impedance change data is performed using the Frinerman formula:SV=Hctcorr.K⁡(sex,age)·Δ⁢⁢RR·Hcorr.2R·α+ββ·Kel·Kw·IBwherein: Hctcorr is the correcting factor depending from Hematocrit; Hct is the Hem... | New N.I. Medical (2011) Ltd. | A61B005/02 A61B005/053 A61B005/029
US08417463 | B2 | 2013-04-09 | Systems and methods for monitoring pulmonary edema dynamics | 2009-07-20 | This document discusses, among other things, systems and methods for measuring the dynamics of pulmonary congestion in heart failure subjects over time to monitor the subjects susceptibility to pulmonary edema, including sensing and receiving information indicative of a bodily pressure and information indicative of pulmonary fluid, and using the transient responses of these measurements to compute parameters related to the dynamics of thoracic fluid accumulation, such as a critical pressure (Pc), a critical time (Tc), or a filtration index (Kfi). | 1. A system for use with a physiologic pressure sensor, the system comprising: a processor comprising: a pressure sensor input configured to receive information indicative of a bodily pressure from the physiologic pressure sensor; andan impedance sensor input configured to receive impedance information indicative of a pulmonary fluid level;wherein the processor is configured to generate a pressure baseline using the information indicative of the bodily pressure received from the physiologic pressure sensor;wherein the processor is configured to generate a pulmonary fluid level baseline using the impedance information indicative of the pulmonary fluid level;wherein the processor is configured to identify a critical pressure (Pc) using the information indicative of the bodily pressure received from the physiologic pressure sensor and the received impedance information indicative of the pulmonary fluid level information, wherein the critical pressure (Pc) is the pressure at which the received impendence falls below the pulmonary fluid level baseline; andwherein the processor is configured to compute a critical time (Tc) using the information indicative of the bodily pressure received from the physiologic pressure sensor and the received impedance information indicative of the pulmonary fluid level, wherein the critical time (Tc) defines a duration beginning when the information indicative of the bodily pressure deviates from the pressure baseline and ending when the pressure reaches the critical pressure (Pc). 2. The system of claim 1, wherein the processor comprises an event detector input configured to receive information about a detected pulmonary fluid overload event, and wherein the processor is configured to identify the critical pressure (Pc) using the information indicative of the bodily pressure received from the physiologic pressure sensor leading up to or during the detected event and the received impedance information indicative of the pulmonary fluid level leading up to or during the detected event. 3. The system of claim 1, wherein the processor comprises an event detector input configured to receive information about a detected pulmonary fluid overload event, and wherein the processor is configured to compute the critical time (Tc) using the information indicative of a bodily pressure received from the physiologic pressure sensor leading up to or during the detected event and the received impedance information indicative of the pulmonary fluid level leading up to or during the detected event. 4. The system of claim 1, wherein the processor is configured to compute an estimated pressure slop (Pes) using the information indicative of the bodily pressure received from the physiologic pressure sensor, and wherein the processor is configured to compute an estimated critical time (Tec) using a current pressure determined using the information indicative of the bodily pressure received from the physiologic pressure sensor, the critical pr... | Cardiac Pacemakers, Inc. | A61B005/00 A61B005/08 A61B005/02 A61B005/04 A61N001/00 G01N033/48 G01N033/50
US08521270 | B2 | 2013-08-27 | Quantitative EEG method to identify individuals at risk for adverse antidepressant effects | 2007-06-05 | Methods, apparatus, and systems are provided for efficiently and accurately identifying individuals at risk for adverse effects from psychotropic or CNS-active treatment. Changes in a brain activity indicator (e.g. EEG cordance) are used to predict the adverse effects of treatment based on an experimentally derived cutoff value. For example, a reliable biological indicator is provided with high predictive capability for identifying, very early in the course of treatment (e.g. | 1. A method of identifying subjects at risk for an adverse effect from treatment with a serotonin reuptake inhibitor, the method comprising: at a computer system comprising a processor, memory, and at least one program stored in the memory and executable by the processor:receiving a first set of physical values that are obtained by measuring one or more physical properties of a subject's brain activity at a first time, wherein the first time is before the subject begins treatment with the serotonin reuptake inhibitor;calculating a first resultant value of a function using the first set of physical values as input values, wherein the function is midline and right frontal cordance (MRFC);receiving a second set of physical values that are obtained by measuring one or more physical properties of a subject's brain activity at a second time, wherein the second time is after the subject begins treatment with the serotonin reuptake inhibitor;calculating a second resultant value of the function using the second set of physical values as input values;determining a change in the function from the first time to the second time by calculating a difference value between the first and second resultant values; andbased on the difference value, determining whether the subject is at risk for a future adverse effect from treatment with the serotonin reuptake inhibitor, wherein the subject's brain activity is measured using an electroencephalograph. 2. The method of claim 1, wherein the subject's brain activity is measured using electroencephalography. 3. The method of claim 2, wherein the one or more physical properties of the subject's brain activity comprise voltages measured at electrodes fp2, af2, f4, f8, fz, and fpz, placed according to the extended international 10-20 system. 4. The method of claim 3, wherein the one or more physical properties of the subject's brain activity consist of voltages measured at electrodes fp2, af2, f4, f8, fz, and fpz, placed according to the extended international 10-20 system. 5. The method of claim 1, wherein the second time is less than or equal to about 48 hours after treatment begins with the serotonin reuptake inhibitor. 6. The method of claim 1, wherein the physical properties of the brain activity consist of theta waves. 7. The method of claim 1, wherein the adverse effect comprises at least one of depressed mood, anxiety, hostility, and suicidal ideation. 8. The method of claim 1, wherein determining whether the subject is at risk for a future adverse effect from treatment with the serotonin reuptake inhibitor includes determining a percentage or range of percentages for a likelihood of a future adverse effect. 9. The method of claim 1, wherein determining whether the subject is at risk for a future adverse effect from treatment with the serotonin reuptake inhibitor includes comparing the difference value to a cutoff value. 10. The method of claim 9, wherein the cutoff value is determined by: acquiring data of ... | The Regents of the University of California | A61B005/04
US08679027 | B2 | 2014-03-25 | Systems and methods for pulse processing | 2008-09-30 | According to embodiments, techniques for using continuous wavelet transforms to process pulses from a photoplethysmographic (PPG) signal are disclosed. The continuous wavelet transform of the PPG signal may be used to identify and characterize features and their periodicities within a signal. Regions, phases and amplitudes within the scalogram associated with these features may then be analyzed to identify, locate, and characterize a true pulse within the PPG signal. Having characterized and located the pulse in the PPG (possibly also using information gained from conventional pulse processing techniques such as, for example, by identifying turning points for candidate pulse maxima and minima on the PPG, frequency peak picking for candidate scales of pulses, etc.), the PPG may be parameterized for ease of future processing. | 1. A method for processing pulses from a PPG signal, comprising: receiving a PPG signal;performing a continuous wavelet transform of the PPG signal to produce a wavelet transformed signal;generating a scalogram from the transformed signal;identifying regions in the scalogram corresponding to desired components of the PPG signal;identifying pulse locations in the PPG signal based on the identified regions; andreplacing at least one pulse in at least one of the identified pulse locations in the PPG signal with a pulse selected from the group comprising a triangular pulse, a square pulse, and a sinusoid pulse. 2. The method of claim 1, further comprising modifying the received PPG signal to produce a processed PPG signal. 3. The method of claim 2, wherein modifying the received PPG signal to produce the processed PPG signal comprises filtering the received PPG signal. 4. The method of claim 2, further comprising performing a continuous wavelet transform of the modified PPG signal to produce a modified wavelet transformed signal and wherein generating the scalogram comprises generating a combined scalogram from the wavelet transformed signal and the modified wavelet transformed signal. 5. The method of claim 1, wherein identifying regions in the scalogram comprises identifying a plurality of pulse locations in the scalogram. 6. The method of claim 5, wherein identifying a plurality of pulse locations in the scalogram comprises determining localized high energy regions at high scales within the scalogram. 7. A system for processing pulses from a PPG signal, comprising: a wavelet processor capable of performing a continuous wavelet transform of the PPG signal to produce a wavelet transformed signal and generating a scalogram from the transformed signal; anda pulse processor capable of identifying regions in the scalogram corresponding to desired components of the PPG signal, identifying pulse locations in the PPG signal based on the identified regions, and replacing at least one pulse in at least one of the identified pulse locations in the PPG signal with a pulse selected from the group comprising a triangular pulse, a square pulse, and a sinusoid pulse. 8. The system of claim 7, further comprising a signal filter capable of modifying the received PPG signal to produce a processed PPG signal. 9. The system of claim 8, wherein the wavelet processor is further capable of performing a continuous wavelet transform of the modified PPG signal to produce a modified wavelet transformed signal and generating a combined scalogram from the wavelet transformed signal and the modified wavelet transformed signal. 10. The system of claim 7, wherein the pulse processor is capable of identifying a plurality of pulse locations in the scalogram. 11. The system of claim 10, wherein the pulse processor is capable of determining localized high energy regions at high scales within the scalogram. 12. Non-transitory computer-readable medium for use in processing pu... | Nellcor Puritan Bennett Ireland | A61B005/02
US20100030050 | A1 | 2010-02-04 | Subcutaneous Glucose Electrode | 2009-08-31 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably three or four-layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer is overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited (third layer). An outer (fourth) layer is biocompatible. | 1. An electrochemical sensor comprising: one or more non-corroding metal or carbon electrodes;a sensing layer comprising an enzyme coupled to each electrode; anda biocompatible layer comprising a biocompatible hydrogel chemically bound to the sensing layer of each electrode. | Abbott Diabetes Care Inc. | A61B005/1473
US20100030044 | A1 | 2010-02-04 | Subcutaneous Glucose Electrode | 2009-08-31 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably three or four-layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer is overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited (third layer). An outer (fourth) layer is biocompatible. | 1. An electrochemical sensor comprising: one or more non-corroding metal or carbon electrodes;a sensing layer comprising an enzyme coupled to each electrode; anda biocompatible layer comprising a biocompatible hydrogel chemically bound to the sensing layer of each electrode. | Abbott Diabetes Care Inc. | A61B005/1473
US20100022862 | A1 | 2010-01-28 | Analyte Sensor with Insertion Monitor, and Methods | 2009-09-30 | A sensor, and methods of making, for determining the concentration of an analyte, such as glucose or lactate, in a biological fluid such as blood or serum, using techniques such as coulometry, amperometry, and potentiometry. The sensor includes a working electrode and a counter electrode, and can include an insertion monitoring trace to determine correct positioning of the sensor in a connector. | 1.-19. (canceled) 20. A sensor for determining a concentration of an analyte, the sensor comprising: a first substrate having a first major surface, a second major surface opposing the first major surface, a proximal end, and a distal end that mates with an analyte meter;a second substrate having a first major surface, a second major surface opposing the first major surface, a proximal end, and a distal end that mates with an analyte meter, wherein the first and second substrate being disposed so that the first major surface of the first substrate is in facing relationship with the first major surface of the second substrate;a working electrode disposed on the first major surface of the first substrate and a conductive trace extending from the proximal end to the distal end of the first substrate; anda counter electrode disposed on the first major surface of the second substrate and conductive trace extending from the proximal end to the distal end of the second substrate;wherein the distal end of the first substrate comprises a cut-out to expose a portion of the conductive trace of the counter electrode on the first major surface of the second substrate, andwherein the distal end of the second substrate comprises a cut-out to expose a portion of the conductive trace of the working electrode on the first major surface of the first substrate. 21. The sensor of claim 20, further comprising a spacer between the first substrate and the second substrate, wherein the spacer material, the first substrate and the second substrate define a sample chamber. 22. The sensor of claim 21 wherein the sample chamber has a volume of no more than about 1 μL. 23. The sensor of claim 21, wherein the sample chamber has a volume of no more than about 0.5 μL. 24. The sensor of claim 21, further comprising an insertion monitor on one of the first and the second major surfaces of one of the first substrate and the second substrate. 25. The sensor of claim 24, wherein the insertion monitor comprises a conductive stripe extending across a width of the sensor strip. 26. The meter of claim 24, wherein said insertion monitor is a conductive dot. 27. The meter of claim 24, wherein said insertion monitor is a conductive grid pattern. 28. The meter of claim 27, wherein said insertion monitor is a conductive stylistic feature. 29. The sensor of claim 24, wherein the insertion monitor comprises a conductive stripe extending across the sensor, the conductive stripe having a predetermined resistance corresponding to calibration information and the conductive stripe is arranged and configured to be read at only two contact points to provide the predetermined resistance, wherein the insertion monitor provides calibration information about the sensor and indicates proper insertion of a sensor into a meter. 30. The sensor of claim 29, wherein the predetermined resistance is one of a plurality of predetermined resistances, and wherein the calibration information is encoded base... | - | A61B005/1468
US20100016744 | A1 | 2010-01-21 | FETAL HEART RATE MONITORING SYSTEM | 2007-04-13 | A fetal heart rate monitoring system includes an ultrasonic transducer for transmitting a signal to a patient and for receiving a reflected heart beat monitoring signal. A monitoring system coupled to the ultrasonic transducer processes the heart beat monitoring signal and produces heart beat information representative of the monitored heart rate. A mechanical or electronic steering system coupled to the ultrasonic transducer controls the location of the signal transmitted to the patent and/or the location of the received heart beat monitoring signal. A control system coupled to the monitor and steering system controls the steering system as a function of the heart beat monitoring signal to optimize the quality of the heart beat information in the received heart beat monitoring signals. | 1. A heart rate monitoring system, including: a transducer for transmitting and/or receiving heart beat monitoring signals;a steering system for controlling the location of the transmitted and/or received heart beat monitoring signals;a monitoring system coupled to the transducer for processing the received heart beat monitoring signals and for producing heart beat information representative of a monitored heart beat;a control unit coupled to the monitoring system and the steering system, for processing the received heart beat monitoring signals and controlling the steering system to optimize quality of the heart beat monitoring signals. 2. The heart rate monitoring system of claim 1 wherein the control unit processes the received heart beat monitoring signals to optimize the quality of the heart beat information in the heart beat monitoring signals. 3. The heart rate monitoring system of claim 1 wherein the steering system includes a mechanical steering system. 4. The heart rate monitoring system of claim 3 wherein the steering system includes a mechanical steering system connected to the transducer. 5. The heart rate monitoring system of claim 1 wherein the steering system includes an electronic steering system. 6. The heart rate monitoring system of claim 5 wherein: the transducer includes an array of transducer elements; andthe electronic steering system selects the transducer elements. 7. The heart rate monitoring system of claim 6 wherein the electronic steering system includes a multiplexer coupled between the control unit, the monitoring system and the array of transducer elements. 8. The heart rate monitoring system of claim 1 wherein the control unit is operable in a search mode and a tracking mode. 9. The heart rate monitoring system of claim 1 wherein the control unit includes a search algorithm and a steering controller. 10. The heart rate monitoring system of claim 1 wherein the control unit is operable in a manual override mode. 11. The heart rate monitoring system of claim 1 wherein: the monitoring system produces a plurality of signals representative the monitored heart beat; andthe control unit controls the steering system as a function of the plurality of signals. 12. The heart rate monitoring system of claim 11 wherein the plurality of signals produced by the monitoring system include an audio signal and a quality level signal. 13. A fetal heart rate monitoring system, including: an ultrasonic transducer for transmitting a signal to a patient and for receiving a reflected heart beat monitoring signal;a monitoring system coupled to the ultrasonic transducer, for processing the heart beat monitoring signal and producing heart beat information representative of the monitored heart rate;a steering system coupled to the ultrasonic transducer, for controlling the location of signal transmitted to the patent and/or the location of the received heart beat monitoring signal with respect to the patient; anda control system c... | Mayo Foundation for Medical Education and Research | A61B005/0444
US20100022851 | A1 | 2010-01-28 | MEASUREMENT PROCESSING APPARATUS OF PHYSIOLOGY SIGNALS | 2009-06-03 | A measurement processing apparatus is disclosed for processing plurality of physiology signals. The measurement processing apparatus includes a multiplex device, an analog-to-digital conversion module and a control signal generator. The multiplex device generates a multiplex signal composed of the physiology signals with output sequence controlled by a first control signal. Furthermore, the multiplex device dispenses each physiology signal with one corresponding multiplex density based on the first control signal. The analog-to-digital conversion module converts the multiplex signal into a digital multiplex signal based on at least one adjustable bias voltage under control of a second control signal. The control signal generator generates the first control signal based on the feature values of the physiology signals. Also, the control signal generator generates the second control signal based on the voltage swing ranges of the physiology signals. The second control signal is synchronized with the first control signal. | 1. A measurement processing apparatus of physiology signals, comprising: a multiplex device for receiving a plurality of physiology signals and outputting a multiplex signal, the multiplex device dispensing each of the physiology signals with a corresponding multiplex density according to a first control signal; andan analog-to-digital conversion module, electrically coupled to the multiplex device for receiving the multiplex signal;wherein the first control signal is generated based on feature values of the physiology signals. 2. The measurement processing apparatus of claim 1, further comprising: a plurality of physiology signal detection devices, electrically coupled to the multiplex device, each of the physiology signal detection devices performing a corresponding physiology detection process so as to generate a corresponding physiology signal of the physiology signals. 3. The measurement processing apparatus of claim 2, further comprising: a plurality of pre-amplifiers, each of the pre-amplifier being electrically coupled between the multiplex device and a corresponding physiology signal detection device of the physiology signal detection devices, for performing a signal amplification operation on a corresponding physiology signal of the physiology signals. 4. The measurement processing apparatus of claim 1, wherein the analog-to-digital conversion module is utilized for converting the multiplex signal into a digital multiplex signal based on at least one bias voltage adjusted according to a second control signal in synchronization with the first control signal; wherein the second control signal is generated based on a plurality of voltage swing ranges of the physiology signals. 5. The measurement processing apparatus of claim 4, wherein the analog-to-digital conversion module comprises: a first bias voltage providing unit for providing a plurality of first bias voltages;a first bias voltage selector, electrically coupled to the first bias voltage providing unit, for selecting a corresponding first bias voltage out of the first bias voltages according to the second control signal; andan analog-to-digital converter, electrically coupled to the multiplex device and the first bias voltage selector for receiving the multiplex signal and the corresponding first bias voltage respectively, for converting the multiplex signal into the digital multiplex signal based on the corresponding first bias voltage. 6. The measurement processing apparatus of claim 5, wherein the analog-to-digital conversion module further comprises: a second bias voltage providing unit for providing a plurality of second bias voltages; anda second bias voltage selector, electrically coupled between the second bias voltage providing unit and the analog-to-digital converter, for selecting a corresponding second bias voltage out of the second bias voltages according to the second control signal, the corresponding second bias voltage being forwarded to the analog-to-digital c... | - | A61B005/00 H03M001/12
US20100049006 | A1 | 2010-02-25 | MEDICAL SIGNAL PROCESSING SYSTEM WITH DISTRIBUTED WIRELESS SENSORS | 2007-02-23 | A wireless medical signal processing system for health monitoring is disclosed which achieves high wireless link reliability/security, low power dissipation, compactness, low cost and supports a variety of sensors for various physiological parameters. The system includes a medical signal processor which communicates with a wireless distributed sensor system as its peripheral for detecting physiological parameters of the person and for providing signals indicative thereof. The medical signal processor wirelessly receives the signals from the distributed wireless sensor system in a multiplexed fashion and processes the signals to provide an indication of the health of the person. The indication of health could relate to a disease state, general health or fitness level of a person. The system also includes a mobile device for receiving the indication of the health of the person to allow for a diagnosis or treatment of the person. | 1.-65. (canceled) 66. A system for wireless healthcare monitoring comprising: (a) a physiological signal monitoring patch comprising a sensor interface that receives signals from sensors coupled to the sensor interface, a signal processor coupled to the sensor interface, a storage element coupled to the signal processor, a radio coupled to the storage element, an antenna coupled to the radio, power dissipation management circuits coupled to the processor, and a battery that provides power to the patch; and(b) a medical signal processor comprising an antenna that receives data from the patch, a radio coupled to the antenna, a storage element coupled to the radio, a processor coupled to the storage element, power dissipation management circuits, and a bus interface that communicates with a device. 67. The system of claim 66 wherein the medical signal processor is incorporated into the device. 68. The system of claim 66 wherein the medical signal processor is a stand-alone unit that is connected to the device. 69. The system of claim 67 wherein the device is a mobile device, smart phone, personal digital assistant, or medical device. 70. The system of claim 66 wherein the patch comprises a custom ASIC wherein the custom ASIC comprises the sensor interface. 71. The system of claim 66 wherein the processor in the patch performs compression to reduce data rate, performs encoding to achieve high reliability, and/or manages buffering to vary the duty cycle of the radio in the patch. 72. The system of claim 66 wherein the medical signal processor decodes the signal received from the patch. 73. The system of claim 66 wherein the medical signal processor authenticates, tests, and/or controls the functionality of one or more patches in the system. 74. The system of claim 66 wherein the medical signal processor dictates the power dissipation of the patches. 75. The system of claim 66 wherein the medical signal processor sends control signals to alter wireless link performance by changing parameters relating to radio functions by instructing the signal processor and/or the radio on the patch. 76. The system of claim 66 wherein the medical signal processor instructs patches to utilize a particular radio or communication scheme. 77. The system of claim 66 wherein the system further comprises a secure server with which the device communicates. 78. The system of claim 77 wherein the secure server links the sensor data with other health information. 79. The system of claim 66 wherein the physiological signals are monitored continuously, periodically, or episodically. 80. A system for wireless healthcare monitoring comprising: (a) a physiological signal monitoring patch comprising a sensor interface that receives signals from sensors coupled to the patch, a signal processor coupled to the sensor interface, a storage element coupled to the signal processor, a radio coupled to the storage element, an antenna coupled to the radio, power dissipation manag... | - | A61B005/00 A61B005/0402
US07641621 | B2 | 2010-01-05 | Elongated intra-lumenal medical device | 2003-08-25 | A medical device includes a coil having a longitudinal axis and a radial axis orthogonal to the longitudinal axis, formed from a wire. The wire includes a cross-section with a centroid, a moment of inertia with respect to an axis running through the centroid and parallel to the longitudinal axis of the coil, and a moment of inertia with respect to an axis running through the centroid and parallel to the radial axis of the coil. The moment of inertia with respect to an axis running through the centroid and parallel to the longitudinal axis of the coil is greater than the moment of inertia with respect to an axis running through the centroid and parallel to the radial axis of the coil. | 1. A medical device comprising: a coil having a longitudinal axis and a radial axis orthogonal to the longitudinal axis, formed from a wire, the wire comprising:(a) a cross-section with a centroid;(b) a moment of inertia with respect to an axis running through the centroid and parallel to the longitudinal axis of the coil; and(c) a moment of inertia with respect to an axis running through the centroid and parallel to the radial axis of the coil, wherein the moment of inertia with respect to an axis running through the centroid and parallel to the longitudinal axis of the coil is greater than the moment of inertia with respect to an axis running through the centroid and parallel to the radial axis of the coil. 2. The medical device according to claim 1, wherein the wire cross-section is a polygonal shape. 3. The medical device according to claim 1, wherein the wire cross-section is a ellipse shape. 4. The medical device according to claim 1, wherein the wire cross-section is an I-Beam shape. 5. The medical device according to claim 1, wherein the wire is formed of a material with a Poisson's ratio from 0.25 to 0.5. 6. The medical device according to claim 1, wherein the coil further comprises a second coil having windings disposed between windings of the coil. 7. The medical device according to claim 1, wherein the wire is a composite wire comprising: (a) a cross-section with a centroid, a wire longitudinal axis parallel to the coil longitudinal axis and a wire radial axis parallel to the coil radial axis;(b) a first material having a first Young's Modulus at the centroid; and(c) a second material having a second Young's Modulus further away from the centroid along the wire radial axis; wherein the second Young's Modulus is greater than the first Young's Modulus. 8. The medical device according to claim 7, wherein the wire cross-section is a circular shape. 9. The medical device according to claim 7, wherein the wire cross-section is a polygonal shape. 10. A medical guidewire comprising: (a) an elongated shaft including a proximal region having a first outer diameter and a distal region having a second outer diameter that is smaller than the first outer diameter;(b) a coil member connected to the elongated shall at the distal end of the proximal region and extending from the distal end of the proximal region over at least a portion of the distal region, the coil member having an inner diameter that is greater than the second outer diameter, wherein the coil has a longitudinal axis and a radial axis orthogonal to the longitudinal axis, formed from a wire, the wire comprising: (i) a cross-section with a centroid;(ii) a moment of inertia with respect to an axis running through the centroid and parallel to the longitudinal axis of the coil; and(iii) a moment of inertia with respect to an axis running through the centroid and parallel to the radial axis of the coil, wherein the moment of inertia with respect to an axis running through the cent... | Boston Scientific Scimed, Inc. | A61B005/00 A61M025/00
US07641630 | B2 | 2010-01-05 | Drainage catheter hub with locking cam | 2005-03-11 | The drainage catheter hub includes a locking cam for securing a suture to maintain an anchor configuration of the distal end of the drainage catheter. The locking cam is rotatable to a first position for securing the suture and a second position for releasing the suture. The locking cam includes a tapered surface which includes a release portion and a securement portion. The release portion allows movement of the suture while the securement portion cooperatively engages the suture in connection with the seat to minimize movement of the suture. | 1. A drainage catheter for use in providing a fluid pathway for draining bodily fluid from a cavity or tissue of a patient, the drainage catheter being configured such that a suture utilized in connection with the drainage catheter selectively secures the distal end of the catheter in an anchor configuration, the drainage catheter comprising: a catheter tube having a proximal end and a distal end, the distal end of the catheter tube being configured to be positioned in tissue, cavity, or other location within the body of the patient having an amount of fluid to be drained;a suture running the length of at least a portion of the catheter tube and being secured to the distal end of the catheter tube to selectively maintain the anchor configuration of the distal end of the catheter;a catheter hub coupled to the proximal end of the catheter tube, the catheter hub comprising: a seat having an inner contact surface and adapted to be in contact with a portion of the suture;a locking cam adapted to be positioned in the seat, the locking cam having an outside diameter adapted to conform to the inner contact surface of the seat to provide rotational movement of the locking cam relative to the seat such that the outside diameter of the locking cam remains in contact with the seat around substantially the entire outside diameter of the locking cam; the locking cam having a tapered camming surface positioned on a portion of the outside diameter of the locking cam such that the tapered camming surface is utilized to cooperatively engage a portion of the suture inside the catheter hub along a substantial length of the inner contact surface of the seat to selectively secure the suture based on the rotational position of the locking cam, wherein the tapered camming surface comprises;a release portion, the release portion being positioned at a given radial position on an outer surface of the locking cam, the release portion extending inwardly from the outside diameter of the locking cam such that the release portion has a predetermined amount of displacement from the inner contact surface of the seat, wherein when the locking cam is rotated to a first rotational position the portion of the suture positioned inside the hub is positioned between the release portion and the inner contact surface of the seat, the release portion providing a sufficient amount of clearance between the release portion and the inner contact surface of the seat to allow the suture to be moved by the user relative to the locking cam and the seat; anda securement portion, the securement portion being positioned at a given radial position on the outer surface of the barrel, the securement portion having a lesser amount of displacement from the inner contact surface of the seat than the release portion, such that when the locking cam is rotated to a second position, the portion of the suture positioned inside the hub is cooperatively engaged between the securement portion of the camming surfac... | Merit Medical Systems, Inc. | A61M031/00 A61M037/00 A61M005/178 A61B005/00
US20100010375 | A1 | 2010-01-14 | DEVICE AND METHOD FOR ANALYZING BODY FLUIDS | 2009-06-06 | A device for analyzing body fluids comprising a lancing element that can puncture a body part along a lancing axis. The lancing element has a collection volume for collecting body fluid obtained by a puncture and a receiving element that can be loaded with body fluid from the collection volume and that can detect a component of the body fluid. The lancing element and the receiving element are arranged in the lancing axis so that they can be moved relative to one another such that the receiving element dips into the collection volume in the lancing direction during a transfer phase following the puncture. | 1. A device for analyzing body fluids which is especially useful for blood sugar tests, comprising a lancing element that moves during a lancing phase along a lancing axis to puncture a body part, the lancing element having a collection volume for collecting body fluid obtained by a puncture and a receiving element that can be loaded with body fluid from the collection volume, the receiving element detecting a component of the body fluid during use of the device, wherein the lancing element and the receiving element are moveable relative to one another along the lancing axis and the receiving element is arranged in an initial position away from the collection volume such that the receiving element does not contact the body part during the puncture, and wherein the receiving element enters the collection volume during a transfer phase following the lancing phase. 2. The device of claim 1, wherein the lancing element has a distal lancing member, the distance between the lancing member and the receiving element being smaller in the transfer phase than in the lancing phase. 3. The device of claim 1, further comprising a lancing drive coupled to the lancing element, the receiving element being movable into the collection volume by a retracting movement of the lancing drive. 4. The device of claim 1, further comprising a feed drive which moves the receiving element into the collection volume. 5. The device of claim 1, wherein the lancing element or the receiving element have a docking structure for a detachable drive. 6. The device of claim 1, further comprising a linear guide in which the lancing element is mounted. 7. The device of claim 1, wherein the lancing element comprises a recess leading to the collection volume, the receiving element engaging the recess. 8. The device of claim 1, wherein the lancing element and the receiving element together comprise a disposable part. 9. The device of claim 8, further comprising a magazine having a plurality of lancing elements and associated receiving elements. 10. The device of claim 9, wherein the receiving element comprises a measuring member for an optical or electrochemical measurement in the collected body fluid. 11. The device of claim 10, wherein the measuring member is coated with a test reagent that reacts with the component of the body fluid while under liquid contact. 12. The device of claim 10, wherein the receiving element is connectable to a measuring unit via a light guide routed in the lancing element. 13. The device of claim 1, wherein the collection volume comprises a slot or channel of the lancing element. 14. The device of claim 1, wherein the collection volume has a capacity of less than 10 nl body fluid. 15. The device of claim 1, wherein, during use of the device, an aliquot of the body fluid is transferred from the collection volume onto a distal front side of the receiving element, the aliquot being less than 5 nl. 16. The device of claim 1, further comprising a lanc... | - | A61B005/151 G01N033/50 A61L002/08
US20100041967 | A1 | 2010-02-18 | ELECTROPHYSIOLOGICAL INTUITION INDICATOR | 2009-08-18 | Systems and methods for electrophysiological detection and measurement of intuition are disclosed. In one embodiment, one or more electrophysiological properties of one or more individuals are monitored and used as an indication of a future event. In one embodiment, the electrophysiological property may include heart rate variability, brain wave activity, respiration pattern, skin conductance level, etc. In another embodiment, a signal averaging technique is used to generate a waveform that may be used as an indicator of future events. | 1-52. (canceled) 53. A method for providing predictive outputs comprising: generating a first classified response curve for the subject based on changes measured in one or more electrophysiological properties of the subject during each of a plurality of trials in which the subject provides a first response type to a stimulus, wherein the first response type corresponds to a randomly determined actual outcome of the plurality of trials;generating a second classified response curve for the subject based on changes measured in one or more electrophysiological properties of the subject during each of the plurality of trials in which the subject provides a second response type to the stimulus, wherein the second response type does not correspond to the randomly determined actual outcome of the plurality of trials;acquiring a newly-acquired response curve for the subject based on measured changes in the one or more electrophysiological properties of the subject during a period of time, prior to an event of interest, in which the subject makes a choice associated with the event of interest;comparing the newly-acquired response curve for the subject to each of the first classified response curve and the second classified response curve; andproviding a predictive output to the subject corresponding to the event of interest based on a result of said comparing. 54. The method of claim 53, wherein the stimulus is a question that has two possible answers, and the first response type corresponds to one of the two possible answers, and the second response type corresponds to the other of the two possible answers. 55. The method of claim 53, wherein the measured changes in the one or more electrophysiological properties of the subject occur, at least partially, during a period of time between when the stimulus is presented to the subject and when the subject responds to the stimulus. 56. The method of claim 53, wherein the result of said comparing corresponds to a confidence level of how closely the newly-acquired response curve matches either of the first classified response curve and the second classified response curve. 57. The method of claim 53, further comprising determining a confidence level based on said comparing, wherein said confidence level corresponding to how closely the newly-acquired response curve matches either of the first classified response curve or the second classified response curve. 58. The method of claim 53, wherein said one or more electrophysiological properties include at least one of heart rate variability, brain wave activity, skin conductance level, respiration pattern, blood pulse amplitude and blood pulse volume. 59. The method of claim 53, wherein generating the first classified response curve comprises generating a first combined classified response curve for a plurality of subjects based on measured changes in one or more electrophysiological properties of the plurality of subjects during the plurality of trials in w... | Quantum Intech, Inc. | A61B005/00
US07653431 | B2 | 2010-01-26 | Arrhythmia discrimination based on determination of rate dependency | 2005-12-20 | Cardiac systems and methods provide for discriminating between supraventricular tachyarrhythmia and ventricular tachyarrhythmia based on a determination that the patient's supraventricular rhythm exhibits rate dependency. One approach involves determining if a patient's supraventricular rhythm exhibits rate dependent morphology. If the patient's supraventricular rhythm is determined to exhibit rate dependent morphology, an implantable device classifies a detected tachyarrhythmia episode based on one or more templates selected from a plurality of rate-indexed templates stored in the device. Determining if the supraventricular rhythm exhibits rate dependent morphology may also include determining one or more rates at which the rate dependent morphology occurs. | 1. A method of discriminating between supraventricular and ventricular tachyarrhythmia, comprising: obtaining a plurality of cardiac templates from supraventricular cardiac waveforms and associating the plurality of cardiac templates with a plurality of cardiac rates to provide a plurality of rate-indexed templates;detecting a tachyarrhythmia episode of a patient;determining an episode cardiac rate for the detected tachyarrhythmia episode;identifying a closest cardiac rate from the plurality of cardiac rates that is closest to the episode cardiac rate;initially selecting from the plurality of rate-indexed templates an initial template associated with the closest cardiac rate;classifying the tachyarrhythmia episode as supraventricular tachyarrhythmia or ventricular tachyarrhythmia based on one or more templates selected from the plurality of rate-indexed templates, staffing with the initial template;creating supraventricular tachyarrhythmia in the patient by atrial pacing at multiple pacing rates;sensing cardiac waveforms during the created supraventricular tachyarrhythmia; anddetermining if the patient's supraventricular rhythm exhibits rate dependent morphology based on morphologies of the cardiac waveforms. 2. The method of claim 1, wherein the atrial pacing is delivered at progressively higher rates. 3. The method of claim 1, wherein the atrial pacing is delivered in accordance with an automated atrial pacing pattern. 4. A method of discriminating between supraventricular and ventricular tachyarrhythmia, comprising: obtaining a plurality of cardiac templates from supraventricular cardiac waveforms and associating the plurality of cardiac templates with a plurality of cardiac rates to provide a plurality of rate-indexed templates;detecting a tachyarrhythmia episode of a patient;determining an episode cardiac rate for the detected tachyarrhythmia episode;identifying a closest cardiac rate from the plurality of cardiac rates that is closest to the episode cardiac rate;initially selecting from the plurality of rate-indexed templates an initial template associated with the closest cardiac rate;classifying the tachyarrhythmia episode as supraventricular tachyarrhythmia or ventricular tachyarrhythmia based on one or more templates selected from the plurality of rate-indexed templates, staffing with the initial template;creating supraventricular tachyarrhythmia by altering patient activity under controlled conditions;sensing cardiac waveforms during the created supraventricular tachyarrhythmia; anddetermining if the supraventricular rhythm exhibits rate dependent morphology based on the morphologies of the sensed cardiac waveforms. 5. A method of discriminating between supraventricular and ventricular tachyarrhythmia, comprising: obtaining a plurality of cardiac templates from supraventricular cardiac waveforms and associating the plurality of cardiac templates with a plurality of cardiac rates to provide a plurality of rate-indexed templates;dete... | Cardiac Pacemakers, Inc. | A61B005/04
US07645234 | B2 | 2010-01-12 | Diagnostic and intervention tools for emergency medical dispatch | 2007-06-13 | A system and method assists an emergency medical dispatcher in responding to emergency calls. A computer implemented emergency medical dispatch protocol includes interrogatories for a dispatcher to ask a caller to generate an appropriate response. A diagnostic tool is provided to determine a vital sign of a patient based on a timer and caller relayed information about the patient. An intervention tool is provided to administer assistance and determine a compression rate based on a timer and caller relayed information. | 1. A computer implemented method performed on a dispatch center computer to assist a dispatcher when responding to an emergency caller regarding a medical emergency of a patient, the method comprising: a dispatch center computer providing an emergency medical dispatch protocol to assist emergency dispatchers, the protocol including a plurality of interrogatories for a dispatcher to ask a caller to generate an emergency medical dispatch response by emergency responders;initiating a diagnostic tool on the dispatch center computer to determine a vital sign of a patient;the diagnostic tool providing a user interface on the dispatch center computer, the user interface configured to enable the dispatcher to initiate a timer indicative of a start time for a caller;in response to the timer initiation, the diagnostic tool running the timer;the diagnostic tool receiving vital sign related data which is input by the dispatcher corresponding to caller relayed information;the diagnostic tool stopping the timer;the diagnostic tool determining a rate or interval of the vital sign related data; andthe diagnostic tool determining a patient diagnosis from the vital sign related data and the determined rate or interval of the vital sign. 2. The computer implemented method of claim 1, further comprising: calculating a determinant value to prioritize an emergency medical response by emergency responders based on the patient diagnosis. 3. The computer implemented method of claim 1, further comprising: the diagnostic tool displaying a recommendation on the user interface based on the patient diagnosis. 4. The computer implemented method of claim 1, further comprising: the user interface providing a close input; andthe diagnostic tool terminating operation in response to the close input. 5. The computer implemented method of claim 1, further comprising: the user interface providing a clear input; andthe diagnostic tool resetting the timer and clearing the patient diagnosis in response to the clear input. 6. The computer implemented method of claim 1, wherein the vital sign is a breathing interval and the diagnostic tool measures the breathing interval in relation to the timer. 7. The computer implemented method of claim 6, further comprising: the user interface providing an input for each patient breath after starting the timer. 8. The computer implemented method of claim 7, further comprising: the diagnostic tool measuring intervals between each patient breath to determine a breathing interval. 9. The computer implemented method of claim 8, further comprising: the diagnostic tool displaying the duration of the intervals on the user interface. 10. The computer implemented method of claim 8, further comprising: the diagnostic tool displaying the breathing interval on the user interface in relation to time. 11. The computer implemented method of claim 8, further comprising: the diagnostic tool determining a pattern based on the breathing interval and displaying ... | - | A61B005/00
US07642338 | B2 | 2010-01-05 | Tacrolimus standard and methods of using same | 2006-06-20 | A composition and kit useful as a tacrolimus standard solution for immunoassays, and methods for making and using same. The composition and kits include a known amount of tacrolimus or a derivative thereof, and a non-specific protein capable of forming a complex with the tacrolimus or derivative thereof. The standard solution may be used to generate calibration curves for an immunoassay or to check the precision of an analytical instrument. | 1. A method for use in determining a concentration of tacrolimus or derivative thereof in a sample, said method comprising: providing a standard solution including a known amount of tacrolimus or a derivative thereof and a non-specific protein capable of forming a water soluble complex with the tacrolimus or derivative thereof, the non-specific protein being at least 95% fatty acid free and the ratio of non-specific protein to tacrolimus being in the range of 25,000:1 to 1000:1 on a weight basis wherein the non-specific protein is albumin or α1-acid glycoprotein; and performing an assay on the standard solution and on the sample to determine a concentration of tacrolimus or derivative thereof in the sample. 2. The method of claim 1, wherein the standard solution further includes an emulsifier. 3. The method of claim 1, wherein the standard solution further includes a hemolysate. 4. The method of claim 1, wherein the solution comprises fatty acid free bovine serum albumin. 5. The method of claim 1, wherein standard solution is free of alkanol solvents. 6. The method of claim 1 wherein the ratio of non-specific protein to tacrolimus is at least 2,000:1 on a weight basis. 7. The method of claim 1 wherein the ratio of non-specific protein to tacrolimus is at least 5,000:1 on a weight basis. 8. The method of claim 1 wherein the ratio of non-specific protein to tacrolimus is at least 10,000:1 on a weight basis. 9. A method for use in determining a concentration of tacrolimus or derivative thereof in a sample, said method comprising: providing a standard solution including a known amount of tacrolimus or a derivative thereof and a non-specific protein capable of forming a water soluble complex with the tacrolimus or derivative thereof, the non-specific protein being at least 95% fatty acid free and the ratio of non-specific protein to tacrolimus being at least 2000:1 on a weight basis wherein the non-specific protein is albumin or α1-acid glycoprotein; and performing an assay on the standard solution and on the sample to determine a concentration of tacrolimus or derivative thereof in the sample. 10. The method of claim 9, wherein the standard solution further includes an emulsifier. 11. The method of claim 9, wherein the solution comprises fatty acid free bovine serum albumin. 12. The method of claim 9, wherein standard solution is free of alkanol solvents. | Siemens Healthcare Diagnostics Inc. | A61K038/12 G01N033/566 A61B005/00
US07662362 | B2 | 2010-02-16 | Nanoparticulate probe for in vivo monitoring of tissue oxygenation | 2004-09-07 | A new class of micro- and nano-particulate paramagnetic spin probes especially useful for magnetic resonance imaging techniques, including electron paramagnetic resonance (EPR) and magnetic resonance imaging (MRI). The probes are lithium phthalocyanine derivative compounds. Also provided are suspensions and emulsions comprising lithium phthalocyanine derivative probes. Also provided are noninvasive methods for measuring noninvasive methods of measuring oxygen concentration, oxygen partial pressure, oxygen metabolism, and nitric oxide concentration in a specific tissue, organ, or cell in vivo or in vitro. | 1. A particulate probe comprising a lithium phthalocyanine derivative of formula 4 or a radical thereof:wherein R is selected from the group consisting of O(CH2)nCH3, S(CH2)nCH3, O(CH2)nCH2OH, O(CH2)nCH2NH2, O(CH2)nCH2SH, and combinations thereof; and n is from 1 to 6. 2. The particulate probe of claim 1, wherein the particulate probe is a nanoparticulate probe or a microparticulate probe. 3. The particulate probe of claim 2 wherein the particulate probe has a size of up to 10 microns. 4. The particulate probe of claim 3 wherein the particulate probe has a size of from 0.22 microns to 10 microns. 5. The particulate probe of claim 3 wherein the particulate probe has a size of less than 0.22 microns. 6. The particulate probe of claim 1 wherein the particulate probe has been derivatized to be a magnetic resonance imaging (MRI) probe. 7. The particulate probe of claim 1 wherein the particulate probe has been derivatized to be an electron spin resonance (ESR) probe. 8. The particulate probe of claim 1 wherein the particulate probe has been derivatized to be an electron paramagnetic resonance (EPR) probe. 9. The particulate probe of claim 1 wherein the particulate probe has been derivatized to be an electron paramagnetic resonance imaging (EPRI) probe. 10. The particulate probe of claim 1 wherein the particulate probe has been derivatized to be a proton electron double resonance imaging (PEDRI) probe. 11. A method of measuring oxygen concentration, oxygen partial pressure, or oxygen metabolism in a specific tissue or organ in a subject, the method comprising the steps of: (a) administering a particulate probe according to claim 1 to the subject; and(b) applying a magnetic resonance (MR) spectroscopy or imaging technique for measuring O2 concentration in tissues or organs of the subject. 12. The method of claim 11, wherein the MR technique is selected from the group consisting of MRI, ESR, EPR, ERPI, and PEDRI. 13. The method of claim 11 wherein the particulate probe is delivered to the subject intravenously, as a suspension or emulsion. 14. The method of claim 11 wherein the particulate probe is implanted in the tissue or organ of interest. 15. The method of claim 11 wherein the phthalocyanine radical particulate probe is attached to a peptide or glycoconjugate that has specific affinity for cell surface markers, wherein the particulate probe acts as a cell migration marker. 16. The method of claim 11 wherein the particulate probe is attached to an antibody, wherein the antibody has an affinity to cell surface proteins that lead as markers of cell migration, cell division, and cell death. 17. The method of claim 11 wherein the particulate probe remains in the subject for at least 180 days. 18. The method of claim 11 wherein the particulate probe is internalized in live cells, for the study of intracellular oxygenation, cellular hypoxia, cellular hyperoxia, cell division, cellular migration, or metastatis. 19. The method of claim 11 wh... | The Ohio State University Research Foundation | A61B005/055 C07B047/00 C07D487/22
US07668585 | B2 | 2010-02-23 | Respiration monitor for computed tomography | 2003-12-19 | A medical diagnostic imaging system (10) includes a diagnostic imaging scanner (12) that acquires imaging data of a medical imaging patient. A reconstruction processor (46) reconstructs at least a portion of the acquired imaging data into an image representation. A pair of electrodes (30, 32) contact a thoracic region of the patient. An electrical meter (34) measures electrical resistance R(t) or another time-varying electrical parameter (70) across the electrode pair (30, 32) during the acquiring of imaging data. A respiration monitor (36) extracts a time-varying respiration characteristic (90, 98, 110, 120) from the measured time-varying electrical parameter (70) to indicate respiratory cycle phase. | 1. A diagnostic imaging system including: a diagnostic imaging scanner that acquires imaging data of a subject in an examination region;a reconstruction processor that reconstructs the acquired imaging data into an image representation;a pair of electrodes adapted to externally contact a thoracic region of the subject;an electrical meter that measures a time-varying electrical parameter across the electrode pair by applying a voltage or current pulse train having a frequency substantially higher than the heart rate across the pair of electrodes;a monitor that extracts a time-varying respiration characteristic from the measured time-varying electrical parameter, the monitor including a differentiator that computes a time derivative of the time-varying electrical parameter; andan imaging controller that receives the respiration characteristic and controls the diagnostic imaging scanner based thereon. 2. The imaging system as set forth in claim 1, wherein the time-varying electrical parameter includes a time-varying resistance, the differentiator computes a first derivative, and the monitor further includes: a respiration state processor that computes the respiration parameter as one of: inhaling corresponding to a positive time derivative of the time varying resistance,exhaling corresponding to a negative time derivative of the time varying resistance, anda breath-hold state corresponding to a substantially zero time derivative of the time-varying resistance. 3. A diagnostic imaging system including: a diagnostic imaging scanner that acquires imaging data of a subject in an examination region;a reconstruction processor that reconstructs the acquired imaging data into an image representation;a pair of electrodes adapted to externally contact a thoracic region of the subject;an electrical meter that measures a time-varying electrical parameter across the electrode pair by applying a voltage or current pulse train having a frequency substantially higher than the heart rate across the pair of electrodes;a monitor that extracts a time-varying respiration characteristic from the measured time-varying electrical parameter, the monitor including a respiratory cycle phase processor that estimates a respiratory cycle phase based on the time varying electrical parameter; andan imaging controller that receives the respiration characteristic and controls the diagnostic imaging scanner based thereon. 4. The imaging system as set forth in claim 3, wherein the time-varying electrical parameter is selected from a group consisting of: a time-varying complex impedance,a time-varying resistance,a time-varying capacitance,time-varying inductance,a time varying current, anda time varying voltage. 5. The imaging system as set forth in claim 3, wherein the diagnostic imaging scanner is a computed tomography scanner. 6. The imaging system as set forth in claim 3, wherein the electrical meter includes: a voltage pulse generator that applies a voltage pulse train to the... | Koninklijke Philips Electronics N.V. | A61B005/05 A61B006/00
US07657135 | B2 | 2010-02-02 | Waveguide and optical motion sensor using optical power modulation | 2008-11-19 | An optical motion sensing device included a sensor frame defining an opening, a sensor pad disposed in the opening, an optical sensing system adapted to detect an amount of movement of the sensor pad in the sensor frame, and an output unit. The optical sensing system includes an optical waveguide, an optical source device, and an optical detector. The optical waveguide is positioned within the sensor frame such that the movement of the sensor pad results in the flexing or compressing of the optical waveguide. The optical source device supplies optical energy to the optical waveguide. The optical detector detects an amount of optical energy exiting the optical waveguide. The output unit is configured to receive a signal indicative of the amount of optical energy exiting the optical waveguide and to generate a measure of the amount of movement of the sensor pad from the received signal. | 1. An optical motion sensing device comprising: a sensor frame defining a cavity and an opening to said cavity;a sensor pad disposed in the opening, the sensor pad adapted to allow for its placement against an anatomical location of a subject;an optical sensing system adapted to detect an amount of downward displacement of the sensor pad relative to the sensor frame, the optical sensing system comprising:(a) an optical waveguide positioned within the cavity of the sensor frame such that the downward displacement of the sensor pad relative to the sensor frame results in the flexing or compressing of the optical waveguide,(b) an optical source device to supply optical energy to a first end of the optical waveguide, and(c) an optical detector to detect an amount of optical energy exiting a second end of the optical waveguide; andan output unit configured to receive a signal indicative of the amount of optical energy exiting the optical waveguide and to generate a measure of the amount of downward displacement of the sensor pad relative to the sensor frame from the received signal. 2. The optical motion sensing device of claim 1, wherein the optical waveguide comprises a siloxane elastomer. 3. The optical motion sensing device of claim 1, wherein the optical waveguide comprises an elastomer selected from the group consisting of polysiloxane, polyurethane, polybutadine rubber, and combinations thereof. 4. The optical motion sensing device of claim 1, further comprising a waveguide support structure comprising a non-compliant surface, the waveguide support supporting at least a portion of the optical waveguide. 5. The optical motion sensing device of claim 1, further comprising a flexible and incompressible support surface that supports the optical waveguide over substantially all of its length. 6. The optical motion sensing device of claim 1, wherein the optical detector is optically coupled to optical waveguide such that the optical detector receives substantially all of the optical energy from the optical source that does not escape from the sides of the sides of the optical waveguide. 7. A method of detecting an amount of localized displacement, the method comprising: transmitting optical energy into a first end of an optical waveguide, the optical waveguide positioned within a sensor frame including a sensor pad such that the downward displacement of the sensor pad relative to the sensor frame results in the compression or flexing of the optical waveguide, wherein the sensor pad is placed against an anatomical location of a subject;detecting, using an optical detector held by the sensor frame, an amount of optical energy exiting a second end of the optical waveguide and generating therefrom a signal indicative of optical energy received, the amount of optical energy exiting the second end of the optical waveguide being reduced in response to the downward displacement of the sensor pad relative to the sensor frame; andgenerating a measure of ... | Tarilian Laser Technologies, Limited | G02B006/00 G01J001/56 A61B005/02
US07645242 | B1 | 2010-01-12 | Composite guidewire with drawn and filled tube construction | 1999-12-22 | The present invention is directed to an intracorporeal device, preferably a guidewire, and method for making the device. The guidewire of the present invention is formed, at least in part, of a composite elongate core formed, at least in part, of precipitation hardened material. The elongate core members of the present invention will have an ultimate tensile strength and modulus of elasticity greater than the same for an identically dimensioned elongate member formed from superelastic NITINOL alone. | 1. A heat-treated elongate member, comprising: a composite elongate core;the composite elongate core including an inner core formed of a precipitation hardened material and a layer formed of a superelastic material; anda second layer concentrically arranged about the layer formed of the superelastic material;wherein the layer is arranged concentrically about the inner core. 2. The elongate member of claim 1 wherein the composite elongate core has a modulus of elasticity of at least 9×106 psi. 3. The elongate member of claim 2 wherein the modulus of elasticity is at least 12×106 psi. 4. The elongate member of claim 3 wherein the modulus of elasticity is at least 15×106 psi. 5. The elongate member of claim 1 wherein the composite elongate core has an ultimate tensile strength of at least 150 ksi. 6. The elongate member of claim 5 wherein the ultimate tensile strength is at least 180 ksi. 7. The elongate member of claim 6 wherein the ultimate tensile strength is at least 200 ksi. 8. The elongate member of claim 1 wherein the precipitation hardenable material is precipitation hardenable stainless steel. 9. The elongate member of claim 8 wherein the precipitation hardenable material is chromium-nickel based single stage martensitic precipitation hardenable stainless steel. 10. The elongate member of claim 8 wherein the precipitation hardenable stainless steel is essentially nickel free. 11. The elongate member of claim 8 wherein the precipitation hardenable stainless steel includes less than about 1% nickel. 12. The heat-treated elongate member of claim 1, wherein the second layer is formed of the precipitation hardened material. 13. A heat-treated elongate member, comprising: a composite elongate core;the composite elongate core including an inner core formed of a precipitation hardened material concentrically surrounded by a first layer formed from superelastic NITINOL and having a proximal section and distal section, the inner core and the first layer being independently formed;a flexible body distinct from the first layer at least partially overlying the distal section; andwherein the precipitation hardenable material comprises at least two materials selected from the group consisting of nickel, cobalt, molybdenum, chromium, tungsten, and iron and the composite elongate core further includes a second layer disposed at least in part concentrically about the first layer and formed from a material similar to the inner core material. 14. A heat-treated elongate member, comprising: a composite elongate core;the composite elongate core including an inner core formed of a precipitation hardened material concentrically surrounded by a first layer formed of a superelastic material and having a proximal section and distal section;a flexible body distinct from the first layer at least partially overlying the distal section; andwherein the precipitation hardenable material comprises at least two materials selected from the group consisting of nic... | Advanced Cardiovascular Systems, Inc. | A61B005/00
US07646199 | B2 | 2010-01-12 | Integrated system of MRI RF loop coils plus spacing fixtures with biocontainment uses | 2007-08-28 | When scanning a patient to generate an image thereof, radio frequency (RF) coil modules are scalably coupled to each other using a plurality of clips to form flat or polygonal coil arrays that are placed on or around the patient or a portion thereof. A user assesses the volume to be imaged, identifies a coil array configuration of suitable size and shape and employs clips of one or more pre-determined angles to construct the identified coil array configuration, which is placed on or about the volume. Coil modules are coupled to a preamplifier interface box (PIB), which provides preamplified coil signal(s) to a patient imaging device, such as an MRI scanner. Small arrays are constructible to accommodate pediatric patients and/or smaller animals. Modules are hermetically sealed, can be sanitized between uses, and discarded at end-of-life. In one aspect, the modular coil array, clips, and PIB are maintained in an isolated contamination zone, separate from the patient imaging device. | 1. A multiply configurable modular coil array system, compatible with magnetic resonance imaging of a subject including: a plurality of radio frequency (RF) receive coil modules;a plurality of rigid clips that releasably connect and retain the coil modules in a final shape in order to form custom geometries and permit dynamic shaping of differently configured modular coil arrays; andan interface box that is connected to each of the plurality of coil modules, which amplifies signals received from each of the coil modules, and outputs the amplified signals into a patient imaging device. 2. The system according to claim 1, wherein each coil module includes an RF receive coil and circuitry communicating with the interface box. 3. The system according to claim 2, wherein the patient imaging device is a magnetic resonance imaging (MRI) device, which executes algorithms performing sensitivity encoding (SENSE) acceleration and constant level appearance (CLEAR) uniformity corrections on signals received from the interface box. 4. The system according to claim 1, wherein the rigid clips are reusable, and include at least two 90° clips, 108° clips, 120° clips, 135° clips, or 180° clips, such that the reusable rigid clips are use to reconfigure the coil modules into an array of differently configured module arrays. 5. The system according to claim 4, wherein the clips are color-coded in order to indicate their respective angles. 6. The system according to claim 4, wherein the clips connect the modules in order to form a polygonal coil array. 7. The system according to claim 4, wherein the coil modules are selectively connected: using 90° clips in order to form all or a portion of a square coil array;using 108° clips in order to form all or a portion of a pentagonal coil array;using 120° clips in order to form all or a portion of a hexagonal coil array;using 135° clips in order to form all or a portion of an octagonal coil array; orusing 180° clips in order to form a flat coil array. 8. The system according to claim 1, wherein each of the coil modules are hermetically sealed. 9. The system according to claim 1, wherein the interface box includes a plurality of ports connecting to individual coil modules, and wherein a plurality of interface boxes are stackable in order to accommodate modular coil arrays having more coil modules than a single interface box alone can accommodate. 10. A magnetic resonance imaging system including the modular coil array system according to claim 1, configured and positioned in order to receive spatially encoded magnetic resonance RF signals. 11. A method of configuring the multiply configurable modular coil array system, of claim 1, including: selecting a size of a volume to be imaged;selecting a coil array configuration in order to fit the volume being imaged;connecting a plurality of coil modules using the rigid clips in order to achieve a selected modular coil array configuration;placing the modular coil array on or ... | Koninklijke Philips Electronics N.V. | G01V003/00 A61B005/055
US08015732 | B2 | 2011-09-13 | Athletic or other performance sensing systems | 2009-10-26 | Articles of footwear and other devices include modules, e.g., for sensing physical and/or physiological characteristics associated with use of the footwear or other devices. Such devices include: (a) a housing, optionally an asymmetrical housing, for receiving an electronic module; (b) release mechanisms for removably mounting the module in the housing; and/or (c) activation/authentication systems for selectively activating the module. Additionally, methods of making articles of footwear or other structures that include such housings, electronic modules, release systems, activation systems, authentication systems, etc. are described. Performance sensing systems and methods of using the above devices for sensing performance characteristics also are described. | 1. An article of footwear, comprising: an upper member at least partially defining a foot-receiving chamber;a sole structure engaged with the upper member; anda housing for releasably receiving an electronic module, the housing being located between a midsole member and an outsole member of the sole structure thereby forming a portion of the sole structure, wherein the housing defines a receptacle, wherein an electronic module receiving chamber of the receptacle is asymmetrical in at least one respect, and wherein the housing is further located between impact-attenuating column structures included as part of the sole structure. 2. An article of footwear according to claim 1, wherein the housing forms a portion of the midsole member of the sole structure. 3. An article of footwear according to claim 1, wherein the housing is formed in a heel portion of the sole structure. 4. An article of footwear according to claim 1, wherein the electronic module receiving chamber is accessible from outside of the foot-receiving chamber. 5. An article of footwear according to claim 1, further comprising: an electronic module releasably engaged at least partially within the housing. 6. An article of footwear according to claim 5, wherein the electronic module provides data relating to at least one physical or physiological parameter associated with use of the article of footwear. 7. An article of footwear according to claim 6, wherein the electronic module provides data relating to at least one of location, movement speed, or movement distance associated with use of the article of footwear. 8. An article of footwear according to claim 6, wherein the electronic module is a speed and distance sensor. 9. An article of footwear according to claim 5, further comprising: a system for securing the electronic module with the housing. 10. An article of footwear according to claim 9, wherein the system includes a cover element engaged with at least one of an impact-attenuating member and the housing, and wherein the cover element extends across the receptacle. 11. An article of footwear according to claim 9, wherein the system includes a retaining device engaged with at least one of an impact-attenuating member and the housing that extends at least partially across the receptacle. 12. An article of footwear according to claim 9, wherein the system includes a fastener for holding the module in place with respect to the housing. 13. An article of footwear according to claim 9, wherein at least a portion of the system directly engages the electronic module. 14. An article of footwear according to claim 9, wherein the system includes a member that extends into a recess, opening, or groove defined in the electronic module. 15. An article of footwear according to claim 1, further comprising: an antenna system engaged with at least one of the upper member or the sole structure. 16. An article of footwear according to claim 15, further comprising: an electronic modu... | NIKE, Inc. | A43B023/00
US08628525 | B2 | 2014-01-14 | Wireless interventional device and a system for wireless energy transmission | 2007-12-17 | An interventional device (12) is configured to be positioned in a body and includes an electrically operable unit (E1, E2) configured to carry out an interaction with the body upon a receipt of electric power. The device further includes a sensor (2) configured for wirelessly receiving electromagnetic energy from a remote source. The sensor is configured as a resonant circuit (2a, 2b) which converts the received electromagnetic energy into the electric power. The electrically operable device may include a diagnostic and/or therapeutic module. | 1. An MR compatible interventional system comprising: an interventional catheter configured to be inserted in a human body located in a body coil of a magnetic resonance imager, the interventional catheter including: an electrically powered unit which carries out an interaction with an interior of the body,a resonant circuit which wirelessly receives RF electromagnetic energy, and converts the received RF electromagnetic energy into electric power,a remote unit remote from the interventional catheter and connected by only resistive wires with the interventional catheter, the resistive wires carrying feedback signals indicative of an amount of the electric power wirelessly delivered to the electrically powered unit by the resonant circuit, the remote unit including: an RF transmitter which generates the RF electromagnetic energy received by the resonant circuit, anda controller which controls the RF transmitter to control an amount of power in the generated RF electromagnetic energy in accordance with the feedback signals. 2. The MR compatible interventional system according to claim 1, wherein the resonant circuit is disposed at a tip of the interventional catheter. 3. The MR compatible interventional system according to claim 1, wherein the electrically powered unit includes at least one of a therapeutic module and a diagnostic module connected to receive the electric power from the resonant circuit. 4. The MR compatible interventional system according to claim 1, wherein the resonant circuit is tuned to a Larmor frequeney of the magnetic resonance imager in which the interventional catheter is positioned. 5. The MR compatible interventional system according claim 1, wherein the resonant circuit is tuned to a frequency, sufficiently above a Larmor frequency of the magnetic resonance imager in which the interventional catheter is positioned such that a power pick up compared to the Larmor frequency is reduced by at least 30 dB. 6. The MR compatible interventional system according to claim 5, wherein the electrically powered unit comprises; an inductive element which produces a local magnetic field upon being energized by said electric power. 7. The MR compatible interventional system according to claim 1, wherein the resonant circuit comprises: a plurality of coils being arranged in a mutually non-planar configuration; andwherein the RF transmitter includes a transmit coil arranged in a non-planar configuration. 8. The MR compatible interventional system according to claim 1, wherein the resonant circuit is disposed at a tip of the interventional catheter such that the electric power is supplied via the received RF electromagnetic energy without wires running along the catheter to supply the electric power. 9. An interventional system comprising: an interventional catheter including; a receive coil disposed at a tip of the interventional catheter and tuned to a preselected frequency such that the receive coil receives RF signals and conve... | Koninklijke Philips N.V. | A61B018/18 A61B019/00 A61B005/05 A61B005/00
US08028450 | B2 | 2011-10-04 | Recipient verification systems and methods of use including recipient identification | 2009-07-31 | A recipient verification system including a band and a label strip. The band includes a base, a trailing shield segment, and a strap. The band defines a band identification portion displaying a band identifier, as well as opposing passages. The trailing shield segment is disposed over the identification portion. The strap extends from the base and terminates at a tail end. The label strip extends from the base and includes at least one removable label displaying the predetermined band identifier. In an initial state, the strap is free of the band identification portion. In the worn state, the strap is wrapped about a wearer's appendage, with the tail end passing through at least one of the passages and a section of the strap maintained along the identification portion. The predetermined band identifier displayed on the identification portion is visible through the trailing shield segment in the worn state. | 1. A recipient verification system comprising: a band including: a base defining: a band identification portion displaying a predetermined band identifier, the band identification portion forming opposing first and second passages,a recipient information portion adjacent the band identification portion for receiving recipient-specific information,a trailing shield segment disposed over the band identification portion,a leading shield segment disposed over the recipient information portion, the leading shield segment having a first end secured to the base and a second end initially free of the base and subsequently bondable to the base for protecting recipient information applied to the recipient information portion,a strap extending from the base and terminating at a tail end; anda label strip extending from the base and including a first removable label displaying the predetermined band identifier;wherein the recipient verification system is transitionable from an initial state in which the strap is free of the band identification portion to a worn state in which the strap is wrapped about a wearer's appendage, the tail end is passed through at least one of the passages, and a section of the strap is maintained along the band identification portion, and further wherein the predetermined band identifier on the band identification portion is visible through the trailing shield segment in the worn state. 2. The recipient verification system of claim 1, wherein the trailing shield segment protects the predetermined band identifier on the base from the environment in the worn state. 3. The recipient verification system of claim 1, wherein the trailing shield segment defines opposing first and second ends, and further wherein the first end of the trailing shield segment is secured to the base and the second end of the trailing shield segment is free of the base in the initial state. 4. The recipient verification system of claim 3, wherein the system is configured such that the second end of the trailing shield segment is bonded to the base in the worn state. 5. The recipient verification system of claim 3 configured such that transitioning from the initial state to the worn state includes passing the tail end through both of the passages and laminating the strap section between the band identification portion and the trailing shield segment. 6. The recipient verification system of claim 1, wherein the worn state includes the strap section maintained above the band identification portion and the band identifier on the band identification portion being visible through the trailing shield segment and the strap section. 7. The recipient verification system of claim 1, wherein the predetermined band identifier displayed on the base is printed on a permanent label contiguously formed as part of the label strip. 8. The recipient verification system of claim 7, wherein the label strip, including the permanent label, is completely formed apart from the ... | Typenex Medical, LLC | A44C005/00
US07867185 | B2 | 2011-01-11 | Graduated compression device for the treatment of circulatory disorders | 2008-07-01 | A method of applying therapeutic compression to a patient's body limb, by: (a) measuring the dimensions of a patient's body limb; (b) determining a desired compression profile for the patient's body limb; and (c) forming a compression device comprising a plurality of tension bands dimensioned to be wrapped around the patient's limb, wherein different tension bands have different widths such that applying an equal tension to the bands results in different compressions being applied along the length of the patient's limb such that the desired compression profile is applied to the patient's limb. | 1. A method of fabricating a compression device to apply therapeutic compression to a particular patient's body limb, comprising: (a) measuring the dimensions of the particular patient's body limb at a plurality of different lengths along the body limb;(b) determining a desired compression profile for the particular patient's body limb at the plurality of different lengths along the body limb; and(c) manufacturing a compression device comprising a plurality of tension bands dimensioned to be wrapped around the patient's limb, wherein different tension bands have different widths such that applying an equal tension to the bands results in different compressions being applied along the length of the particular patient's limb such that the desired compression profile is applied to the particular patient's limb, wherein the compression device is manufactured for the particular patient after steps (a) and (b) have been completed. 2. The method of claim 1, wherein measuring the dimensions of a patient's body limb comprises determining the circumference of the body limb at different lengths along the body limb. 3. The method of claim 1, wherein determining a desired compression profile for the patient's body limb comprises determining a graduated compression profile that increases along a portion of the patient's body limb. 4. The method of claim 1, wherein determining a desired compression profile for the patient's body limb comprises determining a graduated compression profile that decreases along a portion of the patient's body limb. 5. The method of claim 1, wherein determining a desired compression profile comprises determining a compression profile that both increases and decreases along a portion of the patient's body limb. 6. The method of claim 1, wherein manufacturing the tension bands of the compression device with increasing bandwidth along the length of the limb decreases the compression along the length of the limb. 7. The method of claim 1, wherein manufacturing the tension bands of the compression device with decreasing bandwidth along the length of the limb increases the compression along the length of the limb. 8. The method of claim 1, wherein determining a desired compression profile for the patient's body limb comprises determining desired compression levels at various distances along the length of the limb and thereby determining the widths of the bands. 9. The method of claim 1, wherein the bands are designed with different widths based upon the desired compression profile. 10. The method of claim 1, further comprising: (d) wrapping the tension bands of the compression device around the patient's limb; and(e) applying equal tension to the bands, such that the different bands apply different pressures along the length of the limb, thereby applying the desired compression profile to the patient's limb. 11. The method of claim 1, further comprising: (d) re-measuring the dimensions of the particular patient's body limb at a ... | CircAid Medical Products | A61L015/00
US08708922 | B2 | 2014-04-29 | Electrically isolated catheter with wireless sensors | 2007-12-19 | A local power-delivery/data-reception unit is installed within an insertion end of a sealed catheter. The local power-delivery/data-reception unit wirelessly powers a separately sealed sensor that is attached to the insertion end and configured for wirelessly sending a data signal to the local power-delivery/data-reception unit. The catheter may further feature a remote power-delivery/data-reception unit disposed within the handle and configured for wirelessly communicating with the local power-delivery/data-reception unit and a controller for controlling the sensor. | 1. A catheter, comprising: a sealed catheter body including a handle and an insertion end for inserting into a patient, anda sealed sensor adjoined to said insertion end, the sensor and catheter body being sealed from each other, anda local power-delivery/data-reception unit incorporated at the insertion end and configured for wirelessly emitting a radio signal that powers the sensor, said sensor being configured for wirelessly sending a data-containing radio signal to the local power-delivery/data-reception unit. 2. The catheter of claim 1, further comprising an integrated circuit incorporating the sensor such that the sealing seals the integrated circuit, the integrated circuit being flexible. 3. The catheter of claim 2, wherein said end includes an end of an electrode, the integrated circuit being rolled around the electrode end. 4. The catheter of claim 1, wherein said unit is electrically insulated from the sealed sensor. 5. The catheter of claim 1, further comprising one cable disposed longitudinally within the catheter body to extend between said unit and a controller for controlling the sensor, said cable being a coaxial cable. 6. The catheter of claim 1, further comprising a remote power-delivery/data-reception unit disposed within the handle and configured for wirelessly communicating with the local power-delivery/data-reception unit and a controller for controlling the sensor. 7. The catheter of claim 1, further comprising an integrated circuit incorporating the sensor and incorporating an electro-static antenna, said unit also incorporating an electro-static antenna, for communicating wirelessly with the electro-static antenna in the integrated circuit. 8. The catheter of claim 1, wherein the adjoining fixes so that the sealed sensor remains fixed to said end during said inserting and during withdrawal of the catheter from the patient. 9. A catheter, comprising: a sealed catheter body including a handle and an insertion end for inserting into a patient, anda sealed sensor adjoined to said insertion end, the sensor and catheter body being sealed from each other, anda local power-delivery/data-reception unit incorporated at the insertion end and configured for wirelessly emitting an electromagnetic signal that powers the sensor, said sensor being configured for wirelessly sending an electromagnetic data signal to the local power-delivery/data-reception unit,further comprising, for said returning by the sensor, a magnetic loop antenna having an inner coil and an outer coil, each coil having two loops in respective parallel planes, the inner loop being nested within the outer loop. 10. The catheter of claim 9, wherein the inner and outer coils are electrically isolated from each other and disposed to communicate energy with one another by mutual inductance. 11. The catheter of claim 9, wherein said unit comprises an antenna formed into a single coil, said coil having a single loop, said single loop residing in a single plane. 1... | Koninklijke Philips N.V. | A61B005/04
US08515509 | B2 | 2013-08-20 | Multi-stream emitter for noninvasive measurement of blood constituents | 2009-08-03 | The present disclosure relates to an emitter that is suitable for a noninvasive blood constituent sensor. The emitter is configured as a point optical source that comprises a plurality of LEDs that emit a sequence of pulses of optical radiation across a spectrum of wavelengths. In some embodiments, the plurality of sets of optical sources may each comprise at least one top-emitting LED and at least one super luminescent LED. In some embodiments, the emitter comprises optical sources that transmit optical radiation in the infrared or near-infrared wavelengths suitable for detecting glucose. In order to achieve the desired SNR for detecting analytes like glucose, the emitter may be driven using a progression from low power to higher power. In addition, the emitter may have its duty cycle modified to achieve a desired SNR. | 1. A multi-stream emitter for a noninvasive, physiological device configured to transmit optical radiation in a tissue site, said emitter comprising: a set of optical sources comprising: a top-emitting diode affixed to a first surface, anda side-emitting diode affixed to a second surface, wherein the first surface is approximately perpendicular to the second surface; anda driver configured to: drive the top-emitting diode to transmit first infrared optical radiation at a first power into a tissue site of a patient, wherein the top-emitting diode is configured to emit the first infrared optical radiation in a first direction substantially perpendicular to the first surface; anddrive the side-emitting diode to transmit second infrared optical radiation at a second power into the tissue site of the patient, wherein the side-emitting diode is configured to emit the second infrared optical radiation in a second direction substantially parallel to the second surface, and wherein the second power is greater than the first power; enabling measurement of an analyte in the tissue site of the patient. 2. The emitter of claim 1, wherein the side-emitting diode comprises a super luminescent light emitting diode. 3. The emitter of claim 1, wherein the side-emitting diode comprises a laser diode. 4. The emitter of claim 1 further comprising a submount that configures the set of optical sources as a single point source. 5. The emitter of claim 4 wherein the submount is configured to position the set of optical sources within a diameter of about 2 mm of each other. 6. The emitter of claim 4 wherein the submount is configured to position the set of optical sources within a diameter of about 4 mm of each other. 7. The emitter of claim 1, wherein the top-emitting diode is configured to emit the first infrared optical radiation at a wavelength of approximately 900 to approximately 1350 nm. 8. The emitter of claim 1, wherein the side-emitting diode is configured to emit the second infrared optical radiation at wavelengths of approximately 1600 to approximately 1800 nm. 9. The emitter of claim 1, further comprising a heat sink thermodynamically coupled to the set of optical sources. 10. The emitter of claim 1, further comprising at least one thermistor. 11. The emitter of claim 10, wherein the at least one thermistor is configured to detect a temperature of at least one of the optical sources. 12. The emitter of claim 10, wherein the at least one thermistor is configured to detect a temperature of the tissue site. 13. The emitter of claim 1, wherein the set of optical sources comprises a plurality of side-emitting diodes. 14. The emitter of claim 1, wherein the set of optical sources comprises a plurality of top-emitting diodes. 15. An emitter of a noninvasive optical sensor, the emitter configured to transmit optical radiation into a tissue site to be detected after attenuation by said tissue site and processed to determine one or more physiological para... | Cercacor Laboratories, Inc. | A61B005/00
US08439686 | B2 | 2013-05-14 | Device, system, and method for treating psychiatric disorders | 2009-07-10 | Provided are a device, system, and method for treating a patient having posttraumatic stress disorder (PTSD) or other such psychiatric disorder. An event being recounted by the patient is recorded and a sensory environment of the patient is controlled so as to administer exposure therapy to the patient. Patient mental state may be monitored during administration of the exposure therapy to the patient, and the exposure therapy may be varied in correspondence to the monitored mental state of the patient. In one embodiment, hierarchy assembly means creates a treatment hierarchy based on input from the patient, and hierarchy navigation means navigates within the treatment hierarchy in response to the mental state of the patient. | 1. A computer-based system for treating a patient having a psychiatric disorder, the system comprising: a memory; anda processor configured to perform the steps of: recording a patient's recounting of an event that was experienced by the patient;parsing or dividing the recounted event into one or more stimuli for subsequent playback to the patient;generating a mental state metric for each of the one or more stimuli, wherein each mental state metric is indicative of a mental state of the patient at a time when an associated stimulus was being recounted; andcausing a selected one of the stimuli to be played back to the patient one or more times,wherein a mental state of the patient is monitored as the selected one of the stimuli is played back one or more times, andwherein a different one of the stimuli is selected for subsequent play back to the patient based on the monitored mental state of the patient. 2. The computer-based system of claim 1, further comprising: a patient interface for playing back one or more of the stimuli, including providing or directing visual and audio exposure; anda sensor for monitoring the mental state of the patient, including sensing a level of patient anxiety or distress. 3. The computer-based system of claim 1, wherein the processor is further configured to perform the steps of: associating, with each of the stimuli, a first mental state metric value representative of the monitored mental state of the patient when the patient recounted the stimulus; andassociating, with each of the stimuli, a second mental state metric value representative of the monitored mental state of the patient when the stimuli was played back to the patient. 4. The computer-based system of claim 3 wherein the processor is furthermore configured to perform the step of: varying a sequence of the stimuli to be played back to the patient in correspondence to the second mental state metric value until the second mental state metric value indicates habituation of the patient as compared with the first mental state metric value. 5. The computer-based system of claim 3 Wherein the processor is furthermore configured to perform the step of: varying an intensity of at least one of the stimuli in correspondence to the second mental state metric value until the second mental state metric value indicates habituation of the patient as compared with the first mental state metric value. 6. The computer-based system of claim 3 wherein the processor is further configured to cause the system to provide a course of imaginal exposure therapy. 7. The computer-based system of claim 6 wherein the imaginal exposure therapy is implemented by playing back a recording in the voice or image of the patient. 8. The computer-based system of claim 3 adapted to provide a course of in vivo exposure therapy. 9. The computer-based system of claim 8 wherein the in vivo exposure therapy is implemented through use of a portable computing device that guides the patient to ap... | - | A61B005/00 G09B005/06 G09B001/00
US20100016683 | A1 | 2010-01-21 | SMART INSTALL | 2008-01-10 | When installing components of a remote patient health monitoring system (14′), a generic off-the-shelf set top box (STB) (12) and home gateway (30) are configured by inserting a configuration carrier (42) into each device. The configuration carrier (42) has stored thereon configuration routines and information provided by the remote health monitoring service provider. Configuration information includes, without being limited to, network IDs and channel information, hardware ID information, and the like. The home gateway (30) is connected to the STB (12) using WiFi network link, and to a central server (20) over an Ethernet link. The home gateway (30) is further coupled to one or more patient monitoring devices using a wireless link, such as Bluetooth. | 1. A system for installing and configuring a remote healthcare system (14′), including: a set top box (STB) (2);a home gateway (4) having a first reader (48) thereon; anda configuration carrier (6) that stores configuration information provided by a healthcare service provider for configuring both the STB (2) and the home gateway (4);wherein the configuration carrier (6) is inserted into the first reader (48) of the home gateway (4) and the configuration information is read and stored to persistent memory (50) to configure the home gateway (4) to establish an Internet link to a central server (8) and to establish at least one of a wireless or power line communication access point for the STB (2). 2. The system according to claim 1, wherein the STB (2) includes a second reader (40) into which the configuration carrier (6) is inserted, and which stores the configuration information to persistent memory (44) to configure the STB (2) and to establish at least one of a wireless or power line communication link with the home gateway (4). 3. The system according to claim 2, wherein the system (14′) monitors health status of a patient in a remote location. 4. The system according to claim 1, wherein the STB (2) is coupled to a graphical user interface (GUI) (10) that presents information to a patient. 5. The system according to claim 4, wherein the GUI (10) is a television set. 6. The system according to claim 4, wherein the patient enters information into the STB (2) using a remote control (12). 7. The system according to claim 4, wherein the configuration carrier additionally stores information associated with at least one of user profile, language preference, resources, or software, for configuring the GUI (10) and identifying the patient to the central server (8). 8. The system according to claim 1, further including at least one monitoring device coupled to one of the home gateway (4) and the STB (2) via a wireless communication link, wherein the at least one monitoring device monitors a physiological condition of the patient. 9. The system according to claim 8, wherein the at least one monitoring device is at least one of an electronic scale, a blood pressure (BP) monitor, a pulse rate monitor, and an SpO2 monitor. 10. The system according to claim 8, wherein the at least one monitoring device is configurable by the configuration carrier to generate the wireless communication link. 11. The system according to claim 2, wherein the STB (2) is a personal computer having a built-in GUI (10), and wherein the patient employs at least one of a mouse and a keyboard to input information into the STB (2). 12. The system according to claim 2, wherein the configuration carrier (6) is at least one of a universal serial bus (USB) portable memory stick, a DVD, a CD-ROM, a smartcard, a key card, a bar-coded card, a near field communication card, or a magnetic strip card. 13. The system according to claim 2, wherein the configuration component (6) inclu... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/00 G06F003/048 G06F003/01
US20100010595 | A1 | 2010-01-14 | APPARATUS FOR THERMAL TREATMENT OF TISSUE | 2007-08-23 | An apparatus for thermal treatment of tissue comprises a compact manipulator (37) for positioning and orienting an energy radiator (50) for heat treatment of biological tissues. The energy radiator (50) is adapted to emit energy in the direction of a focal axis into a focal volume. The manipulator (37) can adjust the position and orientation of a suspension body (20) in a plane. The plane wherein the suspension body can be maneuvered is parallel to a support face of the apparatus. The energy radiator (50) is suspended from the suspension body (20). The energy radiator (50) can be manipulated with five independent degrees of freedom relative to said support face. | 1. An apparatus for thermal treatment of tissue, provided with a support face for supporting at least a portion of a patient's body and comprising: an energy radiator mounted to a holder, for directing energy along a focusing axis into a focal volume, anda manipulator comprising:a manipulator transmission unit including a suspension body,a transmission driver unit mounted to the manipulator transmission unit, the transmission driver unit comprising at least one transmission driver andthe holder suspended from the suspension body, wherein the manipulator transmission unit is driveable by the transmission driver unit and the holder is driveable by the manipulator transmission unit, wherein the manipulator transmission unit has a first transmission subunit for translating and rotating the suspension body in a plane substantially parallel to the support face and has a second transmission subunit for moving the energy radiator along the focusing axis and for rotating the energy radiator around two distinct axes perpendicular to the focusing axis. 2. The apparatus as claimed in claim 1, wherein the suspension body comprises distant portions and the first transmission subunit comprises moveably guided abutments and abutment guides, each of the distant portions being rotatably connected to at least one of the moveably guided abutments and each moveably guided abutment being guidably supported by at least one of the abutment guides. 3. The apparatus as claimed in claim 2, wherein the first transmission subunit comprises at least one transmission body for cooperation with one of the moveably guided abutments and being coupled to a separate transmission driver of the transmission driver unit. 4. The apparatus as claimed in claim 3, wherein the distant portions each have a first axis of rotation with respect to the moveably guided abutment to which the respective distant portion is connected and wherein the transmission body cooperating with the respective moveably guided abutment comprises a first threaded portion, said portion having a first length axis, said first length axis intersecting with said first axis of rotation. 5. The apparatus as claimed in claim 4, wherein at least a part of at least one of the moveably guided abutments is elastically deformed to establish pretension between the moveably guided abutment and the first threaded portion comprised by the transmission body cooperating with the respective moveably guided abutment. 6. The apparatus as claimed in claim 1, wherein the holder comprises three levers, moveably suspended from the suspension body. 7. The apparatus as claimed in claim 6, wherein the second transmission subunit comprises three mechanisms, each mechanism cooperating with one of the three levers and each lever being moveably connected to one of the three mechanisms. 8. The apparatus as claimed in claim 7, wherein the transmission driver unit comprises three further transmission drivers, each mechanism being coupled to o... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/055 A61N007/02
US20100023101 | A1 | 2010-01-28 | BIOCOMPATIBLE COMPOSITES | 2005-02-02 | The present invention relates to biocompatible composites, in particular biocompatible nanotube composites in the form of a fiber mat and/or film structure, comprising nanotubes and at least one biomolecule. The invention also relates to a process for preparing a biocompatible composite involving (i) forming a dispersing media comprising nanotubes and at least one biomolecule; and either (ii) introducing the dispersing media of step (i) into a coagulating media optionally comprising at least one biomolecule so as to form a continuous fiber; or (iii) filtering the dispersing media of step (i). Alternatively, the process involves (i) forming a dispersing media comprising nanotubes; and (ii) introducing the dispersing media of step (i) into a coagulating media comprising at least one biomolecule so as to form a continuous fiber. The biocompatible composite is useful as a medical device, preferably in a bio-electrode, bio-fuel cell or substrates for electronically stimulated bio-growth. | 1. A biocompatible composite that is formed into a fibre, mat and/or film structure, comprising nanotubes and at least one biomolecule. 2. A biocompatible composite according to claim 1, wherein the biomolecule is selected from one or more of the group consisting of biological electrolytes, nucleic acids, polyaminoacids, proteins, enzymes, polysaccharides, lipids and/or hormones. 3. A biocompatible composite according to claim 2, wherein the biological electrolyte is selected from one or more of the group consisting of hyaluronic acid (HA), chitosan, heparin, chondroitin sulphate, polyglycol acid (PGA), polylactic acid (PLA), polyamides, poly-2-hydroxy-butyrate (PHB), polycaprolactone (PCL), poly(lactic-co-glycolic) acid (PLGA), protamine sulfate, polyallylamine, polydiallyldimethylammonium, polyethyleneimine, eudragit, gelatin, spermidine, albumin, polyacrylic acid, sodium alginate, polystyrene sulfonate, carrageenin and/or carboxymethylcellulose. 4. A biocomptabile composite according to claim 2, wherein the nucleic acid is selected from one or more of the group consisting of DNA, cDNA, RNA, oligonucleotide, oligoribonucleotide, modified oligonucleotide, modified oligoribonucleotide and/or peptide nucleic acid (PNA) or hybrid molecules thereof. 5. A biocompatible composite according to claim 2, wherein the polyamino acid is selected from one or more of the group consisting of poly-L-lysine, poly-L-arginine, poly-L-aspartic acid, poly-D-glutamaic acid, poly-L-glutamaic acid, poly-L-histidine and/or poly-(DL)-lactide. 6. (canceled) 7. (canceled) 8. (canceled) 9. (canceled) 10. (canceled) 11. A biocompatible composite according to claim 1, wherein the biomolecule is selected from one or more of the group consisting of hyaluronic acid (HA), chitosan, heparin, chondroitin sulphate, DNA and/or poly-L-lysine. 12. A biocompatible composite according to claim 1, wherein the nanotubes are selected from one or more of the group consisting of carbon nanotubes, metal oxide nanotubes and/or peptidyl nanotubes. 13. (canceled) 14. A biocompatible composite according to claim 1, wherein the nanotubes are SWNTs and/or MWNTs. 15. A biocompatible composite according to claim 1, wherein the composite is a fibre selected from the group consisting of DNA-SWNT-chitosan fibres, HA-SWNT-chitosan fibres, chitosan-SWNT-chondroitin sulphate fibres and chitosan-SWNT-heparin fibres or a film selected from the group consisting of chitosan-SWNT-films, DNA-SWNT films and poly-1-lysine-SWNT films. 16. A biocompatible compatible composite according to claim 1, wherein said biomolecule is present in an amount in the range of 10-50% based on the total weight of the composite. 17. A biocompatible composite according to claim 1, wherein said composite further comprises an additive in an amount in the range of 1 to 50% based on the total weight of the composite, wherein said additive is selected from the group consisting of drugs, growth factors, hormones, antibiotics,... | University of Wollongong | A61N001/05 A61K038/00 A61K031/7088 A61F002/00 A61B005/0408
US20100049172 | A1 | 2010-02-25 | Sensing and interactive drug delivery | 2009-06-29 | An interactive drug delivery system includes a drug delivery module, an optical probe, a local controller, and an optional central controller. The drug delivery module is constructed and arranged to deliver selected amounts of a drug into a subject. The optical probe is constructed and arranged to detect in a selected tissue region of the subject a manifestation caused by the delivered drug. The local controller is constructed and arranged to receive data from or transmit data to the optical probe and the drug delivery module. The local controller is arranged to correlate optical data, received from the optical probe, to selected data and provide signals to the drug delivery module for adjusting the amounts of the drug to be delivered into the subject. | 1. an interactive drug delivery system comprising: a drug delivery module constructed and arranged to deliver selected amounts of a drug into a subject;an optical probe constructed and arranged to detect in vivo in a selected tissue region of said subject a manifestation caused by the delivered drug; anda local controller constructed and arranged to receive data from or transmit data to said optical probe and said drug delivery module; said local controller being arranged to correlate optical data received from said optical probe to selected data and provide signals to said drug delivery module for adjusting the amounts of the drug to be delivered into said subject. 2. The system of claim 1 further including a central controller remotely located from and in wireless communication with said local controller, central controller constructed and arranged to receive data from said local controller and provide to said local controller data for adjusting the amounts of said drug to be delivered into said subject. 3. The system of claim 1 further including a temperature probe coupled to said local controller. 4. The system of claim 3 wherein said temperature probe uses an infra-red beam to measure a local temperature of a tissue region. 5. The system of claim 1 further including an electrical probe coupled to said local controller. 6. The system of claim 1 wherein said optical probe is a time-resolved (TRS) system. 7. The system of claim 1 wherein said optical probe is a phase modulation system. 8. The system of claim 1 wherein said optical probe is a phased array. 9. The system of claim 1 wherein said optical probe is phase cancellation system. 10. The system of claim 1 wherein said optical probe is an amplitude cancellation system. 11. The system of claim 1 wherein said drug delivery module includes an electrical controller designed to regulate the rate of the delivered drug based on the signals from said local controller. 12. An interactive drug delivery method comprising: delivering by a drug delivery module selected amounts of a drug into a subject;optically detecting in vivo in a selected tissue region of said subject a manifestation caused by the delivered drug by an optical probe;receiving data from or transmitting data to said optical probe and said drug delivery module by a local controller;correlating data received from said optical probe to selected data stored in said local controller; andproviding signals to said drug delivery module for adjusting the amounts of the drug to be delivered into said subject. 13. The interactive drug delivery method of claim 12 further including regulating the rate of the delivered drug based on signals from said local controller. 14. The interactive drug delivery method of claim 12 including measuring tissue temperature. 15. The interactive drug delivery method of claim 14 wherein said measuring temperature includes using an infra-red beam for measuring a local temperature of a tissue region. 16... | - | A61M005/00 A61B005/01
US20100049096 | A1 | 2010-02-25 | SYSTEM FOR FALL PREVENTION AND A METHOD FOR FALL PREVENTION USING SUCH A SYSTEM | 2007-11-09 | System for fall prevention for a user, comprising a number of sensors (2) attachable to at least one lower body segment (3), wherein said sensors (2) are adapted to measure movement of said at least one lower body segment (3) and to translate the movement into a signal (S), the system further comprising a control (12) adapted to receive the signal (S) from said respective sensors (2), wherein in use the control (12) observes the signal (S) as an actual sequence of postures of said at least one lower body segment (3) and compares the actual sequence with a predetermined sequence of postures as a function of time, (the predetermined sequence relating to a low risk of falling,) wherein the control (12) is adapted to determine a high risk of falling when the actual sequence deviates from the predetermined sequence (to a certain degree). The invention further relates to a method for fall prevention using such a system for fall prevention. | 1. A system for fall prevention for a user, comprising a number of sensors (2) attachable to at least one lower body segment (3), wherein said sensors (2) are adapted to measure movement of said at least one lower body segment (3) and to translate the movement into a signal (S), the system further comprising a control (12) adapted to receive the signal (S) from said respective sensors (2), where in use, the control (12) observes the signal (S) as an actual sequence of postures of said at least one lower body segment (3) and compares the actual sequence with a predetermined sequence of postures as a function of time, wherein the control (12) is adapted to determine a high risk of falling when the actual sequence deviates from the predetermined sequence. 2. The system according to claim 1, wherein the posture of the lower body segment (3) is determined by the position of at least two lower body segment parts (6, 7) relative to each other. 3. The system according to claim 2, wherein the lower body segment parts comprise an ankle (8), a foot (9), a knee (5), a lower leg (6), an upper leg (7), hip (4) of a similar lower body segment (3) and/or a trunk. 4. The system according to claim 1, wherein a comparison of the actual sequence of postures to the predetermined sequence of postures of the lower body segment (3) is performed with aid of an adaptive algorithm (11). 5. The system according to claim 1, configured to monitor a muscle strength or power of muscles of the lower body segment (3), and configured to use a detected muscle strength or power in the determining of the high risk of falling. 6. The system according to claim 1, wherein the predetermined sequence of postures of the lower body segment (3) is determined by measuring successive lower body segment (3) postures during normal movement of the user and the amount of variation therein. 7. The system according to claim 6, wherein the deviation of the actual sequence of postures in relation to the predetermined sequence of postures is based on the increase or decrease in variation in the sequence as a function of time. 8. The system according to claim 6, wherein the high risk of falling is determined by a deviation threshold that is estimated from a mean and the variation by classifying the actual sequence of postures. 9. The system according to claim 1, wherein the system is adapted to provide a warning signal to the user, during walking, when the high risk of falling has been determined. 10. The system according to claim 1, wherein the system comprises a memory (10) for storing the sequence of postures of the at least one lower body segment (3). 11. The system according to claim 1, wherein the system is self-learning by adaptation of the predetermined sequence of postures in case of changing conditions of the user. 12. The system according to claim 1, wherein the predetermined sequence of postures is determined by entering parameters into the control (12) preceding using the system ... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/103
US20100040544 | A1 | 2010-02-18 | METHOD FOR MANAGING PATIENTS WITH CANCER | 2007-11-20 | The invention relates to methods for managing patients with cancers, more specifically to the planning of tumor treatment and monitoring treatment response, said management comprises MR imaging and/or MR spectroscopy with an MR imaging medium which comprises hyperpolarized 13C-pyruvate. | 1. Method of management of patients with cancer having a tumor, said method comprises a) performing on the patient 13C-MR detection of the tumor and optionally of surrounding organs and/or the surrounding area using an imaging medium that comprises hyperpolarized 13C-pyruvate wherein signals of 13C-pyruvate and its 13C-containing metabolites alanine, lactate and optionally bicarbonate are detected and;b) using information obtained in step a) optionally together with other information obtained in investigations carried out to confirm diagnosis of cancer to determine the stage of disease. 2. Method according to claim 1 comprising as a further step c) selecting the treatment of the cancer. 3. Method according to claim 2 comprising the further steps d) carrying out treatment; and e) repeating step a) to monitor the early treatment response, and/or treatment response and/or the recurrence of the cancer. 4. Method according to claim 1 wherein said patient is a patient with prostate cancer and 13C-MR detection in step a) is a 13C-MR detection of the prostate. 5. Method of management of patients with non-metastatic cancer having a tumor, said method comprises a) acquiring MR images and/or MR spectra of the tumor and optionally of surrounding organs and/or the surrounding area;b) performing thermal ablation treatment of the tumor using MR guided high intensity focused ultrasound; andc) following the treatment, performing on the patient 13C-MR detection of the tumor and optionally of surrounding organs and/or the surrounding area using an imaging medium that comprises hyperpolarized 13C-pyruvate wherein signals of 13C-pyruvate and its 13C-containing metabolites alanine, lactate and optionally bicarbonate are detected to evaluate the treatment of step b). 6. Method according to claim 5 comprising a further step d) repeating step c) to monitor the early treatment response, and/or treatment response and/or the recurrence of the cancer. 7. Array of devices in a MR scanner room comprising a MR scanner, a patient table, a DNP polarizer, an injector and a transducer that generates focused ultrasound beams. 8. Array according to claim 7 wherein the patient table is docked to the MR scanner, the DNP polarizer is for producing a MR imaging medium comprising a hyperpolarized imaging agent, preferably comprising hyperpolarized 13C-pyruvate, and the injector is connected to said DNP polarizer for injecting the MR imaging medium into a patient on said patient table. 9. Array according to claim 7 wherein said transducer is incorporated into said patient table and said patient table further comprises an automated positioning system. 10. Medical algorithms for patients with cancer having a tumor, preferably for patients with prostate cancer, corresponding to the methods of the invention. 11. Medical algorithms according to claim 10 wherein said medical algorithms are on a computer readable storage medium. | - | A61K051/00 A61B005/055
US20100010321 | A1 | 2010-01-14 | MEDICAL SYSTEM WHICH INCLUDES A BACKPACK POUCH | 2009-07-09 | A medical system includes a folding backpack pouch adapted to be carried by a person, a patient monitoring module, and a laptop computer adapted to display a physiological parameter measured by the patient monitoring module. The patient monitoring module and the laptop computer are held by corresponding first and second panels of the pouch. | 1. A medical system comprising: a) a folding backpack pouch adapted to be carried by a person and including first and second panels together having an open position wherein the first and second panels are disposed side by side and having a closed position wherein one of the first and second panels overlies the other of the first and second panels;b) a no-tool-user-replaceable patient monitoring module held by the first panel and adapted to measure at least one physiological parameter of a patient; andc) a no-tool-user-removable laptop computer held by the second panel, wherein the laptop computer is operatively connectable to the patient monitoring module and is adapted to display the measured at-least-one physiological parameter of the patient. 2. The medical system of claim 1, also including a no-tool-user-replaceable patient gas-exchange module held by the first panel and adapted to exchange at least one gas with the patient, wherein the laptop computer is operatively connectable to the patient gas-exchange module. 3. A medical system comprising: a) a folding backpack pouch adapted to be carried by a person and including first and second panels together having an open position wherein the first and second panels are disposed side by side and having a closed position wherein one of the first and second panels overlies the other of the first and second panels;b) a no-tool-user-replaceable patient monitoring module held by the first panel and including an ECG (electrocardiogram) medical unit, an SpO2 (saturation of peripheral oxygen) medical unit, and a blood pressure medical unit each operatively connectable to a patient;c) a no-tool-user-replaceable patient gas-exchange module held by the first panel and including an oxygen-delivery medical unit and a CO2 (carbon dioxide) monitoring medical unit each operatively connectable to the patient;d) a first inter-module connecting cable affixed to the backpack pouch, having a plug connected to the patient monitoring module, and having a plug connected to the patient gas-exchange module;e) a battery unit held by the first panel and adapted to power the patient monitoring module and the patient gas-exchange module; andf) a no-tool-user-removable laptop computer held by the second panel, wherein the laptop computer is operatively connectable to the patient monitoring module and is adapted to display data from the ECG, SpO2, and blood pressure medical units, wherein the laptop computer is operatively connectable to the patient gas-exchange module through the patient monitoring module, and wherein the laptop computer is adapted to display data from the CO2 monitoring medical unit and to control the oxygen-delivery medical unit to deliver oxygen to the patient. 4. The medical system of claim 3, wherein the second panel includes a pocket for holding the laptop computer, and also including a computer cable having a plug connectable to the laptop computer and having a plug connectable to the patient monitori... | Ethicon Endo-Surgery, Inc. | A61B005/00 A61M016/00 A61M005/00 H02J007/00
US07657312 | B2 | 2010-02-02 | Multi-site ventricular pacing therapy with parasympathetic stimulation | 2003-11-03 | A method and device for delivering multi-site ventricular pacing therapy in conjunction with parasympathetic stimulation for reducing ventricular wall stress. Such reduction in ventricular wall stress is useful in reversing or preventing the ventricular remodeling which can occur in heart failure patients. | 1. An implantable device for delivering cardiac function therapy to a patient, comprising: a plurality of pacing channels for delivering pacing pulses to multiple ventricular sites;a parasympathetic stimulation channel for stimulating parasympathetic nerves innervating the heart;a sensor for measuring cardiac output;an exertion level sensor for measuring the patient's exertion level;a controller for controlling the delivery of pacing pulses to the multiple ventricular sites in accordance with a programmed pacing mode;wherein the controller is programmed to deliver multi-site ventricular pacing therapy in conjunction with parasympathetic stimulation and wherein the controller is further programmed to deliver the multi-site ventricular pacing therapy with a pacing pulse output sequence that pre-excites at least one myocardial region relative to other regions for reducing ventricular wall stress;wherein the controller is further programmed to deliver the multi-site ventricular pacing in accordance with a demand pacing mode that counteracts slowing of the heart rate by the parasympathetic stimulation so that no slowing of the heart rate occurs; and,wherein the controller is programmed to compute a function that maps exertion levels to minimum cardiac output values considered to be adequate for a particular exertion level and is further programmed to cease the delivery of parasympathetic stimulation if a presently measured cardiac output is below the minimum cardiac output indicated as adequate by the computed function. 2. The device of claim 1 wherein the cardiac output sensor is a trans-thoracic impedance measuring circuit. 3. The device of claim 1 wherein the controller is programmed to deliver parasympathetic stimulation only when cardiac output is above a specified limit value. 4. The device of claim 1 wherein the controller is programmed to modulate the delivery of parasympathetic stimulation in accordance with the measured exertion level. 5. The device of claim 4 wherein the controller is programmed to deliver parasympathetic stimulation only when the measured exertion level is below a specified limit value. 6. The device of claim 4 wherein the exertion level sensor is a minute ventilation sensor. 7. The device of claim 4 wherein the exertion level sensor is an accelerometer. 8. The device of claim 1 wherein the controller is programmed to compute the function that maps exertion levels to minimum cardiac output values considered to be adequate for a particular exertion level using a look-up table. 9. A method for operating an implantable cardiac device in order to deliver therapy to a patient, comprising: stimulating parasympathetic nerves innervating the heart in order to reduce ventricular wall stress;delivering multi-site ventricular pacing therapy with a pacing pulse output sequence that pre-excites at least one myocardial region relative to other regions;delivering pacing pulses to multiple ventricular sites in accordance with a de... | Cardiac Pacemakers, Inc. | A61N001/00 A61B005/0402
US07662113 | B2 | 2010-02-16 | Fingertip tracker | 2005-11-07 | A method, apparatus, system, and article of manufacture provide the ability to utilize fingertip tracking technology during the course of a medical examination. A sensor is mounted on a fingertip of a finger of a first person. Such a sensor is communicatively coupled to a computer to record measurements. The sensor is calibrated using a known target. The finger having the sensor is then engaged to perform a medical examination directly on a second person. Measurements are obtained from the sensor on the fingertip during the medical examination. Data is obtained (e.g., via calculations or interpolation) based on the measurements. | 1. A computer implemented method for conducting a medical exam comprising: (a) mounting a sensor on a fingertip of a finger of a first person, wherein the sensor is communicatively coupled to a computer;(b) calibrating the sensor using a known target;(c) engaging the finger having the sensor to perform a medical examination directly on a second person;(d) obtaining measurements from the sensor on the fingertip during the medical examination, wherein the measurements: (i) relate to a position of the sensor with respect to an anatomical element of the second person; and(ii) comprise a distance between a sensor on a left hand and a sensor on a right hand of the first person; and(e) obtaining data based on the measurements. 2. The method of claim 1, wherein the sensor comprises an electromagnetic motion sensor. 3. The method of claim 1, further comprising correcting all data using a best-fit calibration function prior to calibrating the sensor. 4. The method of claim 1, wherein the calibrating comprises: inserting the sensor into a calibration device prior to mounting the sensor on the fingertip, wherein the calibration device comprises one or more known points;storing coordinates of the one or more known points based on the sensor to provide one or more locations in space;touching the one or more known points in the calibration device with the sensor mounted on the fingertip to obtain coordinates of the sensor;subtracting the coordinates of the one or more known points from the coordinates of the sensor to provide a vector pointing from the sensor to the known point at the fingertip; anddetermining a matrix based on the vector that will rotate the vector to account for a position of the fingertip. 5. The method of claim 4, further comprising: determining if an accuracy of the coordinates of the sensor are within a defined threshold; andrepeating the touching and determining steps until the accuracy of the coordinates of the sensor are within the defined threshold. 6. The method of claim 1, wherein multiple sensors are used on two or more fingertips of the user. 7. The method of claim 1, wherein the measurements comprise a distance between two sensors on a single hand of the first person. 8. The method of claim 1, wherein the measurements comprise a distance between a midpoint of two sensors on the right hand to a midpoint of two sensors on the left hand. 9. The method of claim 1, wherein the measurement comprises a depth of an object on the second person. 10. The method of claim 1, wherein the data obtained comprises a plot of the measurements on a display device of the computer. 11. The method of claim 1, wherein: the measurements comprise reference points of a first limb that is serving as a reference to a fractured bone on a second limb of the second person; andthe data is displayed on a display device of the computer system;the first person utilizes the fingertip having the sensor to perform a fracture reduction on the fractured bone o... | California Institute of Technology | A61B005/117 A61B005/103
US07647094 | B2 | 2010-01-12 | Method for deriving standard 12-lead electrocardiogram, and electrocardiograph using the same | 2005-05-16 | In order to derive a standard 12-lead electrocardiogram, ones of limb leads and chest leads which constitute a lead system for the standard 12-lead electrocardiogram using 10 electrodes are selected. Electrocardiographic waveforms corresponding to the selected ones of the limb leads and the chest leads are obtained, as measured electrocardiographic waveforms, with electrodes attached on a living body. Electrocardiographic waveforms corresponding to a remaining ones of the limb leads and the chest leads are calculated, as non-measured electrocardiographic waveforms, with a prescribed transformation matrix and the measured electrocardiographic waveforms. | 1. A method of deriving a standard 12-lead electrocardiogram, comprising: selecting ones of limb leads and chest leads which constitute a lead system for the standard 12-lead electrocardiogram using 10 electrodes;obtaining electrocardiographic waveforms corresponding to the selected ones of the limb leads and the chest leads, as measured electrocardiographic waveforms, with electrodes attached on a living body; andcalculating electrocardiographic waveforms corresponding to a remaining ones of the limb leads and the chest leads, as non-measured electrocardiographic waveforms, with a single prescribed transformation matrix and the measured electrocardiographic waveforms;wherein the transformation matrix includes information indicative of a relationship between the selected ones of the chest leads and the remaining ones of the chest leads which is obtained in advance from past results of the standard 12-lead electrocardiogram using 10 electrodes; andwherein the relationship is an average relationship which is obtained from the past results of one specific person or any plurality of persons. 2. The derivation method as set forth in claim 1, wherein the relationship is obtained such that an ST segment of an electrocardiogram has the best accuracy in the calculated electrocardiographic waveforms. 3. The derivation method as set forth in claim 1, wherein the relationship is obtained by a least squares method. 4. An electrocardiograph for deriving a standard 12-lead electrocardiogram, comprising: a plurality of electrodes, adapted to be attached on portions of a living body corresponding to selected ones of limb leads and chest leads which constitute a lead system for the standard 12-lead electrocardiogram using 10 electrodes;a processor, which calculates electrocardiographic waveforms corresponding to a remaining ones of the limb leads and the chest leads, as non-measured electrocardiographic waveforms, with a single prescribed transformation matrix and electrocardiographic waveforms corresponding to the selected ones of the limb leads and the chest leads which are obtained with the electrodes as measured electrocardiographic waveforms; anda display, which displays the measured electrocardiographic waveforms and the non-measured electrocardiographic waveforms; wherein:the transformation matrix includes information indicative of a relationship between the selected ones of the chest leads and the remaining ones of the chest leads which is obtained in advance from past results of the standard 12-lead electrocardiogram using 10 electrodes; andthe relationship is an average relationship which is obtained from the past results of one specific person or any plurality of persons. 5. The electrocardiograph as set forth in claim 4, wherein the relationship is obtained such that an ST segment of an electrocardiogram has the best accuracy in the calculated electrocardiographic waveforms. 6. The electrocardiograph as set forth in claim 4, wherein the relationsh... | Nihon Kohden Corporation Daming Wei | A61B005/04
US08221418 | B2 | 2012-07-17 | Endoscopic instrument for tissue identification | 2009-02-05 | The present disclosure relates to various apparatus, systems and methods of identifying and treating tissue using at least one electrical property of tissue. Provided is a method for identifying and treating tissue, the method including providing a electrosurgical treatment device including an electrode assembly for measuring one or more electrical properties of a target tissue, the electrode assembly being mounted on a distal end thereof, measuring the one or more electrical characteristics of the target tissue, comparing the measured electrical property values of the target tissue against electrical property values of known tissue types, identifying a tissue type of the target tissue, adjusting an energy delivery configuration of the electrosurgical treatment device to the type of target tissue, and activating the electrosurgical treatment device to treat the target tissue. | 1. A method for identifying and treating tissue, the method comprising the steps of: providing an electrosurgical treatment device including an electrode assembly configured to measure one or more electrical properties of a target tissue;applying an electrical signal having at least one of a range of frequencies and a range of wave configurations to the target tissue;measuring the one or more electrical properties of the target tissue in response to the electrical signal;comparing values of the measured electrical properties of the target tissue against values of electrical properties of known tissue types;identifying a type of the target tissue based on the comparison;adjusting at least one of the range of frequencies and the range of wave configurations of the electrical signal based on the identified tissue type;applying the adjusted electrical signal to the target tissue;measuring the one or more electrical properties of the target tissue in response to the adjusted electrical signal;comparing values of the measured electrical properties of the target tissue type against values of electrical properties of the tissue type;identifying a condition of the target tissue based on the comparison;adjusting an energy delivery configuration of the electrosurgical treatment device to the type and condition of the target tissue; andactivating the electrosurgical treatment device to perform a treatment on the target tissue. 2. A method according to claim 1, wherein the electrode assembly is mounted on a distal end of the electrosurgical treatment device and includes a plurality of electrodes. 3. A method according to claim 1, further including the steps of measuring at least one of the electrical properties of the target tissue following the treatment; and determining the effectiveness of the treatment. 4. A method according to claim 1, wherein the electrode assembly includes a base having an electrode extending coaxially therethrough. 5. A method according to claim 4, where the coaxially extending electrode is operably connected to a high frequency generator. 6. A method according to claim 5, wherein the high frequency generator is capable of generating a frequency between 30 MHz and 30 GHz. | TYCO Healthcare Group LP | A61B018/18
US08216139 | B2 | 2012-07-10 | Signal processing for continuous analyte sensor | 2009-09-23 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A method for estimating an analyte value from a continuous analyte sensor, the method comprising: receiving a data stream from the continuous analyte sensor for a first time period;estimating at least one analyte value for a second time period based at least in part on the data stream for the first time period; anddetermining, using electronic circuitry, a variation of the at least one estimated analyte value based at least in part on a clinical risk analysis of the at least one estimated analyte value, wherein the variation comprises boundaries that represent a range of possible variations of the at least one estimated analyte value for the second time period. 2. The method of claim 1, wherein the second time period is a future time period. 3. The method of claim 1, wherein the analyte is glucose. 4. A method for estimating an analyte value from a continuous analyte sensor, the method comprising: receiving a data stream from the continuous analyte sensor for a first time period; estimating at least one analyte value for a second time period based at least in part on the data stream for the first time period; anddetermining, using electronic circuitry, a variation of the at least one estimated analyte value based at least in part on physiological patterns of the host's analyte data over time, wherein the variation comprises boundaries that represent a range of possible variations of the at least one estimated analyte value for the second time period. 5. The method of claim 4, wherein the second time period is the same as and/or overlaps the first time period. 6. The method of claim 4, wherein the analyte is glucose. 7. A method for estimating an analyte value from a continuous analyte sensor, the method comprising: receiving a data stream from the continuous analvte sensor for a first time period;estimating at least one analyte value for a second time period based at least in part on the data stream for the first time period;determining, using electronic circuitry, a variation of the at least one estimated analyte value based at least in part on a statistical variation from the at least one estimated analyte value, wherein the variation comprises boundaries that represent a range of possible variations of the at least one estimated analyte value for the second time period; andproviding an output to a display and/or to an external device based at least in part on the variation of the estimated analyte value. 8. The method of claim 7, wherein the step of analyzing a variation is further based at least in part on a physiological variation from the at least one estimated analyte value. 9. The method of claim 7, wherein the step of analyzing a variation is further based at least in part on a clinical risk analysis of the at least one estimated analyte value. 10. The method of claim 7, wherein the estimated analyte value is a real-time analyte sensor value. 11. The method of claim 7, wherein the estimated analyte value is a future analyte ... | DexCom, Inc. | A61B005/00 A61B005/05
US08320982 | B2 | 2012-11-27 | Multi-wavelength optical devices and methods of using same | 2007-12-21 | Multi-wavelength optical apparatus includes an optical emitter, and an energy transition layer positioned adjacent to the optical emitter. The energy transition layer generates multi-wavelength electromagnetic radiation when monochromatic light from the optical emitter passes therethrough. The energy transition layer includes a plurality of luminescent films, and each film is configured to luminesce at a respective different wavelength range when monochromatic light from the optical emitter passes therethrough. The plurality of luminescent films may be arranged in contacting face-to-face relationship or may be arranged in an array. The luminescent films may include rare-earth doped oxides, phosphors, metal-doped oxides, rare-earth doped nitrides, nanostructures, and/or nanostructured films, etc. The optical emitter may be a light emitting diode (LED), a laser diode (LD), an organic light-emitting diode (OLED), a resonant cavity light emitting diode (RCLED), and/or an edge-emitting diode (EELED). | 1. A multi-wavelength optical apparatus that is positionable adjacent to a body of a being, comprising: an optical emitter that emits light along a light path;an energy transition layer positioned in the light path, wherein the energy transition layer comprises a plurality of luminescent films, and wherein at least two of the films are configured to luminesce at different time-dependent optical luminescence lifetimes when light passes therethrough; andan optical detector that receives electromagnetic radiation generated by the interaction between luminescent energy from the energy transition layer and a region of the body. 2. The apparatus of claim 1, wherein the detector is a pulse oximeter detector configured to detect a ratio of oxygenated to deoxygenated blood in a body of a subject. 3. The apparatus of claim 2, wherein the apparatus is configured to receive a finger of a subject between the energy transition layer and the detector. 4. The apparatus of claim 2, wherein the apparatus is configured to receive an earlobe, ear canal, tragus, antitragus, meatus and/or concha of a subject between the energy transition layer and the detector. 5. The apparatus of claim 1, wherein the apparatus comprises a transdermal drug delivery device, wherein the optical emitter emits light through the skin of a subject, and wherein the detector is configured to detect an optical signal passing through or reflecting from the skin and/or at least one blood vessel that is relevant to a drug molecule concentration delivered to the subject. 6. The apparatus of claim 1, wherein the detector is a spectroscopic detector configured to detect biometric identification data. 7. The apparatus of claim 1, wherein the plurality of luminescent films are arranged in contacting face-to-face relationship. 8. The apparatus of claim 1, wherein the optical emitter comprises an array of mesas that generate the light and the apparatus generates an optical power density that is directly proportional to the number of mesas. 9. The apparatus of claim 1, wherein the plurality of luminescent films comprise rare-earth doped oxides, phosphors, metal-doped oxides, rare-earth doped nitrides, nanostructures, and/or nanostructured films. 10. The apparatus of claim 1, wherein the optical emitter comprises a light emitting diode (LED), a laser diode (LD), an organic light-emitting diode (OLED), a resonant cavity light emitting diode (RCLED), and/or an edge-emitting diode (EELED). 11. A method of noninvasively measuring a physiological property of a being, comprising: illuminating a portion of a body of the being with multi-wavelength electromagnetic radiation generated by an energy transition layer, wherein the energy transition layer comprises a plurality of luminescent films, wherein at least two of the films are configured to luminesce at different time-dependent optical luminescence lifetimes when light from an optical emitter passes therethrough; anddetecting electromagnetic energy r... | Valencell, Inc. | A61B005/00
US08588890 | B2 | 2013-11-19 | Magnetic resonance imaging apparatus | 2009-07-22 | A magnetization vector is detected for each pixel position in an imaging region including at least part of a subject, the magnetization vector being excited to have a phase difference between a flow portion and a static portion or between a normal portion and an abnormal portion. A pixel value at each pixel position is determined as a value proportional to an absolute value of the amplitude of the magnetization vector detected for each of the pixel positions. On s the basis of a real part or phase of the magnetization vector detected for each of the pixel positions, the determined pixel value is corrected so that the difference of the pixel value increases between (a) the flow portion or the abnormal portion and (b) the static portion or the normal portion. | 1. A magnetic resonance imaging (MRI) apparatus comprising: an MRI gantry including a static magnet, a set of gradient magnets, at least one radio frequency (RF) coil, an RF transmitter, an RF receiver and control circuits including at least one computer configured todetect a magnetization vector having an amplitude and relative phase angle which correspond to orthogonal real and imaginary vector components for each of a large number of pixel positions in an imaging region including at least part of a subject, the magnetization vector being excited to have a phase angle difference between (a) pixel positions from a flow portion of the imaging region in which a fluid flows and pixel positions from a static portion of the imaging region in which tissues are static or (b) between pixel positions from a normal tissue portion of the imaging region and pixel positions from an abnormal tissue portion of the imaging region having a difference in susceptibility from the normal tissue portion;determine a first pixel value for each pixel position as a value proportional to an absolute value of the amplitude of the magnetization vector detected for that pixel position; andcorrect, on the basis of a real part of the magnetization vector detected for each pixel position, the first pixel value to generate a corrected second pixel value by enhancing magnetization vector phases to make the phase angle for pixel positions representing a blood vessel or abnormal tissue more closely approaches ±180° while the phase angle for pixel positions representing other static or normal background portions approaches zero so that the contrast difference of the corrected second pixel values is increased between (a) the flow portion or the abnormal portion and (b) the static portion or the normal portion. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the control circuits are configured to effect use of an MRI sequence having a shortened echo time as zeroth gradient moment nulling (GMN). 3. The magnetic resonance imaging apparatus according to claim 1, wherein the control circuits are configured to effect: obtaining a real part Re of a background phase φback from a complex signal Re+jIm which is found from the magnetization vector,generating a cosine filter on the basis of the real part Re of the background phase φback, andapplying the cosine filter to an amplitude image based on the first pixel values to obtain an image in which the real part Re of the background phase φback is corrected. 4. The magnetic resonance imaging apparatus according to claim 1, wherein the control circuits are configured to effect: obtaining a background phase φback attributed to the static portion or the normal portion for each of the pixel positions,calculating a corrected phase φcor for each pixel position as a phase in which the background phase φback is excluded from the phase of the magnetization vector detected for each pixel position, andcorrecting the pixel value ... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | A61B005/055
US08551027 | B2 | 2013-10-08 | Assist device | 2009-06-15 | In the assist device, the output of an actuator is adjusted so that it becomes lower as a state of charge of a battery exceeds a first threshold value, and also so that it becomes lower as the state of charge of the battery becomes lower. The output of the actuator is controlled on the basis of the state of charge of the battery, so that it is possible to make the agent recognize the change in the state of charge of the battery by changing the output of the actuator, and as a result, to avoid the situation where the agent feels strange with respect to stopping the operation of the assist device following thereto. | 1. An assist device including a power source, an actuator, and a control device for controlling an operation of the actuator by adjusting a supplied electricity from the power source to the actuator, for assisting a motion of or relieving a body burden of an agent by acting an output generated by the operation of the actuator to a body of the agent via a contact element; wherein the control device comprises:a state monitoring means which measures a state of charge of the power source or an available electricity thereof;a first control means which adjusts the output of the actuator, on condition that the state of charge of the power source or the available electricity thereof is equal to or less than a first threshold value, so that a motion assist rate of the assist device becomes lower than in a case where the state of charge of the power source or the available electricity thereof exceeds the first threshold value and also decreases by a greater degree as the state of charge of the power source or the available electricity thereof becomes lower; anda second control means which stops the operation of the actuator, on condition that the state of charge of the power source or the available electricity thereof is equal to or less than a second threshold value which is smaller than the first threshold value,wherein the first control means decreases the motion assist rate of the assist device by shortening a first period in which the output of the actuator is relatively high among a motion period of the agent, and on the other hand extending a second period in which the output of the actuator is lower than in the first period,wherein the assist device comprises a motor as the actuator, andwherein the first control means breaks the movement of the agent by controlling the output of the motor during the second period to be in reverse polarity to the output of the motor during the first period, so as to generate a regenerative electric power by the motor. 2. The assist device according to claim 1, wherein in the situation where the state of charge of the power source or the available electricity thereof is equal to or less than the first threshold value, the first control means decreases the motion assist rate of the assist device by decreasing the output of the actuator during a second period other than a first period among a motion period of the agent than in the situation where the state of charge of the power source or the available electricity thereof exceeds the first threshold value. 3. The assist device according to claim 2, wherein the first control means adjusts a motion state variable based on either one of or both of an oscillation and an angular velocity of a periodical motion of the agent, by adjusting a distribution of the first period and the second period in the motion period of the agent. 4. The assist device according to claim 3, wherein the assist device is configured to assist a walking motion of the agent by assisting a cyclic m... | Honda Motor Co., Ltd. | A61F005/37 A61F005/00 A61B019/00 A63B024/00 A63B021/005 A61H007/00 A61H001/02 A61H005/00
US08473024 | B2 | 2013-06-25 | Flexible headset for sensing brain electrical activity | 2008-08-12 | A headset for detecting brain electrical activity is disclosed. The headset may have a flexible band having a first end and a second end. The flexible band may have at least one expansible region permitting a distance between the first end and the second end to selectably vary. The headset may also have flexible circuitry in the flexible band. The flexible circuitry may be operatively connected to at least one sensor configured to sense brain electrical activity. The headset may further have a stimulus emitter coupled to the flexible band. | 1. A device for detecting brain electrical activity of a subject, the device comprising: a base unit; anda headset that communicates with the base unit, the headset including: a flexible band having a first end, a second end, and a center portion;a first and a second securing device respectively adjacent to the first and second ends of the flexible band; wherein the band includes a first expansible region between the first end and the center portion, and a second expansible region between the second end and the center portion, and wherein the first and the second expansible regions are configured to adaptively stretch or contract to accommodate different forehead sizes when the first and the second securing devices are anchored to the subject's ears;flexible circuitry in the band, the flexible circuitry including at least one sensor; anda stimulus emitter coupled to the band; and further wherein the headset is configured to automatically place the at least one sensor in an accurate position on the forehead when the first and the second securing devices are anchored to the subject's ears. 2. The device of claim 1, wherein the stimulus emitter includes one or more piezoelectric transducers. 3. The device of claim 1, wherein the first and the second expansible regions each include a flexure or corrugation in the band. 4. The device of claim 1, wherein the at least one sensor includes a right temple region sensor, a right forehead sensor, an upper middle forehead sensor, a lower middle forehead sensor, a left forehead sensor, and a left temple region sensor. 5. The device of claim 4, wherein the at least one sensor includes at least one of a right mastoid sensor, a left mastoid sensor, a right earlobe sensor, or a left earlobe sensor. 6. The device of claim 1, wherein the base unit includes: a pulse oximeter;a user interface configured to receive user commands;software and hardware configured to analyze brain electrical activity data received from the at least one sensor; anda display configured to display results from the data analysis. 7. The device of claim 1, wherein the base unit is a handheld unit. 8. The device of claim 1, wherein the base unit is configured to communicate with at least one of an external computer, a printer, or an external memory. 9. The device of claim 1, wherein the headset includes a chip with a unique identifier. 10. The device of claim 1, wherein the base unit is integrated into the headset, and wherein the base unit and the headset are powered by an onboard battery. 11. A headset for detecting brain electrical activity of a patient, comprising: a flexible band having a first end, a second end, and a center portion;a first sensor coupled to the center portion and configured to sense brain electrical signals;a second sensor coupled adjacent to one of the first end or the second end and configured to sense brain electrical signals; andsecuring devices located at the first end and the second end; andflexible circu... | Brainscope Company, Inc. | A61B005/0478 A61B005/0484
US08489177 | B2 | 2013-07-16 | Fiducial marker and method for gamma guided stereotactic localization | 2008-07-16 | A fiducial marker for use in a gamma-guided stereotactic localization system for imaging a suspected cancer and guiding a physician in the removal of tissue samples for biopsy. The fiducial marker includes a fiducial source that can be accurately located in a positioning system and used to correlate the location of the positioning system with the detector and therefore the region-of-interest. The fiducial can be made radioactive such that it can be seen by the gamma camera. The fiducial marker enables the accurate positioning of other hardware in proximity to the object to be viewed. | 1. A fiducial marker in combination with a gamma detector for determining the location of a region-of-interest in a stationary object to be imaged comprising: a gamma guided positioning system located adjacent to and held rigidly in place on one side of said object to be imaged;a stereotactic imaging system mounted on an opposite side of said object to be imaged;said gamma guided positioning system including a grid localization plate having a plurality of grid openings therein;a gamma emitting fiducial source rigidly mounted in said gamma guided positioning system such that said fiducial source and said region-of-interest can be imaged simultaneously with said stereotactic imaging system;said stereotactic imaging system including a gamma camera having a gamma camera crystal and two oppositely viewing slant-hole collimators having parallel slant holes, wherein the collimators are coplanar to each other, configured to be positionable in a first image position and a second image position, and rigidly joined end to end thereby enabling said stereotactic imaging system to provide simultaneous views of a mutual overlap volume within said object;an algorithm for correlating the location of said gamma guided positioning system with said gamma camera and providing the location of said region-of-interest;said oppositely viewing slant-hole collimators including a left side slant-hole collimator viewing at a 20 degree angle to the right and a right side slant-hole collimator viewing at a 20 degree angle to the left;an arrangement for sliding said slant-hole collimators with respect to and in close proximity to said stationary object between the first image position and the second image position;said fiducial source including a radioactive substance having a specific activity of no less than 1 microcurie and no greater than 25 microcurie; andsaid fiducial source enabling accurate manipulation of said positioning system to support and guide diagnostic and surgical equipment to said region-of-interest. 2. The combination of claim 1 wherein said positioning system includes a biopsy needle guide and an obturator positioned above said grid localization plate; anda gamma emitting marker source in said obturator. 3. The combination of claim 2 wherein said grid openings in said grid localization plate arranged in rows and columns; andsaid grid localization plate including a cavity for accepting said fiducial source. 4. The combination of claim 1 wherein said gamma emitting fiducial source is Ce-139. 5. The combination of claim 1 wherein said gamma emitting fiducial source is an isotope having a gamma ray energy of 160 keV. 6. A method of locating and placing equipment at a region-of-interest in a stationary object to be imaged with a stereotactic imaging system including the steps of: providing a container and a positioning system, said positioning system including a grid localization plate having a plurality of grid openings therein, said grid localization plat... | Dilon Technologies, Inc. | A61B006/02
US09687171 | B2 | 2017-06-27 | Magnetic resonance imaging apparatus and magnetic resonance imaging method | 2009-08-11 | A magnetic resonance imaging apparatus includes a trigger generating unit, a blood flow image generating unit and a control unit. The trigger generating unit acquires blood flow information of an object and generates a trigger based on the blood flow information. The blood flow image generating unit acquires imaging data with using the trigger and generates blood flow image data. The control unit controls so as to repeatedly perform a probe sequence for acquiring the blood flow information and an imaging sequence for acquiring the imaging data alternately. | 1. A magnetic resonance imaging (MRI) apparatus comprising: static and gradient magnetic field generators, at least one radio frequency (RF) coil configured to be coupled to an imaging volume, RF transmitter and receiver circuits coupled to the RF coil and at least one programmable computer connected to control said gradient magnetic field generator, said RF transmitter and receiver circuits, the at least one programmable computer being configured toprovide a probe sequence and acquire blood flow information of an object within an imaging scan to acquire magnetic resonance signals from the object and to generate a trigger for an imaging sequence, also within said imaging scan, using the blood flow information acquired in the probe sequence, wherein a data acquisition region of the probe sequence is set to a local region including at least a target blood vessel, and the blood flow information is acquired from magnetic resonance signals obtained from the target blood vessel;acquire imaging data from the object during said imaging sequence initiated by the trigger and to generate blood flow image data using the acquired imaging data; and repeatedly and alternately perform, during said imaging scan, (a) said probe sequence acquiring the blood flow information and (b) the imaging sequence acquiring the imaging data,wherein the local region to which the probe sequence is applied is located downstream of a region to which the imaging sequence is applied. 2. The magnetic resonance imaging apparatus of claim 1, wherein the at least one programmable computer being further configured to acquire a signal from blood flow in a target blood vessel, and an absolute value of intensity of the signal or a luminance value of image data generated from the signal as the blood flow information. 3. The magnetic resonance imaging apparatus of claim 1, wherein the at least one programmable computer being further configured to acquire a velocity of blood flow in a target blood vessel as the blood flow information. 4. The magnetic resonance imaging apparatus of claim 1, wherein the at least one programmable computer being further configured to acquire a temporal phase shift amount of a signal from blood flow in a target blood vessel as the blood flow information. 5. The magnetic resonance imaging apparatus of claim 4, wherein the at least one programmable computer being further configured to acquire a difference in phase between a signal corresponding to each time phase and a reference signal as the temporal phase shift amount. 6. The magnetic resonance imaging apparatus of claim 4, wherein the at least one programmable computer being further configured to acquire a difference in phase between signals adjacent to each other in time as the temporal phase shift amount. 7. The magnetic resonance imaging apparatus of claim 1, wherein the at least one programmable computer being further configured to acquire the blood flow information by NMR RF (nuclear magnetic resonance ... | TOSHIBA MEDICAL SYSTEMS CORPORATION | A61B005/05 A61B005/055 A61B005/00 G01R033/563 A61B005/026 G01R033/3415 G01R033/567
US11000484 | B2 | 2021-05-11 | Methods for diagnosing and treating a heart condition in a patient | 2009-07-24 | Various methods for diagnosing and treating heart conditions are disclosed herein. According to one method, nitric oxide is delivered to a patient's lungs when the patient has one or more symptoms of a heart condition. The patient is diagnosed as suffering from a heart condition when a symptom of a heart condition subsides after nitric oxide delivery. | 1. A method, comprising: identifying a chest pain consistent with a myocardial infarction in a patient;continuously delivering at least 20 ppm of nitric oxide to the patient's lungs for 5 to 10 seconds via a health care kit, the health care kit including a liquid N2O4 source, connected to an air pump supplying nitric oxide gas by forcing air over the liquid N2O4 source to introduce gaseous NO2 into the air so that the air containing the introduced gaseous NO2 contacts a surface active material that converts the gaseous NO2 in the air to gaseous nitric oxide, and a cannula which delivers the gaseous nitric oxide to the patient's lungs in the absence of NO2;determining that the chest pain abated while the nitric oxide is being delivered to the patient's lungsceasing delivery of nitric oxide to the patient; andidentifying a return of the chest pain after ceasing delivery of nitric oxide to the patient's lungs, the patient being treated for the myocardial infarction based on the chest pain abating while the nitric oxide is being delivered and returning after ceasing delivery of the nitric oxide. 2. The method of claim 1, wherein nitric oxide is delivered with air or oxygen enriched air. 3. The method of claim 1, wherein the chest pain is accompanied by shortness of breath, nausea, sweating, lightheadedness, or any combination thereof. 4. The method of claim 1, wherein the nitric oxide is delivered during a first time period before identifying the return of the chest pain, the method further comprising: delivering nitric oxide during a second time period after identifying the return of the chest pain. 5. The method of claim 1, wherein determining that the chest pain abated occurs is less than 100 seconds after delivering nitric oxide begins. 6. The method of claim 1, wherein identifying the return of the chest pain occurs less than 100 seconds after ceasing delivery of nitric oxide. 7. The method of claim 1, wherein a period of time between identifying a chest pain consistent with myocardial infarction and an initiation of the delivery of nitric oxide is less than 10 hours. | Vero Biotech LLC | A61K031/04 G01N033/68
US20100004564 | A1 | 2010-01-07 | WOUND MEASURING DEVICE | 2006-07-25 | This invention relates to a wound measuring device, comprising a transparent, flexible and bendable sheet (1) with at least a first edge (3), a second edge (4) and a third edge (5) and a cross-ruled field (7) located on the transparent sheet (1), wherein at least adjacent one edge (5) there is a ruler (12A) and the cross-ruled field (7) comprises a number of equally sized identically shaped fields (13, 14), wherein at least one corner (6) formed by two of said edges (3, 5) form an angle (α) which is less than 90°. | 1-13. (canceled) 14. Wound measuring device comprising: a transparent, flexible and bendable sheet having at least a first edge, a second edge and a third edge, and a cross-ruled field located on the transparent sheet, wherein at least adjacent one edge there is a ruler and the cross-ruled field comprising a number of equally sized identically shaped fields, at least one corner formed by two of said edges forms an angle (α) which is less than 90°. 15. Wound measuring device according to claim 14, wherein said shaped fields are arranged to form at least some identical groups and that within each group there is at least one field that is visually distinguishable from a neighbouring field. 16. Wound measuring device according to claim 14, further comprising a fourth edge. 17. Wound measuring device according to claim 14, wherein at least one edge of said cross-ruled field is in the form of a second ruler. 18. Wound measuring device according to claim 14, wherein said cross-ruled field comprises a first cross-ruled edge, a second cross-ruled edge, a third cross-ruled edge and a forth cross-ruled edge. 19. Wound measuring device according to claim 14, wherein said angle (α) is greater than, or equal to 10°. 20. Wound measuring device according to claim 14, wherein said angle (α) is between 10° and 80°. 21. Wound measuring device according to claim 14, wherein said angle (α) is about 60°. 22. Wound measuring device according to claim 14, wherein the sheet has resilient, polymeric properties. 23. Wound measuring device according to claim 14, wherein said fields form at least one selected from the group consisting of rectangles, triangles, and hexagons. 24. Wound measuring device according to claim 15, wherein the groups form squares. 25. Wound measuring device according to claim 24, wherein said fields form a number of smaller squares within each group. 26. Wound measuring device according to claim 24, wherein said groups each contain at least four of said fields. 27. Wound measuring device according to claim 24, wherein at least one group equals to one cm2 or one in2. 28. Wound measuring device according to claim 14, wherein at least one of said edges is smaller than 80 mm. 29. Wound measuring device according to claim 14, wherein the cross-ruled field is in the form of a rectangle. 30. Wound measuring device according to claim 29, wherein a first field border is about 1.5-2.5 times the length of a perpendicularly extending second border. 31. Method for the measurement of wounds by using a wound measuring device comprising a transparent, flexible and bendable sheet having at least a first edge, a second edge and a third edge, and a cross-ruled field located on the transparent sheet, wherein at least adjacent one edge there is a ruler and the cross-ruled field comprising a number of equally sized identically shaped fields, at least one corner formed by two of said edges forms an angle (α) which is less than 90°, the method comprising: ... | - | A61B005/107
US20100049074 | A1 | 2010-02-25 | CATEGORIZING SONGS ON A PHYSIOLOGICAL EFFECT | 2006-01-24 | A system (100) for categorizing songs on a physiological effect of the song on a user includes a storage medium (120) for storing songs, a memory (140) for storing for a predetermined physiological state of a user an associated predetermined first selection criterion, a rendering system (150) for rendering songs, and a sensor (130) for determining a biological parameter of the user representative of a physiological state of the user. A processor (110) tests a selected song for a physiological effect. The selected song is rendered a plurality of times. Each time, the sensor obtains at least one measurement, and the first criterion is used to determine a respective indicator indicating whether rendering of the song has a positive effect. Next, a second predetermined criterion is used to, based on a plurality of the determined indicators, determine a suitability of the song for bringing the user closer to the predetermined physiological state. | 1-18. (canceled) 19. A system (100) for categorizing songs based on a physiological effect of the song on a user; the system including: a storage medium (120) for storing a plurality of songs;a memory (140) for storing an associated predetermined first selection criterion for at least one predetermined physiological state of a user, said memory also storing for each of at least one predetermined physiological state, an associated playlist of songs for bringing the user closer to a particular predetermined physiological state;a rendering system (150) for rendering songs retrieved from the storage medium;a sensor (130) for determining a biological parameter of the user representative of a physiological state of the user; anda processor (110) configured for: repeatedly causing the selected song to be rendered,using the sensor to obtain at least one measurement during said rendering,using the first predetermined criterion to determine, in dependence on the measurement, a respective indicator indicating whether rendering of the song has a positive effect on bringing the user closer to the predetermined physiological state; andusing a second predetermined criterion to determine, based on a plurality of the determined indicators for the song, a suitability of the song for bringing the user closer to the predetermined physiological state. 20. A system as claimed in claim 19, wherein the processor is programmed to use the sensor at least twice during the rendering of the song to obtain at least two time-sequential measurements of the biological parameter; the first criterion being dependent on a change of the at least two measurements. 21. A system as claimed in claim 19, wherein the sensor is a brain-wave sensor for determining a frequency of brainwaves of the user. 22. A system as claimed in claim 19, wherein the sensor is of a contact-free type. 23. A system as claimed in claim 20, wherein the predetermined physiological state is asleep and the first criterion includes whether a brainwave frequency has decreased during rendering of the song. 24. A system as claimed in claim 20, wherein the predetermined physiological state is alert and the first criterion includes whether a brainwave frequency has increased during rendering of the song. 25. A system as claimed in claim 19, wherein the first criterion is dependent on an actual physiological state being measured during the rendering 26. A system as claimed in claim 20, wherein the first criterion includes whether the measured physiological state is the predetermined physiological state and whether no substantial change has been detected between the at least two measurements during the rendering. 27. A system as claimed in claim 19, wherein the second criterion includes that a predetermined successive number of times during the rendering a same song all respective indicators have indicated that rendering of the song has a positive effect on bringing the user closer to the predetermined physiologi... | KONINKLIJKE PHILIPS ELECTRONICS, N.V. | A61B005/0484 G06F017/00 G06F003/048
US07643867 | B2 | 2010-01-05 | Fiducial marker devices, tools, and methods | 2003-10-17 | This document discusses, among other things, fiducial marker devices, tools, and methods. One example illustrates a combined computed tomography (CT) imagable fiducial locator head, an integral bone screw, and an integral divot for receiving a positioning wand of an image-guided surgical (IGS) workstation. A further example includes a fluid/gel-absorbing coating or cover into which a magnetic resonance (MR) imagable fluid is introduced, thereby permitting both CT and MR imaging. Protective caps and collars may be used to protect the fiducial marker from mechanical impact and/or to guide the fiducial marker during affixation. A bull's-eye or other template is used to select a center of a substantially spherical fiducial marker head on an image, such as for use during patient registration. | 1. An apparatus comprising: a fiducial marker including: an imagable fiducial locator head that is locatable by an imaging system, the head being at least semispherical;a conical receptacle that is sized and shaped for engaging a locator instrument of a positioning system, the receptacle is threaded and defined within the imagable fiducial locator head and wherein the conical receptacle has a base at an edge defined by a surface of the locator head and an apex at a center of the locator head to permit access to a center of the imagable fiducial locator head; anda bone screw shaft that extends directly outward from the imagable fiducial locator head, at least a portion of the bone screw shaft is configured for being secured to a bone; anda protective cap sized and shaped for protecting the fiducial marker, wherein the protective cap has a disk-like portion with a threaded portion extending from the disk-like portion that is threadable into the receptacle and a cylindrical circumferential skirt extending from the disk-like portion in the same direction as the threaded portion. 2. The apparatus of claim 1, in which the protective cap engages the head. 3. An apparatus comprising: a fiducial marker including:an imagable fiducial locator head that is locatable by an imaging system, the head being at least semispherical;a conical receptacle that is sized and shaped for engaging a locator instrument of a positioning system, the receptacle defined within the imagable fiducial locator head and configured to permit access to a center of the imagable fiducial locator head;a bone screw shaft that extends directly outward from the imagable fiducial locator head, at least a portion of the bone screw shaft is configured for being secured to a bone; anda protective cap sized and shaped for protecting the fiducial marker;wherein the protective has a proximal disk portion, a cylindrical circumferential portion extending from the proximal disk portion, and a distal base flange extending from the cylindrical circumferential portion;wherein the cylindrical circumferential portion extends a distance from the proximal disk portion such that the proximal disk portion extends over the imagable fiducial locator head and the distal base flange;an adhesive portion positioned at the distal base flange for adhering the distal base flange to a scalp. 4. An apparatus comprising: a fiducial marker including:a substantially spherical imagable fiducial locator head that is locatable by an imaging system, the imagable fiducial locator head including at least one slot extending through a side wall of the imagable fiducial locator head to be engaged by a driving instrument for driving the fiducial marker to secure it to a bone;a conical receptacle that is sized and shaped for engaging a locator instrument of a positioning system, the receptacle integrated with the imagable fiducial locator head and configured to permit access to a center of the imagable fiducial locator head;a bone... | Medtronic, Inc. | A61B005/05
US20100022854 | A1 | 2010-01-28 | Implantable Chip Medical Diagnostic Device for Bodily Fluids | 2009-07-16 | An implantable microchip that is attached to a source of bodily fluids such as a vein, capillary, small artery or other fluid source such as lymph fluid or urine where the fluid flows through the microchip. The microchip can contain a micro-laboratory with reagent sources and micro-canal test chambers. The microchip can contain a readout mechanism where test data is command and/or readout to an external unit. Test results can be detected with an on-chip fluorescence or light detector or an external detector. | 1. An implantable microchip laboratory comprising: a microchip laboratory having around a million or more micro-canal test chambers, a plurality of micro-valves, a fluid-router and a plurality of on-chip reagents on an implantable substrate, wherein said microchip laboratory performs a plurality of diagnostic tests, each test being performed in one of said micro-canal test chambers, a particular micro-canal test chamber not being reused after a test;a fluid connection port adapted for fluid communication with at least one human bodily fluid when said microchip laboratory is implanted;a processor also on said substrate, said processor controlling said micro-canal test chambers, micro-valves, fluid-routine and reagents in order to perform said plurality of diagnostic tests and determine results from each test so run;a communications module non-invasively accessible when said microchip laboratory is implanted, wherein test results from each of said plurality of tests are communicated from said processor on said substrate to a remote communication unit;a chargeable power supply also mounted on said substrate, said power supply powering at least said processor, said power supply adapted to be non-invasively recharged. 2. The implantable microchip laboratory of claim 1 wherein said micro-canals have diameters of around 10 micrometers. 3. The implantable microchip laboratory of claim 1 wherein said micro-canals of lengths of around 0.5 mm. 4. The implantable microchip laboratory of claim 1 wherein said chargeable power supply is charged using a radio wave. 5. The implantable microchip laboratory of claim 1 wherein said chargeable power supply is charged using light. 6. The implantable microchip laboratory of claim 1 wherein said micro-chip laboratory is grafted to a capillary, vein, source of lymph fluid or source of urine. 7. The implantable microchip laboratory of claim 1 wherein said communication module contains a radio transceiver. 8. The implantable microchip laboratory of claim 1 wherein fluorescent reagents are used to read out diagnostic tests in said micro-canals. 9. The implantable microchip laboratory of claim 8 further comprising a light source mounted on said substrate behind said micro-canals. 10. the implantable microchip laboratory of claim 8 wherein said microchip laboratory is read out with a hand-held wand. 11. An implantable microchip laboratory comprising: a microchip laboratory having around a million or more micro-canal test chambers, a plurality of micro-valves, a fluid-router and a plurality of on-chip reagents on an implantable substrate, wherein said microchip laboratory performs a plurality of diagnostic tests, each test being performed in one of said micro-canal test chambers, a particular micro-canal test chamber not being reused after a test;a fluid connection port adapted for fluid communication with at least one human bodily fluid when said microchip laboratory is implanted;a processor also on said substrate, ... | - | A61B005/00
US20100010335 | A1 | 2010-01-14 | METHOD AND APPARATUS FOR DIAGNOSING CANCER USING ELECTROMAGNETIC WAVE | 2009-07-09 | A cancer diagnostic apparatus using electromagnetic waves includes: a plurality of antennas configured to transmit an electromagnetic wave signal to a human body and receive the electromagnetic wave signal from the human body; a switching unit configured to switch the antennas to a transmit mode or a receive mode; a signal generator configured to generate an electromagnetic wave signal to be transmitted to an antenna set to the transmit mode; a signal converter configured to convert an electromagnetic wave signal received from an antenna set to the receive mode into an intermediate frequency signal; and a controller including a switching controller to control the switching unit and a data processor to process a permittivity distribution image of the human body from the intermediate frequency signal. | 1. A cancer diagnostic apparatus using electromagnetic waves, comprising: a plurality of antennas configured to transmit an electromagnetic wave signal to a human body and receive the electromagnetic wave signal from the human body;a switching unit configured to switch the antennas to a transmit mode or a receive mode;a signal generator configured to generate an electromagnetic wave signal to be transmitted to an antenna set to the transmit mode;a signal converter configured to convert an electromagnetic wave signal received from an antenna set to the receive mode into an intermediate frequency signal; anda controller including a switching controller to control the switching unit and a data processor to process a permittivity distribution image of the human body from the intermediate frequency signal. 2. The cancer diagnostic apparatus of claim 1, wherein the switching unit comprises: a plurality of mode switches connected to the respective antennas to switch the antennas to the transmit mode or the receive mode;a transmit channel switch, including a input terminal connected to the signal generator and a plurality of output terminals connected to the respective mode switches; anda receive channel switch, including a plurality of input terminals connected to the mode switches and a output terminal connected to the signal converter. 3. The cancer diagnostic apparatus of claim 2, wherein the switching controller controls the switching unit to set one of the mode switches to the transmit mode in sequence one by one with the remaining mode switches set to the receive mode, and controls the receive channel switch to receive the electromagnetic wave signal from the mode switches set to the receive mode in sequence. 4. The cancer diagnostic apparatus of claim 1, wherein the signal converter comprises: a first amplifier configured to amplify the electromagnetic wave signal from the switching unit;a second amplifier configured to amplify a local oscillating signal;a mixer configured to mix the amplified electromagnetic wave signal and the amplified local oscillating signal; anda low-pass filter configured to filter the mixed signal and output the filtered signal to the data processor. 5. The cancer diagnostic apparatus of claim 4, wherein the signal converter further comprises a power amplifier configured to amplify electromagnetic wave signal from the signal generator and output the amplified electromagnetic wave signal through the switching unit. 6. A cancer diagnostic method using electromagnetic waves, comprising: setting one of antennas to a transmit mode and the remaining to a receive mode;transmitting an electromagnetic wave signal to a human body through the antenna set to the transmit mode;receiving an electromagnetic wave signal from the human body through the antennas set to the receive mode in sequence one by one; andgenerating a permittivity distribution image using the received electromagnetic wave signal. 7. The cancer diagnostic metho... | ELECTRONICS AND TELECOMMUNICATIONS RESEARCH INSTITUTE | A61B005/05
US20100016741 | A1 | 2010-01-21 | Heart rate monitor | 2009-07-02 | A waterproof heart rate monitor device, method, and system are disclosed that transmit heart rate information to a user. In operation one or more infrared sensor sits against skin of a user to measure his or her heart rate. An internal computer, or micro-processor, that is coupled to one or more sensors then calculates the number of beats/minute (b/m) that a user's heart is beating and generates output signals to an output unit such as an ear plug or via bone conduction transducer to provide a user an audio representation of his or her heart rate. | 1. A device comprising: a) a light sensor for measuring changes in light due to blood flow in the skin;b) a micro-processor for calculating heart rate based on changes in the light detected by the infrared sensor and generating output signals; andc) output means for converting the outputs to audio signals representative of the heart rate. 2. The device of claim 1, further comprising means for positioning the infrared sensor at or near the temple of the user. 3. The device of claim 2, wherein the means for positioning the infrared sensor at or near the temple of the user comprises a clip for attaching the sensor to a strap of a goggle or other support. 4. The device of claim 1, wherein the output means comprises a bone conduction transducer for transmitting audio signals through a portion of the user's head. 5. The device of claim 4, wherein the output means comprises an ear plug. 6. The device of claim 1, further comprising a second infrared sensor for measuring background noise and wherein the micro-processor generates a correction factor for calculating the heart rate based on the background noise. 7. The device of claim 1, wherein the monitor can automatically sense and measure the user's heart rate by simply having the user turn on the waterproof heart monitor during swimming. 8. The device of claim 1, further comprising a user interface for controlling modes of operation. 9. The device of claim 1, wherein the monitor is powered by an internal lithium-ion rechargeable battery and further comprising a USB port for recharging the battery. 10. The device of claim 1, wherein the monitor's battery is recharged through a USB port that can be plugged into a personal computer or wall adaptor. 11. The device of claim 1, further comprising a light sensor/memory unit for storing the history of the user's heart rate during workouts. 12. A method comprising: a) measuring changes in light through the skin of the user; b) calculating a heart rate of the user based on the change in the light detected by the light sensor; and c) generating an audio signal representative of the heart rate. 13. The method of claim 12 wherein measuring change in light through the user's skin comprises placing an infrared sensor at or near the user's skin. 14. The method of claim 12, wherein calculating the heart rate comprises measuring background noise and subtracting the background noise from light measured through the skin. 15. The method of claim 14, wherein the second infrared sensor measures background noise, and the micro-processor generates a correction factor for calculating the heart rate based on background noise. 16. The method of claim 12, further comprising generating audio signals through a transducer attached to the ear or temple via bone conduction. 17. The method of claim 12, wherein the bone conduction transducer communicates the heart rate to the user through vibrations to the user's temple bone and to the user's inner ear where sounds are tra... | - | A61B005/024
US20100036263 | A1 | 2010-02-11 | Circulatory monitoring systems and methods | 2008-08-07 | Systems and methods are described for obtaining and acting upon information indicative of circulatory health and related phenomena in human beings or other subjects. | 1. A medical or veterinary monitoring system comprising: circuitry for obtaining an indication of an apparent movement of an abnormal structure within a vasculature; andcircuitry for transmitting a notification partly based on a direction of the apparent movement and partly based on an additional indication of the abnormal structure. 2. The medical or veterinary system of claim 1, in which the circuitry for obtaining an indication of an apparent movement of an abnormal structure within a vasculature comprises: circuitry for obtaining information about an object apparently adjacent an implant. 3. The medical or veterinary system of claim 1, in which the circuitry for obtaining an indication of an apparent movement of an abnormal structure within a vasculature comprises: circuitry for obtaining data indicating that the abnormal structure has grown upstream. 4. The medical or veterinary system of claim 1, in which the circuitry for obtaining an indication of an apparent movement of an abnormal structure within a vasculature comprises: circuitry for activating an energy emitter within a proximity of the vasculature. 5. The medical or veterinary system of claim 1, in which the circuitry for obtaining an indication of an apparent movement of an abnormal structure within a vasculature comprises: circuitry for detecting one or more reflective properties of an apparent clot in the vasculature. 6. The medical or veterinary system of claim 1, in which the circuitry for transmitting a notification partly based on a direction of the apparent movement and partly based on an additional indication of the abnormal structure comprises: circuitry for causing a notification mode to be selected partly based on an indication of hemodynamic instability, partly based on the direction of the apparent movement, and partly based on an apparent position of the abnormal structure relative to an anatomical feature. 7. The medical or veterinary system of claim 1, in which the circuitry for transmitting a notification partly based on a direction of the apparent movement and partly based on an additional indication of the abnormal structure comprises: circuitry for causing a selection of a criterion partly based on an apparent position of an item in the vasculature. 8. The medical or veterinary system of claim 1, in which the circuitry for obtaining an indication of an apparent movement of an abnormal structure within a vasculature comprises: circuitry for detecting the apparent movement of the abnormal structure within the vasculature via a series of two or more ultrasound images. 9. The medical or veterinary system of claim 1, in which the circuitry for obtaining an indication of an apparent movement of an abnormal structure within a vasculature comprises: circuitry for facilitating an indication of whether a detected item is apparently normal. 10. The medical or veterinary system of claim 1, in which the circuitry for transmitting a notification partly based on a dir... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/02
US20100039112 | A1 | 2010-02-18 | CIRCULATOR | 2008-03-26 | A circulator suitable for use in an external field, particularly of a magnetic resonance apparatus, has a planar ferrite structure. Instead of a permanent magnet, a field conducting device is provided that is designed to strengthen or weaken an external field, and/or at least one electromagnet is provided for generating an operating field. The operating field proceeds perpendicularly through the ferrite structure. | 1.-19. (canceled) 20. A circulator for use in an external magnetic field of a magnetic resonance apparatus, comprising: a flat ferrite structure; andan operating field generator located relative to said flat ferrite structure that generates an operating field proceeding perpendicularly through said ferrite structure, said operating field generator being selected from the group consisting of a field conductor device that amplifies or attenuates said external magnetic field, and at least one electromagnet. 21. A circulator as claimed in claim 20 comprising homogenization elements located at a top and bottom of said ferrite structure that homogenize and perpendicularly align said operating field in said ferrite structure. 22. A circulator as claimed in claim 20 comprising a shield housing that surrounds said ferrite structure and conducts said external magnetic field around said ferrite structure. 23. A circulator as claimed in claim 22 wherein said operating field generator is an electromagnet, and wherein said shield housing forms at least a portion of a magnetic circuit that closes field lines of a magnetic field generated by said electromagnet. 24. A circulator as claimed in claim 23 comprising a field conductor element, forming a portion of said magnetic circuit, located at a side of said ferrite structure opposite said electromagnet, or located between said electromagnet and said shield housing. 25. A circulator as claimed in claim 20 wherein said operating field generator is an electromagnet, and comprising a yolk that closes field lines of a magnetic field generated by said electromagnet. 26. A circulator as claimed in claim 20 wherein said operating field generator is a field conductor device, and wherein said field conductor device comprises at least one field conductor element located coaxially relative to said ferrite structure that amplifies or attenuates said external magnetic field in said ferrite structure, said field conductor element being at least partially conically shaped. 27. A circulator as claimed in claim 20 comprising a magnetic circuit that closes field lines of a magnetic field generated by said operating field generator within an air gap, and wherein a width of said air gap is adjustable. 28. A circulator as claimed in claim 27 comprising a magnetic field sensor that measures said external magnetic field and generates a sensor output signal corresponding thereto, and wherein said magnetic circuit is supplied with said sensor output signal and adjusts the width of said air gap dependent on said sensor output signal. 29. A circulator as claimed in claim 27 comprising a temperature sensor that senses a temperature of an environment in which said ferrite structure is located, and generates a sensor output signal corresponding thereto, said magnetic circuit being supplied with said sensor output signal and adjusting the width of said air gap dependent thereon. 30. A circulator as claimed in claim 20 comprising a di... | - | A61B005/05 H01P001/38
US07645241 | B2 | 2010-01-12 | Device for sampling bodily fluids | 2005-04-20 | A bodily fluid sampling device comprising a lancet which can be advanced against a skin-piercing site to produce a droplet of bodily fluid. A transport medium is positioned sufficiently close to the lancing medium to retain a droplet of bodily fluid therebetween by capillary action. One of the two mediums has hydrophobic material on its surface and the other hydrophilic. Various mechanisms are provided for relatively displacing the lancing and transport medium to move the droplet of bodily fluid away from the skin-piercing end. Furthermore, multiple lancet assemblies may be provided in a barrel or disk configuration and indexed to the displacing mechanisms for sampling a skin site. | 1. Apparatus for sampling bodily fluids comprising: an elongated lancing medium having a skin piercing end and a stop, said skin piercing end adapted to be displaced against a skin site for making an incision and producing a droplet of bodily fluid from the skin site;a transport medium with a first stepped portion and a second stepped portion, said transport medium being sufficiently close to said lancing medium to retain the droplet of bodily fluid therebetween by capillary action, said stop being adapted to selectably contact one of the stepped portions while making an incision, the distance said skin piercing end extends beyond said transport medium with the stop contacting the first stepped portion being different from the distance said skin piercing end extends beyond said transport medium with the stop contacting the second stepped portion; andan actuating device connected to said lancing medium and said transport medium, wherein said actuating device displaces said lancing medium and said transport medium away from said skin site to move the droplet of bodily fluid retained by said transport medium and said lancing medium away from said skin site. 2. Apparatus as claimed in claim 1 wherein at least one of said lancing medium and said transport medium is hydrophilic and the other is hydrophobic. 3. Apparatus as claimed in claim 2 wherein said transport medium is formed from hydrophilic material. 4. Apparatus as claimed in claim 2 wherein said transport medium is coated with hydrophilic material. 5. Apparatus as claimed in claim 2 wherein said transport medium is coated with polyamide hydrophilic material. 6. Apparatus as claimed in claim 2 wherein said transport medium is formed from polyamide hydrophilic material. 7. Apparatus as claimed in claim 1 wherein said transport medium comprises a discrete element displaceable relative to said lacing medium. 8. Apparatus as claimed in claim 7 wherein said lancing medium is circular in cross section and said transport medium discrete element is a half-round element embracing a portion of the circumference around said lancing medium. 9. Apparatus as claimed in claim 7 wherein said lancing medium comprises a needle with a circular cross section and said transport medium discrete element comprises an element having an opening there through embracing the circumference of said needle and through which said lancing medium extends. 10. Apparatus as claimed in claim 1, further comprising: a rotating device connected to said lancing medium, said rotating device adapted to rotate said lancing medium relative to said transport medium to align the stop with different steps of said stepped portion and vary the distance said skin piercing end extends past said transport medium while making an incision. 11. Apparatus as claimed in claim 1, further comprising an actuating device connected to said lancing medium and said transport medium, wherein said actuating device displaces said lancing medium and sai... | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00 A61B017/14 A61B017/32
US07641611 | B2 | 2010-01-05 | Method and apparatus for remote blood alcohol monitoring | 2006-04-25 | A tamper proof transdermal alcohol content monitoring device is made up of analog and digital sides which are securely attached to the human subject to be monitored. The analog side has a sampling circuit which draws a measured insensible skin perspiration sample from the skin of the subject and measured with an electrochemical fuel cell. A distance measurement of the device from the skin of the subject and temperature of the sample are monitored along with the transdermal alcohol content, and converted to digital signals which are transmitted to a modem when the monitor is in proximity to the modem. The signals are stored in the modem and uploaded to a central monitoring station. Automatic alerts may be sent from the central monitoring station to a supervising agency. The supervising agency may also access the information through secured dedicated websites via the Internet. | 1. A monitor network for use with a modem and a continuous remote blood alcohol monitoring device adapted to be attached to a human subject, the monitor network comprising: a communication server, wherein said communication server receives raw data over a communication link from the modem, wherein said raw data is gathered by the continuous remote blood alcohol monitoring device adapted to be attached to the subject and transmitted to the modem;a situation analyzer connectable to said communication server, wherein said situation analyzer parses said raw data through a predetermined set of rules into a plurality of messages;a workflow instructions connectable to said situation analyzer, wherein said situation analyzer queries said workflow instructions and a historical data relating to similar said plurality of messages for an action to be applied to each of said plurality of messages;a supervising agency/subject database connectable to said situation analyzer for storing said raw data, for storing information on the human subject, a supervising agency, and a supervising person monitoring the human subject, and for storing a predetermined notification method by said supervising person; anda notification server connectable to said situation analyzer, wherein when said situation analyzer determines that said action to be applied to one of said plurality of messages requires an immediate notification of said supervising person, said situation analyzer will communicate said one of said plurality of messages to said notification server, and further wherein said notification server queries said supervising agency/subject database to determine said predetermined notification method, and executes said predetermined notification method to communicate one of said plurality of messages to said supervising person. 2. The monitor network according to claim 1 further comprising: an alert manager connectable to said situation analyzer and to said workflow instructions, wherein said situation analyzer passes each of said plurality of messages to said alert manager, and further wherein said alert manager categorizes each of said plurality of messages into one of a plurality of alert management categories for presentation to a monitoring person at a monitoring station. 3. The monitor network according to claim 2 wherein said plurality of alert management categories further comprises: a review/report;a take action;a snooze alert;a log action required; anda change the status of alert. 4. The monitor network according to claim 1 wherein said supervising agency/subject database further comprises: a call center data for a supervising agency associated with the human subject;a treatment center information for a treatment center associated with said supervising agency;a historical record of all previously received raw data; anda patient data for the human subject, said patient data further comprising a monitoring schedule, a communication schedule, a device information... | Alcohol Monitoring Systems, Inc. | A61B005/00
US20100042012 | A1 | 2010-02-18 | Diagnostic device for remote sensing and transmitting biophysiological signals | 2008-08-15 | A diatrophic, bio-physiological interface is self-contained with onboard intensification, filtering, and signal processing and is wirelessly enabled (idio-electrode), with multiple sensory system for bio-physiological measurements, described herein utilizes spatially resolved potential profiles from a cluster of mini electrodes to form constituent sets comprising mini sensorial electrodes. The sets of sub electrodes containing the clusters are jointly optimized to attain measurable gradient of some diagnostic value. The present invention provides a distinct lead-free single electrode that is rotationally invariant with onboard Digital Signal Processor for arrhythmia detection, source encoding, and passive and active wireless transmission. Additionally, in one aspect of the present invention the lead-free idio-electrode bio-physiological adapter allows for utmost clinical operational freedom and dramatically obviates the needs for leads of any length that invariably encumber the acquisition and performance of electrocardiogram recordings of any sort. | 1. A macro-electrode device for remote sensing of a biophysiological signal comprising; a substrate, said substrate comprising a plurality of sub-electrodes, said substrate forming one end of said macro-electrode;a power source, said power source is removably coupled to said substrate; and,a processing unit, said processing unit removably coupled to said power source;wherein said substrate, power source and processing unit form an integrated, unitary device. 2. The macro-electrode of claim 1, wherein at least one sub-electrode is a receiver and at least one sub-electrode is an explorer. 3. The macro-electrode of claim 2, wherein at least one sub-electrode is a ground sub-electrode. 4. The macro-electrode of claim 1, wherein each sub-electrode is connected to the power source. 5. The macro-electrode of claim 1, wherein the power source is a battery. 6. The macro-electrode of claim 6, wherein the battery is rechargeable. 7. The macro-electrode of claim 1, wherein the power source has a power connection to the processing unit and a data transfer connection from each sub-electrode to the processing unit. 8. The macro-electrode of claim 1, wherein the processing unit comprises a means for acquiring data, a means for optimizing the biophysiological signal, a means for detecting an anomaly in the biophysiological signal, a means for transmitting the biophysiological signal, a means for storing data. 9. The macro-electrode of claim 8, wherein the processing unit further comprises means for transmitting and receiving speech. 10. The macro-electrode of claim 8, wherein at least two macro-electrode acquire a biophysiological signal and one macro-electrode is the master-electrode and the remaining macro-electrodes are slave-electrodes. 11. The macro-electrode of claim 10, wherein the biophysiological signals acquired are synchronized and the slave-electrode transmits data to the master-electrode and the master-electrode transmits the synchronized signal. 12. The macro-electrode of claim 1, wherein the biophysiological source is selected from the group consisting of skeletal muscle tissue, brain tissue, the eye, neurological tissue, nerve tissue, heart muscle, exposed brain tissue and epithelium tissue. 13. The macro-electrode of claim 1, wherein the substrate comprises a circuit board containing an amplifier. 14. A method of remote sensing of a biophysiological signal with a macro-electrode comprising the steps of: acquiring the biophysiological signal;filtering the biophysiological signal;selecting the permutation of sub-electrodes that optimizes the filtered biophysiological signal wherein the optimized signal results in a baseline signal; and,wirelessly transmitting the baseline signal to a receiver. 15. The method of claim 14, wherein the biophysiological signal is acquired from the group consisting of skeletal muscle tissue, brain tissue, the eye, neurological tissue, nerve tissue, heart muscle, exposed brain tissue and epithelium tissue. ... | - | A61B005/04
US20100010362 | A1 | 2010-01-14 | REUSEABLE SNORE/AIR FLOW SENSOR | 2009-09-15 | A sensor for use with a polysomnograph in a sleep lab setting is made reusable by laminating a PVDF film and associated lead contacts within a flexible, moisture-impervious plastic envelope that is hermetically sealed about its periphery. Lead terminals within the envelope are adhered to the metalized surfaces of the PVDP film using a conductive adhesive which inhibits dislodgement of the leads from the sensor even with rough handling and cleaning. | 1-6. (canceled) 7. A method of making a reusable snore/air flow sensor comprising the steps of: (a) providing a film of material exhibiting pyroelectric and piezoelectric properties;(b) metalizing opposed major surfaces of said film;(c) cutting the metalized film of step (b) to a desired, predetermined shape configuration;(d) providing a pair of lead wires, each having a metal tab conductively secured to one end of said wires;(e) affixing the metal tabs individually to opposite sides of the metalized film with a conductive adhesive;(f) overlaying the assembly of step (e) on said opposite sides with a double-sided adhesive tape;(g) adhering a layer of moisture impervious, flexible plastic to the double-sided adhesive tape on each of the opposite sides of the assembly of step (f) where the layers of plastic have the same shape configuration as the metalized film; and(h) bonding perimeter portions of the layers of flexible plastic one another to create a moisture impervious seal. 8. The method as in claim 7 wherein the perimeter portions are bonded in a thermal process. 9. The method as in claim 7 wherein the perimeter portions are bonded in a chemical process. 10. The method as in claim 7 wherein the film is PVDF. 11. The method as in claim 10 wherein the predetermined shape configuration is generally Y-shaped. 12. The method of claim 10 wherein the predetermined shape configuration is generally rectangular, but with one rounded end. | DYMEDIX CORPORATION | A61B005/087
US20100022850 | A1 | 2010-01-28 | TRANSDUCER ARRAY FOR SENSING PHYSIOLOGICAL INFORMATION | 2008-07-28 | A physiologic parameter transducer array for a subject support includes at least one transducer configured for placement between the subject support and a subject being supported by the subject support. The at least one transducer senses an event indicative of a physiologic parameter of the subject that corresponds to the event. A signal processing device, in electrical communication with the at least one transducer, generates a signal indicative of the physiologic parameter. An identification component stores a unique identifier of the subject. The unique identifier associates the generated signal with the subject. A display component displays indicia indicative of the event. | 1. A physiologic parameter transducer array for a subject support, comprising: at least one transducer configured for placement between the subject support and a subject being supported by the subject support, wherein the at least one transducer senses an event indicative of a physiologic parameter of the subject that corresponds to the event;a signal processing device, in electrical communication with the at least one transducer, that generates a signal indicative of the physiologic parameter;an identification component that stores a unique identifier of the subject, wherein the unique identifier associates the generated signal with the subject; anda display component that displays indicia indicative of the event. 2. The transducer array of claim 1, wherein the physiologic parameter includes one of heart rate, respiration rate, respiration intensity, temperature, position, a pressure point, activity, weight and wetness 3. The transducer array of claim 1, wherein the event results in a change in an electrical characteristic of the transducer array. 4. The transducer array of claim 3, wherein the transducer array includes a coil, and the electrical characteristic is an inductance of the coil which changes as a function of the a change in a length of the coil in response to a deflection induced by a subject. 5. The transducer array of claim 3, wherein the electrical characteristic is a resistance or a capacitance of the transducer array. 6. The transducer array of claim 1, wherein the identification component includes a radio frequency identification (RFID) tag, and the unique identifier is stored in and obtained from the RFID tag 7. The transducer array of claim 6, wherein the signal processor employs the unique identifier to determine if the event corresponds to the subject 8. The transducer array of claim 1, wherein the at least one transducer includes an accelerometric transducer. 9. The transducer array of claim 1, wherein the at least one transducer includes a coil with an inductance that is a function of a length of the coil, and the length of the coil is a function of the event. 10. The transducer array of claim 1, wherein at least one of the signal processing devices and the at least one transducer forms a part of the subject support. 11. The transducer array of claim 1, further including a signal dampening component disposed between the at least one transducer and the subject support, wherein the dampening component dampens signals generated by events other than the event indicative of the physiologic parameter. 12. The transducer array of claim 1, wherein the transducer array includes a material that is flexible, and transducer array flexes and stretches to accommodate different subject supports. 13. The transducer array of claim 1, wherein the at least one transducer includes at least two electrically conductive layers with a plenum chamber disposed therebetween so as to form a capacitor. 14. The transducer array of claim 1... | Analogic Corporation | A61B005/00
US20100042015 | A1 | 2010-02-18 | Whole blood capillary collection system | 2008-08-14 | A whole blood capillary collection system is described and includes an enclosure located remotely from a blood collection apparatus. A DC motor-driven vacuum pump is located within the enclosure and is operatively connected to one end of the flexible tube so as to create suction therein. The enclosure includes a switch for controlling the operation of the vacuum pump and also includes a low battery light indicator. By locating the enclosure remote from the collection apparatus, the noise level at the puncture site is reduced. The enclosure may be conveniently placed in the user's pocket or placed on a supporting surface. | 1. A whole blood capillary collection system, comprising: a portable enclosure having an outside surface and an internal compartment;a battery in said internal compartment;a DC motor in said internal compartment;a vacuum pump operatively driven by said DC motor;a switch at said outside surface of said enclosure operatively electrically connected to said battery and said DC motor;an elongated, flexible tube having intake and discharge ends;said discharge end of said tube being operatively connected to said vacuum pump which suction is created in said tube;said switch being operatively coupled to said DC motor for controlling the operation of said DC motor and said vacuum pump;an a blood capillary collection apparatus operatively connected to the intake end of said tube remote from said enclosure adapted to collect blood at a puncture site. 2. The system of claim 1 wherein said enclosure is sized to permit it to be placed in a pocket of the person collecting the blood sample. 3. The system of claim 1 wherein said enclosure is configured so as to permit said enclosure to be placed on a supporting surface remote from the puncture site. 4. The system of claim 1 wherein said enclosure defines a generally rectangular block and includes a rectangular top wall member, a rectangular bottom wall member and opposite side walls and end walls joining said top and bottom wall members. 5. The system of claim 1 further including a low battery indicator on said enclosure. 6. The system of claim 1 further including a hydrophobic filter positioned between said discharge end of said tube and said vacuum pump. | - | A61B005/151
US20100041958 | A1 | 2010-02-18 | Computational system and method for memory modification | 2009-07-01 | Systems and methods are described relating to detecting an indication of an undesired memory of an individual and presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual. | 1. A system, comprising: means for detecting an indication of an undesired memory of an individual; andmeans for presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual. 2. The system of claim 1, wherein means for detecting an indication of an undesired memory of an individual comprises: means for utilizing a web search engine. 3-4. (canceled) 5. The system of claim 1, wherein means for detecting an indication of an undesired memory of an individual comprises: means for utilizing speech recognition. 6. The system of claim 1, wherein means for detecting an indication of an undesired memory of an individual comprises: means for accepting an indication of an undesired memory from the individual. 7. The system of claim 1, wherein means for detecting an indication of an undesired memory of an individual comprises: means for accepting an indication of an undesired memory from a health care provider. 8. The system of claim 1, wherein the means for presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual comprises: means for presenting an indication of at least one prescribed memory dampening agent. 9-12. (canceled) 13. The system of claim 1, wherein the means for presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual comprises: means for presenting the indication to at least one output device. 14-15. (canceled) 16. The system of claim 1, wherein the means for presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual comprises: means for presenting the indication to at least one third party. 17-18. (canceled) 19. The system of claim 1, wherein means for presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual comprises: means for presenting an indication of at least one of a beta-blocker, scopolamine, a benzodiazepine, rohypnol, or a kinase inhibitor as the memory-dampening agent. 20. The system of claim 1, wherein means for presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual comprises: means for presenting an indication of at least one of propranolol, clonazepam, ZIP, or UO126 as the memory-dampening agent. 21. The system of claim 1, wherein means for presenting an indication of at least one memory-dampening agent at least partially based on the detecting an indication of an undesired memory of an individual comprises: means for presenting an indication of a memory-dampening agent administration by at least one of oral dosage ... | Searete LLC | A61B005/16 G06F017/30 G10L015/00 A63F009/24
US07657302 | B2 | 2010-02-02 | Guided radiation therapy system | 2003-11-25 | A system and method for accurately locating and tracking the position of a target, such as a tumor or the like, within a body. In one embodiment, the system is a target locating and monitoring system usable with a radiation delivery source that delivers selected doses of radiation to a target in a body. The system includes one or more excitable markers positionable in or near the target, an external excitation source that remotely excites the markers to produce an identifiable signal, and a plurality of sensors spaced apart in a known geometry relative to each other. A computer is coupled to the sensors and configured to use the marker measurements to identify a target isocenter within the target. The computer compares the position of the target isocenter with the location of the machine isocenter. The computer also controls movement of the patient and a patient support device so the target isocenter is coincident with the machine isocenter before and during radiation therapy. | 1. A control system for use in connection with a radiation delivery system having a radiation source that produces a radiation beam, a patient support, and a drive system that moves the patient support, the control system comprising: a wireless marker configured to be implanted in the patient at a site relative to a treatment target wherein the marker is energized or excited by an excitation source positioned exterior of the patient;sensors that obtains position information about the location and/or orientation of the marker; anda computer operatively coupled to the sensor, the computer having a computer operable medium containing instructions that cause the computer to (a) receive the position information from the sensors, (b) determine an actual location of the treatment target, (c) compute a displacement between the treatment target and a beam isocenter of the radiation beam, and (d) automatically control the radiation beam and/or move the patient support based on the computed displacement between the treatment target and the beam isocenter. 2. The control system of claim 1 wherein the instructions contained by the computer operable medium further cause the computer to compute the displacement between the treatment target and the beam isocenter in three dimensions based on the actual location of the treatment target at least every at least twelve times per minute while the sensor obtains the position information. 3. The control system of claim 1 wherein the instructions contained by the computer operable medium further cause the computer to operate the drive system to move the table so that the treatment target remains in an acceptable range of the beam isocenter. 4. The system of claim 1 wherein the instructions contained by the computer operable medium further cause the computer to operate the radiation source to interrupt the radiation beam when the computed displacement is not within an acceptable range. 5. A control system for use in connection with a radiation delivery system having a radiation source that produces a radiation beam, a patient support, and a drive system that moves the patient support, the control system comprising: a wireless marker configured to be implanted in the patient at a site relative to a treatment target wherein the marker is energized by an external excitation source;sensors that obtains position information about the location and/or orientation of the marker; anda computer operatively coupled to the sensor, the computer having a computer operable medium containing instructions that cause the computer to perform the method of (a) sensing the marker to obtain position information related to an actual position of the marker, (b) determining an actual location of the treatment target based on the sensed position information of the marker, and (c) controlling the radiation beam and/or movement of the patient automatically via a computer based on the actual position of the treatment target while sensing the mar... | Calypso Medical Technologies, Inc. | A61B005/05 A61B019/00 A61N005/10
US07666394 | B2 | 2010-02-23 | Nanoparticles having reactive ester groups covalently coupled thereto and related methods | 2008-11-05 | A particle having a core and a surface to which are covalently coupled a plurality of reactive ester groups. Methods for using the particle to prepare a particle having one or more of a plurality of a targeting agent, a diagnostic agent, or a therapeutic agent covalently coupled thereto. | 1. A particle, comprising: (a) a core having a surface; and(b) a plurality of silane moieties, wherein the silane moiety is covalently coupled to the surface, and wherein the silane moiety has the formula:wherein,R2 is selected from the group consisting of substituted and unsubstituted C1-C6 straight chain and branched alkyl and substituted and unsubstituted C6-C12 aryl;n is 1, 2, 3, 4, 5, or 6; andm is an integer from about 10 to about 1000. 2. The particle of claim 1, wherein the core comprises a magnetic material. 3. The particle of claim 1, wherein the core comprises a metal oxide selected from the group consisting of ferrous oxide, ferric oxide, silicon oxide, polycrystalline silicon oxide, aluminum oxide, germanium oxide, zinc selenide, tin dioxide, titanium dioxide, indium tin oxide, gadolinium oxide, and mixtures thereof. 4. The particle of claim 1, wherein the core comprises a material selected from the group consisting of silicon nitride, stainless steel, titanium, nickel titanium, and mixtures thereof. 5. The particle of claim 1, wherein R2 is 2,2,2-trifluoroethyl. 6. The particle of claim 1, wherein n is 3. 7. The particle of claim 1, wherein m is 10. 8. The particle of claim 1 having a diameter of from about 2 nm to about 200 nm. 9. A nanoparticle, comprising: (a) a core comprising a magnetic material and having a surface; and(b) a plurality of silane moieties, wherein each silane moiety is covalently coupled to the surface, and wherein the silane moiety has the formula:wherein,R2 is selected from the group consisting of substituted and unsubstituted C1-C6 straight chain and branched alkyl and substituted and unsubstituted C6-C12 aryl;n is 1, 2, 3, 4, 5, or 6; andm is an integer from about 10 to about 1000. 10. The nanoparticle of claim 9, wherein the core comprises a metal oxide selected from the group consisting of ferrous oxide, ferric oxide, silicon oxide, polycrystalline silicon oxide, aluminum oxide, germanium oxide, zinc selenide, tin dioxide, titanium dioxide, indium tin oxide, gadolinium oxide, and mixtures thereof. 11. The nanoparticle of claim 9, wherein R2 is 2,2,2-trifluoroethyl. 12. The nanoparticle of claim 9, wherein n is 3. 13. The nanoparticle of claim 9, wherein m is 10. 14. The nanoparticle of claim 9 having a diameter of from about 2 nm to about 200 nm. | University of Washington | A61B005/055
US07654956 | B2 | 2010-02-02 | Transcutaneous analyte sensor | 2005-03-10 | The present invention relates generally to systems and methods for measuring an analyte in a host. More particularly, the present invention relates to systems and methods for transcutaneous measurement of glucose in a host. | 1. An insertion kit, comprising: a sensor and an introducer needle, wherein the sensor and the introducer needle are configured and arranged with a shape, such that the sensor is retained within a lumen of the introducer needle by a friction fit of multiple contact points separated by non-contact points between the sensor and the lumen of the introducer needle, such that longitudinal motion of the sensor with respect to the introducer needle is minimized prior to sensor insertion; and an applicator configured to cause the needle to be inserted through a skin of a host, wherein the applicator is further configured to position at least a portion of the sensor in contact with a biological fluid of the host. 2. The insertion kit of claim 1, wherein the applicator is further configured to substantially completely withdraw the introducer needle from the host after the sensor is transcutaneously positioned in fluid contact with the biological fluid of the host. 3. The insertion kit of claim 1, wherein the applicator includes a spring biased mechanism. 4. The insertion kit of claim 1, wherein the shape is a bend or serpentine shape. 5. An analyte sensor system, comprising: an analyte sensor and an applicator comprising an introducer needle, wherein the applicator is configured to position at least a portion of the sensor in contact with a biological fluid of a host by insertion of, and subsequent withdrawal of, the introducer needle through the skin of the host, wherein the analyte sensor is configured and arranged to be held by a friction fit of the surface of the sensor with an inner lumen of the introducer needle in order to maintain the sensor longitudinally within the introducer needle prior to withdrawal of the introducer needle from the host and minimize motion of the sensor during withdrawal of the introducer needle from the host. 6. The sensor of claim 5, wherein the sensor has a slight bend or serpentine shape when the sensor is held within the introducer needle. 7. A method, comprising the steps of: longitudinally maintaining an analyte sensor in a first configuration of the sensor within an introducer needle during insertion of the introducer needle into a host, wherein the first configuration uses a frictional force caused by a shape of the sensor against a surface of the introducer needle to maintain the sensor within the introducer needle; and withdrawing the introducer needle from the host, wherein the step of withdrawing comprises releasing the sensor from the first configuration to a second configuration of the sensor that is different from the first configuration after release of the sensor from the introducer needle. 8. The method of claim 7, wherein the step of withdrawing comprises positioning the sensor to be in fluid contact with a biological fluid of a host. | DexCom, Inc. | A61B005/00
US09050045 | B2 | 2015-06-09 | Psychological state estimation device | 2008-03-05 | Disclosed is a psychological state estimation device for estimating the psychological state of a subject. A psychological state estimation device for estimating the psychological state of a subject includes bioinformation acquiring means for acquiring bioinformation of the subject, perception response acquiring means for acquiring the perception response of the subject to an external event, influence degree acquiring means for acquiring the degree of influence on a change in the psychological state of the subject on the basis of the acquired perception response, and estimating means for estimating the psychological state of the subject on the basis of the bioinformation and the degree of influence on the change in the psychological state. | 1. A psychological state estimation device for estimating the psychological state of a subject, comprising: bioinformation acquiring means for acquiring skin potential, heart rate and blood pressure as bioinformation of the subject;perception response acquiring means for acquiring a perception response, which is a response of a subject to a perception of an external event when receiving an external event having a possible impact on the psychological state of the subject, on the basis of the skin potential and heart rate;influence degree acquiring means for acquiring the degree of influence on a change in the psychological state of the subject on the basis of the perception response acquired by the perception response acquiring means; andestimating means for estimating the psychological state of the subject on the basis of the bioinformation acquired by the bioinformation acquiring means and the influence degree acquired by the influence degree acquiring means,wherein the perception response acquiring means adopts a time at which the external event is received as a start point to set an interval in which bioinformation for determining the perception response is to be extracted, and with respect to the skin potential and heart rate, extracts a skin potential and heart rate, respectively, within the interval starting from the time, to set a no response as the perception response when the skin potential within the interval is less than a reference potential, to set an orienting response as the perception response when the skin potential within the interval is equal to or greater than the reference potential and the heart rate within the interval is less than a reference heart rate, and to set a defensive response as the perception response when the skin potential within the interval is equal to or greater than the reference potential and the heart rate within the interval is equal to or greater than the reference heart rate,wherein the influence degree acquiring means sets the degree of influence on the change in the psychological state depending on the external event and a perception response coefficient depending on the perception response of either the no response, the orienting response or the defensive response acquired by the perception response acquiring means, and acquires the degree of influence on the basis of the degree of influence on the change in the psychological state depending on the external event with the perception response coefficient, andwherein the estimating means: refers to an evaluation table to set a frustration level corresponding to a heart rate range and a blood pressure range by extracting a heart rate of a preset interval from the heart rate to judge a value range of the heart rate of the interval on the basis of preset thresholds, and extracting a blood pressure of a preset interval from the blood pressure to judge a value range of the blood pressure of the interval on the basis of preset thresholds,when the degree of... | TOYOTA JIDOSHA KABUSHIKI KAISHA | G09B019/00 A61B005/00 A61B005/16 A61B005/18 A61B005/0205 A61B005/021 A61B005/024
US08128562 | B2 | 2012-03-06 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2009-10-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. In a glucose sensing system comprising a glucose sensor configured for producing an output representative of glucose concentration in a host and processing circuitry configured to process the output, a method for processing data from the output, the method comprising: generating, using a processing module, at least one calibrated data stream of glucose concentration values from the glucose sensor output;calculating, using the processing module, a range of glucose concentration values based at least in part on at least one calibrated historical data stream value, wherein the range of glucose concentration values comprises at least one limit value;replacing, using the processing module, at least one calibrated data stream value with the at least one limit value when the calibrated data stream value is outside the limit value; anddisplaying the limit value to the user of the glucose sensing system. 2. The method of claim 1, wherein the at least one limit value is based at least in part on physiology. 3. The method of claim 1, wherein the at least one limit value is defined by a cone of possible values extending from at least one historical data stream value. 4. The method of claim 3, wherein the cone is derived from one or both of physiologically acceptable glucose concentration rate and acceleration values. 5. The method of claim 1, wherein the at least one limit value is derived from a linear extrapolation. 6. The method of claim 1, comprising alerting a user of the glucose sensing system to upcoming hypoglycemic or hyperglycemic events. 7. The method of claim 1, comprising detecting a signal artifact, and reducing or replacing or replacing a portion of the calibrated data stream containing the signal artifact with the at least one limit value. 8. A glucose sensor system for monitoring a glucose concentration in a host, the system comprising: a continuous glucose sensor configured to produce a data stream;a processing module configured to: generate at least one calibrated data stream of glucose concentration values from the data stream;calculate a range of glucose concentration values based at least in part on at least one calibrated historical data stream value, wherein the range of glucose concentration values comprises at least one limit value; andreplace the calibrated data stream value with the at least one limit value when the calibrated data stream value is outside the data stream value range; anda display;wherein the system is configured to display at least one of the calibrated data stream value or the limit value to a user of the glucose sensing system. 9. The system of claim 8, wherein the processing module is housed in at least one of a sensor electronics housing or a receiver electronics housing. 10. The system of claim 8, wherein the processing module is configured to calculate the limit value is based at least in part on physiology. 11. The system of claim 8, wherein the processing module is configured to define the lim... | DexCom, Inc. | A61B005/05
US08282580 | B2 | 2012-10-09 | Data rejection for posture state analysis | 2009-04-30 | The disclosure describes techniques for data rejection for posture state analysis. The techniques may include storing posture state data sensed by a medical device for a patient, rejecting any portion of the posture state data that was stored during a session that was shorter than a session threshold, and generating posture state output for the patient based on a portion of the posture state data that was not rejected. | 1. A method comprising: storing, in a memory, posture state data sensed by a medical device for a patient during a therapy session;rejecting at least a portion of the posture state data that was stored during the therapy session if the therapy session was shorter than a session threshold; andgenerating, by a processor, posture state output for the patient based on a portion of the posture state data that was not rejected. 2. The method of claim 1, further comprising retrieving the stored posture state data from the medical device with a programmer, wherein: rejecting the portion of the posture state data comprises rejecting a portion of the retrieved posture state data in the programmer; andgenerating posture state output for the patient comprises generating posture state output for the patient based on a portion of the retrieved posture state data that was not rejected in the programmer. 3. The method of claim 1, wherein the therapy session is a time between two successive programming sessions in which the medical device received changes in therapy parameters. 4. The method of claim 1, wherein the session threshold is one of one hour, twenty four hours, one week, or one month. 5. The method of claim 1, wherein the session threshold is less than ten percent of a longest therapy session duration in which the posture state data was stored. 6. The method of claim 1, wherein rejecting the portion of the posture state data further comprises deleting the rejected portion of the posture state data from the memory. 7. The method of claim 1, wherein generating posture state output further comprises generating at least one of proportional posture information and sleep quality information. 8. The method of claim 1, further comprising delivering electrical stimulation therapy from the medical device to a patient during the therapy session while storing the posture state data. 9. A system comprising: a memory configured to store posture state data sensed by a medical device for a patient during a therapy session; anda processor configured to reject at least a portion of the posture state data when the therapy session was shorter than a session threshold, and generates posture state output for the patient based on a portion of the posture state data that was not rejected. 10. The system of claim 9, further comprising a telemetry circuit, wherein the processor is configured to retrieve the posture state data from the memory via the telemetry circuit, and wherein the processor is configured to: reject the portion of the retrieved posture state data that was stored during a therapy session when the therapy session was shorter than a session threshold in the programmer;andgenerate posture state output for the patient based on a portion of the retrieved posture state data that was not rejected in the programmer. 11. The system of claim 9, wherein the therapy session is a time between two successive programming sessions in which the medical device received... | Medtronic, Inc. | A61B005/00 A61B005/103 A61B005/117 A61N001/00
US09161714 | B2 | 2015-10-20 | Specificity of analyte detection in etalons | 2008-08-21 | A system and method is provided for detecting concentration of an analyte in a fluid. A first container includes (i) an optical cavity detection region, (ii) a reservoir for one or more modifiers of one or more optical properties of the optical cavity detection region, and (iii) a set of one or more bounding regions through which objects in the fluid can transfer into the container. The optical cavity detection region and the reservoir define separate areas of the first container. The movement of the one or more modifiers between the reservoir and the optical cavity detection region is responsive to changes in concentration of the analyte. A second container includes an optical cavity detection region, and a set of one or more bounding regions through which objects in the fluid can transfer into the container. Also provided are optical components for guiding light into the optical cavity detection regions of the first and second containers. | 1. A system for detecting concentration of an analyte in a fluid, the system comprising: a first container including (i) an optical cavity detection region having a first side and a second side opposite the first side, the first side and the second side at least partially bounded by light reflective components, and further configured to permit passage of light therethrough, (ii) a reservoir for one or more modifiers of one or more optical properties of said optical cavity detection region, the reservoir configured to define boundaries of the optical cavity detection region, and the reservoir extending substantially from the first side to the second side of the optical cavity detection region and (iii) a set of one or more bounding regions through which objects in the fluid can transfer into the container; the optical cavity detection region and the reservoir defining separate areas of the first container; the movement of said one or more modifiers between said reservoir and said optical cavity detection region being responsive to changes in concentration of said analyte;a second container including an optical cavity detection region having a first side and a second side opposite the first side, the first side and the second side at least partially bounded by light reflective components, and further configured to permit passage of light therethrough, and a set of one or more bounding regions through which objects in the fluid can transfer into the container;an illumination source located a distance external from the first container and the second container and configured to generate the light passing through the optical cavity detection region; anda detector component located a distance external from the first container and the second container on a side opposite the illumination source and configured to detect the light exiting from the optical cavity detection region. 2. A system according to claim 1, which is implantable or partially implantable in or underneath skin in contact with interstitial fluid. 3. A system according to claim 1, wherein the external illumination source and the external detector component are located in relationship to each other such that at least some of the light from the external illumination source input into the first container and the second container and then output from the first container and the second container is detected by the external detector component. 4. A system according to claim 1, wherein said illumination source is one of a UV illumination source providing light with wavelengths between 220 and 320 nanometers or a near IR illumination source providing light with wavelengths between 750 and 1500 nanometers. 5. A system according to claim 1, wherein said illumination source is one or more of a vertical cavity surface emitting laser (VCSEL), a solid state laser, a Distributed Feed Back (DFB) laser, a Distributed Bragg Reflector (DBR) laser or a resonant cavity light emitting diode (LED). 6. A sys... | Palo Alto Research Center Incorporated | A61B005/00 A61B005/1459 A61B005/145 G01N021/33 G01N021/359
US08852172 | B2 | 2014-10-07 | Ingestible electronic capsule and in vivo drug delivery or diagnostic system | 2007-11-05 | An ingestible electronic capsule (100) for introduction into the bodily lumen comprises a circuit board with at least one electronic component (4, 5, 7, 8, 9, 11, 12, 13, 14). The circuit board is formed from a flex foil (10), thereby reducing the number of components and improving the robustness and reliability of the ingestible electronic capsule (100). | 1. An ingestible electronic capsule for introduction into a bodily lumen, comprising: a housing;a continuous flex foil formed of a uniform material and having a first portion extending generally in a longitudinal direction, the flex foil further having a second portion that extends from the first portion at an angle relative to the longitudinal direction; an electrically conducting layer integrated into the second portion;a plurality of electronic components disposed directly on the flex foil and spaced from each other along the longitudinal direction on the first portion, the continuous flex foil being bent along at least one bend line disposed in the lateral direction between adjacent electronic components to form a bent flex foil and the second portion being bent relative to the first portion;a reservoir sealed relative to the first portion of the housing, the reservoir containing a drug for expulsion from the capsule via a dispensing hole; anda release mechanism movable to empty the drug through the dispensing hole, the release mechanism being electrically connected to at least one of the plurality of electronic components formed on the flex foil. 2. The ingestible electronic capsule according to claim 1, wherein the electrically conducting layer is an antenna. 3. The ingestible electronic capsule according to claim 2, wherein the antenna is wrapped around a remainder of the flex foil. 4. The ingestible electronic capsule according to claim 1, wherein the at least one electronic component is attached to the flex foil with a solder bump. 5. The ingestible electronic capsule according to claim 4, wherein at least part of remaining open space between the flex foil and the at least one electronic component is filled with a moldable resin. 6. The ingestible electronic capsule according to claim 1, wherein a power supply unit is mounted on the flex foil with a conducting adhesive. 7. The ingestible electronic capsule according to claim 1, wherein the flex foil is bended along at least one bending line. 8. The ingestible electronic capsule according to claim 1, wherein the ingestible electronic capsule comprises a system in package with a microprocessor a transceiver, a sensor and a passive electronic device. 9. The ingestible electronic capsule according to claim 1, wherein the release mechanism is a movable piston for emptying the drug reservoir, which movable piston is electrically and mechanically connected to the flex foil. 10. In vivo drug delivery or diagnostic system comprising an ingestible electronic capsule according to claim 1 and an external device arranged to communicate with the at least one electronic component. 11. The ingestible capsule of claim 1, further comprising a socket configured to receive a plug, wherein the flex foil includes a side arm electrically connecting the socket to the flex foil. 12. The ingestible capsule of claim 11, wherein the side arm is removable. 13. The ingestible capsule of claim 9, wherein t... | Medimetrics Personalized Drug Delivery | A61B005/07 A61B001/04 A61B005/00 A61B001/00
US20100004523 | A1 | 2010-01-07 | Active, radiating low frequency implantable sensor and radio tag system | 2006-04-12 | A low frequency implantable sensor and radio tag system, includes a sensor device which in turn includes: storage storing information including information identifying the device; an apparatus, coupled to the transceiver, for measuring a body condition for transmission to a reader; a transceiver, coupled to the storage, the transceiver operating at a frequency sufficiently low to operate near or within water; and an antenna, coupled to the transceiver, communicating with an external reader. | 1. A low frequency implantable sensor and radio tag system, comprising: a sensor device comprising:storage storing information including information identifying the device;an apparatus, coupled to a transceiver, for measuring a body condition for transmission to a reader;a transceiver, coupled to the storage, the transceiver using induction as a primary communication mode operating at a frequency lower than or equal to 300 kilo Hertz; andan antenna, coupled to the transceiver, communicating with an external reader. 2. The system of claim 1 further comprising a reader to be worn by a user of the implantable device outside the body of the user. 3. The system of claim 1 wherein the transceiver operates at a frequency sufficiently low to operate near steel. 4. The system of claim 1 wherein the reader comprises a loop antenna to read low frequency signals from the sensor. 5. The system of claim 4 wherein the reader loop antenna reads signals from the sensor when located in the same room as the sensor. 6. The system of claim 2 wherein the reader is connected to a network, and data from the sensor can be captured and maintained at a remote site. 7. The system of claim 6 wherein the network is a wide area network. 8. The system of claim 1 wherein the sensor comprises an apparatus for measuring physical parameters relating to the user. 9. The system of claim 8 wherein the physical parameters include at least one of a number of flexes of a joint, the angle of flexion of the joint, and statistics associated with walking. 10. The system of claim 1 wherein the sensor comprises a glucose detector. 11. The system of claim 1 wherein the sensor comprises a radiation monitor. 12. The system of claim 1 wherein the sensor comprises an electrocardiogram monitor. 13. (canceled) 14. The system of claim 1 wherein the frequency of operation is between 30 and 300 kilo Hertz. | Visible Assets, Inc. | A61B005/145 A61B005/0402 G08B001/08
US20100036219 | A1 | 2010-02-11 | Blood Sampler | 2008-12-15 | A blood sampler is adapted for penetrating epidermis of skin in order to collect blood, and includes a plate, a thin-film electrode unit, and a plurality of spaced apart penetrating members. The plate has an upper surface that is indented to form a recess adapted for containing the blood. The thin-film electrode unit is provided in the plate and has a reactive section that is exposed at a bottom of the recess. The penetrating members are provided on the reactive section of the thin-film electrode unit. Each of the penetrating members has a bottom end that is wire-bonded to the reactive section of the thin-film electrode unit, and a top end that is not higher than the upper surface of the plate. | 1. A blood sampler adapted for penetrating epidermis of skin in order to collect blood, said blood sampler comprising: a plate having an upper surface that is indented to form a recess adapted for containing the blood;a thin-film electrode unit provided in said plate and having a reactive section that is exposed at a bottom of said recess; anda plurality of spaced apart penetrating members provided on said reactive section of said thin-film electrode unit, each of said penetrating members having a bottom end that is wire-bonded to said reactive section of said thin-film electrode unit, and a top end that is not higher than said upper surface of said plate. 2. The blood sampler as claimed in claim 1, wherein said plate is made from a polymer, apiece of paper, or a metallic foil. 3. The blood sampler as claimed in claim 1, wherein said thin-film electrode unit includes a first electrode and a second electrode that is disposed apart from said first electrode, said first electrode and said second electrode being adapted to be electrically conducted by the blood. 4. The blood sampler as claimed in claim 1, wherein said penetrating members are made from gold or a gold-containing material. 5. The blood sampler as claimed in claim 1, wherein each of said penetrating members includes a base at said bottom end, and a needle extending from said base to said top end. 6. The blood sampler as claimed in claim 5, wherein adjacent ones of said needles have therebetween a smallest distance that is equal to a diameter of said base. 7. The blood sampler as claimed in claim 5, wherein said needle of each of said penetrating members has a width that ranges from 20 to 100 micrometers and a length that ranges from 500 to 2000 micrometers. 8. The blood sampler as claimed in claim 1, further comprising an enzyme that is provided in said recess and that is adapted for reacting with the blood so as to change color. | I SHOU UNIVERSITY | A61B005/1468 A61B005/15 C12Q001/25
US20100010357 | A1 | 2010-01-14 | Disposable air bag for a blood pressure measuring device and a method of making the same | 2008-07-09 | An disposable air bag for measuring blood pressure having a diaphragm and a nipple mounted therein, the diaphragm having a plurality of waves for allowing the diaphragm to expand easily out of the air bag and a clip for disconnecting and connecting the nipple to and from one end of an air hose the other end of which is connected to an electronic blood pressure measuring device.;Also disclosed is a blood pressure measuring device having ABS storing means for storing the stretching characteristics of an air bag as a function of the air volume in the air bag, AABPOA calculating means for calculating the actual pressure on the artery, and systolic/diastolic algorithm calculating means for determining the systolic and diastolic blood pressures partially based on the AABPOA calculating means. | 1. A diaphragm for an air bag for measuring blood pressure which comprises: a central portion; andat least one wave portion integrally formed with said central portion along the periphery said central portion. 2. A diaphragm as defined in claim 1, wherein: said central portion is thinner than the thickness of said wave portion. 3. A diaphragm as defined in claim 1, wherein: the thickness of said diaphragm is thinnest along a central area of said central portion. 4. A diaphragm as defined in claim 1, wherein: the thickness of said diaphragm is thinnest along a central area of said central portion and gradually increases in thickness from said central area of said central portion outwards towards said wave portion. 5. A diaphragm as defined in claim 1, wherein: the thickness of said diaphragm varies along the surface thereof, the thickness of said diaphragm gradually increasing from a central area in said central portion towards said wave portion, so that when an air bag in which said diaphragm is mounted in is inflated with air, said central area in said central portion expands outwardly first followed by the central portion around said central area followed by the unfurling of said wave portion, so that said wave portion not only allows said central portion to easily move outwards of said air bag, but also prevents said central portion from escaping laterally along a patients' arm when said central portion is pressed against an artery in said arm. 6. A diaphragm as defined in claim 1, wherein: said at least one wave portion comprises a plurality of concentric waves formed around each other, so that when an air bag in which said diaphragm is mounted in is inflated with air, said central portion expands outwardly first followed by the unfurling of said concentric waves. 7. A diaphragm as defined in claim 1, wherein: said diaphragm is formed of an elastic material. 8. A diaphragm as defined in claim 1, wherein said diaphragm is formed of rubber. 9. A diaphragm as defined in claim 1, wherein said diaphragm is formed of latex. 10. A diaphragm as defined in claim 1, wherein said diaphragm is formed of silicone. 11. A diaphragm as defined in claim 1, wherein: said central portion and said wave portion are oval in shape, the longer side of said oval central portion being longer than the distance between the radius bone and the digital tendon in a persons writs. 12. A diaphragm as defined in claim 1, wherein: said central portion and said wave portion are substantially rectangular in shape, the longer side of said rectangular central portion being longer than the distance between the radius bone and the digital tendon in a persons wrist. 13. A diaphragm as defined in claim 1, wherein: said wave portion is 5 mm high and has a pitch of 1-5 mm. 14. A diaphragm as defined in claim 1, wherein: the thickness of the thinnest part of said central portion is 0.03 mm. 15. A diaphragm as defined in claim 1, wherein said diaphragm further comprises: an ... | - | A61B005/022
US20100022910 | A1 | 2010-01-28 | RAPID ALLERGY TEST | 2008-07-25 | A flexible and disposable allergen-containing patch is provided with markings transferable to the skin. The patch may also include piercing elements to become an all-in-one system. A reusable actuator enables a user with little training or experience to administer multiple test injections with the disposable patch. A scope allows magnified viewing of a test site after the skin test is provided, the scope incorporating the patch as a releasable component. An all-in-one applicator assembly that can be bent to fit a convex engagement surface of a rolling actuator is also provided. | 1. A disposable allergy-testing patch, said patch comprising: a flexible body having a plurality of packets, each housing an allergen and at least one piercing element, said at least one piercing element being movably supported within each said packet, and said piercing element including retaining means for retaining a portion of said allergen when at least a portion of said piercing element is forced through each said packet; andat least one transferable marking material adjacent each said packet and disposed on a surface of said flexible body for marking a patient's body upon close contact therewith. 2. The patch of claim 1, wherein said patch further includes actuation means for moving said at least one piercing element at least partly out of each said packet to enable a portion of said piercing element, including said retaining means, to extend outwardly from said body. 3. The patch of claim 2, wherein each said actuation means comprises a tubular section made from an elastomeric material and is supported by at least a portion of said piercing element, each said tubular section being configured to deform upon compressive contact of said surface of said patch with the skin of a patient, thereby permitting said piercing element to move in relation to said tubular portion into direct contact with said skin. 4. The patch of claim 1, wherein said patch includes a removable portion covering said surface. 5. An allergy-testing apparatus comprising: a disposable patch comprising a flexible body having at least one packet housing an allergen;at least one transferable marking material adjacent said at least one packet and disposed on a surface of said flexible body for marking a patient's skin upon close contact therewith; andat least one piercing element movably contained within said patch and configured to pierce a portion of said at least one packet, said at least one piercing element including retaining means for retaining a portion of said allergen, wherein said apparatus is configured to advance said retaining means towards said surface of said flexible body and said at least one transferable marking material into close contact with said patient's skin. 6. The apparatus of claim 5, including a removable strip covering said surface of said flexible body. 7. An allergy-testing apparatus comprising: an actuating device comprising a compact body, an actuator mechanism and an engagement surface, said actuator mechanism being operatively connected to at least a portion of said engagement surface such that operation of said actuator mechanism causes spatial displacement of said at least a portion of said engagement surface relative to said compact body;a packet assembly comprising at least one packet housing an allergen, said packet assembly being releasably attached to said engagement surface of said actuating device; andat least one piercing element having a fluid retaining end, wherein spatial displacement of said engagement surface caused by s... | Welch Allyn, Inc. | A61B005/00
US20100030069 | A1 | 2010-02-04 | TRIPLE-MODALITY IMAGING SYSTEM | 2007-10-17 | The invention relates to a triple-modality imaging system, wherein an optical imaging detector for acquiring optical tomography imaging data, a single photon emission computered tomography (SPECT) detector for acquiring SPECT data, and an x-ray source and detector for acquiring x-ray data are arranged to acquire optical imaging data, SPECT data, and x-ray data of an imaged object simultaneously, the optical imaging detector, the SPECT detector, and the x-ray detector having spatially over-lapping fields-of-view (FOV). | 1-10. (canceled) 11. A triple-modality imaging system, wherein an optical imaging detector (1) for acquiring optical tomography imaging data, a single photon emission computed tomography (SPECT) detector (3) for acquiring SPECT-data, and an x-ray source (5) and detector (6) for acquiring x-ray data are arranged to acquire the optical tomography imaging data, the SPECT data, and the x-ray data of an imaged object simultaneously, the optical imaging detector (1), the SPECT detector (3), and the x-ray detector (6) having spatially overlapping fields-of-view (FOV) (20, 18, 19). 12. The triple-modality imaging system according to claim 11, wherein the optical imaging detector (1), the SPECT detector (3), and the x-ray source (5) and detector (6) are mounted on a common rotatable gantry (7), the gantry (7) being designed like an optical bench. 13. The triple-modality imaging system according to claim 11, wherein the imaged object is placed at a rotation axis (8) of a gantry (7) on which the optical imaging detector (1), the SPECT detector (3), and the x-ray source (5) and detector (6) are mounted, and wherein the fields-of-view (20, 18, 19) of the optical imaging detector (1), the SPECT detector (3), and the x-ray detector (6) over-lapping spatially at the rotation axis (8). 14. The triple-modality imaging system according to claim 11, wherein the optical imaging detector (1), the SPECT detector (3), the x-ray source (5) and detector (6), and the imaged object are arranged within a light-tight and gamma-ray shielded compartment (10). 15. The triple-modality imaging system according to claim 11, further comprising a light source (12) for illuminating the imaged object. 16. The triple-modality imaging system according to claim 11, wherein the optical imaging detector (1), the SPECT detector (3), the x-ray detector (6), and the x-ray source (5) are modules, which are spatially rearrangeable with respect to one another. 17. The triple-modality imaging system according to claim 11, comprising a layer (15) in front of the SPECT detector (3), which essentially reflects photons of bioluminescent and/or fluorescent markers, and which essentially transmits photons of radioactive markers. 18. A triple-modality imaging method comprising simultaneously acquiring optical tomography imaging data with an optical imaging detector, single photon emission computed tomography (SPECT) data with a SPECT detector and x-ray data with an x-ray detector of an imaged object, which is placed within spatially overlapping fields-of-view (FOV) of the optical imaging detector, the SPECT detector and the x-ray detector. 19. The method according to claim 18, comprising rotating the optical imaging detector, the SPECT detector, and the x-ray detector on a common rotatable gantry around a rotation axis, the imaged object being located at the rotation axis. 20. The method according to claim 18, including the steps of reconstructing optical images, SPECT images and x-ray images b... | DKFZ DEUTSCHES KREBSFORSCHUNGSZENTRUM | A61B005/05 G01T001/161
US20100036280 | A1 | 2010-02-11 | METHOD OF UTILISING MEASUREMENTS OF THRESHOLD OF PAIN | 2007-09-10 | A method of determining the sympathetic tone and/or level of stress and/or level or warning system sensitivity includes the steps of: measuring an applied stimulation at a threshold value of the stimulation in one or more sympathetic tone-neutral points and measuring an applied stimulation at the same threshold value in one or more sympathetic tone-dependent points. The invention further relates to a system for applying and measuring a stimulation, and the use of a system for applying and measuring a stimulation for determining the sympathetic tone including the steps of: measuring an applied stimulation at a threshold value of the stimulation at one or more sympathetic tone-neutral points and measuring an applied stimulation at the same threshold value of the stimulation at one or more sympathetic tone-dependent points. | 1-123. (canceled) 124. A method of determining the status of warning system sensitivity (WSS)/defense reaction/reflex sensitivity (DRS) in a subject, including the steps of: measuring an applied stimulation at a threshold value of the stimulation in one or more sympathetic tone-neutral points and measuring an applied stimulation at the same threshold value in one or more sympathetic tone-dependent points, or measuring an applied stimulation at a threshold value in one or more sympathetic tone-dependent points and optionally comparing said threshold value to a predetermined or pre-established calibration threshold value. 125. A method according to claim 124, wherein the threshold value of the stimulation is the stimulation's nociception threshold value. 126. A method according to claim 124, wherein the applied stimulation is provided by means of an applied mechanical stimulation, such as an applied compressive force and, an applied vibrating force, or wherein the applied stimulation is provided by means of an applied thermal stimulation, such as an applied heat or cold source, or wherein the applied stimulation is provided by means of an applied electrical stimulation, such as an alternating current or direct current, or wherein the applied stimulation is provided by means of an applied radiation, such as an applied infrared, visible and/or ultraviolet light or combined spectra thereof, or wherein. 127. A method according to claim 124, wherein the applied stimulation is provided by means of an applied chemical stimulation. 128. A method according to claim 127, wherein the applied chemical stimulation is provided by means of an applied organic or inorganic compound. 129. The method according to claim 124, wherein the applied stimulation in the sympathetic tone dependent point is direct stimulation of a at least one polymodal sensor cell and wherein the threshold value in the one or more sympathetic tone-dependent points is the firing threshold of said at least one polymodal sensor cell. 130. The method according to claim 129, wherein the direct stimulation is selected from AC stimulation, DC stimulation, stimulation with a neurotransmitter, and chemical stimulation. 131. A method according to claim 124, wherein the determination is performed by means of a system for measuring the applied stimulation. 132. A method according to claim 124, wherein the measuring of an applied stimulation at a threshold value of the stimulation in one or more sympathetic tone-neutral points is carried out anteriorly on the upper side of the clavicle and/or posteriorly on the spinal column corresponding to TH 10-11 and/or on a finger and/or on a toe. 133. A method according to claim 124, wherein the measuring of an applied stimulation at a threshold value of the stimulation is carried out in one or more sympathetic tone-dependent points at one or more locations on or in the skin which innervationally correspond(s) to the nerve supply to the heart of the sympat... | ULL METER A/S | A61B005/00
US20100022918 | A1 | 2010-01-28 | SENSING SYSTEM EMPLOYING MEDICAL MANIPULATOR AND PRESSING FORCE MEASURING DEVICE AND ITS PROGRAM | 2007-03-28 | Sensing of the pressing force of a balloon against internal body tissues is made possible using the balloon inserted into a gap in the body and pressing the internal body tissues around that gap without inserting a sensor, or the like, separately into the gap. A sensing system comprises an surgical assist manipulator provided to be able to travel by a reaction produced while it pressure-displaces the internal body tissues, a pressure sensor for measuring the pressure of fluid used to operate the surgical assist manipulator, and a pressing force measuring device for determining a pressing force applied to the internal body tissues by the surgical assist manipulator based on the measurement by the pressure sensor. The surgical assist manipulator includes a balloon which can be expanded/contracted by fluid being supplied internally, and the pressing force measuring device determines a function employing the internal pressure when the balloon is pressed as a parameter based on the measurement of fluid amount supplied to the balloon, and determines the pressing force by substituting the internal pressure at pressing measured by the pressure sensor into that function. | 1. A sensing system employing a medical manipulator which has a balloon inserted in a gap in a living body and expandable and contractible by fluid supplied thereinto, and presses an internal tissue around the gap by expansion of the balloon, the sensing system comprising: an internal pressure measuring sensor which can measure an internal pressure of the balloon; and a pressing force measuring device which measures a pressing force of the balloon to the internal tissue based on the internal pressure measured by the internal pressure measuring sensor,wherein the pressing force measuring device comprises: a function determining section which determines a function using, as a parameter, a pressing internal pressure, which is the internal pressure of the balloon when the internal tissue is pressed, based on an amount of fluid supplied into the balloon; and a pressing force operating section which calculates the pressing force by assigning, to the function, the pressing internal pressure measured by the internal pressure measuring sensor. 2. The sensing system employing the medical manipulator according to claim 1, the function determining section preliminarily stores a function for determining a predetermined correction coefficient and constant using the amount of fluid supply as a parameter therein, uses a measured value of the amount of fluid supply to determine the correction coefficient and the constant, and adds the constant to the pressing internal pressure multiplied by the correction coefficient to determine a function for determining the pressing force. 3. The sensing system employing the medical manipulator according to claim 1, the pressing force measuring device further comprises a no loading pressure operating section which uses the amount of fluid supply to calculate a no loading internal pressure of the balloon when the balloon is not in contact with the internal tissue, wherein the function determining section determines the function based on the no loading internal pressure calculated by the no loading pressure operating section. 4. A sensing system employing a medical manipulator which has a plurality of balloons bendably connected to each other and expandable and contractible by fluid supplied thereinto, and can adjust a pressing force to an internal tissue due to expansion of the balloon to advance in a gap in a living body using a reaction force from the internal tissue, the sensing system comprising: an internal pressure measuring sensor which can measure an internal pressure of the balloon; and a pressing force measuring device which measures the pressing force based on the internal pressure measured by the internal pressure measuring sensor,wherein the pressing force measuring device comprises: a function determining section which determines a function using, as a parameter, a pressing internal pressure of the balloon when the internal tissue is pressed, based on an amount of fluid supplied into the balloon; and a pressing... | WASEDA UNIVERSITY | A61B005/103
US20100016707 | A1 | 2010-01-21 | IMAGING SYSTEM | 2007-07-30 | A medical diagnostic system, the system comprising:;a) a non-ultrasound imaging module configured to generate a volumetric vascular map;;b) a controller configured to translate a 3D position co-ordinate from the map into an aiming instruction;;c) a Doppler ultrasound unit adapted to aim a transducer responsive to the aiming instruction; and;d) a registration module configured to register a position of the transducer with respect to the volumetric map. | 1. A medical diagnostic system, the system comprising: a) a non-ultrasound imaging module configured to generate a volumetric vascular map;b) a controller configured to translate a 3D position co-ordinate from the map into an aiming instruction; andc) a trans-cranial motion-detecting ultrasound unit adapted to aim a narrow beam of a transducer responsive to the aiming instruction. 2. A system according to claim 1, comprising: e) a target identification module. 3. A system according to claim 2, wherein the target identification module operates automatically to identify targets in the volumetric map. 4. A system according to claim 2, wherein the target identification module responds to a user indication of targets in the volumetric map. 5. A system according to claim 2, wherein the controller formulates the aiming instruction responsive to a target identified by the target identification module. 6. A system according to claim 1, wherein the imaging module employs computerized tomography (CT). 7. A system according to claim 1, wherein the imaging module employs magnetic resonance induction (MRI). 8. A system according to claim 1, wherein the imaging module employs CT angiography (CTA). 9. A system according to claim 1, wherein the imaging module employs MRI angiography (MRA). 10. A system according to claim 1, wherein the vasculature includes cerebral vasculature. 11. A system according to claim 1, wherein the motion-detecting ultrasound unit is a transcranial Doppler ultrasound unit (TCDUS). 12. A system according to claim 1, wherein the controller is adapted to scan said beam in space according to said map. 13. A system according to claim 1, comprising a medication providing system whose activity is coordinated with said motion-detecting ultrasound unit. 14. A medical diagnostic method, the method comprising: a) generating a volumetric map of a portion of a vasculature system using a non-ultrasound imaging modality;b) registering a position of a motion-detecting ultrasound transducer with respect to the volumetric map;c) identifying a target on the map by a 3D position co-ordinate;d) translating the 3D position co-ordinate into an aiming instruction;e) aiming a pencil beam of a trans-cranial motion-detecting ultrasound transducer at the target responsive to the aiming instruction; andf) acquiring hemodynamic data pertaining to the target. 15. A method according to claim 14, wherein the identifying is performed by analytic circuitry. 16. A method according to claim 14, wherein the identifying is performed by a user. 17. A method according to claim 14, wherein the volumetric map is generated from computerized tomography (CT) data. 18. A method according to claim 14, wherein the volumetric map is generated from magnetic resonance induction (MRI) data. 19. A method according to claim 14, wherein the volumetric map includes blood flow data. 20. A method according to claim 19, wherein the blood flow data is provided by CT angiography ... | KETER MEDICAL LTD | A61B005/055 A61B008/00 A61B005/05
US20100036233 | A1 | 2010-02-11 | Automatic Methods for Combining Human Facial Information with 3D Magnetic Resonance Brain Images | 2008-08-08 | The present invention relates to compositions and methods for provided fused images of external and internal body scans by fusing data generated from color external laser scans with internal magnetic resonance (MR) images. In particular, the present inventions relate to software for fusing and displaying images produced by fusing color 2.5D laser scans of the human face with 3D magnetic resonance (MR) brain images. Further, these fused images may be used for clinical and research applications, such as in methods for neurosurgical planning procedures. | 1. A system, comprising software, wherein said software is configured to: i) receive a first data set obtained from a digital laser scanner;ii) receive a second data set obtained from a magnetic resonance imaging (MRI) device; andiii) fuse said first data set and said second data set into a fused data set. 2. The system of claim 1, wherein said software is further configured to display said fused data set. 3. The system of claim 2, wherein said first data set comprises a human head and neck scan. 4. The system of claim 3, wherein said first data set further comprises a human face scan. 5. The system of claim 4, wherein said second data set comprises at least one brain anatomy image. 6. The system of claim 4, wherein said second data set comprises at least one brain activation image. 7. The system of claim 5, wherein said fused data comprises said face and head scan superimposed with said brain anatomy image. 8. The system of claim 6, wherein said fused data comprises said a face and head scan superimposed with a brain activation image. 9. The system of claim 1, wherein said digital laser scanner device is a facial scanner device. 10. The system of claim 1, wherein said second data set obtained from a magnetic resonance imaging (MRI) device is in real-time. 11. The system of claim 1, wherein said first data set and said second data set sample are obtained from the same subject. 12. The system of claim 5, wherein said anatomy image comprises an abnormal cell image. 11. The system of claim 6, wherein said activation image comprises an abnormal activation image. 13. A system, comprising: a) a digital laser scanner device,b) a magnetic resonance imaging (MRI) device,c) a first data set and a second data set, wherein said first data set comprises sample data obtained by a digital laser scanner device and said second data set comprises sample data obtained by a magnetic resonance imaging (MRI) device,d) software, wherein said software is configured to: i) receive a first data set obtained from a digital laser scanner;ii) receive a second data set obtained from a magnetic resonance imaging (MRI) device; andiii) fuse said first data set and said second data set into a fused data set. 14. The system of claim 13 wherein said software is further configured to display said fused data set. 15. A method of generating fused sample data, comprising, a) providing, i) a first data set obtained from a digital laser scanner device,ii) a second data set obtained from a magnetic resonance imaging (MRI) device, andb) combining said first sample data set and said second sample data set so as to generate a fused sample data set. 16. The fused sample data set generated according to claim 15. 17. A method of fusing sample data, comprising, a) providing, i) a first data set obtained from a digital laser scanner device,ii) a second data set obtained from a magnetic resonance imaging (MRI) device, andiii) a software package configured to fuse sample data sets,... | Michigan State University | A61B005/055 G06F007/14 G06F017/30 G09G005/00 G06K009/00
US20100049070 | A1 | 2010-02-25 | DEVICE OF IMPROVING AUTOMATIC REAL-TIME HEART RATE VARIABILITY ANALYSIS USING ACCELEROMETER | 2009-08-18 | The present invention provides a mobile ECG device by using accelerometer to improve real-time automatic analysis of heart rate variability, comprising: a data storage module; a microprocessor (MCU) module in which data are stored and accessed in the data storage module; an analog-to-digital converter module in which the data are transmitted to the microprocessor (MCU) module; an accelerometer peripheral hardware module in which the data are transmitted to the analog-to-digital converter module; and an electrocardiogram peripheral hardware module in which the data are transmitted to the analog-to-digital converter module. | 1. A mobile ECG device by using accelerometer to improve real-time automatic analysis of heart rate variability, comprising: a data storage module;a microprocessor (MCU) module in which data are stored and accessed in the data storage module;an analog-to-digital converter module in which the data are transmitted to the microprocessor (MCU) module;an accelerometer peripheral hardware module in which the data are transmitted to the analog-to-digital converter module; andan electrocardiogram peripheral hardware module in which the data are transmitted to the analog-to-digital converter module. 2. The device of claim 1, further comprising a power supply module. 3. The device of claim 2, wherein the power supply module is connected to an external power source or a battery or a combination of the power source and battery. 4. The device of claim 1, wherein the accelerometer peripheral hardware module further comprises a tri-axial accelerometer. 5. The device of claim 1, wherein the ECG peripheral hardware module further comprises an electrocardiogram and electrodes. 6. The device of claim 1, wherein the analog-to-digital converter module converses analog signal to digital signal and is used for operations in the microprocessor (MCU) module. 7. The device of claim 1, wherein the microprocessor (MCU) module further comprises: R-wave detection computation where the data are used for computation of heart rate variability parameters;computation of heart rate variability parameter where the data are used for integrating determination computation;movement determination computation where the data are used for integrating determination computation;integrated determination computation where the data are used for data storage peripheral control; anddata storage peripheral control where the data are stored and accessed in the data storage module. 8. The device of claim 7, wherein the R-wave detection computation is the R-wave interval of the heart beat cycle. 9. The device of claim 7, wherein the parameter computation of heart rate variability is the heart beat interval obtained using ECG or pulse measurement to obtain time domain and frequency spectrum parameters of continuous heart beat in velocity change. 10. The device of claim 7, wherein the movement determination computation uses the data from the accelerometer to process the posture and level of exercise intensity of the current state. 11. The device of claim 7, wherein the integrated determination computation utilizes personal movements as parameters to calibrate the errors in heart rate variability caused by different postures and the HF and LF of various level of exercise intensity. 12. The device of claim 7, wherein the data storage peripheral control is a control chip. 13. The device of claim 1, wherein the data storage module is a flash memory. | HUNGKUANG UNIVERSITY | A61B005/0402
US20100041979 | A1 | 2010-02-18 | INTERVENTIONAL IMMOBILIZATION DEVICE | 2009-10-08 | Interventional immobilization devices used to immobilize a body part and then, during a medical procedure, orient a medical device to treat located tissue within the body part are provided. The devices are designed to immobilize body tissue while preserving, or substantially preserving, the three-dimensional or volumetric integrity of the immobilized tissue. The device enables real time (RT) imaging-guided interventional (IGI) capabilities when the devices are coupled with medical imaging systems, such as magnetic resonance imaging (MRI) systems. | 1. An interventional device comprising: a platform having an upper surface and configured to rotate about a longitudinal axis running perpendicular to the upper surface of the platform;a probe positioner mounted to the platform, wherein the probe positioner is configured to rotate about the longitudinal axis independently of a radio frequency coil;a cup-shaped plate configured to receive a body part, wherein the cup-shaped plate is mounted to the platform; andthe radio frequency coil mounted to the platform, wherein the radio frequency coil is configured to rotate about the longitudinal axis as the platform rotates, and further wherein the radio frequency coil is configured for use with a medical imaging system. 2. The interventional device of claim 1, further comprising a probe guide configured to receive a medical device, wherein the probe guide is connected to the probe positioner in a manner that allows the probe guide to move along the probe positioner in a direction perpendicular to the upper surface of the platform. 3. The interventional device of claim 2, further comprising the medical device mounted in the probe guide. 4. The interventional device of claim 2, wherein the medical device comprises at least one of a biopsy instrument, a therapy probe, a catheter, an ultrasonic device, a trans-cannular device, an excavating tool, an electrical stimulation device, an anesthesia delivery device, a tissue marker placement device, a drug delivery device, a chemical delivery device, and a tumor excision device. 5. The interventional device of claim 4, wherein the therapy probe comprises at least one of a laser ablation probe, a radiofrequency ablation probe, a direct current ablation probe, and a high frequency ultrasound probe. 6. The interventional device of claim 2, wherein the medical device comprises a biopsy needle. 7. The interventional device of claim 2, wherein the medical device comprises a cryo-ablation therapy probe. 8. The interventional device of claim 2, wherein the probe guide is connected to the probe positioner through a pivotal connector, such that the probe guide is configured to pivot relative to a plane which is parallel to the upper surface of the platform. 9. The interventional device of claim 1, wherein the probe positioner moves along a rotary track. 10. The interventional device of claim 1, wherein the interventional device is made of a non-magnetic material. 11. The interventional device of claim 1, wherein the radio frequency coil is mounted to the cup-shaped plate. 12. The interventional device of claim 1, wherein the platform is mounted to a base. 13. The interventional device of claim 1, wherein the cup-shaped plate is configured to rotate independently of the probe positioner. | - | A61B005/055
US07654963 | B2 | 2010-02-02 | Device for visually indicating a blood pressure | 2004-01-14 | An indicator device for visually indicating a pressure of blood inside a blood vessel includes: a body, the body comprising a duct extending in the body and having a sealed proximal end; a distal end portion adapted to be positioned inside the blood vessel and including a liquid inlet opening in fluid communication with the duct; and a window including an at least semi-transparent section configured to enable visual observation of blood entering into the duct via the inlet opening when the inlet opening is located inside the blood vessel. | 1. An indicator device for visually indicating a pressure of blood inside a blood vessel, comprising: a body comprising a passage passing through the body, the body further comprising a duct extending in the body and having a hemostatically sealed blood accommodating chamber;an insertion tube comprising a distal end adapted to be positioned inside the blood vessel and comprising a fluid communication pathway between an uncovered liquid inlet opening near a distal end of the insertion tube and the duct, wherein the inlet opening is formed on an outer surface of the insertion tube, the insertion tube further comprising an opening at an extreme end of the distal end portion;a window comprising an at least semi-transparent section configured to enable visual observation of blood entering into the duct via the inlet opening when the inlet opening is located inside the blood vessel; andan elongated member;wherein the fluid communication pathway is formed between the insertion tube and the elongated member;and wherein the body and the insertion tube are adapted to permit the elongated member to be threaded in a substantially straight path there through between the distal end of the insertion tube and a proximal end of the body and wherein an outer dimension of the elongated member is substantially equal to an inner dimension of the insertion tube at the distal end of the insertion tube and said outer dimension and said inner dimension are configured such that flow of blood between said outer dimension and said inner dimension is prevented when the elongated member is inserted into the insertion tube. 2. The system as claimed in claim 1, wherein the elongated member comprises a guide wire. 3. The system as claimed in claim 1, wherein the duct opens into the chamber via an aperture having a spill-over edge, the aperture being located at a level above a bottom surface of the blood accommodating chamber, whereby return flow of blood back into the duct is prevented. 4. The system as claimed in claim 1, wherein the blood accommodating chamber is located in the body, and wherein the body further comprises the insertion tube extending distally of the body. 5. The system as claimed in claim 4, wherein the inlet opening is located on a side of the insertion tube. 6. The system as claimed in claim 1, wherein the duct extends vertically to an aperture opening into the blood accommodating chamber. 7. The system as claimed in claim 1, wherein the duct extends horizontally above the blood accommodating chamber to an aperture opening into the blood accommodating chamber. 8. The system as claimed in claim 1, wherein the duct exhibits a varying cross-section over its length. 9. The system as claimed in claim 1, wherein the indicator system is configured such that blood flows from a region outside of the insertion tube, through the inlet opening, and into an interior of the insertion tube. 10. The system as claimed in claim 9, wherein the indicator system is conf... | Radi Medical Systems AB | A61B005/02 A61B005/00 A61M005/178
US07649353 | B2 | 2010-01-19 | Magnetic resonance imaging method and apparatus using a circularly polarized RF resonator detecting right and left circularly polarized components of MR signals | 2006-08-22 | In an MR imaging method and apparatus which MR images with improved signal intensity, improved signal-noise ratio, improved contrast and improved image homogeneity can be acquired, the polarization state of the magnetic field of the RF pulses radiated into the measurement subject and of the resonance signals emitted by the measurement subject are distorted by the interaction with electrically-active materials of the measurement subject. In the transmission branch of the RF system the RF pulses emitted by a transmission coil are pre-distorted with regard to their polarization state. The sensitivity of the reception branch is optimized such that it is capable of detecting resonance signals independent of their polarization state. | 1. A method for generating a magnetic resonance image, comprising the steps of: placing an examination subject in a homogenous static magnetic field;operating an RF resonator in a transmission mode to irradiate the examination subject with RF pulses to excite nuclear spins in the examination subject, and operating said RF resonator in a reception mode to detect magnetic resonance signals arising from the excited nuclear spins, and in at least one of said transmission mode and said reception mode, operating said RF resonator as a circularly polarized RF resonator;when operating said RF resonator as a circularly polarized RF resonator in said transmission mode, supplying said RF resonator from two separate transmission systems to irradiate the measurement subject with said RF pulses, by generating, in each of said separate transmission systems, a temporally-oscillating linear polarized electromagnetic field, and adjusting a phase relationship of the respective linearly polarized magnetic fields to cause a sum field produced in the volume of interest by a superimposition of the respective linear polarized electromagnetic fields to exhibit a right circularly polarized electromagnetic field component having a maximum amplitude locally at a site of the nuclear spins to be excited;when operating said RF resonator as a circularly polarized RF resonator in said reception mode, connecting said RF resonator to two separate reception systems, by adjusting respective reception sensitivities of the separate reception systems to detect a maximum of right circularly polarized signal components in said magnetic resonant signals with a first of said separate reception systems and a maximum of left circularly polarized signal components in said magnetic resonance signals with a second of said separate reception systems; andprocessing the magnetic resonance signals detected by said first and second of said separate reception systems in said reception mode to produce a magnetic resonance image therefrom. 2. A method as claimed in claim 1 comprising, in said transmission mode, optimizing a polarization state of the radiated RF pulses to produce a right circularly polarized component that is high and substantially constant in substantially an entirety of a volume of interest within said measurement subject. 3. A method as claimed in claim 1 comprising, in said transmission mode, optimizing a polarization state of the radiated RF pulses to produce a right circularly polarized field with a maximum amplitude within a specified region of a volume of interest in the measurement subject. 4. A method as claimed in claim 3 wherein the step of optimizing the polarization state of the radiated RF pulses comprises: designating N regions within said volume of interest, wherein N is a whole number;successively for each of said N regions, optimizing the polarization state of the radiated RF pulses in the respective N regions and successively detecting magnetic resonance signals, ... | Siemens Aktiengesellschaft | G01V003/00 A61B005/055
US07641612 | B1 | 2010-01-05 | Blood loss detection for hemodialysis system | 2006-01-17 | A blood loss detection device for a hemodialysis system having a blood intake line, a dialysis machine including a pump, with the dialysis machine in communication with the blood intake line, and a blood return line in communication with the dialysis machine. The detection device includes a patch having an exterior edge and an opening through the patch forming a target for receipt of the return blood line. A slit through the patch extends between the exterior edge and the opening. A primary loop circuit terminates at two ends at the slit. An early warning alarm circuit loop terminates at two ends at the slit. An emergency shut-down circuit loop terminates at two ends at the slit. A modular connector connects the patch to an alarm activated by the early warning alarm circuit and to a switch mechanism activated by the shut-down circuit in order to shut down the dialysis machine and the pump. | 1. A blood loss detection device for a hemodialysis system having a blood intake line, a dialysis machine including a pump, said dialysis machine in communication with said blood intake line, and a blood return line in communication with said dialysis machine, which detection device comprises: a patch having an external boundary edge circumnavigating said patch;a central opening through said patch so that said patch surrounds said central opening, wherein said central opening is within said boundary edge and forms a target for receipt of a return access site and said blood return line so that said opening of said patch surrounds and encloses said target;a slit through said patch extending between said boundary edge and said central opening;a primary circuit loop circumnavigating said central opening terminating at two ends at said slit;an early warning alarm circuit loop circumnavigating said central opening terminating at two ends at said slit;an emergency shut-down circuit loop circumnavigating said central opening terminating at two ends at said slit, said emergency shut-down circuit loop independent from said early warning alarm circuit loop wherein said primary circuit loop, said early warning alarm circuit loop, and said emergency shut-down circuit loop are in a radial plane; anda modular connector connecting said patch to an alarm activated by said early warning alarm circuit loop and to a switch mechanism actuated by said emergency shut-down circuit loop to shut down said dialysis machine and pump wherein said early warning alarm circuit loop is activated by a small blood leak and wherein said emergency shut-down circuit loop is activated by a larger blood loss. 2. A blood loss detection device as set forth in claim 1 wherein said modular connector includes a male connector and a female connector and an optional safety loop circuit to sense when said male and female connectors are joined. 3. A blood loss detection device as set forth in claim 1 wherein said emergency shut-down circuit is connected to said dialysis machine including said pump. 4. A blood loss detection device as set forth in claim 1 wherein said patch includes an adhesive base for connection to the skin of a patient. 5. A blood loss detection device as set forth in claim 4 wherein said patch includes a top layer. 6. A blood loss detection device as set forth in claim 1 wherein said patch boundary edge is any geometric shape that surrounds and forms an enclosure of the access site. 7. A method of blood loss detection for a hemodialysis system having a blood intake line, a dialysis machine including a pump, said dialysis machine in communication with said blood intake line, and a blood return line in communication with said dialysis machine, which method comprises: placing a patch having an external boundary edge circumnavigating said patch and a central opening through said patch forming a target for a blood return access site so that said patch surrounds said central... | - | A61B005/0265 A61B005/00 A61M005/00
US07643881 | B2 | 2010-01-05 | Neurostimulation with activation based on changes in body temperature | 2006-01-20 | Methods and devices for detecting and/or predicting onset of an undesirable physiological event or neural state, such as an epileptic seizure, facilitate rapid intervention with a treatment therapy such as neurostimulation or drug therapy. The methods and devices involve monitoring the patient's body temperature, preferably by an implanted temperature sensor, to detect a change in a body temperature parameter over a first time period, which may comprise an increase or decrease in the patient's body temperature, time rate of change of body temperature, moving average temperature, increase or decrease in body temperature exceeding a target temperature, elevated or reduced body temperature, or temperature stability parameter over a first time period. The sampling period may be changed manually or depending upon a sensed body parameter, including temperature. When a parameter change is detected that exceeds a threshold, neurostimulation therapy is delivered to a neural structure of the patient. | 1. A method of providing electrical neurostimulation therapy to a cranial nerve of a patient comprising: generating a time series of temperature data representative of a body temperature of the patient;determining from said data at least two body temperature parameters for the patient over at least one predetermined time period, wherein said at least two body temperature parameters include first and second body temperature parameters;comparing the at least two body temperature parameters with respective thresholds, wherein said respective thresholds include first and second thresholds; andproviding an electrical signal directly to a cranial nerve of the patient contingent upon:(i.) said first body temperature parameter exceeding the first threshold; and(ii.) said second body temperature parameter exceeding the second threshold. 2. The method of claim 1 wherein said at least two body temperature parameters are selected from the group consisting of a time rate of change of temperature over the first predetermined time period, a moving average temperature over the first predetermined time period, a rise in body temperature exceeding a first magnitude over the first predetermined time period, a drop in body temperature exceeding a first magnitude over the first predetermined time period, a body temperature exceeding a first target temperature for the first predetermined time period, a body temperature below a first target temperature for first predetermined time period, a minimum rate of change of body temperature during the first predetermined time period, a maximum rate of change of body temperature during the first predetermined time period, and a temperature stability parameter for the first predetermined time period. 3. The method of claim 2 wherein said temperature stability parameter comprises a duration that body temperature is within a first temperature range, a duration that body temperature is outside a first temperature range, a difference between minimum and maximum body temperatures measured during the first predetermined time period, a standard deviation of body temperature measurements taken over the first predetermined time period, a difference between the minimum and maximum rates of change of body temperature during the first predetermined time period, and standard deviation of rates of change of body temperature during the first predetermined time period. 4. The method of claim 1, wherein providing an electrical signal directly to a cranial nerve comprises providing an electrical signal to a vagus nerve, a trigeminal nerve, or a glossopharyngeal nerve. 5. The method of claim 1 wherein said step of providing the electrical signal comprises performing an action selected from the group consisting of: applying the electrical signal directly to said cranial nerve in response to determining that the at least two body temperature parameters exceed the respective first and second thresholds, andchanging an electrical parameter of said... | Cyberonics, Inc. | A61B005/01 A61N001/36
US07651596 | B2 | 2010-01-26 | Cellulosic-based interference domain for an analyte sensor | 2006-01-18 | The present invention relates generally to devices for measuring an analyte in a host. More particularly, the present invention relates to devices for measurement of glucose in a host that incorporate a cellulosic-based interference domain. | 1. An electrochemical analyte sensor for measuring an analyte in a host, the sensor comprising: at least one electroactive surface; anda membrane system disposed on the electroactive surface, wherein the membrane system comprises:an electrode domain comprising a hydrophilic polymer;an interference domain disposed more distally from the electroactive surface than the electrode domain and comprising cellulose acetate and cellulose acetate butyrate, wherein the interference domain is configured to resist passage therethrough of at least one interferent, and wherein an equivalent glucose signal response of the at least one interferent is less than about 30 mg/dL;an enzyme domain disposed more distally from the electroactive surface than the interference domain; anda resistance domain disposed more distally from the electroactive surface than the enzyme domain, the resistance domain comprising a semipermeable membrane that controls a flux of oxygen and glucose to the underlying enzyme domain. 2. The sensor of claim 1, wherein the equivalent glucose signal response of the at least one interferent is less than about 20 mg/dL. 3. The sensor of claim 1, wherein the equivalent glucose signal response of the at least one interferent is less than about 10 mg/dL. 4. The sensor of claim 1, wherein the equivalent glucose signal response of the at least one interferent is substantially about 0 mg/dL. 5. The sensor of claim 1, wherein the interferent is selected from the group consisting of acetaminophen, ascorbic acid, dopamine, ibuprofen, salicylic acid, tolbutamide, creatinine, uric acid, ephedrine, L-dopa, methyl dopa and tetracycline. 6. The sensor of claim 1, wherein the at least one interferent is selected from the group consisting of ephedrine, tolbutamide and uric acid. 7. The sensor of claim 1, wherein the at least one interferent is selected from the group consisting of acetaminophen, Creatinine, Salicylic acid and Tolazamide. 8. The sensor of claim 1, wherein the at least one interferent is selected from the group consisting of Ascorbic acid, Ibuprofen, and Tetracycline. 9. The sensor of claim 1, wherein the interference domain comprises a layer comprising cellulose acetate and cellulose acetate butyrate. 10. The sensor of claim 1, wherein the interference domain comprises a first layer comprising cellulose acetate butyrate and a second layer comprising cellulose acetate. 11. The sensor of claim 1, wherein the hydrophilic polymer comprises polyvinylpyrrolidone. 12. The sensor of claim 1, wherein the sensor is adapted for subcutaneous insertion. 13. The sensor of claim 1, wherein the sensor is adapted for transcutaneous insertion. 14. The sensor of claim 1, wherein the sensor is adapted for intravascular insertion. 15. A method for forming an electrochemical sensor for measuring an analyte in a host, the method comprising the steps of: providing at least one electroactive surface;forming a membrane system on the electroactive surface, whe... | DexCom, Inc. | G01N027/327 A61B005/05
US07661427 | B2 | 2010-02-16 | Proportional pressure assist ventilation controlled by a diaphragm electromyographic signal | 2003-04-02 | A closed loop system uses (a) the intensity of the diaphragm electromyogram (EMG) for a given inspiratory volume; (b) the inspiratory volume for a given EMG intensity; or (c) a combination of (a) and (b); in view of controlling the level of gas flow, gas volume or gas pressure delivered by a mechanical (lung) ventilator. The closed loop ventilator system enables for automatic or manual adjustment of the level of inspiratory support in proportion to changes in the neuro-ventilatory efficiency such that the neural drive remains stable at a desired target level. An alarm can also be used to detect changes in neuroventilatory efficiency in view of performing manual adjustments. | 1. A method for substantially removing a common noise signal portion from electrical signals produced by an electrical signal source having a polarity reversal region with a centre, said electrical signals having inverse polarities on opposite sides of the centre, said method comprising: sensing a first electrical signal of a first polarity through a first electrode pair located on one side of the centre;sensing a second electrical signal of a second polarity through a second electrode pair located on the other side of said centre;sensing a third electrical signal through a third electrode pair located between said first electrode pair and said second electrode pair;combining the first and second electrical signals into a combination signal; andcombining the combination signal and the third electrical signal into an output signal,whereinsaid first and second polarities are opposite polarities;combining the first and second electrical signals comprises subtracting one of the first and second electrical signals from the other of said first and second electrical signals to produce a combination signal substantially free from said common noise signal portion; andcombining the combination signal with the third electrical signal comprises adding the combination signal and the third electrical signal to produce said output signal still substantially free from said common noise signal portion and thereby substantially prevent output signal loss. 2. A method for substantially removing a common noise signal portion as recited in claim 1, further comprising differentially amplifying said first electrical signal, differentially amplifying said second electrical signal, and differentially amplifying said third electrical signal. 3. A method for substantially removing a common noise signal portion as recited in claim 1, wherein sensing of said first, second and third electrical signals comprises forming with said first, second and third electrode pairs a series of electrodes having an axis extending through the centre of the polarity reversal region substantially in the direction of polarity reversal. 4. A method for substantially removing a common noise signal portion as recited in claim 2, further comprising applying said output signal substantiality free from said common noise signal portion to a ventilatory assistance system. 5. A method for substantially removing a common noise signal portion as recited in claim 2, wherein sensing of said first, second and third electrical signals comprises sensing first, second and third electromyographic signals, respectively, from at least one muscle of a patient, said at least one muscle constituting the electrical signal source. 6. A device for substantially removing a common noise signal portion from electrical signals produced by an electrical signal source having a polarity reversal region with a centre, said electrical signals having inverse polarities on opposite sides of the centre, said device comprising:... | Universite de Montreal | A61B005/04 F16K031/02
US07662106 | B2 | 2010-02-16 | Respiratory profile parameter determination apparatus | 2002-10-15 | A method of sampling one or more respiratory profile characteristics and monitoring a variety of respiratory profile parameters. The sampled respiratory characteristics include respiratory flow rate, respiratory pressure, and partial pressure of at least one constituent of a patient's respiration. The method detects patient breaths, determines whether each breath is a spontaneous breath or a ventilator-induced breath and calculates various respiratory profile parameters based on the sampled measurements. The method displays the respiratory profile parameters in graphic and numeric forms. Preferably, the method allows a user to select the displayed respiratory profile parameters. | 1. Apparatus for monitoring and displaying data obtained from respiration of an individual, comprising: at least one respiratory monitor;at least one processing element in communication with the at least one respiratory monitor and configured to: evaluate a signal generated by the at least one respiratory monitor while monitoring respiration of the individual;calculate at least one respiratory parameter based on the signal; andestimate at least one blood gas parameter based on at least one of the signal and the at least one respiratory parameter; anda display element in communication with at least one of the at least one processing element and the at least one respiratory monitor, and including an indicator comprising a graph including at least one bar, each bar comprising a plurality of sections representing a plurality of complementary parameters of the at least one respiratory parameter or the at least one blood gas parameter obtained at substantially the same time, a first section illustrating a first parameter corresponding to breathing effected by the individual and a second section illustrating a second parameter corresponding to breathing effected by a mechanical ventilator. 2. The apparatus of claim 1, wherein the indicator also comprises a numeric component. 3. The apparatus of claim 1, wherein each bar includes the first section at one end and the second section, adjacent to the first section and discrete therefrom, at another, opposite end. 4. The apparatus of claim 3, wherein the first section represents a first respiratory parameter and the second section represents a second respiratory parameter. 5. The apparatus of claim 4, wherein the first respiratory parameter comprises the number of breaths effected by the mechanical ventilator and the second respiratory parameter comprises the number of breaths effected by the individual. 6. The apparatus of claim 1, wherein the at least one respiratory monitor comprises a respiratory flow monitor. 7. The apparatus of claim 1, wherein the at Least one respiratory monitor comprises a respiratory gas monitor. 8. The apparatus of claim 7, wherein the respiratory gas monitor comprises a respiratory carbon dioxide monitor. 9. The apparatus of claim 1, wherein the at least one respiratory monitor comprises a respiratory flow monitor and a respiratory gas monitor. 10. The apparatus of claim 9, wherein the respiratory gas monitor comprises a respiratory carbon dioxide monitor. 11. A display element of apparatus for monitoring and displaying data obtained from respiration of an individual, the display element comprising: an indicator configured to display a graph in response to a signal from at least one of a respiratory monitor and a processing element of the apparatus, the graph including a plurality of objects, each object of the plurality of objects corresponding to a particular point or period of time, at least one object of the plurality of objects including a plurality of discrete sec... | Ric Investments, LLC. | A61B005/08
US08090427 | B2 | 2012-01-03 | Methods for ultrasound visualization of a vessel with location and cycle information | 2009-10-26 | Methods for displaying a vessel with the aid of intravascular ultrasound images and current location and cycle information provided by detectors and sensors. A sequence of intravascular ultrasound images are stored and indexed-by location and by heartbeat and breathing phase. During a subsequent medical intervention, ultrasound images that correspond to the current phase and location of an object of interest such as a catheter can be selected and displayed on a monitor in real time along with an x-ray image of the vessel, the display showing the geometric position of the ultrasound images with reference to the x-ray image. | 1. A method for displaying a patient's vessel and an object of interest in real time during a medical intervention in the patient, the method comprising the steps: a) recording and storing in a physical memory in a run-up previous to the medical intervention a sequence of intravascular ultrasound vessel cross section images of the vessel from an intravascular ultrasound pull back sequence;b) recording and storing in the same physical memory angiograms of the vessel obtained from an x-ray device while injecting contrast agent;c) indexing the ultrasound vessel cross section images and angiograms at the time of their recording by respective locations of their recording in the vessel;d) indexing the ultrasound vessel cross section images and angiograms at the time of their recording by respective values of a cyclic heartbeat of the patient's heart;e) indexing the ultrasound vessel cross section images and angiograms at the time of their recording by respective values of a cyclic breathing of the patient;f) detecting a current location in the vessel of the object of interest in real time during the medical intervention using an x-ray or NMR (nuclear magnetic resonance) image device to provide a current location image;g) detecting a current value of the cyclic heartbeat using an electrocardiogram;h) detecting a current value of the cyclic breathing using a breathing sensor, and recording the current values of the cyclic heartbeat and the cyclic breathing along with the current location image;i) selecting three or more ultrasound vessel cross section images from the sequence of stored ultrasound vessel cross section images, wherein the selected images have locations of their recording that lie within the current location image and have the same current values of the cyclic heartbeat and the cyclic breathing as the current location image; andj) displaying the selected ultrasound vessel cross section images concurrently with the current location image or a stored angiogram corresponding to the current location image, showing the locations of the recording in the vessel of each of the selected ultrasound vessel cross section images with reference to the current location image or stored angiogram with symbols. 2. The method of claim 1, wherein the object of interest is a catheter tip. 3. The method of claim 1, wherein the object of interest is an anatomical vessel structure of the vessel. 4. The method of claim 1, wherein the medical intervention is a catheter examination and the vessel is the patient's heart. 5. The method of claim 1, wherein the vessel is a heart vessel. 6. A RAM or hard disk storing instructions executable by a computer, the instructions operable to perform a method for displaying a patient's vessel and an object of interest in real time during a medical intervention in the patient, the method comprising the steps: a) recording and storing in physical memory in a run-up previous to the medical intervention a sequence of intravascula... | Koninklijke Philips Electronics N.V. | A61B005/05 A61B008/00
US08323218 | B2 | 2012-12-04 | Generation of proportional posture information over multiple time intervals | 2009-04-30 | The disclosure describes techniques for generation of proportional posture information over multiple time intervals. The techniques may include obtaining posture state data sensed by a medical device for a patient during delivery of therapy by the medical device, determining durations for which the patient occupied each of a plurality of posture states based on the posture state data, generating proportional posture information for a plurality of different time intervals based on the durations, wherein the proportional posture information for each of the time intervals indicates proportional amounts of the respective time interval in which the patient occupied the posture states, and presenting the proportional posture information to a user via a user interface. | 1. A method comprising: obtaining, by at least one processor, posture state data sensed by a medical device for a patient during delivery of therapy by the medical device;determining, by the at least one processor, durations for which the patient occupied each of a plurality of posture states based on the posture state data;generating, by the at least one processor, proportional posture information for a plurality of different time intervals based on the durations, wherein the proportional posture information for each of the time intervals indicates proportional amounts of the respective time interval in which the patient occupied the posture states, wherein each of the time intervals has a length corresponding to a therapy session between programming sessions, and wherein the therapy session defines a period of time during which posture responsive therapy was delivered to the patient by the medical device; andpresenting, by a user interface, the proportional posture information to a user. 2. The method of claim 1, wherein presenting the proportional posture information further comprises presenting a graph that graphically presents the proportional posture information. 3. The method of claim 2, wherein presenting the graph further comprises presenting a posture bar for at least one of the plurality of time intervals, wherein the posture bar includes bar segments indicating proportional amounts of the respective time interval in which the patient occupied each of the posture states. 4. The method of claim 1, further comprising presenting posture state durations of each of the posture states for at least one of the plurality of time intervals. 5. The method of claim 1, further comprising generating a posture state frequency value for each of the posture states for at least one of the plurality of time intervals, wherein the posture state frequency value represents a number of times the patient occupied each of the plurality of posture states during the respective time interval. 6. The method of claim 5, further comprising displaying the posture state frequency value in correlation with the proportional posture information. 7. The method of claim 1, wherein each of the plurality of time intervals corresponds to a time between successive programming sessions for the medical device, and wherein at least some of the time intervals have different lengths. 8. The method of claim 1, wherein the plurality of posture states comprises an upright posture state, a lying front posture state, a lying back posture state, a lying right posture state, and a lying left posture state. 9. The method of claim 1, wherein presenting the proportional posture information further comprises presenting the proportional posture information via a user interface associated with at least one of a patient programmer or a clinician programmer. 10. The method of claim 1, wherein the posture state data comprise a first set of posture state data sensed by the medical device i... | Medtronic, Inc. | A61B005/103 A61B005/117
US08255046 | B2 | 2012-08-28 | Detecting worsening heart failure based on impedance measurements | 2008-07-31 | Techniques for processing impedance data to provide an early warning for heart failure decompensation are described. An example device may be configured to measure intrathoracic impedance values, and increment an index when a determined impedance is less than a reference impedance. The incrementing may be based on the difference between the reference impedances and the determined impedance. In some examples, the amount of incrementing is reduced based on a variability of the impedances, or increased over time so long as the index remains above a threshold, e.g., zero. In some examples, the manner is which the reference impedances are determined changes over time to, for example, address rapid changes in impedance after device or system implantation. In some examples, the index is compared to a threshold to determine whether to provide an alert. In some examples, two thresholds are used to provide hysteresis. | 1. A method comprising: periodically determining, using a medical device, an impedance of a patient based on measured impedances;comparing each of the determined impedances to a respective reference impedance, wherein the respective reference impedance for each of the determined impedances is determined based on a plurality of previously determined impedances;generating an index based on the comparisons between the determined impedances and the reference impedances;determining a variability value based on the plurality of previously determined impedance values;modifying the index over time based on the variability value and the comparisons between the determined impedances and the reference impedances;comparing the index to at least one threshold; anddetermining whether to provide an alert based on the comparison of the index to the at least one threshold. 2. The method of claim 1, wherein modifying the index based on the comparisons and the variability value comprises: determining a modification based on one of the comparisons; andreducing the modification based on the variability value. 3. The method of claim 2, wherein the variability value is scaled by a time dependent value that reduces an affect of the variability value on the modification over time. 4. The method of claim 3, wherein modifying an index over time based on the comparisons between the determined impedances and the reference impedances comprises: increasing the index by differences between the reference impedances and the determined impedances less the scaled variability value so long as the determined impedances are less than the reference impedances. 5. The method of claim 1, wherein determining the variability value comprises: determining a plurality of absolute differences between successive ones of the previously determined impedances; anddetermining a median of the differences as the variability value. 6. The method of claim 1, further comprising: determining a value that increases as a function of time over which the index is greater than one of the thresholds; andadding the value to the index when the index is greater than the threshold. 7. The method of claim 6, wherein adding the value to the index when the index is greater than the threshold comprises adding the value to the index when the index has been greater than the threshold for a threshold duration. 8. The method of claim 1, further comprising determining a subsequent one of the reference impedances based on a comparison of a current one of the determined impedances to a current one of the reference impedances, wherein determining the subsequent one of the reference impedances comprises: increasing the current one of the reference impedances by a predetermined increment value if the current one of the determined impedances is greater than the current one of the reference impedances;decreasing the current one of the reference impedances by a predetermined decrement value if the current one of the determi... | Medtronic, Inc. | A61B005/05
US07888934 | B2 | 2011-02-15 | Magnetic resonance imaging apparatus | 2007-07-13 | With a view to attaining a high quality of a diagnostic image there is provided a magnetic resonance imaging apparatus wherein an image is formed on the basis of magnetic resonance signals received from a subject with an electromagnetic wave transmitted thereto in a static magnetic field. The magnetic resonance imaging apparatus is provided with plural receiving coil units for receiving the magnetic resonance signals, the receiving coil units each comprising a coil body configured to receive a corresponding one of the magnetic resonance signals and output an electrical signal and an optical modulator configured to use an electric field of the electric signal outputted from the coil body directly as a modulation signal. | 1. A magnetic resonance imaging apparatus for generating an image on the basis of magnetic resonance signals received from a subject with an electromagnetic wave transmitted thereto in a static magnetic field, comprising: a plurality of receiving coil units for receiving the magnetic resonance signals, the receiving coil units each comprising a coil body configured to receive a corresponding one of the magnetic resonance signals and output an electrical signal and an optical modulator configured to use an electric field of the electrical signal outputted from the coil body directly as a modulation signal for an optical signal, wherein the optical signal is transmitted through an optical fiber; anda uniting device configured to mutually superimpose and unite the optical signals from each of the plurality of receiving coil units and cause the optical signals to be transmitted in a united state through one optical fiber, the types of optical signals transmitted through the one optical fiber corresponding to a number of receiving coil units in the plurality of receiving coil units. 2. A magnetic resonance imaging apparatus according to claim 1, wherein the optical modulator modulates an optical signal inputted into the optical modulator in accordance with the modulation signal and outputs the modulated optical signal to the exterior of the optical modulator. 3. A magnetic resonance imaging apparatus according to claim 2, wherein the optical modulator is an optical modulator utilizing an electro-optical effect. 4. A magnetic resonance imaging apparatus according to claim 3, wherein the electro-optical effect is Pockels effect. 5. A magnetic resonance imaging apparatus according to claim 4, wherein the optical modulator uses any of BaTiO3, KH2PO4 (KHP), KD2PO4 (KDP) and LiNbO3 crystals. 6. A magnetic resonance imaging apparatus according to claim 3, wherein the electro-optical effect is Kerr effect. 7. A magnetic resonance imaging apparatus according to claim 6, wherein the optical modulator uses the crystal of KTa1-xNbxO3. 8. A magnetic resonance imaging apparatus according to claim 2, wherein the optical modulator is an external modulation type optical modulator. 9. A magnetic resonance imaging apparatus according to claim 8, wherein a light source of the optical signals is a DFB semiconductor laser. 10. A magnetic resonance imaging apparatus according to claim 8, wherein a light source of the optical signals is a single semiconductor laser using a wavelength converter. 11. A magnetic resonance imaging apparatus according to claim 2, wherein the optical modulator is a direct modulation type optical modulator. 12. A magnetic resonance imaging apparatus according to claim 11, wherein the optical modulator is a DFB semiconductor laser with an EA modulator installed therein. 13. A magnetic resonance imaging apparatus according to claim 4claim 1, wherein the optical signals are transmitted through the one optical fiber in accordance with a wav... | GE Medical Systems Global Technology Company, LLC | G01V003/00
US07996064 | B2 | 2011-08-09 | System and method for placing and determining an appropriately sized surgical implant | 2009-10-19 | Images taken before a surgical procedure by an imager are displayed by a computer and overlaid with graphical representations of instruments and/or surgical implants. The graphical representations can be used to determined appropriate sizes of the surgical implants and trajectories for placement. Also, the images can be used in the operating room. The graphical representations can be updated in real-time to correspond to movement of the instruments in the operating room. | 1. A system for placing a surgical implant into a patient comprising: a computer processor;an input system to input dimensions of the surgical implant;a tracking sensor for sensing three-dimensional position information of a surgical instrument to which the surgical implant is removably coupled, the tracking sensor transmitting the position information to the computer processor;a display in communication with the computer processor and operable to display an image;a memory in communication with the computer processor, the memory including computer instructions that when executed by the computer processor cause the computer processor to: overlay a graphical representation of the surgical implant on the image prior to moving the surgical implant into the patient to ensure appropriate dimensions of the surgical implant;generate an icon representing the surgical instrument and the mounted surgical implant; andoverlay the icon on the image to represent the real-time position of at least one of the surgical instrument or the surgical implant relative to the image and the patient;generate a first and a second two-dimensional digitally reconstructed radiograph (DRR) image from a three-dimensional image of an anatomy of the patient;generate a two-dimensional representation of the appropriate size of the surgical implant for display relative to the first and the second digital reconstructed radiograph image; andwherein the overlay of the icon is used to determine an appropriately sized surgical implant based upon the two-dimensional representation relative to both the first and second digital reconstructed radiograph image. 2. The system of claim 1, wherein the image is further a two-dimensional x-ray image. 3. The system of claim 1, wherein the input system to input dimensions includes an input display communicating with the computer processor for displaying a three-dimensional image of anatomy of the patient through which a physician may select an appropriate size and placement for the surgical implant. 4. The system of claim 1, wherein the memory in communication with the computer processor further includes computer instructions that when executed by the computer processor cause the computer processor to generate a two-dimensional representation of the appropriate size of the surgical implant for display relative to a direct two-dimensional image of the patient; wherein the determining an appropriately sized surgical implant is a manual determination of whether the surgical implant is appropriately sized based upon a display of the generated two-dimensional representation of the appropriate size of the surgical implant for display relative to the first and second digital reconstructed radiograph image or the generated two-dimensional representation of the appropriate size of the surgical implant for display relative to the direct two-dimensional image of the patient. 5. The system of claim 4, wherein the surgical implant is a spinal fusion cage and ... | Medtronic Navigation, Inc. | G03B042/02 A61B017/70
US08679025 | B2 | 2014-03-25 | Method for determination of cardiac output | 2007-08-03 | A method for determination of cardiac output from a recording of an arterial pressure waveform (10). The method comprises the steps of measuring the patient's peripheral arterial pressure relative to time in order to estimate a central pressure waveform (CPW) and calculating the patient's cardiac output from the CPW using the Water Hammer formula (as defined herein). | 1. A method for determination of cardiac output from a recording of an arterial pressure waveform, the method comprising the following steps: a) measuring the patient's peripheral arterial pressure relative to time and using a processor to estimate a central pressure waveform (CPW) from the patient's peripheral arterial pressure;b) estimating, using the processor, from the CPW, forward Pressure pulse height (Pf), flow waveform and Pulse Wave Velocity (PWV) in the ascending aorta;c) calculating a peak velocity fluctuation V in the ascending aorta using the following Water Hammer formula: V=Pf/(ρ* PWV), where ρ is the density of blood;d) calibrating an aortic flow waveform from the estimated flow waveform using V; ande) calculating the patient's cardiac output from the calibrated aortic flow waveform. 2. The method as claimed in claim 1, further including calculating Peak Flow Rate (PFR) by multiplying the velocity fluctuation V by the cross sectional area of the aorta, as follows:PFR=V×π⁡(D2)2where D is the aortic diameter. 3. The method as claimed in claim 1, further including calculating a Mean Flow from the calibrated aortic flow waveform. 4. The method as claimed in claim 3, wherein the Mean Flow is derived from an analysis of the CPW calibrated using the Water Hammer formula. 5. The method as claimed in claim 4, further including respectively calculating and estimating: aortic reflection time (ART); aortic PWV; and ejection duration (ED) from the CPW. 6. The method as claimed in claim 5, further including calculating the Mean Flow by integrating the calibrated aortic flow waveform. 7. The method as claimed in claim 6, further including calculating Stroke Volume using the following formula: Stroke Volume=Mean Flow×ED. 8. The method s claimed in claim 7, further including calculating Cardiac Output using the following formula: Cardiac Output=Stroke Volume×Heart Rate. 9. The method as claimed in claim 8, further including estimating the patient's aortic arterial distance from: the patents physical characteristics; a ratio of the carotid to femoral arterial distance; or measuring the superficial distance between sternal notch to the level of the umbilicus. 10. The method as claimed in claim 9, wherein the patient's carotid to femoral arterial distance is estimated form the patient's height, weight and body mass index (BMI). 11. The method as claimed in claim 9, wherein a direct measurement of the distance between the carotid and femoral artery is made. 12. The method as claimed in claim 9, wherein the patient's carotid to femoral arterial distance is estimated from superficial body surface measurements. 13. The method as claimed in claim 9, wherein the ascending aortic pulse wave velocity (PWV) is calculated by dividing the aortic arterial distance by the aortic pulse reflection time. 14. The method as claimed in claim 1, wherein the patient's peripheral arterial pressure measurement is taken at a single site. 15. The method as clai... | AtCor Medical Pty Ltd | A61B005/02
US08758241 | B2 | 2014-06-24 | Electronic module with keyed connection to a wearable garment for monitoring physiological functions and method of use | 2009-07-15 | A plurality of sensors are embedded in a form fitting garment similar to exercise togs such that the sensors are held in contact with or close proximity to the body. The sensors are connected via a plurality wires to an electronics module which is unintrusive being literally in its ultimate configuration the size of a credit card. A range of thickness, from 6 mm (6 credit cards) down to 1 mm or less, is possible for the module inclusive of a rechargeable lithium polymer battery. The electronics module can be easily removed for garment maintenance (laundering). | 1. An electronics module connected to a plurality of sensors via a plurality of wires for use in a garment to monitor physiological functions, the electronics module comprising: a thin polymer battery;a first thin, multilayer substrate placed on and electrically connected to the battery;a plurality of thinned integrated circuits placed on the first substrate;a second thin, multilayer substrate placed on an opposite side of the battery from the first substrate, the second substrate including an antenna patterned thereon and a first inductive coil patterned thereon for inductively recharging the battery and reprogramming the electronics module; andan attachment mechanism for connecting the electronics module to the plurality of wires, whereinthe attachment mechanism is keyed to ensure a proper connection,the attachment mechanism is permanently installed in a pocket formed in the garment, andthe electronics module is programmed to select or deselect ones of the sensors. 2. The electronics module as recited in claim 1, wherein a portion of the first substrate is extended and wrapped around the battery to contact the battery terminals. 3. The electronics module as recited in claim 1, wherein leads on the battery are extended and wrapped around the first substrate to electrical contacts thereon. 4. The electronics module as recited in claim 1, the first substrate having a plurality of resistive and capacitive layers formed therein thereby providing more surface area for active devices contained in the plurality of thinned integrated circuits. 5. The electronics module as recited in claim 1, further comprising an antenna. 6. The electronics module as recited in claim 5, wherein the antenna is patterned on a layer of the first substrate that is disposed farthest away from the battery. 7. The electronics module as recited in claim 1, further comprising means for wirelessly transmitting physiological data received from the plurality of sensors. 8. The electronics module as recited in claim 1, wherein the thinned integrated circuits further include circuitry configured to enable the electronics module to communicate with a base station configured for receiving the electronics module after the electronics module is detached from the plurality of wires to download the physiological data therefrom. 9. The electronics module as recited in claim 1, further comprising a solar cell, the solar cell being thinned and placed on the battery. 10. A method for monitoring physiological functions comprising: embedding a plurality of sensors in a garment;connecting a plurality of wires to the plurality of sensors;connecting an electronics module to the plurality of wires; andprogramming the electronics module to select or deselect ones of the sensors, the electronics module comprising: a thin polymer battery;a first thin, multilayer substrate placed on and electrically connected to the battery; a plurality of thinned integrated circuits placed on the first substra... | The Johns Hopkins University | A61B005/04 H05K007/00 A61B005/05 A61B005/00 G06K019/00 H05K003/00
US08398606 | B2 | 2013-03-19 | Folded perineal pad | 2008-02-28 | A multi-layered, absorbent perineal pad is disclosed that has a generally ovate cross-section. In disclosed embodiments, the pad has an absorbent outer layer, antimicrobial intermediate layer, and absorbent inner layer. At least two opposing edges of the pad are sealed to one another so that the absorbent inner layer encloses an internal cavity within the pad. The pad may be used to absorb liquids from perineal structures (such as vaginal or anal liquids) or administer agents to perineal structures. Liquids that are absorbed in the pad can be subjected to diagnostic assays. | 1. A method of collecting biological liquid from a subject, comprising: providing a multiple layer absorbent pad sized for retention in a perineal space of a wearer, the multiple layers of the pad comprising an inner absorbent layer, an outer absorbent layer, and an intermediate layer between the inner and outer absorbent layers, wherein the inner and outer absorbent layers are denser than the intermediate layer, and the pad forms an empty internal cavity within the pad, wherein liquid in the internal cavity is absorbed through the inner absorbent layer into the intermediate layer and liquid external to the device is absorbed through the outer absorbent layer into the intermediate layer; andplacing the pad in a perineal space of the wearer so that liquid is absorbed by the pad from the perineal space into the pad,wherein the empty internal cavity is circumscribed in at least one plane by the inner absorbent layer and the intermediate layer circumscribes the inner absorbent layer in a continuous manner. 2. The method of claim 1, wherein the intermediate layer is an anti-microbial layer. 3. The method of claim 2, wherein the intermediate anti-microbial layer is treated with an antibacterial substance, or consists essentially of rayon or silk which have inherent anti-microbial activity. 4. The method of claim 3, wherein the intermediate anti-microbial layer consists essentially of rayon. 5. The method of claim 1, further comprising performing a diagnostic test of the liquid absorbed by the pad. 6. The method of claim 1, wherein the pad and cavity are elongated and generally taper in width from a top of the pad to a bottom of the pad, and a ratio of average fiber density of outer layer:intermediate layer:inner layer is about 1-1.5 to 0.25-0.5 to 1.75-2.25, and a ratio of thicknesses of outer layer:intermediate layer:inner layer is about 0.5-1.5 to 8-12 to 1-2. 7. The method of claim 1, wherein the pad is a folded pad that is sealed together along at least two opposing edges to form the internal cavity. 8. The method of claim 1, wherein the pad has an open end that communicates with the cavity, and the pad further comprises a flap that is folded from an edge of the pad into the cavity to form a protective layer. 9. The method of claim 1, wherein the pad has a generally ovate cross-section with an enlarged width top edge and a reduced width bottom edge, and the pad is placed in the perineal space with the enlarged width top edge against a body of the user. 10. The method of claim 1, further comprising: transporting the collected biological liquid in the pad; andextracting at least some of the collected biological liquid from the pad for diagnostic analysis. 11. A method of making a perineal absorbent pad for retention in a perineal space of a wearer, comprising: providing an elongated multi-layered pad comprising at least one absorbent layer and an anti-microbial layer, the pad further comprising a length, a width, a thickness, a longitudinal... | QuantrRx Biomedical Corporation | A61F013/20
US20100036238 | A1 | 2010-02-11 | DEVICE AND METHOD FOR ASSESSING EXTENSION OF A DEPLOYABLE OBJECT | 2009-06-12 | A medical instrument assembly is disclosed. The medical instrument has proximal and distal ends. A receiver coil being mounted proximate to a distal end of the medical instrument, the receiver coil being electrically connectable at a proximal end of the medical instrument. A deployable object being disposed and movable within the medical instrument. The deployable object bearing a high magnetic permeability material located proximate to the receiver coil. The high magnetic permeability material and the receiver coil combining to form an inductive element having an inductance that varies in a predetermined manner with the position of the deployable device relative to the receiver coil. A determination of the inductance being performed at the proximal end of the delivery catheter indicates the extension of the deployable device through the medical instrument. | 1. An inductive measurement system for a medical instrument comprising: a receiver coil mounted proximate to a distal end of the medical instrument, the receiver coil being electrically connectable to an extension analysis system at a proximal end of the medical instrument;a deployable object disposed and movable within the medical instrument, the deployable object bearing a high magnetic permeability material located proximate to the receiver coil, the high magnetic permeability material and the receiver coil combining to form an inductive element having an inductance that varies in a predetermined manner with the position of the deployable device relative to the receiver coil; andan electromagnetic emitter for emitting an electromagnetic wave that induces a current in the receiver coil, the inductance of the inductive element being measurable by the extension analysis system based on the induced current, whereby the extension of the deployable device through the catheter can be determined. 2. The system of claim 1, wherein the medical instrument being a catheter and the deployable object being a needle. 3. The system of claim 2, wherein the high magnetic permeability material includes one or more ferrous cores. 4. The system of claim 3, wherein the one or more ferrous cores being located proximate a distal tip of the needle. 5. The system of claim 2, wherein the receiver coil forms a loop around a lumen of the delivery catheter. 6. The system of claim 1, wherein the medical instrument being a chronic total occlusion catheter and the deployable object being a chronic total occlusion crossing guidewire. 7. The system of claim 1, wherein the deployable object includes three ferrous cores. 8. A medical instrument assembly comprising: a medical instrument having proximal and distal ends;a receiver coil mounted proximate to a distal end of the medical instrument, the receiver coil being electrically connectable at a proximal end of the medical instrument;a deployable object disposed and movable within the medical instrument, the deployable object bearing a high magnetic permeability material located proximate to the receiver coil, the high magnetic permeability material and the receiver coil combining to form an inductive element having an inductance that varies in a predetermined manner with the position of the deployable device relative to the receiver coil, whereby a determination of the inductance at the proximal end of the delivery catheter indicates the extension of the deployable device through the medical instrument. 9. The system of claim 8, wherein the medical instrument being a catheter and the deployable object being a needle. 10. The system of claim 9, wherein the high magnetic permeability material being located near a distal tip of the needle. 11. The system of claim 9, wherein the high magnetic permeability material comprises a ferrous core. 12. The system of claim 8, wherein the receiver coil being a navigation coil. 13. ... | Medtronic, Inc. | A61B005/05
US20100010369 | A1 | 2010-01-14 | NERVOUS SYSTEM MONITORING METHOD | 2009-03-31 | A method for monitoring the response of a nervous system of a body to a stimulus. The method comprises collecting a set of voltage measurements between selected areas on a surface of the body whilst current is being passed between selected regions of the surface of the body. The set of voltage measurements is collected over a predetermined measurement period, the predetermined measurement period is based upon a time after application of the stimulus, and the collected voltage measurements are compared with reference measurements to determine normal or abnormal response of the nervous system. | 1. A method for monitoring the response of a nervous system of a body to a stimulus, the method comprising: providing a plurality of electrodes on a surface of the body and passing current between selected areas of the surface of the body by passing current between at least one pair of electrodes of said plurality of electrodes, said current being provided by a current source external to said body;collecting voltage measurements between selected ones of said electrodes while said current is passed between said at least one pair of electrodes;receiving as input a time delay; andprocessing at least some of said collected voltage measurements to determine a response of the nervous system to the stimulus;wherein said at least some of said collected voltage measurements are collected during a measurement period beginning after a delay based upon the input time delay following occurrence of said stimulus. 2. A method according to claim 1, wherein said processing is based upon a subset of the collected voltage measurements. 3. A method according to claim 1, wherein current is sequentially passed between a plurality of pairs of electrodes and voltage measurements are collected between selected ones of said electrodes while current is passed between each pair of said plurality of pairs of electrodes. 4. A method according to claim 3, wherein current is repeatedly sequentially passed between the plurality of pairs of electrodes and voltage measurements are collected between selected ones of said electrodes while current is passed between each pair of said plurality of pairs of electrodes. 5. A method according to claim 1, further comprising applying said stimulus. 6. A method according to claim 5, wherein voltage measurements are collected independently of stimulus application. 7. A method according to claim 1, wherein the stimulus is selected from the group consisting of a visual stimulus, an auditory stimulus, a sensory stimulus and a transcranial magnetic stimulus. 8. A method according to claim 1, wherein said processing comprises producing an image representing the distribution of impedance within the body. 9. A method according to claim 1, wherein said processing comprises comparing said at least some of said collected voltage measurements with reference data. 10. A method according to claim 1, wherein the stimulus occurs spontaneously. 11. A method according to claim 1, wherein the stimulus is a feature of an environment in which the body is located. 12. A method for monitoring nervous system response to a stimulus, said method comprising: (a) injecting electrical current for a first time period through at least a first pair of electrodes of a plurality of electrodes affixed to a subject;(b) during said first time period, measuring electrical voltage between selected ones of said plurality of electrodes;(c) subsequent to said first time period, injecting electrical current for another time period through at least another pair of said plur... | THE UNIVERSITY OF MANCHESTER | A61B005/05
US20100030085 | A1 | 2010-02-04 | SYSTEM FOR MONITORING AND ANALYSING CARDIORESPIRATORY SIGNALS AND SNORING | 2007-04-26 | Extraction of components of the a signal captured by an accelerometer, obtaining information about physiological data such as cardiac, respiratory and snoring activity. The extracted signal components are useful for the diagnosis of different types of abnormal respiratory phenomena during sleep (apneas, hypopneas and respiratory efforts associated to micro-arousals). | 1-15. (canceled) 16. A method of monitoring and processing cardio-respiratory and snoring signals, the method comprising: securing an acceleration sensor to skin of a patient at the suprasternal notch;collecting from the sensor a continuous acceleration signal that includes acceleration components indicative of cardiac, respiratory and snoring parameters;filtering the signal:triple-processing the filtered signal to extract cardiac, respiratory and snoring information contained in the filtered signal; anddisplaying the extracted information to facilitate medical diagnosis. 17. The method of claim 16, further comprising analyzing the displayed information to determine a medical condition of the patient. 18. The method of claim 16, wherein extracting the cardiac information comprises focusing on [0.5, 3 Hz] frequencies, extracting the respiratory information comprises focusing on [0, 0.5 Hz] frequencies, and extracting the snoring information comprises focusing on an upper spectrum of the signal. 19. The method of claim 16, wherein triple-processing the filtered signal to extract respiratory information comprises performing low-pass filtering with a cutoff frequency of about 0.7 Hz. 20. The method of claim 19, further comprising subjecting the low-pass filtered signal to a calculation step in the time domain, based on zero crossing detection, to determine respiratory rate. 21. The method of claim 16, wherein extracting the cardiac information comprises applying a derivative operator to the filtered signal to obtain derivation values;applying a quadratic operator to the derivation values to obtain absolute values;integrating the absolute values to obtain an integrated waveform;low-pass filtering the integrated waveform to eliminate noises; and thencalculating and averaging intervals between heartbeats, as determined from the filtered waveform. 22. The method of claim 21, wherein calculating and averaging intervals between heartbeats includes identifying peaks of the cardiac signal and corresponding intervals; anddetermining heart rate and heart rate variability. 23. The method of claim 22, further comprising deriving data reflecting sympathetic, parasympathetic and baroreflex sensor activity from the extracted cardiac information. 24. The method of claim 16, wherein extracting the snoring information comprises applying a band-pass filter to the filtered acceleration signal, the band-pass filter having cutoff frequencies corresponding to voice frequencies. 25. The method of claim 24, further comprising identifying snoring intervals from the band-pass filtered signal. 26. The method of claim 16, wherein the extracted information is displayed remotely, for review or diagnosis by a person remote from the patient. 27. The method of claim 16, performed during surgery on the patient, for monitoring patient status during surgery. 28. A cardio-respiratory and snoring signal monitoring and analysis system, comprising a single acceleration sensor t... | Universidad de Cadiz | A61B005/0205
US20100044224 | A1 | 2010-02-25 | BIOLOGICAL FLUID ANALYSIS SYSTEM | 2007-09-12 | A system apparatus (101) includes an integrated sensing and separation component (108) for the monitoring and analysis of clinical species present in biological fluids. In the preferred embodiment this is by use of a polyHIPE membrane (109) with a chemically active sol-gel (110), and a plurality of micro-electrode arrays adapted for selectively adjusting sensitivity. | 1-33. (canceled) 34. A sensor device for the detection of one or more target species comprising an analyte monitoring means, said analyte monitoring means comprising a multiplicity of microelectrode arrays whereby said microelectrode arrays are configured such that they are selectively connectable to each other in order to alter the number of micro-electrode arrays available for the detection of a particular target species. 35. A sensor device as claimed in claim 34, wherein said analyte monitoring means, further comprises a selectively permeable membrane associated with at least one said microelectrode array, said selectively permeable membrane being adapted to substantially simultaneously select which matter is able to pass through said membrane and selectively and reversibly bind a target species. 36. A sensor device as claimed in claim 34, further comprising a sample collection means. 37. A sensor device as claimed in claim 36, wherein the electrode is arranged with respect to the membrane such that collected sample cannot be passed to the electrode without first contacting the membrane. 38. A sensor device as claimed in claim 35, wherein the selectively permeable membrane is adapted to pass only matter of predefined size. 39. A sensor device as claimed in claim 35, wherein the selectively permeable membrane is adapted to selectively pass or retain a species, a charged moiety, a micelle or a biological/biochemical entity. 40. A sensor device as claimed in claim 35, wherein the selectively permeable membrane comprises at least one pore, wherein at least one said pore is at least partially filled with an ion sensing means. 41. A sensor device as claimed in claim 40, wherein the ion sensing means is a chemically modified solgel. 42. A sensor device as claimed in any one of the previous claims wherein the analyte monitoring means is configured to detect ions. 43. A sensor device as claimed in claim 36, wherein the sample collection means and analyte monitoring means are in physical contact to allow substantially simultaneous collection and monitoring of the fluid. 44. A sensor device as claimed in claim 36, wherein the sample collection means is gauze adapted to lake up biological fluid. 45. A sensor device as claimed in claim 34, wherein the analyte monitoring means comprises one or more screen printed electrodes. 46. A sensor device as claimed in claim 45, wherein the electrodes are Ag/AgCl screen printed electrodes modified to act as ion selective electrodes. 47. A sensor device as claimed in claim 34, comprising an internal reference electrode and an external reference electrode with an internal filling layer therebetween. 48. A sensor device as claimed in claim 47, wherein the internal filling layer is a chemically modified hydrogel. 49. A sensor device as claimed in claim 48, wherein the hydrogel is selected from polyvinyl alcohol (PVA) and polyvinyl pyrrolidone (PVP). 50. A sensor device as claimed in claim 35, wherein sai... | - | G01N027/26 A61B005/00 G01N027/333
US07654968 | B1 | 2010-02-02 | Placental blood extractor | 2006-07-28 | A placental blood extractor consisting of an outside box having a plurality of plastic bags mounted within the top and bottom sections of the box and having an opening in the bottom section, a sterile bag having a central closure on a first side and a leader for an umbilical cord on a second side for holding a placenta mounted between the top and bottom plastic bags and a collector of blood system attached to the opening in the bottom, the collector of blood system consists of a pump, a manometer, a computer and a vibrator for gentle shaking of the collector of blood system. | 1. A placental blood extractor for collection of umbilical and placental blood, said blood extractor comprising: a round outside box, said box having a top section and a bottom section and having an opening in said bottom section,a plurality of inflatable compression chambers concentrically mounted within said top section and a corresponding plurality of inflatable compression chambers concentrically mounted within said bottom section, each of said compression chambers having a valve attached thereto,a sterile bag for holding a placenta, said bag being mounted between the compression chambers of said bottom section and said top section upon closure of said top and bottom sections, said sterile bag having a central closure on a first side and a leader for an umbilical cord on a second side, said leader mounted within said opening in said bottom section,a blood collection system attached to said leader, said system having a transparent, firm plastic box for insertion of the sterile collection bags for the collected blood, said collection bags having commercially available anti-coagulates for preventing coagulation of the collected blood, said box having a negative pressure to facilitate the flow of blood from the umbilical cord,a pump attached to each of said plurality of compression chambers for sequentially applying pressure to each of said concentrically mounted compression chambers within said top section and said bottom section,wherein sequentially applying pressure comprises always having the pressure in the next inwardly central compression chamber lower than in the adjacent outer compression chamber,a manometer for determining the pressure to be applied to the compression chambers,a computer for maintaining the levels of the pressure applied to different compression chambers and the time of these applied pressures as the blood is being extracted from the placenta; and for controlling the pressure within said system, anda vibrator mounted within said system for gentle shaking of said blood collection system. | - | A61B005/00 B65D081/00
US07650179 | B2 | 2010-01-19 | Computerized workflow method for stent planning and stenting procedure | 2005-12-09 | In a computerized workflow method for stent planning and conducting a stenting procedure, characteristics of a lesion to be stented are determined from a 3D planning image of the region and selection of an actual stent for stenting the lesion is made with computer-assisted analysis of the lesion based on the characteristics. A virtual stent is electronically generated based on the actual stent, and, using the virtual stent, a best position for the actual stent, for effectively stenting the lesion, is determined. A real time 2D image of the lesion-containing region is displayed during the stenting procedure, with the virtual stent included therein at the aforementioned best position. A physician manually guides the actual stent relative to the lesion during the stenting procedure until the position of the actual stent, as seen in the displayed real time 2D image, coincides with the virtual stent in that image. | 1. A method for computer-assisted stenting, including stent planning, comprising the steps of: acquiring image data from an internal region of a subject comprising a lumen afflicted with a lesion and reconstructing an at least 3D image of said region from said image data, said 3D image containing and representing characteristics of said lesion;conducting a stenosis analysis of said lesion characteristics using only said 3D image with at least a portion of said stenosis analysis being automatically electronically conducted by computerized analysis of said lesion characteristics, to obtain a stenosis analysis result;based on said stenosis analysis result, selecting an actual stent, from among a plurality of available, existing stents of different sizes, to treat said lesion;in a computer, generating a virtual simulated stent corresponding to said actual stent;electronically overlaying only said virtual stent on said 3D image and, in a computer, automatically electronically determining a best position of said virtual stent, relative to said lesion, for effectively treating said lesion as well as automatically electronically determining an angulation of a C-arm of an imaging system that will be used during a stenting procedure to implant said actual stent; andconducting said stenting procedure to implant said actual stent, including generating a real time 2D image of said internal region using said imaging system with said C-arm thereof at said angulation, and combining only said virtual stent at said best position in said real time 2D image, and manually guiding said actual stent in said stenting procedure using said real time 2D image to bring said actual stent to an actual position in said internal region coinciding with said best position. 2. A method as claimed in claim 1 wherein the step of reconstructing an at least 3D image of said internal region comprises reconstructing a 4D image of said internal region from said image data. 3. A method as claimed in claim 1 wherein the step of acquiring said image data comprises acquiring image data from said internal region using a biplane or a monoplane C-arm imaging system. 4. A method as claimed in claim 3 wherein the step of acquiring said image data using a biplane or a monoplane C-arm imaging system comprises: injecting contrast agent into said internal region and acquiring first image data from said internal region in a first plane with said C-arm in a first angulation position; andwith said contrast agent in said internal region, acquiring second image data from said internal region in a second plane with said C-arm in a second angulation position;and wherein the step of reconstructing an at least 3D image from said image data comprises:reconstructing a first 2D image of said internal region from said first image data set and reconstructing a second 2D image of said internal region from said second image data set, and displaying each of said first and second 2D images;in each of the displayed ... | Siemens Aktiengesellschaft | A61B005/05
US07668588 | B2 | 2010-02-23 | Dual-mode physiologic monitoring systems and methods | 2006-03-03 | In some embodiments, a wearable physiologic monitor comprises an application-specific integrated circuit (ASIC) including signal conditioning circuitry, a real-time clock, digital control logic, and mode-selection logic for setting an operating mode of the ASIC to a stand-alone mode or a peripheral mode. In the stand-alone mode, the digital control logic periodically stores data packets including multiple sensor data types in a digital memory such as a removable flash memory card. In the peripheral mode, the data packets are transmitted to a microcontroller for processing. The monitor includes sensors such as electrocardiogram (ECG) electrodes, accelerometers, and a temperature sensor, some of which may be integrated on the ASIC. The same basic chip design may be used in the stand-alone mode in disposable patches, and in the peripheral mode in bedside devices. The operating mode may be chosen at monitor manufacture, by connecting input pins to mode-selection logic levels. | 1. A physiologic monitoring system comprising: a battery;a plurality of electrocardiogram electrodes;a digital memory; andan integrated circuit connected to the battery, the plurality of electrocardiogram electrodes, and the digital memory, comprising: a real-time clock for generating a set of real time indicators;amplification and filtering circuitry for amplifying and filtering a set of physiologic electrical signals received from the plurality of electrocardiogram electrodes to generate a set of filtered electrical signals;digital control logic connected to the real-time clock and the amplification and filtering circuitry; andmode-selection logic connected to the digital control logic, for setting an operating mode of the integrated circuit to a mode selected from a stand-alone mode and a peripheral mode;wherein, in the stand-alone mode, the digital control logic is configured to generate a set of time-stamped physiologic data packets from the set of real-time indicators and a set of digital electrocardiogram data derived from the filtered electrical signals, and transmit the time-stamped data packets for storage in the digital memory; andwherein, in the peripheral mode, the integrated circuit is configured to transfer physiologic data derived from the digital electrocardiogram data to a programmable microcontroller. 2. The system of claim 1, wherein the operating mode of the integrated circuit is set to the stand-alone mode. 3. The system of claim 1, wherein the operating mode of the integrated circuit is set to the peripheral mode. 4. The system of claim 3, further comprising the programmable microcontroller connected to the integrated circuit. 5. The system of claim 1, further comprising a set of integrated circuit mode-selection input pins connected to the mode-selection logic, wherein the mode-selection logic is configured to set the operating mode of the integrated circuit according to a set of logic levels of the mode-selection input pins. 6. The system of claim 5, wherein the mode-selection input pins are bonded to a set of mode-selection logic levels. 7. The system of claim 1, wherein the peripheral mode is a self-clocked peripheral mode, and wherein in the self-clocked peripheral mode the digital control logic is configured to generate the set of time-stamped physiologic data packets from the set of real-time indicators and the digital electrocardiogram data, and transmit the set of time-stamped data packets to the programmable microcontroller. 8. The system of claim 1, wherein the peripheral mode is a passive peripheral mode, and wherein in the passive peripheral mode the digital control logic is substantially disabled. 9. The system of claim 1, wherein the integrated circuit is integrated in a wearable patch configured to be attached to a subject's skin. 10. The system of claim 9, wherein the plurality of electrodes are integrated in the wearable patch. 11. The system of claim 9, wherein the patch comprises an encapsulant... | PhysioWave, Inc. | A61B005/04
US07657308 | B2 | 2010-02-02 | System and methods for performing dynamic pedicle integrity assessments | 2005-02-18 | The present intention involves systems and related methods for performing dynamic pedicle integrity assessments involving the use of neurophysiology. | 1. A method for performing pedicle integrity assessments, comprising the steps of: (a) establishing electrical communication between a stimulation element and an interior of a pedicle hole with an electrical coupling mechanism having a plunger with an electrical contact surface electrically coupled to a stimulation source during at least one of pilot hole formation, pilot hole preparation, and pedicle screw introduction;(b) applying a stimulation signal to said stimulation element; and(c) monitoring to assess whether nerves adjacent said pedicle are innervating as a result of the step of applying said application of stimulation signal to said stimulation element. 2. The method of claim 1 and further, including the further step of providing an insulation member through which said stimulation element is passed. 3. The method of claim 2 and further, wherein step (a) of establishing electrical communication between said stimulation element and said interior of said pedicle hole is accomplished by at least one of: (i) disposing a K-wire through said insulation member such that said K-wire contacts a pedicle screw during introduction into said pedicle hole; (ii) disposing said K-wire through said insulation member such that said K-wire contacts said interior of said pedicle hole during hole formation; (iii) disposing a tap member through said insulation member such that said tap member contacts said pedicle hole during hole preparation; and (iv) disposing a pedicle screw through said insulation member such that said pedicle screw contacts said pedicle hole during screw introduction. 4. The method of claim 1 and further, comprising the step (d) of displaying at least one of alpha-numeric and graphical information to a user regarding whether nerves adjacent said pedicle innervate as a result of applying said stimulation signal to said stimulation element. 5. The method of claim 1 and further, wherein said method may be performed in at least one of a percutaneous procedure, an open medical procedure, and a minimally invasive medical procedure. 6. The method of claim 1 and further, comprising the additional step of providing an insulation member trough which the stimulation element is passed, the insulation member comprising, a flexible hollow elongated sheath having a proximal end, a distal end, an interior lumen, and a substantially enclosed seal tip attached to said distal end of said hollow elongated member, wherein said seal tip includes an expandable aperture for receiving said stimulation element. 7. A system for performing pedicle integrity assessments, comprising: a medical instrument for use in at least one of pedicle hole formation, hole preparation, and pedicle screw insertion at a pedicle target site, said pedicle target site in the general location of neural structures;a sensor configured to detect a voltage response from muscles associated with said neural structures;control unit coupled to said medical instrument and the sensor, the c... | Nuvasive, Inc. | A61B005/04 A61B005/05
US08050751 | B2 | 2011-11-01 | Periodic beat detection to detect artifacts in a cardiac electrogram | 2008-07-31 | Techniques for determining whether artifacts are present in a cardiac electrogram are described. According to one example, a medical device senses a cardiac electrogram via electrodes. The medical device determines a derivative, e.g., a second order derivative, the electrogram. The medical device detects beats within the derivative, e.g., by comparing a rectified version of the derivative to one or more thresholds determined based on a maximum of the rectified derivative. The medical device determines whether the beats are periodic, and determines whether artifacts are present in the cardiac electrogram based on the determination of whether the beats are periodic. The medical device may further determine whether tachyarrhythmia is present and/or whether the cardiac rhythm of the patient is treatable based on the determination of whether the beats are periodic. For example, the medical device may determine that an electrogram is not treatable when the beats are periodic. | 1. A method comprising: determining, using a processor, a derivative of a cardiac electrogram of a patient;detecting, using the processor, beats within the derivative;determining, using the processor, whether the beats are periodic; anddetermining, using the processor, whether artifacts are present in the cardiac electrogram based on the determination of whether the beats are periodic, wherein determining whether artifacts are present comprises determining whether a tachyarrhythmia is present in the patient based on the determination of whether the beats are periodic, and further comprising classifying the cardiac electrogram as treatable using a tachyarrhythmia detection technique, wherein determining whether a tachyarrhythmia is present comprises reclassifying the treatable cardiac electrogram as nontreatable when the beats are periodic. 2. The method of claim 1, wherein determining a derivative comprises determining a second order derivative. 3. A method comprising: determining, using a processor, a derivative of a cardiac electrogram of a patient;detecting, using the processor, beats within the derivative;determining, using the processor, whether the beats are periodic; anddetermining, using the processor, whether artifacts are present in the cardiac electrogram based on the determination of whether the beats are periodic, wherein detecting beats comprises:comparing the derivative to a first threshold to detect small beats; andcomparing the derivative to a second threshold to detect large beats,wherein the first threshold is smaller than the second threshold. 4. The method of claim 3, further comprising: rectifying the derivative;identifying a maximum of the derivative; anddetermining the first threshold and the second threshold based on the maximum,wherein comparing the derivative to the first and second thresholds comprises comparing the rectified derivative to the first and second thresholds. 5. The method of claim 3, wherein determining whether the beats are periodic comprises: determining whether the large beats are periodic;determining whether the small beats are periodic if the large beats are not periodic; anddetermining that the beats are periodic if either the large or small beats are periodic. 6. The method of claim 1, wherein determining whether the beats are periodic comprises determining beat to beat intervals between successive beats. 7. The method of claim 6, wherein determining whether the beats are periodic comprises: comparing the beat to beat intervals to a minimum beat to beat interval threshold; anddetermining that beats are not periodic if any of the beat to beat intervals is less than the threshold. 8. A method comprising: determining, using a processor, a derivative of a cardiac electrogram of a patient;detecting, using the processor, beats within the derivative;determining, using the processor, whether the beats are periodic; anddetermining, using the processor, whether artifacts are present in the cardiac ele... | Medtronic, Inc. | A61B005/0464
US08025404 | B2 | 2011-09-27 | Method and apparatus for monitoring eye tremor | 2009-09-30 | An apparatus consistent with the present invention comprises a sensor for receiving a signal representing eye tremor and a processor for monitoring eye tremor while receiving the signal. A method consistent with the present invention includes receiving a signal representing eye tremor, comparing the received signal representing eye tremor to at least one reference value, and classifying a patient's brain stem function using the comparison of the received signal representing eye tremor to at least one reference value. An embodiment consistent with the present invention includes an ocular micro tremor (OMT) sensor and associated signal processing hardware and software for clinical analyses. | 1. A system for classifying a patient's brain stem function using eye tremor, comprising: a sensor for receiving a signal representing eye tremor;a processor for comparing said received signal representing eye tremor to at least one reference value; anda classifier for classifying the patient's brain stem function using said comparison of said received signal representing eye tremor to at least one reference value; anda filter for filtering microsaccades. 2. The system of claim 1, wherein the classifier further comprises: a classifier for determining the patient's depth of anesthesia. 3. The system of claim 1, wherein the classifier further comprises: a classifier for determining the patient's coma prognosis. 4. The system of claim 1, wherein the processor further comprises: means for determining said at least one reference value from said received signal. 5. The system of claim 1, wherein the processor further comprises: means for determining said at least one reference value from a system that did not produce said received signal. 6. The system of claim 1, wherein the processor further comprises: means for determining said at least one reference value from a signal produced by an electroencephalogram (EEG) monitor. 7. The system of claim 1, wherein the processor further comprises: means for determining said at least one reference value from an electroencephalogram (EEG)-based monitor. 8. The system of claim 1, wherein the processor further comprises: means for determining said at least one reference value from an electroencephalogram (EEG) index. 9. The system of claim 1, wherein the processor further comprises: means for determining said at least one reference value from a signal related to auditory evoked potential. 10. The system of claim 1, wherein the classifier further comprises: means for determining the patient's depth of coma. 11. The system of claim 1, wherein the classifier further comprises: means for monitoring the patient while in a coma state. 12. The system of claim 1, wherein the classifier further comprises: means for determining the patient's brain stem viability. 13. The system of claim 1, wherein the classifier further comprises: means for monitoring motor neuron disease in the patient. 14. The system of claim 1, wherein the classifier further comprises: means for analyzing the patient's sleep pattern. 15. The system of claim 1 wherein the classifier further comprises: means for assessing the patient's combat readiness. 16. The system of claim 1 wherein the classifier further comprises: means for determining when the patient transitions between consciousness and unconsciousness. 17. The system of claim 1, further comprising: a reducer for reducing artifacts from electromagnetic interference. 18. The system of claim 1, further comprising: a discriminator for selecting an eye tremor signal window. | EyeTect, L.L.C. | A61B003/14 A61B005/103 A61B005/117
US08298159 | B2 | 2012-10-30 | Disposable diagnostic article | 2009-09-30 | A disposable diagnostic article comprising a lancing element that can be pierced into a body part and a collection area formed on the lancing element for body fluid obtained by the puncture is disclosed. The lancing element can comprise at least two bent parts folded towards each other such that the at least two bent parts can partially delimit the collection area. | 1. A disposable diagnostic article for testing an analyte in a body fluid sample obtained from a body part, the article comprising: a lancing element for puncturing the body part, wherein the lancing element comprises; at least two bent parts folded towards each other, wherein the at least two bent parts are joined together at their ends to form a lancing member, anda bending zone shaped by the bending and joining of the at least two bent parts, wherein the at least two bent parts and the bending zone have a uniform cross-section, anda collection area formed on the lancing element for collecting body fluid obtained by a puncture, wherein the at least two bent parts at least partially delimit the collection area. 2. The disposable diagnostic article according to claim 1, wherein the collection area is a capillary gap between the at least two bent parts. 3. The disposable diagnostic article according to claim 1, wherein the collection area is continuously open on one side or continuously open on both sides along its longitudinal borders and pointing crosswise to a lancing direction. 4. The disposable diagnostic article according to the claim 1, wherein the at least two bent parts have opposing limiting surfaces for the collection area. 5. The disposable diagnostic article according claim 1, wherein the bending zone is formed at a proximal end of the lancing element. 6. The disposable diagnostic article according to claim 1, wherein the at least two bent parts are joined together by a welding spot. 7. The disposable diagnostic article according to claim 1, wherein the lancing element is formed by bending a single piece of flat material. 8. The disposable diagnostic article according to claim 7, wherein the single piece of flat material is a sheet-metal blank. 9. The disposable diagnostic article according to claim 1, further comprising, a test element to detect the analyte in the collected body fluid. 10. The disposable diagnostic article according to claim 9, wherein the test element is inserted into the lancing element as a prefabricated insert. 11. The disposable diagnostic article according to claim 9, wherein the lancing element comprises an opening for inserting the test element into the collection area. 12. The disposable diagnostic article according to claim 11, wherein the lancing element comprises an opening area in the proximal bending zone for inserting the test element into the collection area. 13. The disposable diagnostic article according to claim 9, further comprising, a proximal end of the collection area forms a retainer for the test element. 14. The disposable diagnostic article according to claim 9, wherein the test element comprises a sensor area formed by a reagent layer which responds to the analyte, wherein the sensor area delimits a flow cross-section of the collection area. 15. The disposable diagnostic article according to claim 9, wherein the test element has at least one optical signal conductor, wherein t... | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00
US08449470 | B2 | 2013-05-28 | System for repetitive measurements of cardiac output in freely moving individuals | 2009-08-20 | A system for evaluating the cardiovascular system parameters using indicator dilution and non-invasive or minimally invasive detection and calibration methods are disclosed. Intravascular indicators are stimulated, and emissions patterns detected for computation of cardiac output, cardiac index, blood volume and other indicators of cardiovascular health. | 1. A cardiac output system for repetitive cardiac output measuring of a freely moving individual comprising: a) a dye calibration and delivery system (CDS) configured to be worn by the individual comprising: i. a calibration station configured to hold at least one calibration cartridge and to draw blood from the individual; andii. a dye delivery station configured to hold at least one dye delivery cartridge and to administer a predetermined amount of dye to the individual's blood stream;b) a sensing device configured to be worn by the individual comprising: i. at least one illumination source configured to rest against the individual's skin and to excite the administered dye in the blood stream; andc) at least one photodetector configured to rest against the individual's skin and to detect fluorescence intensity reflected from or transmitted by the illuminated dye; andd) an external device configured to wirelessly communicate with at least the CDS and the sensing device, and to determine the cardiac output of the individual based on the detected fluorescence intensity and the information indicative of the concentration of the dye in the blood that was drawn from the individual. 2. The system of claim 1 further comprises hardware and software configured to operate the system and the CDS, the sensing device, and the at least one photodetector. 3. The system of claim 1, wherein the system further includes a dye mixing station configured to hold the calibration and/or dye delivery cartridges. 4. The system of claim 3, wherein the cartridges are configured to contain a plurality of separately sealed chambers having different contents. 5. The system of claim 4, wherein the contents include any one of dye, sterile water or sterile saline. 6. The system of claim 5, wherein the mixing station is configured to automatically mix the contents of the cartridges by an ultrasonic process. 7. The system of claim 3, wherein the cartridges are configured to be of a transparent material. 8. The system of claim 1, wherein the calibration station includes at least one element that is configured to illuminate the calibration cartridge having a dye concentration and sense the fluorescence intensity transmitted from or reflected by the dye. 9. The system of claim 8, wherein the external device is configured to calibrate the system based on the sensed fluorescence intensity. 10. The system of claim 1, wherein the external device is configured to enable the CDS remotely to administer the dye to the individual. 11. The system of claim 10, wherein the CDS is configured to administer the dye at predetermined times. 12. The system of claim 11, wherein the dye is administered repetitively based on the individual's measured heart rate. 13. The system of claim 12, wherein the dye is administered while the individual is on an exercise equipment. 14. The system of claim 10, wherein the CDS is configured to administer a predetermined amount of dye to the individual bas... | Alfred E. Mann Institute for Biomedical Engineering at the University of Southern California | A61B005/02 A61B006/00
US08622920 | B2 | 2014-01-07 | Breath test device and method | 2006-11-12 | The disclosure relates to breath test devices and methods for the evaluation of liver functional and metabolic capacity or to assess liver heath and/or degree of liver injury. Specifically, there is provided a method of evaluating a liver condition, the method includes on-line monitoring a metabolic product of methacetin, a salt or a derivative of methacetin, in a subject's breath after administering to the subject isotope labeled methacetin, a salt or a derivative thereof in water solution form. There is also provided a device for evaluating a liver condition, the device includes one or more sensors adapted to monitor on-line an isotope level of a metabolic product of labeled methacetin, or a salt or a derivative of methacetin in a subject's breath and a controller adapted to sample measurements of the one or more sensors at a continuous mode. | 1. A method of quantitatively evaluating a liver function and/or a degree of liver injury in a subject, the method comprising: providing solution samples of isotope-labeled methacetin or a salt or a derivative thereof wherein the solution samples are produced in a process comprising: mixing, in a pre-determined amount of water, a pre-determined amount of isotope-labeled methacetin or the salt or the derivative thereof, to produce a solution;analyzing the solution to measure an amount of the methacetin or the salt or the derivative thereof dissolved in the solution; andproviding for administration solution samples wherein, according to the analysis of the solution, at least 90% of the pre-determined amount of methacetin, or the salt or the derivative thereof, is dissolved in the solution;administering the analyzed solution sample to a subject in need thereof;following the administering, measuring a change in isotope ratio of a metabolic product of the isotope-labeled methacetin, or the salt or the derivative thereof, in the subject's breath using a breath sensor adapted to monitor the metabolic product of the isotope-labeled methacetin or the salt or the derivative thereof; andquantitatively evaluating the liver function and/or the degree of liver injury based on the measured change in the isotope ratio. 2. The method of claim 1, wherein the at least 90% of the pre-determined amount of methacetin, or the salt or the derivative thereof comprises over 99% of the pre-determined amount of isotope-labeled methacetin, the salt or the derivative thereof. 3. The method of claim 1, wherein administration comprises oral, intravenous, an implemented device, or any combination thereof. 4. The method of claim 1, wherein the isotope-labeled methacetin, or the salt or the derivative thereof comprises methacetin or salt or derivative thereof labeled with carbon-13, carbon-14, oxygen-18 or any combination thereof. 5. The method of claim 1, further comprising administering to the subject a test meal. 6. The method of claim 1, further comprising: on-line monitoring a metabolic product of the isotope-labeled methacetin, the salt or the derivative thereof, in the subject's breath after administering the analyzed solution to the subject. 7. The method of claim 6, further comprising monitoring CO2 in breath. 8. The method of claim 7, further comprising analyzing at least one breath related parameter obtained by measuring the change in the isotope ratio of the metabolic product of the isotope-labeled methacetin, the salt or the derivative thereof, in combination with at least one breath related parameter obtained by monitoring CO2 in breath. 9. The method of claim 8, further comprising analyzing at least one breath related parameter obtained by measuring the change in the isotope ratio of the metabolic product of the isotope-labeled methacetin, the salt or the derivative thereof, in combination with at least one physiological and/or medical parameter. 10. The me... | Exalenz Bioscience Ltd. | A61K051/00 A61B005/097
US08958885 | B2 | 2015-02-17 | Posture state classification for a medical device | 2009-04-30 | One embodiment relates to a medical device including a sensor to provide signals indicative of a detected posture state of a patient. A memory stores one or more defined vectors, each defined vector being associated with a tolerance describing a relationship with the defined vector. A processor determines a respective similarity between the detected vector and each of one or more of the defined vectors and classifies a posture state of the patient based on whether any similarity has a relationship to the respective defined vector that is described by the associated tolerance. In one embodiment, the similarity is determined without regard to a coordinate system of the patient. Another embodiment relates to determining the similarity based on at least one of an inner product, a length of the defined posture vector and a length of the detected vector. | 1. An implantable medical device, comprising: a sensor configured to provide signals indicative of a detected posture vector of a patient;a memory configured to store one or more defined posture vectors and to associate each defined posture vector with a tolerance describing a relationship with the defined posture vector; anda processor configured to determine, without any derivation of an angle, a respective similarity between the detected posture vector and each of one or more of the defined posture vectors and to classify a posture of the patient based on whether any similarity has a relationship to the respective defined posture vector that is described by the associated tolerance. 2. The device of claim 1, wherein the processor is further configured to determine each similarity based on at least one of an inner product between the detected posture vector and a respective one of the defined posture vectors, a squared length of the detected posture vector, and a squared length of the defined posture vector. 3. The device of claim 1, wherein the processor is configured to determine the defined posture vector to which the detected posture vector is closest. 4. The device of claim 1, wherein the processor is configured to determine each similarity as a distance between the detected vector and a defined vector. 5. The device of claim 4, wherein the processor is further configured to map the distance to obtain a value ranging between a selected minimum value and a selected maximum value, and to use the obtained value to classify the posture of the patient. 6. The device of claim 1, wherein multiple ones of the defined posture vectors are associated with a same posture. 7. The device of claim 6, wherein the multiple ones of the defined posture vectors associated with a same posture describe multiple regions in space interrelated by one or more logical functions. 8. The device of claim 7, wherein at least some of the multiple regions in space overlap with one another. 9. The device of claim 8, wherein the multiple regions in space include one or more toroids and one or more cones. 10. The device of claim 1, wherein the processor is configured to determine each similarity without regard to a coordinate system of the patient. 11. The device of claim 1, wherein the processor is further configured to classify the posture of the patient based on a first set of the one or more defined posture vectors, and to re-classify the posture of the patient based on a second set of the one or more defined posture vectors, each of one or more defined posture vectors in the second set being a sub-classification of a posture in which the patient is first classified. 12. The device of claim 1, wherein the processor is further configured to process multiple ones of the defined posture vectors to obtain a virtual posture vector, and wherein the virtual posture vector is one of the one or more defined posture vectors used to classify the posture of the patient. ... | Medtronic, Inc. | A61N001/00 A61N001/365 A61B005/11 A61B005/00
US08366707 | B2 | 2013-02-05 | Systems and methods for epicardial navigation | 2008-01-23 | Systems and methods for navigating a catheter along the epieardial surface of the heart are disclosed. At least some of the embodiments disclosed herein are useful for epieardial ablation, Various embodiments permit ablation of the epieardial surface of the heart using an external ablation catheter in the pericardial space and an internal guide catheter within the heart. Such a configuration allows the clinician to precisely target for ablation specific locations on the cardiac tissue. | 1. A system for use with a vacuum source for ablating a tissue of a heart, comprising: an engagement catheter comprising a proximal end, a distal end, and first and second lumens extending between the proximal end and the distal end;an external catheter comprising a distal end, an ablation contact positioned at or near the distal end of the external catheter, and a first magnet positioned at or near the distal end of the external catheter, wherein the external catheter is configured such that the external catheter is capable of at least partial insertion into the second lumen of the engagement catheter;an internal catheter comprising a distal end and a first magnet positioned at or near the distal end of the internal catheter, the internal catheter devoid of an ablation contact; anda vacuum port located at the proximal end of the engagement catheter, the vacuum port being operatively connected to the first lumen of the engagement catheter and capable of operative connection to the vacuum source;wherein the first lumen of the engagement catheter includes a suction port located at or near the distal end of the engagement catheter, the suction port being configured to removably attach to a targeted tissue on the interior of a wall of the heart, such that the suction port is capable of forming a reversible seal with the targeted tissue when the vacuum source is operatively attached to the vacuum port,wherein the internal catheter is configured for engaging the external catheter when a portion of the external catheter is positioned on or adjacent to an epicardial surface of the heart and a portion of the internal catheter is positioned on or adjacent to an endocardial surface of the heart, such that manipulation of the internal catheter is capable of directing the external catheter, andwherein the system is capable of enlarging a pericardial space between the targeted tissue and a pericardial sac that surrounds the heart by retracting the targeted tissue away from the pericardial sac; andwherein the internal catheter is configured for engaging the external catheter when a portion of the external catheter is positioned on or adjacent to the epicardial surface of the heart and a portion of the internal catheter is positioned on or adjacent the endocardial surface of the heart, such that manipulation of the internal catheter is capable of directing the external catheter so that the external catheter is operable to remove or destroy tissue at the first surface of the tissue at a location directed by the internal catheter; andwherein the external catheter is further operable to remove or destroy tissue at the epicardial surface at a second location directed by the internal catheter and along the epicardial surface between the first location and the second location without moving the internal catheter. 2. The system of claim 1, wherein: the internal catheter further comprises a second magnet;the first magnet of the internal catheter comprises a first elect... | CVDevices LLC | A61B018/14
US08398561 | B2 | 2013-03-19 | Motivational spirometry system and method | 2009-10-12 | A method is disclosed for motivating a user of a spirometry system having a display screen, an air-tube, a flow measurement sensor operatively associated with the air-tube, a display screen, and an associated processor. A motivational animation is displayed on the display screen having a first aspect that reflects a relative evaluation of a determined peak flow rate to a goal value therefor and having a second aspect that reflects a relative evaluation of a determined total flow volume to a goal value therefor. The processor repeatedly updates the motivational animation during the period of use of the spirometer to reflect the degree of achievement of both of the goal values by the user. | 1. A method for motivating a user of a spirometer during a period of use, said method comprising the steps of: generating a motivational animation having a first aspect that reflects a relative evaluation of a determined peak flow rate to a goal value therefore and having a second aspect separate from said first aspect that reflects a relative evaluation of a determined total flow volume to a goal value therefore;displaying said motivational animation to said user based on a common animation theme and wherein each of the goal values of peak flow rate and total flow volume result in a tangible result in said motivational animation; andrepeatedly updating the motivational animation display during the period of use of the spirometer to reflect the degree of achievement of both of the goal values by the user and in which the second aspect is not displayed until a target flow rate represented by said first aspect is first achieved. 2. The method as recited in claim 1 wherein the user exhales through the spirometer during the period of use. 3. The method as recited in claim 1 wherein the user inhales through the spirometer during the period of use. 4. A method for motivating a user of a spirometer having an air tube, said method comprising the steps of: measuring a flow of air through the air tube induced by the user during a period of use of the spirometer apparatus;processing the air flow measurements to determine a peak flow rate and a total flow volume for the period of use;evaluating the determined peak flow rate relative to a goal value therefor;evaluating the determined total flow volume relative to a goal value therefore; anddisplaying a motivational animation to the user, said motivational animation being based on a common animation theme and in which the goal values of the peak flow rate and the total flow volume result in a tangible result in said motivational animation, said animation having a first aspect that reflects the relative evaluation of the determined peak flow rate to the goal value therefore and having a second aspect separate from said first aspect that reflects the relative evaluation of the determined total flow volume to the goal value therefore and in which the second aspect is not displayed until a target flow rate represented by said first aspect is first achieved. 5. The method of claim 4 further comprising repeatedly updating the motivational animation display during the period of use of the spirometer apparatus to the degree of achievement by the user of both of the goal values. 6. The method of claim 5 further comprising selecting the goal values for peak flow rate and total flow volume based on the demographics of the user and normative parameters from a clinical population study. 7. The method of claim 6 wherein selecting the goal values for peak flow rate and total flow volume based on the demographics of the user and normative parameters from a clinical study comprises selecting the goal values for peak flow... | Welch Allyn, Inc. | A61B005/08 A63B023/18
US08361535 | B2 | 2013-01-29 | Latent print development apparatus | 2009-08-17 | A kit, an apparatus, and method are provided for developing a latent print. The apparatus comprises a pellet comprising a substrate impregnated with a sublimation compound, a receptacle, an adapter, and an end cap. The receptacle is open at a first end and detachably connected to a heat source at a second end. The receptacle is designed to receive and accommodate the pellet through the first end of the receptacle. The adapter defines an annular space for transmission of heat from the heat source to the pellet. The first end of the adapter is connected to the second end of the receptacle and the second end of the adapter is in communication with the heat source. The end-cap comprising a mesh window in an annular space of the end-cap allows passage of the fumes released from the pellet through the open end of the end cap towards the latent print for development. | 1. A method of developing a latent print, comprising the steps of: providing an apparatus for developing said latent print, comprising: a generally cylindrical receptacle configured to define an annular space, said receptacle having a first end and a second end, said first end connected to an end cap, said second end having a coaxially threaded extension for detachably attaching said second end of said receptacle to a heat source via an adapter, wherein said first end of said receptacle is configured to receive and accommodate a pellet, and wherein said pellet is impregnated with a sublimation compound on a substrate;said adapter generally cylindrically shaped and configured to define an annular space for transmission of heat from said heat source to said pellet via said adapter, and wherein said annular space is configured to accommodate a heat diffuser;said end-cap generally cylindrically shaped with a pivot point defined at said first end of said receptacle, a handle fixedly attached to said end-cap at a position diametrically opposite to said pivot point, and an open end in communication with said first end of said receptacle, wherein said end cap comprises a mesh window in an annular space of said end-cap to allow passage of said fumes released from said pellet through said open end of said end cap towards said latent print, and wherein said end-cap comprises a ball nose of a ball nose spring locking element, wherein said ball nose spring locking element is positioned near said handle of said end cap;attaching said receptacle to said heat source through said second end of said receptacle, wherein said receptacle receives said pellet;introducing said pellet into said receptacle through said first end of said receptacle;closing said first end of said receptacle using said end-cap, wherein said end-cap is pivotally connected to said first end of said receptacle at said pivot point on said first end of said receptacle;applying heat from said heat source to said receptacle through said heat diffuser, wherein said pellet introduced into said receptacle releases fumes through said open end of said end-cap towards said latent print via said mesh window for enabling said development of said latent print, and wherein said heat diffuser absorbs the heat from the heat source and precludes flames from said heat source from crossing the heat diffuser to burn the substrate; andproviding a removal mechanism for removing said pellet from said receptacle when said sublimation compound in said pellet is depleted, wherein said handle is pre-extended in length to enable safe opening of said end cap for removing said pellet depleted of said sublimation compound from said receptacle. 2. The method of claim 1, wherein the step of attaching said receptacle to said heat source comprises engaging said coaxial threaded extension of said receptacle to an internally threaded extension of said adapter in communication with said heat source. 3. The method of claim 1, fur... | - | A61B005/117 G06K009/22 C23C016/00
US08905925 | B2 | 2014-12-09 | Cardiac rehabilitation using patient monitoring devices | 2009-06-10 | Systems and methods permit remotely-monitored rehabilitation of a patient. A system can comprise a patient monitor configured to monitor a patient's physiological data. The patient monitor can monitor using a first monitoring mode. A patient communication device can be configured to present, to the patient, an option to perform an exercise regimen, and to receive a response indicating whether the patient will perform the exercise regimen. A configuration module coupled to the patient communication device can be configured to activate a second monitoring mode when the response indicates that the patient will perform the exercise regimen. When the exercise regimen is complete, the first monitoring mode can be re-established. When the response indicates that the patient will not perform the exercise regimen, the response can be recorded as a negative response, and the option to perform the exercise regimen can be presented again. | 1. A system comprising: a patient monitor configured to monitor a patient's physiological data using a first monitoring mode;a patient communication device configured to: present, to a patient, an option to perform an exercise regimen, the option including at least one response option indicating the patient will perform the exercise regimen and at least one response option indicating the patient will not perform the exercise regimen; andreceive a response to the presented option from the patient indicating whether the patient will perform the exercise regimen; anda configuration module, coupled to the patient communication device, the configuration module configured to, when the response to the presented option indicates that the patient will perform the exercise regimen: activate a second monitoring mode of the patient monitor;detect that the exercise regimen is complete; andre-establish the first monitoring mode after the exercise regimen is complete; andwhen the response to the presented option indicates that the patient will not perform the exercise regimen: receive from the patient, a reason why the patient will not perform the exercise regimen;record the response to the presented option as a negative response with the received reason; andrepeat presenting the option to perform the exercise regimen to the patient. 2. The system of claim 1, wherein the patient monitor is configured to monitor the patient's physiological data repeatedly at intervals of time. 3. The system of claim 1, wherein the patient monitor includes an implantable medical device. 4. The system of claim 1, wherein the patient monitor is configured to: receive the patient's physiological data at a central system; andstore the patient's physiological data at the central system. 5. The system of claim 1, wherein the configuration module is configured to detect that the exercise regimen is complete by receiving an indication from the patient that the patient has completed exercising. 6. The system of claim 1, wherein the configuration module is configured to use patient physiological data sensed in the second monitoring mode to determine that the patient has completed exercising. 7. The system of claim 1, comprising: a clinician communication device, coupled to the patient communication device, and configured to: present, to a clinician, the patient's physiological data captured during the first and second monitoring modes;receive a clinician message from the clinician, the clinician message being in response to the presentation of the patient's physiological data; andcommunicate with patient communication device to revise at least one of the exercise regimen or the second monitoring mode based on the clinician message. 8. The system of claim 7, wherein the patient communication device is configured to revise the exercise regimen by revising at least one of: an intensity of an exercise, a frequency of an exercise, a duration of an exercise, a type of exercise, or a speci... | Cardiac Pacemakers, Inc. | A61B005/00 A61B005/22 A61B005/0205 A61N001/372 A61N001/362 A61N001/00 G06F019/00
US09662045 | B2 | 2017-05-30 | Generation of sleep quality information based on posture state data | 2009-04-30 | This disclosure describes techniques for generation of sleep quality information based on posture state data. The techniques may include obtaining posture state data sensed by a medical device for a patient, generating sleep quality information based on lying posture state changes indicated by the posture state data, and presenting the sleep quality information to a user via a user interface. | 1. A method comprising: delivering therapy, during a time period, via a medical device to a patient in response to determined posture states of the patient, wherein the determined posture states include a plurality of different lying posture states;obtaining, with a computing device, posture state data sensed by the medical device for the patient;generating, with at least one processor of the computing device, sleep quality information based on lying posture state changes indicated by the posture state data, wherein the posture state data indicates lying posture state changes during the time period in which the medical device delivers the therapy to the patient in response to the determined posture states of the patient; andpresenting the sleep quality information to a user via a user interface of the computing device. 2. The method of claim 1, wherein the medical device comprises an implantable medical device, and wherein obtaining posture state data comprises obtaining the posture state data from the implantable medical device. 3. The method of claim 2, wherein the implantable medical device is configured to sense the posture state data, and determine a posture state occupied by the patient based on the posture state data, and deliver electrical stimulation to the patient based on the determined posture state of the patient. 4. The method of claim 1, wherein presenting the sleep quality information further comprises presenting a sleep quality graph that graphically presents the sleep quality information. 5. The method of claim 1, further comprising generating the sleep quality information based on a number of changes from the plurality of different lying posture states during the time period. 6. The method of claim 1, wherein generating the sleep quality information comprises generating sleep quality information for each of a plurality of time intervals during the time period, wherein the sleep quality information for each of the time intervals indicates a number of changes from the plurality of different lying posture states during the respective time interval. 7. The method of claim 1, wherein generating sleep quality information comprises generating the sleep quality information based on the lying posture state changes indicated by the posture state data, and wherein the sleep quality information indicates a number of changes from each of the plurality of different lying posture states during the time period. 8. The method of claim 7, wherein the number of changes includes a daily average number of the changes over the time interval, and wherein the time interval is one of a week, a month, a year, or a time between successive programming sessions for the medical device. 9. The method of claim 1, wherein the sleep quality information indicates a number of changes from each of the plurality of different lying posture states during the time period, and wherein the number of changes is one of a number of changes from each of the lying po... | Medtronic, Inc. | A61B005/00 A61B005/11 A61N001/372 A61N001/05 A61N001/36 A61N001/365
US20100049021 | A1 | 2010-02-25 | DEVICES, SYSTEMS, METHODS AND TOOLS FOR CONTINUOUS ANALYTE MONITORING | 2009-08-20 | One aspect of the invention provides an analyte monitor including a sensing volume, an analyte extraction area in contact with the sensing volume adapted to extract an analyte into the sensing volume, and an analyte sensor adapted to detect a concentration of analyte in the sensing volume. The sensing volume is defined by a first face, a second face opposite to the first face, and a thickness equal to the distance between the two faces. The surface area of the first face is about equal to the surface area of the second face and the extraction area is about equal to the surface area of the first and second face of the sensing volume. The analyte sensor includes a working electrode in contact with the sensing volume, the working electrode having a surface area at least as large as the analyte extraction area, and a second electrode in fluid communication with the sensing volume. | 1. An analyte monitor comprising: a sensing volume defined by a first face, a second face opposite to the first face, and a thickness equal to the distance between the two faces, wherein the surface area of the first face is about equal to the surface area of the second face;an analyte extraction area in contact with the sensing volume and adapted to extract an analyte into the sensing volume, wherein the extraction area is about equal to the surface area of the first and second face of the sensing volume;an analyte sensor adapted to detect a concentration of analyte in the sensing volume, the analyte sensor comprising: a working electrode in contact with the sensing volume, the working electrode having a surface area at least as large as the analyte extraction area, anda second electrode in fluid communication with the sensing volume. 2. The analyte monitor of claim 1, wherein the extraction area is an area of the analyte monitor that is further adapted to contact skin of a patient. 3. The analyte monitor of claim 1, wherein the ratio of an area of the first face of the sensing volume to the thickness is at least 10 to 1. 4. The analyte monitor of claim 1, wherein the extraction area is in contact with the first face of the sensing volume and the working electrode is in contact with the second face of the sensing volume. 5. The analyte monitor of claim 1, wherein the second electrode is not in contact with the sensing volume. 6. The analyte monitor of claim 1, wherein the second electrode is a reference electrode and the analyte monitor further comprising a counter electrode in fluid communication with the sensing volume. 7. The analyte monitor of claim 1, wherein the extraction area comprises a plurality of tissue piercing elements, each tissue piercing element comprising a distal opening, a proximal opening and an interior space extending between the distal and proximal openings. 8. The analyte monitor of claim 7, wherein the sensing volume comprises a sensing fluid and is in fluid communication with the proximal openings of the tissue piercing elements. 9. The analyte monitor of claim 1, wherein the sensing volume comprises a sensing fluid and the analyte sensor is adapted to detect a concentration of analyte in the sensing fluid. 10. The analyte monitor of claim 1, wherein the analyte sensor is an electrochemical sensor. 11. The analyte monitor of claim 1, wherein the surface area of the working electrode is in the range of 2 mm2 to 100 mm2. 12. The analyte monitor of claim 11, wherein the surface area of the working electrode is in the range of 10 mm2 to 50 mm2. 13. The analyte monitor of claim 1, wherein the thickness of the sensing volume is in the range of 50 microns to 3000 microns. 14. The analyte monitor of claim 1, wherein the extraction area is equal to the surface area of the first face of the sensing volume. 15. The analyte monitor of claim 1, wherein the extraction area is the same size and shape as the first face o... | - | A61B005/1468
US20100041976 | A1 | 2010-02-18 | CRYOSTAT FOR REDUCED CRYOGEN CONSUMPTION | 2008-10-31 | A cryostat has a cryogen vessel retained within an outer vacuum container, a thermally insulating jacket surrounding the outer vacuum container and insulating it from ambient temperature. | 1. A cryostat comprising a cryogen vessel retained within an outer vacuum container (OVC), an active refrigeration unit that cools the OVC, and a thermally insulating jacket surrounding the OVC and insulating the OVC from ambient temperature. 2. A cryostat according to claim 1, wherein the jacket is mechanically bonded to the OVC. 3. A cryostat according to claim 1, wherein the thermally insulating jacket integrates the function of a looks cover. 4. A cryostat according to claim 3, wherein the jacket comprises a solid material having molded channels for accommodating auxiliary equipment. 5. A cryostat according to claim 1, wherein the thermally insulating jacket comprises a flexible material. 6. A cryostat according to claim 1, wherein the thermally insulating jacket comprises a loose material. 7. A cryostat according to claim 6, wherein the loose material is contained within flexible pouches. 8. A cryostat according to claim 1, wherein the thermally insulating jacket provides damping against mechanical shocks experienced by the system during transit. 9. A cryostat according to claim 1, wherein the thermally insulating jacket protects the OVC from contaminants such as sea water during transit. 10. A cryostat according to claim 1, wherein the thermally insulating jacket is integrated with a pallet for transporting the cryostat. 11. A cryostat comprising a cryogen vessel retained within an outer vacuum container (OVC), the cryogen vessel having a vent path allowing cryogen gas to escape from the cryogen vessel, and a thermally insulating jacket surrounding the OVC and insulating the OVC from ambient temperature, and a cooling pipe in thermal contact with the OVC, that directs cryogen gas escaping through the vent path through the cooling pipe, to cool the OVC. 12. A cryostat according to claim 11, wherein the cooling pipe is interposed between the thermally insulating jacket and the OVC. 13. A cryostat according to claim 12, wherein the cooling pipe is detachable from the OVC. 14. A cryostat according to claim 11, wherein the cooling pipe encircles the OVC at least once. 15. A cryostat according to claim 11, wherein the cooling pipe follows a serpentine path over a surface of the OVC. 16. A cryostat according to claim 11, wherein the cooling pipe is flexible, and is contained within a flexible wrapper, which is wrapped around the OVC and overlain with the jacket. 17. A cryostat according to claim 11, wherein the cooling pipe is applied to a cylindrical surface of the OVC. 18. A cryostat according to claim 10, wherein the cooling pipe is mechanically bonded to the OVC. 19. A cryostat according to claim 11, wherein the jacket (30) is mechanically bonded to the OVC. 20. A cryostat according to claim 11, wherein the thermally insulating jacket integrates the function of a looks cover. 21. A cryostat according to claim 20, wherein the jacket comprises a solid material having molded channels for accommodating auxiliary equipment. 22.... | - | A61B005/05 F25B019/00 H01F006/04
US07657125 | B2 | 2010-02-02 | Time-lapsing data methods and systems | 2004-10-22 | Time-lapsing mirror methods and related systems. | 1. A method comprising: accepting input related to at least one remotely-captured patient image wherein said accepting input related to at least one remotely-captured patient image is performed via at least one of a machine, article of manufacture, or composition of matter; andpresenting to a medical expert one or more historical images related to at least a part of the at least one remotely-captured patient image in a time-lapse context, wherein said presenting to a medical expert one or more historical images related to at least a part of the at least one remotely-captured patient image in a time-lapse context is performed via at least one of a machine, article of manufacture, or composition of matter and wherein said presenting to a medical expert one or more historical images related to at least a part of the at least one remotely-captured patient image in a time-lapse context further comprises: identifying one or more anatomical landmarks demarcating the at least a part of the at least one remotely-captured patient image in a time-lapse context;obtaining one or more remotely-captured images having the one or more anatomical landmarks; andpresenting the one or more remotely-captured images having the one or more anatomical landmarks. 2. The method of claim 1, wherein said accepting input related to at least one remotely-captured patient image further comprises: accepting input of at least a part of a mirror image. 3. The method of claim 1, wherein said accepting input related to at least one remotely-captured patient image further comprises: accepting input of at least a part of an image representation. 4. The method of claim 1, wherein said accepting input related to at least one remotely-captured patient image further comprises: accepting input from data storage. 5. The method of claim 4, wherein said accepting input from data storage further comprises: accepting data over at least one of a wired connection channel or a wireless communication channel. 6. The method of claim 4, wherein said accepting input from data storage further comprises: retrieving data from the data storage in response to an input indicative of at least a part of a presented image. 7. The method of claim 1, wherein said presenting to a medical expert one or more historical images related to at least a part of the at least one remotely-captured patient image in a time-lapse context further comprises: presenting at least one of a specified region of the image or a recognized anchor point associated with the image. 8. The method of claim 1, wherein said presenting to a medical expert one or more historical images related to at least a part of the at least one remotely-captured patient image in a time-lapse context further comprises: locating one or more images having the at least a part of the at least one remotely-captured patient image in a time-lapse context. 9. The method of claim 1, wherein said presenting the one or more remotely-captured images having the o... | Searete LLC | G06K009/54 G06K009/00 A61B005/05
US20100041973 | A1 | 2010-02-18 | CATHETER RADIO FREQUENCY ADAPTER FOR WIRELESS COMMUNICATION | 2009-07-23 | A catheter system for wireless communication with an electrophysiological (EP) mapping system. The catheter system comprises a catheter, a catheter adapter, and a radio frequency receiver module. The catheter includes a plurality of mapping electrodes including a tip electrode disposed on a distal portion of the elongated body, the mapping electrodes detecting electrocardiograph (ECG) signals; and a reference electrode being disposed on the elongated body at a distance from the plurality of mapping electrodes such that the reference electrode substantially does not detect electrocardiograph signals. The catheter includes a handle. The catheter adapter is attached to the handle. The catheter adapter includes an RF transmitter module for receiving, processing, and transmitting the detected ECG signals. The reference electrode provides a reference signal to the radio frequency (RF) transmitter module. The RF receiver module receives the transmitted ECG signals. The RF receiver module is coupled to the EP mapping system. | 1. A catheter system for wireless communication with an electrophysiological (EP) mapping system, the catheter system comprising: a catheter comprising: an elongated body having a distal end, and a proximal end;a plurality of mapping electrodes including a tip electrode being disposed on a distal portion of the elongated body, the plurality of mapping electrodes detecting electrocardiograph (ECG) signals; anda reference electrode being disposed on the elongated body at a distance from the plurality of mapping electrodes such that the reference electrode substantially does not detect electrocardiograph (ECG) signals; anda handle;a catheter adapter attached to the handle, the catheter adapter including a radio frequency (RF) transmitter module for receiving, processing, and transmitting the detected ECG signals;wherein the reference electrode provides a reference signal to the radio frequency (RF) transmitter module; andan RF receiver module for receiving the transmitted ECG signals, the RF receiver module being coupled to the EP mapping system. | - | A61B005/042
US07648469 | B2 | 2010-01-19 | Method and apparatus for penetrating tissue | 2008-01-31 | These and other objects of the present invention are achieved in a body fluid sampling system for use on a tissue site that includes an electrically powered drive force generator. A penetrating member is operatively coupled to the force generator. The force generator moves the member along a path out of a housing having a penetrating member exit, into the tissue site, stops in the tissue site, and withdraws out of the tissue site. An analyte detecting member is positioned to receive fluid from a wound created by the penetrating member. The detection member is configured to determine a concentration of an analyte in the fluid using a sample of less than 1 μL of the fluid. | 1. A body fluid sampling system for use on a tissue site, the system comprising: an electrically powered drive force generator, wherein the drive force generator is selected from a voice coil drive force generator and a rotary voice coil drive force generator;a penetrating member operatively coupled to said force generator, said force generator moving said member along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site;an analyte detecting member positioned to receive fluid from a wound created by said penetrating member, said detecting member configured to determine a concentration of an analyte in the fluid using a sample of less than 1 μL of the body fluid;a programmable processor operatively programmed with software that has control instructions for controlling the drive force generator to provide controlled impact of the penetrating member to and through the tissue site, to provide controlled retraction velocity of the penetrating member from the tissue site, and to provide controlled dwell time of the penetrating member at the tissue site, with the penetrating member being withdrawn from the tissue site at a slower speed than a speed of initial penetration into the tissue site, and to determine a concentration of the analyte in the fluid using less than 1 μL of the fluid, wherein the control instructions are selected based on at least one of, penetration depth of the penetrating member and velocity of the penetrating member, the programmable processor providing a tissue dwell time that is related to an amount of skin deformation as the penetrating member tries to puncture a surface of the tissue site. 2. The system of claim 1, further comprising a penetrating position sensor that includes an interference encoder. 3. The system of claim 1, further comprising: a tissue stabilizer device coupled to the housing. 4. The system of claim 3, wherein the tissue stabilizer device is configured to enhance fluid flow from a target tissue. 5. The system of claim 3, wherein the tissue stabilizer device creates a stretching of a skin surface. 6. The system of claim 3, wherein the tissue stabilizer device includes a plurality of protrusions. 7. The system of claim 3, wherein the tissue stabilizer device creates a vacuum around a target tissue. 8. The system of claim 3, wherein the tissue stabilizer device is configured to apply a force to a target tissue and cause the target tissue to press in an inward direction relative to the housing member. 9. The system of claim 3, wherein the tissue stabilizing device generates a stimulation at a target tissue. 10. The system of claim 1, further comprising: a seal formed by a fracturable material that seals the penetrating member and positioned at a place of at least one of a distal port or a proximal port. 11. The system of claim 1 further comprising a vacuum source to provide a low pressure environment to d... | Pelikan Technologies Inc. | A61B005/00 A61B010/00 B65D081/00 A61B017/14 A61B017/32
US07666137 | B2 | 2010-02-23 | Method for analyzing mitochondrial function | 2008-04-10 | A method for analyzing the mitochondrial function of a subject involves: conducting a first test wherein the average oxygen consumption of the subject during exercise is measured while the subject is breathing air having a first oxygen concentration, preferably room air having a concentration of 20-21% by weight; conducting a second test wherein the average oxygen consumption of the subject during exercise is measured while the subject is breathing air having a second oxygen concentration, preferably air having an oxygen concentration of 30-33% by weight; measuring first and second energy quotients obtained from the first and second tests, respectively; and then using the energy quotients to calculate the subject's Mitochondrial Factor. The Mitochondrial Factor is directly proportional to the subject's mitochondrial efficiency and function. If conducted periodically, the method can be used to determine whether the subject's mitochondrial efficiency is increasing or decreasing over time. The rate of increase or decrease will indicate the subject's risk of developing one or more age-related diseases. | 1. A method for analyzing the mitochondrial function of a subject, comprising: 1) obtaining age, body fat percentage, weight and sex information from the subject;2) calculating the subject's predicted maximum oxygen consumption at a specified respiratory exchange ratio using an analyzer, the predicted maximum oxygen consumption being based on the subject's sex, body fat percentage, weight and age in years, wherein the age in years of a subject over a predetermined age is a default age, and the age in years of a subject under a predetermined age is the subject's actual age;3) coupling the subject to a pulmonary breath analyzer via a mouthpiece, and conducting a first test while the subject is breathing air having a first concentration of oxygen, the first test comprising: a) measuring the subject's average oxygen consumption at the specified respiratory exchange ratio when the subject is exercising; andb) calculating a first energy quotient for the subject by dividing the subject's average oxygen consumption measured in step 3a) by the subject's predicted maximum oxygen consumption to obtain a first energy quotient for the subject;4) conducting a second test while the subject is breathing air having a second concentration of oxygen administered by use of the mouthpiece, the first and second concentrations of oxygen being such that the amount of oxygen delivered to the subject's mitochondria in the second test is 50% higher than the amount of oxygen delivered to the subject's mitochondria in the first test, the second test comprising: a) measuring the subject's average oxygen consumption at the specified respiratory exchange ratio when the subject is exercising; andb) calculating a second energy quotient for the subject by dividing the subject's average oxygen consumption measured in step 4a) by the subject's predicted maximum oxygen consumption; and5) calculating a Mitochondrial Factor for the subject using the equation:Mitochondrial⁢⁢Factor=(EQ⁢⁢2-EQ⁢⁢1)×200EQ⁢⁢1EQ1 being the first energy quotient and EQ2 being the second energy quotient, wherein the Mitochondrial Factor is directly proportional to the subject's mitochondrial efficiency. 2. The method of claim 1, wherein the first oxygen concentration is about 20-21% by weight, and the second oxygen concentration is about 30-33% by weight. 3. The method of claim 1, wherein the first oxygen concentration is about 21% by weight, and the second oxygen concentration is about 31.5-32.5% by weight. 4. The method of claim 1, wherein the specified respiratory exchange ratio is about 1.00. 5. A method for determining whether a subject's mitochondrial efficiency is increasing or decreasing over time, comprising repeating the method of claim 1 on a periodic basis, wherein an increase or decrease of a subject's Mitochondrial Factor over time indicates, respectively, an increase or decrease of the subject's mitochondrial efficiency over time. 6. A method according to claim 5, wherein the method of claim ... | - | A61B005/00
US07647085 | B2 | 2010-01-12 | Method of and apparatus for investigating tissue histology | 2005-06-13 | Apparatus for monitoring the presence of one or more chromophores in a tissue sample, comprises a light source for projecting light to illuminate an area of such tissue sample, a photo-receptor for receiving light remitted by the illuminated area of tissue, and spectroscopic analyzer means for monitoring the remitted light, a comparator having means for comparing variations in the intensity and spectral characteristics of the remitted light with respect to the intensity and spectral characteristics of the projected light at different wavelengths and with a record of the intensity and spectral characteristics of light remitted by a reference sample of such tissue and means for emitting a control signal in response to any such variations. Methods of analyzing tissue histology, especially skin histology, are discussed, and a mathematical model is proposed for the analysis and comparison of the remitted light with a reference sample. | 1. Apparatus for monitoring the presence of individual chromophores in a tissue sample, said apparatus comprising: a light source configured to illuminate an area of a tissue sample with different wavelengths of light;a photo-receptor configured to receive light re-emitted by an area of a tissue sample illuminated by said light source; anda spectroscopic analyzer configured to utilize the spectral characteristics of re-emitted light received by said photo-receptor and the spectral characteristics of light projected by said light source to identify the presence of said one or more chromophores within the area of the illuminated tissue sample from which re-emitted light is received by said photo-receptor, wherein said spectroscopic analyzer comprises:records of the spectral characteristics of light re-emitted by reference samples of known structure containing identified chromophores;said spectroscopic analyzer being configured to compare the spectral characteristics of re-emitted light received by said photo-receptor with said records to identify the record including spectral characteristics most closely matching the characteristics of said re-emitted light received by said photo-receptor and to output as an identification of individual chromophores present in the area of the illuminated tissue sample from which re-emitted light is received by said photo-receptor an identification of the chromophores associated with said spectral characteristics by said record. 2. Apparatus in accordance with claim 1, wherein said records comprise records of the spectral characteristics of light remitted by reference samples of epithelial or epithelial and sub-epithelial tissue. 3. Apparatus in accordance with claim 2, wherein said records comprise records of the spectral characteristics of light remitted by reference samples of skin. 4. Apparatus in accordance with claim 1, wherein said records comprise records of the spectral characteristics of light remitted by reference samples of normal healthy tissue. 5. Apparatus in accordance with claim 1, wherein said records comprise records of calculated estimated spectral characteristics of light re-emitted by samples of tissue of known structure containing identified chromophores. 6. Apparatus in accordance with claim 1 wherein said light source is operable to selectively illuminate an area of a tissue sample with different wavelengths of light. 7. Apparatus in accordance with claim 6, wherein said light source further comprises a set of filters for selective substitution into the tissue-incident light path in order to selectively illuminate an area of tissue with different wavelengths of light. 8. Apparatus in accordance with claim 1, wherein said light source further comprises a polarizing filter arranged to polarize the light which illuminates said area of tissue. 9. Apparatus in accordance with claim 8, further comprising a cross-polarizing filter arranged to cross-polarize light remitted from said area of ... | Astron Clinica Limted | A61B005/05
US07650178 | B2 | 2010-01-19 | Magnetic field sensor-based navigation system to track MR image-guided interventional procedures | 2004-04-30 | A magnetic sensor navigation system is disclosed, which includes: (a) a magnetic sensor array, including one or more magnetic non-Faradaic sensors; (b) a sensor interface operatively connected to receive a signal from each sensor of the array, wherein the sensor interface operates to generate a processed signal from one or more signals received from each sensor of the array; and (c) a processing and control unit connected to receive the processed signal, wherein the processing and control unit operates to calculate a spatial position of the sensor array using the processed signal when the sensor array is placed within a magnetic resonance environment of an MR scanner. | 1. A magnetic sensor navigation system comprising: (a) a magnetic sensor array, including one or more magnetic non-Faradaic sensors;(b) a sensor interface operatively connected to receive a first signal from each sensor of the array, wherein the sensor interface operates to generate a processed signal from one or more first signals received from each sensor of the array;(c) an MR scanner having a magnetic resonance environment; and(d) processing and control unit connected to receive the processed signal, wherein the processing and control unit operates to calculate a spatial position of the sensor array using the processed signal when the sensor array is placed within the magnetic resonance environment of the MR scanner, wherein the MR scanner comprises i. a magnet generating a first homogenous static magnetic field along a z-axis of a measuring space of the MR scanner;ii. a gradient magnetic field coil assembly for generating a second gradient magnetic field in the measuring space, wherein the gradient magnetic field coil assembly includes three gradient magnetic field coils disposed so as to be linearly independent of one another; andiii. an RF coil for generating a third RF magnetic field in the measuring space,wherein the magnetic resonance environment comprises the first homogenous static magnetic field generated by the magnet of the MR scanner and the second gradient magnetic field generated by the gradient magnetic field coil assembly of the MR scanner and the third RF magnetic field generated by the RF coil of the MR scanner, and wherein the one or more magnetic non-Faradaic sensors use the first homogenous magnetic field and the second gradient magnetic field of the magnetic resonance environment to generate the one or more first signals. 2. A magnetic sensor navigation system as recited in claim 1, further comprising an interventional device, wherein the one or more magnetic non-Faradaic sensors are detachably attached to the interventional device. 3. A magnetic sensor navigation system as recited in claim 2, further comprising a base member, wherein the magnetic sensor array is fixed to the base member and the base member is detachably attached to the interventional device. 4. A magnetic sensor navigation system as recited in claim 2, wherein the one or more magnetic non-Faradaic sensors are selected from the group consisting of Hall sensors and magneto-optical sensors. 5. A magnetic sensor navigation system as recited in claim 4, wherein the magnetic sensor array comprises three Hall sensors arranged mutually orthogonal to each other. 6. A magnetic sensor navigation system as recited in claim 4, wherein the magnetic sensor array comprises a single three-axis Hall sensor. 7. A magnetic sensor navigation system as recited in claim 4, wherein the magnetic sensor array consists of three Hall sensors arranged mutually orthogonal to each other. 8. A magnetic sensor navigation system as recited in claim 4, wherein the magnetic sensor ... | University of Basel | A61B005/05
US07650176 | B2 | 2010-01-19 | Physiological stress detector device and system | 2006-12-19 | A non-invasive device and a system for monitoring and measuring blood saturation and heart pulse rate of a baby or infant is provided. The device includes a housing unit configured to be integrated within apparatus, which is attachable proximate to a limb being measured. The housing unit includes at least one light source, providing light directed toward the surface of the limb, a light detector spaced apart from the light source and sensitive to intensity levels of the light reflected from the limb and a processing unit for processing the intensity signals received from the light detector for producing output signals. The device may determine the level of the blood constituent and may also use this level for monitoring and/or to activate an alarm when the level falls outside a predetermined range. | 1. A non-invasive device for monitoring and measuring blood saturation and heart pulse rate of a baby or infant, comprising: a housing unit configured to be integrated within an apparatus, said apparatus being configured to be attachable proximate to a limb being measured, said apparatus comprising one of a group of devices including a bracelet, a sock, a ribbon and a diaper, said housing unit comprising: at least one light source, providing light directed toward the surface of said limb, the light being reflected from said limb;a light detector spaced apart from said at least one light source and being sensitive to intensity levels of said reflected light for producing intensity signals in accordance therewith; anda processing unit for processing said intensity signals received from said light detector for producing output signals, said processing unit comprising: first and second amplifiers for amplifying said intensity signals, each in accordance with a respective first and second gain amplification factor; anda processor for automatically determining said first and second gain amplification factors in adjustable fashion;wherein during a first stage, said first and second amplifiers amplify a DC signal component of said intensity signals in accordance with predetermined first and second gain amplification factors, said amplified DC signal component being subtracted from said intensity signals at an input of said first amplifier, to isolate an AC signal component of said intensity signals,and wherein during a second stage, said second amplifier amplifies said isolated AC signal component in accordance with said adjustably-determined second gain amplification factor,said processing unit producing output signals in accordance with said isolated AC signal component and said DC signal component and calculating in accordance therewith, at least one blood constituent level. 2. The device according to claim 1, wherein said processing unit comprises mapping means for mapping the intensity of said AC signal component along the surface of said limb to detect regions of said limb having a reduced blood flow. 3. A non-invasive device for monitoring and measuring blood saturation and heart pulse rate of a baby or infant, comprising: a housing unit configured to be integrated within an apparatus, said apparatus being configured to be attachable proximate to a limb being measured, said apparatus comprising one of a group of devices including a bracelet, a sock, a ribbon and a diaper, said housing unit comprising: at least one light source, providing light directed toward the surface of said limb, the light being reflected from said limb;a light detector spaced apart from said at least one light source and being sensitive to intensity levels of said reflected light for producing intensity signals in accordance therewith; anda processing unit for processing said intensity signals received from said light detector for producing output signals,wherein processin... | SPO Medical Equipment Ltd. | A61B005/1455 A61B005/02
US07666150 | B2 | 2010-02-23 | Blood and interstitial fluid sampling device | 2004-04-29 | A device and method for lancing a patient, virtually simultaneously producing and collecting a small fluid sample from a body. The device comprises a blood collection system including a lancing needle (16), drive mechanism (11), kneading or vibration mechanism (25), optional suction system (7), and sample ejection mechanism. The device is preferably sized to be hand-held in one hand and operable with one hand. The device can optionally contain integral testing or analysis component (83) for receiving the sample and providing testing or analysis indication or readout for the user. A method involves piercing the skin at a rapid rate, kneading the surrounding area by ultrasonic action, piezoelectric or mechanical oscillation to stimulate the blood flow from the wound, drawing the fluid using a pumping system. | 1. A tissue penetration sampling device for collecting blood from the skin of a patient, the device comprising: a sampling module having a sampling site, the sampling site having a sample input port,a sample reservoir in fluid communication with the sample input port,a tissue penetrating member maintained within the sampling module, the tissue penetrating member having a tip adjacent the sample input port, anda driver coupled to the tissue penetrating member to drive the tip through the opening to penetrate the skin when the driver is actuated,the device being configured to provide actuation of the driver, penetration of the skin, application of an electric current to the skin to stimulate blood flow, collection of the blood, and movement of the blood to the sample reservoir. 2. The tissue penetration sampling device of claim 1, further comprising at least one electrode adapted to contact the skin and deliver the electric current to the skin. 3. The tissue penetration sampling device of claim 2, wherein at least one electrode is a stimulator ring. 4. The tissue penetration sampling device of claim 3, wherein the stimulator ring oscillates. 5. The tissue penetration sampling device of claim 2, wherein at least one electrode is a guard, and wherein the tissue penetrating member has a retracted position and the guard is adapted to contact the skin while preventing the tissue penetrating member from contacting the skin in the retracted position. 6. The tissue penetration sampling device of claim 5, further comprising a stimulator ring adapted to contact the skin and oscillate. 7. The tissue penetration sampling device of claim 2, wherein the sample reservoir is an at least one electrode. 8. The tissue penetration sampling device of claim 2, wherein the tissue penetrating member is an at least one electrode. 9. The tissue penetration sampling device of claim 1, wherein the device is configured to provide low voltage DC current to the skin. 10. The tissue penetration sampling device of claim 9, wherein the DC voltage is in the range of one (1) millivolt to twelve (12) volts. 11. The tissue penetration sampling device of claim 9, wherein the DC current is pulsed. 12. The tissue penetration sampling device of claim 1, wherein the device is configured to provide low voltage AC current to the skin. 13. A method of penetrating a target tissue, comprising: providing a tissue penetrating system with a penetrating member, and an analyte detecting member coupled to a sample chamber;advancing the penetrating member through the target tissue;withdrawing the penetrating member from the target tissue;drawing an adjustable sample of body fluid from the target tissue; andreceiving the volume of body fluid in the sample chamber. 14. The method of claim 13, further comprising: drawing three (3) microliters (μL) of body fluid. 15. The method of claim 13, further comprising: drawing eight (8) microliters (μL) of body fluid. 16. A skin penetrating system, c... | Roche Diagnostics Operations, Inc. | A61B005/00
US08351667 | B2 | 2013-01-08 | Methods of contrast enhancement for images having blood vessel structures | 2008-08-15 | A method of contrast enhancement for improved visualization of diagnostically important tissue structures, such as blood vessels. A texture analysis algorithm is applied to identify regions with a high likelihood of disease. Mathematical morphology operations are applied to identify areas of high and low brightness (intensity). The low intensity areas are then subtracted, and controllably variable amounts of the high intensity areas are added, controlled by a selectable tuning parameter, to produce an image with controllably variable visualization enhancement. | 1. A method of contrast enhancement for an image having blood vessel structures, comprising: identifying a texture region having high spatial variations in brightness in said image;defining high intensity areas and low intensity areas in said texture region;controllably variably increasing contrast between said high intensity areas and said low intensity areas and controlling generated noise by subtracting only said low intensity areas from said texture region and using only a single user-controlled tuning parameter to controllably variably add said high intensity areas to said texture region without controlling said subtracting of said low intensity areas;whereby visualization of said blood vessel structures is controllably variably enhanced. 2. A method according to claim 1, wherein said defining step comprises applying a top-hat filter using only a first structuring element to define said high intensity areas in said texture region, and using a bottom-hat filter having a second structuring element to define said low intensity areas in said texture region. 3. A method according to claim 2, wherein said first structuring element is disk shaped and said second structuring element is disk shaped and has a diameter slightly exceeding the diameter of blood vessels in said blood vessel structures. 4. A method according to claim 1, wherein said identifying step is performed by segmenting said image using intensity variations in said image. 5. A method of contrast enhancement for an image having blood vessel structures, comprising: defining high intensity areas and low intensity areas in said image;controllably variably increasing contrast between said high intensity areas and said low intensity areas and controlling generated noise by subtracting only said low intensity areas from said image and using only a single user-controlled tuning parameter to controllably variably add said high intensity areas to said image, without controlling said subtracting of said low intensity areas;whereby visualization of said blood vessel structures is controllably variably enhanced. 6. A method according to claim 5, wherein said defining step is performed by applying a top-hat filter using only a first structuring element to define said high intensity areas, and using a bottom-hat filter having a second structuring element to define said low intensity areas. | STI Medical Systems, LLC | G06K009/18
US08150118 | B2 | 2012-04-03 | Image recording apparatus, image recording method and image recording program stored on a computer readable medium | 2007-03-26 | There is provided an image recording apparatus for assisting an observer who observes an observed person. The image recording apparatus includes an image-capturing section that captures an image of the observed person, an observation result input section that inputs a result of the observation done on the observed person by the observer, an observer's point of sight position judging section that judges whether the observer looks at the observed person when the image-capturing section captures the image of the observed person, an image extracting section that extracts an image of the observed person for which the observer's point of sight position judging section judges that the observer looks at the observed person, a physical status quantifying section that quantifies a physical status of the observed person by analyzing the image of the observed person which is extracted by the image extracting section, for an observation result which is input by the observation result input section and corresponds to the image of the observed person which is extracted by the image extracting section, and a physical status data storing section that stores thereon physical status quantified data which is generated by the quantification by the physical status quantifying section, in association with the corresponding observation result. | 1. An image recording apparatus for assisting an observer who observes an observed person, comprising: an image-capturing section that captures the image of the observed person;an observation result input section that inputs a result of the observation done on the observed person by the observer;an observer's point of sight position judging section that judges whether the observer looks at the observed person when the image-capturing section captures the image of the observed person;an image extracting section that extracts an image of the observed person for which the observer's point of sight position judging section judges that the observer looks at the observed person;a physical status quantifying section that quantifies a physical status of the observed person by analyzing the image of the observed person which is extracted by the image extracting section, for an observation result which is input by the observation result input section and corresponds to the image of the observed person which is extracted by the image extracting section; anda physical status data storing section that stores thereon physical status quantified data which is generated by the quantification by the physical status quantifying section, in association with the corresponding observation result. 2. The image recording apparatus as set forth in claim 1, further comprising: an observer's point of sight measuring section that measures a point of sight of the observer when the image-capturing section captures the image of the observed person; andan image storing section that stores the observation result input by the observation result input section and observer's point of sight position information indicating the point of sight of the observer which is measured by the observer's point of sight measuring section, in association with the image of the observed person which is captured by the image-capturing section, whereinthe observer's point of sight position judging section judges, based on the observer's point of sight position information stored on the image storing section, whether the observer looks at the observed person when the image-capturing section captures the image of the observed person which is stored on the image storing section in association with the observer's point of sight position information, andthe image extracting section extracts the image of the observed person which is stored on the image storing section in association with the observer's point of sight position information based on which the observer's point of sight position judging section judges that the observer looks at the observed person. 3. The image recording apparatus as set forth in claim 1, further comprising: an image input section that inputs an image of a person under observation;a physical status determining section that determines a physical status of the person under observation by analyzing the image of the person under observation which is input by the image input sectio... | Fujifilm Corporation | G06K009/00
US07950971 | B2 | 2011-05-31 | Electrical connector apparatus and methods | 2008-08-08 | Size-specific electrical connector apparatus optimized for e.g., biomedical applications, and which prevents mating to the wrong sized terminals. In one embodiment, the connector apparatus comprises a device having a two piece assembly adapted to pivotally engage and lock a sensor such as those used in impedance cardiography (ICG). The conductor is shaped so as to provide a high degree of specificity in sensor connection as well as signal stability and reliability. In one variant, the pivot comprises a spring-loaded pivot having no external stops and the shape of a sensor engaging aperture is defined by a metal insert associated with one element of the assembly and the pivot is permanently formed within the connector. Thus, a user is obstructed from changing the properties of the connector to accept smaller or larger sensors. Methods of manufacturing and operating the connector are also disclosed. In one embodiment, an opening in the metal insert (212) is sized to exclude oversized sensors. In another embodiment, a V-shaped stop (308) on one of the elements (304) acts to limit closure and to prevent a tight grip upon an undersized sensor. | 1. An electrical connector adapted for use with a biomedical sensor of a particular size, said connector comprising: a first actuation element having at least one electrically conductive element associated therewith;a second actuation element, at least a portion of said second actuation element adapted to be mated with said first actuation element, said mating comprising forming at least one feature for receiving at least a portion of said sensor; anda pivot element adapted to enable said actuation elements to pivot with respect to one another, said pivoting causing an effective diameter of said feature for receiving said at least a portion of said sensor to adjust;wherein said pivot element is configured to ensure that, upon said pivoting of said actuation elements to a resting position without the presence of the sensor, said effective diameter of said feature is maintained at a pre-set amount corresponding to the particular size of the biomedical sensor. 2. The electrical connector of claim 1, wherein said pivot element is adapted to ensure that said effective diameter of said feature is said pre-set amount via one or more internal stopping features. 3. The electrical connector of claim 2, wherein said one or more internal stopping features are not accessible to a user of said connector when said actuation elements are mated. 4. The electrical connector of claim 1, wherein said pre-set amount comprises approximately 0.125 inches. 5. The electrical connector of claim 1, wherein said pre-set amount comprises approximately 0.155 inches. 6. The electrical connector of claim 1, wherein said mating of said actuation elements further comprises forming at least one channel adapted to receive said pivot element, and wherein said connection of said actuation elements by said pivot element comprises permanently receiving said pivot element within said channel. 7. The electrical connector of claim 6, wherein said receipt of said pivot element in said channel is made permanent by an asymmetric shaped portion of said pivot element such that once said pivot element is received within said channel, said asymmetric portion substantially prevents said pin element from being removed. 8. The electrical connector of claim 1, wherein said at least one electrically conductive element associated with said first actuation element comprises a metallic disk molded onto a surface thereof, said disk placed on said first actuation element so as to comprise at least a portion of said feature for receiving at least a portion of said sensor. 9. The electrical connector of claim 8, wherein said disk is at least partly chamfered along a contact surface thereof. 10. The electrical connector of claim 1, wherein said first actuation element further comprises at least one resistor disposed therein, said resistor adapted to dissipate energy. 11. The electrical connector of claim 10, wherein said at least one resistor does not dissipate more than ten percent (10%) of the to... | CardioDynamics International Corporation | H01R004/48
US09211155 | B2 | 2015-12-15 | Non-thermal ablation system for treating BPH and other growths | 2009-08-19 | Systems and methods for non-thermal ablation of tissue are provided. A non-implantable minimally invasive system for treatment of tissue in a body via direct current ablation is provided including a catheter, a plurality of electrodes for deployment through the catheter, a power source for applying power to the electrodes, and a fixation element for maintaining the catheter in a treatment position during treatment of the tissue. A minimally invasive method for treating tissue in a body via direct current ablation is provided including inserting a catheter into the body such that a portion of the catheter remains outside of the body, deploying a fixation element to fix the catheter in a treatment position, deploying a plurality of electrodes through the catheter, applying power to the plurality of electrodes, using the electrodes to apply a current to the tissue, and removing the catheter from the body. | 1. A minimally invasive method for treating prostate tissue of a prostate in a body via direct current ablation from a urethra transecting the prostate, comprising: inserting a catheter into the urethra, wherein a portion of the catheter remains outside of the body when the catheter is in a treatment position proximate the prostate;deploying a fixation element associated with the catheter to fix the catheter in the treatment position;deploying a plurality of electrodes through the catheter into the prostate, wherein at least one of the plurality of electrodes is deployed into at least one of a lateral posterior quadrant of the prostate tissue relative to the urethra, a lateral anterior quadrant of the prostate tissue relative to the urethra, or combinations of the prostate tissue relative to the urethra thereof, and wherein the electrodes are separated by an angular spacing;determining input treatment parameters, wherein the input treatment parameters include a quantity of charge to the prostate tissue;inputting the input treatment parameters into at least one of a power source and a processor;using a power source to apply direct current between approximately 10 and 50 mA and a power of less than 3200 mW to each of the plurality of electrodes to deliver the quantity of charge, wherein the direct current applied is based on the input treatment parameters;wherein upon application of the direct current, at least one of the plurality of electrodes imparts at least one of a high pH and a low pH in a treatment zone proximate to each electrode of the plurality of electrodes to create a necrotic zone in the treatment zone, wherein the application of the direct current is substantially non-thermal;comparing a measured treatment parameter with at least one of the input treatment parameters; andstopping the application of direct current to the treatment zone once the measured treatment parameter reaches the at least one input treatment parameter, wherein the at least one input treatment parameter is reached and the application of direct current to the plurality of electrodes is stopped between approximately 15 and 45 minutes with the catheter in the treatment position. 2. The minimally invasive method of claim 1, wherein the plurality of electrodes are the fixation element and deploying the fixation element and deploying the plurality of electrodes comprises one step of deploying the plurality of electrodes. 3. The minimally invasive method of claim 1, wherein the treatment zone is created between a bladder neck and a verumontanum of the prostate and wherein the minimally invasive method treats benign prostatic hyperplasia. 4. The minimally invasive method of claim 1, further comprising determining a prostate size to determine the input treatment parameters. 5. The minimally invasive method of claim 4, wherein an ultrasound, a cystoscopy, an MRI, a PET, or a CT is used to determine the prostate size. 6. The minimally invasive method of claim 4, further... | Prostacare Pty Ltd. | A61B018/18 A61B018/14 A61B005/055 A61B008/08 A61B019/00 A61B018/00 A61B018/12
US08521271 | B2 | 2013-08-27 | Brain wave identification method adjusting device and method | 2007-10-04 | In a system having an interface that utilizes an electroencephalogram, there is realized a calibration which prevents overlooking of calibrational stimulations concerning the electroencephalogram and which is highly accurate.;An electroencephalogram interface system (1) includes an interface section (100) which acquires an electroencephalogram signal from a user, distinguishes a P3 component of a visual event-related potential contained in the electroencephalogram signal after a manipulation menu is presented, and operates a device based on the distinguished P3 component. An electroencephalogram distinction method adjustment apparatus (10) is used for adjusting a distinction method in the interface section (100). Based on a database (13) defining a correlation between a P3 component of an event-related potential obtained from a stimulation to a modality other than the visual sense and a P3 component of a visual event-related potential and on a P3 component of an event-related potential after a stimulation is presented, the apparatus (10) derives a characteristic quantity of the user concerning the P3 component of the visual event-related potential, and adjusts the electroencephalogram distinction method in the interface section (100) based on the characteristic quantity. | 1. In an electroencephalogram interface system having an output section configured to visually present a manipulation menu for a device and to present a non-visual stimulation to a modality other than the visual sense,a biological signal measurement section configured to acquire an electroencephalogram signal from a user, andan electroencephalogram interface section configured to distinguish, by using a prestored parameter of visual event-related potential, a P3 component of a visual event-related potential contained in an electroencephalogram signal after the manipulation menu is presented, and to operate the device based on the distinguished P3 component,an electroencephalogram distinction method adjustment apparatus to be used for adjusting a distinction method in the electroencephalogram interface section, comprising:a database defining a correlation between a P3 component of an event-related potential obtained from a stimulation to a modality other than the visual sense and a P3 component of a visual event-related potential;a stimulation presentation section configured to present the non-visual stimulation via the output section;an analysis section configured to analyze an event-related potential contained in the electroencephalogram signal after the non-visual stimulation is presented; anda distinction method adjustment section configured to derive at least one of an amplitude and a latent period of a P3 component as a characteristic quantity of the user concerning the P3 component of a visual event-related potential based on a P3 component of the analyzed event-related potential from the non-visual stimulation and the correlation in the database, and to adjust the electroencephalogram distinction method in the electroencephalogram interface section based on the characteristic quantity. 2. The electroencephalogram distinction method adjustment apparatus of claim 1, wherein the stimulation presentation section presents an auditory stimulation as the non-visual stimulation. 3. The electroencephalogram distinction method adjustment apparatus of claim 1, wherein the stimulation presentation section presents a somatic stimulation as the non-visual stimulation. 4. The electroencephalogram distinction method adjustment apparatus of claim 1, wherein, the electroencephalogram interface section distinguishes presence or absence of a P3 component of a visual event-related potential of the user based on a threshold value which is set as the parameter; andthe distinction method adjustment section adjusts the threshold value based on the characteristic quantity. 5. The electroencephalogram distinction method adjustment apparatus of claim 1, wherein, the electroencephalogram interface section distinguishes presence or absence of a P3 component of a visual event-related potential of the user based on a template waveform of visual event-related potential which is set as the parameter; andthe distinction method adjustment section adjusts the template bas... | Panasonic Corporation | A61B005/0476 A61B005/0484
US09713701 | B2 | 2017-07-25 | Using multiple diagnostic parameters for predicting heart failure events | 2008-07-31 | Techniques for using multiple physiological parameters to provide an early warning for worsening heart failure are described. A medical device monitors a primary diagnostic parameter that is indicative of worsening heart failure, such as intrathoracic impedance or pressure, and one or more secondary diagnostic parameters. The medical device detects worsening heart failure in the patient based on the primary diagnostic parameter when an index that is changed over time based on the primary diagnostic parameter value is outside a range of values, termed the threshold zone. When the index is within the threshold zone, the medical device detects worsening heart failure in the patient based on the one or more secondary diagnostic parameters. Upon detecting worsening heart failure, the medical device may, for example, provide an alert that enables the patient to seek medical attention before experiencing a heart failure event. | 1. A method comprising: monitoring, by a processor, at least one primary diagnostic parameter and at least one secondary diagnostic parameter of a patient, wherein the primary and secondary diagnostic parameters are associated with worsening heart failure;changing, by the processor, an index value over time based on the primary diagnostic parameter, wherein the index indicates worsening heart failure of the patient;determining, by the processor, whether the index is within a threshold zone defined by a lower threshold and an upper threshold;detecting, by a processor, worsening heart failure in the patient when the index is greater than the upper threshold;anddetermining, by the processor, whether worsening heart failure is detected in the patient based on the secondary diagnostic parameter when the index is within the threshold zone. 2. The method of claim 1, further comprising determining that worsening heart failure is not detected in the patient when the index is less than the lower threshold. 3. The method of claim 1, wherein monitoring at least one secondary diagnostic parameter comprises monitoring a plurality of secondary diagnostic parameters, andwherein determining whether worsening heart failure is detected in the patient based on the secondary diagnostic parameter comprises determining whether worsening heart failure is detected in the patient when any of the secondary diagnostic parameters indicates worsening heart failure. 4. The method of claim 1, wherein the primary diagnostic parameter comprises intrathoracic impedance. 5. The method of claim 1, wherein the primary diagnostic parameter comprises cardiovascular pressure. 6. The method of claim 1, wherein the at least one secondary diagnostic parameter comprises one or more of atrial fibrillation (AF) burden, ventricular rate during AF, ventricular fibrillation (VF) burden, ventricular rate during VF, atrial tachyarrhythmia (AT) burden, ventricular rate during AT, ventricular tachyarrhythmia (VT) burden, or ventricular rate during VT. 7. The method of claim 1, wherein the at least one secondary diagnostic parameter comprises one or more of activity level, heart rate variability, percentage of cardiac resynchronization pacing, night heart rate, difference between day heart rate and night heart rate, heart rate turbulence, heart rate deceleration capacity, or baroreflex sensitivity. 8. The method of claim 1, wherein the at least one secondary diagnostic parameter comprises one or more of respiratory rate, respiratory depth, or respiratory pattern. 9. The method of claim 1, wherein the at least one secondary diagnostic parameter comprises one or more of weight, blood pressure, metrics of renal function, medication history, or history of heart failure hospitalization. 10. The method of claim 1, wherein the threshold zone is static over time. 11. The method of claim 1, further comprising changing the threshold zone over time. 12. The method of claim 11, wherein changing the th... | Medtronic, Inc. | A61N001/365 A61M037/00 A61B005/046 A61B005/0464 A61B005/053 A61B005/11
US20100030110 | A1 | 2010-02-04 | Glucose meter with er:yag laser lancing device | 2006-04-21 | A personal portable painless lancing device and a blood glucose meter using the same, which can lance at the minimum without any pain and use it to measure the blood glucose. The minimum painless lancing device using an Er:YAG laser includes an Er:YAG laser beam-generating section in which the laser beam is generated to punch skin and radiated in a single direction; an adapter installed at the front end of the Er:YAG laser beam-generating section so that the laser beam can be locally concentrated on a skin punch portion; a skin adhesion cap installed at the front end of the adapter so that it can be easily detached, and can reduce the contamination of the Er:YAG laser beam-generating section, which is produced at the time of the lancing, and the danger due to the reuse of it among the patients of lancing; and a case to accommodate the Er:YAG laser beam-generating section and the adapter therein. | 1. A painless minimum lancing device using an Er:YAG laser, comprising: an Er:YAG laser beam-generating section to generate a laser beam required to punch skin;an adapter installed at the front end of the Er:YAG laser beam-generating section so that the laser beam can be locally concentrated on a skin punch portion; anda case to accommodate the Er:YAG laser beam-generating section and the adapter therein. 2. The painless minimum lancing device according to claim 1, further comprising a skin adhesion cap detachably installed at the front end of the adapter 3. The painless minimum lancing device according to claim 2, wherein the skin adhesion cap includes: a body penetratingly formed internally so that the laser beam can pass through, secured detachably to the front end of the adapter at a rear end portion thereof, and inclined at a front end thereof so that it can be in close contact with the skin; andan inspection film installed at the rear end of the body so as to inspect if the skin adhesion cap is used, and to prevent the dispersion of the blood occurring at the time of the lancing and the contamination of the adapter due to the smoke, and to remove the problem of infection of the patient due to reuse of the skin adhesion cap. 4. The painless minimum lancing device according to claim 3, wherein the front end portion of the body of the skin adhesion cap is inclined by 68° with respect to the laser beam passing through the inside of the body. 5. The painless minimum lancing device according to claim 1, wherein the Er:YAG laser beam-generating section includes: a reflector having a receiving space formed at the inside;an Er:YAG crystal disposed in the receiving space of the reflector;a total reflection mirror installed at the rear surface of the Er:YAG crystal;a semi-permeable mirror installed at the front end of the Er:YAG crystal;a flash lamp installed at the inside of the receiving space in such a manner as to be spaced apart by a certain interval from the Er:YAG crystal;front and rear surface securing covers respectively installed at both front and rear ends of the reflector so that both ends of the Er:YAG crystal and the flash lamp can be supported; anda front and rear surface securing holder, inserted into the outside of both sides of the front and rear ends of the reflector, to integrally secure the front and rear surface securing cover and the reflector. 6. The painless minimum lancing device according to claim 2, wherein the Er:YAG laser beam-generating section includes: a reflector having a receiving space formed at the inside;an Er:YAG crystal disposed in the receiving space of the reflector;a total reflection mirror installed at the rear surface of the Er:YAG crystal;a semi-permeable mirror installed at the front end of the Er:YAG crystal;a flash lamp installed at the inside of the receiving space in such a manner as to be spaced apart by a certain interval from the Er:YAG crystal;front and rear surface securing covers respective... | - | A61B005/151
US07662102 | B2 | 2010-02-16 | Method of measuring instantaneous arterial blood pressure and compliance and device thereof | 2008-02-19 | A method of measuring the instantaneous arterial blood pressure and compliance and the device thereof are disclosed. The method identifies a critical depth under the skin. At this critical position, the arterial mean blood pressure is equivalent to the mean blood pressure obtained on the skin. The variation in the arterial mean blood pressure is tracked and controlled to maintain the equivalence between the mean blood pressure measured on the skin and the arterial mean blood pressure. In this case, the compliance of the muscle tissues remains invariant. Therefore, the variation measured on the skin is simply the behavior of the blood vessel. Using the skin-artery correlation, the invention can obtain the instantaneous blood pressure after some computation and the dynamic compliance at the critical position. | 1. A method for measuring the arterial instantaneous blood pressure and compliance, comprising: a searching step: finding the critical depth λs when Ps= Pb and determining Ps and λs using an actuator to impose a pressure on the skin, a pressure sensor to measure a skin mean blood pressure Ps, and a displacement sensor to detect the depth λs under the pressure, then using the pressure sensor to measure the peak-to-peak pressure value, where the above-mentioned procedure is repeated until the maximum peak-to-peak value is reached, with the corresponding depth being the critical depth λs and the mean blood pressure measured at the skin being equal to the arterial mean blood pressure Pb; wherein Ps is the mean blood pressure measured at the skin, Pb is the arterial mean blood pressure, and λs is the critical depth of the skin;a controlling step: using the pressure sensor to monitor the blood pressure change ΔPs transmitted from the artery to the skin when the arterial blood pressure changes by ΔPb at the critical depth λs, using the actuator to change the depressed depth of the skin for generating the corresponding displacement change Δλs so that the blood pressure change at the skin ΔPs approaches zero, in which case the muscle tissue compliance remain invariant, the tissue stiffness Kt being equal to the tissue impedance Zt and thus making the mean blood pressure at the skin Ps and the arterial mean blood pressure Pb remain equivalent, where the displacement change Δλs of the actuator detected by the displacement sensor is the variation in the arterial blood vessel diameter; wherein ΔPb is the difference between the arterial blood pressure and the mean blood pressure, Δλs is the displacement change from the critical depth λs for the arterial mean blood pressure change of ΔPb;an identifying step: simulating the artery vessel as a second-order system with mass, damping, and stiffness and performing system identification to obtain blood vessel parameters Mv, Dv, and Kv as well as the artery impedance Zv;an estimating step: transmitting the detected signal to a controlling unit for computations, using the mean blood pressure at the skin Ps, the artery impedance Zv, and the arterial blood vessel diameter change Δλs to estimate the arterial instantaneous blood pressure {circumflex over (P)}b, and using the arterial instantaneous blood pressure {circumflex over (P)}b and the diameter displacement change Δλv to make a plot of the arterial dynamic compliance at the critical position. 2. The method of claim 1, wherein the arterial blood vessel diameter change is given byΔ⁢⁢λs=-1/Zv⁢Δ⁢⁢Pb=Δ⁢⁢λv. 3. The method of claim 1, wherein the arterial instantaneous blood pressure is estimated using the formula {circumflex over (P)}b= Pb+ΔPb= Ps−ZvΔλs. 4. The method of claim 1, wherein the arterial dynamic compliance is computed usingC=[ⅆ(Δλs)/ⅆP^b]⁢¦P_b=P_s. 5. The method of claim 1, wherein when the arterial mean blood pressure Pb changes by Δ Pb so that the diame... | E-Med Biotech, Inc. | A61B005/02
US20100023344 | A1 | 2010-01-28 | System and device for maintaining physiological levels of steroid hormone in a subject | 2008-10-23 | A system, device, or method is provided which maintains a substantially physiological level or a substantially physiological cyclic pre-menopausal level of one or more steroid hormones in a mammalian subject in need thereof. The system, device, or method includes providing to the subject at least one treatment regimen including replacement therapy for the one or more steroid hormones or metabolites or modulators thereof, wherein the at least one treatment regimen is determined based on pre-menopausal cyclic steroid hormone levels of the subject and on current cyclic steroid hormone levels of the subject, wherein the at least one treatment regimen is configured to maintain the subject's one or more steroid hormones or metabolites or modulators thereof at substantially physiological cyclic pre-menopausal levels. The system, device, or methods can be used for the treatment of a disease or condition in the mammalian subject. | 1.-23. (canceled) 24. A system comprising: at least one computer program included on a computer-readable medium for use with at least one computer system wherein the computer program includes a plurality of instructions including,one or more instructions for determining at least one treatment regimen including replacement therapy for one or more steroid hormones or metabolites or modulators thereof in a subject, based on steroid hormone levels prior to disease diagnosis in the subject and on current steroid hormone levels in the subject. 25. The system of claim 24, further comprising one or more instructions for inputting information associated with the steroid hormone levels prior to disease diagnosis in the subject; and one or more instructions for inputting information associated with the current cyclic steroid hormone levels in the subject. 26. The system of claim 24, wherein the at least one treatment regimen is configured to maintain a substantially physiological level of one or more steroid hormones in the subject in need thereof. 27. The system of claim 24, further comprising one or more instructions for determining the one or more steroid hormones levels in the subject during a treatment period. 28. The system of claim 24, further comprising one or more instructions for determining at least one second treatment regimen adjusted to maintain the subject's one or more steroid hormones or one or more metabolites or modulators thereof at substantially physiological levels of the subject. 29. A system comprising: circuitry for determining at least one treatment regimen including replacement therapy for one or more steroid hormones or metabolites or modulators thereof in a subject, based on steroid hormone levels prior to disease diagnosis in the subject and on the current steroid hormone levels in the subject. 30. The system of claim 29, further comprising circuitry for inputting information associated with the steroid hormone levels prior to disease diagnosis in the subject; and circuitry for inputting information associated with the current cyclic steroid hormone levels in the subject. 31. The system of claim 29, wherein the at least one treatment regimen is configured to maintain a substantially physiological level of one or more steroid hormones in the subject in need thereof. 32. The system of claim 29, further comprising one or more instructions for determining the one or more steroid hormones levels in the subject during a treatment period. 33. The system of claim 29, further comprising one or more instructions for determining at least one second treatment regimen adjusted to maintain the subject's one or more steroid hormones or one or more metabolites or modulators thereof at substantially physiological levels of the subject. 34. A device comprising: a system including at least one computer program included on a computer-readable medium for use with at least one computer system and wherein the computer program includes a pl... | Searete LLC, a limited liability corporation of the State of Delaware | G06Q050/00 A61B005/00 A61M001/00
US20100004522 | A1 | 2010-01-07 | Continuously Wearable Compact Blood Glucose Measuring Device | 2009-06-30 | A wearable blood glucose measuring device based device and kit that incorporates non-continuous, invasive, electrochemical glucose test strip technologies for measuring blood glucose concentrations. The device can be inconspicuously worn due to its compact nature and integration into commonly worn articles such as watches, bracelets, armbands, wristbands, and the like. All-in-one embodiments include integrated test strip management including automatic test strip dispensing/ejection as well as incorporating spring loaded lancet apparatus capable of the dispensing, arming and activating functions enabling the user to provide a blood sample for glucose testing. Some embodiments do not integrate, nor contain any lancing devices/lancets. Other embodiments contain secure storage compartments integrated into the wearable blood glucose measuring device where the user can choose to store loose glucose test strips, calibration strips, lancets, cleaning wipes, and any combination thereof. | 1. A continuously wearable device adapted to receive a compact blood glucose measuring device comprising a housing that accommodates a compact blood glucose measuring device using invasive, electrochemical, single use glucose test strip based technology, wherein the housing is adapted to receive a fastening means that cooperates with the housing to provide a means for mounting the continuously wearable device on a limb. 2. The continuously wearable device of claim 1, wherein said continuously wearable device further comprises a storage means, whereby said storage means provides a holding area for blood glucose accessories. 3. The continuously wearable device of claim 2, wherein said storage means is substantially airtight. 4. The continuously wearable device of claim 1, wherein said continuously wearable device further comprises a clock. 5. The continuously wearable device of claim 2, wherein said storage means is comprised of at least two discrete storage areas. 6. The continuously wearable device of claim 1, wherein said continuously wearable device further comprises a casual memory storage device that cooperates with USB readers. 7. The continuously wearable device of claim 1, wherein said compact blood glucose measuring device is removably attachable to the continuously wearable device. 8. A continuously wearable compact blood glucose measuring device comprising a housing containing a compact blood glucose measuring device using invasive, electrochemical, single use glucose test strip based technology disposed therein, wherein said housing is adapted to receive a fastening means, whereby said fastening means, in cooperation with said housing, provides a means for mounting said continuously wearable device on a limb. 9. The continuously wearable compact blood glucose measuring device of claim 8, wherein said housing further comprises a storage means whereby said storage means provides a holding area for blood glucose accessories. 10. The continuously wearable compact blood glucose measuring device of claim 9, wherein said storage means is substantially airtight. 11. The continuously wearable compact blood glucose measuring device of claim 8, wherein said continuously wearable compact blood glucose measuring device further comprises a clock. 12. A continuously wearable compact blood glucose measuring device comprising a housing having a compact blood glucose measuring device using invasive, electrochemical, single use glucose test strip based technology disposed therein, wherein said housing is adapted to receive a fastening means, whereby said fastening means, in cooperation with said housing, provides a means for mounting said continuously wearable device on a limb and wherein said housing, in cooperation with said compact blood glucose measuring device, has a means for automatically engaging a glucose test strip such that said glucose test strip is ready to receive a blood sample. 13. The continuously wearable compact blood gluco... | - | A61B005/1468
US20100030101 | A1 | 2010-02-04 | Method And System For Acquiring Loudness Level Information | 2009-06-05 | The present invention relates to methods for acquiring hearing threshold values in a subject comprising providing a sound to the subject; providing a monitor for monitoring a movement response by the subject; monitoring the movement response by the subject to the sound; and progressively reducing the volume of the sound until no movement response by the subject to the sound occurs. The present invention further relates to methods and systems for acquiring loudness level information for a subject comprising providing a sound to a subject having a sensation level; providing a first sound to the subject having a sensational level; providing a monitor for monitoring a movement response by the subject; monitoring the movement response by the subject to the sound; and determining an elapsed reaction time between the sound and the movement response by the subject to the sound. | 1. A method for acquiring hearing threshold values in a subject comprising: providing a sound to the subject;providing a monitor for monitoring a movement response by the subject;monitoring the movement response by the subject to the sound; andprogressively reducing the volume of the sound until no movement response by the subject to the sound occurs. 2. The method of claim 1, wherein monitoring the movement response by the subject to the sound comprises monitoring head turn response of the subject. 3. The method of claim 1, wherein monitoring the movement response by the subject to the sound comprises monitoring eye response of the subject. 4. The method of claim 1, wherein providing the monitor comprises providing a mechanical rate monitor to the subject. 5. The method of claim 1, wherein providing the monitor comprises providing a video image monitor to the subject. 6. The method of claim 5, wherein the video image monitor comprises: an imaging device for providing a video image of the subject; anda processor running software adapted to detect the movement response of the subject by evaluating the video image of the subject. 7. The method of claim 1, wherein providing the monitor comprises providing an electrophysiological recording instrument. 8. The method of claim 7, wherein the electrophysiological recording instrument is an electromyograph. 9. The method of claim 7, wherein the electrophysiological recording instrument is an electrooculograph. 10. A method for acquiring loudness level information for a subject comprising: providing a first sound to the subject having a sensation level;providing a monitor for monitoring a movement response by the subject;monitoring the movement response by the subject to the first sound; anddetermining an elapsed reaction time between the first sound and the movement response by the subject to the first sound. 11. The method of claim 10, further comprising correlating the elapsed reaction time to the sensation level of the first sound. 12. The method of claim 11, further comprising providing a second sound to the subject having a sensation level, monitoring the movement response by the subject to the second sound, determining an elapsed reaction time between the second sound and the movement response by the subject to the second sound, correlating the elapsed reaction time to the sensation level of the second sound, and determining a rate of change between the elapsed reaction time to the first sound and the elapsed reaction time to the second sound. 13. The method of claim 10, further comprising determining a loudness parameter based on the elapsed reaction time. 14. The method of claim 10, wherein monitoring the movement response by the subject to the first sound comprises monitoring head turn response of the subject. 15. The method of claim 12, wherein monitoring the movement response by the subject to the second sound comprises monitoring head turn response of the subject. 16. The method... | - | A61B005/00
US20100022908 | A1 | 2010-01-28 | System and Method for Interfacing Cellular Matter with a Machine | 2009-09-14 | A system for interfacing cellular matter with a machine comprising a transponder implantable in cellular matter and configured to detect a first signal. The system also includes a receiver external to the cellular matter and configured to receive a second signal in response to the transponder detection of the first signal. A transmitter is also included in the system to transmit a third signal to a machine in response to the receiver receipt of the second signal. | 1-57. (canceled) 58. A implantable transponder system comprising: an electronic device;a first transponder, communicably connected to said electronic device and providing transponder signals to said electronic device;a second transponder, communicably connected to said electronic device and providing feedback signals to said electronic device subsequent to said transponder signals. 59. The system of claim 58, wherein said electronic device is a computer. 60. The system of claim 58, wherein said first transponder includes an inductance to communicate by magnetic coupling. 61. The system of claim 58, wherein said transponder signals are provided in response to biological signals. 62. The system of claim 58, wherein said feedback signals are provided in response to cellular matter signals. 63. The system of claim 58, wherein said first transponder includes an antenna to communicate by electric field coupling. 64. The system of claim 58, wherein said first transponder is coated with a bioprotectant. 65. The system of claim 58, wherein said first transponder has a volume less than one cubic millimeter. 66. The system of claim 58, further comprising a receiver, communicably connected to said first transponder and to said electronic device, wherein said receiver provides a receiver signal to said electronic device in response to receiving a transponder signal from said first transponder. 67. The system of claim 66, further comprising a transmitter, communicably connected to said receiver and to said electronic device, wherein said transmitter provides a transmitter signal to said electronic device in response to receiving a receiver signal from said receiver. 68. The system of claim 58, wherein said electronic device is a wheelchair. 69. The system of claim 58, wherein said electronic device is a telephone. 70. The system of claim 58, wherein said electronic device is an automobile. 71. The system of claim 58, wherein said electronic device is a speech synthesizer. 72. The system of claim 58, wherein said electronic device is a robotic actuator. 73. The system of claim 58, wherein said electronic device is a motor. 74. The system of claim 58, wherein said electronic device is a navigational system. 75. The system of claim 58, wherein said electronic device is a graphical display unit. 76. The system of claim 58, wherein said electronic device is a music synthesizer. 77. The system of claim 58, wherein said electronic device is a mechanized apparatus. 78. The system of claim 58, wherein said first transponder is a plurality of transponders. 79. The system of claim 58, wherein said second transponder is a plurality of transponders. 80. The system of claim 61, wherein said biological signal is a transmembrane current. | BOARD OF REGENTS, THE UNIVERSITY OF TEXAS SYSTEM | A61B005/0476 A61B005/0484
US20100036218 | A1 | 2010-02-11 | FINGERTIP BLOOD OXYGEN SATURATION MEASURING APPARATUS | 2007-02-08 | A blood oxygen saturation measuring apparatus comprising an upper shell, a lower shell, and an elastic restoring structure, characterized in that a mark is arranged on a lateral side surface of the upper and/or the lower shell; and the mark is provided to be in alignment with the vertical position of a photoelectric sensing chip which is arranged inside the shell. According to the present invention, the humanized mark is adopted, which can observe the measured position of the measured finger. Even if in the case of weak light environment or a lazy eye user, the accurate and reliable blood oxygen saturation data can be acquired. | 1. A blood oxygen saturation measuring apparatus comprising an upper shell, a lower shell, and an elastic restoring structure, wherein a mark is arranged on a lateral side surface of the upper and/or the lower shell; and the mark is provided to be in alignment with a vertical position of a photoelectric sensing chip which is arranged inside the shell. 2. The blood oxygen saturation measuring apparatus according to claim 1, wherein the mark is arranged on the lateral side surface of the shell in a convex or concave manner. 3. The blood oxygen saturation measuring apparatus according to claim 1, wherein the mark is a fluorescence mark. 4. The blood oxygen saturation measuring apparatus according to claim 2, wherein the mark in the convex or concave manner is a fluorescence mark. 5. The blood oxygen saturation measuring apparatus according to claim 1, wherein the blood oxygen saturation measuring apparatus further comprises a detection sensing chip, an internal power source, an internal photoelectric sensing chip, a detecting, processing and controlling circuit, and a display screen. | BEIJING CHOICE ELECTRONIC TECHNOLOGY CO., LTD. | A61B005/00
US20100030058 | A1 | 2010-02-04 | Methods of Treating Wrinkles, Developing Wrinkle Treatments And Evaluating Treatment Efficacy, Based On Newly Discovered Similarities Between Wrinkles And Skin Wounds | 2009-07-09 | Disclosed are methods of measuring an endogenous wrinkle electric field and methods of characterizing topical applications, or other treatments, in terms of their effect on the endogenous electric field of wrinkles. Also included, are methods of developing topical applications, or other treatments that have a desired effect on the endogenous electric field of a wrinkle. The methods disclosed herein, are based on a new model of wrinkle physiology that exploits previously unknown similarities between skin wounds and wrinkles. | 1. A method of characterizing the electric field that arises specifically from a skin wrinkle, comprising: measuring the endogenous wrinkle electric field;noting the intensity and/or polarity of the electric potential that arises from the wrinkle; andassociating a more negative reading with a more severe wrinkle or a less negative reading with a less severe wrinkle. 2. The method of claim 1 wherein a vibrating probe device is used to measure the endogenous wrinkle electric field. 3. The method of claim 2 wherein the vibrating probe device is a non-contact device. 4. The method of claim 3 wherein the wrinkle is scanned with the vibrating probe device at the wrinkle center and at locations within about 5 mm on either side thereof. 5. The method of claim 3 wherein a vibrating probe is a flat gold disc, not larger than 500 μm in diameter. 6. The method of claim 3 wherein a vibrating probe vibrates perpendicular to the skin surface and has a total displacement of not more than about 50 μm. 7. The method of claim 1 wherein a surface probe device is used to measure the endogenous wrinkle electric field. 8. A method of treating wrinkles, based on electric field properties of wrinkles, comprising applying to a wrinkle, a treatment that tends to reverse the electric field polarity of the wrinkle. 9. A method of preventing wrinkles, based on electric field properties of wrinkles, comprising: identifying a section of skin in need of protective treatment; andapplying to the section, a treatment that causes the section of skin to retain its electric field polarity. 10. A method of characterizing the progression of a skin wrinkle, based on electric field properties of wrinkles, comprising: making a first measurement of the endogenous wrinkle electric field;thereafter, waiting an amount of time with or without treating the wrinkle;making a second measurement of the endogenous wrinkle electric field; andcomparing the two measurements for changes in intensity and/or polarity. 11. A method of evaluating the efficacy of a skin wrinkle treatment, based on the electric field properties wrinkles, comprising: identifying a wrinkle for treatment;making a first measurement of the endogenous wrinkle electric field;thereafter, applying a treatment to the wrinkle;thereafter, making a second measurement of the endogenous wrinkle electric field; andcomparing the two measurements for changes in intensity and/or polarity. 12. A method of comparing the efficacy of two skin wrinkle treatments, based on the electric field properties wrinkles, comprising: selecting a first wrinkle and a second wrinkle;making a first measurement of the endogenous wrinkle electric field of the first wrinkle;thereafter, applying a first treatment to the first wrinkle;thereafter, making a second measurement of the endogenous wrinkle electric field of the first wrinkle;subtracting the first measurement from the second for the first wrinkle;making a first measurement of the endogenous wrinkle e... | - | A61B005/05 A61M035/00
US20100030103 | A1 | 2010-02-04 | Measuring device for physiological parameters | 2009-07-17 | In a measuring device for physiological parameters with an implantable fluid chamber accommodating brain fluid, with a sensor device communicating with the brain fluid in the fluid chamber via a membrane and comprising a sensor for one or more physiological parameters, electronic components and a telemetry device for wireless transmission of signals of the sensor, in order to improve the manufacture and handling, it is proposed that the fluid chamber comprise at least two housing parts which are adapted to be fitted together in a sealed manner, thereby forming a closed interior, and which allow access to the interior when not fitted together, and that the sensor device be arranged in a measurement chamber closed on all sides, which is configured as a component that is adapted for independent handling and is insertable in a defined position in the interior of the fluid chamber. | 1. Measuring device for physiological parameters, comprising an implantable fluid chamber accommodating brain fluid, a sensor device communicating with the brain fluid in the fluid chamber via a membrane and having a sensor for one or more physiological parameters, electronic components and a telemetry device for wireless transmission of signals of the sensor, wherein the fluid chamber comprises at least two housing parts which are adapted to be fitted together in a sealed manner, thereby forming a closed interior, and which allow access to the interior when not fitted together, and wherein the sensor device is arranged in a measurement chamber closed on all sides, which is configured as a component that is adapted for independent handling and is insertable in a defined position in the interior of the fluid chamber. 2. Measuring device in accordance with claim 1, wherein the interior is connected via at least one inlet to a probe via which brain fluid can flow from a measuring point of the brain into the interior. 3. Measuring device in accordance with claim 1, wherein the interior has at least one outlet for brain fluid. 4. Measuring device in accordance with claim 1, wherein the fluid chamber comprises a first housing part with a bottom and adjoining side walls and a second housing part which is configured as a cover that is adapted to be placed in a sealed manner on the first housing part. 5. Measuring device in accordance with claim 1, wherein one region of the wall of the fluid chamber is of flexible construction whereas the remaining regions of the wall are of rigid construction. 6. Measuring device in accordance with claim 5, wherein the flexible wall region is formed by a flexible membrane. 7. Measuring device in accordance with claim 6, wherein the flexible membrane consists of silicone. 8. Measuring device in accordance with claim 5, wherein the flexible membrane is inserted into a sealing ring which together with the membrane forms a housing part which is adapted to be fitted together with a second housing part in order to form the fluid chamber. 9. Measuring device in accordance with claim 1, wherein the measurement chamber is inserted into the fluid chamber in such a way that it is surrounded on all sides by brain fluid. 10. Measuring device in accordance with claim 9, wherein fixing projections which are part of the fluid chamber and/or the measurement chamber are arranged between the fluid chamber and the measurement chamber for securing the measurement chamber in the fluid chamber. 11. Measuring device in accordance with claim 1, wherein the measurement chamber consists of metal. 12. Measuring device in accordance with claim 1, wherein the measurement chamber has rigid walls. 13. Measuring device in accordance with claim 1, wherein the fluid chamber consists of a sterilizable plastic material. 14. Measuring device in accordance with claim 13, wherein the fluid chamber consists of polyetheretherketone. | Aesculap AG | A61B005/00
US20100030046 | A1 | 2010-02-04 | Subcutaneous Glucose Electrode | 2009-08-31 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably three or four-layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer is overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited (third layer). An outer (fourth) layer is biocompatible. | 1. An electrochemical sensor comprising: one or more non-corroding metal or carbon electrodes;a sensing layer comprising an enzyme coupled to each electrode; anda biocompatible layer comprising a biocompatible hydrogel chemically bound to the sensing layer of each electrode. | Abbott Diabetes Care Inc. | A61B005/1473
US07663367 | B2 | 2010-02-16 | Shaped MRI coil array | 2007-08-15 | An MRI rf coil array is comprised of a large number of separate coil elements that are supported on a substrate that is shaped to the contour of the anatomy being imaged. The coil elements overlap each other to reduce mutual inductance and their location is determined by tiling the surface of the substrate with regular, substantially same sized polygons. The center of each coil element is aligned with the center of a polygon. By using a mixture of different polygons, such as hexagons and pentagons, an arrangement of coil elements may be formed that cover a surface with non-zero Gaussian curvature where each coil is overlapped with its neighbors such that their mutual inductance is nulled. | 1. A magnetic resonance imaging coil array configured to acquire data from a subject having an exterior contour extending along multiple curves, the coil array comprising: a plurality of conductive coil elements, wherein the coil elements are arranged so that their centers are aligned with the centers of corresponding polygons that define contiguous regions on a supporting structure, configured to substantially contour the surface of the subject having the exterior contour extending along multiple curves, andwherein the structure has non-zero Gaussian curvature and the shapes of the polygons are selected such that the coil elements closely follow the contour of the supporting structure. 2. The coil array of claim 1 in which each coil element is formed by a loop of wire fastened to a substrate. 3. The coil array of claim 1, wherein each coil element overlaps at least one other coil element. 4. The coil array of claim 3, wherein each coil element overlaps with all adjacent coil elements. 5. The coil array of claim 1, wherein the coil elements are substantially the same size. 6. The coil array of claim 5, wherein the ratio of the circumferences of any two coil elements is equal to or greater than 3/5, and less than or equal to 5/3. 7. The coil array of claim 1, wherein the coil elements are formed on a substrate and the shape of the substrate is a helmet. 8. The coil array of claim 1, wherein the polygonal regions include a plurality of hexagonal polygons and at least one non-hexagonal polygon. 9. The coil array of claim 1, wherein the location of each coil element is based on a tiling pattern of regular polygons. 10. The coil array of claim 1, wherein each coil element is sized to fit within its corresponding polygonal region. 11. The coil array of claim 1, wherein at least one polygon is a pentagon. 12. The coil array of claim 1, wherein the polygonal regions comprise a combination of hexagonal and pentagonal regions. 13. The coil array of claim 1, wherein the supporting structure conforms closely to a surface of at least a portion of a human head. 14. The coil array of claim 1, wherein the polygonal regions comprise a combination of hexagonal, pentagonal and heptagonal regions. 15. The coil array of claim 1, wherein the polygonal regions comprise a geodesic tiling. 16. The coil array of claim 1 in which the conductive coil elements are connected together by a substrate. 17. The coil array of claim 1 in which the conductive coil elements are connected together by a flexible substrate. 18. The coil array of claim 4 in which the conductive coil elements are connected together by fasteners at points where adjacent conductive coil elements overlap. 19. A coil array configured for use in a magnetic resonance imaging (MRI) system in order to acquire signals from a region of interest in a subject, the combination comprising: a substrate having a surface shaped in order to curve around the region of interest extending in a plurality of c... | The General Hospital Corporation | G01V003/00 A61B005/055
US07647235 | B1 | 2010-01-12 | System and method for reducing the placebo effect in controlled clinical trials | 2004-03-31 | A method and system for performing a clinical trial having a reduced placebo effect is disclosed. The method includes randomizing study participants into three or more treatment groups and performing a first phase of testing on the groups. In a typical embodiment, the first phase of testing includes administering an active treatment to a first group, and administering a placebo to a second group and to a third group. Responders and non-responders are determined for each group. A second phase of testing is then performed. The second phase of testing includes administering the placebo to non-responders in the first group, administering the active treatment to non-responders in the second group, and administering the placebo to non-responders in the third group. The data from the first phase of testing and from the second phase of testing is pooled and analyzed to determine response rates to active treatment and placebo. | 1. A method of performing a clinical trial comprising: performing a first phase of testing, said first phase of testing including administering an active treatment to a first group of a plurality of treatment groups of study participants, and administering a placebo to a remainder of said plurality of treatment groups of study participants, wherein said study participants have been randomized with a processor into said plurality of treatment groups;determining whether each participant in each of said treatment groups is a responder or a non-responder; andperforming a second phase of testing, said second phase of testing including administering said placebo to at least one non-responder in at least one group, and administering said active treatment to at least one non-responder in at least one group,wherein data from said first phase of testing and from said second phase of testing are analyzed with a processor,wherein said analyzing comprises determining an effect of active treatment, andwherein said determining an effect of active treatment is performed in accordance with the formula:h=w⁡(n3,1n⁡(1-2⁢a)-(n1,3+n2,3)2⁢na)+(1-w)⁢(n2,1n2,1+n2,2-n1,1n1,1+n1,2),where h is a value representative of effectiveness of the treatment, w is a weighting factor, n is the total number of study participants, n1,1 is the number of participants who were non-responders to placebo in the first phase and were responders to placebo in the second phase, n1,2 is the number of participants who were non-responders to placebo in the first phase and were non-responders to placebo in the second phase, n1,3 is the number of participants who were responders to placebo in the first phase and were responders to placebo in the second phase, n2,1 is the number of participants who were non-responders to placebo in the first phase and were responders to treatment in the second phase, n2,2 is the number of participants who were non-responders to placebo in the first phase and were non-responders to treatment in the second phase, n2,3 is the number of participants who were responders to placebo in the first phase and were responders to treatment in the second phase, n3,1 is the number of participants who were responders to treatment in the first phase, and a is a randomization fraction. 2. A method of performing a clinical trial comprising: performing a first phase of testing, said first phase of testing including administering an active treatment to a first group of a plurality of treatment groups of study participants, and administering a placebo to a remainder of said plurality of treatment groups of study participants, wherein said study participants have been randomized with a processor into said plurality of treatment groups;determining whether each participant in each of said treatment groups is a responder or a non-responder; andperforming a second phase of testing, said second phase of testing including administering said placebo to at least one non-responder in at least one ... | The General Hospital Corporation | G06Q010/00 G06Q050/00 A61B005/00 G06F019/00
US07648470 | B2 | 2010-01-19 | Elasticity measuring device for biological tissue | 2002-08-20 | We disclose an elasticity-measuring device which can be inserted into a canal part of living body and which is capable of quantitatively measuring the elasticity of the biological tissue of inner side of canal part. The device consists of a probe base (5) and probes (7). The probes (7) are secured to probe base (5) and driven to press onto and return from biological tissue. According to the stress or hardness of the biological tissue measured by sensors on probes (7) and to the deviation between the probes (7) and the probe base (5), we can decide the elasticity of the biological tissue of inner side of canal part quantitatively. | 1. An elasticity measuring device for being inserted into a canal part of a human body and for measuring elasticity of the inner side of the canal part of the human body, said device comprising: a probe base for being inserted into the canal part of the human body;a probe arranged around said probe base, which is located near the inner side of the canal part of the human body when the device is inserted into the canal part and is driven to press onto and return from the biological tissue;a resilient arm member having one end and an other end, said one end supporting said probe thereon and said other end being firmly fixed to said probe base, said resilient arm member comprising a plurality of spring members, said probe comprising a plurality of probes being symmetrically arranged around said probe base through corresponding spring members;a stress detection sensor provided on at least one of the plurality of probes, for detecting hysteresis of the stress applied to the biological tissue based on the repulsion from the biological tissue when said at least one of the plurality of probes is driven to press onto and return from the biological tissue; anda deviation detection sensor for detecting the hysteresis of changes in distance of said stress detection sensor with respect to said probe base;wherein the elasticity of the biological tissue is measured based on the stress and deviation magnitude characteristics when the at least one of the plurality of probes is driven to press onto and return from the biological tissue. 2. An elasticity measuring device for biological tissue according to claim 1, in which said deviation detection sensor comprises a pair of light emitting element and light receiving element, said light emitting element being secured on a surface of said probe base and said light receiving element being secured on at least one of the plurality of the spring members so as to oppose to each other. 3. An elasticity measuring device for biological tissue according to claim 1, in which said stress detection sensor comprises a distortion gauge. 4. The device of claim 1 wherein the probe base comprises an elongated bar. 5. The device of claim 1 wherein the probe base comprises an outer surface, and wherein the plurality of the probes is configured to move substantially perpendicularly to the outer surface of the probe base. 6. The device of claim 1 wherein each of the plurality of the springs members comprises a plate spring. 7. The device of claim 1 further comprising a sleeve, wherein the sleeve and the probe base are configured to move relative each other in a longitudinally axial direction, and wherein the relative movement of the sleeve and the probe base moves the plurality of the probes in substantially a perpendicular direction relative the longitudinally axial direction. 8. The device of claim 1 wherein the probe base comprises an elongated structure that extends along a longitudinal axis, and wherein the plurality of the p... | Nihon University | A61B005/103 A61B005/117
US08178342 | B2 | 2012-05-15 | In vivo flow cytometry system and method | 2009-01-07 | The present invention provides methods and systems for performing in vivo flow cytometry. In one embodiments, selected circulating cells of interest of a subject are labeled with fluorescent probe molecules. The labeled cells are irradiated in vivo so as to excite the fluorescent probes, and the radiation emitted by the excited probes is detected, preferably confocally. The detected radiation is then analyzed to derive desired information, such as relative cell count, of the cells of interest. | 1. A system for performing in vivo cytometry in a live subject, comprising: a source for generating radiation having one or more wavelength components configured to excite selected fluorescent probe molecules capable of binding to circulating cells of a selected type of said subject upon introduction into the subject's circulatory system;an optical system for imaging radiation received from said source onto a focal plane in which a portion of a vessel of said subject is positioned so as to be illuminated by said imaged radiation, said illuminating radiation exciting one or more of said fluorescent probe molecules which are bound to said circulating cells flowing through said vessel portion;a detector configured to confocally detect fluorescence radiation emitted by said excited fluorescent probe molecules which are bound to said circulating cells flowing through said vessel portion, andan analysis module coupled to said detector configured to analyze said fluorescence radiation so as to obtain information for quantification of said circulating cells. 2. The system of claim 1, further comprising a first spatial filter optically coupled to said source such that radiation from said source is directed to said optical system through said spatial filter. 3. The system of claim 2, wherein said spatial filter comprises a pinhole. 4. The system of claim 2, further comprising a second spatial filter optically coupled to said detector and said optical system, said second spatial filter positioned so as to be conjugate with respect to said first spatial filter, wherein said optical system focuses said fluorescence radiation onto said second pinhole. 5. The system of claim 4, wherein said second spatial filter comprises a pinhole. 6. The system of claim 4, wherein each of said first and second pinholes comprises an adjustable-sized slit having a height and a width, wherein the height is larger than the width, said slit being capable of adjustment so as to vary confocality. 7. The system of claim 6, further comprising a cylindrical lens disposed between said radiation source and said slit of the first pinhole to modify cross-sectional shape of a radiation beam received from said source to one having an elongated shape characterized by substantially perpendicular long and short axes. 8. The system of claim 6, wherein said slit of the first pinhole is oriented such that its height is substantially parallel with said long axis of the beam's cross-section. 9. The system of claim 6, wherein said slit of the first pinhole is configured such that said beam passing through said slit of the first pinhole illuminates said vessel portion with the beam's long axis substantially traversing a width of said vessel portion. 10. The system of claim 1, wherein said analysis module comprises a filter for substantially removing high frequency noise. 11. The system of claim 10, wherein said analysis module comprises an amplifier for amplifying signals generated by the de... | The General Hospital Corporation | C12M001/34
US08340746 | B2 | 2012-12-25 | Motif discovery in physiological datasets: a methodology for inferring predictive elements | 2009-07-16 | The application relates a methodology and apparatus for identifying predictive patterns for acute clinical events in the absence of prior knowledge. Principles of conservation are used to identify activity that consistently precedes an outcome in patients, and describe a two-stage process that allows us to more efficiently search for such patterns in large datasets. This is achieved by first transforming continuous physiological signals from multiple patients into symbolic sequences, and by then searching for patterns in these reduced representations that are strongly associated with an outcome. | 1. A method of detecting patterns in a collection of raw physiological time-series information comprising the steps of: recording raw physiological time-series information from a patient using a device; segmenting the raw physiological information from the device into a plurality of components using a processor; grouping the components into a plurality of time-series information classes using a processor; assigning a symbolic representation to each information class using a processor; searching, using a processor, for patterns of these symbolic representations in a database; and determining the physiological condition associated with the symbolic representation using a processor. 2. The method of claim 1 wherein the pattern may be a sequence of representations. 3. The method of claim 2 wherein the sequence of representations occurs more often than expected given the distribution of symbolic representations. 4. The method of claim 2 wherein the pattern is a sequence of representations that occur more often than expected in patients who have an event. 5. The method of claim 2 wherein the pattern is a sequence of representations that occur more often than expected in patients who have an event relative to patients who do not have an event. 6. The method of claim 1 wherein the pattern may be the entropy of the representations. 7. The method of claim 1 further comprising the step of compressing multiple equivalent time portions that contain normal raw physiological time-series information. 8. The method of claim 1 further comprising the step of removing multiple equivalent time portions that contain normal raw physiological time-series information. 9. The method of claim 1 further comprising the step of assigning a significant event to a predictive conserved sequence of symbolic representations. 10. The method of claim 1 wherein the raw physiological time-series information is an ECG and the equivalent time portion is a heartbeat. 11. An apparatus for detecting patterns in a physiological signal comprising: a module for dividing the physiological signal into a plurality of equivalent time portions; and a module for assigning a symbol to each portion of the plurality of equivalent time portions. 12. The apparatus of claim 11 further comprising a module for recording the physiological signal from a patient. 13. The apparatus of claim 11 further comprising a module for removing multiple equivalent time portions that comprise normal physiological signals. 14. The apparatus of claim 11 further comprising a module for compressing multiple equivalent time portions that comprise normal physiological signals. 15. The apparatus of claim 11 further comprising a module for assigning a significant event to a predictive substantially conserved sequence of symbols. 16. The apparatus of claim 11 wherein the module for recording a physiological signal is an ECG and the equivalent time portion is a heartbeat. 17. A method of detecting patterns in raw ti... | Massachusetts Institute of Technology | A61B005/04
US08155732 | B2 | 2012-04-10 | ECG system for use in ECG signal measurement of intra-cardiac ECG using a catheter | 2007-11-09 | An ECG system is used in ECG signal measurement of intra-cardiac ECG using a catheter. The system includes a lead signal switch for disconnecting an amplifier input terminal from a reference point and reconnecting the input terminal to a patient limb electrode. An ECG signal processor amplifies the difference of a chest lead signal electronically connected to a catheter, and a patient limb signal provided by the reconnecting of the input terminal. | 1. An ECG system for use in ECG signal measurement of intra-cardiac ECG using a catheter, comprising: an external lead configuration having at least two limb electrodes with a common node connection reference point therebetween and at least a left leg electrode lead connection point,a lead signal switch for disconnecting a first input terminal of an amplifier from the common node connection reference point and reconnecting said input terminal to said left leg electrode lead connection point; andan ECG signal processor including said amplifier having said first input terminal and a second input terminal, said second input terminal being electrically connected to a catheter ECG electrode via a chest lead signal conductor for amplifying the difference of: a chest lead signal sensed by the catheter ECG electrode, anda patient limb signal provided by said reconnecting said input terminal. 2. The system of claim 1 wherein said chest lead signal electronically connected to said catheter ECG electrode is an unused chest lead signal. 3. The system of claim 1 wherein said chest lead signal electrode comprises a V1, V2, V3, V4, V5 or V6 signal electrode of a conventional 12 lead ECG signal set. 4. The system of claim 1 wherein said lead signal switch is implemented in at least one of, (a) hardware and (b) software. 5. The system of claim 1 wherein the reference point is the Wilson Star point. 6. The system of claim 1 wherein the lead signal switch reconnects the input terminal of the amplifier to the patient limb electrode without impairing the derivation of the Einthoven signals. 7. The system of claim 1 wherein the lead signal switch reconnects the input terminal of the amplifier to the patient limb electrode while maintaining the connections of other electrodes connected to said reference point. 8. An ECG system for use in ECG signal measurement of intra-cardiac ECG using a catheter, comprising: an external lead configuration having at least two limb electrodes with a common node connection reference point therebetween and at least two limb lead electrode connection points,a lead signal switch for disconnecting a first input terminal of an amplifier from a common node connection reference point and re-connecting said input terminal to one of said at least two limb lead electrode connection points in response to a user command entered via an ECG signal display system user interface without impairing the signals of other electrodes connected to said common node connection reference point; andan ECG signal processor including said amplifier having said first input terminal and a second input terminal being electrically connected to a catheter ECG electrode via a chest lead signal conductor for amplifying the difference of, a chest lead signal sensed by the catheter ECG electrode, anda patient limb signal provided by said re-connecting said input terminal. 9. The system of claim 8 wherein the lead signal switch reconnects the input terminal of the am... | Draeger Medical Systems, Inc. | A61B005/04
US08761855 | B2 | 2014-06-24 | Systems and methods for determining oxygen saturation | 2008-10-10 | According to embodiments, techniques for using continuous wavelet transforms and spectral transforms to determine oxygen saturation from photoplethysmographic (PPG) signals are disclosed. According to embodiments, a first oxygen saturation may be determined from wavelet transformed PPG signals and a second oxygen saturation may be determined from spectral transformed PPG signals. An optimal oxygen saturation may be determined by selecting one of the first and the second oxygen saturation or by combining the first and the second oxygen saturation. According to embodiments, a spectral transform of PPG signals may be performed to identify a frequency region associated with a pulse rate of the PPG signal. A continuous wavelet transform of the PPG signals at a scale corresponding to the identified frequency region may be performed to determine oxygen saturation from the wavelet transformed signal. | 1. A method for determining oxygen saturation from photoplethysmographic (PPG) signals, comprising: using a processor for: receiving a first PPG signal and a second PPG signal;calculating a first oxygen saturation based at least in part on continuous wavelet transforms of the first PPG signal and the second PPG signal;calculating a second oxygen saturation based at least in part on spectral transforms of the first PPG signal and the second PPG signal; anddetermining an optimal oxygen saturation based at least in part on the first oxygen saturation and the second oxygen saturation. 2. The method of claim 1 wherein performing the continuous wavelet transform of the first PPG signal and the second PPG signal comprises generating a first wavelet transformed signal and a second wavelet transformed signal. 3. The method of claim 2 wherein calculating the first oxygen saturation comprises: identifying a pulse ridge in the first or the second wavelet transformed signals;identifying a first value corresponding to an amplitude of the pulse ridge in the first wavelet transformed signal;identifying a second value corresponding to an amplitude of the pulse ridge in the second wavelet transformed signal; andperforming a ratio of the first value and the second value, wherein the first oxygen saturation is calculated based at least in part on the ratio. 4. The method of claim 2 wherein calculating the first oxygen saturation comprises: identifying a scale in the first wavelet transformed signal and the second wavelet transformed signal that corresponds to pulse rate;using the scale to identify a first amplitude in the first wavelet transformed signal;using the scale to identify a second amplitude in the second wavelet transformed signal; andperforming a ratio of the first amplitude and the second amplitude, wherein the first oxygen saturation is calculated based at least in part on the ratio. 5. The method of claim 1 wherein determining the optimal oxygen saturation comprises performing a weighted average of the first oxygen saturation and the second oxygen saturation. 6. The method of claim 1 wherein determining the optimal oxygen saturation comprises selecting the first oxygen saturation or the second oxygen saturation based at least in part on a confidence indicator. 7. A system for determining oxygen saturation from photoplethysmographic (PPG) signals, comprising: a wavelet processor capable of calculating a first oxygen saturation based at least in part on continuous wavelet transforms of a first received PPG signal and a second received PPG signal;a spectral processor capable of calculating a second oxygen saturation based at least in part on spectral transforms of the first received PPG signal and the second received PPG signal; andan SpO2 processor capable of determining an optimal oxygen saturation based at least in part on the first oxygen saturation and the second oxygen saturation. 8. The system of claim 7 wherein the wavelet processor is capa... | Nellcor Puritan Bennett Ireland | A61B005/00
US08792957 | B2 | 2014-07-29 | Multiple electrode composite systems in electrocardiogram detection systems | 2009-08-05 | A biomedical sensor system is disclosed that includes a plurality of electrodes and a contiguous adhesive material that is in contact with each of the plurality of electrodes. In certain embodiments a method is provided that includes the step of applying a first surface of adhesive material to a patient wherein the adhesive material includes at least two electrodes on second surface thereof that is opposite the first surface. The method also includes the step of receiving a time varying signal a first electrode of the at least two electrodes at a first location such that the time varying signal is not received at a second electrode of the at least two electrodes. | 1. A biomedical sensor system comprising a plurality of electrodes on a first side of a contiguous adhesive material that is in contact with each of the plurality of electrodes, wherein said contiguous adhesive material includes polar material dispersed with the contiguous adhesive material such that the polar material facilitates providing an electrical signal through the contiguous adhesive material in a direct path between a biomedical signal and one of the plurality of electrodes yet does not facilitate the providing of the electrical signal through the contiguous adhesive material in a path between the biomedical signal and other of the plurality of electrodes that do not directly oppose the biomedical signal. 2. The biomedical sensor system as claimed in claim 1, wherein said adhesive material is a dielectric material. 3. The biomedical sensor system as claimed in claim 1, wherein said multiple electrodes are coupled to a signal analysis system. 4. The biomedical sensor system as claimed in claim 3, wherein said signal analysis system selects a subset of the plurality of electrodes to be used for signal analysis. 5. The biomedical sensor system as claimed in claim 4, wherein said signal analysis system analyzes electrical signals from the subset of the plurality of electrodes to determine whether a patient has an abnormal heartbeat. 6. The biomedical sensor system as claimed in claim 1, wherein said system includes six electrodes in contact with the adhesive material. 7. The biomedical sensor system as claimed in claim 6, wherein said sensor system is adapted for performing ECG analyses. 8. The biomedical sensor system as claimed in claim 1, wherein said electrodes and the adhesive material are mounted on a carrier substrate prior to application to a patient. 9. The biomedical sensor system as claimed in claim 1, wherein said adhesive material has a thickness of less than about 200 microns. 10. The biomedical sensor system as claimed in claim 1, wherein said adhesive material has a thickness of between about 25 microns and about 100 microns. 11. The biomedical sensor system as claimed in claim 1, wherein at least some of said plurality of electrodes are spaced from one another by a distance of less than about 25,000 microns. 12. The biomedical sensor system as claimed in claim 1, wherein at least some of said plurality of electrodes are spaced from one another by a distance of less than about 2,500 microns. 13. A biomedical sensor system comprising a plurality of electrodes on a first side of a contiguous adhesive material that is in contact with each of the plurality of electrodes, wherein aligned polar material within the contiguous adhesive material that is between a first electrode of the plurality of electrodes and a biomedical signal on an opposing second surface of the contiguous adhesive material facilitates the providing of an electrical signal that is representative of the biomedical signal through the contiguous adhes... | FLEXcon Company, Inc. | A61B005/0402 H01B001/00 A61B005/0478 A61B005/04
US09364173 | B2 | 2016-06-14 | Signal processing for continuous analyte sensor | 2009-09-23 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A method for processing sensor data from a continuous analyte sensor, comprising: receiving sensor data from a continuous analyte sensor for a first time period, wherein the sensor data comprises at least one sensor data point;processing, using one or more processors, the at least one sensor data point to obtain a range of target analyte values for a future time period; andproviding, using the one or more processors, output, wherein providing output comprises the one or more processors selecting whether or not to display an object representative of the range of target analyte values on a user interface based at least in part on a variability analysis of the sensor data using the one or more processor. 2. The method of claim 1, wherein, the target glucose value is based on an individual physiological pattern. 3. The method of claim 1, wherein processing the at least one sensor data point to obtain a range of target analyte values for a future time period comprises determining whether the target analyte value or range of target analyte values is based on a clinical acceptability. 4. The method of claim 1, wherein the continuous analyte sensor is a continuous glucose sensor. 5. The method of claim 1, wherein processing the at least one sensor data point to obtain a range of target analyte values for a future time period comprises determining the range of target analyte values based at least in part on at least one parameter that is customizable by the user. 6. The method of claim 1, wherein processing the at least one sensor data point to obtain a range of target analyte values for a future time period comprises determining the range of target analyte values based at least in part on an amount and/or timing of a medicament delivery and/or caloric consumption. 7. The method of claim 1, wherein processing the at least one sensor data point to obtain a range of target analyte values for a future time period comprises determining the range of target analyte values based at least in part on a host's physiological pattern. 8. The method of claim 1, wherein processing the at least one sensor data point to obtain a range of target analyte values for a future time period comprises determining the range of target analyte values based at least in part on the sensor data. 9. The method of claim 1, wherein processing the at least one sensor data point to obtain a range of target analyte values for a future time period comprises determining the range of target analyte values based at least in part on the variation analysis of the sensor data using a clinical, physiological, and/or statistical variation analysis. 10. The method of claim 1, wherein providing output comprises displaying the range of target analyte values along a path of target analyte values. 11. The method of claim 1, wherein providing output comprises simultaneously displaying a real-time analyte value and the range of target analyte values on the user interface. 12. The method of cl... | DexCom, Inc. | A61B005/145 A61B005/1486 A61B005/1495 A61B005/1468 A61B005/00
US09705239 | B2 | 2017-07-11 | Flexible electrode assembly and apparatus for measuring electrophysiological signals | 2007-08-30 | A flexible electrode assembly comprises a central portion having electrical contacts disposed thereon and a plurality of elongated portions extending radially outwards from the central portion. One or more of the elongated portions has an electrode disposed in the vicinity of a distal end thereof. The electrodes are electrically coupled to respective ones of the electrical contacts disposed on the central portion. An apparatus for measuring electrophysiological signals in a human or animal body that incorporates the flexible electrode assembly is also disclosed. | 1. An apparatus for measuring electrophysiological signals in a human or animal body, said apparatus comprising: a support structure comprising a central portion and a plurality of elongated portions extending radially outwards from said central portion, the central portion having a plurality of openings passing entirely through the support structure at respective interfaces between the central portion and the elongated portions; and a flexible electrode assembly comprising a central region having a plurality of electrical contacts, a plurality of elongated portions extending from the central region and each having a distal end spaced from the central region with an electrode disposed thereon, each electrode at the distal end coupled to one of the electrical contacts in the central region, said flexible electrode assembly mounted on said support structure such that said elongated portions of said flexible electrode assembly each extend through the openings and along a respective elongated portion of said support structure, wherein the electrode assembly is removably and repeatedly attached to the support structure; and a flexible printed circuit board disposed on at least one side of at least one elongated portion of the support structure and connecting said electrical contacts and said electrode, the flexible circuit board having a plurality of electrically-conductive through holes. 2. The apparatus of claim 1, wherein said plurality of elongated portions of said support structure are resiliently deformable. 3. The apparatus of claim 1, wherein said flexible electrode assembly is of unitary structure. 4. The apparatus of claim 1, wherein said flexible electrode assembly comprises a flexible electrode assembly in accordance with claim 1. 5. The apparatus of claim 1, further comprising an amplifier module for amplifying electrophysiological signals detected by said electrodes. 6. The apparatus of claim 5, wherein said electrodes are electrically connected to said amplifier by way of compressed contact of electrically conducting surfaces. 7. The apparatus of claim 2, further comprising attachment means for attaching said apparatus to a human or animal body. 8. The apparatus of claim 7, wherein said attachment means comprises a strap which, when under tension, causes said elongated portions of said support structure to deform such that contact between said electrodes and said human or animal body is substantially maintained. 9. The apparatus of claim 1, wherein said flexible electrode assembly is disposable. 10. An apparatus for measuring electrophysiological signals in a human or animal body, said apparatus comprising: an electrode assembly having a plurality of electrodes for detecting electrophysiological signals, the electrode assembly further including a central portion having a plurality of electrical contacts and a plurality of elongated portions extending radially outwards from the central portion, one or more of said elongated port... | OBJECTIVISION LIMITED | A61B005/04 H01R013/52 A61B005/0408
US20100016712 | A1 | 2010-01-21 | Method and Device for Visually Assisting a Catheter Application | 2008-02-26 | A method and a device for visually assisting a catheter application on the heart of a patient using an image of the patient obtained by a C-arm X-ray device and using electroanatomical mapping data of the patient obtained by an electromagnetic position detection system and mapping system. The C-arm X-ray device and the electromagnetic position detection system and mapping system are calibrated in relation to each other, by determining a co-ordinate transformation between a co-ordinate system assigned to the C-arm x-ray device and/or a co-ordinate system assigned to the image generated by the C-arm X-ray device and a co-ordinate system assigned to the electromagnetic position detection system and mapping system. The position of the patient is determined during the detection of the image and/or during the detection of the electroanatomic mapping-data and is at least indirectly assigned to the image and/or the electroanatomic mapping-data. | 1.-56. (canceled) 57. A method for visually assisting a catheter application on a heart of a patient, comprising: acquiring an image of the patient by a C-arm X-ray device;acquiring electroanatomical mapping data of the patient by an electromagnetic position-detection and mapping system;determining a coordinate transformation between a coordinate system of the C-arm x-ray device and a coordinate system of the electromagnetic position-detection and mapping system;calibrating the C-arm X-ray device and the electromagnetic position-detection and mapping system in relation to each other based on the coordinate transformation;determining a position of the patient during the acquisition of the image and the acquisition of the electroanatomical mapping data; andassigning the position of the patient to the image and the electroanatomical mapping data. 58. The method as claimed in claim 57, wherein a patient supporting device is assigned to the C-arm X-ray device,wherein the electromagnetic position-detection and mapping system comprises a catheter with a sensor and a transmitter for generating an electromagnetic alternating field,wherein the transmitter is arranged on the C-arm X-ray device or on the patient supporting device and is detachable from the C-arm X-ray device or from the patient supporting device,wherein the transmitter is arranged on the C-arm of the C-arm X-ray device, andwherein a positioning and orientation sensor of the electromagnetic position-detection and mapping system is arranged on or in the patient supporting device and is detachable from the patient supporting device. 59. The method as claimed in claim 57, wherein the coordinate transformation is determined by a marker that is detectable by the electromagnetic position-detection and mapping system. 60. The method as claimed in claim 59, wherein the marker is an X-ray positive marker and an image of the marker is depicted in at least two X-ray projections taken from different projection angles or in a 3D image. 61. The method as claimed in claim 60, wherein the image of the marker is located manually or automatically in the X-ray projections or in the 3D image by a method of pattern recognition. 62. The method as claimed in claim 59, wherein the marker is touched by a position sensor of the electromagnetic position-detection and mapping system. 63. The method as claimed in claim 59, wherein the marker is an X-ray positive position sensor of the electromagnetic position-detection and mapping system. 64. The method as claimed in claim 59, wherein a phantom comprising the marker is provided and the marker of the phantom is a position sensor of the electromagnetic position-detection and mapping system. 65. The method as claimed in claim 64, wherein the phantom comprises a further marker and a position of the further marker is known relative to the position sensor. 66. The method as claimed in claim 64, wherein a coordinate transformation between a coordinate system of the ph... | - | A61B006/12 A61B005/055 G06K009/00
US20100049023 | A1 | 2010-02-25 | Non-invasive method and system for determining physiological characteristics | 2008-08-20 | A non-invasive method of determining a physiological characteristic, comprising providing at least one physiological sensor that is adapted to measure at least one physiological characteristic at a target measurement site on a subject's body, heating an extended tissue region on the subject's body, whereby blood perfusion of the tissue region is enhanced, and measuring at least one physiological characteristic with the physiological sensor during or within a predetermined period after heating the extended tissue region. | 1. A method of determining a physiological characteristic, comprising the steps of: providing at least one physiological sensor that is adapted to measure at least one physiological characteristic at a target measurement site on a subject's body;disposing said physiological sensor proximate said target measurement site on said subject's body;heating a tissue region on said subject's body, whereby blood perfusion of said tissue region is enhanced, said tissue region including said target measurement site and extending beyond said target measurement site; andmeasuring at least one physiological physiological characteristic at said target measurement site with said physiological sensor during or within a predetermined period after said heating of said tissue region. 2. The method of claim 1, wherein said heating of said tissue region is sufficient to induce an optimal homeostatic reflex, whereby said tissue region blood perfusion is enhanced, without burning said subject. 3. The method of claim 1, wherein said tissue region comprises the entire ear of said subject and said target measurement site comprises the earlobe of said heated ear. 4. The method of claim 1, wherein said tissue region comprises the entire ear and adjoining tissue region of said subject and said target measurement site comprises the earlobe of said heated ear. 5. The method of claim 1, wherein said tissue region comprises the entire arm of said subject and said target measurement site comprises a finger on said heated arm. 6. The method of claim 1, wherein said tissue region comprises a hand of said subject and said target measurement site comprises a finger on said heated hand. 7. A method of determining a physiological characteristic, comprising the steps of: providing a plurality of physiological sensors that are adapted to measure at least one physiological characteristic at target measurement sites on a subject's body;disposing a first physiological sensor proximate a first target measurement site on said subject's body;disposing a second physiological sensor proximate a second target measurement site on said subject's body;heating a first tissue region on said subject's body, whereby blood perfusion of said first tissue region is enhanced, said first tissue region including said first target measurement site and extending beyond said first target measurement site; andmeasuring at least one physiological characteristic with said first and second physiological sensors during or within a predetermined period after said heating of said first tissue region. 8. The method of claim 7, wherein said heating of said first tissue region is sufficient to induce an optimal homeostatic reflex, whereby said first tissue region blood perfusion is enhanced, without burning said subject. 9. The method of claim 7, wherein said first tissue region comprises the entire ear of said subject and said first target measurement site comprises the earlobe of said heated ear. 10. The method o... | - | A61B005/1491
US20100010333 | A1 | 2010-01-14 | Bipolar, Non-Vectorial Electrocardiography | 2009-09-07 | An improved method for registering the changes in electrical potentials present on the surface of the body in association with the contraction of the heart by recognizing that today's art hypothesis on the genesis of such potentials are unsustainable. The new “Bipolar non-Vectorial Leads” are obtained by paring a distal “Common or Positive Electrode” placed on the left leg with an “Exploring or Negative Electrode” placed near the myocardium on areas where the electrical potentials generated by the different structures of the myocardium are prevalent. The approximate twelve leads will sample all the areas were each myocardial structure is prevalent. The leads so obtained are to be analyzed as generated on the surface of the myocardium and conducted throughout through the body to the entire surface by the muscular masses that are in close contact with the different structures of the myocardium. The final report besides the printed electrocardiographic traces includes all the digital data sets, obtained by the electrocardiograph, saved on a digital disk. | 1. A method of registering Bipolar Non-Vectorial Electrocardiographic leads, comprising of the following steps: a. placing approximately twelve “Exploring or Negative Electrodes” on a surface of a body were changes of electrical potential are prevalent for each myocardial structure,b. connecting each of said “Exploring or Negative Electrodes” to a negative terminal of their respective individual amplifier,c. placing a “Common or Positive Electrode” on a left leg,d. connecting each said “Common or Positive Electrodes” to a positive terminal of all amplifiers,e. placing one “Ground Electrode” on a right leg, andf. connecting said “Ground Electrode” to a ground terminal of all the amplifiers,whereby said registered Bipolar non-Vectorial electrocardiographic leads will facilitate, understanding normal and pathological physiological processes, and diagnosing normal and pathological processes associated with myocardial systole and diastole. 2. A method of analyzing said Bipolar Non-Vectorial Electrocardiographic Leads, comprising of the following steps: a. recognizing that a change in electrical potential on said surface of said body synchronized with the contraction of a myocardium is generated by a monophasic electrical potential present on a surface of a structure of said myocardium,b. recognizing that said change in electrical potential on the surface of the body is propagated throughout through the body by a muscular mass that is in close contact with said structure of the myocardium,c. recognizing that the structure of the myocardium propagates throughout through the body said monophasic electrical potentials present on said surface of the structure of the myocardium to a specific area of the surface of the body at a specific time, amplitude and morphology,whereby said analysis facilitates the recognition of the normal processes and facilitates recognition and localization of: abnormal rhythms, alterations of conduction of impulses along the Bundle of His, specific coronary alterations involved in angina and myocardial infarctions, myocardial aneurisms. 3. A method of analyzing a wave of a Bipolar Non-Vectorial Electrocardiographic Lead trace, comprising of the following step: a. recognizing that said wave of said Bipolar Non-Vectorial Electrocardiographic trace is generated by a difference between specific times, amplitudes and morphologies of monophasic electrical potentials present on a surface of different structures of a myocardium,whereby said analysis of the waves facilitates the recognition, localization, and myocardial origin of the abnormal Bipolar Non-Vectorial Electrocardiographic Lead. 4. A method of analyzing a segment of a Bipolar non-Vectorial Electrocardiographic trace, comprising of the following step: a. recognizing that said segment of said Bipolar non-Vectorial Electrocardiographic trace is generated by a difference between specific times, amplitudes and morphologies of the monophasic electrical potentials present on the s... | - | A61B005/0402
US20100016753 | A1 | 2010-01-21 | Systems and Methods for Portable Neurofeedback | 2009-07-15 | Systems and methods for portable neurofeedback are disclosed herein. In an embodiment, a neurofeedback system includes sensors for capturing, recording and transmitting brain electrical activity; and a portable device for receiving, processing and displaying in real-time the brain electrical activity, wherein the brain electrical activity is provided visually on the portable device as visual feedback. In an embodiment, the portable device further includes speakers for implementing auditory feeedback based on the recorded brain electrical activity. In an embodiment, the portable device further includes at least one vibration feedback motor for implementing vibratory feedback based on the recorded brain electrical activity. | 1. A neurofeedback system comprising: sensors for capturing, recording and transmitting brain electrical activity; anda portable device for receiving, processing and displaying in real-time the brain electrical activity, wherein the brain electrical activity is provided visually on the portable device as visual feedback. 2. The neurofeedback system of claim 1 wherein the sensors are electrodes. 3. The neurofeedback system of claim 2 wherein the electrodes transmit the brain electrical activity wirelessly to the portable device. 4. The neurofeedback system of claim 2 wherein the electrodes transmit the brain electrical activity to the portable device via one or a plurality of frequency channels. 5. The neurofeedback system of claim 2 wherein the electrodes support a frequency range from about 0.25 to about 100 Hz. 6. The neurofeedback system of claim 1 wherein the portable device is an iPhone®. 7. The neurofeedback system of claim 1 wherein the portable device includes a display screen for providing the visual feedback. 8. The neurofeedback system of claim 7 wherein the visual feedback indicates one of a first state or a second state of a user of the system. 9. The neurofeedback system of claim 8 wherein the first state represents a focused state of the user. 10. The neurofeedback system of claim 8 wherein the second state represents a distracted state of the user. 11. The neurofeedback system of claim 9 wherein when the user is in the focused state, the visual feedback displayed on the portable device is of a first type. 12. The neurofeedback system of claim 10 wherein when the user is in the distracted state, the visual feedback displayed on the portable device is of a second type. 13. The neurofeedback system of claim 1, wherein the portable device further comprises speakers for implementing auditory feedback based on the recorded brain electrical activity. 14. The neurofeedback system of claim 1, wherein the portable device further comprises at least one vibration feedback motor for implementing vibratory feedback based on the recorded brain electrical activity. 15. The neurofeedback system of claim 1, wherein the system is used to improve a number of conditions and disorders selected from the group consisting of attention-deficit/hyperactivity disorder, chronic pain, acute pain, memory problems, learning disorder, focusing disorder, sleep disorder, tourette's syndrome, pre-menstrual syndrome, migraines, seizures, extreme stress, anxiety attacks, autism and depression. 16. The neurofeedback system of claim 1, wherein the system is used for peak performance training. 17. A method of neurotherapy comprising: providing a neurofeedback system to a user, wherein the neurofeedback system includes sensors for capturing, recording and transmitting brain electrical activity and a portable device for receiving, processing and displaying in real-time the brain electrical activity;configuring the portable device to select a neurofeedback p... | - | A61B005/0482
US20100010368 | A1 | 2010-01-14 | SYSTEMS AND METHODS FOR TREATING AORTIC LESIONS | 2009-07-15 | Devices, systems, and methods for the localization of body lumen junctions and other intraluminal structure are disclosed. Various embodiments permit clinicians to identify and locate lesions and/or anatomical structures within a lumen and accurately place leads and/or devices within a lumen, through determining the intralumen conductance and/or cross-sectional area at a plurality of locations within the body lumen. | 1. A method for treating a lesion within an aorta, comprising the steps of: providing a system, comprising: an expandable tubular member having a first open end, a second open end, an outer surface, and a hollow interior, both the outer surface and the hollow interior extending between the first and second open ends and the expandable tubular member capable of moving between a first closed position and a second open position,a catheter removably disposed within the interior of the tubular member and having a proximal end and a distal end, the distal end of the catheter for placement into a aorta, the catheter comprising a first electrode and a second electrode, each of the first and second electrodes having a proximal end and a distal end, the distal ends of the first and second electrodes located between the proximal and distal ends of the catheter, anda processor connected to the first and second electrodes of the catheter, the processor capable of collecting conductance data to determine a profile of the aorta, the conductance data collected at a plurality of locations within the aorta and determined at each of the plurality of locations when the distal ends of the first and second electrodes are immersed in a fluid within the aorta;delivering the catheter and the expandable tubular member to the aorta;determining a profile of the aorta through operation of the catheter and the processor;using the profile of the aorta to identify the location of a lesion within the aorta;disposing at least a portion of the tubular member within the lesion; andcoupling the expandable tubular member with the aorta by moving the expandable tubular member to the second open position;whereby the expandable tubular member provides a fluid passageway through the lesion. 2. The method of claim 1, wherein the processor further comprises a computational program capable of detecting a bifurcation in the aorta and distinguishing the bifurcation from a lesion. 3. The method of claim 1, wherein the profile of the aorta is determined in real time. 4. The method of claim 1, wherein the lesion comprises an aneurysm. 5. The method of claim 1, further comprising the step of using the profile of the aorta to identify the dimensions of the lesion. 6. The method of claim 5, wherein the step of coupling the expandable tubular member with the aorta further comprises: coupling the outer surface adjacent to the first open end of the expandable tubular member to the aorta in a first location outside of the dimensions of the lesion; andcoupling the outer surface adjacent to the second open end of the expandable tubular member to the aorta in a second location outside of the dimensions of the lesion. 7. The method of claim 5, wherein the second location is in an antegrade position relative to the first location. 8. The method of claim 1, wherein the step of operating the system to determine a profile of the aorta further comprises the steps of: providing constant electrical current... | DTherapeutics, LLC (a California Limited Liability Company) | A61B005/053 A61M029/00
US20100030100 | A1 | 2010-02-04 | Microneedle Device For Diagnosis Of Allergy | 2008-02-05 | There is provided an inexpensive microneedle device for diagnosis of an allergy, enabling one to perform a clear skin test with simple operation for a short period of time in the diagnosis of allergies. The microneedle device (1) comprises: a microneedle substrate (2), microneedles (3) formed on the substrate at a density of 100 to 10000 needles per 1 cm2, being capable of piercing the skin 50 μm to 500 μm deep, and allergen-holding means (4) formed on the microneedles, holding at least one allergen. The microneedles (3) can be prepared by the use of a non-metallic synthetic or natural resin material. | 1. A microneedle device for diagnosis of an allergy, comprising: a substrate,microneedles formed on the substrate at a density of 100 to 10000 needles per 1 cm2, being capable of piercing the skin 50 μm to 500 μm deep, andallergen-holding means placed on the microneedles, holding at least one allergen. 2. The microneedle device for diagnosis of an allergy according to claim 1, wherein the microneedles are prepared by the use of a non-metallic synthetic or natural resin material. 3. The microneedle device for diagnosis of an allergy according to claim 2, wherein the non-metallic synthetic or natural resin material is a biodegradable polymer such as polylactic acid, polyglycolide, polylactic acid-co-polyglycolide, capronolactone, polyurethane, or polyanhydride; or hyaluronic acid, pullulan, dextran, dextrin or chondroitin sulfate that are all polysaccharides; or polycarbonate, polymethacrylic acid, ethylenevinylacetate, polytetrafluoroethylene or polyoxymethylene that are all non-degradable polymers. 4. The microneedle device for diagnosis of an allergy according to claim 1, wherein the allergen is a peptide(s) or a protein(s). 5. The microneedle device for diagnosis of an allergy according to claim 1, wherein the allergen-holding means is a coating on the microneedle surface with the allergen by using a coating carrier. 6. The microneedle device for diagnosis of an allergy according to claim 5, wherein the coating carrier includes a water-soluble polymer having a hydroxy group. 7. The microneedle device for diagnosis of an allergy according to claim 6, wherein the coating carrier includes a water-soluble polymer having a hydroxy group, the polymer being selected from the group consisting of polyethylene oxide, polyhydroxymethylcellulose, polyhydroxypropylcellulose, polyhydroxypropylmethylcellulose, polymethylcellulose, dextran, polyethylene glycol, and polyvinyl alcohol. 8. The microneedle device for diagnosis of an allergy according to claim 6, wherein the coating carrier includes a polysaccharide. 9. The microneedle device for diagnosis of an allergy according to claim 8, wherein the polysaccharide includes a water-soluble polymer selected from the group consisting of pullulan, carmellose sodium, chondroitin sulfate, hyaluronic acid, dextran, and Gum arabic. 10. The microneedle device for diagnosis of an allergy according to claim 5, wherein the coating is in a state of being anchored to the microneedle surface. 11. The microneedle device for diagnosis of an allergy according to claim 2, wherein the allergen is a peptide(s) or a protein(s). 12. The microneedle device for diagnosis of an allergy according to claim 3, wherein the allergen is a peptide(s) or a protein(s). 13. The microneedle device for diagnosis of an allergy according to claim 2, wherein the allergen-holding means is a coating on the microneedle surface with the allergen by using a coating carrier. 14. The microneedle device for diagnosis of an allergy according to cl... | HISAMITSU PHARMACEUTICAL CO., INC. | A61B005/00
US07660622 | B2 | 2010-02-09 | Method and system for positioning a tracking sensor for optimal accuracy | 2004-10-07 | The present invention provides a method for positioning an electromagnetic tracking sensor. The method includes attaching a sensor to an anatomy, obtaining an image of the anatomy, indicating a size of each of a plurality of volumes on the image, where the volumes include a smaller and larger volume, and making a determination of whether a region of interest is encompassed by one or more of the volumes. The present invention also provides a system for positioning an electromagnetic tracking sensor. The system includes a sensor attached to a patient anatomy, an imaging modality obtaining an image of the anatomy, an image processor indicating a size of each of a plurality of volumes on the image, where the plurality of volumes includes a smaller and larger volume, and an operator making a determination of whether a region of interest is encompassed by one or more of the volumes. | 1. A method for positioning an electromagnetic tracking sensor, said method including: determining a region of interest in a patient anatomy, wherein said region of interest represents an area of reduced electromagnetic interference;attaching said sensor to said patient anatomy;obtaining an image of said anatomy;indicating a size of each of a plurality of volumes on said image, said plurality of volumes including a predetermined smaller and larger volume, wherein each volume in said plurality of volumes represents at least one distance from said sensor, wherein said larger volume contains said smaller volume;making a determination of whether said region of interest is encompassed by one or more of said volumes; andmoving said sensor based on at least said making a determination step such that the region of interest is encompassed by at least one of said volumes. 2. The method of claim 1, wherein said volumes are concentric with said sensor located at a center of said volumes. 3. The method of claim 1, wherein said smaller volume represents a distance from said sensor in which said region is preferred to be located. 4. The method of claim 1, wherein said moving step includes moving said sensor if said region is not located within said smaller volume. 5. The method of claim 1, wherein said moving step includes moving said sensor if said region is not located within said larger volume. 6. The method of claim 1, wherein said volumes include at least one sphere. 7. The method of claim 1, wherein said indicating step includes drawing a two-dimensional representation of at least one of said volumes in said image. 8. The method of claim 1, further including: obtaining a second image of said anatomy;indicating a size of each of a second plurality of volumes on said second image, said second plurality of volumes including a second smaller volume and second larger volume; andmaking a second determination of whether said region is encompassed by one or more of said second volumes. 9. The method of claim 8, further including: moving said sensor based on at least one or more of said making a determination step and said making a second determination step. 10. The method of claim 8, wherein said image and said second image are mutually orthogonal. 11. A system for positioning an electromagnetic tracking sensor, said system including: a sensor adapted to attach to a patient anatomy;an imaging modality adapted to obtain an image of said anatomy;an image processor adapted to indicate a size of each of a plurality of volumes on said image, said plurality of volumes including a predetermined smaller and larger volume, wherein each volume in said plurality of volumes represents at least one distance from said sensor, wherein said larger volume contains said smaller volume, wherein said image processor is further adapted to indicate a region of interest in said patient anatomy to aid in making a determination of whether said region of interest is encompassed ... | General Electric Company | A61B005/05
US07647088 | B2 | 2010-01-12 | Reconstruction method for images of the beating heart | 2006-09-07 | Projection views of the moving heart and stationary background tissues are acquired and processed to provide corresponding moving tissue projection views. An average image is reconstructed in a conventional manner and a moving tissue image is reconstructed using a highly constrained backprojection method and a composite image formed from selected moving tissue projection views. The average image is then combined with the moving tissue image. The method is disclosed in a cardiac gated MRI scan. | 1. A method for producing an image of a moving subject positioned in a field of view (FOV) of a medical imaging system, the steps comprising: a) acquiring with the medical imaging system a set of projection views of the subject for each of a series of image frames;b) producing from the projection views acquired in step a) corresponding moving tissue projection views in which substantially only moving tissues are depicted;c) producing from more than one set of the projection views acquired in step a) an average image;d) reconstructing a composite image for one of said image frames using moving tissue projection views corresponding to the projection views acquired for said image frame plus other selected moving tissue projection views;e) reconstructing a moving tissue image for said image frame by: e)i) back projecting the moving tissue projection views corresponding to the projection views acquired for said image frame and weighting the value backprojected into each image pixel by the value of the corresponding pixel in the composite image; ande)ii) summing the backprojected values for each image pixel; andf) combining the moving tissue image and the average image. 2. The method as recited in claim 1 in which step c) includes reconstructing the average image from a sufficient number of projection views acquired in step a) that static tissues in the FOV are depicted with few artifacts. 3. The method as recited in claim 1 in which step b) includes: b)i) reprojecting the average image to produce average image reprojections; and b)ii) subtracting average image reprojections from corresponding acquired projection views to produce corresponding moving tissue projection views. 4. The method as recited in claim 3 is which substantially all the projection views acquired in step a) are used to reconstruct the average image in step c). 5. The method as recited in claim 4 in which the moving tissue projection views are produced in step b) by reprojecting the average image at each of the view angles of the projection views acquired in step a). 6. The method as recited in claim 1 in which the acquisitions in step a) are cardiac gated and the series of image frames depict the moving subject at a succession of cardiac phases. 7. The method as recited in claim 6 in which the projection views acquired in step a) are acquired over a plurality of heartbeats. 8. The method as recited in claim 7 in which the medical imaging system is a magnetic resonance imaging system. 9. The method as recited in claim 1 in which said other moving tissue projection views are selected in step d) by: d)i) determining how much the subject has moved while acquiring each set of projection-views relative to the subject's position when the projection views for said image frame were acquired; andd)ii) selecting moving tissue projection views that were acquired when the subject was within a predetermined distance from said image frame position. 10. The method as recited in claim 1 in ... | Wisconsin Alumni Research Foundation | A61B005/055
US07660616 | B1 | 2010-02-09 | Implantable multi-wavelength oximeter sensor | 2005-09-20 | Implantable multi-wavelength oximetry sensors, which can be used to monitor a patient's blood oxygen saturation level, are described. The sensor includes an implantable sensor housing within which are located a plurality of light sources that each transmits light of a different wavelength. Also within the sensor housing are one or more surfaces that are configured to combine the light from the plurality of light sources into a beam of light for transmission through a portion of the housing (e.g., a window) through which light can exit and enter the housing. Additionally, a light detector is located within the sensor housing, to detect light scattered by blood back into the housing. This description is not intended to be a complete description of, or limit the scope of, the invention. Other features, aspects, and objects of the invention can be obtained from a review of the specification, the figures, and the claims. | 1. An implantable multi-wavelength oximetry sensor, comprising: an implantable sensor housing defining a window through which light can pass;a first light source, within said sensor housing, to transmit light having a first wavelength;a second light source, within said sensor housing, to transmit light having a second wavelength; anda plurality of prisms within said sensor housing, each prism including a plurality of surfaces;a dichroic surface formed on a first surface of one of said prisms;wherein said dichroic surface reflects light of the first wavelength and passes light of the second wavelength;wherein said first light source is mounted to a second surface of a first one of said prisms directly, or with a first lens therebetween such that said first light source is mounted to one side of said first lens and another side of said first lens is mounted to the second surface of said first one of said prisms;wherein said second light source is mounted to a second surface of a second one of said prisms directly, or with a second lens therebetween such that said second light source is mounted to one side of said second lens and another side of said second lens is mounted to the second surface of said second one of said prisms; andwherein said dichroic surface is angled relative to said first light source and said second light source such that light of the first wavelength transmitted by said first light source is reflected to travel in generally a same direction as light of the second wavelength transmitted by said second light source that passes through said dichroic surface; andwherein said surface on which said dichroic surface is formed is different than both said surface to which said first light source is mounted and said surface to which said second light source is mounted. 2. The sensor of claim 1, wherein said plurality of prisms are bonded together to form a three dimensional structure having inner surfaces where a pair of said prisms meet, and wherein said dichroic surface is formed on one of said inner surfaces. 3. The sensor of claim 2, wherein said three dimensional structure also has outer surfaces, and wherein said first and second light sources are mounted to two of said outer surfaces of said three dimensional structure. 4. The sensor of claim 2, wherein said plurality of prisms are bonded together by an optical cement or another clear epoxy. 5. The sensor of claim 3, wherein said three-dimensional structure, formed by said plurality of prisms bonded together, is cubical, rectangular or trapezoidal. 6. The sensor of claim 3, wherein light of the first wavelength that is reflected by said dichroic surface, and light of the second wavelength that passes through said dichroic surface, exit said three dimensional structure through one of said outer surfaces, which said one of said outer surfaces is concave to increase an exit angle. 7. The sensor of claim 1, wherein said window is positioned relative to said dichroic surface su... | Pacesetter, Inc. | A61B005/1455
US20100036287 | A1 | 2010-02-11 | Method and Device for Detecting Parameters for the Characterization of Motion Sequences at the Human or Animal Body | 2008-12-05 | The invention relates to a bending sensor for detecting function parameters for the characterization of motion sequences at the human or animal body. The bending sensor comprises a fixing element 20, in particular a fixing plaster, for fixing the bending sensor on the skin of the human or animal body. Furthermore, a bending-sensitive detector 10 for detecting bending parameters of the bending sensor is provided. The detected bending parameters such as, for instance, the bending angle, the bending rate, and/or the bending acceleration are stored in a data memory 30. The fixing element is extensible and comprises an extensible cavity for accommodating a measuring sensor of the detector. The measuring sensor is fixed at a reference point of the fixing element in the cavity. | 1. A bending sensor for detecting function parameters for the characterization of motion sequences at the human or animal body, comprising: a fixing element (20), in particular a fixing plaster, for fixing the bending sensor on the skin of the human or animal body;a bending-sensitive detector (10) for detecting bending parameters of the bending sensor, in particular a bending angle, a bending rate, and a bending acceleration, anda data memory (30) for storing the detected bending parameters; andcharacterized in that the fixing element is extensible and comprises an extensible cavity for accommodating a measuring sensor of the detector, wherein the measuring sensor is fixed at a reference point of the fixing element in the cavity. 2. The bending sensor according to claim 1, wherein the fixing element (20) comprises an elastic bottom layer with a biocompatible thermally activatable adhesive layer for application on the skin, and an elastic top layer, wherein the cavity is formed between the top layer and the bottom layer. 3. The bending sensor according to claims 1 or 2, wherein the measuring sensor comprises an optical fiber with a transmission function that changes as a function of a bending of the optical fiber. 4. The bending sensor according to claim 3, wherein the measuring sensor comprises a strain gauge with an electric impedance that changes with an extension or compression of the strain gauge. 5. The bending sensor according to claim 4, wherein the detector comprises a detection device, wherein the detection device detects a change of the electric impedance of the strain gauge (60, 70) or a change of the transmission function of the optical fiber. 6. The bending sensor according to claim 5, wherein the detection device is adapted to detect the change of the electric impedance of a strain gauge (60, 70) or the change of the transmission function of the optical fiber with a predetermined scanning frequency of, for instance, 100 Hz. 7. The bending sensor according to claim 6, comprising an electronic data memory for storing the digitized electronic signals output by the detection device. 8. The bending sensor according to any of claim 1, wherein the detector comprises a substrate that guarantees tensile strength and is elastically bendable on which the measuring sensor is fixed. 9. The bending sensor according to claim 8, wherein the substrate is manufactured of spring steel. 10. The bending sensor according to claims 8, wherein the detector comprises a plurality of measuring sensors that are fixed on opposite sides of the substrate. 11. The bending sensor according to any of the preceding claims 1, wherein the detector (10) comprises a plurality of measuring sensors for detecting the bending parameters in respectively different measurement zones (80). 12. The bending sensor according to claim 11, wherein the measuring sensors are arranged in a cascaded or overlapping manner. 13. The bending sensor according to any of claims 12, ... | - | A61B005/11 G01N003/20 G01N003/08 G01V007/00
US20100041968 | A1 | 2010-02-18 | IMAGE CAPTURE IN COMBINATION WITH VITAL SIGNS BEDSIDE MONITOR | 2008-03-26 | When monitoring a patient (32), objective data is readily collected using a vital signs monitor (VSM) (12) coupled to the patient (32) via a plurality of sensors (20). Subjective data, such as the appearance of a rash, wound or dressing, facial pallor or flush, facial expression indicative of pain, and the like, are additionally captured using an imager (22) coupled to the VSM (12). The imager (22) can be a two-dimensional barcode reader that captures a digital image of the patient (32) or portion thereof and relays the image data to the VSM (12). Images and patient vital sign data are then stored to an electronic medical record (14) and presented to a user or physician. Additionally, a standard reference color plate (30) with an optional barcode (34) can be placed on or near the patient (32), and the patient (32) and SCP (30) can be imaged. The imaged SCP (30) is then employed as a reference to perform color correction to permit a reviewing physician to evaluate the image for diagnosis. | 1. A system for concurrently capturing vital sign data and an image of a patient, including: a vital signs monitor that receives patient parameter data and image data;an imager, coupled to the VSM, that captures an image of a patient and transmits image data to the VSM; andone or more sensors that monitor at least one patient vital sign and transmit patient parameter data describing the at least one vital sign to the VSM. 2. The system according to claim 1, further including a standard color plate (SCP) that is associated with the image of the patient. 3. The system according to claim 2, wherein the SCP has at least one portion colored to approximate the normal skin tone of the patient. 4. The system according to claim 2, further including a barcode positioned on the SCP, the barcode includes information related to the color of the SCP. 5. The system according to claim 4, wherein the imager includes a barcode reader that captures the image of the patient, or a portion thereof, and transmits the image data to the VSM for storage in an electronic medical record (EMR) or display to a physician. 6. The system according to claim 5, wherein the barcode reader reads the barcode, and wherein the VSM employs information obtained from the barcode to calibrate at least one of color, contrast, or brightness in the image. 7. The system according to claim 1, wherein the imager includes a barcode reader that captures the image of the patient, or a portion thereof, and transmits the image data to the VSM for storage in an electronic medical record (EMR) or display to a physician. 8. The system according to claim 1, wherein the image of the patient is an image of at least one of the patient's face, a wound on the patient, a wound dressing on the patient, or a catheter insertion point. 9. The system according to claim 8, wherein the image of the patient's face is evaluated for information related to at least one of skin tone or expression of pain, for diagnosis of the patient's condition. 10. The system according to claim 1, wherein the imager is a digital camera that captures images of the patient, or a portion thereof, and transmits the image data to the VSM for storage to an EMR or presentation to a physician, or both. 11. The system according to claim 2, wherein the memory includes: a routine or means for capturing the image of the patient and the SCR;a routine or means for receiving the image data and optionally receiving patient parameter data at the VSM;a routine or means for correcting for color distortion in the image at the VSM; anda routine or means for storing the image and optional patient parameter data at the VSM and displaying the stored data to a user. 12. A method of capturing vital sign and image data using the system of claim 1, including: connecting the patient to the VSM using the one or more sensors;capturing an image of the patient using the imager;receiving the image data and patient parameter data at the VSM;storing the image d... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/00 A61B006/00
US07662104 | B2 | 2010-02-16 | Method for correction of posture dependence on heart sounds | 2005-01-18 | A system to monitor heart sounds. The system comprises an implantable heart sound sensor operable to produce an electrical signal representative of at least one heart sound, a heart sound sensor interface circuit coupled to the heart sound sensor to produce a heart sound signal, an implantable posture sensor operable to produce an electrical signal representative of a patient's posture, and a controller circuit, coupled to the heart sound sensor interface circuit and the posture circuit. The controller circuit is operable to measure at least one heart sound in correspondence with at least one sensed patient posture. | 1. A system comprising: a device comprising: an implantable heart sound sensor operable to produce an electrical signal representative of at least one heart sound, the heart sound associated with mechanical activity of a patient's heart;a heart sound sensor interface circuit coupled to the heart sound sensor to produce a heart sound signal;an implantable posture sensor operable to produce an electrical signal representative of a patient's posture;a memory circuit; anda controller circuit, coupled to the heart sound sensor interface circuit and the posture circuit, wherein the controller circuit is configured to: determine a patient posture using the posture sensor signal;measure the at least one heart sound in correspondence with at least one corresponding determined patient posture;adjust the heart sound measurement by using the corresponding determined patient posture to reduce or remove variation in the heart sound measurement due to patient posture; andprovide the adjusted heart sound measurement to a user or an automated process. 2. The system of claim 1, wherein the controller is configured to determine the patient is in an upright posture and to measure the heart sound when the controller determines the patient is in the upright posture. 3. The system of claim 1, wherein the controller circuit is configured to: determine the patient is in each of a plurality of postures;measure the heart sound in correspondence with the determined postures; andreduce or remove variation in the heart sound measurement due to varying posture by relating a heart sound measurement according to a determined posture. 4. The system of claim 3, wherein the controller is configured to store at least one heart sound measurement in correspondence with or according to a determined patient posture. 5. The system of claim 3, wherein the controller is configured to reduce or remove variation in the heart sound information due to varying posture by applying a scaling factor to at least one heart sound measurement, wherein the scaling factor is a function of a patient posture. 6. The system of claim 1, wherein the device further includes a clock circuit and the controller circuit is operable to measure heart sounds at a time of day when a patient is in a recumbent posture. 7. The system of claim 1, wherein the controller circuit is operable to use the electrical signal provided by the implantable heart sound sensor to determine that mechanical interference is below a threshold level before a heart sound measurement is executed. 8. The system of claim 1, wherein the device further includes at least one implantable cardiac signal sensing circuit operable to detect at least one intrinsic cardiac signal, and wherein the controller is operable to measure the at least one heart sound in correspondence with at least one determined patient posture when a patient heart rate is below a specified heart rate threshold value. 9. The system of claim 1, wherein the device further... | Cardiac Pacemakers, Inc. | A61B005/02
US07648464 | B1 | 2010-01-19 | Detecting ischemia using an implantable cardiac device based on morphology of cardiac pressure signal | 2006-07-19 | Methods and systems are presented for using an ICD to detect myocardial ischemia. In one embodiment, a method includes sensing a signal indicative of cardiac pressure, measuring a height of the sensed signal at a peak amplitude of the sensed signal, and measuring a duration of the sensed signal. The method further includes indicating an ischemia based on a comparison of a ratio of the height to the duration with a predetermined value. In another embodiment, a method includes sensing a signal indicative of cardiac pressure, determining a derivative signal that is a first derivative of the sensed signal, measuring a maximum positive value of the derivative signal, and measuring a maximum negative value of the derivative signal. The method further includes indicating an ischemia based on a comparison of a ratio of the maximum positive value to the maximum negative value with a predetermined value. | 1. A method for detecting ischemia, comprising: (a) sensing via an implantable cardiac-rhythm-management device (ICRMD) a signal indicative of cardiac pressure;(b) measuring via a processor associated with the ICRMD a height of the sensed signal at a peak amplitude of the sensed signal;(c) measuring via the processor a duration of the sensed signal;(d) determining via the processor a ratio of the height to the duration; and(e) indicating via the processor an ischemia based on a comparison of the determined ratio to a predetermined value of the ratio. 2. The method of claim 1, wherein step (b) comprises: measuring a baseline value of the sensed signal; andmeasuring the height of the sensed signal from the baseline value to the peak amplitude. 3. The method of claim 1, wherein step (c) comprises measuring the duration of the sensed signal at an amplitude that is a percentage of the peak amplitude. 4. The method of claim 3, wherein step (c) comprises measuring the duration of the sensed signal at an amplitude that is about fifty percent of the peak amplitude. 5. The method of claim 1, wherein step (e) comprises indicating an ischemia based on a comparison of the ratio to a previously stored value of the ratio. 6. The method of claim 5, wherein step (e) comprises indicating an ischemia when the ratio is less than the previously stored value of the ratio. 7. A system for detecting ischemia, comprising: means for sensing a signal indicative of cardiac pressure;means for measuring a height of the sensed signal at a peak amplitude of the sensed signal;means for measuring a duration of the sensed signal;means for determining a ratio of the height to the duration; andmeans for indicating an ischemia based on a comparison of the determined ratio to a predetermined value of the ratio. 8. A system for detecting ischemia, comprising: a sensor that generates a signal indicative of cardiac pressure; andan ischemia detector that: measures a height of the sensed signal at a peak amplitude of the sensed signal;measures a duration of the sensed signal;determines a ratio of the height to the duration; andindicates an ischemia based on a comparison of the determined ratio to a predetermined value of the ratio. 9. The system of claim 8, wherein the ischemia detector measures a baseline value of the sensed signal and measures the height of the sensed signal from the baseline value to the peak amplitude. 10. The system of claim 8, wherein the ischemia detector measures the duration of the sensed signal at an amplitude that is a percentage of the peak amplitude. 11. The system of claim 10, wherein the ischemia detector measures the duration of the sensed signal at an amplitude that is about fifty percent of the peak amplitude. 12. The system of claim 8, wherein the ischemia detector indicates an ischemia based on a comparison of the ratio to a previously stored value of the ratio. 13. The system of claim 12, wherein the ischemia detector indicates an ischemia ... | Pacesetter, Inc. | A61B005/021
US07666390 | B2 | 2010-02-23 | DTPA derivatives and metal complexes thereof, and radiation source and contrast agent including the metal complexes | 2006-04-25 | The present invention relates to diethylenetriamine pentaacetic acid (DTPA) derivates and metal complexes thereof, and radiation sources and contrast agents including the metal complexes. More particularly, the present invention relates to DTPA derivatives, which contain iodine and are useful as contrast agents for diagnosing renal function, metal complexes of the DTPA derivatives with 99mTc, 166Ho, 111In, 90Y, 153Sm, 186Re, 188Re, 68Ga, or 177Lu, which are useful as liquid radiation sources for treating vascular stenosis and contrast agents for diagnosing renal function, and radiation sources and contrast agents including the metal complexes. The DTPA derivatives and metal complexes thereof are safe because the metal complexes are excreted via the kidneys and bladder within several minutes from the time of use as liquid radiation sources for treating vascular stenosis, and are cost-effective because they are potentially useful as contrast agents and as agents for diagnosing renal function. | 1. A metal complex represented by Chemical Formula 4,whereinR is selected from the group consisting of:wherein, R1 is H, CH3 or C2H5; andM is selected from the group consisting of Y, Re, Lu, Ho, Dy, and Sm. 2. A method for treating vascular stenosis in a mammal comprising administering a liquid radiation source comprising a metal complex represented by Chemical Formula 3 to a mammal having vascular stenosis,whereinR is selected from the group consisting of:wherein, R1 is H, CH3 or C2H5; andM is selected from the group consisting of 90Y, 186Re, 188Re, 177Lu, 166Ho, and 153Sm. 3. A method for treating vascular stenosis in a mammal comprising administering a liquid radiation source comprising a metal complex represented by Chemical Formula 4 to a mammal having vascular stenosis,whereinR is selected from the group consisting of:wherein,R1 is H, CH3 or C2H5; andM is selected from the group consisting of 90Y, 186Re, 188Re, 177Lu, 166Ho, and 153Sm. 4. The metal complex of claim 1, wherein Y is 90Y, Re is 186/188Re, Lu is 177Lu, Ho is 166Ho, and Sm is 153Sm. | Korea Atomic Energy Research Institute | A61K051/00 A61K049/04 A61M036/14 A61B005/055 C07F005/00
US20100049009 | A1 | 2010-02-25 | APPARATUS AND ARRANGEMENT FOR CAPTURING OR EVALUATING MEDICAL DATA ON BOARD A MEANS OF TRANSPORT | 2009-08-19 | An apparatus for capturing and/or evaluating medical data on board a means of transport, and a system including the apparatus. The apparatus includes at least one interface configured to capture or input medical data and a device configured to transmit medical data, audio data, or video data from the apparatus to or from a ground station situated outside the means of transport via data transmission means. The apparatus has a first and second mode of operation for transmitting data to or from the ground station. The first mode of operation is designed to use a first bandwidth sufficient for video transmission on the data transmission means. The second mode of operation is designed to use a second bandwidth narrower than the first bandwidth or a discontinuous data transmission mode on the data transmission means. | 1. An apparatus for capturing or evaluating medical data on board a means of transport, comprising: at least one interface configured to capture or input medical data; anda device configured to transmit medical data, audio data, or video data from the apparatus to or from a ground station situated outside the means of transport via a data transmitter,wherein the apparatus has a first and second mode of operation for transmitting data to or from the ground station, the first mode of operation designed to use a first bandwidth sufficient for video transmission on the data transmitter and the second mode of operation designed to use a second bandwidth narrower than the first bandwidth or a discontinuous data transmission mode on the data transmitter. 2. The apparatus according to claim 1, wherein the at least one interface comprises connections for sensors that measure vital functions of the body. 3. The apparatus according to claim 1, wherein the at least one interface is further configured to capture audio or video data or receive input of data by an ancillary person. 4. The apparatus according to claim 1, wherein the device comprises a wireless or wired network connection for connecting to a network in the means of transport. 5. The apparatus according to claim 1, wherein the first bandwidth is at least 64 kBit/s. 6. The apparatus according to claim 1, wherein the second bandwidth is at least 16 kBit/s. 7. The apparatus according to claim 1, wherein the apparatus is configured to automatically use the first mode of operation when a video transmission is performed. 8. The apparatus according to claim 1, further comprising an identification device which, when the first mode of operation is used, identifies the apparatus to an operator of the data transmitter and requests an appropriate bandwidth. 9. The apparatus according to claim 1, wherein the apparatus has a third mode of operation in which no data transmission from or to a ground station takes place. 10. The apparatus according to claim 1, further comprising a label printer. 11. A system for capturing or evaluating medical data on board a means of transport, comprising: an apparatus according to claim 1 on board the means of transport;a network configured to transmit data from the apparatus to a first transmission/reception device in the means of transport on board the means of transport;a ground station having a second transmission/reception device designed for data communication with the first transmission/reception device in the means of transport;devices in the ground station for data communication with a device configured to provide telemedical services. 12. A system according to claim 11, wherein the network on board the means of transport is a wireless network. 13. A system according to claim 11, wherein data communication between the first transmission/reception device in the means of transport and the ground station is effected using a satellite link. 14. A system accordi... | Lufthansa Technik AG | A61B005/00 H04W004/00
US07641614 | B2 | 2010-01-05 | Wearable blood pressure sensor and method of calibration | 2006-08-22 | Methods and apparatus for measuring arterial blood pressure at an extremity of a subject. Arterial blood pressure is derived from a circulatory measurement performed on an extremity of a subject and the circulatory measurement is normalized to account for the instantaneous vertical displacement of the extremity. The vertical displacement of the extremity relative to the heart of the subject is obtained using the angular orientation of the subject's extremity. An improved photoplethysmograph can discriminate light traversing the extremity from ambient light on the basis of differential response. The apparatus may have a conducting polymer actuator for applying pressure to the extremity of the subject. A pulsatile waveform from the photoplethysmographic signal may be obtained at a plurality of externally applied pressures to calibrate the photoplethysmograph. | 1. A method for performing a circulatory measurement on an extremity of a subject, the method comprising: measuring a plethysmographic signal of a plethysmographic sensor disposed on the extremity of the subject at a plurality of heights of the sensor, relative to the heart of the subject, and at a plurality of externally applied pressures at each height of the sensor;while measuring the plethysmographic signal, using a solid protrusion to externally apply a focal pressure to the extremity of the subject, the focal pressure having a value that may vary;determining a relationship characterizing variation of the plethysmographic signal with the plurality of heights of the sensor and the plurality of externally applied pressures; andderiving an arterial blood pressure feature based on at least the relationship characterizing variation of the plethysmographic signal with the plurality of heights of the sensor and the plurality of externally applied pressures, and on an instantaneously measured plethysmographic signal and height of the sensor relative to the heart of the subject. 2. A method in accordance with claim 1, further comprising: modeling the relationship of a derived arterial pressure to the plethysmographic signal by fitting parameters of a specified functional relationship. 3. A method in accordance with claim 1, further comprising: modeling the relationship of a derived mean arterial pressure to a peak plethysmographic signal by fitting parameters of a specified functional relationship. 4. A method in accordance with claim 1, wherein the plethysmographic signal is a photoplethysmographic signal. 5. A method in accordance with claim 1, wherein the focal pressure is additionally varied by means of an inflatable cuff. 6. A method in accordance with claim 1, wherein the focal pressure is additionally varied by means of a band using a polymer actuator. 7. A method in accordance with claim 1, wherein measurement of a plurality of heights of the sensor is derived from at least one angular displacement of the extremity. 8. A method in accordance with claim 7, wherein the at least one angular displacement of the extremity is obtained by means of an accelerometer. 9. A method in accordance with claim 1, further comprising: removing spurious data points from the plethysmographic signal. 10. A method in accordance with claim 1, further comprising: removing oscillations between beats of the heart from the plethysmographic signal. 11. A method in accordance with claim 1, further comprising: deriving a systolic pressure by multiplying a peak plethysmographic signal by a specified factor. 12. A method in accordance with claim 11, wherein the factor is approximately 0.5. 13. A method in accordance with claim 1, further comprising: deriving a diastolic pressure by multiplying a peak plethysmographic signal by a specified factor. 14. A method in accordance with claim 13, wherein the factor is between 0.75 and 0.85. 15. A method in accordance w... | Massachusetts Institute of Technology | A61B005/00
US07668590 | B1 | 2010-02-23 | Methods and devices for determining exercise diagnostic parameters | 2006-02-10 | A device, such as an implantable cardiac device, and methods for determining exercise diagnostic parameters of a patient are disclosed. Specifically, a maximum observed heart rate of a patient during exercise can be identified when an activity level and a heart rate measurement of the patient exceed predetermined thresholds. Included are methods for filtering out premature heartbeats or noise from the maximum heart rate determination. Methods of determining other exercise parameters, such as workload are also disclosed. The device includes hardware and/or software for performing the described methods. | 1. In an implantable cardiac device, a method for determining workload of a patient during exercise, comprising: (a) monitoring a changing heart rate of the patient and producing heart rate measurements;(b) monitoring activity level of the patient; and(c) determining workload of the patient using at least one heart rate measurement when the activity level exceeds an activity threshold. 2. The method of claim 1, wherein step (c) comprises determining workload of the patient when the activity level exceeds the activity threshold for a predetermined period of time. 3. The method of claim 1, wherein step (c) comprises determining workload of the patient using heart rate measurements over a predetermined period of time. 4. The method of claim 1, further comprising determining one or more of the following: heart rate intensity, percent oxygen consumption (% VO2) reserve, metabolic equivalents (METS), and percentage METS. 5. An implantable cardiac device for determining workload of a patient during exercise, comprising: means for monitoring a changing heart rate of the patient and producing heart rate measurements;means for monitoring activity level of the patient;means for determining work of the patient using at least one heart rate measurement when the activity level exceeds an activity threshold; andmeans for transmitting the work of the patient to an external device. 6. A computer readable medium having instructions that when executed on a processor of an implantable medical device (IMD) have cardiac pacing capabilities cause the IMD to: (a) monitor a changing heart rate of the patient and produce heart rate measurements;(b) monitor activity level of the patient; and(c) determine workload of the patient using at least one heart rate measurement when the activity level exceeds an activity threshold, wherein workload is calculated as the product of heart rate intensity and duration, wherein duration is the period of time during which the activity level exceeds an activity threshold. 7. The computer readable medium of claim 6, wherein the instructions cause the processor to determine workload of the patient using at least one heart rate measurement when the activity level exceeds an activity threshold for at least a predetermined period of time. 8. The computer readable medium of claim 6, wherein intensity is VO2observed. 9. The computer readable medium of claim 6, wherein intensity is % VO2 reserve. 10. The computer readable medium of claim 6, wherein the instructions further cause the processor to: (d) identify a heart rate as the maximum observed heart rate using the IMD; and(e) store the identified maximum observed heart rate in the IMD, when the following conditions occur: (i) the activity level exceeds an activity threshold for a predetermined period of time, (ii) a heart rate measurement is greater than a stored heart rate measurement, and (iii) a difference between the heart rate measurement and the stored heart rate measurement does ... | Pacesetter, Inc. | A61B005/04
US07664541 | B2 | 2010-02-16 | Multi-channel magnetic resonance imaging reconstruction method for water-fat separation | 2006-02-28 | In a multi-channel magnetic resonance imaging reconstruction method for water-fat separation, one in-phase image and two opposed-phase images are acquired with multiple reception coils in respective channels. The sensitivity distribution of the coils of the respective channels is calculated. The images of respective channels are merged dependent on the sensitivity distribution. A phase difference between the two opposed-phase images is calculated. At lease one characteristic region of the in-phase image is detected, which is used as a criterion for phase correction. The phases of the opposed-phase images are corrected and images of water and fat are calculated. The method provides stable and reliable imaging, has a short reconstruction time and also solves the problem that images of water and fat may be exchanged. | 1. A multi-channel magnetic resonance imaging reconstruction method for water-fat separation, comprising the steps of: (i) acquiring one in-phase image and two opposed-phase images using multiple reception coils in respective channels of a magnetic resonance imaging;(ii) calculating a sensitivity distribution of the coils of the respective channels;(iii) merging the images of respective channels on the basis of the sensitivity distribution;(iv) calculating a phase difference between the two opposed-phase images;(v) detecting at least one characteristic region of the in-phase image, which is used as a criterion for phase correction; and(vi) correcting phases of the merged opposed-phase images and calculating images of water and fat therefrom. 2. A method as claimed in claim 1, wherein step (ii) further comprises calculating the sum of squares of moduli of the in-phase images of respective channels, which is then extracted:SOS⁡(x,y)=∑i=1n⁢S0i⁡(x,y)2,wherein S0i(x,y) is the in-phase image of channel i, i=1 . . . n, and n is the total number of the channels. 3. A method as claimed in claim 2, wherein step (ii) further comprises selecting a channel signal having a relatively high SNR S01(x,y) among the in-phase images and normalizing the modulus thereof:S0′⁡(x,y)=S0I⁡(x,y)S0I⁡(x,y). 4. A method as claimed in claim 2, wherein step (ii) further comprises selecting a linear combination of respective channel signals and normalizing the modulus thereof:S0′⁡(x,y)=∑i=1n⁢S0i⁡(x,y)·Wi∑i=1n⁢S0i⁡(x,y)·Wi,wherein Wi is a weighting. 5. A method as claimed in claims 3, wherein step (ii) further comprises calculating the sensitivity distribution of coils of respective channels as:Pi⁡(x,y)=S0i⁡(x,y)·[S0′⁡(x,y)]•SOS⁡(x,y),wherein [ ]• represents the conjugate. 6. A method as claimed in claim 1 comprising, in step (ii), utilizes a corrected space matched filter method to calculate the sensitivity distribution of coils of respective channels, wherein signals of each pixel and its neighborhood of images of respective channels are used to calculate the signal and noise correlation matrix, and the sensitivity distribution of coils of respective channels is derived by calculating the characteristic value and characteristic vector of the signal correlation matrix. 7. A method as claimed in claim 5 comprising, in said step (iii), merging the images of respective channels to obtain merged images having an optimized SNR, and at the same time, reserving phases of the images as:S0⁡(x,y)=∑i=1n⁢[Pi⁡(x,y)]•·S0i⁡(x,y)S1⁡(x,y)=∑i=1n⁢[Pi⁡(x,y)]•·S1i⁡(x,y)S2⁡(x,y)=∑i=1n⁢[Pi⁡(x,y)]•·S2i⁡(x,y),wherein S1i(x,y) and S2i(x,y) represent the two opposed-phased images, respectively. 8. A method as claimed in claim 7 comprising, in said step (iv), determining said phase difference between the two opposed-phase images is: φ(x,y)=angle└S2(x,y)·[S1(x,y)]•┘,wherein angle[ ] represents the operator of calculating the phase angle. 9. A method as claimed in claim 8 comprising, using a pha... | Siemens Aktiengesellschaft | A61B005/05 G01V003/00
US07650193 | B2 | 2010-01-19 | Lead assembly with porous polyethylene cover | 2005-06-10 | This document discusses, among other things, a lead assembly including a porous polyethylene cover. In an example, the cover includes sections that have differing pore sizes. In an example, a section of the cover near a distal end portion of a lead assembly includes pores that are large enough to allow tissue ingrowth. In another example, a lead assembly includes two or more polyethylene covers having different porosities. | 1. A lead assembly comprising: a lead body;a conductor extending through the lead body;an electrode coupled to the conductor;a porous polyethylene cover extending over the electrode; andwherein the cover includes a first section having tissue ingrowth allowing pores and a second section having tissue ingrowth inhibiting pores, and the first section is at a different longitudinal location than the second section. 2. The lead assembly of claim 1, wherein the cover is formed from at least one piece of polyethylene wrapped around at least a portion of the electrode, the first section of the cover wrapped with a first tension, and the second section wrapped with a second tension that is different from the first tension. 3. The lead assembly of claim 2, wherein the piece of porous polyethylene is laser-sintered to itself. 4. The lead assembly of claim 3, wherein the piece of porous polyethylene is spirally wrapped around the electrode. 5. The lead assembly of claim 1, where the first section is formed from a first piece of porous polyethylene and the second section is formed from a second piece of porous polyethylene, the first piece of porous polyethylene having larger pores than the second piece of porous polyethylene. 6. The lead assembly of claim 5, wherein the first piece of porous polyethylene is joined to the second piece of porous polyethylene. 7. The lead assembly of claim 1, wherein the second section of the cover is proximate a distal end portion of the lead assembly. 8. The lead assembly of claim 1, wherein the cover extends over substantially all of the lead body. 9. The lead assembly of claim 1, wherein the porous polyethylene cover is a piece of mechanically stretched ultra high molecular weight polyethylene wrapped around at least a portion of the electrode. 10. The lead assembly of claim 9, wherein the piece of mechanically stretched ultra high molecular weight polyethylene is wrapped around at least a portion of the lead body. 11. The lead assembly of claim 9, wherein the piece of mechanically stretched ultra high molecular weight polyethylene is hydrophilic. 12. A lead assembly comprising: a lead body;a conductor extending through the lead body;an electrode coupled to the conductor;a porous polyethylene cover extending over the electrode, the cover includes a first section having tissue ingrowth allowing pores and a second section having tissue ingrowth inhibiting pores;the first section having larger pores than the second section of porous polyethylene, and the first section is joined with the second section and is a continuous cover over the electrode, and the first section is at a different longitudinal location than the second section. 13. The lead assembly as recited in claim 12, wherein the pores are filled with conductive substance. 14. The lead assembly as recited in claim 12, wherein the first section is fused with the second section. 15. The lead assembly as recited in claim 12, wherein the porous polyethylene... | Cardiac Pacemakers, Inc. | A61N001/04 A61B005/04
US07641619 | B2 | 2010-01-05 | Barometric pressure correction based on remote sources of information | 2007-09-19 | The invention includes systems and methods for converting absolute pressure data acquired in anatomical environments to gauge pressure data using an implant configured to monitor pressure. The implant is configured to communicate with an external controller, which is configured to communicate with a remote microprocessor that includes real-time barometric pressure data for one or more geographic locations. | 1. A system for determining gauge pressure within a patient's body, comprising: an implantable medical device including at least one sensor configured to sense absolute pressure at a location within the body;an acoustic transducer operable to wirelessly transmit absolute pressure information to at least one device located inside or outside of the body, the absolute pressure information including an absolute pressure value corresponding to the absolute pressure sensed by the implantable medical device;a switch coupled between the acoustic transducer and an energy storage device, the energy storage device adapted to supply operating current to an electrical circuit within the implantable medical device upon the acoustic transducer receiving an acoustic wake-up signal from an acoustic energy source, the switch configured to actuate the implantable medical device between an inactive state, in which the operating current is limited from flowing from the energy storage device to the electrical circuit, and an active state that allows operating current to flow from the energy storage device to the electrical circuit; anda remote device operable to: receive the absolute pressure value; andcalculate a gauge pressure value based at least in part on the absolute pressure value. 2. The system of claim 1, wherein the remote device is further configured to calculate the gauge pressure value based on position data corresponding to a geographic location of the patient's body. 3. The system of claim 2, wherein the gauge pressure value is derived from the barometric pressure at the geographic location of the body. 4. The system of claim 2, wherein the position data includes altitude data from an altimeter. 5. The system of claim 2, wherein the position data includes data from a GPS unit. 6. The system of claim 2, wherein the position data includes a telephone number or zip code. 7. The system of claim 1, wherein the acoustic transducer is a part of the implantable medical device. 8. The system of claim 1, wherein the acoustic transducer is a separate component from the implantable medical device. 9. The system of claim 1, wherein the at least one sensor includes a pressure sensor. 10. The system of claim 1, wherein the implantable medical device is a biosensor. 11. The system of claim 1, wherein the remote device is an external monitor. 12. The system of claim 1, wherein the implantable medical device includes an internal controller configured to receive the gauge pressure value. 13. A method for measuring gauge pressure in a patient's body, comprising: transmitting an acoustic wake-up signal to an implantable medical device located within the body, the implantable medical device including an acoustic transducer and a switch coupled between the acoustic transducer and an energy storage device, the energy storage device adapted to supply operating current to an electrical circuit within the implantable medical device, the acoustic wake-up signal configu... | Remon Medical Technologies, Ltd. | A61B005/00
US07939806 | B2 | 2011-05-10 | Nuclear medicine imaging system having attenuation correction | 2008-08-05 | A nuclear medicine (NM) imaging system having attenuation correction (AC) has a stationary portion having an NM rotating portion mounted thereon. At least one gamma camera head is mounted on the NM rotating portion proximate a front side of the stationary portion and is configured to be rotated about a central axis. An AC portion comprises an X-ray source mounted on an AC rotating portion. The AC rotating portion is mounted proximate a back side of the stationary portion. The X-ray source is configured to be rotated about the central axis. | 1. A nuclear medicine imaging system having attenuation correction, comprising: a stationary portion having front and back sides and an opening that extends between the front and back sides along a central axis;a nuclear medicine (NM) rotating portion rotatably mounted to the stationary portion, the NM rotating portion extending along the central axis from the front side of the stationary portion, through the opening, and projecting beyond the back side, at least one gamma camera head mounted on the NM rotating portion proximate the front side of the stationary portion and configured to be rotated about the central axis; andan attenuation correction (AC) rotating portion mounted proximate to the back side of the stationary portion, an X-ray source mounted on the AC rotating portion and configured to be rotated about the central axis. 2. The system of claim 1, wherein the AC rotating portion is mounted to the NM rotating portion, the AC rotating portion being spaced apart from the back side of the stationary portion. 3. The system of claim 2, further comprising one of an NM retractor or an NM slip ring mounted proximate to the back side of the stationary portion, said one of an NM retractor or an NM slip ring being located between the back side of the stationary portion and the AC rotating portion. 4. The system of claim 1, further comprising an AC bearing mounted to the NM rotating portion proximate the back side, the AC rotating portion being configured to be mounted to the AC bearing. 5. The system of claim 1, wherein the front and back sides face in opposite directions along the central axis, the system further comprising an NM bearing mounted to the front side of the stationary portion for interconnecting the NM rotating portion to the stationary portion. 6. The system of claim 1, further comprising one of an NM retractor or an NM slip ring mounted proximate to the back side of the stationary portion, and a communicative connection that is configured to route at least one of power, image data, or commands between the NM rotating portion on the front side of the stationary portion and said one of the NM retractor or NM slip ring. 7. The system of claim 1, further comprising an AC slip ring mounted between the X-ray source and the stationary portion, the AC slip ring configured to communicate at least one of power, image data, or commands between the AC rotating portion and a brush block. 8. The system of claim 1 further comprising an AC slip ring mounted between the X-ray source and the back side of the stationary portion, the AC slip ring comprising a brush block mounted on the NM rotating portion. 9. The system of claim 1, further comprising: an NM retractor mounted to the back side of the stationary portion; andan AC slip ring mounted between the X-ray source and the back side of the stationary portion, the AC slip ring comprising a brush block mounted on the NM retractor. 10. The system of claim 1, further comprising a motor confi... | General Electric Company | G01N023/08 G01T001/166 G01T001/167
US08202232 | B2 | 2012-06-19 | Lancing system | 2009-02-27 | The invention relates to a lancing system having a lancet carrier tape (3) supporting a plurality of lancets (4), a transport means (5, 11, 17) serving to move the lancet carrier tape (3) in a feed or transport direction along a transport path and to thereby move lancets (4) of the lancet carrier tape (3) one after the other to a puncturing position, and a lancing drive (6) for causing a lancet (4) that occupies a puncturing position to perform a puncturing movement. It is provided according to the invention that during a puncturing movement the lancing drive (6) moves a lancet (4), that has been brought to a puncturing position, in the puncturing direction together with a portion of the lancet carrier tape (3) on which that lancet (4) is supported and that once a lancet (4) has been moved to the puncturing position at least one element of the transport means (5, 11, 17), arranged behind the puncturing position in the transport direction, performs a movement before or during the puncturing movement of that lancet (4). | 1. Lancing system having a lancet carrier tape that supports a plurality of lancets arranged transversely to its longitudinal direction;a transport means serving to move the lancet carrier tape in a transport along a transport path and to thereby move lancets of the lancet carrier tape one after the other to a puncturing position; anda lancing drive for causing a lancet that occupies a puncturing position to perform a puncturing movement, whereinduring a puncturing movement the lancing drive moves a lancet, that has been brought to a puncturing position, in the puncturing direction together with a portion of the lancet carrier tape on which that lancet is supported and thatonce a lancet has been moved to the puncturing position at least one element of the transport means, arranged behind the puncturing position in the transport direction, performs a compensating movement before or during the puncturing movement of that lancet; andwherein the lancet remains fixed to the portion of the lancet carrier tape during the puncturing movement, wherein the compensating movement allows the lancet and the portion of the lancet carrier tape to move together during the puncturing movement. 2. The lancing system as defined in claim 1, wherein the transport means comprises a winding device that successively brings lancets to the puncturing position by winding up the lancet carrier tape by steps. 3. The lancing system as defined in claim 2, wherein the movement of the at least one part of the transport means is a reverse rotary step of the winding device, and that the winding device performs a forward rotary step to move a lancet to the puncturing position, and then performs the reverse rotary step to unwind a length of tape from the winding device for a puncturing movement of the lancet that has been moved to the puncturing position. 4. The lancing system as defined in claim 1, wherein the reverse rotary step is effected during the puncturing movement. 5. The lancing system as defined in claim 1, wherein a driving element, which en-gages the winding device during the forward rotary step to drive the winding de-vice, and is uncoupled from the winding device for the reverse rotary step. 6. The lancing system as defined in claim 1, wherein the transport means comprises at least one guide element, arranged behind the puncturing position in the transport direction, which is located in a bend of the transport path and about which the lancet carrier tape is bent. 7. The lancing system as defined in claim 6, wherein the guide element can be moved between a first position in which the lancet carrier tape is bent about the guide element at a first bending angle, and a second position in which the lancet carrier tape is bent to a lesser degree or not at all, the guide element being moved toward the second position once a lance has been located in the puncturing position. 8. The lancing system as defined in claim 6, wherein two tape-guiding elements which enclose the... | Roche Diagnostics Operations, Inc. | A61B005/00
US08317716 | B2 | 2012-11-27 | Diagnosis system of deficient and forceful pulse | 2007-09-28 | Disclosed herein is a system for diagnosing a deficient pulse and an forceful pulse. The system includes a pulse diagnotic device, a deficient pulse and forceful pulse determining device, and an output device. The pulse diagnotic device measures pulse condition information at an examinee's Cun (˜\f˜) Gu (H), and Chi (,R) pulse-taking locations on his or her wrist using one or more pulse-taking sensors. The deficient pulse and forceful pulse determining device is operably connected to the pulse diagnotic device, analyzes the pulse pressure information measured by the pulse diagnotic device, calculates a quantified deficiency/forceful coefficient, and determines whether a pulse of interest is a deficient pulse or an forceful pulse. The output device is connected to the determining device and displays results of the determination. | 1. A system for diagnosing a deficient pulse and a forceful pulse, comprising: a pulse diagnostic device for measuring pulse condition information at an examinee's Cun , Gu , and Chi pulse-taking locations on his or her wrist using one or more pulse-taking sensors;a deficient pulse and forceful pulse determining device operably connected to the pulse diagnostic device, and configured to analyze the pulse pressure information measured by the pulse diagnostic device, to calculate a quantified deficiency/forceful coefficient and to determine whether a pulse of interest is a deficient pulse or an forceful pulse and an output device connected to the determining device and configured to display results of the determination,wherein the deficient pulse and forceful pulse determining device includes storage means for receiving and storing pulse pressure information, measured by the pressure sensor of the pulse diagnostic device, which comes into contact with a pulse taking location;graphing means for receiving the stored pulse pressure information and converting variation in pulse pressure over time into a graph;selection means for selecting ‘maximum pulse pressure during systole hP’, ‘maximum pulse pressure during diastole hD’, ‘inflection point time period between a first peak and a second peak during systole tQ’, ‘systole time tC’, ‘wave width WP’, ‘systole area AS’ and ‘diastole area AD’, and ‘wave angle θP’ from the graph as calculation variables;deficiency/forceful coefficient calculation means for calculating a deficiency/forceful coefficient using the selected calculation variables;deficient and forceful pulse determination means for determining whether an examinee's pulse is a deficient pulse or an forceful pulse based on the calculated deficiency/forceful coefficient; andtransmission means for transmitting results of the deficient and forceful pulse determination to the output device,wherein the ‘maximum pulse pressure during systole hP’ is defined as a pulse pressure-axis maximum value of a first peak during systole, the ‘maximum pulse pressure during diastole hD’ is defined as a pulse pressure-axis maximum value of a peak during diastole, the ‘inflection point time period between a first peak and a second peak during systole tQ’ is defined as a difference between a time-axis value of an inflection point Q between the first peak and the second peak during systole and a time-axis value of a point S at which the first peak during systole is created, the ‘systole time tC’ is defined as a difference between a time-axis value of an inflection point C between the second peak during systole and the peak during diastole and a time-axis value of a point S at which a pulse pressure wave form during systole starts, the ‘wave width WP’ is defined as a difference ‘b-a’, when the x-axis coordinate value of a leftmost point among the points at which a straight line extending through the pulse pressure values that correspond to ⅔ of the representative pulse h... | Korea Institute of Oriental Medicine | A61B005/02
US08577452 | B2 | 2013-11-05 | Method of acquiring a physiological response | 2007-07-12 | A method for acquiring a physiological response from a test subject, particularly an auditory brainstem response, by presenting a plurality of stimuli, such as a Maximum Length Sequence, detecting electrophysiological signals in response to the stimuli, generating a recover signal based on the stimuli and determining the physiological response from the electrophysiological signals and the recovery signal. | 1. A method for acquiring a physiological response, the method being performed by an automated auditory brainstem response (A-ABR) hardware apparatus and comprising the steps of: presenting to a test subject, by the A-ABR hardware apparatus, a plurality of stimuli with variable inter-stimulus intervals;detecting, by the A-ABR hardware apparatus, an electrophysiological signal from the test subject in response to the stimuli;generating, by the A-ABR hardware apparatus, a recovery signal based on the stimuli, where an amplitude and latency of the recovery signal are modified according to the inter-stimulus intervals of the stimuli; anddetermining, by the A-ABR hardware apparatus, a physiological response signal using both the electrophysiological signal and the recovery signal. 2. The method of claim 1 wherein the stimuli are presented cyclically. 3. The method of claim 1 wherein the plurality of stimuli is in excess of 1,000 auditory stimuli presented to the test subject's ear. 4. The method of claim 1 wherein the stimuli are auditory stimuli and the physiological response is an auditory brainstem response (ABR). 5. The method of claim 1 wherein the auditory stimuli are selected from one of: broad band clicks; tone bursts; noise bursts; or chirp stimuli. 6. The method of claim 1 wherein the plurality of stimuli is a Maximum Length Sequence and the length of the sequence is longer than the expected length of the electrophysiological signal. 7. The method of claim 1 wherein generating a recovery signal based on the auditory stimuli comprises application of a comb filter. 8. The method of claim 1 wherein the step of generating a recovery signal produces a unipolar recovery sequence. 9. The method of claim 1 wherein the step of generating a recovery signal produces a bipolar recovery sequence. 10. The method of claim 1 wherein the step of generating a recovery signal further includes the step of determining the amplitude and latency by reference to functional relationships determined from a control group with known physiological response. 11. The method of claim 1 wherein the step of determining a physiological response signal includes performing artefact detection to remove signals contaminated by artefacts. 12. The method of claim 1 wherein the step of determining a physiological response signal includes calculating an ensemble average of signals. 13. The method of claim 12 further including the step of cross-correlating the ensemble average with the recovery signal to obtain the physiological response. 14. The method of claim 1 wherein the method is applied to both ears of a single test subject simultaneously. 15. The method of claim 14 wherein the plurality of stimuli is a Maximum Length Sequence and the length of the sequence is longer than twice the expected length of the electrophysiological signal. 16. The method of claim 1 wherein the test subject is a human neonate. 17. The method of claim 1 wherein the plurality of stimuli a... | The University of Queensland | A61B005/04
US08532747 | B2 | 2013-09-10 | Biopsy marker delivery device | 2008-08-22 | A biopsy device for delivering a biopsy marker to a biopsy site is disclosed. The biopsy device can include a marker deployer having a unitary endpiece disposed in a distal end of the cannula. The distal endpiece can be a molded component which includes a distal tip, a marker deployment ramp, and a marker engaging element that aids in retaining the marker in the cannula until the marker is meant to be deployed. | 1. A biopsy marker delivery device comprising: a cannula having an internal lumen extending from a proximal end of the cannula to the distal end of the cannula, and a marker exit formed in a sidewall of the cannula proximal to the distal end of the cannula, wherein the cannula defines a longitudinal axis extending therethrough, wherein the distal end of the cannula defines an opening;a distal tip member closing the distal end of the cannula, wherein a proximal portion of the distal tip member extends proximally through the opening defined by the distal end of the cannula and along the interior of the sidewall of the cannula such that wherein the cannula extends distally beyond a proximal portion of the distal tip, wherein the proximal portion of the distal tip member is positioned within the cannula, wherein a distal portion of the distal tip member is positioned distal to the distal end of the cannula;a ramp positioned adjacent the marker exit, wherein the ramp defines a first included angle with the longitudinal axis, wherein the ramp forms a first planar surface, wherein the ramp is formed by the proximal portion of the distal tip member;a marker engaging element disposed within the internal lumen of the cannula, wherein at least a portion of the marker engaging element is disposed distally of a proximal end of the marker exit, wherein the marker engaging element comprises a tapered proximal end defining a second included angle with the longitudinal axis, wherein the tapered proximal end forms a second planar surface, wherein the marker engaging element extends proximally from the ramp, wherein the marker engaging element extends proximally along the sidewall of the cannula, wherein the marker engaging element has a thickness of at least twice the thickness of the sidewall of the cannula, wherein the marker engaging element is formed by the proximal portion of the distal tip member; andat least one marker disposed within the internal lumen of the cannula;wherein the cannula is transparent and the distal tip member is radiopaque. 2. The device of claim 1 wherein the ramp and marker engaging element comprise a unitary component. 3. The device of claim 1 wherein the distal tip member, ramp, and marker engaging element comprise a unitary component. 4. The device of claim 1 wherein the distal tip member, ramp, and marker engaging element comprise a unitary molded component. 5. The device of claim 1 wherein the cannula comprises a relatively flexible thin walled tube, and wherein the marker engaging element comprises a relatively stiff member disposed within the internal lumen. 6. The device of claim 1 wherein the marker engaging element comprises a step having a uniform thickness along at least a portion of its axial length. 7. The device of claim 1 wherein the distal tip member is relatively more radiopaque than the cannula. 8. The device of claim 7 wherein the distal tip member, ramp, and the marker engaging element are relatively more rad... | Devicor Medical Products, Inc. | A61B006/00
US08439838 | B2 | 2013-05-14 | Inserter for transcutaneous sensor | 2007-06-07 | The invention relates to an inserter for a transcutaneous sensor comprising a sensor part e.g. for registration of the bloods content of glucose. The inserter comprises a needle unit comprising a needle hub and a carrier body, and a sensor housing. The sensor housing and the needle hub are releasably connected and when they are connected, the insertion needle is placed along the sensor e.g. surrounding the sensor wholly or partly. The carrier body guides the movement relative to the housing between a retracted and an advanced position. When released the needle unit and the sensor housing are forced by a spring unit to an advanced position where the needle and sensor are placed subcutaneously. The object of the invention is to provide a disposable inserter for a transcutaneous sensor which inserter is easy and safe for the user to handle during use and to dispose of after use. | 1. An inserter for a transcutaneous sensor comprising: a housing, the housing comprising housing guiding members on an internal surface of the housing;a carrier body, the carrier body comprising carrier body guiding members on an external surface of the carrier body, wherein the carrier body guiding members and the housing guiding members guide movement of the carrier body relative to the housing between a retracted and an advanced position;a needle hub comprising an insertion needle for piercing of the skin and an unreleasable locking member for interacting with a corresponding locking member on the carrier body;a sensor housing comprising a sensor part configured for subcutaneous placement, the sensor housing and the needle hub being releasably connected to each other, wherein the insertion needle is adjoined to the sensor part when the sensor housing and the needle hub are connected, wherein the insertion needle has a proximal and distal end along a longitudinal axis, wherein the sensor part has a proximal and distal end along the longitudinal axis therebetween, and wherein the needle at least partly surrounds the sensor part along the longitudinal axis; anda spring unit connected to a release member, the release member being activatable so that the spring unit moves the sensor housing, the needle hub and the carrier body to an advanced position for subcutaneous placement of the needle and sensor part. 2. The inserter according to claim 1, wherein the needle hub and the carrier body are created as a single unit. 3. The inserter according to claim 1, wherein the needle hub comprises openings and the carrier body is provided with projections corresponding to the openings in the needle hub. 4. The inserter according to claim 1, wherein a needle unit comprising the needle hub and the carrier body after insertion can be placed in a retracted position. 5. The inserter according to claim 1, wherein a lower base of the housing is formed with a projecting part. 6. The inserter according to claim 5, wherein the projecting part forms an angle with the longitudinal direction of the insertion needle. 7. The inserter according to claim 1, wherein the inserter further comprises a stopper. 8. The inserter according to claim 7, wherein the stopper comprises at least one end of a track for a flange. 9. The inserter as claimed in claim 1, wherein the housing forms at least a part of the delivery packing for the inserter. 10. The inserter as claimed in claim 1, wherein the housing is provided with a substantially rigid cover. 11. The inserter as claimed in claim 1, wherein the spring unit is fastened to the housing in a first position and to the carrier body or the needle hub in a second position, and the first position is situated closer to a front end of the housing than the second position when the spring unit is unbiased. 12. The inserter as claimed in claim 1, wherein the spring unit is an elastic O-ring. 13. The inserter as claimed in claim 1, w... | Unomedical A/S | A61B005/00
US08556819 | B2 | 2013-10-15 | Treatment device with memory-supported control means | 2007-09-11 | The therapy device stores the time profiles of the blood pressure and of other influencing variables during preceding therapy procedures. The time profile of the blood pressure and of other influencing variables during the actual therapy is also recorded. An analytical comparison is then made between the time profiles of the blood pressure and of other influencing variables during the actual measurement procedure and those during preceding therapy procedures. Depending on the comparison, a prospective blood pressure control is carried out in order to prevent drops in blood pressure in a subsequent time period. | 1. A therapy device for extracorporeal blood treatment, comprising: a blood-pressure measuring device and a memory-supported controller that controls blood pressure of a patient using settable treatment parameters, wherein said memory-supported controller is configured to perform the following functions during a current therapy procedure:a) detecting and storing a temporal development of the blood pressure during the current therapy procedure;b) analytically comparing the temporal development of the blood pressure up to a current point of time during the current therapy procedure with temporal developments of the blood pressure of preceding therapy procedures up to a respective point of time corresponding to the current point of time of the current therapy procedure;c) selecting the temporal developments of the blood pressure of the preceding therapy procedures that are most similar up to the respective point of time to the temporal development of the blood pressure of the current therapy procedure up to the current point of time of the current therapy procedure;d) analyzing the selected temporal developments of the blood pressure of the preceding therapy procedures to detect in the selected temporal developments of the blood pressure of the preceding therapy procedures a decrease in the blood pressure beyond the respective point of time;e) choosing at least one of the selected temporal developments of the blood pressure of the preceding therapy procedures having the decrease in the blood pressure beyond the respective point of time; andf) preventing a decrease in the blood pressure of the current therapy procedure, wherein the temporal development of the at least one the selected temporal developments of the blood pressure of the preceding therapy procedures beyond the respective point of time used as a guide value to prospectively control the blood-pressure in time periods beyond the current point of time of the current therapy procedure. 2. The therapy device according to claim 1, wherein said controller is configured to repeatedly perform functions (b)-(e). 3. The therapy device according to claim 1, wherein said controller is operative to decide from which preceding therapy procedures the guide value is to be taken over for the time period beyond the current point of time of the current therapy procedure. 4. The therapy device according to claim 1, wherein temporal developments of blood-pressure-relevant influencing variables and the settable treatment parameters are detected and stored; andat least one of said temporal developments of blood-pressure-relevant influencing variables from the selected preceding therapy procedures is used as an additional guide value in time periods following the current point of time of the current therapy procedure. 5. The therapy device according to claim 1, wherein the detection of a similarity of preceding therapy procedures to the current therapy procedure is based on the time periods from the start of... | B. Braun Avitum AG | A61B005/02 B01D011/00 B01D061/00
US09301712 | B2 | 2016-04-05 | Method and apparatus for continuous measurement of motor symptoms in parkinson's disease and essential tremor with wearable sensors | 2009-07-28 | Disclosed embodiments include an apparatus for generating a plurality of movement impairment indices from one or more kinematic signals to characterize movement disorders. Additionally we disclose methods for generating a plurality of movement impairment indices from one or more kinematic signals obtained from one or more kinematic sensors, said methods implemented in a digital computer with one or more processors in order to characterize movement disorders based on spectral analysis, regularity metrics, and time-frequency analysis. | 1. A method for generating a plurality of movement impairment indices from one or more kinematic signals obtained from one or more kinematic sensors, said method implemented in a digital computer with one or more processors in order to characterize movement disorders, said method comprising: (a) collecting said one or more kinematic signals from said one or more kinematic sensors in a wearable inertial monitoring system configured for continuous measurement of movement;(b) processing said one or more kinematic signals to reject artifacts and noise resulting in one or more enhanced kinematic signals;(c) finding a measure of power spectral density of said one or more enhanced kinematic signals; and(d) calculating one or more impairment indices to characterize movement disorders in Parkinson's disease or essential tremor from said measure of power spectral density. 2. The method of claim 1, further comprising detrending said one or more enhanced kinematic signals prior to finding said measure of power spectral density of said one or more enhanced kinematic signals to eliminate low frequency components of said one or more kinematic signals. 3. The method of claim 2, further comprising generating one or more total magnitude bandlimited signal powers from said one or more enhanced kinematic signals. 4. The method of claim 3, wherein said one or more total magnitude bandlimited signal powers correspond to a total magnitude of acceleration or total rotational rate within a given frequency band characterizing specific impairment conditions. 5. The method of claim 4, further comprising postprocessing said plurality of impairment indices from said measure of power spectral density to generate a plurality of smoother impairment indices for continuous quantification of tremor, dyskinesia, and bradykinesia as a function of time based on said one or more kinematic sensors. 6. The method of claim 1, wherein calculating said plurality of impairment indices from said measure of power spectral density involves computing a plurality of regularity metrics. 7. The method of claim 6, wherein said plurality of regularity metrics are chosen from the group consisting of spectral flatness, eigenvalue spread, Lempel-Ziv complexity, Approximate Entropy, Sample Entropy, Multi-Scale entropy, nonlinear metrics, and complexity metrics. 8. The method of claim 1, wherein finding said measure of power spectral density of said one or more enhanced kinematic signals includes generating a time-frequency analysis plot. 9. The method of claim 8, wherein said time-frequency analysis plot is a spectrogram showing spectral changes as a function of time. 10. The method of claim 1, wherein finding said measure of power spectral density of said one or more enhanced kinematic signals involves using a power spectral estimation method chosen from the group consisting of periodogram, modified periodogram, Barletts's method, Welch's method, Blackman-Tukey's method, autoregressive spectrum... | PORTLAND STATE UNIVERSITY APDM, INC. | A61B005/11 A61B005/00
US09776008 | B2 | 2017-10-03 | Posture state responsive therapy delivery using dwell times | 2009-04-30 | Techniques related to classifying a posture state of a living body are disclosed. One aspect relates to sensing at least one signal indicative of a posture state of a living body. Posture state detection logic classifies the living body as being in a posture state based on the at least one signal, wherein this classification may take into account at least one of posture and activity state of the living body. The posture state detection logic further determines whether the living body is classified in the posture state for at least a predetermined period of time. Response logic is described that initiates a response as a result of the body being classified in the posture state only after the living body has maintained the classified posture state for at least the predetermined period of time. This response may involve a change in therapy, such as neurostimulation therapy, that is delivered to the living body. | 1. A medical device system, comprising: a sensor configured to sense at least one signal indicative of a posture state of a living body;posture state detection logic configured to classify the living body as being in a posture state based on the at least one signal and to determine whether the living body is classified in the posture state for at least a predetermined period of time; andresponse logic configured to initiate a response as a result of the living body being classified in the posture state, the response being initiated only after the living body has maintained the classified posture state for at least the predetermined period of time. 2. A medical device system, comprising: a sensor configured to sense at least one signal indicative of a posture state of a living body;posture state detection logic configured to classify the living body as being in a posture state based on the at least one signal and to determine whether the living body is classified in the posture state for at least a predetermined period of time;response logic configured to initiate a response as a result of the living body being classified in the posture state, the response being initiated only after the living body has maintained the classified posture state for at least the predetermined period of time; anda therapy delivery module configured to deliver therapy to the living body, and wherein the response is a change in the therapy delivered by the therapy delivery module. 3. The system of claim 2, wherein the posture state detection logic comprises at least one of a posture classifier configured to classify the posture of the living body and an activity classifier configured to classify the activity state of the living body, and wherein the classification of the living body as being in a posture state is based on at least one of the posture classification and the activity state classification. 4. The system of claim 2, wherein the posture state detection logic comprises dwell time logic having one or more timers, each configured to measure a respective dwell time that commences at the time the living body is classified in the posture state; and wherein upon expiration of each of the one or more timers, the response logic initiates a respective response as a result of the body being classified in the posture state only if the living body has not exited the posture state during the respective dwell time. 5. The system of claim 2, wherein the posture state detection logic comprises dwell time logic configured to automatically select a dwell time as the predetermined period of time, selection being based on the posture state in which the living body is classified. 6. The system of claim 5, wherein the dwell time logic is further configured to select the dwell time based on a posture state in which the living body was previously classified. 7. The system of claim 2, wherein the posture state detection logic comprises dwell time logic configured to automatically... | MEDTRONIC, INC. | A61N001/00 A61N001/365 A61B005/11 A61B005/00
US20100016763 | A1 | 2010-01-21 | INTRALUMINAL FLUID PROPERTY STATUS SENSING SYSTEM AND METHOD | 2009-07-17 | Implementations of an intraluminal fluid property status sensing system and method locate an acoustic transducer within a lumen of a biological creature to transmit ultrasound through intraluminal fluid to be reflected or otherwise affected by the fluid with subsequent reception by the same transducer. Reflection or interaction of the ultrasound with an intraluminal fluid depends upon one or more properties of the intraluminal fluid so can be used to determine status of such properties. | 1. A system comprising: for positioning within a biological lumen containing a fluid, a sensing component including an oscillator and a transducer electrically coupled with the oscillator to receive electrical signaling from the oscillator, the electrical signaling to an original acoustic signal and transmit the original acoustic signal into the fluid, the transducer configured to receive and convert a first acoustic signal into a first electrical signal from the original acoustic signal being attenuated and reflected off of scatterers found in the fluid a first distance range from the transducer, the transducer configured to receive and convert a second acoustic signal into a second electrical signal from the original acoustic signal being attenuated and reflected off of scatterers found in the fluid a second distance range from the transducer, to a second electrical signal, the sensing component configured to transmit the first electrical signal as a first external signal and to transmit the second electrical signal as a second external signal; andan external component communicatively linked to the sensing component to receive the first electrical signal and the second electrical signal, the external component including a processor to output fluid property status based upon attenuation of the original acoustic signal as indicated by the first external signal and the second external signal received by the external component. 2. The system of claim 1 wherein the sensing component is communicatively linked to the external component through a wired connection. 3. The system of claim 1 wherein the sensing component is communicatively linked to the external component through a wireless connection. 4. The system of claim 3 wherein the wireless connection is one of the following: an RF magnetic link and an acoustic link. 5. The system of claim 1 wherein the sensing component is configured to be positioned within the lumen by one of the following: insertion and implantation. 6. The system of claim 1 wherein the processor is configured to determine concentration of biological cells within the fluid as included with the fluid property status. 7. The system of claim 1 wherein the processor is configured to determine fluid property status based upon attenuation values of pulse-echo events. 8. The system of claim 7 wherein the processor is configured to determine attenuation values of the pulse-echo events through the fluid as blood to derive the fluid property status to include a hematocrit value. 9. The system of claim 7 wherein the processor is configured to average values of a series of successive pulse-echo events to obtain amplitude data to calculate attenuation of the original acoustic signal in the fluid. 10. The system of claim 7 wherein the processor is configured to process amplitudes at two or more ranges to average values of the series of successive pulse-echo events. 11. A system comprising: for positioning within a biological lumen o... | - | A61N001/00 A61B005/00
US20100036214 | A1 | 2010-02-11 | APPARATUS, METHOD AND SYSTEM FOR DETERMINING A PHYSIOLOGICAL CONDITION WITHIN A MAMMAL | 2009-08-05 | A system, method and apparatus for determining a physiological condition within a mammal based on pH and/or temperature measurements. In one aspect, the invention is directed to a system for measuring a physiological condition within a mammal comprising: an apparatus comprising an elongated housing extending along a longitudinal axis; a pH sensor for measuring pH within a body lumen of the mammal, the pH sensor located at a distal portion of the elongated housing; and a removable cap for enclosing the distal portion of the elongate housing, the cap comprising a base having a first channel in a top surface of the base; a lid having a second channel in a bottom surface of the lid; and wherein the lid is pivotably mounted to the base so as to be pivotable between: (i) an open position wherein the distal portion of the elongated housing can be laid into the first channel by movement in a direction substantially perpendicular to the longitudinal axis; and (ii) a closed position wherein the first and second channels form a heretically sealed cavity about the distal portion of the elongated housing. | 1. A system for measuring a physiological condition within a mammal, comprising: an apparatus comprising: an elongated housing extending along a longitudinal axis from a proximal end to a distal end;a pH sensor for measuring pH within a body lumen of the mammal, the pH sensor located at a distal portion of the elongated housing;a temperature sensor for measuring temperature within the body lumen, the temperature sensor located at the distal portion of the elongated housing; anda handle having a second housing for manipulating the elongated housing, the second housing having a top surface, a bottom surface and a plurality of lateral surfaces bounding the top and bottom surfaces, the elongated housing protruding from one of the lateral surfaces of the second housing;a removable cap for enclosing the distal portion of the elongate housing, the cap comprising: a base having a first channel in a top surface of the base;a lid having a second channel, in a bottom surface of the lid; andwherein when the removable cap is in a closed position, the lid is positioned atop the base so that the bottom surface of the lid opposes the top surface of the base and the first and second channels form a cavity having an open end and a closed end; andwherein when the removable cap is secured to the elongated housing in the closed position, the distal portion of the elongate housing nests within the cavity and a remaining portion of the elongated housing protrudes from the open end of the cavity. 2. The system of claim 1 wherein the lid is pivotably mounted to the base so as to be pivotable between: (1) an open position wherein the first channel is exposed; and (2) the closed position. 3. The system of claim 2 wherein the lid is pivotably mounted to the base by a hinge. 4. The system of claim 1 wherein the removable cap is constructed of a material that prevents light from entering the cavity when the removable cap is in the closed position and secured to the elongated housing. 5. The system of claim 4 wherein the removable cap is constructed of an opaque plastic. 6. The system of claim 1 wherein a bottom edge of the elongate housing is non-coplanar with the bottom surface of the second housing, and wherein a bottom surface of the base is coplanar with the bottom surface of the second housing when the removable cap is secured to the elongated housing. 7. The system of claim 6 wherein the base comprises a plurality of legs, the legs forming the bottom surface of the base. 8. The system of claim 1 further comprising an annular rim located about the open end of the cavity that provides a hermetic seal between the elongate housing and the removable cap when the removable cap is secured to the elongated housing in the closed position, a first portion of the annular rim located in the first channel and a second portion of the annular rim located in the second channel. 9. The system of claim 8 wherein the annular rim is constructed of a compressible material. 10. The ... | - | A61B005/01 A61B005/145
US20100022909 | A1 | 2010-01-28 | METHODS AND DEVICES FOR MEASURING CORE BODY TEMPERATURE | 2008-02-15 | A core body temperature measurement device includes temperature sensor (12, 12f, 12b), a head-mountable mechanical frame or pad (42, 62, 64, 72, 82, 102, 112) configured to operatively couple the temperature sensor with skin overlaying an arterial blood rich superficial region (STA, PAA) disposed near to an auricle and outside of an ear canal or other skin (CAR) overlaying the carotid artery or a major arterial branch thereof, and a readout controller (10, 48, 68, 90) configured to acquire a temperature measurement using the temperature sensor and to output a core body temperature based on the acquired temperature measurement. | 1. A core body temperature measurement device comprising: a temperature sensor;a head-mountable mechanical frame or pad configured to operatively couple the temperature sensor with skin overlaying an arterial blood-rich superficial region disposed near to an auricle and outside of an ear canal; anda readout controller configured to acquire a temperature measurement using the temperature sensor and to output a core body temperature based on the acquired temperature measurement. 2. The core body temperature measurement device as set forth in claim 1, wherein the temperature sensor comprises a plurality of temperature sensors each configured to operatively couple with a different portion of skin overlaying an arterial blood-rich superficial region disposed near to the auricle and outside of the ear canal, and the readout controller is configured to acquire temperature measurements using the plurality of temperature sensors and to output a core body temperature based on the temperature measurements. 3. The core body temperature measurement device as set forth in claim 1, wherein the head-mountable mechanical frame comprises: an eyeglasses frame having left and right bends generally passing over respective left and right auricles and each terminating in an earpiece, the temperature sensor being disposed at or near at least one of: the left bend and operatively coupling with skin overlaying a portion of a superficial temporal artery anterior to the left auricle;the right bend and operatively coupling with skin (STA) overlaying a portion of a superficial temporal artery anterior to the right auricle;the left earpiece and operatively coupling with skin overlaying a portion of an artery ascending posterior to the left auricle, andthe right earpiece and operatively coupling with skin overlaying a portion of an artery ascending posterior to the right auricle. 4. The core body temperature measurement device as set forth in claim 3, wherein the head-mountable mechanical frame further comprises: a support for the temperature sensor; anda spring bias coupling the support to the eyeglasses frame, the spring bias pressing the temperature sensor against skin overlaying an arterial blood-rich superficial region disposed near to the auricle and outside of the ear canal. 5. The core body temperature measurement device as set forth in claim 1, wherein the head-mountable mechanical frame comprises: a behind-the-head pillow having extensions configured to loop over left and right auricles, the temperature sensor being disposed on one or both extensions in thermal communication with skin overlaying an arterial blood-rich superficial region disposed near to the auricle around which the extension loops. 6. The core body temperature measurement device as set forth in claim 5, wherein the readout controller is disposed on or in the pillow. 7. The core body temperature measurement device as set forth in claim 1, wherein the head-mountable mechanical frame comprises: a h... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/01
US07641615 | B2 | 2010-01-05 | Method and device for removing oscillatory artefacts from invasive blood pressure measurement data | 2006-09-29 | The invention relates to a method for removing oscillatory artefacts caused by pressure waves in a fluid-filled trans-luminal catheter from invasive blood pressure measurement data, including the steps of setting up a model for the measurement data, fitting the model to the measurement data, identifying the model components, representing the oscillatory artefact, and removing the oscillatory artefact from the measurement data. The invention also relates to a respective device for acquiring and correcting invasive blood pressure measurement data. | 1. A method for removing an oscillatory artefact caused by a pressure wave in a fluid-filled trans-luminal catheter from an invasive blood pressure measurement data of a patient, comprising: generating a model for the measurement data by a data processing module;transforming the measurement data and the model of the measurement data to a Z-transformed domain by the data processing module;fitting the model to the measurement data in the Z-transformed domain by the data processing module;identifying a model component representing the oscillatory artefact by the data processing module; andremoving the oscillatory artefact from the measurement data by the data processing module,wherein the model is a high-order auto-regression model that generates the measurement data when provided with a white noise, andwherein the model is a 9th order auto-regression model. 2. The method as claimed in claim 1, wherein a damping coefficient and a natural frequency of the oscillatory artefact are estimated. 3. The method as claimed in claim 2, wherein the damping coefficient and the natural frequency of the oscillatory artefact are estimated from a measured step response of the catheter. 4. The method as claimed in claim 2, wherein the model component representing the oscillatory artefact is identified based on the damping coefficient and the natural frequency of the oscillatory artefact. 5. The method as claimed in claim 1, wherein the measurement data comprises a series of points sampled at a predetermined sampling rate. 6. The method as claimed in claim 1, wherein a catheter pole originating from the oscillatory artefact is: identified in the Z-transformed domain of the model,separated from a heart-related pole, andused to remove the oscillatory artefact from the measurement data. 7. The method as claimed in claim 6, wherein a first approximation of a position of the catheter pole identified in the Z-transformed domain is calculated based on an estimated damping coefficient and natural frequency of the oscillatory artefact. | Siemens Aktiengesellschaft | A61B005/02
US07653426 | B2 | 2010-01-26 | Endoscopic image pickup method and magnetic resonance imaging device using the same | 2002-07-12 | An endoscope-like image taking method includes providing at least one peculiar index, which can be discriminated from other portions on an MR image, at the tip of a catheter, previously inserting a metal guide wire for guiding the catheter into a body cavity of a patient inserting the catheter into the body cavity along the guide wire, executing an MR imaging sequence of a plurality of sliced images intersecting the guide wire, reconstructing three-dimensional image data based upon the nuclear magnetic resonance signals, which are received by the guide wire, and determining the tip position and the inserting direction of the catheter by detecting the peculiar index provided at the tip of the catheter based upon the three-dimensional image data, and reconstructing the center projected image using the three-dimensional image data and setting the tip position and the inserting direction of the catheter as a view point and a line-of-sight direction and displaying the center projected image on a display means. | 1. A method in a magnetic resonance (MR) imaging apparatus, said method comprising: providing at least one peculiar index, which can be discriminated from other portions on art MR image, at the tip of a catheter;previously inserting a metal guide wire for guiding the catheter into a body cavity of a patient into which the catheter is inserted;executing, by said magnetic resonance imaging apparatus, an MR imaging sequence of a plurality of sliced images intersecting the guide wire while inserting the catheter into the body cavity along the guide wire;receiving the nuclear magnetic resonance signals, which are generated from the patient when the sequence is executed, by using the guide wire as antenna means to receive said nuclear magnetic resonance signals;reconstructing, by said magnetic resonance imaging apparatus, three-dimensional image data using the nuclear magnetic resonance signals;determining the tip position and the insetting direction of the catheter by detecting the peculiar index provided at the tip of the catheter based on any one of the three-dimensional image data and image data imaged by an MR imaging method different from that of the three-dimensional image data;setting the tip position and the inserting direction of the catheter as a view point and a line-of-sight direction, respectively;acquiring, by said magnetic resonance imaging apparatus, three-dimensional image data at a plurality of sliced positions set forward of the tip of the catheter; andreconstructing, by said magnetic resonance imaging apparatus, a center projected image being observed from the view point by projecting the three-dimensional image data acquired at the sliced positions set forward of the tip of the catheter along a line-of-sight direction to a projection plane and displaying the center projected image on a display means. 2. A method in a magnetic resonance (MR) imaging apparatus, said method comprising: providing at least one peculiar index, which can be discriminated from other portions on an MR image, at the tip of a catheter;previously inserting it metal guide wire for guiding the catheter into a body cavity of a patient into which the catheter is inserted;inserting the catheter into the body cavity along the guide wire;executing, by said magnetic resonance imaging apparatus, an MR imaging sequence of a plurality of sliced images intersecting the guide wire;reconstructing, by said magnetic resonance imaging apparatus, three-dimensional image data based upon the nuclear magnetic resonance signals, which are generated from the patient when the sequence is executed and are received by using the guide wire as antenna means to receive said nuclear magnetic resonance signals, and determining the tip position and the inserting direction of the catheter by detecting the peculiar index provided at the tip of the catheter based on the three-dimensional image data;setting the tip position and the inserting direction of the catheter as a view point and a line... | Hitachi Medical Corporation | A61B005/05
US07663365 | B2 | 2010-02-16 | Magnetic resonance imaging apparatus and analysis method for fat suppression effect in magnetic resonance imaging apparatus | 2008-09-24 | A magnetic resonance imaging apparatus includes an imaging unit which performs imaging more than once with respect to an imaging target while changing a central frequency of a fat suppression pulse, a generation unit which generates a plurality of images based on magnetic resonance signals obtained by imaging performed more than once, and a calculation unit which calculates factor information of spatial inhomogeneity of a fat suppression effect based on the plurality of images. | 1. A magnetic resonance imaging apparatus (MRI) comprising: an imaging unit which performs MR imaging more than once with respect to an imaging target while changing a central frequency of a fat suppression pulse sequence utilized in an MRI scan each time imaging is performed;a generation unit which generates a plurality of reconstructed MR images based on magnetic resonance signals obtained by MR imaging being performed more than once; anda calculation unit which calculates factor information, which contributes to spatial inhomogeneity of a fat suppression effect based on the plurality of reconstructed MR images, and storing or outputting said factor information. 2. The magnetic resonance imaging apparatus according to claim 1, wherein the calculation unit obtains information related to inhomogeneity of a static magnetic field and information indicative of the fat suppression effect, which has no relation to a frequency shift of a fat suppression spectrum as the factor information contributing to the spatial inhomogeneity of the fat suppression effect. 3. The magnetic resonance imaging apparatus according to claim 2, wherein the calculation unit calculates a shift amount between two frequency spectrums obtained at a magnetic field central position and a position other than the magnetic field central position based on the plurality of MR images as the information related to inhomogeneity of the static magnetic field. 4. The magnetic resonance imaging apparatus according to claim 1, wherein the MR imaging unit performs MR imaging more than once while changing a polarity of at least one of a slice selective gradient magnetic field, a frequency read-out gradient magnetic field, and a phase encoding gradient magnetic field, and the calculation unit calculates at least one of information related to eddy magnetic field and information related to inhomogeneity of a static magnetic field based on a plurality of MR images obtained by the gradient magnetic field having a changed polarity as the factor information which contributes to the spatial in homogeneity of the fat suppression effect. 5. The magnetic resonance imaging apparatus according to claim 1, wherein the calculation unit calculates information related to inhomogeneity of a radio-frequency pulse as the factor information which contributes to the spatial inhomogeneity of the fat suppression effect. 6. The magnetic resonance imaging apparatus according to claim 1, further comprising a display unit which generates and displays an image representing spatial inhomogeneity of the fat suppression effect in a contour pattern as the factor information which contributes to the spatial in homogeneity of the fat suppression effect. 7. The magnetic resonance imaging apparatus according to claim 1, further comprising a correction unit which corrects at least one of a central frequency of a fat suppression pulse at the time of actual imaging and a gradient magnetic field pulse at the time of application... | Kabushiki Kaisha Toshiba Toshiba Medical Systems Corporation | G01V003/00 A61B005/055
US07657295 | B2 | 2010-02-02 | Medical sensor and technique for using the same | 2005-08-08 | A unitary sensor is provided that includes a frame upon which electrical and optical components may be disposed and a coating, such as an overmolded coating provided about the frame. The unitary sensor includes a spring, such as a flat spring or torsion spring, provided on the frame that provides a closing force for the sensor. The sensor may be placed on a patient's finger, toe, ear, and so forth to obtain pulse oximetry or other physiological measurements. | 1. A sensor assembly, comprising: a frame;an emitter disposed on the frame;a detector disposed on the frame;a biasing component attached to an end of the frame; anda coating molded over at least the frame to form a sensor assembly, wherein the coating provides a biasing force to supplement the biasing component. 2. The sensor assembly of claim 1, wherein the sensor assembly comprises at least one of a pulse oximetry sensor, a sensor for measuring a water fraction, or a combination thereof. 3. The sensor assembly of claim 1, comprising one or more signal transmission structures connected to at least one of the emitter or the detector. 4. The sensor assembly of claim 1, wherein the biasing component comprises at least one of a formed flat spring or a torsion spring. 5. The sensor assembly of claim 1, wherein the frame comprises a single molded part. 6. The sensor assembly of claim 1, wherein the frame comprises two or more parts joined to form the frame. 7. The sensor assembly of claim 1, wherein the biasing component comprises a polymeric material. 8. The sensor assembly of claim 1, wherein the coating comprises a thermoplastic elastomer. 9. The sensor assembly of claim 8, wherein the thermoplastic elastomer comprises at least one of a thermoplastic polyolefin, a thermoplastic vulcanizate alloy, thermoplastic polyurethane, silicone, or a combination thereof. 10. The sensor assembly of claim 1, wherein the coating comprises a conformable material. 11. The sensor assembly of claim 1, wherein no moving parts of the frame remain exposed from the coating. 12. The sensor assembly of claim 1, wherein the sensor assembly is fluid tight. 13. The sensor assembly of claim 1, wherein the coating is chemically bonded to the frame. 14. The sensor assembly of claim 1, comprising a pulse oximeter monitor configured to receive data from the sensor assembly. 15. The sensor assembly of claim 1, comprising a multi-parameter monitor configured to receive data from the sensor assembly. 16. The sensor assembly of claim 1, comprising at least one integrated circuit device. 17. The sensor assembly of claim 1, comprising a cable comprising one or more integrated circuits. 18. The sensor assembly of claim 1, comprising an adapter cable connected to one or more signal transmission structures of the sensor assembly. 19. The sensor assembly of claim 1, wherein the emitter comprises one or more light emitting diodes. 20. The sensor assembly of claim 1, wherein the detector comprises a photodetector. 21. A sensor assembly, comprising: a frame;an emitter disposed on the frame;a detector disposed on the frame; anda coating provided over at least the frame to form a sensor assembly, wherein the coating provides a closing force that acts to close one end of the sensor assembly in the absence of an opposing force. 22. The sensor assembly of claim 21, wherein the sensor assembly comprises at least one of a pulse oximetry sensor, a sensor for measuring a water frac... | Nellcor Puritan Bennett LLC | A61B005/145
US07653424 | B2 | 2010-01-26 | Apparatus for measuring blood glucose concentrations | 2006-03-16 | An apparatus for measuring ocular and/or blood glucose levels comprises (a) an irradiating means (10) for irradiating light onto the eye (1) of a user from outside the cornea of the eye to excite an ocular glucose sensor in contact with an ocular fluid, said sensor being able to emit a total fluorescence having first and a second wavelength bands;(b) an optical path splitting means (11) for splitting said total fluorescence into a first fluorescence and a second fluorescence, said first fluorescence and said second fluorescence traveling along first and second optical paths;(c) a first detecting means (14) located in the first optical path;(d) a second detecting means (17) located in the second optical path;(e) a calculating means for calculating the intensity ratio of the first fluorescence to the second fluorescence and for determining an ocular glucose concentration in the ocular fluid; and(f) an arithmetic means for converting the ocular glucose concentration into a blood glucose concentration. | 1. An apparatus for measuring blood glucose levels, comprising: (a) an irradiating means for irradiating light onto the eye of a user from outside the cornea of the eye to excite an ocular glucose sensor, wherein said ocular glucose sensor is in contact with an ocular fluid and can emit a total fluorescence having a first and a second wavelength bands upon irradiation with said irradiating means;(b) an optical path splitting means for splitting said total fluorescence having both bands into a first fluorescence having said first wavelength band and second fluorescence having said second wavelength band, wherein said first fluorescence travels along a first optical path and said second fluorescence travels along a second optical path;(c) a first detecting means located in the first optical path for detecting the intensity of the first fluorescence at a first wavelength;(d) a second detecting means located in the second optical path for detecting the intensity of the second fluorescence at a second wavelength;(e) a calculating means for calculating the intensity ratio of the first fluorescence to the second fluorescence and for determining based on the calculated intensity ratio an ocular glucose concentration in the ocular fluid of the user according to a predetermined calibration table or calibration curve; and(f) an arithmetic means for converting the ocular glucose concentration determined by the calculating means into a blood glucose concentration by referring to a predetermined correlation between blood glucose concentrations and ocular glucose concentrations. 2. An apparatus of claim 1, further comprising a displaying means for visually displaying the blood glucose concentration determined by the arithmetic means. 3. An apparatus of claim 1, further comprising a transmitting means for transmitting a signal to a pump, wherein said signal contains instructions for the pump to inject an amount of insulin into a tissue of the user. 4. An apparatus according to claim 1, wherein said irradiating means emits a light of third and fourth wavelengths which are different. 5. An apparatus of claim 2, further comprising a transmitting means for transmitting a signal to a pump, wherein said signal contains instructions for the pump to inject an amount of insulin into a tissue of the user. 6. An apparatus according to claim 2, wherein said irradiating means emits a light of third and fourth wavelengths which are different. 7. An apparatus according to claim 3, wherein said irradiating means emits a light of third and fourth wavelengths which are different. 8. A kit comprising: at least three solutions of known glucose concentrations which are different, and an apparatus, wherein the apparatus comprises (a) an irradiating means for irradiating light onto the eye of a user from outside the cornea of the eye to excite an ocular glucose sensor, wherein said ocular glucose sensor is in contact with an ocular fluid and can emit a total fluorescence having... | Eyesense AG | A61B005/1455
US07666149 | B2 | 2010-02-23 | Cassette of lancet cartridges for sampling blood | 2002-10-28 | A cassette containing cartridges for sampling blood from a patient. The cassette includes a container for storing a plurality of cartridges and at least one cartridge in the container. The cartridge includes a cartridge case and a lancet. The lancet has a tip and is housed in the cartridge case. The lancet can be driven to extend the tip outside the cartridge case for lancing the skin of the patient to yield blood. The container has a compartment that contains at least one cartridge. A cartridge from the compartment can be loaded onto a glucometer that drives the lancet in the cartridge to lance the skin of a patient. | 1. An apparatus for sampling blood from a patient, comprising: a support member;a plurality of penetrating members housed in a radial pattern on said support member, each of said penetrating members having a tip, wherein the penetrating member can be driven to extend outwardly to penetrate tissue of a patient to yield blood for collection for analysis;at least one test region receiving said blood for analysis; andwherein each of said penetrating members and said at least one test region is housed in a cylindrical case, said case being coupled to the support member and configured to be coupled to a penetrating member driver. 2. An apparatus for sampling blood from a patient, comprising: a support member;a plurality of penetrating members housed in a radial pattern on said support member, each of said penetrating members having a tip, wherein the penetrating member can be driven to extend outwardly to penetrate tissue of a patient to yield blood for collection for analysis;at least one test region receiving said blood for analysis; andwherein the support member includes a plurality of slots for holding a plurality of cases, each case containing at least one penetrating member and one test region. 3. An apparatus for sampling blood from a patient, comprising: a support member;a plurality of penetrating members housed in a radial pattern on said support member, each of said penetrating members having a tip, wherein the penetrating member can be driven to extend outwardly to penetrate tissue of a patient to yield blood for collection for analysis:at least one test region receiving said blood for analysis; anda calibration cartridge having a function of calibration or identification of the analytical characteristics of the test region. | Peliken Technologies, Inc. | A61B005/00 B65D081/00
US20100036360 | A1 | 2010-02-11 | STENT GRAFT DELIVERY SYSTEM | 2009-04-24 | A system for treating an aneurysm comprises an elongate flexible shaft and an expandable member. An expandable scaffold is disposed over the expandable member and may be expanded from a collapsed configuration to an expanded configuration. A double-walled filling structure is disposed over the scaffold and has an outer wall and an inner wall. The filling structure is adapted to be filled with a hardenable fluid filing medium so that the outer wall conforms to an inside surface of the aneurysm and the inner wall forms a substantially tubular lumen to provide a path for blood flow. In the expanded configuration the scaffold engages the inner wall of the filling structure. A tether is releasably coupled with the filling structure and the flexible shaft thereby constraining axial movement of the structures relative to each other. | 1. A system for treating an aneurysm in a blood vessel, the system comprising: an elongate flexible shaft having a proximal region and a distal region;a first double-walled filling structure disposed over the distal region of the shaft, the filling structure having an outer wall and an inner wall, wherein the filling structure is adapted to be filled with a hardenable fluid filing medium so that the outer wall conforms to an inside surface of the aneurysm and the inner wall forms a first substantially tubular lumen to provide a path for blood flow; andat least a first expandable scaffold disposed adjacent the filling structure, the first scaffold radially expandable within at least a portion of the tubular lumen of the filling structure,wherein the filling structure is separate from the first scaffold and axially separated therefrom. 2. The system according to claim 1, wherein the first scaffold is proximal to the filling structure. 3. The system according to claim 1, wherein the first scaffold is distal to the filling structure. 4. The system according to claim 1, wherein a spacing exists between one end of the first scaffold and one end of the filling structure. 5. The system according to claim 1, wherein the first scaffold is slidably received by the filling structure such that the first scaffold and the filling structure are concentric with one another. 6. The system according to claim 1, further comprising a sheath disposed at least partially over the filling structure. 7. The system according to claim 1, further comprising a sheath disposed at least partially over the scaffold. 8. The system according to claim 1, further comprising a tapered sheath disposed at least partially over the flexible shaft. 9. The system according to claim 8, wherein the tapered sheath comprises axially oriented slits near a distal end thereof. 10. The system according to claim 1, further comprising a pusher tube disposed at least partially over the flexible shaft and slidably engageable with the first double-walled filling structure. 11. The system according to claim 1, further comprising a first tether coupled with the filling structure, the tether adapted to guide movement of the first double-walled filling structure relative to the first scaffold axially along the shaft. 12. The system according to claim 11, wherein the first tether extends between the proximal and distal regions of the flexible shaft. 13. The system according to claim 11, wherein the filling structure is slidably engaged with the first tether. 14. The system according to claim 11, further comprising a second tether coupled with the filling structure. 15. The system according to claim 14, wherein the second tether extends between the proximal and distal regions of the flexible shaft. 16. The system according to claim 11, wherein the filling structure comprises one or more eyelets. 17. The system according to claim 11, further comprising a nosecone coupled with the distal region... | Nellix, Inc. | A61M031/00 A61M029/00 A61B005/00
US08378811 | B2 | 2013-02-19 | Mobile wireless customizable health and condition monitor | 2009-09-28 | A wearable sensor may be provided. The wearable sensor may include a timing mechanism; a plurality of sensor pairs coupled to the timing mechanism wherein each sensor pair comprises an emitter and a detector; and a microcontroller coupled in communication to the timing mechanism and the plurality of sensor pairs. | 1. A wearable sensor comprising: a timing mechanism;a plurality of sensor pairs coupled to said timing mechanism wherein each sensor pair comprises an emitter and a detector;a microcontroller coupled in communication to said timing mechanism and said plurality of sensor pairs; anda band that facilitates coupling said plurality of sensor pairs to a user,wherein said plurality of sensor pairs are positioned in an array on at least one of said band and said timing mechanism, andwherein said microcontroller is adapted to determine which of said plurality of sensor pairs detect an acceptable signal from the user. 2. A wearable sensor in accordance with claim 1 further comprising an accelerometer coupled to said microcontroller. 3. A wearable sensor in accordance with claim 1 wherein said timing mechanism comprises a watch. 4. A wearable sensor in accordance with claim 1 wherein said emitter comprises a light emitting device. 5. A wearable sensor in accordance with claim 1 wherein said detector comprises a photodiode. 6. A monitoring system, comprising: at least one wearable monitor coupled to a user, said at least one wearable monitor comprising:a timing mechanism;a plurality of sensor pairs coupled to said timing mechanism wherein each sensor pair comprises an emitter and a detector; anda microcontroller coupled in communication to said timing mechanism and said plurality of sensor pairs;at least one gateway having a second transceiver;at least one device located at a remote location; andat least one server coupled in communication using a wide area network to the at least one gateway and the at least one device, said at least one wearable monitor is coupled in wireless communication to said at least one gateway, said plurality of sensor pairs facilitate sensing a heart rate of the user and transmit the heart rate to said at least one server; anda band that facilitates coupling said plurality of sensor pairs to a user,wherein said plurality of sensor pairs are positioned in an array on at least one of said band and said timing mechanism, andwherein said microcontroller is adapted to determine which of said plurality of sensor pairs detect an acceptable signal from the user. 7. A system in accordance with claim 6 further comprising an accelerometer coupled to said microcontroller, said accelerometer facilitates detecting movement of said wearable sensor. 8. A system in accordance with claim 6 wherein said timing mechanism comprises a watch. 9. A system in accordance with claim 6 wherein said emitter comprises a light emitting device. 10. A system in accordance with claim 6 wherein said detector comprises a photodiode. 11. A method of sensing a heart rate using a wearable monitor, said wearable monitor comprising: a timing mechanism;a memory;a plurality of sensor pairs coupled to the timing mechanism wherein each sensor pair comprises an emitter and a detector;a microcontroller coupled in communication to the timing mechanism, the memory and ... | Aframe Digital, Inc. | G08B001/08 G08B023/00 A61B005/00
US08162857 | B2 | 2012-04-24 | Limb movement monitoring system | 2007-10-08 | A limb movement monitoring system, comprising a motion sensor, a garment comprising spatially addressable photonic textiles and spatially resolving pressure-sensitive textiles and furthermore a muscular activity sensor, a processing unit processing the data from the sensors and issuing illumination commands to the spatially addressable photonic textiles and a database. There is furthermore a process for monitoring limb movement by the system. | 1. Limb movement monitoring system, comprising: a motion sensor (1),a garment comprising spatially addressable photonic textiles (2) and spatially resolving pressure-sensitive textiles (3),a processing unit (5) which receives data from the motion sensor (1) and from the spatially resolving pressure-sensitive textiles (3) and which issues illumination commands to the spatially addressable photonic textiles (2),a database (6) in communication with the processing unit (5), whereinthe database (6) stores data from the motion sensor (1) resolved into individual spatial descriptors and time,the database (6) stores data from the pressure-sensitive textiles (3) resolved into at least the categories of x and y coordinates and pressure,the database (6) furthermore comprises comparison motion data resolved into individual spatial descriptors and time,the processing unit (5) compares the motion data from the motion sensor to the comparison motion data in the database (6), andthe processing unit (5) issues illumination commands to the spatially addressable photonic textiles (2) based upon the deviation of the motion data from the comparison motion data and the deviation of the pressure data from a preset threshold. 2. System according to claim 1, further comprising a muscular activity sensor (4), wherein the processing unit (5) also receives data from the muscular activity sensor (4), the database (6) also stores data from the muscular activity sensor (4) and the processing unit (5) also issues illumination commands to the spatially addressable photonic textiles (2), based upon the deviation of the muscular activity data from a preset threshold. 3. System according to claim 1, further comprising a user interface (7) for accessing data from the database. 4. System according to claim 1, further comprising an external feedback system (8) communicating motion and/or pressure data. 5. System according to claim 1, wherein the spatially resolving pressure-sensitive textiles (3) comprise piezoresistive fabrics. 6. System according to claim 1, wherein the signals from the at least one motion sensor (1), from the pressure-sensitive textiles (3) and/or the commands to the spatially addressable photonic textiles (2) are transmitted wirelessly. 7. System according to claim 1, further comprising a glove comprising spatially resolving pressure-sensitive textiles, wherein the glove transmits pressure data to the processing unit (5). 8. System according to claim 1, further comprising sensors selected from the group comprising skin perspiration sensors, pulse sensors, blood pressure sensors and/or blood oxygen level sensors, wherein the sensors further provide input to the processing unit (5). 9. Process for monitoring limb movement by a system according to claim 1, comprising the steps of: a) calibrating the at least one motion sensor (1) by bringing the at least one motion sensor (1) into a predetermined position and registering the output signal from the at least on... | Koninklijke Philips Electronics N.V. | A61B005/103 A61B005/117
US07925323 | B2 | 2011-04-12 | Radial electrode array | 2009-10-01 | A sensor array apparatus for monitoring medical signals includes a first flexible substrate defining a central focal point and a second flexible substrate associated with the central focal point. A plurality of medical electrodes are disposed on the periphery of the first flexible substrate and a reference electrode is disposed on the second flexible substrate. A connector is in electrical communication with the medical electrodes and the reference electrode and adapted to connect to an electronic system. Bio-electric information is monitored between one of the medical electrodes and the reference electrode. | 1. A sensor array apparatus for monitoring medical signals, which comprises: a first flexible substrate defining a central focal point;a plurality of medical electrodes disposed on the periphery of the first flexible substrate;a second flexible substrate associated with the central focal point;a reference electrode disposed on the second flexible substrate;at least one connector in electrical communication with the medical electrodes and the reference electrode, the at least one connector adapted to connect to an electronic system,wherein bio-electric information is monitored between at least one of the medical electrodes and the reference electrode. 2. The sensor array apparatus according to claim 1 wherein the first flexible substrate generally define a shape selected from a group consisting of a circle, triangle, square, rectangle, polygon, or oval. 3. The sensor array apparatus according to claim 1 wherein the medical electrodes are disposed at substantially equal predetermined radial distances with respect to the central focal point. 4. The sensor array apparatus according to claim 1 wherein the medical electrodes are generally symmetrically arranged on the first flexible substrate about an axis of symmetry whereby a medical electrode on one side of the axis of symmetry has a corresponding medical electrode on the opposite side of the axis of symmetry. 5. The sensor array apparatus according to claim 1 wherein the reference electrode defines at least one reference aperture formed in the second flexible substrate wherein the at least one reference aperture is dimensioned to partially encapsulate a body tissue. 6. The sensor array apparatus according to claim 5 wherein the reference electrode extends around the inner perimeter of the at least one reference aperture. 7. The sensor array apparatus according to claim 6 wherein the plurality of medical electrodes are each spaced at substantially equal distances with respect to the reference electrode. 8. The sensor array apparatus according to claim 1 wherein the first flexible substrate and the second flexible substrate are in mechanical communication. 9. The sensor array apparatus according to claim 8 wherein the first flexible substrate and the second flexible substrate are in a substantially fixed relationship with respect to each other. 10. The sensor array apparatus according to claim 9 wherein the reference electrode extends around the inner perimeter of the reference aperture and the reference electrode and each of the plurality of medical electrodes are adapted to be along the skin of a patient during use, wherein each of the plurality of medical electrodes are a predetermined distance along the skin relative to the reference electrode, whereby the predetermined distance for each electrode is substantially the same. 11. The sensor array apparatus according to claim 1 wherein at least one of the plurality of medical electrodes further includes: at least one monopolar medical elec... | Tyco Healthcare Group LP | A61B005/04
US08082031 | B2 | 2011-12-20 | Neurofeedback system | 2009-08-04 | An apparatus and process for low energy neurofeedback includes computer and display means on a general purpose computer to calculate and display dominant brain wave frequencies as measured by an electroencephalogram. The general purpose computer is also operative to compute and display a new frequency that is offset from the dominant frequency by 5 to 20 Hz. A therapeutically beneficial low power RF field is generated in the EEG leads merely by the computation and display of the new frequency without the need to deliberately power the EEG leads. | 1. A method for low energy neurofeedback treatment, the method consisting essentially of the steps of: a) providing an EEG amplifier in signal communication with a patient,b) providing a general purpose computer in signal communication with the EEG amplifier to receive, analyze, and display EEG signals from the patient and to calculate a prospective treatment regimen,c) acquiring one or more analog EEG signals from the patient in the EEG amplifier,d) converting the one or more analog EEG signals to a digital format,e) determining the dominant frequency of the one or more analog EEG signals in the general purpose computer,f) calculating the prospective treatment regimen in the general purpose computer, said step of calculating further comprising:i) providing a frequency offset value,ii) calculating a first prospective treatment frequency by summing the dominant frequency and the frequency offset value,g) redetermining the dominant frequency each 1/16 of a second from the last 16 measurements and further recalculating the prospective treatment regimen for each redetermined dominant frequency. 2. The method for low energy neurofeedback treatment according to claim 1 further comprising a step of displaying the dominant frequency, the frequency offset, and the prospective treatment frequency on a monitor associated with the general purpose computer. 3. The method for low energy neurofeedback treatment according to claim 2 wherein the step of displaying the dominant frequency, the frequency offset, and the prospective treatment frequency on the monitor associated with the general purpose computer is repeated every 1/16 of a second for a time interval of at least one second. 4. The method of claim 3 wherein the calculated prospective treatment frequency is not feedback in an amplified form to the EEG leads. 5. The method of claim 2 wherein the calculated prospective treatment frequency is not feedback in an amplified form to the EEG leads. 6. The method of claim 1 wherein the calculated prospective treatment frequency is not feedback in an amplified form to a plurality of EEG leads. 7. The method of claim 1 wherein the calculated prospective treatment frequency is not feedback in an amplified form to a plurality of EEG leads or as a photic stimulation. 8. An apparatus for low energy neurofeedback treatment, the apparatus consisting essentially of: a) a plurality of EEG leads, wherein the leads are adapted to be attached to the patient,b) an EEG amplifier in signal communication with said EEG leads, the EEG amplifier consisting essentially of means to amplify signal received from the leads attached to the patient,c) a computer in signal communication with said EEG amplifier to acquire and display an EEG signal,d) at least one of the EEG amplifier and computer comprising computational means to determine the dominant frequency in each EEG lead pair,e) the computer further comprising:i) a user interface for providing a predetermined offset frequency,... | Ochslabs, Inc. | A61B005/04
US08958866 | B2 | 2015-02-17 | System and method to analyze blood parameters using magnetic resonance imaging | 2009-03-27 | A system and method for accurately producing MR images of selected vascular compartments includes employing a control scan and a tag scan, each including velocity selective modules that suppress signal from blood flowing faster than a given cutoff velocity, to acquire control and tag sets of NMR data that may be subtracted to produce a compartment-specific MR image that is substantially free of information from stationary tissues and blood outside the selective vascular compartments. Accordingly, physiological parameters, such as oxygen saturation (SaO2), oxygen extraction fraction (OEF), and cerebral metabolic rate of oxygen (CMRO2), can be generated from the compartment-specific images. Further still, kinetic curves of oxygen exchange can be created, thus providing detailed insight into oxygen exchange dynamics. | 1. A method of producing an image of a subject using an MRI system, the method comprising the steps of: a) directing the MRI system to perform a first pulse sequence that includes the following sequential steps: a1) applying a first velocity-selective pulse to suppress NMR signal from blood flowing faster than a first cutoff velocity through a plurality of vascular compartments of the subject;a2) providing an inflow time delay to allow blood to flow into selected vascular compartments of the plurality of vascular compartments;a3) applying a second velocity-selective pulse to suppress NMR signal from blood flowing faster than a second cutoff velocity through the plurality of vascular compartments;b) acquiring a tag set of NMR data from the subject after step a) and at a tag acquisition time after step a1);c) directing the MRI system to perform a second pulse sequence that includes the following sequential steps: c1) applying a third velocity-selective pulse to suppress NMR signal from blood flowing faster than a third cutoff velocity through the plurality of vascular compartments of the subject;c2) providing the inflow time delay to allow blood to flow into selected vascular compartments of the subject;c3) optionally applying a fourth velocity selective pulse to suppress NMR signal from blood flowing faster than a fourth cutoff velocity through the plurality of vascular compartments;d) acquiring a control set of NMR data from the subject after step c) and at a control acquisition time after step c1) that is equal to the tag acquisition time;e) reconstructing a tag image from the tag set of NMR data and a control image from the control set of NMR data; andf) subtracting the tag image from the control image to produce the image of the subject,wherein the third cutoff velocity and the fourth cutoff velocity are different velocities, andwherein the image of the subject indicates a characteristic of the blood flowing within the selected vascular compartments and is substantially free of information from both stationary tissue and blood outside the selected vascular compartments. 2. The method as recited in claim 1 wherein step a) includes applying an inversion pulse at a selected time during the inflow time delay to null recovering NMR signals. 3. The method as recited in claim 1 wherein the first velocity-selective pulse is spatially non-selective. 4. The method as recited in claim 1 further comprising the step of performing velocity bracketing to suppress NMR signal from blood having a velocity outside a selected velocity range. 5. The method as recited in claim 1 wherein step a3) includes eliminating signal from blood that has accelerated above the second cutoff velocity during the inflow time delay. 6. The method as recited in claim 1 wherein a physiological parameter is determined from the image that indicates a characteristics of the blood flowing within the selected vascular compartments. 7. The method as recited in claim 1 further includi... | The General Hospital Corporation | A61B005/05 G06K009/00 G01V003/00 A61B005/055 G01R033/56 G01R033/563 A61B005/145
US10238311 | B2 | 2019-03-26 | Devices, systems, and methods to evaluate cardiovascular function | 2008-01-22 | Devices and methods are disclosed which relate to the detection of cardiovascular efficiency and risk of disease. The rate of volumetric change of the heart can be determined by measuring the parallel conductance across electrodes attached to the heart. Measurements from a lumen would consider the total conductance. The rate of volumetric or lumen cross-section area change can then be compared to an average model to determine the health of a patient. | 1. A device for determining the index of a heart and/or vessel function, the device comprising: an impedance device comprising a first excitation electrode, a second excitation electrode, a first detection electrode, and a second detection electrode, the impedance device configured for percutaneous intravascular delivery and at least partial placement within a pericardial space and operable to obtain conductance data to determine at least one total conductance measurement when placed therein, the at least one total conductance measurement indicative of conductance of a heart wall, conductance of a ventricle adjacent to the heart wall, and conductance within the pericardial space, wherein the at least one total conductance measurement is obtained from within the pericardial space and without a determination of parallel conductance; anda conductance reader, the conductance reader in connection with the impedance device and operable to calculate the at least one total conductance measurement using the conductance data obtained from the first detection electrode and the second detection electrode and a max rate of change of a cross sectional area of the heart and/or vessel based on the calculated at least one total conductance measurement;whereby the max rate of change of the cross sectional area is sufficient to assess an index of a heart and/or vessel function without the determination of parallel conductance. 2. The device of claim 1, wherein the conductance reader comprises a data acquisition and processing system, the data acquisition and processing system comprising: a processor;a storage medium operably connected to the processor, the storage medium capable of receiving and storing the conductance data; anda program stored upon the storage medium, the program operable by the processor upon the conductance data to compare the conductance data to a rate of volumetric change of a heart and/or vessel. 3. The device of claim 2, wherein the processor is operable to compare the conductance data from conductance acquired from an epicardial surface of a heart. 4. The device of claim 2, wherein the processor is operable to compare the conductance data from conductance acquired from a lumen surface of a heart. 5. The device of claim 2, wherein the processor is operable to evaluate the maximum rate of volumetric change of the heart. 6. The device of claim 2, wherein the processor compares the conductance data from conductance acquired from the outer surface of a vessel. 7. The device of claim 2, wherein the processor compares the conductance data from conductance acquired from the lumen surface of a vessel. 8. The device of claim 2, wherein the device is configured to be positioned upon an epicardial surface of a vessel, and wherein the conductance reader is operable to detect conductance from the epicardial surface of the vessel. 9. The device of claim 2, wherein the device is configured to be positioned upon a lumen surface of a vessel, and wher... | CVDevices, LLC | A61B005/053
US20100030089 | A1 | 2010-02-04 | Methods and systems for monitoring and modifying a combination treatment | 2008-10-10 | Methods, computer program products, and systems are described that include monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent, identifying at least one characteristic of the artificial sensory experience that is proximate to a desired change in the at least one attribute of the individual, and/or modifying at least one of the artificial sensory experience or the bioactive agent at least partly based on the identifying at least one characteristic of the artificial sensory experience that is proximate to the desired change in the at least one attribute of the individual. | 1. A computer-implemented method, comprising: monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent;identifying at least one characteristic of the artificial sensory experience that is proximate to a desired change in the at least one attribute of the individual; andmodifying at least one of the artificial sensory experience or the bioactive agent at least partly based on the identifying at least one characteristic of the artificial sensory experience that is proximate to the desired change in the at least one attribute of the individual. 2-40. (canceled) 41. A system, comprising: means for monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent;means for identifying at least one characteristic of the artificial sensory experience that is proximate to a desired change in the at least one attribute of the individual; andmeans for modifying at least one of the artificial sensory experience or the bioactive agent at least partly based on the identifying at least one characteristic of the artificial sensory experience that is proximate to the desired change in the at least one attribute of the individual. 42. The system of claim 41, wherein means for monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent comprises: means for dispensing a bioactive agent to affect the at least one attribute of an individual. 43. The system of claim 41, wherein means for monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent comprises: means for receiving data from an automated medical device. 44. The system of claim 41, wherein means for monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent comprises: means for monitoring at least one of physical activity, body weight, body mass index number, heart rate, blood oxygen level, or blood pressure temporally associated with an artificial sensory experience. 45. The system of claim 41, wherein means for monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent comprises: means for monitoring a neurophysiological activity. 46. (canceled) 47. The system of claim 41, wherein means for monitoring at least one attribute of an individual during an artificial sensory experience while the individual is under the influence of a bioactive agent comprises: means for recording at least one monitored attribute of the individual. 48. The system of claim 41, wherein means for identifying at least one characteristic of the artificial senso... | Searete LLC, a limited liability corporation of the state of Delaware | A61B005/024 A61B005/00 A61M037/00
US20100041963 | A1 | 2010-02-18 | Cerebral and Cardiovascular Disease Prevention Using Optical -Magnetic Resonance Hybrid Imaging | 2008-08-15 | The invention provides a computer implemented method for calculating a chronic disease risk index for a patient undergoing a first examination, the method comprising providing a first result of the first examination wherein the first examination is selected from a group of medical imaging method; providing a second result of a second examination wherein the second examination being related to the examination of the patient's eye; Processing in combination the first result and the second such as to calculated a combined result wherein the combined result relates the first and the second result; Classifying the combined result; Calculating the risk index for the patient. The risk index is based on the combined result or on the classified combined result. | 1. A computer-implemented method for calculating a chronic disease risk index for a patient undergoing a first examination, the method comprising: providing a first result of the first examination, wherein the first examination is selected from a group of medical imaging methods;providing a second result of a second examination, the second examination being related to an examination of the patient's eye;processing the first result in combination with the second result such as to calculate a combined result, wherein the combined result relates to both the first and the second result;classifying the combined result;calculating the risk index for the patient, the risk index being based on the combined result or on the classified common result. 2. Computer-implemented method according to claim 1, wherein providing the first result or providing the second result comprise: carrying out an additional second examination concomitantly to the first examination. 3. Computer-implemented method according to claim 1, wherein the first examination and the second examination are performed in combination or are performed separately. 4. Computer-implemented method according to claim 1, wherein the first result or the second result comprise at least: detecting the first result and the second result of at least a previous examination;matching those results from the group of the first result and the second result which have the same detection times. 5. Computer-implemented method according to claim 1, wherein processing of the first result or processing the second result comprise at least: extracting relevant features from the first result or from the second result;statistically processing the first result or the second result;selecting relevant parts of the first result or the second result or segmenting an image in relevant image segments;accessing a database for using rules, the rules being adapted to relate the first result to the second result in order to form a combined result;accessing a database for comparing the first result or the second result with reference values for evaluating a conclusion with respect to the chronic disease risk index for the patient. 6. Computer-implemented method according to claim 1, wherein the method is based on predefinable patterns wherein the first result or the second result is compared with the predefinable patterns or properties of reference objects. 7. Computer-implemented method according to claim 1, wherein the method is based on automatic image recognition. 8. Computer-implemented method according to claim 1, wherein the method is performed automatically. 9. Computer-implemented method according to claim 1, wherein the method comprises at least one of the following: storing an intermediate result of the method;storing the chronic disease risk index as a final result;storing a conclusion, wherein the conclusion is based on the chronic disease risk index. 10. Computer-implemented method according to claim 1, where... | Siemens Aktiengesellschaft | A61B005/00
US07664542 | B2 | 2010-02-16 | Registering intra-operative image data sets with pre-operative 3D image data sets on the basis of optical surface extraction | 2006-07-28 | The invention relates to a method for registering intra-operative image data set with pre-operative 3D image data set, including: spatially calibrating an optical 3D sensor system with an intra-operative imaging modality,intra-operatively detecting the surface of an examination area of interest with the 3D sensor system to produce an intra-operative surface mask,intra-operatively recording the area of interest for examination with the intra-operative modality at least partly containing the intra-operative surface mask to obtain an intra-operative image data set,computing the same surface from the pre-operative 3D image data set containing the detected surface to obtain a pre-operative surface mask,registering the intra-operative and pre-operative surface mask with each other,determining a mapping specification between pre-operative 3D image data set and intra-operative image data set based on the calibration and the registration, andoverlaying the intra-operative image data set with the pre-operative 3D data set based on the mapping specification. | 1. A method for registering an intra-operative image data set with a pre-operative 3D image data set for a medical examination, comprising: providing a pre-operative 3D image data set for a medical examination, said data set including information descriptive of a surface along a feature of a patient such that an image of the surface is obtainable from the data set;obtaining a pre-operative surface mask for the surface from the pre-operative 3D image data set;positioning an optical 3D sensor system to provide a side view of the surface of the feature;calibrating the optical 3D sensor system with respect to an intra-operative imaging modality;illuminating the surface with a light pattern such that the side view of the illuminated surface indicates displacement of lines in the pattern along the feature;detecting an area of interest including the illuminated surface of the feature with the optical 3D sensor system and obtaining an intra-operative surface mask of the feature based on data acquired with the optical 3D sensor system;recording the area of interest with the intra-operative imaging modality at least partly containing the intra-operative surface mask to obtain the intra-operative image data set;registering the intra-operative surface mask with the pre-operative surface mask;determining a mapping specification between the intra-operative image data set and the pre-operative 3D image data set based on the calibration of the optical 3D sensor system with the intra-operative imaging modality and the registration of the intra-operative surface mask with the pre-operative surface mask; andoverlaying the intra-operative image data set with the pre-operative 3D data set based on the mapping specification. 2. The method as claimed in claim 1, wherein the intra-operative image data set is a 2D fluoroscopic image of an x-ray C-arm. 3. The method as claimed in claim 1, wherein the intra-operative imaging modality is selected from the group consisting of: an fMRT device, a PET device, a SPECT device, and a functional mapping modality. 4. The method as claimed in claim 3, wherein the functional mapping modality is taken from the group consisting of an EEG and an MEG. 5. The method as claimed in claim 1, wherein the pre-operative 3D image data set originates from a morphological imaging modality selected from the group consisting of: an MRT device, a CT device, a C-arm rotation angiography device, and a 3D ultrasound device. 6. The method as claimed in claim 1, wherein the step of detecting is undertaken by illuminating and observing the surface of the area of interest via a white light projector providing a pattern of stripes on the surface, and features of the surface are computed based on displacement of the stripes as viewed in the side view. 7. The method as claimed in claim 6, wherein the detection is undertaken within one second at a resolution of up to 200 μm. 8. The method as claimed in claim 7, wherein the detection is undertaken before t... | Siemens Aktiengesellschaft | A61B005/05 A61B006/00
US20100016749 | A1 | 2010-01-21 | Method and System for the Monitoring of Respiratory Acitivity and for the Treatment of Breathing Disorders Such as Sleep Apnea | 2007-09-19 | A method and system for sensing the vagus nerve in order to monitor respiratory activity, treating breathing disorders, such as, for example, sleep apnea and a generic bio-interfacing platform that may be adapted for either open-loop or closed-loop applications. | 1. A method for monitoring the respiratory activity of a subject, comprising the steps of: recording an electroneurogram signal from the vagus nerve of the subject;amplifying the electroneurogram signal;computing an amplitude envelope of the amplified electroneurogram signal;applying a matched filter to the amplitude envelope; andcomputing a time between successive peaks of the filtered amplitude envelope;wherein the time between successive peaks is indicative of the respiratory activity of the subject. 2. A method for monitoring the respiratory activity of a subject, comprising the steps of: recording an electroneurogram signal from the vagus nerve of the subject;amplifying the electroneurogram signal;computing an amplitude envelope of the amplified electroneurogram signal;applying a matched filter to the amplitude envelope; andcomputing a time between successive peaks of the filtered amplitude envelope;wherein the time between successive peaks is indicative of the respiratory activity of the subject. 3. A method according to claim 1, further comprising the step of displaying the time between successive peaks. 4. A method according to claim 1, wherein the electroneurogram signal is recorded from a portion of the vagus nerve located between the head and the pulmonary branches of the subject. 5. A method according to claim 1, wherein the electroneurogram signal is recorded from a portion of the vagus nerve located in the neck of the subject. 6. A method according to claim 1, wherein the amplitude envelope computing step includes includes the application of a low-pass filter. 7. A method according to claim 1, wherein the amplitude envelope computing step includes applying a filter selected from a group consisting of a matched filter, a simple averaging filter and a finite-impulse-response filter. 8. A method according to claim 1, wherein the amplitude envelope computing step includes includes applying a rectification and bin integration algorithm to the amplified electroneurogram signal. 9. A method according to claim 8, wherein the amplitude envelope computing step includes further includes applying a moving average filter to the amplified electroneurogram signal after the application of the rectification and bin integration algorithm. 10. A method according to claim 9, wherein the amplitude envelope computing step includes further includes optimizing the result of the moving average filter using the solution to the Wiener-Hopf equation. 11. A method according to claim 1, further comprising the steps of: comparing the time between successive peaks with a sleep apnea event threshold;reporting a sleep apnea event should the compared time between successive peaks be greater than the sleep apnea event threshold. 12. A method according to claim 11, wherein the apnea event threshold is set individually for each subject according to the normal respiration rate of the subject during sleep. 13. A method according to claim 11, wherein the apnea ... | VICTHOM HUMAN BIONICS, INC. | A61B005/08
US20100004553 | A1 | 2010-01-07 | Device for Eliminating Liquid Condensation in a Mouthpiece | 2007-12-04 | In a device for elimination of liquid condensation in a mouthpiece intended to be used in the analysis of exhalation air a tube shaped body of a heat conducting material is provided with heating means and is circumvented by a mouthpiece preferably for one time use. | 1.-5. (canceled) 6. A device for eliminating liquid condensation in a mouthpiece for analysis of exhalation air, comprising a tube-shaped body made from a heat conducting material and having a heater and configured to be inserted into a mouth cavity, whereby the mouthpiece is applied so as to enfold the tube-shaped body. 7. The device of claim 6, wherein the heater is enclosed in a wall of the tube-shaped body. 8. The device of claim 7, wherein a temperature of the tube-shaped body is controlled by a temperature sensor arranged in the wall. 9. The device of claim 6, wherein the mouthpiece is configured for one-time use. 10. The device of claim 9, wherein the mouthpiece protrudes beyond the tube-shaped body into the mouth cavity. | NANOPULS AB | A61B005/08
US20100030051 | A1 | 2010-02-04 | Subcutaneous Glucose Electrode | 2009-08-31 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably three or four-layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer is overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited (third layer). An outer (fourth) layer is biocompatible. | 1. An electrochemical sensor comprising: one or more non-corroding metal or carbon electrodes;a sensing layer comprising an enzyme coupled to each electrode; anda biocompatible layer comprising a biocompatible hydrogel chemically bound to the sensing layer of each electrode. | ABBOTT DIABETES CARE INC. | A61B005/1468
US20100016733 | A1 | 2010-01-21 | Assessing Blood Supply to a Peripheral Portion of an Animal | 2007-02-15 | A post-operative, arterial patency monitoring system for monitoring patency of an artery of a limb including: a first probe positioned adjacent a first tissue area of a limb for producing first output signals dependent upon a first arterial blood supply to the first tissue area; a second probe positioned adjacent a second different tissue area of the limb for producing second, different output signals dependent upon a second arterial blood supply to the second tissue area; and means for processing the first output signals to produce a first arterial blood supply metric for the first arterial blood supply and a second, different, arterial blood supply metric for the first arterial blood supply; means for processing the second output signals to produce the first arterial blood supply metric for the second arterial blood supply and the second arterial blood supply metric for the second arterial blood supply; means for using a predetermined function that takes as its arguments at least the first and second arterial blood supply metrics for the first arterial blood supply to produce a combined arterial patency metric for the first blood supply; means for using the predetermined function that takes as its arguments at least the first and second arterial blood supply metrics for the second arterial blood supply to produce a combined arterial patency metric for the second arterial blood supply; means for comparing the combined arterial patency metric for the first arterial blood supply with the combined arterial patency metric for the second arterial blood supply; and means for determining the patency of a first artery that provides the first arterial blood supply using a difference between the combined arterial patency metric for the first arterial blood supply and the combined arterial patency metric for the second arterial blood supply. | 1. A post-operative, arterial patency monitoring system for monitoring patency of an artery of a limb comprising: a first probe positioned adjacent a first tissue area of a limb for producing first output signals dependent upon a first arterial blood supply to the first tissue area;a second probe positioned adjacent a second different tissue area of the limb for producing second, different output signals dependent upon a second arterial blood supply to the second tissue area; andmeans for processing the first output signals to produce a first arterial blood supply metric for the first arterial blood supply and a second, different, arterial blood supply metric for the first arterial blood supply;means for processing the second output signals to produce the first arterial blood supply metric for the second arterial blood supply and the second arterial blood supply metric for the second arterial blood supply;means for using a predetermined function that takes as its arguments at least the first and second arterial blood supply metrics for the first arterial blood supply to produce a combined arterial patency metric for the first blood supply;means for using the predetermined function that takes as its arguments at least the first and second arterial blood supply metrics for the second arterial blood supply to produce a combined arterial patency metric for the second arterial blood supply;means for comparing the combined arterial patency metric for the first arterial blood supply with the combined arterial patency metric for the second arterial blood supply; andmeans for determining the patency of a first artery that provides the first arterial blood supply using a difference between the combined arterial patency metric for the first arterial blood supply and the combined arterial patency metric for the second arterial blood supply. 2. A system as claimed in claim 1, wherein the predetermined function multiplies together at least the first and second arterial blood supply metrics to produce a combined arterial patency metric. 3. A system as claimed in claim 2, wherein the first arterial blood supply metric and the second arterial blood supply metric are normalized before multiplication. 4. A system as claimed in claim 1, wherein the first arterial blood supply metric and the second arterial blood supply metric are different ones of: arterial blood supply signal strength, arterial blood supply signal caliber, peripheral pulse rate, skin color redness and skin temperature. 5. A system as claimed in claim 1, wherein each probe comprises at least a first sensor and a second sensor, wherein the first sensor provides first output signals used to produce the first arterial blood supply metric and wherein the second sensor provides first output signals used to produce the second arterial blood supply metric and wherein the first and second sensors are different ones of: a plethysmograph sensor, a skin temperature sensor and a skin color sensor. 6. A sys... | - | A61B005/0205 A61B005/021
US07653430 | B2 | 2010-01-26 | Method and system for verifying the integrity of normal sinus rhythm templates | 2005-11-28 | A method and system for verifying the integrity of normal sinus rhythm (NSR) templates and updating the NSR template after selected time intervals. At selected time intervals after establishing a NSR template, cardiac complexes are sensed and values for one or more cardiac parameters are measured. The values of the cardiac parameters are compared to predetermined value ranges for NSR cardiac complexes. When the values of the cardiac parameters fall within the predetermined value ranges, values for the differences between the values of the cardiac parameters for the cardiac complexes and the values for the cardiac parameters of the NSR cardiac complexes are calculated. When the values of the differences are greater than one or more threshold values, the NSR template is updated as a function of the sensed cardiac complexes. | 1. A medical device, comprising: a morphology analyzer adapted to extract and measure sensed cardiac complex parameter values;a template generating circuit connected to the morphology analyzer and adapted to receive the sensed cardiac complex parameter values and generate a normal sinus rhythm (NSR) template; anda template comparison circuit connected to the template generating circuit and adapted to compare sensed cardiac complex parameter values to predetermined cardiac parameter values,wherein the template comparison circuit and the template generating circuit cooperate to update the NSR template based on the compared sensed cardiac complex parameter values and the predetermined cardiac parameter values. 2. The medical device of claim 1, wherein the template comparison circuit and the template generating circuit are adapted to cooperate to update the NSR template when the sensed cardiac complex parameter values vary from a current NSR template by a calculated threshold value. 3. The medical device of claim 2, wherein the template comparison circuit is adapted to calculate the threshold value by multiplying parameter values for the current NSR template by a predetermined deviation percentage. 4. The medical device of claim 3, wherein the template comparison circuit is adapted to determine a number of attempts to evaluate the sensed cardiac complex parameter values. 5. The medical device of claim 4, further comprising input circuitry connected to the morphology analyzer and adapted to be connected to at least one electrode on at least one lead. 6. The medical device of claim 1, wherein the template comparison circuit is adapted to compare sensed cardiac complex parameter values to predetermined cardiac parameter values after a selected time interval has elapsed. 7. A device, comprising: input circuitry adapted to be electrically connected to at least one electrode to sense cardiac complexes including normal sinus rhythm (NSR) cardiac complexes;a morphology analyzer circuit connected to the input circuitry and adapted to measure sensed cardiac complex parameter values;a template generating circuit connected to the morphology analyzer circuit and adapted to receive the sensed cardiac complex parameter values and generate a NSR template; anda template comparison circuit connected to the template generating circuit and adapted to compare sensed cardiac complex parameter values to predetermined cardiac parameter values after a selected time interval has elapsed,wherein the template comparison circuit and the template generating circuit cooperate to update the NSR template based on the compared sensed cardiac complex parameter values and the predetermined cardiac parameter values. 8. The device of claim 7, further comprising: a microprocessor adapted to connect to the input circuitry, the morphology analyzer circuit, the template generating circuit, and the template comparison circuit. 9. The device of claim 8, wherein the microprocessor inclu... | Cardiac Pacemakers, Inc. | A61B005/04
US07660619 | B2 | 2010-02-09 | Method for determining the condition of an object by magnetic resonance imaging | 2006-03-18 | A method for determining the condition of an animate or inanimate object by magnetic resonance imaging, MRI, particularly for determining the pathologic condition of rheumatoid arthritis (“arthritis rheumatoides”) and carrying out a patient follow-up. | 1. A method for using magnetic resonance imaging (MRI) to evaluate a selected region of interest of an object, the method comprising: acquiring at a plurality of successive acquisition times an MRI image of the selected region of interest of a sample object having a known condition;generating for the sample object a function representing a mean intensity of each of the acquired MRI images for each of the successive acquisition times;acquiring at a plurality of successive acquisition times an MRI image of the selected region of interest of a test object having an unknown condition, the test object being different from the sample object;generating for the test object a function representing a mean intensity of each of the acquired MRI images for each of the successive acquisition times;evaluating the unknown condition of the test object by comparing the generated function for the test object to the generated function for the sample object; andconstructing a development time table of conditions of the test object so as to monitor the development of a disease. 2. The method of claim 1, further comprising prior to the steps of acquiring the MRI images of the sample object and the test object, respectively: treating the sample object to induce a response indicative of the known condition; andtreating the test object to induce a response indicative of the unknown condition. 3. The method of claim 2, wherein the treating steps comprise injecting a substance into the sample object and test object, respectively. 4. The method of claim 1, wherein the generating the sample object function and the test object function steps each comprises filling a perfusion curve to a plurality of data points, wherein each of the data points comprises a measured mean intensity and a corresponding one of the successive acquisition times. 5. The method of claim 4, wherein the filling the perfusion curve step comprises interpolating between at least two of the data points. 6. The method of claim 1, wherein the acquiring the MRI images steps each comprise acquiring a two-dimensional MRI image of the selected region of interest along a section plane of the sample object and the test object, respectively. 7. The method of claim 1, wherein the acquiring the MRI images steps each comprise acquiring a three-dimensional MRI image of the selected region of interest of the sample object and the test object, respectively. 8. The method of claim 7, wherein the acquiring a three-dimensional MRI image step comprises: acquiring a three-dimensional image data array;defining a section plane of the sample object and the test object, respectively; andselecting image data points from the three-dimensional image data array that fall on the defined section plane. 9. The method of claim 1, further comprising tracking the development of the unknown condition by repeating for each of a plurality of follow-on sessions the acquiring the MRI images of the test object step, the generating the test... | Esaote S.p.A. | A61B005/05
US07647086 | B2 | 2010-01-12 | Apparatus and method for magnetic resonance angiography utilizing flow pulses and phase-encoding pulses in a same direction | 2003-10-06 | A magnetic resonance imaging apparatus includes an RF coil unit which generates RF pulses toward a subject, and which receives an MR signal from the subject. Gradient magnetic field coils generate a gradient magnetic field for slice selection, a gradient magnetic field for phase encoding and a gradient magnetic field for frequency encoding, respectively. An arithmetic unit generates image data on the basis of the MR signal, and a sequence controller controls phase encoding gradient magnetic field coils in order to generate flow pulses for dephasing or rephasing the MR spin of blood flow within the subject, in the same direction as that of the phase encoding gradient magnetic field. | 1. A magnetic resonance imaging (MRI) apparatus for providing MRA image data without use of contrast agent, said apparatus comprising: an RF coil unit which transmits RE pulses toward a subject, and which receives an MR signal from the subject in the absence of contrast agent in said subject;gradient magnetic field coils which generate gradients in a superposed background magnetic field Bo, said gradients being generated in respectively different spatial directions and including (a) a gradient magnetic field for slice selection, (b) a gradient magnetic field for phase encoding, and (c) a gradient magnetic field for frequency encoding;an arithmetic unit which generates MRA image data based on the MR signal including a subtraction image related to a systolic phase and a diastolic phase on the basis of the MR signals acquired during the systolic phase and the diastolic phase;a sequence controller which (a) executes a first MRI scanning process during a systolic phase of a subject's heart cycle and a second MRI scanning process during a diastolic phase of the subject's heart cycle, and (b) controls the phase encoding gradient magnetic field coils to generate magnetic gradient flow pulses for dephasing or rephasing MR spin of a blood flow within said subject in the same spatial direction as the phase encoding gradient magnetic field during at least one of the first and second scanning processes; anda subtraction processing unit which subtracts from each other (a) a systolic phase image based on said first scanning process and (b) a diastolic phase image based on said second scanning process. 2. A magnetic resonance imaging apparatus as in claim 1, wherein the flow pulses are flow compensation pulses or flow spoiled pulses. 3. A magnetic resonance imaging apparatus as in claim 1, wherein said sequence controller controls the frequency encoding gradient magnetic field coils to generate additional flow pulses in the same spatial direction as the frequency encoding gradient magnetic field. 4. A magnetic resonance imaging apparatus as in claim 3, wherein the first-mentioned flow pulses are flow compensation pulses, and the additional flow pulses are also flow compensation pulses. 5. A magnetic resonance imaging apparatus as in claim 3, wherein the first-mentioned flow pulses are flow spoiled pulses, and the additional flow pulses are also flow spoiled pulses. 6. A magnetic resonance imaging apparatus as in claim 3, wherein the first-mentioned flow pulses are flow spoiled pulses, and the additional flow pulses are flow compensation pulses. 7. A magnetic resonance imaging apparatus as in claim 3, wherein the first-mentioned flow pulses are flow compensation pulses, and the additional flow pulses are flow spoiled pulses. 8. A magnetic resonance imaging apparatus as in claim 1, wherein said sequence controller controls the slice selection gradient magnetic field coils to generate a slice encoding gradient magnetic field in the same direction as the slic... | Kabushiki Kaisha Toshiba | A61B005/05
US20100010320 | A1 | 2010-01-14 | MOBILE MEDICAL WORKSTATION AND A TEMPORARILY ASSOCIATING MOBILE COMPUTING DEVICE | 2009-07-07 | A system, method and software instructions that enable a mobile computing device to communicate with and search a vital signs measuring device for any physiological data that has been obtained from one or more patients and that has not yet been communicated to another device, such as to a electronic medical record repository. Further, the data is accessed by and stored into the mobile computing device for review, logging and selective transmission to other devices including such as the electronic medical record repository. | 1. A system for managing physiological data obtained from each of one or more patients, comprising: at least one vital signs measuring device that is configured to obtain and store physiological data associated with one or more patients;a mobile computing device that is configured to communicate with said vital signs measuring device; anda software program executing within said mobile computing device, said program is configured to communicate commands to said vital signs measuring device and configured to search said vital signs measuring device for said physiological data that has been obtained and that has not yet been communicated to another device by said vital signs monitoring device; and where said program is further configured to access and store said data within said mobile computing device. 2. The system of claim 1 wherein said vital signs measuring device and said mobile computing device communicate via a wireless communications channel. 3. The system of claim 1 further including a docking station that is electronically attached to said vital signs measuring device and that interfaces with said mobile computing device and wherein said vital signs measuring device and said mobile computing device communicate via said docking station. 4. The system of claim 1 wherein said mobile computing device transmits data selected by a user of said mobile computing device to said another device upon demand of said user. 5. The system of claim 1 wherein said mobile computing device logs data that is selected by a user of the mobile computing device; 6. The system of claim 1 wherein said mobile computing device periodically attempts to re-transmit said data in response to a transmission error associated with said another device. 7. The system of claim 6 wherein said mobile computing device periodically reminds said user of data that has not yet been transmitted from said mobile computing device. 8. The system of claim 1 wherein said another device is an electronic medical records server. 9. A method for managing physiological data obtained from each of one or more patients, the method comprising: communicating commands to a vital signs measuring device;searching said vital signs measuring device for said physiological data that has been obtained and that has not yet been communicated to another device by said vital signs measuring device;accessing said data within a mobile computing device that is configured to communicate with said vital signs measuring device; andstoring said data within said mobile computing device. 10. The method of claim 1, wherein said vital signs measuring device and said mobile computing device communicate via a wireless communications channel. 11. The method of claim 1, wherein a docking station is electronically attached to said vital signs measuring device and interfaces with said mobile computing device; and wherein said vital signs measuring device and said mobile computing device communicate via said docking st... | - | A61B005/00
US07643859 | B2 | 2010-01-05 | Non-invasive glucose meter | 2007-10-11 | An apparatus and method for measuring a concentration of a substance in an eye using a retro-reflected measurement light beam having a first wavelength at which the substance has a non-zero first absorption coefficient and a retro-reflected reference light beam having a second wavelength at which the substance has a second absorption coefficient which is substantially equal to zero. The apparatus further includes a detector positionable to receive the retro-reflected measurement light beam and the retro-reflected reference light beam. The detector is responsive to light having the first wavelength by generating a measurement signal and responsive to light having the second wavelength by generating a reference signal. The apparatus further includes an electrical circuit coupled to the detector. The electrical circuit is responsive to the measurement signal and the reference signal to measure the concentration of the substance in the eye. | 1. A method of determining the concentration of a substance in an eye, the method comprising: providing a nearly parallel measurement light beam having a first wavelength at which the substance has a non-zero first absorption coefficient;providing a nearly parallel reference light beam having a second wavelength at which the substance has a second absorption coefficient which is substantially equal to zero;irradiating the retina of the eye with the measurement light beam, thereby passing the measurement light beam through the cornea, the pupil, the lens, and the liquid of the eye;irradiating the retina of the eye with the reference light beam, thereby passing the reference light beam through the cornea, the pupil, the lens, and the liquid of the eye;reflecting at least a portion of the measurement light beam from the retina and through the liquid, the lens, the pupil, and the cornea of the eye, thereby producing a retro-reflected measurement light beam having the first wavelength;reflecting at least a portion of the reference light beam from the retina and through the liquid, the lens, the pupil, and the cornea of the eye, thereby producing a retro-reflected reference light beam having the second wavelength;providing a detector responsive to light having the first wavelength and responsive to light having the second wavelength;irradiating the detector with the retro-reflected measurement light beam, the detector generating a measurement signal in response thereto;irradiating the detector with the retro-reflected reference light beam, the detector generating a reference signal in response thereto; anddetermining the concentration of the substance in the eye in response to the measurement signal and the reference signal. 2. The method of claim 1, wherein the detector primarily receives the retro-reflected measurement light beam and the retro-reflected reference light beam without receiving other reflections from the eye. 3. The method of claim 1, wherein the measurement light beam and the reference light beam are focused at a point on the retina and the detector is located such that the retro-reflected measurement light beam and the retro-reflected reference light beam propagate along an optical path of at least 100 mm between the eye and the detector. 4. The method of claim 1, wherein the substance comprises glucose. 5. The method of claim 1, wherein the measurement light beam is provided by a first infrared light source and the reference light beam is provided by a second infrared light source. 6. The method of claim 1, wherein the measurement light beam is focused at a point on the retina. 7. The method of claim 1, wherein the reference light beam is focused at a point on the retina. 8. An apparatus for measuring a concentration of a substance in an eye having a retina, the apparatus comprising: a measurement light source producing a nearly parallel measurement light beam having a first wavelength, the substance having a non-zero first ab... | GlucoVista, LLC | A61B005/145
US07653425 | B2 | 2010-01-26 | Method and system for providing calibration of an analyte sensor in an analyte monitoring system | 2006-08-09 | Method and apparatus for providing calibration of analyte sensor including applying a control signal, detecting a measured response to the control signal, determining a variance in the detected measured response, and estimating a sensor sensitivity based on the variance in the detected measured response is provided. | 1. A method of calibrating an analyte sensor, comprising: applying a control signal to the analyte sensor;detecting a measured response to the control signal;determining a variance in the detected measured response; andestimating, with a processor, a sensor sensitivity based on the variance in the detected measured response. 2. The method of claim 1 wherein the level of the control signal varies in time. 3. The method of claim 1 wherein the control signal includes a square wave signal. 4. The method of claim 3 wherein the square wave signal is applied to a poise voltage. 5. The method of claim 1 wherein detecting the measured response includes determining a work electrode current signal. 6. The method of claim 1 wherein the variance is determined based on comparing the difference between the beginning and end of the half duty cycle of the measured response to the control signal. 7. The method of claim 1 wherein estimating the sensor sensitivity includes retrieving a predetermined sensor sensitivity corresponding to the determined variance in the detected measured response. 8. The method of claim 1 further including determining a validity of the estimated sensor sensitivity. 9. The method of claim 1 further including determining analyte level based on the estimated sensor sensitivity. 10. The method of claim 1 wherein the sensor includes an analyte sensor. 11. An analyte sensor device, comprising: a processor configured to apply a control signal to the analyte sensor, detect a measured response to the control signal, determine a variance in the detected measured response, and estimate a sensor sensitivity based on the variance in the detected measured response. 12. The device of claim 11 wherein the processor is configured to vary the level of the control signal with time. 13. The device of claim 11 wherein the control signal includes a square wave signal. 14. The device of claim 13 wherein the processor is configured to apply the square wave signal to a poise voltage. 15. The device of claim 11 wherein the processor is configured to determine a work electrode current signal of an analyte sensor operatively coupled to the processor. 16. The device of claim 11 wherein the processor is configured to determine the variance based on comparing the difference between the beginning and end of the half duty cycle of the measured response to the control signal. 17. The device of claim 11 wherein the processor is configured to retrieve a predetermined sensor sensitivity corresponding to the determined variance in the detected measured response. 18. The device of claim 11 wherein the processor is operatively coupled to a data receiver unit configured to determine determining a validity of the estimated sensor sensitivity. 19. The device of claim 18 wherein the data receiver unit is configured to determine an analyte level based on the estimated sensor sensitivity. 20. The device of claim 11 wherein the processor is operatively coupled to a... | Abbott Diabetes Care Inc. | A61B005/00 C12M001/00 C12N009/00 C12N011/00 G01N001/00
US07664661 | B2 | 2010-02-16 | Electronic method and system that improves efficiencies for rendering diagnosis of radiology procedures | 2008-02-11 | An electronic method and system for improving radiologists' efficiencies when viewing radiology procedures and rendering diagnosis in a manner that emulates current methods and apparatus. The method of the present invention includes reviewing electronic radiology images and reports contained in a patient's digital master folder (an information object invented to manage the patient's radiology information), comparing images from the current procedure to specific images from prior procedures in a specified order, and dictating the procedure's diagnosis into the digital master folder. The apparatus of the present invention includes a flat panel monitor for the viewing and manipulation of digital master folders, a dictation trackball device for manual and voice enabled operation and navigation of the system, and multiple high-resolution computer monitors functioning as a “light-box” for the viewing of radiology images. | 1. An electronic method of improving the efficiency of a radiologist, comprising the steps of: a) providing a system that includes a computer and a plurality of monitors interfaced with the computer, each monitor for displaying an image;b) using at least one of the monitors to simulate a radiology “light box” for displaying electronic radiology images;c) using at least one of the monitors to simulate a digital graphical representation of a patient's manual master folder comprising the steps of: (i) generating by a specialized software application the digital graphical representation of the patient's manual master folder as an image on the monitor specifically designed for use as a graphical user interface by the radiologist;(ii) generating data fields on the digital graphical representation including patient information, medical procedures information and radiologist information regarding all radiology procedures associated with such patient stored in an information database;(iii) providing information and hyperlinks to radiology reports and images in an electronic layout and color scheme conforming to the layout and color scheme of the patient's manual master folder and tailored to a radiologist practice; and(iv) including hyperlinks within at least one data field to view additional information or images relating to a patient's medical records when clicked;d) using a hyperlink to open the folder displayed in step “c” to display information contained in the folder; ande) using a hyperlink that accesses the folder to display a current radiology image from a current radiology procedure being performed on the patient to permit diagnosis. 2. The method of claim 1 wherein in step “d” a voice activated command is used to open the patient's master folder. 3. The method of claim 1 wherein in step “d” a trackball device is used to open the patient's master folder. 4. The method of claim 1 further comprising the step of providing a combination dictation and trackball device, and wherein in step “d” a user can selectively use either a voice activated command or a trackball to open the patient's master folder. 5. The method of claim 1 further comprising the step of using the computer to interface the monitors and the hyperlink. 6. The method of claim 4 further comprising the step of using the computer to interface the monitors and the combination dictation and trackball device. 7. The method of claim 1 wherein there are two monitors in step “c” that are used to display electronic radiology images. 8. The method of claim 1 wherein the monitor in step “c” that is used to display electronic radiology images is a high resolution monitor. 9. The method of claim 7 wherein the monitors in step “c” that are used to display electronic radiology images is a high resolution monitor. 10. The method of claim 1 wherein in step “b” the image viewed is an ultrasound image. 11. The method of claim 1 wherein in step “b” the image viewed is a magnetic resonance image... | Avreo, Inc. | A61B005/00
US07643863 | B2 | 2010-01-05 | Diffusion tensor and q-space MRI specimen characterization | 2004-07-08 | Magnetic resonance methods include modeling magnetic resonance signals obtained from specimens at low and high q-values to obtain parameters associated with specimen structure and orientation. In evaluation of brain white matter specimens, diffusion within axons can be modeled based on hindered diffusion parallel to an axis of the axon and restricted diffusion perpendicular to the axis. Diffusion exterior to axons can be modeled as hindered diffusion with differing diffusivities parallel to and perpendicular to the axis. Based on extracted parameters and associated model functions, specimen properties such as intra and extra-axonal principal diffusivities and the corresponding principal directions can be estimated. | 1. A computer-implemented method for performing magnetic-resonance analysis of a sample, comprising: obtaining a set of translational diffusion-weighted magnetic resonance signals associated with a plurality of diffusion-weighted field strengths and a plurality of diffusion-weighted field directions;providing the set of translational diffusion-weighted magnetic resonance signals to a processor;using the processor to differentiate a signal portion in the set of signals corresponding to restricted diffusion from a signal portion in the set of signals corresponding to hindered diffusion;estimating at least one parameter associated with translational diffusion in a restricted compartment of the sample based on the signal portion corresponding to restricted diffusion, wherein the estimated parameter associated with the restricted compartment corresponds to at least one of an orientation of the restricted compartment, a spin fraction associated with the restricted compartment, a principal diffusivity associated with the restricted compartment, or a dimension of the restricted compartment; andproviding the estimated parameter. 2. The method of claim 1, wherein the set of signals is a set of image signals. 3. The method of claim 1, wherein the estimated parameter associated with the restricted compartment corresponds to the orientation of the restricted compartment. 4. The method of claim 1, wherein the estimated parameter associated with the restricted compartment corresponds to the spin fraction associated with the restricted compartment of the sample. 5. The method of claim 1, wherein the estimated parameter associated with the restricted compartment is the principal diffusivity associated with the restricted compartment. 6. The method of claim 1, wherein the estimated parameter associated with the restricted compartment corresponds to the dimension of the restricted compartment. 7. The method of claim 1, further comprising estimating at least a second parameter associated with translational spin diffusion in a second restricted compartment based on the signal portion corresponding to restricted diffusion. 8. The method of claim 1, further comprising estimating at least one parameter associated with translational diffusion in a hindered compartment of the sample based on the signal portion corresponding to hindered diffusion. 9. The method of claim 8, wherein the at least one estimated parameter associated with the hindered compartment is a spin fraction associated with the hindered compartment. 10. The method of claim 9, wherein the at least one estimated parameter associated with the hindered compartment is a diffusivity associated with the hindered compartment. 11. The method of claim 1, further comprising: associating the signal portions corresponding to restricted diffusion with diffusion in an intra-axonal volume of the sample; andassociating the signal portions corresponding to hindered diffusion with diffusion in an extra-axonal volu... | - | A61B005/055
US08167824 | B2 | 2012-05-01 | Sensation detection presentation apparatus and walking rehabilitation supporting apparatus | 2007-05-05 | It is intended to feed back any somatic sensation information on given body surface portion in easily understandable form without the use of special sensations, such as visual and auditory sensations. Walking rehabilitation supporting apparatus (10) comprises sensation detection unit (11) for detecting the sole pressure of patient (H); sensation presentation unit (13) for imparting given stimuli to the patient (H); and control unit (15) for controlling the behavior of the sensation presentation unit (13) so as to impart stimuli corresponding to the sole pressure detected by the sensation detection unit (11). The sensation detection unit (11) is disposed so as to be capable of detecting the sole pressure on a paralyzed side, and the sensation presentation unit (13) is fitted to a body surface portion of sensation superior to that of the sole. | 1. A walking rehabilitation supporting apparatus that supports rehabilitation for a patient having walking difficulty to acquire accurate walking behavior, comprising: a sensation detection unit configured for detecting sole pressure of said patient;a sensation presentation unit configured for imparting stimulus to said patient;a control unit for controlling behavior of said sensation presentation unit so as to impart stimulus corresponding to the sole pressure detected by said sensation detection unit,wherein said sensation detection unit is configured for detecting the sole pressure on a paralyzed side and a non-paralyzed side of said patient, andsaid sensation presentation unit is configured to fit a body surface portion of said patient and provide sensation superior to that of said foot sole of said paralyzed side, andsaid control unit is further configured to control the behavior of said sensation presentation unit according to a difference in sole pressure between said paralyzed side and said non-paralyzed side. 2. A walking rehabilitation supporting apparatus that supports rehabilitation for a patient having walking difficulty to acquire accurate walking behavior, comprising: a sensation detection unit configured for detecting sole pressure of said patient;a sensation presentation unit configured for imparting stimulus to said patient; anda control unit for controlling behavior of said sensation presentation unit so as to impart stimulus corresponding to the sole pressure detected by said sensation detection unit;wherein said sensation detection unit is configured for detecting the sole pressure on a paralyzed side and a non-paralyzed side of said patient, said sensation presentation unit is configured to fit a body surface portion of said patient and provide sensation superior to that of said foot sole of said paralyzed side, andsaid control unit is further configured to control the behavior of said sensation presentation unit so as to switch between stimulus corresponding to the sole pressure on said paralyzed side and stimulus corresponding to the sole pressure on said non-paralyzed side and to impart the stimulus to said body surface portion. 3. The walking rehabilitation supporting apparatus according to claim 1 or 2, wherein said sensation detection unit includes two soles configured for placement on said paralyzed side and said non-paralyzed side, respectively, and pressure sensors for detecting sole pressure of the patient for each of a plurality of regions in the respective soles, and said pressure sensors are disposed so as to be capable of detecting the sole pressures of said plurality of regions, said sensation presentation unit includes a stimulus imparting device configured for imparting the stimulus to the skin so that said patient receives touch-pressure sensation, vibratory sensation and/or movement sensation, wherein said stimulus imparting device includes a plurality of pressing members provided correspondingly to said... | Waseda University | A61B005/103 A61B005/117
US08394144 | B2 | 2013-03-12 | System for positioning of surgical inserts and tools | 2007-09-25 | A tracking and positioning system and method to enable the precise positioning of an object or tool relative to its surgical surroundings, and in accordance with preoperative CT images of the operating site. When used for artificial spinal disc positioning, the system comprises a computing system incorporating in memory the preoperative CT data showing the two vertebrae and the predetermined disc position between them; a 3-D target having radio-opaque markers for attaching to one of the vertebrae to define the position of the vertebra in an intra-operative fluoroscope image of the spine; a tool for intra-operative insertion of the artificial disc, and a registration system for relating the intra-operative fluoroscope image to the preoperative CT data, such that the predetermined disc position is displayed in the fluoroscope image of the subject, thereby enabling the surgeon to place the artificial disc accurately in its intended position. | 1. A system for intraoperatively placing an artificial disc at a predetermined position between two vertebrae of a subject, said system comprising: a computing system having a storage medium including a set of preoperative CT data showing said two vertebrae and said predetermined position of said disc between them;at least one three-dimensional target having radio-opaque markers adapted to be attached to at least one of said vertebrae to define the position of said vertebra in at least one intra-operative fluoroscope image of the subject;a tool for intra-operative insertion of said artificial disc; anda registration system for relating said at least one intra-operative fluoroscope image to said set of preoperative CT data, such that said predetermined disc position is displayed in said at least one fluoroscope image of the subject, thereby enabling intraoperative placement of said artificial disc in said predetermined disc position. 2. A system according to claim 1 and wherein said predetermined position of said disc is obtained from predefined geometrical information about the position. 3. A system according to claim 1 and wherein said predetermined position of said disc is obtained by adjusting said disc position in a model obtained from said preoperative CT data, such that facet joints between said vertebrae are in minimally stressed positions. 4. A system according to claim 1 and further comprising at least two pins attached to each of said vertebrae, each of said pins comprising at least two radio-opaque markers disposed at known positions along said pins, such that said pins define the three-dimensional position of said vertebrae in said fluoroscope images of the subject, such that said target may be removed to allow intra-operative placement of said disc in said predetermined position. 5. A system according to claim 1 and further comprising at least one pin attached to each of said vertebra, each of said pins comprising at least three radio-opaque markers disposed at fixed positions along said pins, such that said pins define the three-dimensional position of said vertebrae in said fluoroscope images of the subject, such that said target may be removed to allow intra-operative placement of said disc in said predetermined position. 6. A system according to claim 1 and wherein said insertion tool is either a manually operated tool or a robotically operated tool. 7. A system according to claim 1, wherein said at least one three-dimensional target comprises at least two pins attached to one of said two vertebrae, each of said at least two pins comprising at least two radio-opaque markers disposed along their length, such that the position of said vertebra is defined in at least one intra-operative fluoroscope image of the subject. 8. A method of intraoperative positioning an artificial disc between two vertebrae of a subject, said method comprising the steps of: generating from a preoperative set of CT data, a CT model showing said vert... | Mazor Surgical Technologies Ltd. | A61F002/44
US08494623 | B2 | 2013-07-23 | Method and apparatus for in-vivo physiological monitoring | 2009-06-09 | A method is provided that includes providing a monitoring apparatus including one or more modules within a target cavity or lumen of a body. The one or more modules are provided within the target cavity or lumen in a first state in which the monitoring apparatus is configured to remain within the target cavity or lumen. The method further includes monitoring physiological conditions of the body using one or more sensors within the one or more modules, and providing the one or more modules in a second state in which the monitoring apparatus is configured to exit the target cavity or lumen. | 1. A monitoring apparatus comprising: a plurality of modules connected in series configured to be provided within a target cavity or lumen of a body,wherein one or more of the plurality of modules is configured to continuously monitor physiological conditions of the body, andwherein the plurality of modules are configured to be provided in a first state in which each of the plurality of modules is locked to another one or more of the plurality of modules to form a closed ring shape and said monitoring apparatus is configured to remain within the target cavity or lumen without fixing to a wall of the target cavity or lumen, andin a different second state in which the plurality of modules separate from each other and said monitoring apparatus is configured to exit the target cavity or lumen, and wherein a state indicates a form of a collection of the plurality of modules. 2. The monitoring apparatus according to claim 1, wherein said one or more modules are configured to be provided in a third state in which said monitoring apparatus is configured to be inserted within the target cavity or lumen. 3. The monitoring apparatus according to claim 1, wherein said one or more modules are configured to transform into said second state based upon a signal generated from a component within said one or more modules. 4. The monitoring apparatus according to claim 1, wherein said one or more modules includes a module configured to wirelessly transmit and receive communications from an external device provided outside of the body. 5. The monitoring apparatus according to claim 4, wherein said one or more modules are configured to transform into said second state based upon a signal from the external device. 6. A method comprising: providing a monitoring apparatus including a plurality of modules connected in series within a target cavity or lumen of a body, wherein the plurality of modules are provided within the target cavity or lumen in a first state in which each of the plurality of modules is locked to another one or more of the plurality of modules to form a closed ring shape and the monitoring apparatus is configured to remain within the target cavity or lumen without fixing to a wall of the target cavity or lumen;monitoring physiological conditions of the body using one or more sensors within one or more of the plurality of modules; andproviding the one or more modules in a different second state in which the plurality of modules separate from each other and the monitoring apparatus is configured to exit the target cavity or lumen,wherein a state indicates a form of a collection of the plurality of modules. 7. The method according to claim 6, wherein the monitoring apparatus is implanted within the target cavity or lumen with the one or more modules in a third state in which the monitoring apparatus is configured to be inserted into the target cavity or lumen. 8. The method according to claim 7, wherein the monitoring apparatus is implanted within... | The Board of Trustees of the Leland Stanford Junior University | A61B005/0452
US08114033 | B2 | 2012-02-14 | Sleep determination apparatus | 2007-11-19 | A sleep determination apparatus capable of detecting wakefulness of a subject before entering a hypnagogic state with high accuracy is provided.;When the subject is not yet in the hypnagogic state, it is determined that the subject is asleep when a detected body movement signal continuously remains below a determination threshold for a predetermined time or longer, and that the subject is awake in other situations. When the subject has entered the hypnagogic state, it is determined that the subject is awake when the detected body movement signal continuously remains above the determination threshold for the predetermined time or longer, and that the subject is asleep in other situations. | 1. A sleep determination apparatus comprising: a body movement detection device disposed outside a body of a subject and which is attached to a bedding to detect body movement of the subject and output the body movement as a body movement signal;a first sleep determination device which determines that the subject is asleep when the body movement signal continuously remains below a determination threshold for a predetermined time or longer, and that the subject is awake when a condition that the body movement signal continuously remains below the determination threshold for the predetermined time or longer is not satisfied;a sleep onset determination device which determines that sleep onset of the subject has occurred when the first sleep determination device makes a determination continuously for a predetermined time or longer that the subject is asleep; anda second sleep determination device which determines in place of the first sleep determination device, when the sleep onset determination device determines that sleep onset of the subject has occurred, that the subject is awake when the body movement signal continuously remains above a determination threshold for a predetermined time or longer, and that the subject is asleep when a condition that the body movement signal continuously remains above the determination threshold for the predetermined time or longer is not satisfied,wherein the first sleep determination device, the sleep onset determination device, and the second sleep determination device are disposed in the apparatus. 2. The sleep determination apparatus of claim 1, wherein the predetermined time is 3 minutes. 3. The sleep determination apparatus of claim 1, further comprising: a determination threshold setting device which sets and updates the determination threshold based on a minimum value of the body movement signal in every given time. | Daikin Industries, Ltd. | A61B005/103 A61B005/117
US08219206 | B2 | 2012-07-10 | Dwell time adjustments for posture state-responsive therapy | 2009-04-30 | A medical device detects a patient posture state and adjusts therapy based on the detected posture state. A patient posture state may be transient. To avoid adjusting therapy in response to transient posture states, a dwell time may be imposed. Using the dwell time, only those posture states that are assumed by a patient for a minimum length of time will result in therapy adjustment. To help ensure that the dwell time is set to an appropriate length of time, patient therapy adjustment data may be analyzed. Therapy adjustments made during the dwell time may indicate that the dwell time is too long. Therapy adjustments made prior to the initiation of the dwell time may also indicate that the dwell time is too long. Therapy adjustments made by the patient after expiration of the dwell time may indicate that the dwell time is too short. Based on analysis of the therapy adjustment data, the dwell time can be adjusted to a length that is appropriate for the patient and/or posture state. | 1. A method comprising: delivering therapy to a patient;detecting a posture state transition of a patient;after a time delay following the posture state transition, adjusting the therapy based on the posture state transition;detecting a patient adjustment of the therapy; andadjusting the time delay based on the patient adjustment. 2. The method of claim 1, wherein adjusting the time delay comprises adjusting the time delay if the patient adjustment is detected during or after the time delay. 3. The method of claim 1, wherein adjusting the time delay comprises adjusting the time delay if the patient adjustment is detected during the time delay. 4. The method of claim 1, wherein adjusting the time delay comprises adjusting the time delay if a number of patient adjustments detected during multiple instances of the posture state transition exceeds a threshold. 5. The method of claim 1, wherein adjusting the time delay comprises reducing the time delay if the patient adjustment is detected during the time delay. 6. The method of claim 5, wherein reducing the time delay comprises reducing the time delay based on timing of the patient adjustment. 7. The method of claim 1, wherein adjusting the time delay comprises increasing the time delay if the patient adjustment is detected after the time delay. 8. The method of claim 7, wherein increasing the time delay comprises increasing the time delay based on timing of the patient adjustment. 9. The method of claim 1, further comprising determining whether the patient adjustment is consistent with the adjustment of therapy based on the posture state transition, wherein adjusting the time delay comprises adjusting the time delay based on the patient adjustment and the determination of whether the patient adjustment is consistent with the adjustment of therapy based on the posture state transition. 10. The method of claim 1, wherein the posture state transition comprises a transition from a first posture state to a second posture state, and wherein each of the first and second posture states comprises at least one of a posture or a combination of a posture and an activity level. 11. The method of claim 1, further comprising adjusting the therapy based on a plurality of different posture state transitions after time delays following the respective posture state transitions, wherein at least some of the time delays are different for different posture state transitions. 12. The method of claim 11, wherein adjusting the time delay comprises adjusting different time delays for different posture state transitions based on patient adjustments associated with the respective time delays. 13. The method of claim 1, wherein delivering the therapy to the patient comprises delivering the therapy to the patient from an implantable medical device, wherein the therapy comprises electrical stimulation therapy. 14. A medical device comprising: a posture state module that detects a posture state transition of a patient... | Medtronic, Inc. | A61N001/00
US08597187 | B2 | 2013-12-03 | Hand-held electric field imager for measuring the surface topography of mammalian skin and other epithelial structures | 2009-06-15 | The present invention provides a hand-held, noninvasive diagnostic device for measuring the electric fields in mammalian skin and other epithelial structures. The device includes an outer housing that contacts the skin, providing stability and allowing the device to move along with minor movement of the skin. Recessed within the outer housing is a probe that acts as a sensor to measure the electric field in the skin through an aperture in the bottom surface of the outer housing. By applying a series of known voltages while the probe is vibrating, the skin's local surface potential can be measured and the lateral electric field can be calculated from the spatial distribution of surface potential measurements. Active feedback is used to maintain a constant distance between the probe and the skin surface. | 1. A noninvasive diagnostic device for measuring a surface topography associated with an epithelium of a mammal comprising: a hand-held outer housing having a bottom surface that includes an aperture, wherein the bottom surface is adapted to contact an epithelium of a mammal such that the outer housing moves with a movement of the epithelium;a probe recessed within the outer housing and positioned above said aperture, the probe comprising a conducting plate;a vibration unit attached to said probe configured to vibrate said probe over the epithelium of the mammal, wherein the vibration of the probe defines an oscillation path of the probe;a voltage supply configured to create a voltage bias between the probe and the epithelium;a positioning device having a first motor attached to the hand-held outer housing configured to move the probe in an x direction with respect to the outer housing aperture, and a second motor attached to the probe configured to maintain a substantially constant distance in a z direction between the epithelium and a point on the oscillation path of the probe; anda system for recording or displaying travel of the positioning device second motor in the z direction, wherein said travel represents a surface topography of the epithelium. 2. The device as recited in claim 1, wherein the aperture is a slit. 3. The device as recited in claim 2, wherein said slit is about 1 mm wide. 4. The device as recited in claim 1, further comprising: an electrically transparent material covering the aperture. 5. The device as recited in claim 4, wherein the electrically transparent material comprises a polyvinyl material. 6. The device as recited in claim 5, wherein the electrically transparent material comprises polyethylene. 7. The device as recited in claim 1, wherein a point on the oscillation path of the probe is the closest approach of the probe to the epithelium. 8. The device as recited in claim 1, wherein the substantially constant distance is less than or equal to 500 μm. 9. The device as recited in claim 1, wherein the substantially constant distance is 20-500 μm. 10. The device as recited in claim 1, wherein the substantially constant distance is maintained to within a tolerance of about 6 μm. 11. The device as recited in claim 1, further comprising a microcontroller connected to said positioning device and a meter. 12. The device as recited in claim 1, wherein said bottom surface comprises a lens through which the epithelium is visible. 13. The device as recited in claim 12, further comprising a camera that provides a video image of the epithelium through the lens. 14. The device as recited in claim 1, wherein said vibration unit vibrates said probe at a frequency of 800 Hz to 1200 Hz. 15. The device as recited in claim 1, wherein said outer housing comprises a plastic material. 16. The device as recited in claim 1, wherein the first motor configured to move the probe in the x direction is configured to move the vibrat... | BioElectroMed Corp. | A61B005/103
US09545215 | B2 | 2017-01-17 | Apparatus and method for detecting cardiac events | 2008-07-31 | A method and system which includes a minimally invasive, implantable device with a sensor configured for collecting electrical data associated with cardiac performance, a sensor configured for collecting mechanical data associated with cardiac performance, a sensor for collecting optical data associated with cardiac performance, a sensor for collecting biochemical data associated with cardiac performance, and a processor for deriving cardiac conditions and actuating an alarm upon identifying a cardiac event. | 1. An implantable system for monitoring of cardiac conditions comprising: a plurality of sensors configured to collect physiological data associated with a subject of the implantable system at preset time intervals, wherein the collected physiological data includes real-time electrocardiac signal data, heart sound data, activity level data, and tissue perfusion data; anda processing circuitry configured to: compare the electrocardiac signal data with a threshold electrocardiac criteria for indicating a strong likelihood of a cardiac event;trigger an alarm state if the electrocardiac signal data is not within the threshold electrocardiac criteria;determine the current activity level of the subject from the activity level data if the electrocardiac signal data is within the threshold electrocardiac criteria;determine whether the current activity level is below a threshold activity level;compare the tissue perfusion data with a threshold tissue perfusion criteria for indicating a strong likelihood of a cardiac event if the current activity level is determined to be below a threshold activity level;trigger the alarm state if the threshold tissue perfusion data is not within the threshold tissue perfusion criteria; andtrigger the alarm state if the threshold tissue perfusion data is within the threshold tissue perfusion criteria and the heart sound data indicates that S3 and S4 heart sounds are detected,wherein if the alarm state is not triggered, the processor collects the physiological data associated with the subject is at the expiration of the preset time interval. 2. The implantable system of claim 1, wherein the threshold electrocardiac criteria relates to ST segment deviation from preset values. 3. The implantable system of claim 1, wherein at least one of the plurality of sensors is configured to measure the current activity level as a state relating to at least one of physical exertion and body posture of the subject. 4. The implantable system of claim 1, wherein the threshold tissue perfusion criteria relates to an acceptable level of oxygenation in adjacent body tissue of the subject. 5. The implantable system of claim 4, wherein the acceptable level of oxygenation is determined based on the color of the adjacent body tissue of the subject. 6. The implantable system of claim 1, wherein at least one of the plurality of sensors is configured to measure biomarker data relating to biochemical indicators of a cardiac event, and wherein the processing circuitry is configured to trigger the alarm state if the biochemical indicators are detected. 7. The implantable system of claim 1, wherein the plurality of sensors is configured to: collect respiration data relating to the subject's respiration;wherein the processing circuitry is configured to compare the collected respiration data with a threshold respiration criteria for indicating a strong likelihood of a cardiac event in response to the current activity level being below a threshold acti... | Medtronic, Inc. | A61B005/0205 A61N001/37 A61B005/0452 A61B005/145 A61B005/00 A61B005/021 A61B005/024 A61B005/08 A61B005/11 A61B007/00
US20100016700 | A1 | 2010-01-21 | ANALYTE MEASUREMENT AND MANAGEMENT DEVICE AND ASSOCIATED METHODS | 2009-07-17 | Various embodiments of a diabetes management system are provided. One exemplary system may include an analyte measurement device and a therapeutic agent delivery device. The measurement device includes a measurement unit, display, and first wireless module. The therapeutic agent delivery device has a delivery device housing, delivery mechanism disposed in the housing that delivers a dosage of the agent to the user upon actuation by the user or health care provider, and a second wireless module. The second module, automatically, without prompting from a user or any active input or action by the user, transmits a signal to the first wireless module indicative of: (a) type of therapeutic agent delivered; and (b) amount of therapeutic agent delivered to the user; or (c) type of therapeutic agent device from which the therapeutic agent was administered. Also described are diabetes management devices and methods. | 1. A diabetes management system comprising: an analyte measurement device including: a housing;a processor and memory disposed in the housing;a measurement unit in communication with the processor to provide a numerical value representing generally an amount of analyte in body fluids;a display in communication with the processor to display information relating to analyte and the therapeutic agent;a first wireless module coupled to the processor and memory to store data received by the first wireless module in the memory;a therapeutic agent delivery device including: a delivery device housing;a delivery mechanism disposed in the housing that delivers a dosage of the agent to the user upon actuation by the user or health care provider;a second wireless module disposed in the housing, the second wireless module being configured to automatically, without prompting or any action from a user, transmits a signal to the first wireless module indicative of: (a) type of therapeutic agent delivered; and(b) amount of therapeutic agent delivered to the user; or(c) type of therapeutic agent device from which the therapeutic agent was administered. 2. A diabetes management system comprising: an analyte measurement device including: a housing;a processor and memory disposed in the housing;a measurement unit in communication with the processor to provide a numerical value representing generally an amount of analyte in body fluids;a display in communication with the processor to display information relating to analyte and the therapeutic agent;a first wireless module coupled to the processor and memory to store data received by the first wireless module in the memory;a therapeutic agent delivery device including: a delivery device housing;a delivery mechanism disposed in the housing that delivers a dosage of the agent to the user upon actuation by the user or health care provider;a second wireless module disposed in the housing and coupled to the delivery mechanism, the second wireless module, upon prompting or by an action from a user, transmits a signal to the first wireless module indicative of: (a) type of therapeutic agent delivered; and(b) amount of therapeutic agent delivered to the user; or(c) type of therapeutic agent device from which the therapeutic agent was administered; anda health care provider computer in communication with the analyte measurement device to prescribe one of a plurality of therapeutic agent titration protocols to the memory of the analyte measurement device. 3. A diabetes management device comprising: a housing;a processor and memory disposed in the housing, the memory including a plurality of therapeutic administration protocols loaded into the memory from an external source that relates a dosage administration to one or more analyte amount;a measurement unit in communication with the processor to provide a numerical value representing generally an amount of analyte in body fluids;a display in communication with the processor to ... | LifeScan, Inc. | A61B005/00 G06F003/048 H04L009/32 G08B001/08 G06F015/16 G06N005/02
US20100004531 | A1 | 2010-01-07 | MEASUREMENT CATHETER | 2008-07-07 | Enables a measurement catheter configured to hold a package such as an optical tomographic module and/or blood flow velocity module. May hold any package reduced to a size small enough to fit in a blood vessel. Example packages include an optical transmitter and receiver/detectors that enable internal optical tomographic images of vessels to be captured, for example without rotation of the catheter in the vessel. Alternatively or in combination, a thermal package may be coupled to the measurement catheter that includes a thermal element and detector(s) that enable blood flow velocity to be accurately internally measured that does not require a static catheter position. In addition, the measurement catheter may optionally attach to an interchangeable coupler that allows for the introduction of substitution of packages or any type of catheter end assembly to provide rapid deployment of additional surgical or sensory elements. | 1. A measurement catheter comprising: a catheter;a thermal element coupled with said catheter wherein said thermal element is configured to alter blood temperature as said blood flows through a vessel;a thermal sensor coupled with said catheter at an offset from said thermal element wherein said thermal sensor is configured to measure said blood temperature and wherein a blood flow velocity is calculated as said offset divided by a time difference from a point in time when said blood temperature is altered at said thermal element to a point in time wherein an altered blood temperature is measured at said thermal sensor;an optical transmitter coupled with said catheter; and,an optical sensor coupled with said catheter wherein said optical sensor receives coherent radiation transmitted from said optical transmitter to generate an optical tomographic image of said vessel without rotation of said catheter. 2. The measurement catheter of claim 1 wherein said thermal element is a heating element. 3. The measurement catheter of claim 1 wherein said thermal element is a cooling element. 4. The measurement catheter of claim 1 wherein said thermal sensor comprises a plurality of sensors offset from said thermal element. 5. The measurement catheter of claim 1 wherein said optical transmitter is a broadband light source. 6. The measurement catheter of claim 1 wherein said optical transmitter is a diode. 7. The measurement catheter of claim 1 wherein said optical transmitter is a superluminescent diode. 8. The measurement catheter of claim 1 further comprising an interchangeable coupler configured to mechanically couple a package to said measurement catheter. 9. A measurement catheter comprising: a catheter;a thermal element coupled with said catheter wherein said thermal element is configured to alter blood temperature as said blood flows through a vessel; and,a thermal sensor coupled with said catheter at an offset from said thermal element wherein said thermal sensor is configured to measure said blood temperature and wherein a blood flow velocity is calculated as said offset divided by a time difference from a point in time when said blood temperature is altered at said thermal element to a point in time wherein an altered blood temperature is measured at said thermal sensor. 10. The measurement catheter of claim 9 wherein said thermal element is a heating element. 11. The measurement catheter of claim 9 wherein said thermal element is a cooling element. 12. The measurement catheter of claim 9 wherein said thermal sensor comprises a plurality of sensors offset from said thermal element. 13. The measurement catheter of claim 9 further comprising an interchangeable coupler configured to mechanically couple with said thermal element and said thermal sensor. 14. A measurement catheter comprising: a catheter;an optical transmitter coupled with said catheter; and,an optical sensor coupled with said catheter wherein said optical sensor receives cohe... | - | A61B005/026 A61B005/05
US20100022848 | A1 | 2010-01-28 | PHYSIOLOGICAL STATUS MONITORING/POSITIONING SYSTEM | 2008-08-27 | A physiological status monitoring/positioning system is disclosed, which comprises: a physiological status sensing device, being configured with at least a sensing unit and a signal transmitting unit to be used for detecting and transmitting a physiological attribute of a living entity; a positioning device, being configured with a second signal transceiving unit to be used for transmitting a second signal to a positioning system where the location of the living entity is identified; a calculation device, being configured with a first signal transceiving unit for receiving the physiological attribute while converting the same into a bio index to be transmitted; a signal transceiver, for transmitting/receiving the bio index; and a display device, for displaying the identified location of the living entity as well as the bio index. With the aforesaid system, the physiological status of the living entity as well as its location can be monitored at the same time so that as soon as an abnormal physiological status of the living entity is detected, an immediate action can be taken in response to the emergency. | 1. A physiological status monitoring/positioning system, comprising: a physiological status sensing device, being configured with at least a sensing unit and a signal transmitting unit to be used for detecting and transmitting a physiological attribute of a living entity;a positioning device, being configured with a second signal transceiving unit to be used for transmitting a second signal to a positioning system where the location of the living entity is identified;a calculation device, being configured with a first signal transceiving unit for receiving the physiological attribute while converting the same into a bio index to be transmitted;a signal transceiver, for transmitting/receiving the bio index; anda display device, for displaying the identified location of the living entity as well as the bio index. 2. The system of claim 1, wherein the second signal specifies a coordinate in a coordinate system. 3. The system of claim 1, wherein the second signal is substantially a positioning signal. 4. The system of claim 2, wherein the physiological attribute, the coordinate and the bio index are transmitted through a communication network system. 5. The system of claim 3, wherein the physiological attribute, the coordinate and the bio index are transmitted through a communication network system. 6. The system of claim 4, wherein the signal transceiver and the display unit are configured in a handheld electronic device while enabling the signal transceiver to transmit/receive the coordinate and the bio index through the communication network system and thus to display the same on the display unit. 7. The system of claim 5, wherein the signal transceiver and the display unit are configured in a handheld electronic device while enabling the signal transceiver to transmit/receive the coordinate and the bio index through the communication network system and thus to display the same on the display unit. 8. The system of claim 6, wherein the handheld electronic device is designed to form a closed circuit with other handheld electronic devices in a manner that they are enabled to receive the coordinates and the bio indexes of different living entities from one another. 9. The system of claim 7, wherein the handheld electronic device is designed to form a closed circuit with other handheld electronic devices in a manner that they are enabled to receive the coordinates and the bio indexes of different living entities from one another. 10. The system of claim 4, wherein the calculation device is configured in a center operation system while enabling the same to receive the physiological attribute and then converting the received physiological attribute in to the bio index. 11. The system of claim 5, wherein the calculation device is configured in a center operation system while enabling the same to receive the physiological attribute and then converting the received physiological attribute in to the bio index. 12. The system of claim 6, wherein t... | IMRISS TECHNOLOGY CORPORATION | A61B005/00
US20100030059 | A1 | 2010-02-04 | Method for Measuring Information of Biological Systems | 2008-02-14 | The invention relates to a method for measuring information of biological systems. The aim of the invention is to receive signals using less energy. To achieve this, random generators are used as receivers (B) of low-energy quanta, since the random generators can be regarded and implemented as antennae and receivers of signals of this type. The extensive natural transmission range of low-energy quanta can also be used to receive information from spatially remote systems. | 1-15. (canceled) 16. A method for measuring information of biological systems, said method comprising steps of providing suitable receivers formed as noise generators for receiving and evaluating low-energy quanta LEQ with a frequency in the range between 1 Hz and 100 Hz or very low-energy quanta LSTEQ with a frequency of less than 1 Hz,receiving said low-energy quanta or very low-energy quanta by said receivers,evaluating said received low-energy quanta or very low-energy quanta with the physical relationship between frequency and energy being used in order to determine the energy of the low-energy quanta or very low-energy quanta to be received and in order to use the noise generators as receivers or transmitters of low-energy quanta or very low-energy quanta,time sampling the noise signal generated when receiving low energy quanta or very low energy quanta,selectively filtering out the received low energy quanta or very low energy quanta from the noise signal. 17. (canceled) 18. The method as claimed in claim 16, wherein the received quanta originate from human beings. 19. The method as claimed in claim 16, wherein the received quanta originate from natural systems such as animals, plants, minerals or other materials. 20. The method as claimed in claim 18, wherein the receivers, which are based on noise generators, of low-energy quanta are used for diagnosis of illnesses, or for diagnosis of mental states. 21. The method as claimed in claim 18, wherein the receivers, which are based on noise generators, of low-energy quanta are used for communication with the very seriously ill. 22. The method as claimed in claim 18, wherein the receivers, which are based on noise generators, of low-energy quanta are used to determine the truth of human statements. 23. The method as claimed in claim 16, wherein the received quanta originate from systems which are physically a long distance away, thus making it possible to carry out remote diagnoses of biological systems or remote monitoring of technical systems and installations. 24. The method as claimed in claim 16, wherein the receivers, which are based on noise generators, of low-energy quanta are used for exploration of natural resources. 25. The method as claimed in claim 16, wherein the receivers, which are based on noise generators, of low-energy quanta are used for determination of materials, and these materials can thus be located specifically by calibration of the receivers for the corresponding materials, which makes it possible to select those quanta which those materials permanently emit, from the range of signals. 26. The method as claimed in claim 16, wherein the receivers, which are based on noise generators, of low-energy quanta are used for data communication, in that addressing and calibration are carried out between the transmitters and receivers of quanta, as a result of which the receiver can filter out the quanta sent via the transmitter from the information mixture of its no... | TECDATA AG | A61B005/05 A61B005/00 G01V009/00
US20100010379 | A1 | 2010-01-14 | DEVICES AND METHOD FOR MONITORING THE FORM OF THREE-DIMENSIONAL OBJECTS | 2005-12-13 | Devices for monitoring the deformation of three-dimensional objects thanks to the presence of sensors integrated in the device are described. | 1-14. (canceled) 15. Device for monitoring the form of a three-dimensional object or mechanical deformation field, comprising a flexible substrate;a plurality of sensors being provided on the substrate; andwires provided to connect the sensors with acquisition electronics;wherein the sensors are made of electrically conductive elastomers having piezoresistive effects if mechanically stimulated, wherein the plurality of sensors comprise groups of sensors being formed as consecutive portions of a conductive strip of said conductive elastomer material, wherein the consecutive portions are defined through the wires provided on the substrate contacting the strip, and wherein the wires are made of electrically conductive elastomers. 16. Device according to claim 1, wherein the flexible substrate is extendible. 17. Device according to claim 1, wherein the sensors are distributed on the substrate in a uniform manner and according to the desired topology. 18. Device according to claim 1, wherein the wires are distributed on the substrate according to the desired geometries using appropriate templates. 19. Device according to claim 1, wherein the flexible substrate consists of a fabric in natural or artificial fibers, preferably elastic fibers. 20. Device according to claim 1, in which said elastomers are chosen from the group of elastomers comprising: intrinsic conductive polymers (such as polypyrrole, polyaniline and their derivates) or loaded (with carbon, graphite or metal powders) polymers (such as silicon, natural rubber or polyurethane). 21. Device according to claim 6, in which said elastomer is Elastosil LR 3162 AB, manufactured by Wacker Ltd. 22. Template for the production of a device for monitoring the form of a three-dimensional object or mechanical deformation field, the device having a flexible substrate, a plurality of sensors being provided on the substrate and wires provided to connect the sensors with acquisition electronics; the template comprising a sheet of adhesive material on which a vector drawing is printed, starting from a drawing of the arrangement of the group comprising sensors and wires on a virtual model so that all the factors necessary for their construction can be taken into consideration. 23. Process for preparing a device for monitoring the form of a three-dimensional object or mechanical deformation field using a template, comprising the steps of: cutting the desired template onto an adhesive sheet that is placed on the substrate in which the sensors are to be integrated;application of the elastomer solution to the fabric to be treated by spreading it evenly over the template so that it can be deposited on the substrate along the cuts in the template;placing the coated substrate and template in an oven at a temperature of approximately 120° Celsius for approximately 15 minutes to achieve the evaporation of the solvent, the vulcanization and subsequent adhesion of the elastomer to the substrate;cooling the mixt... | - | A61B005/103
US20100036225 | A1 | 2010-02-11 | SYSTEMS AND METHODS FOR REPLACING SIGNAL ARTIFACTS IN A GLUCOSE SENSOR DATA STREAM | 2009-10-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. A method for processing data from a continuous glucose sensor, the method comprising: generating a data stream from a continuous glucose sensor;detecting a signal artifact in the data stream comprising transient non-glucose dependent signal loss; andaugmenting the data stream to reduce or replace the detected signal loss. 2. The method of claim 1, wherein the augmenting comprises retaining at least one high data stream value associated with a time period and discarding at least one lower signal value associated with the time period. 3. The method of claim 2, comprising discarding at least one high data stream value. 4. The method of claim 2, comprising averaging a plurality of retained high data stream values. 5. The method of claim 2, wherein the at least one high data stream value is retained based at least in part on a preselected ratio. 6. The method of claim 5, wherein the preselected ratio is between the at least one high data stream value and a projected data stream value. 7. The method of claim 1, further comprising replacing at least a portion of the data stream that trends downwardly from an initial value, wherein the downward trend is greater than a predetermined threshold. 8. The method of claim 1, comprising high pass filtering. 9. The method of claim 1, wherein augmenting the data stream comprises generating a filtered signal and a compensation signal. 10. The method of claim 1, comprising scaling the data stream. 11. The method of claim 10, wherein the scaling comprises accessing a look-up table. 12. The method of claim 10, wherein a data stream filtering process comprises the scaling. 13. The method of claim 10, wherein the augmenting comprises the scaling. 14. The method of claim 1, wherein the signal loss has a higher amplitude than system noise. 15. A glucose sensor system for monitoring a glucose concentration in a host, the system comprising: a continuous glucose sensor configured to produce a data stream;a signal artifact detection module configured to detect a signal artifact in the data stream comprising transient non-glucose dependent signal loss; anda signal estimation module configured to augment the data stream to reduce or replace the signal loss. 16. The system of claim 15, wherein the signal estimation module is configured to retain at least one high signal value associated with a time period and discard at least one lower signal values associated with the time period. 17. The system of claim 15, wherein the signal estimation module is configured to replace at least a portion of the data stream that trends downwardly from an initial value, wherein the downward trend is greater than a predetermined threshold. 18. The system of claim 15, comprising a high pass filter. 19. The system of claim 15, wherein the signal estimation module is configured to generate a filtered signal and a compensation signal. 20. The system of claim 15, wherein the signal estimation module is configured to scale the data... | DexCom, Inc. | A61B005/145 G06F019/00
US20100022898 | A1 | 2010-01-28 | MEASUREMENT OF CARDIAC OUTPUT AND BLOOD VOLUME BY NON-INVASIVE DETECTION OF INDICATOR DILUTION | 2009-09-22 | A system for evaluating the cardiovascular system parameters using indicator dilution and non-invasive or minimally invasive detection methods is disclosed. Intravascular indicators are stimulated, and emissions patterns detected for computation of cardiac output, cardiac index, blood volume and other indicators of cardiovascular health. | 1. A method of measuring a parameter of a cardiovascular system in a subject in absolute physical units comprising: a. determining the volume of blood in the cardiovascular system;b. administering to the cardiovascular system a detectable amount of an indicator that fluoresces when illuminated;c. illuminating the indicator;d. measuring the intensity of the fluorescence of the indicator while it is illuminated; ande. determining the magnitude of the parameter in absolute physical units based on the measured intensity of the fluorescence, the administered amount of the indicator and the determined blood volume. 2. The method of claim 1, further including: determining a conversion factor that converts the measured intensity of the fluorescence to the concentration of indicator in the cardiovascular system; and wherein determining the magnitude of the parameter is also based on the conversion factor. 3. The method of claim 1, further including determining the concentration of the indicator in the cardiovascular system based on the amount of the indicator that was administered and the determined volume of blood. 4. The method of claim 3, wherein: measuring the intensity of the fluorescence includes measuring the intensity as a function of time;measuring the intensity of the fluorescence as a function of time results in an initial fast rising portion and a subsequent slow decay portion; anddetermining the conversion factor, wherein the determining the conversion factor includes determining the point on the fast rising portion that is intercepted by an extrapolation of a portion of the slow decay portion. 5. The method of claim 4, wherein the determining the conversion factor is based on the ratio of the determined concentration to the measured intensity of the fluorescence at the intercepted point. 6. A system of measuring a parameter of a cardiovascular system in a subject in absolute physical units comprising: a. an illumination source configured to illuminate an administered indicator in the cardiovascular system of the subject to emit fluorescence;b. a photodetector configured to measure the intensity of the fluorescence from the indicator while it is illuminated;c. a computing system configured to determine the magnitude of the parameter in absolute physical units based on the measured intensity of the fluorescence, the amount of the administered indicator and a determined blood volume of the subject. 7. The system of claim 6, wherein the computing system is configured to determine a conversion factor that converts the measured intensity of the fluorescence to the concentration of the indicator in the cardiovascular system. 8. The system of claim 6, wherein the computing system is further configured to measure the intensity as a function of time so as to create an intensity curve having an initial fast rising portion and a subsequent slow decay portion; and to determine the point on a fast rising portion of the intensity curve that is inter... | Alfred E. Mann Institute for Biomedical Engineering at the University of Southern California | A61B005/02
US20100010323 | A1 | 2010-01-14 | Temperature-compensated in-vivo sensor | 2009-09-21 | A reagent matrix composition disposed on an electrically conductive electrode to form an in-vivo sensor for a predefined analyte includes a first hydrogel layer containing an enzyme that is a substrate for the analyte adjacent the electrically conductive electrode and a composite membrane layer disposed onto the first hydrogel layer where the composite layer includes a membrane hydrogel containing a plurality of microspheres. The plurality of microspheres is made of a material having no or little permeability to the analyte and a substantially high permeability to oxygen. | 1. An in-vivo sensor for measuring an analyte, the sensor comprising: a base insulating layer;an electrically conductive electrode disposed on the base insulating layer;a semi-permeable layer disposed over the electrode;a first hydrogel layer containing an enzyme that is a substrate for the analyte disposed on the semi-permeable layer;a composite membrane layer disposed on the first hydrogel layer wherein the composite layer includes a hydrogel containing a plurality of microspheres, the plurality of microspheres being made of a material having no or little permeability to the analyte and substantially high permeability to oxygen. 2. The sensor of claim 1 further comprising a second hydrogel layer disposed over the composite membrane layer. 3. The sensor of claim 2 wherein the second hydrogel layer further includes a catalase. 4. The sensor of claim 1 wherein the material of the microspheres is polydimethylsiloxane. 5. The sensor of claim 1 wherein the first hydrogel layer is one of a polyurethane or a poly-2-hydroxyethyl methylacrylate. 6. The sensor of claim 2 wherein the second hydrogel layer is one of polyurethane or poly-2-hydroxyethyl methacrylate. 7. The sensor of claim 1 wherein the composite membrane layer has a ratio of the quantity of microspheres to the quantity of membrane hydrogel selected for a predefined dynamic measurement range of the analyte. 8. A reagent matrix composition disposed on an electrically conductive electrode to form an in-vivo sensor for a predefined analyte, the composition comprising: a first hydrogel layer containing an enzyme that is a substrate for the analyte adjacent the electrically conductive electrode; anda composite membrane layer disposed onto the first hydrogel layer wherein the composite layer includes a membrane hydrogel containing a plurality of microspheres, the plurality of microspheres being made of a material having no or little permeability to the analyte and a substantially high permeability to oxygen. 9. The composition of claim 8 further comprising a semi-permeable layer between the first hydrogel layer and the electrically conductive electrode. 10. The composition of claim 8 further comprising a second hydrogel layer disposed onto the composite membrane layer. 11. The composition of claim 10 wherein the second hydrogel layer contains a catalase. 12. The composition of claim 8 wherein the first hydrogel layer, the membrane hydrogel and the second hydrogel layer are one of polyurethane or poly-2-hydroxyethyl methacrylate. 13. The composition of claim 8 wherein the material of the microspheres is polydimethylsiloxane. 14. The composition of claim 8 wherein the composite membrane layer has a ratio of the quantity of microspheres to the quantity of membrane hydrogel selected for a predefined dynamic measurement range of the analyte. 15. A method of making a sensor for in-vivo measurement of an analyte, the method comprising: obtaining an electrically conductive electrode;disposing... | NOVA BIOMEDICAL CORPORATION | A61B005/00
US20100036269 | A1 | 2010-02-11 | Circulatory monitoring systems and methods | 2008-08-07 | Systems and methods are described for obtaining and acting upon information indicative of circulatory health and related phenomena in human beings or other subjects. | 1. A medical or veterinary monitoring system comprising: a physiological support, including at least circuitry for detecting an intensive property of at least an internal portion of a limb of a mammal; andcircuitry for transmitting an indication of whether an apparent trend in the intensive property exceeds a temporal threshold. 2. The medical or veterinary system of claim 1, in which the circuitry for detecting an intensive property of at least an internal portion of a limb of a mammal comprises: circuitry for detecting a thermal normalization rate decrease. 3. The medical or veterinary system of claim 1, in which the circuitry for detecting an intensive property of at least an internal portion of a limb of a mammal comprises: circuitry for activating an electromagnetic radiation emitter adjacent the limb of the mammal. 4. The medical or veterinary system of claim 1, in which the circuitry for transmitting an indication of whether an apparent trend in the intensive property exceeds a temporal threshold comprises: circuitry for causing a determination of whether an apparent transition in the intensive property satisfies a directional criterion. 5. The medical or veterinary system of claim 1, in which the circuitry for transmitting an indication of whether an apparent trend in the intensive property exceeds a temporal threshold comprises: circuitry for causing a determination of whether an apparent transition in the intensive property exceeds a size threshold. 6. The medical or veterinary system of claim 1, in which the circuitry for transmitting an indication of whether an apparent trend in the intensive property exceeds a temporal threshold comprises: circuitry for causing a determination of whether a signal indicates a temporal reflectivity drift affecting energy that passes into the limb. 7. The medical or veterinary system of claim 1, in which the circuitry for transmitting an indication of whether an apparent trend in the intensive property exceeds a temporal threshold comprises: circuitry for causing a determination of whether a signal indicates a temporal transmissivity drift affecting energy that passes into the limb. 8. The medical or veterinary system of claim 1, in which the circuitry for transmitting an indication of whether an apparent trend in the intensive property exceeds a temporal threshold comprises: circuitry for causing a determination of whether a signal manifests a temporal transmissivity drift across one or more microwave or radio frequency ranges. 9. The medical or veterinary system of claim 1, in which the circuitry for transmitting an indication of whether an apparent trend in the intensive property exceeds a temporal threshold comprises: circuitry for determining whether a duration of a temporal drift meets or exceeds the temporal threshold, wherein the temporal threshold exceeds a minute. 10. The medical or veterinary system of claim 9, in which the circuitry for transmitting an indication of whether an appar... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/02
US07668737 | B2 | 2010-02-23 | Method and system for interfacing with a multi-level data structure | 2003-09-12 | The present invention provides a method and system for interfacing with a multi-level data structure by selecting a concept object stored in the multi-level data structure, displaying a first image representing the selected concept object, displaying one or more second images generally above the first image, and displaying a first connector connecting each second image to the first image. Whenever the selected concept object has one or more child concept objects, one or more third images are displayed generally below the first image, and a second connector is displayed connecting each third image to the first image. Whenever the selected concept object has one or more lateral concept objects, one or more fourth images are displayed generally on one or both sides of the first image, and a third connector is displayed connecting each fourth image to the first image. | 1. In a computer system, a computer-implemented method for displaying and creating relationships between different medical sources on a display device, the computer-implemented method comprising: receiving a selection of a medical concept with a computer for display on the display device;in response to the receiving the selection of the medical concept, the computer: displaying a first image in a first window with the display device comprising an alphanumeric string representing the selected medical concept;displaying one or more second images with the display device and along one or more respective geometrical rays originating from a central region of the first image, each second image comprising an alphanumeric string representing a parent concept of the selected medical concept and displaying a first symbol on the display device along each respective geometrical ray originating from the central region of the first image;displaying a billing code comprising an alphanumeric string in a second window adjacent to the first window with the display device, the billing code originating from a first medical source associated with the selected medical concept;displaying a medical code adjacent to the billing code in the second window with the display device, the medical code comprising an alphanumeric string originating from a second medical source that is different from the first medical source and is associated with the selected medical concept;receiving input comprising alphanumeric text through a third window on the display device; andmodifying, removing or creating relationships between said medical concepts. 2. The computer-implemented method of claim 1, wherein the first medical source comprises at least one of International Statistical Classification of Disease and Related Health Problems (ICD) and Physicians' Current Procedural Terminology (CPT) billing codes. 3. The computer-implemented method of claim 1, wherein the second medical source comprises at least one of systemized nomenclature medical reference terminology (SNOMED RI), MeSH, UMLS CUI, and pharmacy terminology. 4. The computer-implemented method of claim 1, wherein the selected medical concept is a first medical concept, the computer implemented method further comprising: receiving input comprising a second medical concept other than the selected first medical concept;creating an association between the received input comprising the second medical concept and the selected first medical concept; andstoring the association between the received input comprising the second medical concept and the selected first medical concept in memory. 5. The computer-implemented method of claim 4, wherein the received medical concept is a child concept relative to the selected medical concept. 6. The computer-implemented method of claim 1, further comprising: receiving a medical term;creating an association between the received medical term and the selected medical concept;storing the associatio... | Health Language, Inc. | A61B005/00
US07652259 | B2 | 2010-01-26 | Apparatus and methods for imaging and attenuation correction | 2003-11-04 | Imaging apparatus, is provided, comprising a first device, for obtaining a first image, by a first modality, selected from the group consisting of SPECT, PET, CT, an extracorporeal gamma scan, an extracorporeal beta scan, x-rays, an intracorporeal gamma scan, an intracorporeal beta scan, an intravascular gamma scan, an intravascular beta scan, and a combination thereof, and a second device, for obtaining a second, structural image, by a second modality, selected from the group consisting of a three-dimensional ultrasound, an MRI operative by an internal magnetic field, an extracorporeal ultrasound, an extracorporeal MRI operative by an external magnetic field, an intracorporeal ultrasound, an intracorporeal MRI operative by an external magnetic field, an intravascular ultrasound, and a combination thereof, and wherein the apparatus further includes a computerized system, configured to construct an attenuation map, for the first image, based on the second, structural image. Additionally, the computerized system is configured to display an attenuation-corrected first image as well as a superposition of the attenuation-corrected first image and the second, structural image. Furthermore, the apparatus is operative to guide an in-vivo instrument based on the superposition. | 1. Imaging apparatus, comprising: a first device which obtains a first image, by a first modality, said image being an ionizing radiation image, wherein said first image is registered to a system of coordinates;a second device which obtains a second, structural image, by a second modality, said structural image being an ultrasonic image; anda computerized system, which comprises: a registrator for co-registering said second, structural image to said system of coordinates, andan attenuation-instruction generator configured to compute a set of attenuation instructions for said first image, based at least on non-uniformities in said second, structural image. 2. The imaging apparatus of claim 1, wherein said computerized system is further configured to compute, based on said a set of attenuation instructions an attenuation-corrected image of said first image. 3. The imaging apparatus of claim 2, wherein said computerized system is further configured to display a superposition of said attenuation-corrected first image and said second, structural image. 4. The imaging apparatus of claim 3, wherein said apparatus further includes an instrument, registered to said system of coordinates and visible on at least one of said first image and said second, structural image, and wherein said computerized system is further configured to guide said instrument in-vivo, based on said superposition. 5. The imaging apparatus of claim 1, wherein said registrator for co-registering said second, structural image to said system of coordinates relies on that said first and second devices share a single position-registration device, for co-registering said second, structural image to said system of coordinates. 6. The imaging apparatus of claim 1, wherein said registrator for co-registering said second, structural image to said system of coordinates relies on that said first and second devices have substantially equal position-registration devices, for co-registering said second, structural image to said system of coordinates. 7. The imaging apparatus of claim 1, wherein said registrator for co-registering said second, structural image to said system of coordinates relies on fiduciary marks visible both on said first image and on said second, structural image, for co-registering said second, structural image to said system of coordinates. 8. The imaging apparatus of claim 1, further comprising an ultrasound transducer operative for focused ablation. 9. The imaging apparatus of claim 1, designed as a rectum probe. 10. The imaging apparatus of claim 1, designed as an endoscope probe. 11. The imaging apparatus of claim 1, designed to be inserted through a trucar valve. 12. The imaging apparatus of claim 1, designed to be mounted on a resectoscope. 13. The imaging apparatus of claim 1, designed to be inserted in a catheter. 14. The imaging apparatus of claim 1, designed for intravascular imaging. 15. The imaging apparatus of claim 1, designed as a handheld, extrac... | Spectrum Dynamics LLC | G01T001/24 A61B005/05
US20100016682 | A1 | 2010-01-21 | PATIENT MONITORING SYSTEM AND METHOD | 2007-12-18 | The present invention relates to a patient monitoring system (1). Further the invention relates to a method for monitoring a patient and to a computer program for a patient monitoring system (1). In order to provide a reliable technique for reducing the number of false alarms in a patient monitoring system a patient monitoring system (1) is suggested, which comprises a first sensor device (2) adapted to acquire (101) a first patient signal (10) corresponding to a first physiological parameter of the patient, a second sensor device (3, 4, 5) adapted to acquire (102, 103, 104) a second patient signal (11, 12, 13) corresponding to a second physiological parameter of the patient, said second patient signal (11, 12, 13) comprising an overlay signal (14, 15, 16) caused by the first physiological parameter of the patient, a processing device (7) adapted to determine (105) from the first patient signal (10) a first value of the first physiological parameter of the patient, to determine (106, 107, 108) from the overlay signal (14, 15, 16) of the second patient signal (11, 12, 13) a second value of the first physiological parameter of the patient, and to analyze (109) the first and second value of the first physiological parameter of the patient, and a control device (8) adapted to control (110) a patient monitor alarm system (9) depending on the result of the analysis. | 1. A patient monitoring system, comprising a first sensor device adapted to acquire a first patient signal corresponding to a first physiological parameter of the patient,a second sensor device adapted to acquire a second patient signal corresponding to a second physiological parameter of the patient, said second patient signal comprising an overlay signal caused by the first physiological parameter of the patient,a processing device adapted to determine from the first patient signal a first value of the first physiological parameter of the patient, to determine from the overlay signal of the second patient signal a second value of the first physiological parameter of the patient, and to analyze the first and second value of the first physiological parameter of the patient, anda control device adapted to control a patient monitor alarm system (9) depending on the result of the analysis. 2. The system as claimed in claim 1, wherein the first physiological parameter of the patient is the patient's respiration and the overlay signal of the second patient signal is a respiration overlay signal. 3. The system as claimed in claim 1, wherein the second physiological parameter of the patient is the patient's invasive blood pressure. 4. The system as claimed in claim 1, wherein the second physiological parameter of the patient is the pulsation of the patient's blood. 5. The system as claimed in claim 1, wherein the second physiological parameter of the patient is the electrical voltage in the patient's heart. 6. A method of monitoring a patient, the method comprising the steps of acquiring a first patient signal corresponding to a first physiological parameter of the patient,determining from the first patient signal a first value of the first physiological parameter of the patient,acquiring a second patient signal corresponding to a second physiological parameter of the patient, said second patient signal comprising an overlay signal caused by the first physiological parameter of the patient,determining from the overlay signal of the second patient signal a second value of the first physiological parameter of the patient,analyzing the first and second value of the first physiological parameter of the patient, andcontrolling the release of a monitoring alarm depending on the result of the analyzing step. 7. The method as claimed in claim 6, wherein in case of the first value being critical an alarm is suppressed, if the second value is non-critical. 8. A computer program for a patient monitoring system, which is adapted to acquire a first patient signal corresponding to a first physiological parameter of the patient, and a second patient signal corresponding to a second physiological parameter of the patient, said second patient signal comprising an overlay signal caused by the first physiological parameter of the patient, the program comprising: computer instructions to determine from the first patient signal a first value of the first physiologica... | KONINKLIJKE PHILIPS ELECTRONICS N. V. | A61B005/00
US07641617 | B2 | 2010-01-05 | Wrist blood pressure gauge | 2007-05-23 | A convenient, inexpensive wrist blood pressure gauge whereby a cuff can easily be adjusted to the height of the heart merely by adding a simple structure. A lenticular sheet (display body (14)) whose visible image varies according to an angle in relation to a sight line (S) of a measurement subject is provided to a wrist blood pressure gauge (1) worn on a wrist (Mt) of the measurement subject (M), wherein a specific image is visible to the eyes (Ma) of the measurement subject when the wrist is set at the correct height in relation to the heart (H), and a different image is visible to the eyes of the measurement subject when the wrist is not in an appropriate height range. | 1. A wrist blood pressure gauge comprising: a cuff configured to be worn on a wrist of a measurement subject;a blood pressure gauge body that calculates a blood pressure value, said gauge body having a display unit for displaying the blood pressure value;a lenticular sheet whose visible predetermined specific content is displayed corresponding to an angle in relation to a sight line of the measurement subject is provided to or in a vicinity of the cuff worn on a wrist of the measurement subject; andan operating button for starting blood pressure measurement is provided on an exterior case of the blood pressure gauge body that is integrally provided to said cuff; andsaid lenticular sheet is disposed on a pressing surface of the operating button, whereinthe lenticular sheet is an additional structure different from the display unit;the lenticular sheet of predetermined specific content is visible to the measurement subject viewing the lenticular sheet when said cuff worn on the wrist is set in an appropriate height range in relation to the measurement subject's heart; andthe lenticular sheet displays another previously defined content other than the specific content to the measurement subject when the cuff is not in an appropriate height range. 2. The wrist blood pressure gauge according to claim 1, wherein said lenticular sheet presents a different color according to a difference in an angle with respect to a sight line of a measurement subject between a state in which said cuff is in an appropriate height range and a state in which said cuff is not in an appropriate height range. 3. The wrist blood pressure gauge according to claim 1, wherein said lenticular sheet is affixed in a location near the cuff and visible to the measurement subject. 4. A wrist blood pressure gauge comprising: a cuff configured to be worn on a wrist of a measurement subject;a blood pressure gauge body that calculates a blood pressure value, said gauge body having a display unit for displaying the blood pressure value; anda lenticular sheet whose visible predetermined specific content is displayed corresponding to an angle in relation to a sight line of the measurement subject is provided to or in a vicinity of the cuff worn on a wrist of the measurement subject; andan operating button for starting blood pressure measurement is provided on an exterior case of the blood pressure gauge body that is integrally provided to said cuff; andsaid lenticular sheet is disposed on a pressing surface of the operating button whereinthe lenticular sheet is an additional structure different from the display unit;the lenticular sheet of predetermined specific content is visible to the measurement subject viewing the lenticular sheet when said cuff worn on the wrist is set in an appropriate height range in relation to the measurement subject's heart;the lenticular sheet displays another previously defined content other than the specific content to the measurement subject when the cuff is... | Japan Precision Instruments Inc. | A61B005/02
US07647083 | B2 | 2010-01-12 | Multiple wavelength sensor equalization | 2006-03-01 | A physiological sensor has intensity compensation introduced along an optical path from emission to detection so as to compensate for unequal tissue attenuation as a function of wavelength. The sensor has emitters configured to transmit optical radiation having multiple wavelengths into a tissue site. At least one detector is capable of receiving the optical radiation after tissue attenuation. An equalization is capable of compensating optical radiation intensity so as to account for differences in tissue attenuation across at least a portion of the multiple wavelengths. | 1. A physiological sensor having optical intensity compensation introduced along an optical path from emission to detection so as to compensate for unequal tissue attenuation as a function of wavelength by a tissue site, the sensor comprising: a plurality of emitters configured to transmit optical radiation having a plurality of wavelengths into the tissue site;at least one detector capable of receiving the optical radiation after tissue attenuation; andan optical filter, wherein the optical filter attenuates at least a portion of a wavelength spectrum so as to compensate for an imbalance in an attenuation of the plurality of wavelengths by the tissue site. 2. The physiological sensor according to claim 1, wherein the optical filter has a transfer function adapted to attenuate only selected wavelengths of the optical radiation. 3. The physiological sensor according to claim 1, wherein the optical filter is positioned proximate the emitters. 4. The physiological sensor according to claim 1 wherein the optical filter is positioned proximate the at least one detector. 5. The physiological sensor according to claim 1, wherein said optical filter includes a filtering medium and a subset of the emitters having IR wavelengths are encapsulated with the filtering medium. 6. The physiological sensor according to claim 5, wherein the emitters of the subset are individually encapsulated. 7. A physiological sensor comprising: an emitter means for emitting optical radiation having a plurality of wavelengths into a tissue site;a detector means for detecting the optical radiation after tissue attenuation, wherein the tissue attenuates each of the plurality of wavelengths differently; andan equalization means for attenuating at least a portion of a wavelength spectrum so as to compensate for an imbalance in the attenuation of the plurality of wavelengths by the tissue. 8. The physiological sensor according to claim 7, wherein the equalization means in an optical filter configured to attenuate optical radiation intensity at selected wavelengths. 9. A physiological sensor having optical intensity compensation introduced along an optical path from emission to detection so as to compensate for unequal tissue attenuation as a function of wavelength, the sensor comprising: first plurality of emitters configured to transmit a first plurality of wavelengths into a tissue site, the tissue site attenuating the first plurality of wavelengths at at least a first attenuation level;a second plurality of emitters configured to transmit a second plurality of wavelengths into the tissue site, the tissue site attenuating the second plurality of wavelengths at at least a second attenuation level different than the attenuation level of the first plurality of wavelengths;at least one detector configured to receive the optical radiation after tissue attenuation; andan optical filtering medium encapsulating the first plurality of emitters but not the second plurality of emitter... | Masimo Laboratories, Inc. | A61B005/1455
US07657306 | B2 | 2010-02-02 | Non-invasive method and device to monitor cardiac parameters | 2005-11-03 | A method of and a device for non-invasively measuring the hemodynamic state of a subject or a human patient involve steps and units of non-invasively measuring cardiac cycle period, electrical-mechanical interval, mean arterial pressure, and ejection interval and converting the measured electrical-mechanical interval, mean arterial pressure and ejection interval into the cardiac parameters such as Preload, Afterload and Contractility, which are the common cardiac parameters used by an anesthesiologist. The converted hemodynamic state of a patient is displayed on a screen as a three-dimensional vector with each of its three coordinates respectively representing Preload, Afterload and Contractility. Therefore, a medical practitioner looks at the screen and quickly obtains the important and necessary information. | 1. A system for retrofitting existing non-invasive cardiac parameter measuring devices to generate invasive cardiac analogues, comprising: an interface unit for receiving predetermined non-invasive cardiac parameters of a subject from the existing non-invasive cardio monitoring devices; anda conversion unit connected to said interface unit for converting the non-invasive cardiac parameters into a plurality of the invasive cardiac analogues based upon a set of predetermined conversion equations, wherein said conversion unit outputs the predetermined invasive cardiac analogues including preload as denoted by P, afterload as denoted by A and contractility as denoted by C. 2. The system for retrofitting existing non-invasive cardio monitoring devices according to claim 1 wherein the existing non-invasive cardiac parameter measuring devices measure the predetermined non-invasive cardiac parameters from a human. 3. The system for retrofitting existing non-invasive cardio monitoring devices according to claim 1 wherein the existing non-invasive cardiac parameter measuring devices measure the predetermined non-invasive cardiac parameters from an animal. 4. The system for retrofitting existing non-invasive cardio monitoring devices according to claim 1 wherein the existing non-invasive cardiac parameter measuring unit further comprises a heart rate monitor for measuring heart rate as denoted by HR, a vibration sensing device for measuring an ejection interval as denoted by EI and a mechanical event M of an electrical-mechanical interval as denoted by E-M, a blood pressure measuring device for measuring mean arterial pressure as denoted by MAP and an electrocardiogram measuring device for measuring an electrical event E of the electrical-mechanical interval. 5. The system for retrofitting existing non-invasive cardio monitoring devices according to claim 4 wherein said vibration sensing device comprises at least one of a Doppler ultrasound device and a fiber optic device. 6. The system for retrofitting existing non-invasive cardio monitoring devices according to claim 4 wherein said conversion unit determines the P, the A and the C based upon the predetermined conversion equations including P=k1(EI*MAP*E-M)+c1,A=k2(MAP*E-M)+c2, andln(C)k3(1/E-M)+c3where k1, k2, c1, c2 and c3 are empirical proportionality constants. 7. The system for retrofitting existing non-invasive cardio monitoring devices according to claim 4 wherein said conversion unit determines the P, the A and the C based upon the predetermined conversion equations including P=k1′((T-EI)*MAP*E-M)+c1′,A=k2′(MAP*E-M)+c2′, andln(C)=k3′(1/E-M)+c3′where k1′, k2′, k3′, c1′, c2′ and c3′ are empirical proportionality constants for a particular one of the subjects and T is a time period of the cardiac cycle for the particular one of the subjects. 8. The system for retrofitting existing non-invasive cardio monitoring devices according to claim 7 wherein the M in the E-M is defined as a time when a se... | - | A61B005/0428
US07647093 | B2 | 2010-01-12 | Apparatus and method for cordless recording and telecommunication transmission of three special ECG leads and their processing | 2004-08-19 | The invention is related to a device and method for cordless recording, telecommunication transmission, and processing of three special ECG leads by the means of the mobile device (1) and diagnostic-calibration center (2) where the reconstruction of standard ECG leads is performed. Reconstruction parameters are determined previously by calibration for each patient. In urgent situation, a patient (3) performs recording of three special ECG leads with the mobile device (1), and sends the memorized data to the diagnostic-calibration device (2) via a cellular telephone (4). The patient can trace the process of recording and sending the data with the help of sound and light indicators. Diagnostic-calibration center (2) is equipped with a PC computer (6) connected to a receiving cellular telephone (7) and calibration ECG device (8) with 14 electrodes. Ten electrodes are grouped into a standard 12-channel ECG cable, while the remaining four are used for recording of three special ECG leads and are grouped into a separate box (9) with integrated electrodes arranged identically as the integrated electrodes on the mobile device (1). | 1. A handheld device adapted for placement against the body of a patient in a measurement position to thereby extract special ECG signals for combining with a predetermined calibration matrix in order to construct an ECG, the handheld device comprising: at least three integrated active electrodes, including first and second integrated active electrodes adapted to electrically contact the chest of the patient in the measurement position, and a third integrated active electrode adapted to electrically contact the left hand of the patient in the measurement position; anda integrated fourth electrode adapted to electrically contact the right hand of the patient in the measurement positionswherein said integrated active electrodes and integrated fourth electrode are disposed in the device such that, when in the measurement position, said contact with the chest of the patient by the first and second integrated active electrodes, with the left hand of the patient by the third integrated active electrode, and with the right hand of the patient by the integrated fourth electrode, is operative to achieve maximum recording point vector orthogonality, to thereby optimize ECG reconstruction. 2. The device of claim 1, wherein the first and second integrated active electrodes are arranged such that in the measurement position they are disposed on the patient between the right and left mamilar lines (linea mamillaris) and lie in a line transverse to the medial line (linea mediana anterior) at an angle ranging from about 30° to about 90°. 3. The device of claim 1, wherein the first and second integrated active electrodes are disposed on a first side of the handheld device and the third integrated active electrode is disposed on a side of the handheld device different from the first side. 4. The device of claim 1, wherein the first and second integrated active electrodes are disposed on a first side of the handheld device and the integrated fourth electrode is disposed on a side of the handheld device different from the first side. 5. The device of claim 1, further comprising a integrated fifth electrode arranged such that in the measurement position, the integrated fifth electrode is in electrical contact with the chest of the patient. 6. The device of claim 5, wherein the first and second integrated active electrodes and the integrated fifth electrode are all disposed on a first side of the handheld device. 7. The device of claim 5, further comprising an amplifier module including first, second and third amplifiers each having associated therewith a first input node adapted for connection to the first, second and third integrated active electrodes, respectively, a reference node adapted for connection to the integrated fourth electrode, and a ground node adapted for connection to the integrated fifth electrode. 8. The device of claim 7, further comprising a control module adapted to receive the outputs of first, second and third amplifiers of the amplifie... | New Cardio, Inc. | A61B005/0402
US08059000 | B2 | 2011-11-15 | Wearable/portable protection for a body | 2009-05-12 | In one embodiment, a particular state of a body is sensed. In response to the sensing, at least one action is taken to modulate a projected adverse interaction between the body or a portion thereof and at least one object in the environment of the body. | 1. A system comprising: means for sensing a particular state of a living body; andmeans for protecting the living body from an object responsive to the sensing, includingmeans for determining one or more protective specifics related to at least one protective cushioning action located substantially at the body based upon accessing stored information associated substantially with an approximation of the living body's mass distribution, andmeans for activating the at least one protective cushioning action with the one or more protective specifics responsive to the means for determining. 2. The system of claim 1 wherein the means for sensing a particular state of a living body is located proximate to the body. 3. The system of claim 1 wherein the means for sensing a particular state of a living body is located remotely from the body. 4. The system of claim 1 wherein the means for sensing a particular state of a living body includes at least means for detecting motion; andmeans for determining whether the motion is likely to be the particular state. 5. The system of claim 1 wherein the means for sensing a particular state of a living body includes at least means for sensing an acceleration from substantially a beginning of a specified time-interval until substantially an end of the specified time-interval. 6. The system of claim 1 wherein the means for sensing a particular state of a living body includes at least means for sensing a direction of motion of the body. 7. The system of claim 1 wherein the means for sensing includes at least determining a location at which to perform the activating based on at least the direction of the sensed motion. 8. The system of claim 1 wherein the means for sensing includes at least means for sensing a vector direction of a motion. 9. The system of claim 1 wherein the means for sensing includes at least means for determining whether a positive indication of the particular state is substantially expected to be a false positive. 10. The system of claim 1 wherein the particular state is associated substantially with at least an acceleration of the body; andthe means for sensing includes at least means for detecting motion; andmeans for determining whether the motion is likely to be substantially due to the acceleration. 11. The system of claim 1 wherein the particular state is associated substantially with at least an approximate positioning of the body; andthe means for sensing includes at least means for determining the approximate positioning. 12. The system of claim 1 wherein the particular state is associated substantially with at least a change in an approximate positioning of the body; and the means for sensing includes at least means for determining the change in the approximate positioning. 13. The system of claim 1 wherein the particular state is associated substantially with at least an approximation of an acceleration of the body; andthe means for sensing includes at least means for determining... | The Invention Science Fund I, LLC | G08B023/00
US08005525 | B2 | 2011-08-23 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2009-10-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. In a glucose sensing system comprising a glucose sensor configured for producing an output current representative of a glucose concentration in a host and processing circuitry configured to process the output, a method for processing data from the output, the method comprising: generating a data stream from the glucose sensor output current;detecting a signal artifact by examining the data stream for high frequency content cycles above a predetermined threshold; andgenerating at least one estimated data stream or data stream portion to reduce or replace the detected signal artifact. 2. The method of claim 1, wherein the data stream comprises raw data stream. 3. The method of claim 2, wherein the raw data stream comprises integrated or averaged signal values. 4. The method of claim 1, wherein the predetermined threshold is expressed in units of cycles/hours. 5. The method of claim 1, wherein the data stream comprises a window of data from a time period of from about 15 minutes to about 60 minutes. 6. The method of claim 1, comprising filtering the data stream. 7. The method of claim 6, comprising filtering the data stream with a recursive filter. 8. The method of claim 6, comprising filtering the data stream with a high pass filter. 9. The method of claim 1, wherein generating the estimated signal is initiated in response to detection of the signal artifact. 10. The method of claim 1, wherein generating the estimated signal is terminated in response to a drop in a severity of the signal artifact to a level below the predetermined threshold. 11. A glucose sensor system for monitoring a glucose concentration in a host, the system comprising: a continuous glucose sensor configured to produce a data stream comprising current representative of glucose concentration in a host;a signal artifact detection module configured to detect a signal artifact by examining the data stream for high frequency content cycles above a predetermined threshold; anda signal estimation module configured to generate an estimated data stream or data stream portion to reduce or replace the signal artifact. 12. The system of claim 11, wherein the predetermined threshold is expressed in units of cycles/hours. 13. The method of claim 11, wherein the data stream comprises a window of data for a time period of from about 15 minutes to about 60 minutes. 14. The system of claim 11, wherein the system comprises a filter. 15. The system of claim 14, wherein the system comprises a recursive filter. 16. The system of claim 14, wherein the system comprises a high pass filter. 17. The system of claim 11, wherein the signal estimation module is configured to initiate signal estimation in response to positive detection of the signal artifact. 18. The system of claim 11, wherein the signal estimation module is configured to terminate signal estimation in response to a drop in severity of the signal artifact. 19. The system of claim 11, wherein the signal artifact detectio... | DexCom, Inc. | A61B005/05 A61B005/00
US08401666 | B2 | 2013-03-19 | Modification profiles for posture-responsive therapy | 2009-04-30 | In general, the disclosure describes techniques for modifying therapy provided to a patient by a medical device. The techniques may be applicable to electrical stimulation therapy or other therapies. Modification of therapy may include adjustment of one or more therapy parameter values that define one or more characteristics of stimulation therapy delivered to a patient. The therapy modification may be based on activity of a patient that is detected by an IMD, such as a change in a detected posture state occupied by the patient. Different therapy modifications may be applied for different changes in detected posture state. An IMD may modify therapy based on a transition from one posture state to another posture state, and apply different modifications for different transitions. In some aspects, the modification may include a profile, such as a ramp up or ramp down in a parameter value over a period of time. The profile may be different for different posture transitions. A ramp slope may be steeper for one transition than for another transition. When a patient transitions from an upright to a lying posture state, for example, the profile may be especially abrupt. In some examples, a modification profile may define a dwell time. | 1. A method comprising: controlling the delivery of therapy to a patient from a medical device according to a posture state of the patient;detecting a first posture state transition of the patient;modifying the therapy according to a first modification profile defined for the first posture state transition, wherein the first modification profile defines a first rate of change for an adjustment of at least one parameter of the therapy delivered to the patient;detecting, subsequently, a second posture state transition of the patient; andmodifying the therapy according to a second modification profile defined for the second posture state transition, wherein the second modification profile defines a second rate of change for an adjustment of the at least one parameter of the therapy delivered to the patient that is different than the first rate of change defined by the first modification profile, andwherein at least one of the controlling, detecting, determining, and modifying are performed via one or more processors. 2. The method of claim 1, wherein the therapy includes electrical stimulation therapy, and wherein the at least one parameter of the therapy comprises at least one stimulation parameter value of the electrical stimulation therapy. 3. The method of claim 1, wherein the first posture state transition comprises a transition from an upright posture state to an upright and active posture state, and the second posture state transition comprises a transition from the upright posture state to a lying down posture state, and wherein the first rate of change is less than the second rate of change. 4. The method of claim 3, wherein the second rate of change defines a substantially immediate change in the at least one parameter from a first value associated with the upright posture state to a second value associated with the lying down posture state. 5. The method of claim 2, wherein the first rate of change includes a third rate of change and a fourth rate of change for the adjustment to the at least one stimulation parameter value. 6. The method of claim 5, wherein the third rate of change defines a first change in the at least one stimulation parameter value from a first value associated with a first posture state to an intermediate value, and the fourth rate of change defines a second change in the at least one stimulation parameter value from the intermediate value to a second value associated with a second posture state. 7. The method of claim 6, wherein the third rate of change in the at least one stimulation parameter value from the first value to the intermediate value is substantially immediate. 8. The method of claim 6, wherein the intermediate value is less than the first value and the second value. 9. The method of claim 6, wherein the intermediate value is associated with an unstable posture state and the second posture state is a stable posture state. 10. The method of claim 1, wherein detecting the first posture state trans... | Medtronic, Inc. | A61N001/365
US08400313 | B2 | 2013-03-19 | Vehicle driver sleep state classification generating device based on Hidden Markov Model, sleep state classification device and warning device | 2008-01-21 | There is provided an arousal state classification model generating device for generating a blink waveform pattern model and an arousal state pattern model based on data on a blink waveform, which are preferably used for accurately estimating an arousal level with respect to unspecified object persons, an arousal state classifying device for classifying the arousal state of an object person, and a warning device. In the arousal state classification model generating device, a first pattern model is generated by learning a statistical model by using as learning data first feature data extracted from the blink data of at least one of the eyes of each object person at the time of blinking and blink waveform identification information. A second pattern model is generated by learning a statistical model by using as learning data second feature data including data on the occurrence ratio of each specific type of blink waveform in the sequence of analysis intervals and arousal state information data in which arousal state information indicating the arousal state of each object person is provided to each sequence of analysis intervals. | 1. An arousal state classification model generating device for generating a statistical model to determine an arousal state of an object person, the arousal state classification model generating device characterized by comprising: learning data storage means for storing first feature data extracted from blink data of at least one eye of each object person at the time of blinking, and blink waveform identification information data in which blink waveform identification information indicating a specific type of blink waveform is provided to each blink in the blink data;blink waveform pattern model generation means for learning a statistical model by using as learning data the first feature data and the blink waveform identification information data stored in the learning data storage means, and generating a first pattern model having as an input the first feature data and having as an output a likelihood for the blink waveform identification information in relation to the first feature data;feature data generation means for generating second feature data including data on an occurrence ratio of each of the specific types of blink waveforms in an analysis interval based on the blink waveform identification information data stored in the learning data storage means; andarousal state pattern model generation means for learning a statistical model by using as learning data the second feature data generated by the feature data generation means and arousal state information data in which arousal state information indicating the arousal state of the object person is provided to each sequence of the analysis intervals, and generating a second pattern model having as an input the second feature data and having as an output a likelihood for the arousal state information in relation to the second feature data. 2. The arousal state classification model generating device according to claim 1, characterized in that the blink data is electro-oculogram (EOG) waveform data or moving picture of eye region. 3. The arousal state classification model generating device according to any one of claims 1 and 2, characterized in that the data on the occurrence ratio is a time variation of the occurrence ratio or a time variation of the number of occurrences. 4. The arousal state classification model generating device according to claim 1, characterized in that an HMM (Hidden Markov Model) is used for the statistical model. 5. The arousal state classification model generating device according to claim 1, further comprising: blink data storage means for storing the blink data of at least one eye of the object person at the time of blinking;feature data extraction means for extracting the first feature data from the blink data acquired from the blink data storage means;first pattern model storage means for storing the first pattern model generated by the blink waveform pattern model generation means;second feature data storage means for storing the second feature data gene... | Asahi Kasei Kabushiki Kaisha | G08B023/00 G08B001/00 G08B021/00 B60Q001/00 G06F019/00
US08989854 | B2 | 2015-03-24 | Retrieving mental images of faces from the human brain | 2008-08-21 | Methods, systems and computer program products for retrieving mental images of faces from the human brain. An exemplary embodiment includes a method for retrieving mental images of faces from a human brain, the method including generating an image on a screen, measuring brain activity data from a brain scanner, varying characteristics of the image on the screen, detecting changes in the brain activity data measured from the brain scanner, and generating a final image on the screen, the final image being related to a subjectively described mental image. | 1. A method for retrieving mental images of faces from a human brain, the method comprising: generating a human facial image on a screen, the human facial image including characteristics based on an initial description of a human face;measuring brain activity data from a brain scanner, the brain activity data being measured from a subject viewing the screen;varying the characteristics of the human facial image on the screen;detecting changes in the brain activity data measured from the brain scanner, the brain activity indicative of recognizing the human face;locking certain characteristics that correspond to brain activity that indicated strong recognition of the certain characteristics to reconstruct a final human facial image;reconstructing a final human facial image from the certain characteristics that have been locked based on the strong recognition evidenced by the strong brain activity; andgenerating a final human facial image on the screen, the final human facial image including updated characteristics based on predetermined brain activity indicative of recognizing the human face;wherein varying characteristics of the human facial image on the screen comprises:in response to varying a first facial characteristic that results in an increase of brain activity as indicated from the brain activity data, fixing the first facial characteristic as part of the final human facial image on the screen;varying a second facial characteristic; andin response to varying a second facial characteristic that results in an increase of brain activity as indicated from the brain activity data, fixing the second facial characteristic as part of the final human facial image on the screen. 2. The method as claimed in claim 1 wherein the brain activity data from the brain scanner is activity from areas of a human brain related to facial recognition. 3. The method as claimed in claim 2 wherein the areas of the human brain include at least one of prefrontal cortex, posterior associative neocortex, medial temporal lobe, amygdala complex, and fusiform gyms. 4. The method as claimed in claim 1 further comprising generating a plurality of facial characteristics to be displayed as the human facial image on the screen. 5. The method as claimed in claim 4 further comprising: for each of the plurality of facial characteristics displayed on the screen:varying the facial characteristic;in response to varying the facial characteristic that results in an increase of brain activity as indicated from the brain activity data, fixing the single facial characteristic as part of the final human facial image on the screen. 6. The method as claimed in claim 2 wherein detecting changes in the brain activity data measured from the brain scanner comprises an increase of brain activity from areas of the human brain related to facial recognition. | International Business Machines Corporation | A61B005/04 G06F003/01
US08565857 | B2 | 2013-10-22 | Catheter and medical assembly | 2007-09-07 | The present invention relates to a catheter (6) comprising: a connector (65, 66) at a proximal side of the catheter for connecting the catheter to an external signal transmission/receiving unit (10) for transmitting and/or receiving signals, an electrode (63, 64) at a distal side of the catheter, and an electrical connection including an electrical wire (61, 62) for electrically connecting the electrode and the connector for the transmission of signals between the electrode and the connector, wherein the electrical connection has a high electrical resistance of at least 1 kΩ, in particular of at least 5 kΩ. Thus, the present invention provides a solution to prevent excessive heating during EP interventions under MR guidance by using highly resistive wires and or lumped resistors as connections within catheters. | 1. A catheter comprising: a connector at a proximal side of the catheter for connecting the catheter to an external signal transmission/receiving unit for communicating signals between the catheter and the signal transmission/receiving unit,an electrode at a distal side of the catheter configured to be introduced into a patient P, andan electrical connection including an electrical wire for electrically connecting the electrode and the connector for the transmission of the signals between the electrode and the connector, wherein the electrical connection has a high electrical resistance of at least 1 kΩ to limit any power deposition in a tissue of the patientfurther comprising a current sensing unit R2 configured to sense the electrical current flowing in the electrical connection at the distal side of the catheter and further configured to transmit a sensing signal indicative of the sensed current to an external current control unit. 2. A catheter as claimed in claim 1, wherein the wire has an electrical resistance of at least 1 kΩ. 3. A catheter as claimed in claim 1, wherein the electrical connection comprises a lumped resistor having a resistance of at least 1 kΩ. 4. A catheter as claimed in claim 1, wherein said electrode is adapted for receiving physiological signals. 5. A catheter as claimed in claim 1, wherein said electrical connection comprises a pair of wires for electrically connecting a pair of electrodes with the connector. 6. A catheter as claimed in claim 1, wherein said current sensing unit comprises a sensing resistor R2 having a small electrical resistance of 1 kΩ for sensing a voltage representing said sensing signal and a pair of sensing wires for transmitting said sensing signal to said external current control unit. 7. A catheter as claimed in claim 1, wherein said wire is substantially made of a metal, a metal alloy, carbon fibers, conductive polymers or other conductive non-magnetic material. 8. A catheter as claimed in claim 1, wherein said wire is made of non-conductive fibers or threads coated with a conductive coating. 9. A catheter as claimed in claim 1, wherein said wire is made of non-conductive fibers or threads filled with a conductive filling or conductive particles. 10. Medical assembly comprising: a signal transmission/receiving unit configured to communicate and process signals; anda catheter including a connector at a proximal side of the catheter connecting the catheter to the signal transmission/receiving unit for communicating the signals between the catheter and the signal transmission/receiving unit,an electrode at a distal side of the catheter configured to be introduced into a patient P, andan electrical connection including an electrical wire for electrically connecting the electrode and the connector for the transmission of the signals between the electrode and the connector, wherein the electrical connection has a high electrical resistance of at least 1 kΩ to limit any power deposition in... | Koninklijke Philips N.V. | A61B005/05
US09568580 | B2 | 2017-02-14 | Identifying white matter fiber tracts using magnetic resonance imaging (MRI) | 2009-07-01 | Systems, methods, and software are used for identifying fibers based at least in part on magnetic resonance imaging. A fiber tract atlas for a nervous system includes atlas voxels that each represent a different volume element of the nervous system; a first atlas voxel represents a first volume element of the nervous system. The fiber tract atlas also includes information on orientations of a first fiber tract in the first volume element of the nervous system. Magnetic resonance data is acquired from the nervous system of a subject. The magnetic resonance data includes data voxels; a first data voxel relates to the first atlas voxel. A diffusion vector is generated for the first data voxel based at least in part on the acquired magnetic resonance data. The fiber tract atlas is used to find a probability that the first data voxel represents the first fiber tract based at least in part on the generated diffusion vector and the information on the orientations of the first fiber tract in the first volume element. | 1. A computer implemented method for determining how well magnetic resonance imaging data obtained from a subject corresponds to a given white matter fiber tract, the method comprising: obtaining a fiber tract atlas for a nervous system, the fiber tract atlas comprising: a plurality of atlas voxels that each represent a different volume element of the nervous system, the plurality of atlas voxels including a first atlas voxel that represents a first volume element of the nervous system, andaveraged information on probabilistic locations and orientations of a first fiber tract in the first volume element of the nervous system, wherein the averaged information on the probabilistic locations and the orientations of the first fiber tract includes a rank-two symmetric atlas diffusion tensor (T) having a first atlas diffusion tensor eigenvector (ν), such that ν′Tν is equal to a largest eigenvalue of T;acquiring magnetic resonance data from the nervous system of a subject, the magnetic resonance data comprising a plurality of data voxels including a first data voxel that relates to the first atlas voxel;calculating a diffusion tensor for the first data voxel;generating a diffusion vector (ν) for the first data voxel, the diffusion vector being a first eigenvector corresponding to a largest eigenvalue of the calculated diffusion tensor for the first data voxel; anddetermining a probability P that the first data voxel of the magnetic resonance data of the subject corresponds to the first fiber tract, based at least in part on the generated diffusion vector (ν) and the averaged information on the probabilistic locations and orientations of the first fiber tract in the first volume element, and without manual inspection of the generated diffusion vector, the orientation information, or magnetic resonance images associated with the magnetic resonance data, wherein the determined probability P is determined by combining the atlas diffusion tensor (T), the first atlas diffusion tensor eigenvector (ν), and the diffusion vector (ν) of the first data voxel according to the expressionv′⁢T_⁢vv′⁢Tv_,wherein ν′ is the transpose of v and ν′ is the transpose of ν. 2. The computer implemented method of claim 1, the method further comprising: producing location information of a fiber tract of interest associated with the first data voxel using fiber density data, wherein the fiber density data is associated with the plurality of atlas voxels and includes streamline fiber counts to indicate a likelihood of an atlas voxel to the fiber tract of interest;resampling the averaged information on the probabilistic locations and the orientations of the first fiber tract from an atlas space to a subject space for the subject, wherein the determined probability that the first data voxel corresponds to the first fiber tract is based in part on resampled location probability and resampled orientation information. 3. The computer implemented method of claim 1, wherein the nervous sy... | The Regents Of The University Of California | A61B005/05 G01R033/563 G01R033/56
US20100004559 | A1 | 2010-01-07 | BIOSENSOR CARTRIDGE, METHOD OF USING BIOSENSOR CARTRIDGE, BIOSENSOR DEVICE, AND NEEDLE INTEGRAL SENSOR | 2007-12-17 | A biosensor cartridge capable of making the collected amount of a sample required for measurement into a small quantity to reduce a user's burden, and easily collecting and measuring a sample of a puncture hole without requiring the operation of bringing a sample collection opening close to the puncture hole is provided. When one end portion of biosensor chip is pushed against subject, elastic body is compressed, and puncturing tool can protrude and perform puncturing. Additionally, if a pushing force is weakened, puncturing tool is extracted from subject by the restoring force of elastic body, and blood (sample) flows out of the puncture hole. In this case, since the puncture hole, and sample collection opening provided at biosensor chip are contained in enclosed half-open space formed by elastic body, even a small amount of blood (sample) can also be easily collected. Additionally, elastic body is formed with a groove which forms a ventilation passage, so that enclosed half-open space can be opened to the atmosphere. | 1. A biosensor cartridge comprising: a biosensor chip, anda puncturing tool fixed to a portion of the biosensor chip and having a tip protruding therefrom, whereinan elastic body having a through hole which contains a sample collection opening provided at a tip of the biosensor chip and a puncture hole formed in a subject by the puncturing tool is provided at a tip of the biosensor chip, andthe elastic body includes a ventilation passage which passes through between an internal wall surface which forms the through hole and an external wall surface of the elastic body. 2. The biosensor cartridge of claim 1, wherein the sample collection opening provided at the tip of the biosensor chip and the puncture hole formed in the subject by the puncturing tool are connected together by the through hole. 3. The biosensor cartridge of claim 1, wherein the elastic body is configured such that the internal diameter of the hole on the side which touches the subject is made larger than the internal diameter of the hole on the side which touches the biosensor chip. 4. The biosensor cartridge of claim 3, wherein the elastic body is configured such that the internal wall surface which forms the through hole is raised toward the tip of the elastic body on the side which touches the subject. 5. The biosensor cartridge of claim 1, wherein the ventilation passage is a groove provided at the end face of the elastic body on the side of the subject. 6. The biosensor cartridge of claim 1, wherein the ventilation passage is an open through hole formed substantially along a radial direction of the elastic body. 7. The biosensor cartridge of claim 1, wherein a plurality of the ventilation passages are formed in the elastic body, and are arranged so as to be symmetrical in the radial direction of the elastic body with the axial center of the elastic body as a center. 8. The biosensor cartridge of claim 1, further comprising: a drive mechanism which punctures a subject with the puncturing tool. 9. The biosensor cartridge of claim 1, wherein the elastic body has stickiness. 10. The biosensor cartridge of claim 1, wherein the elastic body is fixed to the tip of the biosensor chip with its own stickiness, fixed to the tip of the biosensor chip with an adhesive, or fixed to the tip of the biosensor chip with a double-sided tape. 11. A method of using a biosensor cartridge of claim 1, the method comprising the steps of: pushing the elastic body against the subject to compress the elastic body to puncture the subject,pulling out a puncturing tool from the subject by a restoring force of the elastic body, andperforming sample collection while the through hole is opened to the atmosphere. 12. A biosensor device comprising: the biosensor cartridge of claim 1, anda measuring instrument which is connected to detecting electrodes of the biosensor cartridge and obtains information on a sample which has been collected. 13. A needle integral sensor comprising: a needle integral se... | NATIONAL INSTITUTE OF ADVANCED INDUSTRIAL SCIENCE AND TECHNOLOGY SUMITOMO ELECTRIC INDUSTRIES, LTD. | A61B005/151
US20100016706 | A1 | 2010-01-21 | METHODS FOR DIAGNOSIS AND MONITORING OF NEUROLOGIC DISEASES USING MAGNETIC RESONANCE METHODS | 2007-12-07 | A method of diagnosing or monitoring a neurological condition in a subject is described. The method includes performing a first magnetic resonance method on a subject to produce a first data set, performing a second magnetic resonance method on the subject to produce a second data set, and analyzing the first data set and the second data set. | 1. A method of diagnosing, monitoring, or predicting the future severity of a neurological condition in a subject comprising: (a) performing a first magnetic resonance method on said subject to produce a first data set,(b) performing a second magnetic resonance method on said subject to produce a second data set, and(c) analyzing said data first data set and second data set to extract mathematical features from these data sets that are then used to diagnose, monitor, or predict the future severity of a neurological condition or disease in said subject. 2. (canceled) 3. The method of claim 1, wherein said first magnetic resonance method is a magnetic resonance imaging (MRI) method selected from the group consisting of diffusion tensor imaging (DTI), anatomical resonance imaging, magnetization transfer imaging, volumetric measurements of brain, brain structures or lesions, T1 weighted MRI, T2 weighted MRI, T1 weighted MRI with contrast agents, and T2 weighted MRI with contrast agents. 4. The method of claim 1, wherein said second magnetic resonance method is magnetic resonance spectroscopy (MRS). 5.-7. (canceled) 8. The method of claim 4, wherein said magnetic resonance spectroscopy (MRS) is a single voxel method or a multi-voxel method. 9. The method of claim 8, wherein said multi-voxel method is chemical shift imaging. 10. The method of claim 8, wherein said magnetic resonance spectroscopy (MRS) includes analyzing MRS data and measuring levels of one or more metabolites. 11.-12. (canceled) 13. The method of claim 4, wherein said magnetic resonance spectroscopy (MRS) measures A/B wherein A is the amount of a first metabolite and B is the amount of a second metabolite. 14. The method of claim 4, wherein said magnetic resonance spectroscopy (MRS) measures the ratio of C/D wherein C is the sum of the amounts of two or more metabolites, and D is the sum of two or more metabolites, wherein C/D does not equal one. 15.-21. (canceled) 22. The method of claim 1, wherein said neurological condition is multiple sclerosis or Alzheimer's Disease. 23.-25. (canceled) 26. A method of diagnosing, monitoring, or predicting the future severity of a neurological disease or condition in a subject, comprising: (a) performing a first magnetic resonance method on said subject to produce a first data set;(b) performing a second magnetic resonance method on said subject to produce a second data set;(c) repeating step (a), or repeating step (b), or repeating step (a) and step (b), to generate additional data sets; and(d) analyzing said first data set, said second data set and said additional data sets to extract mathematical features that are then used to diagnose, monitor, or predict the future severity of a neurological disease or condition in said subject. 27. The method of claim 26, wherein steps (a)-(d) are performed before and after treatment with a drug or therapy. 28. The method of claim 26, wherein step (a) and step (b) occur prior to treatment of s... | Molecular Image, Inc. | A61B005/055
US20100016710 | A1 | 2010-01-21 | PROSTATE TREATMENT APPARATUS | 2009-07-13 | The presented invention aims to address these and/or additional issues by providing a tracked grid for brachytherapy applications that can adjust to changes in the prostate between original image acquisition. The grid is allowed to be moved around and even rotated under tracked conditions, such that it can be maneuvered to align with any target region. A virtual 3-D grid is displayed on a computer with respect to the actual anatomy of the patient. The tracked grid can be moved relative to the virtual 3-D grid, such that the brachytherapy, cryo therapy or any other type of image guided therapy may be performed in real 3-D or even 4-D, when motion compensation is embedded during the procedure. | 1. An improved system for image guided brachytherapy procedure, comprising: a 3-D TRUS scanning subsystem for 3-D generating a current 3-D image of a patient anatomy;an image combination image for overlaying a previous image of the patient anatomy having a therapy plan therein onto the current 3-D image of patient anatomy;a virtual grid engine for generating a virtual grid on the current image;graphical user interface showing virtual grid, live ultrasound video, 3-D motion compensated image and the overlaid plan with respect to the virtual grid a tracked brachytherapy grid for aligning a needle trajectory with said virtual grid. 2. The system of claim 1 further comprising: motion compensation engine for removing motion artifacts during the procedure. 3. The system of claim 1, further comprising: a cancer atlas for superimposing cancer information onto the current image. | - | A61B005/05
US20100036286 | A1 | 2010-02-11 | DEVICE FOR INDIRECTLY MEASURING OCCLUSAL FORCES | 2008-02-22 | An apparatus, in particular a dental splint, for the measurement of occlusal forces in case of a terminal occlusion of teeth comprises at least one sensor which is arranged outside of an occlusion region of the apparatus. A system for the diagnosis or treatment of bruxism serves to determine the position or the size of occlusal forces on the apparatus. A method allows exact localization and quantitative determination of occlusal forces on the apparatus. | 1. A dental splint for the measurement of occlusal forces in case of a terminal occlusion of teeth, wherein said dental splint is provided for coming into contact with an occlusal surface of the teeth, said dental splint comprising: at least one sensor for determining the measurement of the occlusal forces;a support comprising: two opposing flank regions arrangeable to be perpendicular to the occlusal surface of the teeth; andan occlusion region arrangeable to be parallel to the occlusal surface of the teeth;wherein the support can be fitted onto at least one of the teeth by means of the flank regions and the occlusion region;wherein the at least one sensor is arranged within one of the flank regions of the support and outside of the occlusion region, the support completely enclosing the at least one sensor. 2. The dental splint according to claim 1, wherein the at least one sensor is arrangeable on the outside of a tooth, on the outside of a dental filling, a plombage, a crown or an inlay of a tooth, or on the outside of a dental implant, an artificial tooth or a dental prosthesis. 3. The dental splint according to claim 1, wherein the at least one sensor is selected from the group consisting of an expansion sensor, a bending sensor, a vibration sensor, a tension sensor, a displacement sensor, a pressure sensor, and a force sensor. 4. The dental splint according to claim 1, wherein the at least one sensor is a piezoelectric sensor or a strain gauge. 5. The dental splint according to claim 1, wherein said at least one sensor is a plurality of sensors differing in structure or differing in function. 6. The dental splint according to claim 1, and further comprising an electronic unit for taking up and/or evaluating measured variables registered by the at least one sensor. 7. The dental splint according to claim 6, and further comprising a transmitter unit operatively connected to the at least one sensor for transmitting measured variables registered by the at least one sensor or taken up or evaluated by the electronic unit to a receiver unit. 8. The dental splint according to claim 1, wherein the support comprises at least two different regions, each region having a different hardness, stiffness or strength. 9. The dental splint according to claim 8, wherein the at least one sensor is a plurality of sensors completely arranged within either one, the other or both of the at least two different regions. 10. The dental splint according to claim 8, wherein the at least one sensor is a plurality of sensors arranged adjacent to either one, the other or both of the at least two different regions. 11. A system for the diagnosis or treatment of bruxism, comprising: a dental splint in case of a terminal occlusion of teeth, wherein said dental splint is provided for coming into contact with an occlusal surface of the teeth, said dental splint comprising: at least one sensor for determining the measurement of the occlusal forces; anda support compris... | - | A61B005/11
US20100036210 | A1 | 2010-02-11 | ELASTIC CLOTH DEVICE WITH A FUNCTION OF PHYSIOLOGICAL DETECTION | 2009-03-26 | An elastic cloth device with a function of physiological detection includes two elastic bodies, two fastening elements, and a connecting strap. The two elastic bodies are aligned with an interval in-between and each includes an elastic cloth strap woven with at least one strain sensing unit; wherein the strain sensing unit has a surface emerging from the elastic cloth strap. Each fastening element is combined with one end of the corresponding elastic body and electrically coupled with the corresponding strain sensing unit. The connecting strap is connected with the ends of the elastic bodies so that the elastic bodies are spacedly connected in series. Accordingly, when a physiological strain reaction has an effect on the elastic cloth straps, the straps drive the strain sensing units to produce a physiological strain signal which is electrically transmitted to the fastening elements, thereby the function of physiological detection is achieved. | 1. An elastic cloth device with a function of physiological detection, comprising: an elastic body, including an elastic cloth strap and at least one strain sensing unit elastically woven with the elastic cloth strap, the strain sensing unit having a surface emerging from the elastic cloth strap; anda fastening element with electrical conductivity, combined on the elastic body and electrically coupled with the strain sensing unit;whereby when a physiological strain reaction has an effect on the elastic cloth strap, the elastic cloth strap drives the strain sensing unit to produce a physiological strain signal and the physiological strain signal is electrically transmitted to the fastening element. 2. The elastic cloth device as claimed in claim 1, wherein the strain sensing unit is a plurality of strain sensing conductive yarns. 3. The elastic cloth device as claimed in claim 1, wherein the strain sensing unit has a first portion and a second portion integrally and continuously connected with the first portion, the first portion emerges from a surface of the elastic cloth strap, and the second portion contacts the fastening element. 4. The elastic cloth device as claimed in claim 3, wherein the second portion is submerged in the elastic cloth strap, or submerged in the elastic cloth strap for a distance and then emerges from the surface of the elastic cloth strap, or the second portion directly extends to the fastening element from the surface of the elastic cloth strap. 5. The elastic cloth device as claimed in claim 1, wherein one strain sensing unit or more than one strain sensing units that is non-conductive to each other being disposed side by side are combined on the elastic cloth strap, wherein the one or more strain sensing units are long-strip-shaped. 6. The elastic cloth device as claimed in claim 1, wherein the fastening element is a male fastening element or a female fastening element, or a plurality of male fastening elements or female fastening elements, corresponding to a plurality of strain sensing units, wherein the strain sensing units are long-strip-shaped. 7. An elastic cloth device with a function of physiological detection, comprising: two elastic bodies, adjacent to each other at intervals in turn and each including an elastic cloth strap and at least one strain sensing unit woven with the elastic cloth strap, the strain sensing unit having a surface emerging from the elastic cloth strap;two fastening elements with electrical conductivity, each combined with one end of the corresponding elastic body and electrically coupled with the corresponding strain sensing unit; anda connecting strap, connected with the ends of the elastic bodies so that the elastic bodies are spacedly connected in series;whereby when a physiological strain reaction has an effect on the elastic cloth straps, the elastic cloth straps drive the strain sensing units to produce a physiological strain signal, and the physiological strain signal is ele... | ZENTAN TECHNOLOGY CO., LTD. | A61B005/00
US07655217 | B2 | 2010-02-02 | Near infrared imaging agent | 2006-03-22 | This invention relates to an in-vivo diagnostic method based on near infrared radiation (NIR radiation) that uses water-soluble dyes and their biomolecule adducts, each having specific photophysical and pharmaco-chemical properties, as a contrast medium for fluorescence and transillumination diagnostics in the NIR range, to new dyes and pharmaceuticals containing such dyes. | 1. An in-vivo diagnostic based on near infrared radiation comprising compounds of the general formula I BL—(F—Wn)m (I)wherein L represents a number from 0 to 6,n represents a number from 0 to 10, andm represents a number from 1 to 100;B is a biological detecting unit having a molecular weight of up to 30,000 that bonds to specific cell populations or selectively to receptors, or accumulates in tissues or tumours, or generally stays in the blood, or is a macromolecule that bonds non-selectively;F represents a dye showing maximum absorption in the range of 650 to 1200 nm;as well as their physiologically tolerable salts, wherein F within the general formula I represents a cyanine dye of the general formula IIawherein r represents the numbers 0, 1 or 2, wherein, for r=2, the respective L6 and L7 groups that occur in duplicate may be the same or different, L1 to L7 are the same or different, each independently representing a CH or CR group,where R is a halogen atom, a hydroxy, carboxy, acetoxy, amino, nitro, cyano or sulfonic acid group or an alkyl, alkenyl, hydroxyalkyl, carboxyalkyl, alkoxy, alkoxycarbonyl, sulfoalkyl, alkylamino, dialkylamino or halogenalkyl group containing up to 6 carbon atoms, an aryl, alkylaryl, hydroxyaryl, carboxyaryl, sulfoaryl, arylamino, diarylamino, nitroaryl or halogenaryl group containing up to 9 carbon atoms, or where R represents a bond that bonds to another group R and forms a 4- to 6-member ring together with the interspersed groups L1 to L7, or two of L1 to L7 that are linked via a —CO— group,R3 to R12 are the same or different, each independently representing a hydrogen atom, a group B as defined above or W as defined below, or an alkyl or alkenyl group containing up to 6 carbon atoms or an aryl or aralkyl group optionally carrying an additional group W as defined below, or with 5- to 6- member rings that may be saturated, unsaturated or aromatic, and that may optionally carry an additional group R as defined above, annealed to each pair of adjacent residue R3 to R10 with due regard for the interspersed C atoms,X and Y are the same or different, each independently representing an O, S, Se or Te or a —C(CH3)2—, —CH═CH— or —CR13R14— group,where R13 and R14 independently represent a hydrogen atom, a group B as defined above or W as defined below, or an alkyl or alkenyl group containing up to 6 carbon atoms or an aryl or aralkyl group containing up to 9 carbon atoms, the alkyl, alkenyl, aryl or aralkyl group optionally carrying an additional group W as defined below,wherein W within the general formula I is a carboxy or sulfonic acid group or a carboxyalkyl group or an alkoxycarbonyl or alkoxyoxoalkyl group containing up to 12 carbon atoms, andwherein W takes the position of R4 and R8, R6 and R10, R11 and R12, or combinations thereof. 2. A method of producing a fluorescent image in tissues, the method comprising administering at least one of the compounds of claim 1 to tissues and detecting said compound by means of... | Institut fur Diagnosikforschung GmbH | A61B010/00 A61B005/00 A61B008/00
US07647089 | B2 | 2010-01-12 | Surface identification using microwave signals for microwave-based detection of cancer | 2006-01-26 | A method and a system of determining a surface location defining an interface between an object to image and an antenna is provided. The method uses geometric principles and the fact that an impedance mismatch at the interface results in significant reflection. A propagation time from the surface to each antenna of a plurality of antennas is estimated. The propagation time locates the surface on a circle centered at each antenna and having a radius calculated using the propagation time. A tangent line connects the intersection of adjacent circles with the surface. The surface is estimated to be located at the tangent point where the circle and the shared tangent line between adjacent antennas intersect. Multiple tangent points may be averaged for each interior antenna. A curve is fit to the set of tangent points to provide an estimate of the surface location. | 1. A system for determining a two-dimensional cross section of a surface location that defines an interface between an object to detect and an antenna, the determination of the surface location for use in an imaging system, the system comprising: a plurality of antennas;a receiver coupled to the plurality of antennas, the receiver defining signal data from a signal received by at least one antenna of the plurality of antennas; anda processor operably coupled to the receiver to receive the defined signal data, the processor configured to execute computer-readable instructions configured to cause the system to(a) estimate, from the signal data, a first minimum propagation time for a first signal received at a first antenna, the first signal reflected from a surface, wherein the first antenna is one of the plurality of antennas;(b) calculate a first distance between the surface and the first antenna using the estimated first minimum propagation time;(c) define a first circle centered at the first antenna, the first circle having a first radius equal to the calculated first distance;(d) estimate, from the signal data, a second minimum propagation time for a second signal received at a second antenna, the second signal reflected from the surface, wherein the second antenna is one of the plurality of antennas and is adjacent a transmit location of the first antenna;(e) calculate a second distance between the surface and the second antenna using the estimated second minimum propagation time;(f) define a second circle centered at the second antenna, the second circle having a second radius equal to the calculated second distance;(g) define a tangent point, the tangent point located on the first circle and on a line tangent to both the first circle and the second circle and in a direction of radiation of the first antenna toward the surface;(h) save the defined tangent point in a set of tangent points;(i) repeat steps (a)-(h) for the first antenna being each of the plurality of antennas; and(j) determine a location of the surface by applying a curve fit to the set of tangent points. 2. The system of claim 1, wherein at least one of the plurality of antennas transmits a transmit signal selected from the group consisting of a short broadband pulse, a pulse synthesized from a plurality of discrete frequencies, and a frequency swept pulse. 3. The system of claim 1, wherein at least one of the plurality of antennas transmits a wideband signal. 4. The system of claim 3, wherein the wideband signal has a frequency in the range of one to ten gigahertz. 5. The system of claim 1, wherein at least one of the plurality of antennas transmits a coherent signal. 6. The system of claim 1, wherein the plurality of antennas are arranged to form a rectangular array. 7. A microwave imaging system for detecting an object, the system comprising: a plurality of antennas;a receiver coupled to the plurality of antennas, the receiver defining signal data from a signal receiv... | Wisconsin Alumni Research Foundation | A61B005/05
US07645243 | B2 | 2010-01-12 | Method and device for measuring of oedema | 2005-04-06 | A method and apparatus for measurement and distinguishing between brawny and pitting oedema comprises a measurement head including a tactile sensor to measure a first characteristic of the oedema and a force decay sensor mounted adjacent the tactile sensor, to measure a second characteristic of the oedema so as to distinguish brawny oedema from pitting oedema. The tactile sensor may be a resonance oscillating sensor which may determine frequency shift, and the forced decay sensor may provide information regarding the degree of displaced fluid as a measure of pitting oedema through measurement of decay of force. | 1. A method of measuring oedema in a patient, the method comprising: pressing a surface of an oscillating resonance sensor against the skin surface of the patient to a certain depth;simultaneously reading with a force sensor a force required to depress the skin surface to the certain depth and a frequency of the oscillating resonance sensor over a period of time;determining a difference Δf between the oscillating resonance sensor frequency reading and a known resonance frequency of the oscillating resonance sensor that occurs when the oscillating resonance sensor freely oscillates as a measure of brawny oedema; anddetermining a decay in the force F required to depress the skin surface to the certain depth as a measure of pitting oedema. 2. A method according to claim 1, further comprising: presenting the frequency readings and the force readings. 3. A method according to claim 1, further comprising: calculating a parameter Δf/F that describes a relative relationship between pitting oedema and brawny oedema. 4. A method according to claim 3, wherein simultaneously reading the force and the frequency comprises continuously reading the force and the frequency, and the method further comprises: plotting a curve corresponding to the parameter Δf/F, wherein an inclination of the curve corresponds to a measure of the relative relationship between pitting oedema and brawny oedema. 5. A method according to claim 1, wherein the oscillating resonance sensor surface has a circular cross section. 6. A method according to claim 5, wherein the oscillating resonance sensor surface has a diameter of 15 mm. 7. A method according to claim 1, wherein the period of time comprises 20 seconds. 8. A method according to claim 1, wherein the certain depth comprises 4 mm. 9. A method according to claim 1, wherein the surface of the oscillating resonance sensor is parallel to the skin surface when pressed. 10. A device for measuring oedema in a patient, comprising: a measuring head, the measuring head comprising an oscillating resonance sensor, the oscillating resonance sensor comprising a surface configured to be pressed against a skin surface of the patient, anda force sensor, wherein the oscillating resonance sensor is connected to the force sensor so that the oscillating resonance sensor is operable to oscillate without interference from the force sensor; andan adjustable stand that adjustably supports the measuring head. 11. A device according to claim 10, wherein the control is further configured to calculate a parameter Δf/F that describes a relative relationship between pitting oedema and brawny oedema. 12. A device according to claim 11, wherein control is configured to continuously simultaneously read the force and the frequency, and plot a curve corresponding to the parameter Δf/F, and an inclination of the curve corresponds to a measure of the relative relationship between pitting oedema and brawny oedema. 13. A device according to claim 10, further ... | Bioresonator AB | A61B005/103 A61B005/117
US20100041972 | A1 | 2010-02-18 | APPARATUS AND METHODS | 2007-07-09 | The invention relates to a multi-electrode array comprising a plurality of electrodes, the electrodes being spaced apart from each other by spacer means, the electrodes being secured to the spacer means, the spacer means being encapsulated within a housing. The invention also relates to an implantable device comprising a multi-electrode array of the invention, a method of manufacturing electrodes for use in an array of the invention and a method for manufacturing a multi-electrode array of the invention. The invention also relates to a method for monitoring the effect of a test substance on a biological tissue using a multi-electrode array of the invention. | 1. A multi-electrode array comprising a plurality of electrodes, the electrodes being spaced apart from each other by spacer means, the electrodes being secured to the spacer means, the spacer means being encapsulated within a housing. 2. A multi-electrode array according to claim 1 wherein the spacer means and housing are unitary. 3. A multi-electrode array according to claim 1 wherein the electrodes have a diameter in the range of from about 30 μm to about 200 μm, preferably from about 40 μm to about 150 μm, e.g. 125 μm. 4. A multi-electrode array according to claim 1 wherein the electrodes comprise tungsten, stainless steel, platinum, platinum/iridium, carbon fibres, conductive nanotubes, carbon nanotubes, an Elgiloy® alloy, or a conductive polymer. 5. A multi-electrode array according to claim 1 wherein the electrodes are metal electrodes. 6. A multi-electrode array according to claim 5 wherein the electrodes have shape memory. 7. A multi-electrode array according to claim 6 wherein the electrodes are formed from stainless spring steel. 8. A multi-electrode array according to claim 1 wherein the electrodes are wires. 9. A multi-electrode array according to claim 1 wherein at least one of the electrodes is hollow. 10. A multi-electrode array according to claim 1 wherein the spacer means comprises a plurality of electron microscopy grids, a plurality of fine mesh patches, or a block containing channels. 11. A multi-electrode array according to claim 1 wherein the electrodes are spaced in a regular pattern or a regular repeating pattern. 12. A multi-electrode array according to claim 1 wherein the electrodes are spaced equidistant from each other. 13. A multi-electrode array according to claim 1 wherein one or both of the housing and spacer means are formed from a mouldable material. 14. A multi-electrode array according to claim 13 wherein the mouldable material is an acrylic polymer or a thermoplastic material. 15. A multi-electrode array according to claim 1 wherein the electrodes are releasably secured to the spacer means. 16. A multi-electrode array according to claim 1 wherein the electrodes tips are in a non-planar configuration, or are adjustable to a non-planar configuration. 17. A multi-electrode array wherein each electrode comprises a connector for connection to a driver. 18. A multi-electrode array wherein each electrode comprises a push fit connector. 19. A multi-electrode array according to claim 1 wherein electrodes are provided with an electrical insulating layer or jacket. 20. A multi-electrode array according to claim 1 wherein the electrodes are dual function electrical signal and mechanical transduction detectors. 21. An implantable device comprising a multi-electrode array according to claim 1. 22. A method for monitoring the effect of a test substance on a biological tissue comprising: (a) providing a biological tissue in a medium,(b) contacting the electrode tips of a MEA, in particular according to c... | Lectus Therapeutics Limited | A61B005/0492 A61N001/04 A61N001/05 A61B005/04 H01R043/00 C12Q001/02 B44C001/22 C12M001/34 G01R027/00
US07645263 | B2 | 2010-01-12 | Impedance sensor | 2002-10-21 | A transdermal transport device includes a reservoir for holding a formulation of an active principle, a needle with a bore through which the formulation is transported between the reservoir and a target area of a biological body, and an impedance sensor. The impedance sensor has an electrode positioned to measure the impedance of a portion of the target area between the needle and the electrode to indicate the depth of penetration of the needle into the target area. | 1. A transdermal transport device, comprising: a reservoir for holding a substance;at least two needles including a first needle and a second needle, each needle having a bore for transporting the substance between the reservoir and a target area of a biological body;an impedance sensor to measure electrical impedance of a portion of the target area between the first needle and the second needle, the impedance changing as the needles penetrate through the target area and being indicative of the depth of penetration of the needles into the target area; andan automated control unit responding to the indicated impedance that monitors the changing impedance through the target area and stops the penetration of the needles into the target area when a desired impedance change has been achieved. 2. The transdermal transport device of claim 1 further comprising one or more additional needles. 3. The transdermal transport device of claim 1 wherein the measured impedance after the needle penetrates the skin is an order of magnitude less than the measured impedance before the needle penetrates the skin. 4. The transdermal transport device of claim 3, wherein the impedance drops by over three orders of magnitude when the needle has penetrated to the proper depth. 5. A method to measure penetration into a biological body, comprising: penetrating a target area of the biological body with at least two needles including a first needle and a second needle each having a bore for transporting a substance to or from the target area;measuring electrical impedance of a portion of the target area between the first needle and the second needle to determine when the needles have penetrated to a desired depth into the target area, the impedance changing as the needles penetrate through the target area; andstopping the penetration of the needles, using an impedance responsive automated control unit that monitors the changing impedance through the target area, when a desired impedance change has been achieved. 6. The method of claim 5, wherein the needles are microneedles. 7. The method of claim 5, further comprising penetrating the skin with one or more additional needles, each needle having a bore through which the formulation is transported to and from the target area. 8. The transdermal transport device of claim 1, wherein the needles are microneedles. | Massachusetts Institute of Technology | A61M001/00 A61M031/00 A61B005/05
US07647831 | B2 | 2010-01-19 | Method for measuring pressure inside an anatomical fluid system | 2007-10-22 | A disclosed method determines fluid pressure inside a vessel without compromising the integrity of the vessel. A sensor is positioned in operative communication with the external wall of the vessel such that expansion of the external wall of the vessel exerts a force against the sensor that is directed substantially radially outward with respect to the vessel. A substantially radially inward force is caused to be directed against the sensor in response to the substantially radially outward force exerted by the external vessel wall. The sensor can thus be used to detect the magnitude of the substantially radially outward force.;A disclosed apparatus determines fluid pressure inside a vessel without compromising the integrity of the vessel. The apparatus includes a sensor and a band operatively associated with the sensor and configured to at least partially encircle the vessel so as to retain the sensor in operative communication against the external wall of the vessel. | 1. A method for determining fluid pressure inside a vessel without compromising the integrity of the vessel, comprising the steps of: positioning a sensor in operative communication with the external wall of the vessel such that expansion of the external wall of the vessel exerts a force against the sensor that is directed substantially radially outward with respect to the vessel;causing a substantially radially inward force to be directed against the sensor in response to the substantially radially outward force exerted by the external vessel wall which further comprises the step of inserting the sensor between the external wall of the vessel and an adjacent body of tissue; andusing the sensor to detect the magnitude of the substantially radially outward force. 2. The method of claim 1, wherein the vessel is an artery, and wherein the step of inserting the sensor between the external wall of the vessel and an adjacent body of tissue comprises the step of inserting the sensor just below the adventitia of the artery but outside the media of the artery. 3. The method of claim 1, wherein the vessel is an aneurysm. 4. The method of claim 1, wherein the vessel runs through the visceral pericardium, and wherein the step of inserting the sensor between the external wall of the vessel and an adjacent body of tissue comprises the step of inserting the sensor outside the vessel and within the visceral pericardium. | CardioMEMS, Inc. | G01L007/00 A61B005/00
US07643870 | B2 | 2010-01-05 | Method and system or managing ECGs in a clinical trial | 2007-01-02 | The present disclosure includes a method and system for managing ECGs in a clinical trial including a processor, storage media, ECG acquisition system, user interface device, and the appropriate databases configured to collect and store a set of trial specific drug ECG recordings. The system and method allows a trial administrator to select a set of placebo ECG recordings and a set of known QT prolonged ECG recordings from databases and adjust these sets according to the parameters of the study. Interpretations are made on this superset of ECG recordings and the interpretation is outputted to a user. | 1. A method of managing an ECG clinical trial, the method comprising: collecting a set of trial specific drug ECG recordings from a patient;selecting a set of placebo ECG recordings and a set of known QT prolonged ECG recordings;adjusting the set of placebo ECG recordings and the set of known QT prolonged ECG recordings based on a predetermined set of demographic study parameters;making an interpretation based on an analysis of the set of trial specific drug ECG recordings, the set of placebo ECG recordings and the set of known QT prolonged ECG recordings; andoutputting the interpretation to a user. 2. The method as claimed in claim 1, further comprising storing the set of trial specific drug ECG recordings in a trial specific drug ECG database. 3. The method as claimed in claim 2, further comprising retrieving the set of trial specific drug ECG recordings form the trial specific drug ECG database before making the interpretation. 4. The method as claimed in claim 1, wherein the set of placebo ECG recordings are selected from a placebo ECG database. 5. The method as claimed in claim 1, wherein the set of known QT prolonged ECG recordings are selected from a known QT prolonged ECG database. 6. The method as claimed in claim 1, wherein the step of selecting the set of known QT prolonged ECG recordings further includes sorting a collection of known QT prolonged ECG recordings according to a T-wave state. 7. The method as claimed in claim 1, wherein the set of known QT prolonged ECG recordings are based on a response to Moxifloxacin. 8. The method of claim 1, further comprising maintaining a log of the selected set of placebo ECG recordings and the selected set of known QT prolonged ECG recordings. 9. The method as claimed in claim 1, wherein the selecting step and the step of making the interpretation are made using a user interface device. 10. A system for managing an ECG trial, the system comprising: an ECG acquisition device configured to collect a set of trial specific drug ECG recordings from a patient;a placebo ECG database configured to store placebo ECG database recordings, the placebo ECG database further configured such that a user can select a set of placebo ECG recordings;a known QT prolonged ECG database configured to store known QT prolonged ECG recordings, the known QT prolonged database further configured such that a user can select a set of known QT prolonged ECG recordings, wherein the set of placebo ECG recordings and the set of known QT prolonged ECG recordings are adjusted based on a predetermined set of demographic study parameters, and further wherein a reader makes an interpretation based on an analysis of the set of trial specific drug ECG recordings, the set of placebo ECG recordings, and the known QT prolonged ECG recordings; anda user interface device configured to receive the interpretation from the reader. 11. The system as claimed in claim 10, further comprising a trial specific database configured to store t... | The General Electric Company | A61B005/0402
US08328728 | B2 | 2012-12-11 | Implantable hemodynamic monitor and methods for use therewith | 2008-08-22 | Provided herein are implantable systems that include an implantable photoplethysmography (PPG) sensor, which can be used to obtain an arterial PPG waveform. In an embodiment, a metric of a terminal portion of an arterial PPG waveform is determined, and a metric of an initial portion of the arterial PPG waveform is determined, and a surrogate of mean arterial pressure is determined based on the metric of the terminal portion and the metric of the initial portion. In another embodiment, a surrogate of diastolic pressure is determined based on a metric of a terminal portion of an arterial PPG waveform. In a further embodiment, a surrogate of cardiac afterload is determined based on a metric of a terminal portion of an arterial PPG waveform. | 1. For use in an implantable system including an implantable photoplethysmography (PPG) sensor, a method for determining a surrogate of cardiac afterload (CA) in a subject, comprising: (a) obtaining an arterial PPG waveform measurement using the implantable PPG sensor implanted in the subject;(b) determining a metric of a terminal portion of the arterial photoplethysmography (PPG) waveform, wherein the terminal portion is from a maximum of the arterial PPG waveform to the following minimum of the PPG waveform; and(c) determining the surrogate of CA in the subject as a function of the metric of the terminal portion of the arterial PPG waveform in step b) based on the measured arterial PPG waveform by the implanted PPG sensor. 2. The method of claim 1, wherein the arterial PPG waveform comprises an averaged arterial PPG waveform that is produced by averaging a plurality of cardiac cycles of a PPG signal obtained from the implanted PPG sensor. 3. The method of claim 1, further comprising: repeating steps (a) and (b) and (c) from time to time; andmonitoring changes in CA by monitoring changes in the surrogate of CA as steps (a) and (b) and (c) are repeated from time to time. 4. The method of claim 1, wherein, the metric of the terminal portion comprises at least one of the following: an integral from a maximum to the following minimum of the arterial PPG waveform;an integral from a dicrotic notch to the following minimum of the PPG waveform;a terminal deceleration slope (TDS), which is a downward slope from a maximum to the following minimum of the arterial PPG waveform;a beginning terminal deceleration slope (BTDS), which is a downward slope from a maximum to the following dicrotic notch of the arterial PPG waveform;a total relaxation time (TRT), which is a time from a maximum to the following minimum of the atrial PPG signal; anda terminal deceleration time (TeDT), which is a time from a dicrotic notch to the following minimum of the atrial PPG signal. | Pacesetter, Inc. | A61B005/02 A61B005/04
US08218846 | B2 | 2012-07-10 | Automatic pathway and waypoint generation and navigation method | 2009-05-14 | A navigation system for use in a luminal network of a patient, such as the airways, that is able to analyze a three-dimensional model of the luminal network and automatically determine a pathway from an entry point to a designated target. The system further automatically assigns waypoints along the determined pathway in order to assist a physician in navigating a probe to the designated target. | 1. A method of providing a pathway through a luminal network of a patient from an entry point to a target comprising: collecting data regarding an anatomy of a luminal network;developing a three-dimensional map of said luminal network using said data;displaying a representation of said map on a user interface;selecting a target on said map;providing an algorithm that automatically determines a pathway from a selected entry point to said target;displaying said determined pathway on said user interface;providing an algorithm that automatically assigns and displays waypoints along said pathway. 2. The method of claim 1 wherein collecting data regarding an anatomy of a luminal network comprises taking a plurality of images of said luminal network. 3. The method of claim 2 wherein taking a plurality of images of said luminal network comprises taking a plurality of CT scans of said luminal network. 4. The method of claim 3 wherein developing a three-dimensional map of said luminal network using said data comprises assembling said plurality of CT scans and performing a data fill function to create a seamless three-dimensional model. 5. The method of claim 1 wherein developing a three-dimensional map of said luminal network comprises developing a skeletonized three-dimensional map of said luminal network. 6. The method of claim 1 wherein said algorithm automatically determines said pathway by performing steps comprising: beginning at said selected target;following lumina back to said entry point. 7. The method of claim 6 wherein following lumina back to said entry point comprises; selecting a point within said lumina proximate said target;determining a lumen diameter at said point;selecting a second point outside a plane of said lumen diameter;determining a second diameter at said second point;continuing to select successive points having increasing diameters until said entry point is reached. 8. The method of claim 6 wherein said luminal network comprises airways and developing a three-dimensional map of said luminal network using said data comprises labeling said map with exhalation air flow direction through said luminal network. 9. The method of claim 8 wherein following lumina back to said entry point comprises following said airflow direction labels back to said entry point. 10. The method of claim 1 wherein providing an algorithm that automatically assigns displays waypoints along said pathway comprises providing an algorithm that detects airway intersections. 11. The method of claim 10 wherein providing an algorithm that detects airway intersections further comprises providing an algorithm that detects airway intersections and designates a point spaced apart from said intersection, and on said pathway, as a waypoint. 12. The method of claim 11 wherein providing an algorithm that detects airway intersections and designates a point spaced apart from said intersection, and on said pathway, as a waypoint comprises providing an algorithm th... | superDimension, Ltd. | G06K009/00
US08498688 | B2 | 2013-07-30 | Magnetic resonance device and method | 2008-03-17 | A device for magnetic resonance imaging of a body placed in an examination volume is configured to generate a series of MR spin echo signals from a first nuclear spin species by subjecting at least a portion of the body to an MR imaging pulse sequence. The MR imaging pulse sequence includes a spatially non-selective excitation RF pulse, a plurality of refocusing RF pulses, and a plurality of phase encoding switched magnetic field gradients. The device is further is configured to acquire the MR spin echo signals without application of frequency encoding magnetic field gradients; and to reconstruct an MR image from the acquired MR spin echo signals. | 1. A device for magnetic resonance (MR) imaging of a body placed in an examination volume, the device comprising: a main magnet configured to establish a substantially homogeneous main magnetic field in the examination volume;coils configured to generate switched magnetic field gradients superimposed upon the main magnetic field;an antenna configured to radiate radio frequency (RF) pulses towards the body;a controller configured to control the coils and the antenna to generate the magnetic field gradients and the RF pulses;a receiver for receiving and sampling magnetic resonance signals; anda reconstructor configured to form MR images from the signal samples, the controller being further configured to perform the acts of:generating a series of MR spin echo signals from a first nuclear spin species by subjecting at least a portion of the body to an MR imaging purse sequence comprising a spatially non-selective excitation RF pulse, a plurality of refocusing RF pulses, and a plurality of phase encoding switched magnetic field gradients;after each phase encoding gradient, applying a rewinding gradient to cancel an effect of the phase encoding gradient before application of a next refocusing RF pulse, wherein the rewinding gradient has equal strength as the phase encoding gradient and an opposite direction of the phase encoding gradient;acquiring the MR spin echo signals without application of frequency encoding magnetic field gradients and without application of slice selection magnetic field gradients; andreconstructing an MR image from the acquired MR spin echo signals. 2. The device of claim 1, wherein the reconstructed MR image is a projection image of the portion of the body. 3. The device of claim 2, wherein the controller is further configured to determine a spatial position of a site of accumulation of a contrast agent containing the first nuclear spin species by combining at least two projection images having different projection directions. 4. The device of claim 1, wherein the controller is further configured to apply a phase encoding scheme in the generating act such that k-space is sampled in all three spatial directions. 5. The device of claim 1, wherein the controller is further configured to apply a phase encoding scheme in the generating act such that k-space is sampled along a plurality of radial directions. 6. The device of claim 1, wherein the first spin species has a chemical shift spectrum. 7. The device of claim 6, wherein the reconstructor is further configured to reconstruct an MR image for a predetermined spectral position in the chemical shift spectrum and/or for the integrated intensity of the chemical shift spectrums. 8. The device of claim 6, wherein the first nuclear spin species is Fluorine 19 (19F). 9. The device of claim 6, wherein the reconstructor is further configured to reconstruct MR images selectively for spin species having different chemical shift spectra. 10. The device of claim 1, wherein the contr... | Koninklijke Philips N.V. | A61B005/05
US08385996 | B2 | 2013-02-26 | Multiple wavelength sensor emitters | 2009-04-13 | A physiological sensor is adapted to removably attach an emitter assembly and a detector assembly to a fingertip. The emitter assembly is adapted to transmit optical radiation having multiple wavelengths into fingertip tissue. The detector assembly is adapted to receive the optical radiation after attenuation by the fingertip tissue. The sensor has a first shell and a second shell hinged to the first shell. A spring is disposed between the shells and urges the shells together. An emitter pad is fixedly attached to the first shell and configured to retain the emitter assembly. A detector pad is fixedly attached to the second shell and configured to retain the detector assembly. A detector aperture is defined within the detector pad and adapted to pass optical radiation to the detector assembly. A contour is defined along the detector pad and generally shaped to conform to a fingertip positioned over the detector aperture. | 1. A physiological sensor configured to measure an indication of a physiological characteristic of a living patient, the physiological sensor comprising: a plurality of light emitting sources arranged to impinge light on body tissue of a living patient including at least eight light emitting sources, each of the at least eight light emitting sources emitting light at a different wavelength,wherein at least a subset of the plurality of light emitting sources are further arranged in two or more clusters, each cluster comprising at least two light emitting sources that emit light within a wavelength range that is different from and does not overlap with the wavelength range of the at least two light emitting sources of other clusters,wherein for each cluster, light emitting sources that emit light at proximate wavelengths are adjacent to each other,wherein each light emitting source is activated by addressing at least one of a plurality of rows and at least one of a plurality of columns of an electrical grid and the plurality of light emitting sources are arranged to minimize optical pathlength differences;a detector responsive to transmitted light after attenuation by body tissue of the living patient, the body tissue including pulsating blood, wherein the detector is configured to generate a signal indicative of a physiological characteristic of the living patient; anda sensor housing configured to position the plurality of light emitting sources and the detector with respect to the body tissue of the living patient. 2. The physiological sensor according to claim 1, wherein multiple ones of the plurality of light emitting sources are capable of transmitting light of the same wavelength, the multiple ones simultaneously activated by addressing one of the rows or columns. 3. The physiological sensor according to claim 2, wherein: the plurality of light emitting sources are LEDs, andeach of the LEDs have an anode in common with one of the rows and a cathode in common with one of the columns so that driving one of the rows and one of the columns activates a unique one of the LEDs. 4. The physiological sensor according to claim 3, further comprising: a plurality of row drivers in communication with the rows; anda plurality of column drivers in communication with the columns, wherein a selected row driver sources current to a corresponding row and selected column driver sinks current from a corresponding column so as to activate an addressed one of the LEDs. 5. The physiological sensor according to claim 4, wherein deselected ones of the row drivers pull corresponding rows to a low voltage via a functional switch configured to operably connect deselected rows to a low voltage source and deselected ones of the column drivers pull corresponding columns to a high voltage via a functional switch configured to operably connect deselected columns to a high voltage source so as to substantially block parasitic current from unaddressed ones of the LEDs. 6.... | Cercacor Laboratories, Inc. | A61B005/1455
US08712515 | B2 | 2014-04-29 | Method and apparatus for disease diagnosis and screening using extremely low frequency electromagnetic fields | 2009-08-13 | Novel methods and apparatus for diagnosing or screening disease states in living organisms by the measurement and analysis of extremely low frequency electromagnetic fields, particularly extremely low frequency alternating current. The measurement of such fields is performed at a single point or at several test points on or in the body and compared to one or more reference. Information in the time-varying electromagnetic field is collected, then processed by diagnostic or screening algorithms to provide information about the disease state of the tissue being assessed. | 1. A method of screening for or diagnosing a disease condition in a test subject comprising: passively obtaining electromagnetic frequency signals, as a measure of native electropotentials of the test subject, between a test site and a reference site via at least one sensor electrode and at least one reference electrode;band pass filtering the signals obtained by the sensor electrode to yield sensor extremely low frequency alternating current (ELFAC) signals within a frequency range of 0.001 Hz to 0.1 Hz;bandpass filtering the signals obtained by the reference electrode to yield reference ELFAC signals;determining likelihood of disease based upon said ELFAC signals. 2. The method of claim 1, wherein said ELFAC signals are sent to a processing unit for use in the determining step, and wherein said processing unit outputs a determined result to a display device. 3. The method of claim 1, wherein said determining step comprises comparing one or more ELFAC field waveform measurements selected from periodicity, frequency, peaks, troughs, slopes, fractal dimensions, and chaos. 4. The method of claim 3, wherein said comparing step further comprises: correlating similar sensor and reference data with a low risk of disease state; andcorrelating dissimilar sensor and reference data with a higher risk of disease state. 5. The method of claim 1, wherein the reference electrode is placed on the test subject at a location complimentary or symmetrical to said test site. 6. The method of claim 1 wherein the reference electrode contacts a reference standard solution. 7. The method of claim 1 wherein said determining step uses the noise-minimized ELFAC data, wherein the equilibration period comprises establishing a noise-minimized status by continuously measuring ELFAC waveforms at said sensor electrode and said reference electrode during an equilibration period, wherein said equilibration period ends, and noise-minimized ELFAC data is achieved when substantially all of the measured ELFAC waveforms reach equilibrium. 8. The method of claim 3 whereby the comparing step is not performed unless and until noise-minimized ELFAC data has been achieved. 9. The method of claim 7 wherein noise-minimized ELFAC data are integrated with user input biological subject variables and compared using pattern recognition techniques, and a probability of a presence of a disease state is determined as a result of the pattern recognition. 10. The method of claim 9, wherein said biological subject variables are selected from age of the subject, menopausal stage, day in a menstrual cycle, palpability of a suspicious lesion, size of a suspicious lesion, dimensions of a suspicious lesion, and combinations thereof. 11. The method of claim 9 in which the biological subject variable is the menopausal stage of the subject from which the pattern recognition technique and waveform values to be employed for a given subject are determined. 12. The method of claim 9 in which the biologic... | - | A61B005/04 A61B005/05 A61B010/00
US08700122 | B2 | 2014-04-15 | Skin preparation device and biopotential sensor | 2009-05-04 | The skin preparation device and sensor of the present invention include an array of rigid tines. The tines serve to “self-prepare” the skin at each electrode site. These tines, when pressed against the skin, penetrate the stratum corneum, thereby reducing skin impedance and improving signal quality. A self-prepping device of the present invention is an optimized array of short non-conductive rigid tines in which the individual tines are created in a geometry that allows for a sharp point at the tip when molding, machining or etching is used as a method of fabrication. | 1. A sensor system comprising: a flexible substrate;a plurality of electrodes coupled to the flexible substrate;an adhesive substrate adjacent to each of the electrodes and the flexible substrate; anda plurality of groups of preparation devices, each group of preparation devices coupled to a respective electrode, each of the preparation devices comprising: a base member in the shape of a ring;an array of rigid tines contiguous with the base member, each tine having a sharp tip adapted to penetrate stratum corneum and to reduce skin impedance; anda component configured to deliver a conductive gel below the stratum corneum. 2. The sensor system of claim 1 wherein the preparation devices are arranged such that upon application of pressure to the electrode, the preparation devices angle to adjust to the surface to which the sensor system is applied. 3. A sensor system comprising: a flexible substrate;a plurality of electrodes coupled to the flexible substrate;an adhesive substrate adjacent to each of the electrodes and the flexible substrate; anda plurality of groups of preparation devices, each group of preparation devices coupled to a respective electrode, each of the preparation devices comprising: a base member;an array of rigid tines contiguous with the base member, each tine having a sharp tip adapted to penetrate stratum corneum and to reduce skin impedance;a component configured to deliver a conductive gel below the outermost layer of the stratum corneum; andwherein the base member is formed of multiple stepped levels each level containing a portion of the array of rigid tines. 4. The sensor system of claim 3 wherein the preparation devices are arranged such that upon application of pressure to the electrode, the preparation devices angle to adjust to the surface to which the sensor system is applied. 5. A sensor system comprising: a flexible substrate;a plurality of electrodes coupled to the flexible substrate;an adhesive substrate adjacent to each of the electrodes and the flexible substrate; and a plurality of groups of preparation devices, each group of preparation devices coupled to a respective electrode, each of the preparation devices comprising:a base member; andan array of tines contiguous with the base member, each tine having a pyramidal shape wherein one side of the pyramid is concave. 6. The sensor system of claim 5 wherein the preparation devices are arranged such that upon application of pressure to the electrode, the preparation devices angle to adjust to the surface to which the sensor system is applied. 7. A sensor system comprising: a flexible substrate;a plurality of electrodes coupled to the flexible substrate;an adhesive substrate adjacent to each of the electrodes and the flexible substrate; anda plurality of groups of preparation devices, each group of preparation devices coupled to a respective electrode, each of the preparation devices comprising: a base member; andan array of rigid tines contiguous with the b... | Covidien LP | A61B005/0408 A61B005/04025 A61B017/20
US09179854 | B2 | 2015-11-10 | Three-dimensional localization, display, recording, and analysis of electrical activity in the cerebral cortex | 2007-10-31 | The present invention describes a method and apparatus to localize the electrical signals measured from a subject's scalp surface, preferably in near-real time, and to generate dynamic three-dimensional information of the electrical activity occurring within the cerebral cortex of the brain. In the preferred embodiment, it can produce images that can be immediately inspected and analyzed by an operator in near-real time, resulting in a powerful new cortical imaging modality, which we denote as Dynamic Electrocortical Imaging (DECI). The present invention involves the use of a computer, an electroencephalographic (EEG) amplifier, EEG electrodes, and custom software. It can measure healthy and diseased cortical events and states in both conscious and unconscious subjects. This is useful, as it allows for the diagnosis, monitoring and treatment of cortical disorders, while also furthering the understanding of the human brain and lending use to additional non-medical applications such as in entertainment, education, lie-detection and industry. The invention in one embodiment is implemented using software in conjunction with readily available EEG hardware. Furthermore, this same method can be applied to pre-existing data and when doing so, EEG hardware is not required. Having a practical near-real time 3D imaging system brings a far more accessible technology to doctors, researchers, individuals, and private clinics to better diagnose, monitor, treat and understand many of the conditions and abnormalities of the brain. | 1. A method for near-real time three-dimensional localization of electrical activity in the brain comprising: applying electrodes to a subject's scalp and associated connectors;connecting the electrodes to an amplifier/recorder and connecting the amplifier/recorder to a computer;utilizing said amplifier/recorder to receive electrical signals transmitted through said subject's scalp;capturing the electrical signals from said amplifier/recorder with said computer continuously as to provide respective time varying electrical signals, wherein capturing the electrical signals comprises sampling sets of digital values from the electrodes at an associated sampling rate corresponding to frames of captured electroencephalographic (EEG) data;processing said time varying electrical signals, using said computer, via an inverse solution approximation to provide a series of representations of localized electrical activity originating from said subject's brain on a three-dimensional solution space representing the subject's brain, such that each representation of the series of representations of localized electrical activity comprises a current density vector for each of a plurality of locations within the three-dimensional solution space and being calculated as a product of a corresponding frame of captured EEG data and a transformation matrix; andpresenting the series of representations of the localized electrical activity to a human operator in near-real time. 2. The method of claim 1 wherein the series of representations of the quantified localized electrical activity are presented via a final presentation method selected from any using a single screen display with one window, using a using a single screen display with multiple windows, using a multiple screen display with single windows, using a multiple screen display with each screen having multiple windows, and using a speaker to generate sound. 3. The method of claim 2 wherein said human operator is aided in drawing conclusions about the subject to perform a wide variety of useful tasks. 4. The method of claim 1 wherein said series of representations of the quantified localized electrical activity is displayed as a stream of images. 5. The method of claim 4 wherein the series of representations of the quantified localized electrical activity is presented as the current density vectors mapped to voxels that have been arranged in the shape of the solution space. 6. The method of claim 5 wherein said voxels are displayed graphically utilizing a tool to manipulate the graphical display and analysis of voxels, said tool comprising one of additive blending, spatial filtering based on a selected region of interest, spatial filtering based on neuroanatomical information, and selective display of cortical shells. 7. The method of claim 5 wherein said stream of images is displayed graphically utilizing a tools to modify graphically displayed results, said tool providing one of display of electrode position... |  | A61B005/04 A61B005/0476 A61B005/16
US08506498 | B2 | 2013-08-13 | Systems and methods using induced perturbation to determine physiological parameters | 2008-10-09 | According to embodiments, systems and methods for non-invasive blood pressure monitoring are disclosed. An exciter may induce perturbations in a subject, and a sensor or probe may be used to obtain a detected signal from the subject. The detected signal may then be used to measure one or more physiological parameters of the patient. For example if the perturbations are based on a known signal, any differences between the known signal and the input signal may be attributable to the patient's physiological parameters. A phase drift between the perturbation signal and the detected signal may be determined from a comparison of the scalograms of the exciter location and the sensor or probe location. From the scalogram comparison, more accurate and reliable physiological parameters may be determined. | 1. A method for monitoring blood pressure comprising: inducing a perturbation signal in a subject's body;detecting a signal from the subject's body;transforming the detected signal using a continuous wavelet transform to generate a scalogram, wherein the scalogram comprises three dimensional data, wherein a first dimension of the scalogram corresponds to time, wherein a second dimension of the scalogram corresponds to scale, and wherein a third dimension of the scalogram corresponds to phase;determining, from the scalogram, a phase that corresponds to a particular time and scale;determining, based at least in part on the phase and the perturbation signal, a phase difference; anddetermining, based at least in part on the determined phase difference, a blood pressure measurement. 2. The method of claim 1 wherein determining the phase difference comprises: generating a scalogram of the induced perturbation signal; andcomparing the scalogram of the detected signal and the scalogram of the induced perturbation signal. 3. The method of claim 2 wherein comparing the scalogram of the detected signal and the scalogram of the induced perturbation signal comprises comparing phase information from the same or similar scales at the same or similar times of the scalogram of the detected signal and the scalogram of the induced perturbation signal. 4. The method of claim 3 wherein the same or similar scales of the first and second scalograms comprises at least one scale at the characteristic frequency of the induced perturbation signal. 5. The method of claim 1 wherein determining a blood pressure measurement comprises determining an elapsed time for the induced perturbation signal to propagate a fixed distance. 6. The method of claim 5 wherein determining a blood pressure measurement comprises taking the natural log of the elapsed time. 7. The method of claim 1 wherein determining a blood pressure measurement comprises determining a time difference, T, based at least in part on the determined phase difference and solving a multi-parameter equation that includes the time difference. 8. The method of claim 7 wherein the multi-parameter equation is p=a+b·ln(T)or a mathematical equivalent thereof, where p is the determined blood pressure measurement and a and b are constants. 9. The method of claim 1 further comprising performing at least one calibration of the blood pressure measurement, the calibration based at least in part on a known reference blood pressure measurement. 10. The method of claim 1 wherein inducing a perturbation signal in a subject's body comprises inducing a known perturbation signal in the subject's body using an ultrasonic, acoustic, or mechanical exciter. 11. The method of claim 1 wherein detecting a signal from the subject's body comprises detecting a photoplethysmograph (PPG) signal from the subject's body. 12. A system for monitoring blood pressure comprising: an exciter capable of inducing a perturbation signal in a subject's ... | Nellcor Puritan Bennett Ireland | A61B005/02 G06F019/00
US08565855 | B2 | 2013-10-22 | MRI method | 2009-02-09 | A method of operating a magnetic resonance imaging scanner for imaging the heart of a patient comprising inducing apnoea in the patient; sensing an electrical heart waveform; in response thereto moving the chest wall of the patient to a desired location; and triggering the scanner to image. | 1. A method of operating a magnetic resonance imaging scanner for imaging the heart of a patient comprising inducing apnoea in the patient;sensing an electrical heart waveform;moving the chest wall of the patient to a desired location in response to said electrical heart waveform; andtriggering the scanner to image. 2. The method of claim 1 wherein said step of moving the chest wall of the patient to a desired location comprises providing an enclosure around at least the chest of the patient and applying gas at a predetermined pressure to said enclosure. 3. The method of claim 1 wherein said step of moving the chest wall of the patient to a desired location comprises applying pressure to the outside of the chest to produce expiration. 4. The method of claim 1 wherein the apnoea is induced by hyperventilation. 5. The method of claim 1 wherein the step of inducing apnoea comprises cyclically moving the chest to cause hyperventilation. 6. The method of claim 1 further comprising the step of holding the chest wall at the desired location. 7. A magnetic resonance imaging device comprising a magnetic resonance imaging scanner, a structure in use engaging at least the chest of a patient to define, with said chest, a pressure chamber, a respirator device configured to apply cyclically varying pressures to said chamber to cause cyclic forced respiration of said patient, said forced respiration comprising successive inspiration and expiration periods that cause hyperventilation, and plural electrodes connected to said respirator device and said magnetic resonance imaging scanner configured to sense an electrical waveform of the heart of said patient, and triggering said magnetic resonance imaging scanner in accordance with said waveform, and a discriminator device configured to trigger said forced respiration in synchronism with a predetermined point of said waveform. 8. The magnetic resonance imaging device of claim 7 wherein said discriminator device is configured to determine the time period between successive maximum electrical amplitudes of said waveform, and further comprises variable timing circuitry for setting a triggering instant as a proportion of said time period. 9. The magnetic resonance imaging device of claim 7 wherein said discriminator device is configured to provide a trigger pulse in response to a predetermined characteristic of said waveform. | - | A61B005/05 A61M016/00 A62B007/00
US20100022846 | A1 | 2010-01-28 | Method for Health Monitoring | 2006-11-28 | This invention relates to methods for monitoring the anatomy of an individual, which methods comprise the steps of: (i) subjecting a healthy or asymptomatic individual to non-invasive analysis techniques, such as imaging or acoustic techniques, to obtain images or data on determined cells, tissue(s) or organs, at different time points over a period of more than one month, preferably over years; (ii) collecting the images or data in a database; (iii) comparing the images or data to the images or data previously acquired for the same individual, in essentially the same conditions; wherein said comparison of images or data obtained at different time points allows definition of non-pathological baseline anatomy for this individual. | 1.-17. (canceled) 18. A method for monitoring the health of an individual, which method comprises the steps of: (i) repeatedly subjecting a healthy or asymptomatic individual to non-invasive analysis techniques to obtain images or data of several distinct types of determined cells, tissue(s) or organs, at different time points over a period of more than one month;(ii) collecting the images or data in a database;(iii) comparing the images or data to images or data previously acquired for the same individual, in essentially the same conditions;wherein said comparison of images or data obtained at different time points allows definition of non-pathological baseline images or data for this individual. 19. The method of claim 18, wherein a difference between newly acquired images or data as compared to baseline images or data is potentially indicative of an abnormality or disease that has developed. 20. The method of claim 18, wherein the images or data are collected under various physiological conditions. 21. The method of claim 18, wherein the images or data are collected using imaging or acoustic techniques. 22. The method of claim 18, wherein the images or data are collected at different time points over years. 23. The method of claim 18, wherein the images or data undergo processing techniques to facilitate storage, improve quality or analysis. 24. The method of claim 18, wherein wavelet signal processing techniques are applied to data. 25. The method of claim 18, wherein 3 Dimensional images are analysed. 26. The method of claim 18, wherein data obtained under different body positions are compared and define boundaries of non pathological variability for a single individual. 27. The method of claim 18, wherein data obtained under different physiological conditions are compared and define boundaries of non pathological variability for a single individual. 28. The method of claim 18, wherein data obtained at time intervals during growth of a child are used to predict normal anatomy upon adulthood. 29. The method of claim 18, wherein the images or data from at least two different imaging techniques are combined or compared to improve accuracy. 30. The method of claim 18, wherein the analysis of images and data further takes into account the health history and condition of family members. 31. The method of claim 18, wherein the collected images or data are further correlated to genotypic or phenotypic features of the individual. 32. The method of claim 18, wherein the collected images or data are further correlated to environmental factors. 33. The method of claim 18, wherein step (i) is performed using an instrument-based technique to monitor human microscopic or macroscopic anatomy or physiology, that is installed over several months or years for acquisition of data in the personal house or workplace or school of a given individual in order to generate a database of individual anatomical or physiological data over time. 34. A comp... | - | A61B005/00 G06F017/30
US20100004526 | A1 | 2010-01-07 | ABNORMALITY FINDING IN PROJECTION IMAGES | 2009-06-01 | The present invention includes a utility for determining the severity of a stenosis in a blood vessel. In one aspect, a method for improving DSA image quality includes: (1) registration of the mask and bolus images prior to subtracting procedure to reduce the artifacts from misalignment; (2) enhancement of the registered DSA image by an anisotropic diffusion technique and a nonlinear normalization technique; and (3) detecting the boundary of blood vessel and quantitatively measuring percentage stenosis, which may be done automatically. | 1. A method for use in the quantitative measurement of a blood vessel blockage, comprising: operating an imaging system to obtain a series of mask images and a series of bolus images corresponding to a region of interest;operating a processor to register the mask images to the bolus images to generate a registered mask image and subtract the registered mask image form the bolus image to produce a motion compensated DSA image;operating the processor to process the DSA image to enhance contrast in at least the region of interest of the DSA image;receiving a user input identifying a lumen within the region of interest; andbased on the user input, operating the processor to autonomously identify a boundary of the lumen and calculate a stenosis percentage for the lumen. 2. The method of claim 1, wherein registering the series of mask images and the series of bolus images further comprises averaging said series to produce an average mask image and an average bolus image, wherein the average mask image is registered to the average bolus image. 3. The method of claim 1, wherein registering comprises applying an inverse-consistency constraint to a deformation of the mask frame to the bolus frame. 4. The method of claim 3, wherein the inverse-consistency constraint comprises a B-spline parameterization. 5. The method of claim 1, wherein enhancing the region of inertest comprises: operating the processor to diffuse the DSA image to remove background noise from the DSA image. 6. The method of claim 5, further comprising: operating the processor to perform a non-linear normalization on the DSA image, wherein a contrast between the lumen and the background of the DSA image is increased. 7. The method of claim 1, wherein receiving the user input identifying the lumen within the region of interest comprises one of: receiving at least two user selected points associated with a centerline of said lumen;receiving a plurality of user selected points associated with edges of the lumen. 8. The method of claim 7, wherein operating the processor to autonomously identify a boundary of the lumen and calculate a stenosis percentage for the lumen further comprises at least one of: initial edge detection;edge refinement; andlesion measurement. 9. The method of claim 8, wherein initial edge detection based on the at least two points associated with the centerline of the lumen, comprises defining an initial centerline in the lumen;projecting a plurality of perpendicular lines relative to the centerline; andfor each perpendicular line, identifying a boundary point between the lumen and the background based on gradients calculated along said perpendicular line, wherein a plurality of the boundary points define an initial edge of the lumen. 10. The method of claim 8, wherein edge refinement comprises fitting a contour to the initial edge, wherein the contour is smoothed to define a refined edge. 11. The method of claim 8, wherein lesion measurement comprises: using the ... | EIGEN, INC. | A61B005/05 G06K009/00
US07643864 | B2 | 2010-01-05 | Adaptive imaging parameters with MRI | 2005-02-03 | Systems, methodologies, media, and other embodiments associated with automatically adapting MRI controlling parameters are described. One exemplary method embodiment includes configuring an MRI apparatus to acquire MR signal data using a non-rectilinear trajectory. The example method may also include acquiring MR signals, transforming the MR signals into image data, and selectively adapting the MRI controlling parameters based, at least in part, on information associated with the MR signals. | 1. A system, comprising: a magnetic resonance imaging (MRI) apparatus configured to acquire a magnetic resonance (MR) signal from an object during an MR procedure using a radial trajectory without acquiring information associated with a catheter inserted into the object, the apparatus being controlled, at least in part, by an imaging parameter;a parameter logic configured to selectively automatically adapt, substantially in real time, the imaging parameter during the MR procedure based, at least in part, on analyzing one or more attributes of the object as provided in data associated with the MR signal without analyzing an attribute of a catheter; anda data store configured to store the imaging parameter. 2. The system of claim 1, the imaging parameter being a number of radial projections per MR signal acquisition period of time. 3. The system of claim 2, where the MR procedure is a perfusion MRI procedure, and where the parameter logic is configured to analyze and respond to the data associated with MR signals from the perfusion MRI procedure with respect to one or more of, a contrast change data, a contrast rate of change data, and a contrast saturation data. 4. The system of claim 2, where the MR procedure is a diffusion MRI procedure, and where the parameter logic is configured to analyze and respond to the data associated with MR signals from the diffusion MRI procedure with respect to one or more of, a water diffusion change data, and a water diffusion rate of change data. 5. The system of claim 2, where the MR procedure is a BOLD MRI procedure, and where the parameter logic is configured to analyze and respond to the data associated with MR signals from the BOLD MRI procedure with respect to one or more of, a blood oxygen change data, and a blood oxygen rate of change data. 6. The system of claim 2, where the MR procedure is a perfusion MRI procedure, and where the parameter logic is configured to analyze and respond to the data associated with MR signals from the catheter based MRI procedure with respect to one or more of, a catheter location, a catheter speed, a catheter acceleration, and a catheter orientation. 7. The system of claim 2, where the MRI apparatus is controllable to have a field of view (FOV) smaller than a sample. 8. The system of claim 2, the imaging parameters including one or more of, a temporal resolution parameter, a spatial resolution parameter, a field of view (FOV) parameter, a slice position parameter, a slice orientation parameter, a slice thickness parameter, a slice rotation parameter, a bandwidth parameter, a time to repetition (TR) parameter, a time to echo (TE) parameter, time from inversion (TI) parameter, a signal to noise ratio parameter, a tip angle for the RF excitation parameter, an RF excitation/ADC center frequency parameter, an RF excitation/ADC phase parameter, an active receive coils for data collection parameter, a pulse sequence type parameter, a number of projections per MR signal acquis... | Case Western Reserve University | A61B005/00
US07668738 | B2 | 2010-02-23 | Insurance claim filing system and method | 2001-02-27 | A method and system for enabling a service provider to file insurance claims electronically directly with the payer rather than through a clearinghouse. The system has two components interacting with an existing practice management system of a service provider and with online claim filing systems of insurance payers. The first is a software application installed on the service provider's computer to seek out billing information and create files suitable for filing of formatted claims with the payer's web sites. The second component is a web portal that interacts with both the first component and the payers' web sites to keep the former with formatting requirements and to facilitate filing of the claims electronically. The system uses a software “cookie” to convey instructions to the first component so that it can automatically connect to the appropriate page of each payer's web site. | 1. A method for enabling a practice management system of a service provider to file an insurance claim electronically via a payer's web site, said insurance claim being based on billing data stored on said practice management system, said method comprising the steps of: electronically locating billing data stored on a practice management system of a service provider;electronically eliminating a portion of said billing data intended for payers that do not permit online filing of claims;downloading from a payer's web site to a web portal format instructions for formatting billing data for online filing of insurance claims at said payer's web site, said portal operated by a third party who is not a prayer nor said service provider;determining by said web portal an address at said payer's webs site for filing insurance claims;conveying from said web portal said format instructions and said address to said practice management system of said service provider;said practice management system electronically accessing said payer's web site;said practice management system electronically navigating to said payer's web site to said address where said insurance claims can be filed;then electronically formatting said billing data in accordance with said format instructions to generate insurance claims; andthen uploading said insurance claims to said payer's web site. 2. The method as recited in claim 1, wherein said conveying step further comprises the step of setting a cookie in said practice management system by said web portal, said cookie containing said address of said payer's web site. 3. The method as recited in claim 1, wherein said conveying step further comprises the steps of: comparing format instructions on said practice management system with instructions downloaded from said payer's web site to determine if said format instructions on said practice management software are current; andconveying said downloaded format instructions if said format instructions on said practice management software are not current. 4. The method as recited in claim 1, further comprising the step of polling by said web portal said payer's web site for changes in format instructions. 5. The method as recited in claim 1, wherein said determining step further comprises the steps of: logging onto said payer's web site; andnavigating said payer's web site to locate said address for filing billing files. 6. A system for filing an insurance claim at a web site operated by a payer, said web site permitting electronic filing by a service provider of billing data formatted as an insurance claim said system comprising: a web portal, operated by a third party who is not a payer nor said service provider, programmed to obtain from each web site of plural web sites that permit electronic filing of insurance claims, insurance claim formatting instructions and a web page location for filing said formatted billing data, wherein each web site is operated by a different payer; anda co... | Blue Cross and Blue Shield of South Carolina | A61B005/00
US07860549 | B2 | 2010-12-28 | Method and apparatus for 3-D imaging of internal light sources | 2009-09-29 | The present invention provides systems and methods for obtaining a three-dimensional (3D) representation of one or more light sources inside a sample, such as a mammal. Mammalian tissue is a turbid medium, meaning that photons are both absorbed and scattered as they propagate through tissue. In the case where scattering is large compared with absorption, such as red to near-infrared light passing through tissue, the transport of light within the sample is described by diffusion theory. Using imaging data and computer-implemented photon diffusion models, embodiments of the present invention produce a 3D representation of the light sources inside a sample, such as a 3D location, size, and brightness of such light sources. | 1. A method for modeling an object located inside a mammal, the method comprising: capturing surface light image data for light emitted from a surface portion of a mammal, said light propagating from the object located inside the mammal to the surface portion;converting the surface light image data into photon density just inside the surface of the mammal using sets of modeled surface elements and volume elements; andconstructing a three-dimensional representation of the light source distribution internal to the mammal with a processor using the photon density just inside the surface. 2. The method of claim 1, wherein the three-dimensional representation of the light source distribution is approximated by a set of point light sources. 3. The method of claim 1, further comprising the step of: defining a cost function and a set of constraints for obtaining a solution for the three-dimensional representation of the light source distribution. 4. The method of claim 3, wherein the cost function is related to a sum of source strengths for each point source in the mammal, and the constraints include the following conditions: (i) that the source strengths be positive definite and (ii) that the resulting photon density at the object surface produced by the distribution of point sources be everywhere less than the measured surface photon density. 5. The method of claim 4, wherein obtaining the three-dimensional representation maximizes the cost function subject to the constraints. 6. The method of claim 3, wherein the cost function and constraints are described mathematically by a system of linear equations, and a solution for the three-dimensional representation of the source distribution is obtained using a SIMPLEX method. 7. The method of claim 3, wherein the cost function includes a weighting factor that can be varied to produce a set of solutions for the three-dimensional representation of the source distribution. 8. The method of claim 1, wherein the light source is comprised of bioluminescent or fluorescent emission. 9. The method of claim 1, further comprising the step of: applying a noise threshold to the surface light image data. 10. The method of claim 1, further comprising the step of: placing the mammal on a stage included in an imaging chamber coupled to a camera configured to capture an image of the mammal on the stage. 11. The method of claim 10, further comprising the steps of: moving the stage to one or more other positions in the imaging chamber, wherein the other positions have different angles relative to a fixed datum associated with the camera than the first position; andcapturing additional image sets of the mammal from the other positions using the camera. 12. A method for obtaining a three-dimensional representation of a light source distribution located inside a mammal, the method comprising: capturing surface light image data for light emitted from a surface portion of a mammal by propagating from a light source distr... | Xenogen Corporation | A61B006/03 A61B006/12
US20100030047 | A1 | 2010-02-04 | Subcutaneous Glucose Electrode | 2009-08-31 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably three or four-layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer is overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited (third layer). An outer (fourth) layer is biocompatible. | 1. An electrochemical sensor comprising: one or more non-corroding metal or carbon electrodes;a sensing layer comprising an enzyme coupled to each electrode; anda biocompatible layer comprising a biocompatible hydrogel chemically bound to the sensing layer of each electrode. | Abbott Diabetes Care Inc. | A61B005/1473
US20100036211 | A1 | 2010-02-11 | SYSTEMS AND METHODS FOR MEASURING PHYSIOLOGICAL PARAMETERS OF A BODY | 2007-11-07 | Systems and methods for measuring physiological parameters of a body are described. In some embodiments, the system employs a physiological monitoring device capable of measuring, amplifying, digitizing and transmitting events ranging in frequency from 0 (DC) to 10 kHz. | 1. A method for measuring a physiological parameter of a mammalian body, the method comprising: sampling the physiological parameter with a plurality of contacts positioned at the body to determine two or more sampled values for the physiological parameter;constructing a value for the physiological parameter using the two or more sampled values; andadjusting for a DC offset. 2. The method of claim 1, wherein adjusting for a DC offset comprises applying an offset voltage to a signal corresponding to a sampled value of the physiological parameter. 3. The method of claim 1, further comprising converting the signal from an analog signal to a digital signal. 4. The method of claim 3, further comprising wirelessly transmitting the digital signal to a data station. 5. The method of claim 1 wherein applying an offset voltage comprises applying a dynamic offset voltage. 6. The method of claim 1, further comprising applying an electrical stimulus to the body via one or more of the contacts. 7. The method of claim 1, further comprising implanting the contacts within the body. 8. The method of claim 1, further comprising placing the contacts at a skin of the body. 9. The method of claim 1, further comprising positioning the contacts at a cortical surface of the body. 10. The method of claim 1 wherein sampling the physiological parameter comprises sampling an electrical potential of the body. 11. The method of claim 1 wherein sampling the physiological parameter with the contacts comprises sensing the physiological parameter with a plurality of detectors arranged in an array. 12. The method of claim 1 wherein: sampling the physiological parameter with the contacts comprises at least approximately simultaneously sampling the physiological parameter with first and second contacts; andthe method further comprises at least approximately simultaneously sampling the physiological parameter with third and fourth contacts positioned at the body. 13. The method of claim 1 wherein: sampling the physiological parameter with the contacts comprises sampling the physiological parameter with first and second contacts at a first time interval; andthe method further comprises sampling the physiological parameter with third and fourth contacts at a second time interval different than the first time interval. 14. An apparatus for measuring a physiological parameter of a body, the apparatus comprising: an array of contacts for sampling a physiological parameter of a body; anda processing device electrically coupled to the contacts to receive signals from the contacts, the processing device including a voltage offset component for adjusting for a DC offset, an analog to digital converter for converting the signals from analog signals to digital signals, and an output component for transmitting the digital signals to an external device. 15. The apparatus of claim 14 wherein the voltage offset component, the analog to digital converter, and the output component are di... | WASHINGTON STATE UNIVERSITY | A61B005/00
US20100016676 | A1 | 2010-01-21 | Systems And Methods For Adaptively Filtering Signals | 2008-09-30 | According to embodiments, systems and method are provided for adaptively filtering signals. A combination of filtering techniques is described in which non-wavelet based filtering techniques are used with wavelet-based filtering techniques to filter an input signal. One or more physiological parameters may then be determined from the filtered output signal. | 1. A method for filtering a signal, the method comprising: receiving an input signal;filtering the received input signal using a first adaptive filtering technique to produce a first filtered signal; andfiltering the first filtered signal using a wavelet filtering technique to produce an output filtered signal. 2. The method of claim 1, wherein receiving an input signal comprises receiving at least one photoplethysmograph (PPG) signal. 3. The method of claim 1, wherein filtering the received input signal using a first adaptive filtering technique comprises generating a reference signal from the received input signal. 4. The method of claim 3, wherein filtering the received input signal using a first adaptive filtering technique additionally comprises filtering the received input signal using an adaptive correlation canceller and the generated reference signal. 5. The method of claim 1, wherein filtering the first filtered signal using a wavelet filtering technique comprises transforming the first filtered signal using a continuous wavelet transform. 6. The method of claim 5, wherein filtering the first filtered signal using a wavelet filtering technique comprises generating a scalogram of the continuous wavelet transform of the first filtered signal. 7. The method of claim 6 further comprising identifying at least one artifact in the scalogram. 8. The method of claim 1, wherein filtering the first filtered signal using a wavelet filtering technique comprises generating at least two scalograms from the first filtered signal, wherein at least one of the at least two scalograms is a residual scalogram. 9. The method of claim 1 further comprising determining at least one physiological parameter from the output filtered signal. 10. A method for filtering a signal, the method comprising: receiving an input signal;filtering the received input signal using a wavelet filtering technique to produce a first filtered signal; andfiltering the first filtered signal using an adaptive filtering technique to produce an output filtered signal. 11. A system for filtering a signal, the system comprising: a sensor capable of generating an input signal; anda processor capable of: filtering the input signal using a first adaptive filtering technique to produce a first filtered signal; andfiltering the first filtered signal using a wavelet filtering technique to produce an output filtered signal. 12. The system of claim 11, wherein the sensor is capable of generating at least one photoplethysmograph (PPG) signal. 13. The system of claim 11, wherein the processor is capable of generating a reference signal from the input signal. 14. The system of claim 13, wherein the processor is capable of filtering the input signal using an adaptive correlation canceller and the generated reference signal. 15. The system of claim 11, wherein the processor is capable of transforming the first filtered signal using a continuous wavelet transform. 16. The system of claim 1... | Nellcor Puritan Bennett Ireland | A61B005/00
US20100042016 | A1 | 2010-02-18 | BLOOD SENSOR AND BLOOD INSPECTION DEVICE USING IT | 2007-12-21 | A blood sensor usable both for a blood inspection device that performs puncture with a puncturing needle and for a blood inspection device that performs puncture with laser. In concrete, this invention comprises a sensor unit having a holder and a blood sensor mounted to the holder and a needle unit having a slidable puncturing needle. The needle unit provides a blood sensor block detachably fitted to the inside of the holder of the sensor unit. The blood sensor block is applied to a needle-puncturing type blood inspection device. The sensor unit from which the needle unit is removed is applied to a laser-puncturing type blood inspection device. | 1. A sensor unit comprising: a holder; anda blood sensor attached to the holder,wherein a needle unit with a puncturing needle slidably provided can be attached detachably inside the holder in the sensor unit. 2. The sensor unit according to claim 1, wherein a shape of a cross section of the holder in the sensor unit is a circle, oval or polygon. 3. The sensor unit according to claim 1, wherein at least a portion of the holder in the sensor unit comprises a transparent member. 4. The sensor unit according to claim 1, wherein detecting means that detects contact with skin is provided in the holder in the sensor unit. 5. The sensor unit according to claim 1, wherein a surface of the holder in the sensor unit that contacts with skin is water-repellant. 6. The sensor unit according to claim 1, wherein the blood sensor in the sensor unit comprises: a plurality of detection electrodes; anda reference electrode for specifying the plurality of detection electrodes. 7. The sensor unit according to claim 1, wherein: the blood sensor in the sensor unit comprises a blood storing part provided in practically the center of the blood sensor; anda projecting part projecting from a blood supply channel is provided over the storing part. 8. The sensor unit according to claim 1, wherein the blood sensor and the holder in the sensor unit are formed integrally. 9. A blood sensor block comprising: a sensor unit that comprises:a holder; anda blood sensor attached to the holder; anda needle unit with a puncturing needle provided slidably,wherein the needle unit is attached detachably to the holder in the sensor unit. 10. The blood sensor block according to claim 9, wherein: a first insertion guide is provided in an inner surface of the holder in the sensor unit; andthe needle unit is attached in the holder in the sensor unit along the first insertion guide. 11. The blood sensor block according to claim 9, wherein the puncturing needle of the needle unit is urged in a direction to part from the blood sensor. 12. The blood sensor block according to claim 9, wherein: a second insertion guide is provided in an outer surface of the holder in the sensor unit; andthe sensor unit is attached along the second insertion guide to a blood test apparatus that analyzes blood components based on a signal from the blood sensor. 13. A blood test apparatus comprising: a housing that comprises at one end an insertion part to which the sensor unit according to claim 1 is attached;an electrical circuit section that is provided in the housing and that is connected with a blood sensor in the sensor unit; andpuncturing means that is provided in the housing and that emits laser light to puncture skin. 14. The blood test apparatus according to claim 13, wherein the insertion part abuts on the blood sensor included in the sensor unit, and comprises a connector that electrically connects the electrical circuit section and the blood sensor. 15. The blood test apparatus according to cl... | PANASONIC CORPORATION | A61B005/151
US20100041974 | A1 | 2010-02-18 | Physiological monitoring garment | 2009-06-12 | A physiological monitoring garment includes a fabric portion. At least one elongate stretchable textile data/power bus is disposed the fabric portion. The elongate stretchable textile data/power bus includes a plurality of integral conductors, woven, knitted, or braided along the length thereof. One or more sensors are connected to the elongate stretchable textile data/power bus. | 1. A physiological monitoring garment comprising: a fabric portion;at least one elongate stretchable textile data/power bus disposed on the fabric portion, the elongate stretchable textile data/power bus including a plurality of integral conductors, woven, knitted, or braided along the length thereof; andone or more sensors connected to the elongate stretchable textile data/power bus. 2. The physiological monitoring garment of claim 1 in which at least one of said plurality of integral conductors extends externally from said textile/data power bus for connecting said one or more sensors thereto. 3. The physiological monitoring garment of claim 1 in which the fabric portion is an elastic fabric portion. 4. The physiological monitoring garment of claim 1 in which the plurality of conductors are woven, knitted, or braided in a strain relief pattern. 5. The physiological monitoring garment of claim 4 in which the plurality of conductors include conductors for data transfer, conductors for power transfer, and conductors structured and arranged to form an electric shield for reducing noise. 6. The physiological monitoring garment of claim 4 in which a coating of conductive material is disposed on the elongate stretchable textile data/power bus to form an electrical shield for reducing noise. 7. The physiological monitoring garment of claim 4 in which said strain relief pattern is out of the plane of said stretchable data/power bus. 8. The physiological monitoring garment of claim 4 in which said strain relief pattern is in the plane of said stretchable data/power bus. 9. The physiological monitoring garment of claim 1 in which connectors connect the one or more sensors to the elongate stretchable textile data/power bus. 10. The physiological monitoring device of claim 9 in which the connectors are insulation displacement connectors. 11. The physiological monitoring garment of claim 1 in which a sensor includes a respiration monitoring device. 12. The physiological monitoring garment of claim 11 in which said respiration monitoring device includes: at least a first elongate stretchable textile member,at least first and second conductive components associated with said textile member, andthreads of said textile member providing dielectric separation between the at least first and second conductive components so that the expansion and contraction of the textile member results in a change in the spacing between the first and second conductive components for measuring change in capacitance. 13. The physiological monitoring garment of claim 12 in which the first conductive component includes woven, knitted, or braided conductive threads integral with said first textile member. 14. The physiological monitoring garment of claim 13 in which there are a plurality of elongate stretchable textile members adjacent each other and each includes integral woven, knitted, or braided conductive threads. 15. The physiological monitoring garment of claim 12 i... | - | A61B005/04 A61B005/00
US07666151 | B2 | 2010-02-23 | Devices and methods for passive patient monitoring | 2002-11-20 | Devices, systems and methods provide passive patient monitoring of such parameters as body motion, body position, respiratory rate and/or heart rate. Passive monitoring generally involves a sensor device having at least two piezoelectric sensors, provided on a surface, such as a bed, chair, operating table or the like, so that a patient may be coupled with the device by simply allowing the patient to lie, sit, lean, stand on or wear the surface. In one embodiment, multiple patients in a general care area of a hospital may be continuously monitored via multiple passive monitoring devices. If a patient fails to meet one or more predefined threshold criteria or has a negative physiological trend, the system may activate an alarm. | 1. A system for passively monitoring multiple patients, the system comprising: a plurality of sensor devices, each sensor device comprising: a surface for coupling with a patient; andat least two piezoelectric sensors coupled with the surface for sensing at least one mechanical signal from the patient, each of the sensors comprising a piezoelectric film;a plurality of processors, each processor coupled with one sensor device and being operative for converting the at least one mechanical signal into at least one digital signal and for analyzing the digital signals corresponding to each of the at least two piezoelectric sensors to select the sensors with the highest quality heart rate and respiration rate signals to use for deriving heart rate and respiration rate data pertaining to the patient; andat least one connector coupled with each processor for connecting each processor with a common alarm system for providing an alarm based on heart rate and/or respiration rate criteria for each of the multiple patients to a user at a remote location. 2. A system as in claim 1, wherein each surface comprises a flat pad. 3. A system as in claim 2, wherein the flat pad comprises at least one layer of a resilient foam material. 4. A system as in claim 1, wherein the piezoelectric sensors are disposed along each surface in a pattern to facilitate monitoring of at least one of patient body position, patient movement, patient respiratory rate and patient heart rate. 5. A system as in claim 4, wherein the pattern comprises two or more adjacent, flat piezoelectric sensors. 6. A system as in claim 5, wherein the pattern comprises a grid pattern comprising at least two adjacent rows and at least two columns of piezoelectric sensors. 7. A system as in claim 1, wherein dimensions of each surface allow the surface to be positioned on at least one of a crib and a bed for monitoring an infant. 8. A system as in claim 1, wherein dimensions of each surface allow the surface to be positioned on a bed for monitoring a patient in a hospital, long-term care facility or nursing home. 9. A system as in claim 1, wherein dimensions of each surface allow the surface to be positioned on a chair for monitoring a patient undergoing a procedure under conscious sedation. 10. A system as in claim 1, wherein each sensor device further includes a sheath for containing at least a portion of the surface to provide a protective layer between the surface and the patient. 11. A system as in claim 10, wherein the sheath is water resistant. 12. A system as in claim 10, wherein the sheath is removable from the sensor device after being used to monitor one patient. 13. A system as in claim 1, wherein the piezoelectric sensors are selected from the group consisting of polyvinylidene fluoride film, polyvinylidene fluoride cable, piezoelectric ceramic discs and piezoelectric foam. 14. A system as in claim 1, wherein each surface is provided on a flat surface of a bed, and wherein the at l... | Hoana Medical, Inc. | A61B005/103 A61B005/117 A61B005/08 G08B023/00
US07662133 | B2 | 2010-02-16 | Spinal fluid introduction | 2004-02-23 | A fluid introduction system includes an introducer configured to create a pressure of at least 69 kPa within a spine, and an operator configured to actuate the introducer to introduce fluid into the spine according to a predetermined fluid introduction profile. The system can include a computer readable medium having code for receiving fluid introduction data indicative of a fluid introduction parameter, and for receiving response data indicative of a response of the patient at a time related to a time of the fluid introduction data. A method for introducing fluid includes positioning a first introducer in a first portion of a spine, positioning a second introducer in a second, different portion of the spine and, without removing the first and second introducers, introducing fluid into the first portion of the spine with the first introducer and introducing fluid into the second portion of the spine with the second introducer. | 1. A fluid introduction system, comprising: an introducer configured to introduce fluid into a spine of a patient, the introducer having a flow-rate dependent impedance opposing the introduction of the fluid and comprising a movable element disposed within the introducer, the movable element including a pressure transducer secured with respect to the movable element such that the pressure transducer is in direct contact with fluid in the introducer;an operator configured to actuate the introducer to introduce fluid into the spine of the patient at a constant flow rate that is not controlled by feedback based on a fluid introduction parameter and configured to determine impedance data indicative of the flow rate-dependent impedance based upon pressure and volume of fluid dispensed during an actuation of the introducer prior to insertion of the introducer into the spine;a computer readable medium having code that, when executed by a computer, receives: fluid introduction data indicative of the fluid introduction parameter;response data indicative of pain level of a response of the patient at a time related to a time of the fluid introduction data; andresponse data, input separately from the pain level data, indicative of concordance of the response of the patient at the time related to the time of the fluid introduction data, the concordance data indicating whether the pain level response of the patient is a result of a pain symptom or is a result of pain unrelated to the pain symptom. 2. The fluid introduction system of claim 1, wherein the fluid introduction parameter is a pressure within an intervertebral disc of the patient at the time of the fluid introduction data. 3. The fluid introduction system of claim 1, wherein the fluid introduction parameter is a total amount of fluid introduced into an intervertebral disc of the patient at the time of the fluid introduction data. 4. The fluid introduction system of claim 1, wherein the fluid introduction system is configured to obtain the response data from an observation of the patient. 5. The fluid introduction system of claim 1, wherein the fluid introduction system is configured to obtain the response data upon a response inputted by the patient. 6. The fluid introduction system of claim 1, wherein the introducer is configured to create a pressure of at least 100 kPa within the spine. 7. The fluid introduction system of claim 1, further comprising a sliding device. 8. The fluid introduction system of claim 7, wherein the response data comprises data inputted directly by the patient using the sliding device. 9. The fluid introduction system of claim 1, wherein the response data comprises observed physiological parameters. 10. The fluid introduction system of claim 1, wherein the response data is correlated with actual measurements of disc pressure and a volume of fluid introduced into one or more discs. 11. The fluid introduction system of claim 1, wherein the introducer comprises a need... | Smith & Nephew, Inc. | A61M037/00 A61B005/103 A61B005/117
US07647087 | B2 | 2010-01-12 | Apparatus and methods of cortical surface registration and deformation tracking for patient-to-image alignment in relation to image-guided surgery | 2006-08-02 | A cortical surface registration procedure related to a diagnostic or surgical procedure. In one embodiment, the procedure includes the steps of pre-operatively obtaining a first textured point cloud of the cortical surface of a targeted region of a brain of a living subject, intra-operatively obtaining optically a second textured point cloud of the cortical surface of the brain of the living subject, and aligning the first textured point cloud of the cortical surface to the second textured point cloud of the cortical surface so as to register images of the brain of the living subject to the cortical surface of the living subject. | 1. A cortical surface registration procedure related to a diagnostic or surgical procedure, comprising: (a) pre-operatively obtaining a first textured point cloud of the cortical surface of a targeted region of a brain of a living subject;(b) intra-operatively obtaining optically a second textured point cloud of the cortical surface of the brain of the living subject;(c) aligning the first textured point cloud of the cortical surface to the second textured point cloud of the cortical surface so as to register images of the brain of the living subject to the cortical surface of the living subject; and(d) generating an image of the registered images,wherein at least the second textured point cloud contains intensity patterns representing sulcal-gyrus differences and contrast-enhanced vasculature related to the cortical surface of the brain of the living subject. 2. The procedure of claim 1, wherein the step of pre-operatively obtaining a first textured point cloud comprises the steps of: (a) pre-operatively acquiring an image volume from the brain of the living subject;(b) segmenting the acquired image volume;(b) extracting a point cloud representation of the brain surface geometry from the segmented image volume;(c) performing a ray-casting and voxel intensity averaging on the point cloud representation so as to generate a grayscale encoded brain surface that contains intensity patterns representing sulcal-gyrus differences and contrast-enhanced vasculature; and(d) obtaining the first textured point cloud from the grayscale encoded brain surface. 3. The procedure of claim 2, wherein the image volume of the brain of the living subject comprises image data with respect to the brain surface geometry. 4. The procedure of claim 3, wherein the image data with respect to the brain surface geometry is obtained through the use of at least one of positron emission tomography, electroencephalography, computer tomography, functional magnetic resonance imaging and magnetic resonance imaging. 5. The procedure of claim 1, wherein the step of intra-operatively obtaining optically a second textured point cloud comprises the steps of: (a) optically scanning an exposed brain surface of the living subject during surgery;(b) capturing surface-reflected light from the brain surface of the living subject;(c) acquiring a point cloud representation of the geometry of the cortical surface from the captured surface-reflected light; and(d) color-encoding the acquired point cloud representation with intensity values of a field of view so as to obtain the second textured point cloud of the cortical surface of the at least one targeted region of the brain. 6. The procedure of claim 5, wherein the step of optically scanning an exposed brain surface is performed with an optical device that is capable of obtaining frequency, intensity and geometric data with respect to the cortical surface simultaneously. 7. The procedure of claim 6, wherein the optical device is a laser ran... | Vanderbilt University | A61B005/05
US08312174 | B2 | 2012-11-13 | Protocol converter for wireless patient monitoring | 2008-01-08 | The invention relates to a wireless device for medical data transmission that comprises a wireless interface, a processor, and a medical equipment interface arranged to provide connection to an item of medical equipment. The processor is arranged to communicate with the medical equipment interface and the wireless interface, and is also 5 arranged to perform a protocol adaptation. The protocol adaptation takes place from a protocol of incoming data received at the medical equipment interface to a protocol of data to be transmitted through the wireless interface, or vice versa. The protocol adaptation comprises individual or combined filtering, gathering, partitioning, prioritizing, or discarding the incoming data. The processor is further arranged to read a configuration file containing an 10 adaptation specification that pertains to the protocol adaptation; and to perform the protocol adaptation in compliance with the adaptation specification. The invention also relates to a method for setting up the wireless device. | 1. A wireless communication system for wireless medical data transmission between first and second items of medical equipment which are configured to communicate with each other non-wirelessly with a first communication protocol, the system comprising: a first wireless device including: a first medical equipment interface arranged to provide connection to the first item of medical equipment;a first wireless interface;one or more first processors programmed to perform a protocol adaptation from at least one of: the first communications protocol of incoming data received at the first medical equipment interface to a second communications protocol of data to be wirelessly transmitted by said first wireless interface, and,the second communications protocol of incoming data received at the first wireless interface and the first communications protocol of data to be transmitted through the first medical equipment interface to the first item of medical equipment; and,a second wireless device including: a second medical equipment interface arranged to provide connection to the second item of medical equipment;a second wireless interface which wirelessly communications with the first wireless interface using the second communications protocol;one or more second processors programmed to perform a protocol adaptation from at least one of: the second communication protocol of incoming data received at the second wireless interface and the first communication protocol of data to be transmitted through the second medical equipment interface,the first communication protocol of incoming data received at the second medical equipment interface to the second communication protocol of data to be transmitted through said second wireless interface, and,said protocol adaptations comprising individual or combined filtering, gathering, partitioning, prioritizing, or discarding said incoming data;wherein said first wireless device and said second wireless device communicate through their respective wireless interfaces;wherein the first and second wireless devices further include: first and second parameter memories configured to store a configuration file containing an adaptation specification pertaining to the protocol adaptation;wherein said one or more first and second processors are further programmed to: in a configuration state, receive the configuration file via one of the medical equipment interface and the wireless interface and load the configuration file in the parameter memory;read the configuration file containing the adaptation specification that pertains to said protocol adaptation; andperform said protocol adaptation in compliance with said adaptation specification;wherein said configuration file describes said protocol adaptation by defining at least one mapping between a data pattern at said medical equipment interfaces and a data pattern at said wireless interfaces, said mapping comprising a transformation rule defining a transformation from said data p... | Koninklijke Philips Electronics N.V. | G06F015/16
US08233970 | B2 | 2012-07-31 | Method for processing cardioelectric signals and corresponding device | 2007-07-19 | Method for processing cardioelectric signals and corresponding device. The method enables processing previously sampled cardioelectric signals so as to filter the T wave, thereby improving the visualization of the P wave. The method comprises the following stages: [a] dyadic decomposition the signal into bands by calculating its wavelet transform up to a range of frequencies comprised in a first range of 15 to 150 Hz and preferably 20 to 100 Hz,[b] selection of significant bands with P wave[c] processing of the significant bands by modifying the wavelet coefficients using statistical parametric models, and preferably statistical parametric noise suppression models,[d] weighting the non-significant bands by multiplying them by a weighting function, and[e] reconstructing the signal;The invention also relates to a device for carrying out this method. | 1. A method for processing cardioelectric signals, said method being suitable for being applied to at least one cardioelectric signal from a plurality of cardioelectric signals sampled with cardioelectric signal detection means and said cardioelectric signals being obtained in the standard leads of a patient, the method comprising: dyadically decomposing said signal into a plurality of bands by calculating the wavelet transform of said signal using a mother wavelet, up to a range of frequencies comprised at least in a first range from 15 to 150 Hz;selecting at least one significant band corresponding to the frequencies of a P wave, and a plurality of non-significant bands;processing said at least one significant band, said processing comprising modifying the coefficients of said bands according to a rule based on the use of statistical parametric models;weighting said non-significant bands, said weighting comprising multiplying said non-significant bands by a weighting function; andreconstructing said signal, to obtain a reconstructed signal. 2. A method for processing cardioelectric signals according to claim 1, wherein said weighting function is a constant. 3. A method for processing cardioelectric signals according to claim 2, wherein a first significant band is selected in said selection stage, said first significant band comprising a range of 25 to 70 Hz. 4. A method for processing cardioelectric signals according to claim 1, wherein said selecting comprises selecting a first significant band in the range of 25 to 70 Hz. 5. A method for processing cardioelectric signals according to claim 1, wherein said selecting comprises selecting a second significant band in the range of 75 to 95 Hz. 6. A method for processing cardioelectric signals according to claim 1, wherein said mother wavelet is a wavelet from the Daubechies wavelet family. 7. A method for processing cardioelectric signals according to claim 6, wherein said mother wavelet from said Daubechies wavelet family is a Haar type wavelet. 8. A method for processing cardioelectric signals according to claim 1, wherein said mother wavelet is a wavelet from the Coiflet wavelet family. 9. A method for processing cardioelectric signals according to claim 1, wherein prior to dyadically decomposing said signal said method comprises: selecting a representative RR range from a plurality of RR ranges, detecting the T wave and the P wave;testing the mother wavelets by dyadic decomposition of at least one of said signals by calculating the wavelet transform of said at least one of said signals; andselecting a mother wavelet, frequency bands and a lead or composition of optimum leads based on the minimization of the energy of said previously detected T wave and the maximization of the energy of said P wave, for dyadically decomposing said signal. 10. A method for processing cardioelectric signals according to claim 1, wherein prior to dyadically decomposing said signal, a first bandpass freque... | GEM-MED S.L. | A61B005/0452
US07998071 | B2 | 2011-08-16 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2009-10-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. In a glucose sensing system comprising a glucose sensor configured for producing an output representative of glucose concentration in a host and processing circuitry configured to process the output, a method for processing data from the output, the method comprising: generating at least one data stream from the glucose sensor output;detecting a severity of a signal artifact in the at least one data stream based at least in part on a data stream content: andgenerating, using a signal estimation module, at least one estimated data stream or data stream portion for transmission or display, wherein the at least one estimated data stream is estimated based on one of a plurality of predetermined estimation algorithms, wherein the one of a plurality of predetermined estimation algorithms used for the at least one estimated data stream is selected from the plurality of predetermined estimation algorithms based on a comparison of the severity of the detected signal artifact against one or more predetermined thresholds. 2. The method of claim 1, wherein the data stream comprises a raw data stream. 3. The method of claim 1, wherein the data stream comprises a calibrated data stream of glucose concentration values. 4. The method of claim 1, wherein the signal artifact is a transient non-glucose related signal artifact. 5. The method of claim 1, wherein the detecting comprises analysis of a spectral content of the data stream. 6. The method of claim 1, wherein the detecting comprises a time series analysis. 7. The method of claim 6, wherein the time series analysis is based at least in part on a variance of the data stream over a window of data. 8. The method of claim 7, wherein the window of data comprises a time period selected from the group consisting of about 15 minutes, about 30 minutes, about 45 minutes, and about 60 minutes. 9. The method of claim 1, wherein the detecting comprises a frequency content analysis. 10. The method of claim 9, comprising monitoring the data stream for high frequency content analysis. 11. The method of claim 1, comprising filtering the data stream. 12. The method of claim 11, comprising filtering the data stream with a recursive filter. 13. The method of claim 11, comprising filtering the data stream with a high pass filter. 14. The system of claim 1, wherein generating the estimated data stream or data stream portion is initiated in response to detection of the signal artifact. 15. The system of claim 1, wherein generating the estimated signal is terminated in response to a drop in a severity of the signal artifact to a level below a predetermined threshold. 16. The system of claim 1, wherein the detecting comprises detecting an acceleration of glucose concentration higher than a threshold. 17. The system of claim 1, wherein the detecting comprises detecting a rate of change of glucose concentration higher than a threshold. 18. The method of claim 1, wherein the selected one estimation algorithm uses a ... | DexCom, Inc. | A61B005/00 A61B005/05
US08090428 | B2 | 2012-01-03 | Spinal midline indicator | 2009-11-11 | A spinal midline indicator (10) has a body (14) of radiolucent material for insertion between adjacent vertebrae (18, 20) and a radiographic marker (12) located centrally with the body to indicate the position of the spinal midline (22) in anterior-posterior images when the body is centrally located between the vertebrae. The radiographic marker is typically an elongate metal handle. The body may carry secondary radiographic markers (16) on opposite sides of and equidistant from the handle so that the handle indicates the position of the spinal midline when the body is placed centrally between the vertebrae. | 1. A spinal midline indicator comprising: a body of radiolucent material sized and shaped to fit into an intervertebral space between two adjacent vertebrae; anda radiographic marker formed as an elongated handle attached to and extending from the body along a central axis of the body in an anterior-posterior direction andat least two secondary radiographic markers on opposite sides of the elongated handle and located equidistant from the elongate handle and from lateral edges of the body. 2. A spinal midline indicator according to claim 1, wherein the handle is made of stainless steel or titanium. 3. A spinal midline indicator according to claim 1, wherein the secondary radiographic markers are embedded in the body. 4. A spinal midline indicator according to claim 1, wherein the body has a length in a direction of the handle which is smaller than a width in a direction perpendicular to the length. 5. A spinal midline indicator according to claim 1, wherein the handle extends substantially through the body. 6. A spinal midline indicator to identify a spinal midline, the indicator comprising: a body of radiolucent material sized and shaped to fit into an intervetebral space between two adjacent vertebrae; anda radiographic marker centrally disposed within the body and extending from the body to guide a marking applied to the midline of at least one of the two adjacent vertebrae andat least two secondary radiographic markers on opposite sides of the elongated handle and located equidistant from the elongate handle and from lateral edges of the body. 7. A spinal midline indicator according to claim 6, further comprising a handle is made of stainless steel or titanium. 8. A spinal midline indicator according to claim 6, wherein the secondary radiographic markers are embedded in the body. 9. A spinal midline indicator according to claim 6, wherein the marker extends substantially through the body. 10. A spinal midline indicator according to claim 1, wherein a central axis of the body is equal distance from each of two lateral edges of the body. 11. A spinal midline indicator according to claim 9, wherein the radiographic marker is disposed equal distance from each of two lateral edges of the body. | SpinalMotion, Inc. | A61F002/44 A61B017/56 A61B005/05
US08700114 | B2 | 2014-04-15 | Analyte sensor apparatuses comprising multiple implantable sensor elements and methods for making and using them | 2008-07-31 | Embodiments of the invention provide analyte sensors having optimized elements and/or configurations of elements as well as methods for making and using such sensors. Typical embodiments of the invention include glucose sensors used in the management of diabetes. | 1. An apparatus for monitoring glucose concentrations in a patient, the apparatus comprising: a flexible sensor substrate adapted to secure the apparatus to the patient;a first piercing member coupled to and extending from the flexible sensor substrate, wherein the first piercing member is operatively coupled to a first electrochemical sensor comprising a first electrochemical sensor electrode for measuring glucose concentrations at a first electrochemical sensor placement site;a second piercing member coupled to and extending from the flexible sensor substrate, wherein the second piercing member is operatively coupled to a second electrochemical sensor comprising:a second electrochemical sensor electrode for determining at least one physiological characteristic of the patient at a second electrochemical sensor placement site;a processor capable of characterizing a plurality of signals received from the first and second electrochemical sensors, wherein:the processor is capable of comparing a first signal received from a working electrode in response to a first working potential with a second signal received from a working electrode in response to a second working potential;at least one working potential is 280, 535 or 635 millivolts;the first piercing member disposes the first electrochemical sensor in a first tissue compartment of the patient and the second piercing member disposes the second electrochemical sensor in a second tissue compartment of the patient; andthe first and second piercing members are disposed on the flexible sensor substrate in a configuration selected to avoid a physiological response to the implantation of the first electrochemical sensor from affecting a sensor signal generated by the second electrochemical sensor. 2. The apparatus of claim 1, wherein the first and the second electrochemical sensor comprise: an elongated base layer;a conductive layer disposed on the base layer and comprising a reference electrode, a working electrode and a counter electrode;an analyte sensing layer disposed on the conductive layer;an analyte modulating layer disposed on the analyte sensing layer, wherein the analyte modulating layer comprises a composition that modulates the diffusion of an analyte diffusing through the analyte modulating layer; anda cover layer comprising an aperture positioned on the cover layer so as to facilitate:an analyte present in the patient contacting and diffusing through the analyte modulating layer; andcontacting the analyte sensing layer. 3. The apparatus of claim 1, wherein the first electrochemical sensor comprises a working electrode coated with glucose oxidase and the second electrochemical sensor comprises no working electrode coated with glucose oxidase and the processor is capable of obtaining information on glucose concentrations in the patient by comparing the signals received from said working electrodes so as to identify a background signal that is not based on a glucose concentrations in the p... | Medtronic MinMed, Inc. | A61B005/00 A61B005/05 A61M031/00 C12Q001/54
US08465977 | B2 | 2013-06-18 | Method and apparatus for lighted test strip | 2008-07-22 | A test strip with a sample chamber is secured to a meter. The sample chamber in the portion of the test strip that extends out of the meter is illuminated by transmitting light from a light source inside the meter internally through the test strip towards the sample chamber. By way of analogy, the test strip acts in a fashion similar to a fiber optic cable or optical wave guide by transmitting the light from the meter to the remotely located sample chamber that extends outside the meter. The user is then able to easily see the sample chamber of the test strip in dark conditions so that the user is able to readily align the sample chamber with the drop of fluid on the skin as well as view the sample chamber in order to ensure proper filling. The light also illuminates a test strip slot into which the test strip is inserted. | 1. An apparatus, comprising: a meter including a light source; anda test strip coupled to the meter, the test strip having a sampling end with a sample chamber extending outside the meter, at least a portion of the test strip being made of a light transmissive material to transmit light from the light source to the sampling end, the test strip having an intermediate portion adjacent the sampling end, the intermediate portion including a light source opening in which the light source is received. 2. The apparatus of claim 1, further comprising: the meter including a test strip opening in which the test strip is received; andthe light source being positioned in the meter to illuminate the test strip opening. 3. The apparatus of claim 1, further comprising: an integrated disposable that includes the test strip and a lancet connected to the test strip;the meter including a firing mechanism for firing the lancet into tissue; andthe light source being received in the light source opening to hold the test strip in place as the lancet is fired. 4. The apparatus of claim 1, further comprising: the test strip including a spacer layer that at least in part defines a sample chamber; andthe spacer layer being made of the light transmissive material to transmit light from the light source to illuminate the sample chamber. 5. The apparatus of claim 4, further comprising: the test strip including a cover layer covering the sample chamber, the cover layer being clear for visualization of the sample chamber. 6. The apparatus of claim 1, further comprising: the sampling end of the test strip including a fluorescent material that fluoresces in presence of the light from the light source. 7. The apparatus of claim 1, further comprising: a light source shell covering the light source. 8. The apparatus of claim 1, further comprising: the test strip including a light source edge through which the light from the light source is transmitted to the sampling end. 9. The apparatus of claim 1, further comprising: the light source being configured to emit different colors depending on stage of testing. 10. The apparatus of claim 1, further comprising: wherein the light source received in the light source opening demountably retains the test strip in the meter. 11. An apparatus, comprising: an integrated disposable including a lancet and a test strip coupled to the lancet, the test strip having a sampling end and an intermediate section adjacent the sampling end, the intermediate section including a light source opening, the test strip including a spacer layer made of light transmitting material that defines a sample chamber in the sampling end; anda meter having a light source inserted in the light source opening of the test strip to retain the test strip during lancing, the light source being configured to illuminate the spacer layer for allowing a user to visually detect fill sufficiency of the sample chamber. 12. The apparatus of claim 11, further comprising: the... | Roche Diagnostics Operations, Inc. | G01N033/49
US08613695 | B2 | 2013-12-24 | Highly precise and low level signal-generating drivers, systems, and methods of use | 2009-07-09 | Apparatus and method for providing and using a highly precise and low level driver. In one described embodiment, a driver for a magnetic field generating device is provided, the driver including a first digital to analog converter and a second digital to analog converter; a differential amplifier configured to receive a first signal from the first digital to analog converter and receive a second signal from the second digital to analog converter and output a third signal; and an attenuator to configured to receive the third signal from the differential amplifier. | 1. A driver for a magnetic field generating device comprising: a programmable signal generator configured to receive a control signal from a processor, and based on the control signal, output a driver signal;a first digital to analog converter configured to receive the driver signal and output a first output signal;a second digital to analog converter configured to receive the driver signal and output a second output signal;a differential amplifier configured to receive the first output signal and the second output signal and output a third signal;an attenuator configured to receive the third signal and output a fourth signal; andone or more coils configured to receive the fourth signal and generate a magnetic field to provide a magnetic resonance treatment to a subject. 2. The driver of claim 1, wherein the processor is configured to receive the first and second signals and calibrate the signal generator based on the first and second signals. 3. The driver of claim 1, further comprising at least one relay device connected to at least one of the first or second digital to analog converters. 4. The driver of claim 3, further comprising a voltage reference source connected to the at least one relay device. 5. The driver of claim 4, wherein the at least one relay device is used to relay output from the voltage reference source to the second digital to analog converter. 6. The driver of claim 3, wherein the signal generator is coupled to the at least one relay device. 7. The driver of claim 6, wherein the at least one relay device is used to relay output from the signal generator to the first digital to analog converter. 8. The driver of claim 1, further comprising at least one sensor configured to monitor physical parameters of the subject receiving the magnetic field. 9. The driver of claim 8, wherein the sensor comprises one or more of: an EGK, an EEG, a heart rate monitor, or a blood pressure monitor. 10. The driver of claim 1, further comprising a magnetometer. 11. The driver of claim 10, wherein the magnetometer is configured to sense an ambient magnetic field. 12. The driver of claim 1, wherein the signal generator comprises a low noise, low frequency, sinusoidal signal generator. 13. The driver of claim 1, wherein the control signal comprises one or more waveform parameters. 14. The driver of claim 13, wherein the one or more waveform parameters comprise one or more of frequency and waveform type. 15. The driver of claim 1, wherein the subject comprises a human. | Applied Magnetics, LLC | A61N002/02
US08706189 | B2 | 2014-04-22 | Combined PET/MRI device, component and local coil | 2009-07-08 | A combined PET/MRI device is disclosed. In at least one embodiment, the PET/MRI device includes an MRI unit for exciting nuclear spins in an examination volume and for receiving signals generated by the excitation in the examination volume, and a PET unit with a detector which surrounds the examination volume at least in part and is used for detecting radiation emanating from the examination volume, with, firstly, damping of the radiation emitted by the examination volume and, secondly, undesired interactions with electromagnetic fields of the MRI unit on the components of the PET/MRI device arranged between the examination volume and the detector being avoided due to the material properties and/or structural design of the components. Corresponding components such as, for example, patient couches, bearing or support apparatuses and local coils, are both MRI and PET compatible. | 1. A component of a medical imaging system, arranged between an examination volume and a detector of the imaging system during an examination, the component comprising: materials compatible with both MRI and PET, an outer layer of a first material having a tensile strength of >100 N/mm2 and a modulus of elasticity of >5 kN/m2 which encloses a filler material with a lower density or lower atomic number than the first material, wherein the materials are not electrically conducting, are not ferromagnetic, have a dielectric constant of ∈≦2 and a dielectric loss factor of less than 0.02,the component has an effective total radiation damping coefficient of 0.01 1/cm or less, andthe component is an antenna support for a local coil in the medical imaging system, wherein the filling material is stiffened by a cell structure, the cell structure being a honeycomb structure with walls having a thickness of a few centimeters. 2. The component of claim 1, wherein the walls have a thickness of one millimeter or less. 3. A component of a medical imaging system, arranged between an examination volume and a first detector of the imaging system during an examination, the component comprising: materials compatible with both MRI and PET, an outer layer of a first material having a tensile strength of >100 N/mm2 and a modulus of elasticity of >5 kN/m2 which encloses a filler material with a lower density or lower atomic number than the first material, the outer layer having a thickness of one millimeter or less, wherein the materials are not electrically conducting, are not ferromagnetic, have a dielectric constant of ∈≦2 and a dielectric loss factor of less than 0.02,the component has an effective total radiation damping coefficient of 0.01 1/cm or less, andthe component is an antenna support for a local coil. 4. The component as claimed in claim 3, wherein the materials compatible with both MRI and PET include materials which comprise polyamide fibers and resin. 5. The component as claimed in claim 4, wherein the polyamide fibers and resin include epoxy resin or polyester resin. 6. The component as claimed in claim 3, wherein the materials include materials which comprise aramid fiber tissue. 7. The component as claimed in claim 3, wherein the materials include materials composed of composite materials. 8. The component as claimed in claim 3, wherein the first material is aramid fiber tissue. 9. The component as claimed in claim 8, wherein the filling material comprises rigid or flexible foam. 10. The component as claimed in claim 3, wherein the filling material comprises rigid or flexible foam. 11. The component as claimed in claim 3, wherein the filling material is air. 12. The component as claimed in claim 3, wherein the antenna support is designed as a structure without any obstruction in a region of conductor tracks of an antenna of the local coil. 13. The component as claimed in claim 3, wherein the antenna support comprises windows in a region of ... | Siemens Medical Solutions Siems Aktiengesellschaft | A61B005/055 A61B019/00
US09351668 | B2 | 2016-05-31 | Signal processing for continuous analyte sensor | 2009-10-12 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A glucose sensor system for monitoring glucose concentration in a biological sample of a host, the system comprising: a continuous glucose sensor configured to produce sensor data indicative of a host's glucose concentration, wherein the continuous glucose sensor comprises a sensing region and a sensing membrane located over the sensing region, and wherein at least a portion of the glucose sensor is configured for transcutaneous implantation; andan electronics module configured to receive sensor data from the continuous glucose sensor, the electronics module comprising a processor configured to process sensor data received from the continuous glucose sensor for a first time period so as to generate real-time glucose readings from the sensor data, to detect an approaching clinical risk based at least in part on an acceleration of the host's glucose concentration, and to determine a possibility of avoiding the approaching clinical risk, wherein the electronics module delays activation of an alarm when the possibility of avoiding the approaching clinical risk is determined. 2. The system of claim 1, wherein detecting an approaching clinical risk further comprises determining that a predicted potential glucose concentration value is greater than a high threshold value or less than a low threshold value. 3. The system of claim 1, wherein the alarm comprises an audible signal. 4. The system of claim 1, wherein the alarm comprises a tactile signal. 5. The system of claim 2, wherein the high threshold value and the low threshold value are user configurable. 6. The system of claim 1, wherein the processor module is further configured to calibrate the sensor data using a glucose concentration measured by a reference glucose monitor operatively linked to the electronics module, wherein the reference glucose monitor is integral with the electronics module. 7. The system of claim 1, wherein detecting an approaching clinical risk further comprises predicting a potential glucose concentration for a second time period based at least in part on both an extrapolation of real time glucose readings and on a variation analysis, and wherein the variation analysis comprises a statistical analysis. 8. The system of claim 1, wherein detecting an approaching clinical risk further comprises predicting a potential glucose concentration for a second time period based at least in part on both an extrapolation of real time glucose readings and on a variation analysis, and wherein the variation analysis comprises a clinical analysis. 9. The system of claim 1, wherein detecting an approaching clinical risk further comprises predicting a potential glucose concentration for a second time period based at least in part on both an extrapolation of real time glucose readings and on a variation analysis, and wherein the variation analysis comprises a physiological analysis. 10. The system of claim 1, wherein delay is for a predetermined amount of time. 11. A device for pro... | DexCom, Inc. | A61B005/145 A61B005/1468 A61B005/1486 A61B005/1495 A61B005/00
US20100041965 | A1 | 2010-02-18 | Handheld Sleep Assistant Device and Method | 2009-08-07 | In accordance with one aspect of the present invention, an electrical sleep assistant device, which comprises a handheld case, a display panel, a plurality of biosensors, and a calculation module, is provided. The case has a surface. The display panel is disposed on the surface of the case. The plurality of biosensors are disposed on the surface of the case for collecting a plurality of physiological information. The calculation module is disposed in the case, and coupled to the display panel and the plurality of biosensors. Preferably, a plurality of buttons disposed on the surface of the case, and coupled to the calculation module. | 1. An electrical sleep assistant device, comprising: a handheld case having a surface;a display panel disposed on the surface of the case;a plurality of biosensors disposed on the surface of the case for collecting a plurality of physiological information; anda calculation module disposed in the case, and coupled to the display panel and the plurality of biosensors. 2. An electrical sleep assistant device as claimed in claim 1, further comprising a plurality of buttons disposed on the surface of the case, and coupled to the calculation module. 3. An electrical sleep assistant device as claimed in claim 1, further comprising a power supply device for providing an electrical power to drive the calculation module, the display panel and the plurality of biosensors. 4. An electrical sleep assistant device as claimed in claim 3, wherein the power supply device is a rechargeable battery. 5. An electrical sleep assistant device as claimed in claim 4, wherein the rechargeable battery is one selected from a group consisting of a lithium-ion battery, a nickel-cadmium battery and a Ni-metal-hydride battery. 6. An electrical sleep assistant device as claimed in claim 1, further comprising an earphone socket disposed on the surface of the case, and coupled to the calculation module. 7. An electrical sleep assistant device as claimed in claim 1, further comprising a universal serial bus socket disposed on the surface of the case, and coupled to the calculation module. 8. An electrical sleep assistant device as claimed in claim 1, wherein the calculation module analyzes the plurality of physiological information, outputs an analyzed result to the display panel, and receives the plurality of physiological information from the plurality of biosensors. 9. An electrical sleep assistant device as claimed in claim 1, wherein the calculation module is an electronic device including at least one of a calculation processor and a central processing unit, and a sleep assistant program is installed in the calculation module. 10. An electrical sleep assistant device as claimed in claim 1, wherein the display panel is one selected from a group consisting of a touch-screen panel, a flat-panel display, and a liquid-crystal display. 11. An electrical sleep assistant device as claimed in claim 1, wherein the plurality of biosensors include one selected from a group consisting of a biosensor for heart beat measurement, a biosensor for body temperature measurement, a biosensor for blood pressure measurement, a biosensor for skin conductivity measurement, a biosensor for breath frequency measurement and a combination thereof. 12. An electrical sleep assistant device as claimed in claim 1, wherein the plurality of physiological information include one selected from a group consisting of a heart beat, a body temperature, a blood pressure, a skin conductivity, a breath frequency and a combination thereof. 13. An electrical sleep assistant device as claimed in claim 1, where... | National Taiwan University | A61B005/00
US20100010330 | A1 | 2010-01-14 | WIRELESS MONITOR FOR A PERSONAL MEDICAL DEVICE SYSTEM | 2009-09-22 | A monitor device for a fluid infusion system and its operating, display, and data processing characteristics are described herein. One embodiment of the monitor device is used in an insulin infusion system having an insulin infusion pump and a continuous glucose sensor transmitter. The monitor device is configured as a wireless bedside monitor that wirelessly receives status data from a device in the fluid infusion system, such as the infusion pump or the sensor transmitter. The monitor device supports a number of user interface features, alarm/alert features, and graphical display features, where such features enhance the overall operation and user-friendliness of the monitor device. For example, the monitor device can generate status icons that graphically indicate the time remaining for an exhaustible operating quantity of a device in the infusion system (e.g., a battery charge, a fluid reservoir volume, or a calibration or replacement period). The monitor device can also estimate future measurements of a physiological characteristic of a monitored patient, based upon empirical measurement data received by the monitor device. | 1. A monitor device for a medical device system, the monitor device comprising: a display element;a display controller/driver coupled to the display element and configured to generate a screen for rendering on the display element; anda processing architecture configured to: analyze sensor data obtained by the monitor device, the sensor data indicating empirical measurements of a physiological characteristic of a monitored patient; andestimate future measurements of the physiological characteristic based upon the sensor data; whereinthe screen comprises a predictive graph of the physiological characteristic that graphically indicates the empirical measurements and the future measurements. 2. A monitor device according to claim 1, wherein the predictive graph comprises indicia of a target zone for the physiological characteristic. 3. A monitor device according to claim 2, the processing architecture being configured to predict whether the future measurements will leave the target zone within a period of time. 4. A monitor device according to claim 3, wherein: the processing architecture is configured to obtain an estimated time corresponding to when the future measurements will leave the target zone; andthe screen comprises an indicator of the estimated time. 5. A monitor device according to claim 4, wherein the indicator comprises a countdown timer. 6. A monitor device according to claim 1, further comprising a wireless data communication module configured to wirelessly receive the sensor data from an infusion pump. 7. A monitor device according to claim 1, further comprising a wireless data communication module configured to wirelessly receive the sensor data from a continuous physiological sensor transmitter. 8. A monitor device according to claim 1, wherein: the medical device system is an insulin infusion system having a continuous glucose sensor transmitter; andthe sensor data comprises blood glucose data. 9. A monitor device according to claim 1, further comprising alarm control logic configured to generate a warning alarm when the future measurements indicate an alarm condition for the physiological characteristic of the monitored patient, wherein the display controller/driver is configured to generate an alarm screen corresponding to the warning alarm. 10. A monitor device according to claim 1, the processing architecture being configured to extrapolate the empirical measurements, resulting in extrapolated data, wherein the future measurements are based upon the extrapolated data. 11. A monitor device according to claim 1, the processing architecture being configured to perform curve fitting on the empirical measurements, resulting in curve fitted data, wherein the future measurements are based upon the curve fitted data. 12. A method of operating a monitor device for a medical device system, the method comprising: receiving sensor data that indicates empirical measurements of a physiological characteristic of a monitored patien... | MEDTRONIC MINIMED, INC. | A61B005/145
US20100049061 | A1 | 2010-02-25 | SYSTEMS AND METHODS FOR PROVIDING AUDITORY FEEDBACK DURING CATHETER PLACEMENT | 2009-08-20 | A system for providing auditory feedback may include a blood pressure monitor interface for acquiring pressure waveform data from a pressure transducer indicative of blood pressure readings over time. The system may further include an anatomic location annunciator for generating an audible signal from the pressure waveform data that conveys information about the anatomic location of the pressure transducer within the heart chamber or blood vessel. | 1. A method for providing auditory feedback about the anatomic location of a pressure transducer coupled to a catheter during insertion thereof into a heart chamber or blood vessel of a patient, the method comprising: acquiring pressure waveform data from the pressure transducer indicative of blood pressure readings over time; andgenerating an audible signal from the pressure waveform data that conveys information about the anatomic location of the pressure transducer within the heart chamber or blood vessel. 2. The method of claim 1, wherein the audible signal comprises a tone, and wherein generating comprises: modulating a pitch of the tone according to the pressure waveform data. 3. The method of claim 2, wherein a higher pitched tone represents a higher blood pressure reading in the pressure waveform data and a lower pitched tone represents a lower blood pressure reading in the pressure waveform data. 4. The method of claim 1, wherein generating comprises: matching the pressure waveform data with a pressure waveform signature for particular location within the heart chamber or blood vessel;locating a stored audio signal associated with the matching pressure waveform signature; andoutputting the stored audio signal. 5. The method of claim 4, wherein the pressure waveform signature comprises a stored representation of typical pressure waveform data for the particular location. 6. The method of claim 4, wherein the pressure waveform signature comprises a stored representation of pressure waveform data previously acquired at the particular location. 7. The method of claim 3, wherein matching comprises: determining one or more pressure waveform characteristics of the pressure waveform data; andlocating a pressure waveform signature having the one or more pressure waveform characteristics. 8. The method of claim 7, wherein at least one pressure waveform characteristic comprises a biphasic waveform with a systolic component and a diastolic component. 9. The method of claim 7, wherein at least one pressure waveform characteristic comprises a magnitude of pulse pressure determined by subtracting diastolic pressure from systolic pressure. 10. The method of claim 7, wherein at least one pressure waveform characteristic comprises a dicrotic notch at an end-systolic portion of the pressure waveform data. 11. The method of claim 7, wherein at least one pressure waveform characteristic comprises a continual increase or decrease in diastolic pressure. 12. The method of claim 7, wherein at least one pressure waveform characteristic comprises timing of pressure waveform data relative to ECG waveform data. 13. The method of claim 4, wherein the stored audio signal comprises prerecorded voice prompt describing the anatomic location of the pressure transducer within the heart chamber or blood vessel. 14. The method of claim 4, further comprising: storing a pressure waveform signature corresponding to pressure waveform data acquired by the pressure tr... | MR HOLDINGS (HK) LTD. | A61B005/0215
US20100034783 | A1 | 2010-02-11 | Medical kit and using method thereof | 2006-11-03 | An aseptic/sterile medical kits are comprising a cartilage regeneration kit, a bone regeneration kit or an umbilical cord blood storage kit in a configuration that each process performs according to functionally-specialized kit sets for each step, via division of overall processes into corresponding steps for isolation, culture, collection and storage of cells, and implantation of desired cells into target sites of the body. The cartilage is regenerated by cartilage tissue collection; chondrocyte isolation; chondrocyte medium change and subculture; preparation of chondrocyte therapy product; media for isolation/culture/preparation/cryopreservation of cells; and media for isolation/culture/cryopreservation of cells, using the cartilage regeneration kit. The bone is regenerated by bone marrow collection; osteoblast isolation; osteoblast medium change and subculture; and preparation of osteoblast therapy product, using the bone regeneration kit. Additionally, the umbilical cord blood is stored by umbilical cord blood collection; hematopoietic stem cell isolation; and cryopreservation of hematopoietic stem cells, using the umbilical cord blood storage kit. | 1-31. (canceled) 32. A cartilage regeneration of an aseptic/sterile medical kit comprising a cartilage regeneration kit (10) in a configuration that each process can be performed according to the functionally-specialized kit sets for each step, via division of the overall processes into the corresponding steps for isolation, culture, collection and storage of cells, and implantation of the desired cells into the target sites of the body, wherein the cartilage regeneration kit (10) includes a cartilage tissue collection kit (11), a chondrocyte isolation kit (12), a chondrocyte medium change and subculture kit (13), and a chondrocyte therapy product preparation kit (14), wherein the cartilage tissue collection kit (11) includes a hard tissue collection kit for collection of cartilage tissues, a universal container which is an internal container for tissue delivery, a transport bottle which is an external container for tissue delivery, a blue tip (1 mL) used to take 1 mL of a solution, a styrofoam box which is an internal box for tissue delivery under refrigeration conditions, sponge used for introduction of the transport bottle into the styrofoam box, and a Koolit refrigerant to keep the cold conditions,wherein the chondrocyte isolation kit (12) includes a Pasteur pipette for solution removal, T25 (plug seal) as a cell culture vessel with a plug-seal cap (a 25 cm2 flask), a cell strainer which is placed over a 50 mL centrifuge tube and separates cells from the solution by passing the resulting suspension of cells and culture medium through it, a tube for centrifugation to wash cells, an E-tube used upon mixing trypan blue and the cell-containing medium, a blue tip to take 1 mL of a solution, a pipette for solution transfer, a syringe filter for solution filtration, a T25 cell culture flask, and a cryo tube for cryopreservation of cells,wherein the chondrocyte medium change and subculture kit (13), includes a Pasteur pipette for solution removal, a tube for centrifugation to wash cells, a pipette for solution transfer, a pipette for transfer of a large volume of a solution, a Yellow tip for taking a small volume of a solution, an E-tube used upon mixing trypan blue and the cell-containing medium, a T75 cell culture flask, and a cryo tube for cryopreservation of cells, andwherein the chondrocyte therapy product preparation kit (14) includes a Pasteur pipette for solution removal, a cryo tube for cryopreservation of cells, an E-tube used upon mixing trypan blue and the cell-containing medium, a tube for centrifugation to wash cells, a pipette for solution transfer, a pipette for transfer of a large volume of solution, a cell strainer which is placed over a 50 mL centrifuge tube and separates cells from the solution by passing the resulting suspension of cells and culture medium through it, a filter blue tip to take a solution, a rubber cap and an aluminum cap for stopping a vial, a 1 mL V-vial used to contain cells for preparation of a chondrocyte th... | - | A61K035/12 A61B010/00 B65D081/18 B65B007/28 A61B005/15 C12N005/02 A01N001/02
US20100028259 | A1 | 2010-02-04 | USE OF ORGANIC CATION TRANSPORTERS FOR CANCER DIAGNOSIS AND THERAPY | 2007-04-18 | The present invention provides, for the first time, the finding that organic cation transporters (OCTs) are major determinants of the anticancer activity of platinum-based drugs such as oxaliplatin, and therefore have clinical significance for selecting oxaliplatin as the preferred therapy for a cancer that expresses one or more OCTs, such as colorectal cancer or liver cancer. In addition, the OCT genotype can also be used to predict oxaliplatin response or to select therapy. The present invention also provides methods of treating or inhibiting cancers that expresses one or more OCTs by administering a therapeutically effective amount of a platinum-based drug such as an oxaliplatin analog having an organic non-leaving group with an increased size. The present invention further provides methods of sensitizing a therapy resistant cancer to a platinum-based drug such as oxaliplatin by administering a therapeutically effective amount of a nucleic acid encoding an OCT. Compositions, kits, and integrated systems for carrying out the diagnostic, prognostic, and therapeutic methods of the present invention are also provided. | 1. A method of providing a prognosis for oxaliplatin cancer therapy in a subject, the method comprising the steps of: (a) contacting a sample from the subject with an antibody that specifically binds to OCT protein; and(b) determining whether or not OCT protein is expressed in the sample, thereby providing a prognosis for oxaliplatin cancer therapy. 2. The method of claim 1, wherein the cancer is selected from the group consisting of colorectal cancer, liver cancer, prostate cancer, renal cancer, bladder cancer, ovarian cancer, breast cancer, lung cancer, leukemia, non-Hodgkin's lymphoma, and multiple myeloma. 3. The method of claim 1, wherein the OCT is selected from the group consisting of OCT1, OCT2, OCT3, and combinations thereof. 4. The method of claim 1, wherein the sample is from colon, rectum, liver, kidney, bladder, prostate, ovary, bone, or lymph node. 5. The method of claim 1, wherein the antibody is a monoclonal or polyclonal antibody. 6. The method of claim 1, wherein the method further comprises genotyping the subject to determine an OCT genotype. 7. A method of providing a prognosis for oxaliplatin cancer therapy in a subject, the method comprising the steps of: (a) contacting a sample from the subject with a primer set of a first oligonucleotide and a second oligonucleotide that each specifically hybridize to an OCT nucleic acid;(b) amplifying the OCT nucleic acid in the sample; and(c) determining whether or not the OCT nucleic acid in the sample is expressed in the sample, thereby providing a prognosis for oxaliplatin cancer therapy. 8. The method of claim 7, wherein the cancer is selected from the group consisting of colorectal cancer, liver cancer, prostate cancer, renal cancer, bladder cancer, ovarian cancer, breast cancer, lung cancer, leukemia, non-Hodgkin's lymphoma, and multiple myeloma. 9. The method of claim 7, wherein the OCT is selected from the group consisting of OCT1, OCT2, OCT3, and combinations thereof. 10. The method of claim 7, wherein the sample is from colon, rectum, liver, kidney, bladder, prostate, ovary, bone, or lymph node. 11. The method of claim 7, wherein the first oligonucleotide comprises SEQ ID NO:1 and the second oligonucleotide comprises SEQ ID NO:2. 12. The method of claim 7, wherein the first oligonucleotide comprises SEQ ID NO:3 and the second oligonucleotide comprises SEQ ID NO:4. 13. The method of claim 7, wherein the first oligonucleotide comprises SEQ ID NO:5 and the second oligonucleotide comprises SEQ ID NO:6. 14. The method of claim 7, wherein the method further comprises genotyping the subject to determine an OCT genotype. 15. A method of providing a prognosis for oxaliplatin cancer therapy in a subject by determining the genotype of an OCT gene, the method comprising the steps of: (a) contacting a sample from the subject with an antibody that specifically binds to an OCT protein encoded by a selected allele; and(b) determining whether or not the OCT protein is expressed in... | The Regents of the University of California | A61K049/00 C12Q001/68 G01N033/53 A61B005/05
US20100049080 | A1 | 2010-02-25 | THIXOTROPIC SWALLOW CHALLENGE MEDIUM | 2009-10-23 | A swallow challenge medium is thixotropic for easy swallowing and to provide enough viscosity for effective challenge to peristalsis and has high ionic density for effective impedance measurements by contact with electrodes positioned in a person's esophagus or oropharynx during swallow testing. The medium also has a high surface tension so as not to adhere to or coat the electrodes or probe surfaces. These physical characteristics are stabilized and consistent enough to provide standard for esophageal and/or oropharyngeal function testing and diagnostics. | 1. A swallow challenge medium comprising a thixotropic material having a low shear viscosity in a range of 391,000 to 529,000 centipoises, a medium shear viscosity in a range of 76,500 to 103,500 centipoises, and a high shear viscosity in a range of 11,475 to 15,525 centipoises. 2. The swallow challenge material of claim 1, wherein the material also has conductivity in a range of 4.8 to 6.6 mS/cm. 3. The swallow challenge material of claim 1, wherein the material also has impedance in a range of 340 to 460 ohms. 4. The swallow challenge material of claim 1, wherein the material also has pH in a range of 3.8 to 5.2. | SANDHILL SCIENTIFIC, INC. | A61B005/053
US20100036268 | A1 | 2010-02-11 | Circulatory monitoring systems and methods | 2008-08-07 | Systems and methods are described for obtaining and acting upon information indicative of circulatory health and related phenomena in human beings or other subjects. | 1. A medical or veterinary monitoring system comprising: a sensor;a support module configured to position the sensor in a vicinity of a body portion for a period of at least about an hour;circuitry for receiving from the sensor first data indicative of whether the body portion exhibits one or more clotting symptoms within an earlier part of the period; andcircuitry for receiving from the sensor second data indicative of whether the body portion exhibits one or more clotting symptoms within a later part of the period. 2. The medical or veterinary system of claim 1, in which the sensor comprises: circuitry for determining from the first data and from the second data a direction of an apparent movement of an object. 3. The medical or veterinary system of claim 1, in which the sensor comprises: circuitry for determining from the first data and from the second data whether a duration of an apparent movement exceeds a threshold, wherein the threshold is at least a minute. 4. The medical or veterinary system of claim 1, in which the sensor comprises: circuitry for determining from the first data and from the second data one or more spatial proportions of an object within the body portion. 5. The medical or veterinary system of claim 1, in which the sensor comprises: circuitry for absorbing infrared light from the body portion. 6. The medical or veterinary system of claim 1, in which the support module comprises: a physiological support configured to position the sensor in the vicinity of the body portion for the period, the period being at least about a day. 7. The medical or veterinary system of claim 1, in which the support module comprises: a physiological support configured to maintain contact between the sensor and the body portion for more than an hour. 8. The medical or veterinary system of claim 1, in which the sensor comprises: one or more of an optical sensor or an auditory sensor. 9. The medical or veterinary system of claim 1, in which the sensor comprises: one or more of a thermal sensor or a pressure sensor. 10. The medical or veterinary system of claim 1, in which the sensor comprises: a conductivity sensor. 11. The medical or veterinary system of claim 1, in which the support module comprises: a garment. 12. The medical or veterinary system of claim 1, in which the support module comprises: at least one of a bandage or a compression stocking. 13. The medical or veterinary system of claim 1, in which the support module comprises: at least one a cast or a wrap. 14. The medical or veterinary system of claim 1, in which the support module comprises: a seat. 15. The medical or veterinary system of claim 1, in which the support module comprises: a bed. 16. The medical or veterinary system of claim 1, further comprising: circuitry for detecting a thermal normalization rate decrease from at least the first data. 17. The medical or veterinary system of claim 1, further comprising: circuitry for activating an electromagnetic radiat... | Searete LLC, a limited liability corporation of the State of Delaware | A61B005/026
US08630701 | B2 | 2014-01-14 | Frequency domain monitoring of myocardial electrical stability | 2009-08-28 | Embodiments of the present invention relate to implantable systems, and methods for use therewith, for monitoring myocardial electrical stability. A patient's heart is paced for a period of time using a patterned pacing sequence that repeats every N beats, and an electrical signal is obtained that is representative of a plurality of consecutive beats of the patient's heart while it is being paced using the patterned pacing sequence that repeats every N beats. Myocardial electrical stability is then analyzed using frequency domain techniques that are tailored to the patterned pacing sequence used to pace the patient's heart. In other embodiments, the patient's heart need not be paced. This abstract is not intended to be a complete description of, or limit the scope of, the invention. | 1. A method for detecting an N beat alternans pattern from an electrical signal representative of a plurality of consecutive beats, where N is an integer that is at least 3, the method comprising: (a) dividing the plurality of consecutive beats into a plurality of sets of N consecutive beats;(b) for each set of N consecutive beats, selecting 2 of the N beats to produce a sub-set of 2 beats per set of N consecutive beats;(c) transforming time domain data, associated with the plurality of sub-sets of 2 beats, to frequency domain data; and(d) determining, based on the frequency domain data, whether electrical alternans are present. 2. The method of claim 1, wherein step (d) includes: determining, based on the frequency domain data, an alternans magnitude at 0.5 cycles/beat; anddetermining, based on the alternans magnitude at 0.5 cycles/beat, whether electrical alternans are present. 3. The method of claim 1, wherein the selecting 2 of the N beats, for each set of N consecutive beats at step (b), comprises selecting 2 consecutive beats from each set of N beats. 4. The method of claim 3, wherein the selecting 2 of the N beats, for each set of N consecutive beats at step (b), comprises selecting the first 2 beats from each set of N beats. 5. A system for detecting an N beat alternans pattern from an electrical signal representative of a plurality of consecutive beats, where N is an integer that is at least 3, the method comprising: means for dividing the plurality of consecutive beats into a plurality of sets of N consecutive beats;means for selecting 2 of the N beats to produce a sub-set of 2 beats per set of N consecutive beats, for each set of N consecutive beats;means for transforming time domain data, associated with the plurality of sub-sets of 2 beats, to frequency domain data; andmeans for determining whether electrical alternans are present, based on the frequency domain data. 6. The system of claim 5, wherein the means for determining: determines, based on the frequency domain data, an alternans magnitude at 0.5 cycles/beat; anddetermines, based on the alternans magnitude at 0.5 cycles/beat, whether electrical alternans are present. 7. The system of claim 5, wherein the means for selecting selects 2 consecutive beats from each set of N beats. 8. The system of claim 7, wherein the means for selecting selects the first 2 beats from each set of N beats. 9. A system for monitoring myocardial electrical stability, comprising: one or more pulse generator configured to pace a patient's heart for a period of time using a patterned pacing sequence that repeats every N beats, where N is an integer that is at least 3;one or more sensing circuit configured to obtain an electrical signal that is representative of a plurality of consecutive beats of the patient's heart while the patient's heart is being paced using the patterned pacing sequence that repeats every N beats; anda controller configured to divide the plurality of consecutive beats into ... | Pacesetter, Inc. | A61B005/04
US08688225 | B2 | 2014-04-01 | Posture state detection using selectable system control parameters | 2009-04-30 | Techniques for detecting and classifying a posture state of a patient are disclosed, wherein a posture state includes at least one of a posture and/or an activity state related to motion. In one embodiment, one or more signals indicative of at least one of posture and activity state of the patient may be sensed by a sensor. Control logic may be provided to process the one or more sensed signals in a selectable manner. This selectable manner may be based on current posture state data describing a posture state in which the patient has previously been classified. Alternatively or additionally, this current posture state data may describe a posture state transition previously undergone by the patient. The one or more signals that are so processed may then be used to classify a posture state of the patient. | 1. A medical device system, comprising: a storage device configured to store current posture state data describing a posture state in which a patient has previously been classified;a sensor configured to sense one or more signals indicative of at least one of a posture and an activity state of the patient; andcontrol logic configured to process the one or more signals in a selectable manner that is based on the current posture state data, and after processing of the one or more signals in the selectable manner, to utilize the processed one or more signals to classify a patient as being in a posture state, wherein the posture state in which the patient has previously been classified is a posture state in which the patient was most recently classified. 2. The system of claim 1, wherein the sensor is a three-axis accelerometer. 3. The system of claim 1, wherein the sensed one or more signals are sampled at a sampling rate selected based on the current posture state data. 4. The system of claim 1, wherein the sensed one or more signals are filtered to retain selected frequencies that are selected based on the current posture state data. 5. The system of claim 1, wherein the control logic includes an M-of-N filter to identify a posture state indicated by M of N sampled posture states, and wherein at least one of M and N is selected based on the current posture state data. 6. The system of claim 1, wherein the control logic includes an episode detector to determine whether a processed one of the one or more signals crosses a signal level identified by a transition threshold and does not re-cross the transition threshold for at least a time period identified by a transition duration, wherein at least one of transition threshold and transition duration are selected based on the current posture state data. 7. The system of claim 1, further including response logic configured to initiate a response based on the posture state indicated by the processed one or more signals. 8. The system of claim 7, wherein the response is further based on the current posture state data. 9. The system of claim 7, wherein the control logic includes dwell time logic configured to impose a dwell time before the response logic initiates the response. 10. The system of claim 9, wherein the dwell time is selected based on the current posture state data. 11. The system of claim 7, wherein the response involves a therapy delivered to the patient. 12. The system of claim 1, wherein the control logic updates the current posture state data to indicate the posture state indicated by the processed one or more signals. 13. The system of claim 1, wherein the control logic processes the one or more signals to derive a vector indicative of at least one of posture and activity state. 14. The system of claim 1, wherein the control logic processes the one or more signals to derive an activity level indicative of activity state. 15. A method, comprising: receiving a signal from a s... | Medtronic, Inc. | A61N001/00
US08660625 | B2 | 2014-02-25 | Signal processing systems and methods for analyzing multiparameter spaces to determine physiological states | 2008-10-01 | The present disclosure relates to signal processing systems and methods, and more particularly, to systems and methods for analyzing multiparameter spaces to determine changes in a physiological state. In embodiments, a first signal and a second signal may be obtained, from which a first plurality of values of a physiological parameter may be determined. At least one of the signals also may be used to generate a scalogram derived at least in part from the signal. A second plurality of values may be determined based at least in part on a feature in the scalogram. The first and second plurality of values may then be associated, and a physiological state may be analyzed using the associated first and second values. In an embodiment, the signals may be PPG signals and the associated first and second values may include a parameter scatter plot that may permit a user to determine changes in a patient's ventilation state over time. | 1. A method for analyzing a physiological state, comprising: obtaining a first signal and a second signal;determining a first plurality of values of a physiological parameter using at least the first signal and the second signal;deriving a scalogram from the first signal;determining a second plurality of values based at least in part on a feature in the scalogram;associating the first plurality of values with the second plurality of values; andanalyzing the physiological state based at least in part on the associated first and second plurality of values. 2. The method of claim 1, wherein the first signal and the second signal are photoplethysmograph signals from a user. 3. The method of claim 2, wherein the analyzing the physiological state comprises analyzing a ventilation state of the user. 4. The method of claim 2, wherein the physiological parameter corresponds at least in part to blood oxygen saturation of the user. 5. The method of claim 1, wherein the determining a second plurality of values based at least in part on a feature in the scalogram comprises selecting a plurality of scale values based at least in part on a respiration ridge in the scalogram. 6. The method of claim 1, wherein the determining a second plurality of values based at least in part on a feature in the scalogram comprises deriving a plurality of respiration rate values based at least in part on a respiration ridge in the scalogram. 7. The method of claim 1, wherein the associating the first plurality of values with the second plurality of values comprises deriving a parameter scatter plot. 8. The method of claim 7, wherein the analyzing the physiological state further comprises analyzing whether the parameter scatter plot is located within a threshold region. 9. The method of claim 8, wherein the threshold region is calibrated based at least in part upon at least a first portion of the first signal and at least a first portion of the second signal. 10. The method of claim 8, wherein the analyzing the physiological state further comprises calculating a slope value from the parameter scatter plot. 11. The method of claim 10, wherein the slope value having a lower value than a threshold value indicates a generally poor physiological state. | Covidien LP | A61B005/026 A61B005/0452
US07951080 | B2 | 2011-05-31 | On-body medical device securement | 2009-10-30 | Devices and methods for maintaining a medical device on-body are provided. Embodiments include medical device securement systems having first and second on-body securement elements. Also provided are systems and kits for use maintaining a medical device on-body. | 1. A method for maintaining a control unit of an analyte monitoring system on a body part of a subject using an on-body securement element comprising an adhesive surface and a non-adhesive surface that is opposite to the adhesive surface, the on-body securement element being protected by a first removable layer adjacent the adhesive surface and a second removable layer adjacent the non-adhesive surface, the method comprising: removably attaching the analyte monitoring system control unit to a body part of a patient;after attaching the control unit to the body part, positioning the on-body securement element about the control unit so that the adhesive surface contacts the control unit; andremoving the first and second removable layers from the on-body securement element. 2. The method of claim 1, further comprising attaching the on-body securement element to a mounting unit. 3. The method of claim 1, further comprising controlling a continuous analyte monitoring device with the control unit. 4. The method of claim 3, wherein the analyte is glucose. 5. The method of claim 3, wherein the continuous analyte monitoring device is a glucose sensor, wherein at least a portion of the sensor is positionable under the skin of the subject. 6. The method of claim 1, further comprising maintaining rigidity of the on-body securement element by at least one of the removable layers. 7. The method of claim 1, wherein the on-body securement element further comprises a hole therein, and wherein positioning of the on-body securement element about the control unit comprises positioning the securement element hole around the control unit. 8. The method of claim 1, wherein at least the first removable layer comprises a plurality of release liners. 9. The method of claim 8, further comprising removing at least a first release liner of the first removable layer after attaching the control unit to the body part and prior to positioning the on-body securement element about the control unit. 10. The method of claim 9, further comprising removing at least a second release liner of the first removable layer after positioning the on-body securement element about the control unit. 11. The method of claim 1, wherein upon positioning the on-body securement element about the control unit, the adhesive surface faces the body part and the non-adhesive surface faces away from the body part. 12. A method for maintaining a control unit of an analyte monitoring system on a body part of a subject, the method comprising: attaching a first adhesive surface of a first securement element to a body part;attaching a second adhesive surface of the first securement element to the control unit, wherein the second adhesive surface is opposite the first adhesive surface;removing at least one removable liner from an adhesive surface of a second securement element, wherein the second securement element further comprises a hole therein; andpositioning the second securement element about the ... | Abbott Diabetes Care Inc. | A61B005/00
US08383058 | B2 | 2013-02-26 | Quantitative analysis of a biological sample of unknown quantity | 2009-09-22 | Disclosed is a method for testing a modified specimen such as a dried blood spot, plasma or serum specimen, for an analyte of interest, such as cholesterol. In accordance with the disclosed subject matter, the level of the analyte of interest in the medium from which the modified specimen was obtained (e.g., from a patient's blood) is determined based on the level of an analyte in a solution formed from the modified specimen and on the level of at least one normalizing analyte. The analyte and normalizing analyte each may be an ion, compound, biochemical entity, or property of the specimen. Also disclosed are a fluid collector and a fluid collection device. | 1. A kit comprising: a fluid collector with an absorbent substrate coated with a saccharide, said substrate comprising a mat of glass fibers at least substantially coated with polyvinyl alcohol, said fibers defining a plurality of pores, the pores in said mat having a pore size effective to at least substantially prevent lysing of red blood cells while permitting at least substantial separation of serum from red blood cells via differential wicking and a superstrate, said fluid collector being generally fixed with respect to said superstrate, said superstrate having an aperture defining a blood receiving opening and permitting access to said fluid collector, and further comprising a desiccant, said desiccant being present in an amount effective to provide a desiccating protective effect on a blood fluid specimen. 2. A method for collecting a specimen from a patient, comprising: providing a fluid collection device with a fluid collector, said fluid collector comprising an absorbent substrate coated with a saccharide, said substrate comprising a mat of glass fibers at least substantially coated with polyvinyl alcohol, said fibers defining a plurality of pores, the pores in said mat having a pore size effective to at least substantially prevent lysing of red blood cells while permitting at least substantial separation of serum from red blood cells via differential wicking; and a superstrate fixed with respect to the fluid collector; and allowing said patient to bleed onto said collector until at least a predetermined adequate amount of blood has been deposited onto said collector, further comprising sending the fluid collection device to a remote location for testing. 3. The method according to claim 2, comprising sealing the collector in a barrier film pouch. 4. The method according to claim 3, said barrier film pouch comprising a laminar structure that includes a polyester film and an aluminum foil film. 5. The method according to claim 3, further comprising adding a desiccant to said barrier film pouch, said desiccant being present in an amount effective to provide a desiccating protective effect on a blood fluid specimen. 6. The method according to claim 5, said desiccant comprising silica. 7. The method according to claim 2, further comprising receiving a results reporting form after sending said fluid collection device to a remote location for testing. 8. The method according to claim 2, further comprising receiving a results reporting form after sending said fluid collection device to a remote location for testing. 9. A method for collecting a specimen from a patient, comprising: providing a fluid collection device with a fluid collector, said fluid collector comprising an absorbent substrate coated with a saccharide, said substrate comprising a mat of glass fibers at least substantially coated with polyvinyl alcohol, said fibers defining a plurality of pores, the pores in said mat having a pore size effective to at least substantiall... | Home Access Health Corporation | B01L003/00 G01N033/555
US09327129 | B2 | 2016-05-03 | Blended posture state classification and therapy delivery | 2009-04-30 | Techniques for classification of a posture state of a patient using multiple posture state definitions, and for delivering therapy according to the patient's classified posture state. Detected parameter values describing a patient's posture state are compared to posture state definitions. This comparison is used to determine similarity values describing how similar the patient's posture state is to each of the posture states described by the posture state definitions. Weighting factors may be determined from the similarity values and used to weight therapy parameter values that are associated with each of the posture state definitions. The resulting weighted therapy parameter values may be used to derive a blended therapy parameter value for use in delivering therapy to the patient. The patient's posture state may be expressed in terms of a blending of the multiple posture state definitions. | 1. A method, comprising: obtaining a detected parameter value for use in classifying a posture state of a patient;for each of multiple posture state definitions, obtaining a respective defined parameter value and a respective therapy parameter value;determining, for each of a plurality of the posture state definitions, a similarity between the detected parameter value and the respective defined parameter value;determining a blended therapy parameter value based on the similarities and respective therapy parameter values for the plurality of the posture state definitions; anddelivering therapy to the patient via a medical device according to the blended therapy parameter value. 2. The method of claim 1, wherein the detected parameter value is indicative of a posture of the patient, each of the defined parameter values is indicative of a respective defined posture, and wherein determining, for each of a plurality of the posture state definitions, a similarity comprises determining, for each of the plurality of the posture state definitions, how close the posture of the patient is to the defined posture that is indicated by the respective defined parameter value. 3. The method of claim 2, wherein the detected parameter value comprises a detected posture vector, wherein the respective defined parameter value comprises a respective defined posture vector, and wherein determining, for each of a plurality of the posture state definitions, a similarity further comprises determining, for each of the plurality of the posture state definitions, a cosine of an angle between the detected posture vector and the respective defined posture vector. 4. The method of claim 1, wherein determining, for each of a plurality of the posture state definitions, a similarity between the detected parameter value and the respective defined parameter value comprises selecting the plurality of the posture state definitions. 5. The method of claim 1, further comprising determining a respective weighting factor W for each of the similarities and wherein determining a blended therapy parameter value comprises using the respective weighting factor to weight a respective therapy parameter value. 6. The method of claim 5, wherein each weighting factor is determined using a weighting function. 7. The method of claim 6, wherein determining, for each of a plurality of the posture state definitions, a similarity between the detected parameter value and the respective defined parameter value comprises determining a distance between the detected parameter value and the respective defined parameter value, and further comprising selecting the weighting function to map the distance to a value between 0 and a predetermined maximum positive number. 8. The method of claim 5, wherein each of the plurality of posture state definitions is included in a set I and the respective therapy parameter value is a therapy parameter value P, and wherein the blended therapy parameter value is determine... | Medtronic, Inc. | A61N001/00 A61N001/365 A61B005/11 A61B005/00
US09675265 | B2 | 2017-06-13 | Feeding tube | 2007-12-07 | The invention is related to a feeding tube (1, 11, 15) in particular for total parental nutrition and/or medicine dosing including at least one inner tubing (9), at least one lumen (2, 13), at least one sensing element, in particular an electrode (3), the sensing element being connected to at least one monitoring device by a wiring (4, 12, 17) for internal monitoring of a patient's vital functions. The wiring (4, 12, 17) is at least partially wounded in tight contact with a surface of the inner tubing (9). Wherein the inner tubing (9) is at least partially surrounded by an outer tubing (8) in order to cover the inner tubing (9) and/or the wiring (4, 12, 17). | 1. A feeding tube configured for insertion in an esophageal opening including: a flexible inner tubing including at least one lumen configured for parenteral nutrition and/or medicine dosing;at least one sensing element disposed along one end of an outside surface of the inner tubing and configured for internally sensing a patient's vital functions;electrically conductive wire connected to the at least one sensing element and to at least one monitoring device, and disposed on the outside surface of the inner tubing;an outer tubing which covers the inner tubing and the electrically conductive wire, a surface of the outer tubing having a spiral shape; andwherein the inner tubing has recesses pre-formed on the inner tubing surface that match with the dimensions of and receive the at least one sensing element and the wire. 2. The feeding tube according to claim 1, wherein the outer tubing includes a shrinkable tube which is shrunk fit over the inner tubing and wire by shrinking. 3. A feeding tube configured for insertion in an esophageal opening including: an inner tubing including at least one inner lumen configured for parenteral nutrition and/or medicine dosing and at least one outer lumen disposed adjacent an outside surface of the inner tubing with an opening extending continuously along the at least one outer lumen between the outer lumen and an outside surface of the inner tubing;at least one sensing element disposed adjacent one end on the outside surface of the inner tubing and configured for internally sensing a patient's vital functions;wire connected to the at least one sensing element and to at least one monitoring device, and disposed in the outer lumen of the inner tubing;an outer tubing which covers the inner tubing and the wire; andwherein the inner tubing is sufficiently flexible that radial pressure widens the opening to facilitate positioning the wire in the outer lumen. 4. The feeding tube according to claim 3, wherein the sensing element includes a coil of the wire on said inner tubing. 5. The feeding tube according to claim 3, wherein the inner and outer tubing are formed of a polymer. 6. The feeding tube according to claim 3, wherein the at least one sensing element includes an electrode, and the electrode includes at least one of: a thin metal film; ora highly conductive polymer. 7. The feeding tube according to claim 3, wherein the at least one sensing element includes an inert metal alloy. 8. A feeding tube, comprising: at least one sensing element which internally senses a patient's vital functions;wire connected to the at least one sensing element and to a contact;a tubing which includes a plurality of lumens, wherein a first lumen is formed in the center of the tubing and configured for feeding, and at least a second lumen formed between the first lumen and a surface of the tubing, and the wire being disposed in the second lumen, and the at least one sensing element being disposed on a surface of the tubing at a d... | KONINKLIJKE PHILIPS N.V. | A61B005/04 A61B005/042 A61J015/00
US20100022916 | A1 | 2010-01-28 | Method and Apparatus for Collecting and Preparing Biological Samples for Testing | 2008-07-24 | A kit and a method are disclosed for collecting and preparing a biological sample for testing where the sample is to be mixed with a buffer prior to being tested. | 1. A method for collecting and preparing a biological sample to be tested, comprising: obtaining the sample with a swab, the swab having a sorbent mounted on the end of a stick;obtaining a dropper bottle assembly having a dropper cap with a spout, a threaded base and a bottle with a threaded neck, the bottle containing a buffer solution;opening the dropper bottle assembly by unscrewing and removing the cap with the spout;inserting the sorbent end of the swab into the bottle;breaking the stick leaving the sorbent and a portion of the stick in the bottle;closing the dropper bottle assembly by re-screwing the cap onto the threaded neck;agitating the bottle to mix the sample with the buffer solution;opening the dropper cap to expose the spout;inverting the bottle; anddispensing at least a portion of the mixed sample-buffer solution contents of the bottle through the spout onto a first location of said test device. 2. A method according to claim 1, wherein: said agitating is accomplished by shaking the bottle multiple times. 3. A method according to claim 1, wherein: the stick has a weakened portion at a particular place to facilitate said breaking. 4. A method according to claim 1, wherein: said dropper cap has a hinged cover. 5. A method according to claim 4, wherein: said opening the dropper cap comprises rotating the hinged cover. 6. A method according to claim 1, further comprising: obtaining a second dropper bottle containing buffer solution; anddispensing at least a portion of said buffer solution of said second dropper bottle onto a second location different than said first location on said test device. 7. A kit for collecting and preparing a biological sample for testing, said kit comprising: a sterile swab having a stick with a sorbent at one end of the stick, said stick having a weakened portion to facilitate breaking the stick at a pre-selected location;a dropper bottle assembly, said assembly including a screw cap having a spout, a locking hinged cover and a threaded base and a first bottle with a threaded neck, said first bottle containing a liquid buffer and said screw cap being coupled to said threaded neck, whereinthe weakened portion of the stick located a distance from the end of the sorbent approximately equal to the height of the first bottle. 8. A kit according to claim 7, further comprising: a second dropper bottle containing a liquid buffer. 9. A kit according to claim 8, further comprising: an alcohol swab, a safety lancet, and a bandage. 10. A kit according to claim 8, further comprising: an immunoassay device having a first opening identified for receiving drops from said first bottle and a second opening identified for receiving drops from said second bottle. 11. A kit according to claim 7, further comprising: an immunoassay device having an opening for receiving drops from said first bottle. | - | A61B005/00 A61B010/00
US07660624 | B2 | 2010-02-09 | Method and x-ray system for detecting position changes of a medical implant | 2004-04-12 | In a method and apparatus to detect position changes of an implant x-ray detectable markers are arranged in the environment of the implant, and 2D x-ray exposures are obtained at temporal intervals of a region containing the implant in which a distribution of the markers as well as of marked points of the implant are visible in each 2D x-ray exposure. A first of the 2D x-ray exposures is obtained at a first point in time from one projection direction, and a second of the 2D x-ray exposures is obtained at a second point in time from another projection direction. The distribution of the markers and marked points is determined in the first and second 2D x-ray exposure by an evaluation device, and from the distributions a degree of probability is calculated that the distribution of the first 2D x-ray exposure and the distribution of the second 2D x-ray exposure are projections of the same three-dimensional distribution of markers and marked points. An automated detection of position changes by implants with reduced x-ray exposure for the patient is facilitated. | 1. A method for detecting position changes of a medical implant in a patient, comprising the steps of: disposing a plurality of x-ray detectable markers in an anatomical environment of an implanted medical implant, said markers being spatially separated from said medical implant and said medical implant having x-ray detectable points thereon different from said markers;obtaining one first 2D x-ray exposure, from a first projection direction, of a region of the patient containing said implant and said anatomical environment at a first point in time, in which a first distribution of said markers and said points is detectable;obtaining one second 2D x-ray exposure, from a second projection direction different from said first projection direction, of said region at a second point in time, in which a second distribution of said markers and said points is detectable, said second point in time being temporally separated from said first point in time such that a positional change of said implant in said environment may have occurred;electronically detecting said first and second distributions respectively in said first and second 2D x-ray exposures; andin a processor, automatically calculating from said first and second distributions, with no 3D reconstruction of spatial positions of said markers and points, a degree of probability that said first and second distribution represent projections of the same three-dimensional distribution of said markers and said points and, from said degree of probability, determining whether a positional change of said implant in said environment has occurred and, from said processor, emitting an indication that identifies only whether said positional change has occurred or not occurred. 2. A method as claimed in claim 1 wherein the step of determining whether said positional change of the implant has occurred comprises defining a degree of probability threshold, and automatically electronically generating said indication that a positional change of said implant has occurred if said degree of probability does not exceed said threshold. 3. A method as claimed in claim 1 comprising detecting said first and second distributions by digital image processing of the respective first and second 2D x-ray exposures. 4. A method as claimed in claim 1 comprising: at said first point in time, also obtaining a first calibrated 2D x-ray exposure of said region;from said first 2D x-ray exposure and said first calibrated 2D x-ray exposure, calculating a first 3D distribution of said markers and points for said first point in time;at said second point in time, obtaining a second calibrated 2D x-ray exposure of said region;from said second 2D x-ray exposure and said second calibrated 2D x-ray exposure, calculating a second 3D distribution of said markers and said points for said second point in time; andcomparing said first 3D distribution and said second 3D distribution and, from said comparison, calculating a magnitude of said positiona... | Siemens Aktiengesellschaft | A61B005/05
US07651468 | B2 | 2010-01-26 | Urological instrument for assessing a urine flow | 2003-12-15 | A urological instrument (1) for assessing a urine flow, includes a receiving member (5) for the urine flow which, at least during operation, is in open liquid communication with a urine guide (6). The urine guide (6) leads to an inlet of flow indicator element (7) for assessing at least a magnitude of the urine flow. The receiving member (5) has a first, relatively compact storage state and a second, expanded ready-to-use state and is designed and adapted to allow it to be brought manually from the storage state to the ready-to-use state. At least in the ready-to-use state the receiving member comprises a cavity open on at least one side for receiving the urine flow therein. | 1. A urological measuring instrument for determining a magnitude of urine flow, the instrument comprising: a receiving member for receiving the urine flow; anda urine guide in communication with the receiving member, comprising an inlet and an outlet for the urine flow, the urine guide further comprising a flow indicator for measuring the magnitude of the urine flow, the flow indicator in communication with the inlet and the outlet,wherein the instrument has a compacted storage state and an expanded ready-to-use state, and in the compacted storage state the receiving member, urine guide, and flow indicator are substantially folded onto each other with respective external surfaces facing each other and are substantially flush with each other, and in the ready-to-use state, the receiving member is open on at least one side, comprises a cavity for receiving the urine flow therein, and is in liquid communication with the inlet of the urine guide, andthe flow indicator comprises a plurality of outflow openings and side channels arranged in a side wall of the urine guide through which the urine flow occurs, whereby the number of the outflow openings through which the urine flow occurs is a measure of the magnitude of the urine flow. 2. The instrument according to claim 1, further comprising a collecting device, wherein at least in the expanded ready-to-use state, the collecting device is in liquid communication with the urine guide on an end opposite the receiving member, the collecting device comprises a closing member having a normally-closed position, and the collecting device is expandable from the compacted storage state to the expanded ready-to-use state. 3. The instrument according to claim 2, wherein the collecting device is connected to the outlet of the urine guide. 4. The instrument according to claim 2, wherein the collecting device comprises a collection bag. 5. The instrument according to claim 2, wherein the collecting device further comprises a volume indicator for measuring a urine volume received therein. 6. The instrument according to claim 2, wherein in the compacted storage state at least one of the receiving member and the collecting device covers the urine guide on one side of the flow indicator. 7. The instrument according to claim 1, further comprising an irreparable breakable seal adapted to hold the instrument in the compacted storage state and allowing the instrument to be brought into the expanded ready-to-use state only after the seal is broken. 8. The instrument according to claim 1, further comprising a registration member arranged on an outer side of the instrument for manual registration of one or more values determined with the instrument. 9. The instrument according to claim 8, wherein the registration member comprises a removable self-adhesive label. 10. The instrument according to claim 9, wherein the label is arranged at the position of the flow indicator, leaving a transparent window at the position of a... | IQ + Investments N.V. | A61B005/00 B65D081/00
US07660623 | B2 | 2010-02-09 | Six degree of freedom alignment display for medical procedures | 2003-01-30 | A display and navigation system for use in guiding a medical device to a target in a patient includes a tracking sensor, a tracking device and display. The tracking sensor is associated with the medical device and is used to track the medical device. The tracking device tracks the medical device with the tracking sensor. The display includes indicia illustrating at least five degree of freedom information and indicia of the medical device in relation to the five degree of freedom information. | 1. A navigation system for guiding a medical device to a target in a patient during a medical procedure, said navigation system comprising: a display; anda workstation operably connected to the display and configured to process information from:a tracking sensor associated with the medical device and operable to be used to track the medical device, anda tracking device operable to track the medical device with the tracking sensor to update a position of the medical device as the medical device is guided,the workstation, in response to the position of the medial device, configured to provide information for displaying on the display:a first indicia identifying a first degree of freedom information;a second indicia identifying a second degree of freedom information;a third indicia identifying a third degree of freedom information;a fourth indicia identifying a fourth degree of freedom information; anda fifth indicia identifying a fifth degree of freedom information;a sixth indicia identifying a sixth degree of freedom information; andwherein said first, second, third, fourth, fifth and sixth indicia are illustrated on said display to assist in guiding the medical device to the target, the first, second, and third indicia are X, Y, and Z axes, respectively, with said X axis represented by a horizontal bar, said Y axis represented by a vertical bar crossing said X axis, and said Z axis represented by a separate bar, with a position of the medical device as the medical device is guided represented by an arrow on each of the displayed X axis, Y axis, and Z axis, and the fourth, fifth, and sixth indicia are rotation about the X, Y, and Z axes, respectively. 2. The navigation system as defined in claim 1 wherein the indicia are illustrated in a coordinate system selected from a group comprising a Cartesian coordinate system, a spherical coordinate system, a polar coordinate system, and a combination thereof. 3. The navigation system as defined in claim 1 wherein at least one of the indicia illustrates a safe zone in relation to said corresponding degree of freedom information. 4. The navigation system as defined in claim 3 wherein each indicia illustrates a safe zone. 5. The navigation system as defined in claim 3 wherein said safe zone is illustrated by way of color, audible tones, cross-hatching and a combination thereof. 6. The navigation system as defined in claim 1 wherein the workstation further displays on the display an indicia of a tip portion of the medical device and an indicia of a hind portion of the medical device depicting the relative position and orientation of the medical device. 7. The navigation system as defined in claim 1 wherein rotation about the X axis is illustrated by a straight bar graph having an arrow identifying the amount of rotation, wherein the rotation about the Y axis is illustrated by a straight bar graph having an arrow illustrating the amount of rotation about the Y axis and the rotation about the Z axis is il... | Medtronic Navigation, Inc. | A61B005/05 G06T015/00
US07666144 | B2 | 2010-02-23 | Methods and apparatus for determining cardiac output and left atrial pressure | 2006-02-21 | Method and apparatus are introduced for determining proportional cardiac output (CO), absolute left atrial pressure (LAP), and/or other important hemodynamic variables from a contour of a circulatory pressure waveform or related signal. Certain embodiments of the invention provided herein include the mathematical analysis of a pulmonary artery pressure (PAP) waveform or a right ventricular pressure (RVP) waveform in order to determine beat-to-beat or time-averaged proportional CO, proportional pulmonary vascular resistance (PVR), and/or LAP. The invention permits continuous and automatic monitoring of critical hemodynamic variables with a level of invasiveness suitable for routine clinical application. The invention may be utilized, for example, to continuously monitor critically ill patients with pulmonary artery catheters installed and chronically monitor heart failure patients instrumented with implanted devices for measuring RVP. | 1. A method for determining an average left atrial pressure (LAP) using a cardiac cycle waveform signal, the method comprising: converting, by an analog-to-digital converter of an apparatus, a cardiac cycle waveform signal of a patient into a digital cardiac signal, the digital cardiac signal comprising a plurality of digital cardiac signal values;selecting, from a buffer system of the apparatus, two or more digital cardiac signal values from a first interval of the digital cardiac signal;fitting, by a processing unit of the apparatus, a nonlinear function to match the two or more digital cardiac signal values within the first interval;extrapolating, by the processing unit, a plurality of values of the nonlinear function for a second interval of the nonlinear function, the second interval of the nonlinear function corresponding to an interval of the digital cardiac signal that is different from the first interval; anddetermining, by the apparatus, one or more average left atrial pressure (LAP) values based on a mean value of the plurality of values of the extrapolated nonlinear function over a portion of the second interval. 2. The method of claim 1, further comprising determining, by the apparatus, a proportional cardiac output (CO) based at least in part on subtracting the one or more average LAP values from one or more digital cardiac signal values. 3. The method of claim 2, wherein the mean value is further based on a predetermined time constant, the predetermined time constant related to a duration of time within the second interval. 4. The method of claim 1, further comprising determining, by the apparatus, a proportional cardiac output (CO) based on a computed mean value that is responsive to a result calculated by subtracting the one or more average LAP values from the plurality of values of the extrapolated nonlinear function. 5. The method of claim 4, further comprising determining, by the apparatus, the proportional CO based at least in part by dividing the computed mean value by a duration of the second interval. 6. The method of claim 1, wherein the cardiac cycle waveform is a pulmonary artery pressure (PAP) signal of the patient. 7. The method of claim 1, wherein the cardiac cycle waveform is a right ventricular pressure (RVP) signal of the patient. 8. The method of claim 1, wherein the cardiac cycle waveform is a systemic arterial pressure (SAP) signal of the patient. 9. The method of claim 1, wherein the nonlinear function further comprises an exponential component. 10. The method of claim 9, wherein the nonlinear function is a mono-exponential function. 11. A method for determining an average left atrial pressure (LAP) using a cardiac cycle waveform signal, the method comprising: receiving a cardiac cycle waveform signal, the cardiac cycle waveform signal representing a cardiac output of a patient;converting the cardiac cycle waveform signal into a digital cardiac signal, the digital cardiac signal comprising a pluralit... | Board of Trustees Operating Michigan State University Massachusetts Institute of Technology | A61B005/02
US20100004565 | A1 | 2010-01-07 | SENSOR ARRANGEMENT FOR HOME REHABILITATION | 2007-12-18 | A sensor arrangement for home rehabilitation in particular after a stroke comprising at least two sensors. The sensors are attached to the user's body. Each sensor includes—a receiver for receiving a first signal being generated from a source outside the sensor 10 arrangement—a sensor processing unit processing the first signal and initializing a second signal upon reception of the first signal, the second signal including information regarding the identity of the sensor—a transmitter for transmitting the second signal to a central processing unit 15 On receipt of the first and/or the second signal a time stamp is generated for each sensor for a determination of the location of each sensor relative to a source of the first signal trough comparison of the different time stamps. | 1. A sensor arrangement (1) for home rehabilitation in particular after a stroke comprising at least two sensors (1a, 1b, 1c) to be attached to a user's body, each sensor (1a, 1b, 1c) including: a receiver for receiving a first signal being generated from a source outside the sensor arrangement (1);a sensor processing unit processing the first signal and initializing a second signal upon reception of the first signal, the second signal including information regarding the identity of the sensor (1a, 1b, 1c)a transmitter for transmitting the second signal to a central processing unit (5) wherein upon receipt of the first and/or the second signal a time stamp is generated for each sensor (1a, 1b, 1c) for a determination of the location of each sensor (1a, 1b, 1c) relative to a source of the first signal trough comparison of the different time stamps. 2. A sensor arrangement (1) according to claim 1, characterized in that the central processing unit (5) comprises a time stamp generator and a time stamp is generated in the central processing unit (5) each time a second signal has been received each time stamp being separately stored in the central processing unit (5) together with information regarding the identity of the sensor (1a, 1b, 1c). 3. A sensor arrangement (1) according to claim 1, characterized in that each sensor (1a, 1b, 1c) comprises a time stamp generator, the time stamp generators being synchronized with each other, wherein a time stamp is generated and stored in the sensor processing unit immediately after receipt of a first signal. 4. A sensor arrangement (1) according to claim 1, characterized in that the first and/or the second signal is an acoustic signal and/or an impulse signal and/or a light signal. 5. A sensor arrangement (1) according to claim 1, characterized in that the sensor processing unit compares the signals received by the receiver with a predetermined signal corresponding to the first signal, wherein a second signal is only generated by the sensor processing unit if the signal being received by the receiver matches with the predetermined signal. 6. A sensor arrangement (1) according to claim 1, characterized in that the central processing unit (5) comprises a communication means to issue instructions how to proceed with the initializing of the sensor arrangement (5). 7. Sensor arrangement (1) according to claim 1, characterized in that each sensor (1a, 1b, 1c) generates a different second signal upon receipt of the first signal to thereby simultaneously indicating its identity to the central processing unit (5). 8. Sensor arrangement (1) according to claim 1, characterized in that the sensors (1a, 1b, 1c) exchange their time stamps with each other in an all-to-all communication to determine their position relative to the source of the first signal. 9. A sensor arrangement (1) according to claim 8, characterized in that the sensors (1a, 1b, 1c) complete an ordering process after exchangement of their time stam... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/103
US07643868 | B2 | 2010-01-05 | Method and apparatus for determining the position of the tibial exit point of the anterior cruciate ligament | 2005-03-21 | In order to establish an optimum position on the upper side of the tibia for the exit point of an anterior cruciate ligament replacement, it is proposed that, with the leg bent, the position of a number of points along the front edge of the notch is detected and in this way their position and progression on the femur are also determined, that, with the leg straight, the course of this front edge of the notch is projected onto a projection plane which is perpendicular to the longitudinal axis of the tibia in the region of the upper side of the tibia, and that, within the region enclosed by this projected curve, the exit point is selected such that it maintains at least a distance corresponding to the radius of the implant from the projected curve. Furthermore, an apparatus for carrying out such a method is proposed. | 1. Method for determining a position on an upper side of a tibia of an exit point of an implant replacing an anterior cruciate ligament of a knee joint of a leg, in which positional data of the tibia and a femur are determined by means of a navigation system, in which geometrical form data of the femur in a region of the femur facing an upper side of the tibia are recorded and in which the form data are processed together with the positional data of the tibia and the femur in a data-processing system to calculate positional data for the exit point, comprising: with the leg bent, detecting respective positions of a number of points along a front edge of a notch formed by condyles of the femur such that position and progression of the points on the femur are determined,with the leg straight, projecting a course of the front edge of the notch in the form of a curve onto a projection plane which is perpendicular to a longitudinal axis of the tibia in a region of the upper side of the tibia,within a region enclosed by this projected curve, selecting the exit point such that the exit point maintains at least a distance corresponding to a radius of the implant from the projected curve, andselecting a position of the exit point in which the exit point lies as far as possible in a front part of the region enclosed by the projected curve and as close as possible to the projected curve. 2. Method for determining a position on an upper side of a tibia of an exit point of an implant replacing an anterior cruciate ligament of a knee joint of a leg, in which positional data of the tibia and a femur are determined by means of a navigation system, in which geometrical form data of the femur in a region of the femur facing an upper side of the tibia are recorded and in which the form data are processed together with the positional data of the tibia and the femur in a data-processing system to calculate positional data for the exit point, comprising: with the leg bent, detecting respective positions of a number of points along a front edge of a notch formed by condyles of the femur such that position and progression of the points on the femur are determined,with the leg straight, projecting a course of the front edge of the notch in the form of a curve onto a projection plane which is perpendicular to a longitudinal axis of the tibia in a region of the upper side of the tibia, andwithin a region enclosed by this projected curve, selecting the exit point such that the exit point maintains at least a distance corresponding to a radius of the implant from the projected curve;wherein: a circular area of a diameter corresponding to a diameter of the implant used is projected into the projection plane, this circular area is displaced in an area enclosed by the projected curve such that the circular area does not intersect the projected curve, andin one of multiple positions of the circular area achieved by such displacement, a center point of the circular area is select... | Aesculap AG | A61B005/05
US07657307 | B2 | 2010-02-02 | Method of and apparatus for classifying arrhythmias using scatter plot analysis | 2005-10-31 | A method of classifying arrhythmias using scatter plot analysis to define a measure of variability of a cardiac rhythm parameter such as for example, without limitation, R-R interval, A-A interval, and the slope of a portion of a cardiac signal, is disclosed. The variability measurement is derived from a scatter plot of a cardiac rhythm parameter, employing a region counting technique that quantifies the variability of the cardiac rhythm parameter while minimizing the computational complexity. The method may be employed by an implantable medical device or system, such as an implantable pacemaker or cardioverter defibrillator, or by an external device or system, such as a programmer or computer. The variability measurement may be correlated with other device or system information to differentiate between atrial flutter and atrial fibrillation, for example. The variability information may also be used by the device or system to select an appropriate therapy for a patient. | 1. A computer-readable medium programmed with instructions for performing a method of classifying arrhythmias, the medium comprising instructions for causing a programmable processor to: acquire a cardiac signal;derive a recurring cardiac rhythm parameter from the cardiac signal and store the parameter as a series of values;plot a point on a scatter plot for each pair of successive stored values, whereinan abscissa of the point is a first value of a pair of successive stored values, andan ordinate of the point is a second value of the pair of successive stored values,and wherein the scatter plot is divided into a plurality of regions;derive a measure of variability by counting sum up the number of regions that have one or more points plotted therein; andcompare the measure of variability to a threshold. 2. The medium of claim 1, the instructions further causing the processor to continue plotting a point on a scatter plot until at least a specified number, N, of points have been plotted such that a window of N plotted points is available for the step of deriving a measure of variability. 3. The medium of claim 1 wherein the plurality of regions includes generally rectangular portions of the scatter plot. 4. The medium of claim 1 wherein the plurality of regions includes generally curved portions of the scatter plot. 5. The medium of claim 1 wherein the recurring cardiac rhythm parameter is derived from R-R intervals of the cardiac signal. 6. The medium of claim 1 wherein the recurring cardiac rhythm parameter is derived from P-P intervals of the cardiac signal. 7. The medium of claim 1 wherein the recurring cardiac rhythm parameter is derived from an amplitude of a portion of the cardiac signal. 8. The medium of claim 7 wherein the recurring cardiac rhythm parameter is a slope of a portion of the cardiac signal. 9. The medium of claim 1, the instructions further causing the processor to select a therapy based on the result of comparing the measure of variability to a threshold. 10. The medium of claim 1, the instructions further causing the processor to measure a rate derived from a recurring cardiac rhythm parameter. 11. The medium of claim 10, the instructions further causing the processor to compare the rate to a threshold. 12. The medium of claim 11, the instructions further causing the processor to classify an arrhythmia based on the steps of comparing the measure of variability to a threshold and comparing the rate to a threshold. 13. The medium of claim 12 wherein the arrhythmia is classified as atrial fibrillation (AF) when the rate and variability meet specified AF criteria. 14. The medium of claim 12 wherein the arrhythmia is classified as atrial flutter (AFL) when the rate and variability meet specified AFL criteria. 15. The medium of claim 1, the instructions further causing the processor to divide the number of regions counted containing one or more plotted points by the number of points in the window, N, to calculate a ... | Medtronic, Inc. | A61B005/04
US07658717 | B2 | 2010-02-09 | Pulmonary evaluation device | 2004-08-10 | Pulmonary evaluation device has sensor means adapted to sense fluctuations in a user's lung operation and feedback means, driven by said sensor means, for determining successive values representative of the user's lung fluctuations and adapted to translate said values into appropriate lung-evaluating information. Optionally, the sensor means comprises or forms part of an item suitable to be worn by or carried adjacent the user so as to follow body movements caused by the user's lung operation to evaluate the user's lung operation. | 1. A pulmonary volume evaluation device comprising: an item worn over the user's body for following body movements caused by the user's lung operation;said item comprising:a front panel corresponding to the user's front;a rear panel corresponding to the user's back;an upper aperture sized and shaped to allow the user's head to be outside the item when worn;a lower aperture sized and shaped to allow the user's legs to be outside the item when worn;said front panel extending from said upper aperture to said lower aperture and being sized and shaped to substantially entirely cover the anterior chest wall and at least the upper abdomen;a sensor for sensing fluctuations in a user's lung operation; andfeedback means, driven by said sensor, for determining successive values representative of the user's lung fluctuations and for translating said values into appropriate lung-evaluating information;wherein said item has at least one chamber located in both said front and rear panels formed between an inner wall and an outer wall, said at least one chamber having a substantially enclosed volume of gas disposed therein, said at least one chamber being sized and shaped so as to substantially entirely cover the chest wall and at least the upper abdomen of the user's body, said inner wall being spaced from said outer wall throughout the entire lung region and being substantially flexible to remain in contact with the entire lung region in order to follow, in use, the displacement of the entire lung region; and said outer wall is substantially rigid in order to remain in position during the displacement of the lung region; whereby said inner wall and said outer wall combine to compress said volume of gas as said flexible inner wall is pushed towards said rigid outer wall during inspiration as the lungs expand and to decompress said volume of gas as said flexible inner wall relaxes during expiration as the lungs contract; and said sensor senses changes in pressure within said chamber throughout inspiration and expiration. 2. A device according to claim 1, comprising a seal for sealing said at least one chamber; whereby the volume of gas contained by said at least one chamber remains constant and as the body displaces during respiration, a measurable change in internal chamber pressure occurs as the chamber's wall displaces. 3. A device according to claim 1, incorporating an array of chambers locating a chamber over a separate region of the user's lung. 4. A device according to claim 1, wherein said at least one chamber comprises two chambers each of which correspond to a lung. 5. A device according to claim 1, wherein said at least one chamber comprises four chambers each of which correspond to one of an upper rib region and a lower rib region of a lung. 6. A device according to claim 1, wherein said feedback means comprises at least one of a microprocessor, a computer, and a data logger. 7. A device according to claim 1, wherein said item incorporates at l... | - | A61B005/08 A61B005/103 A61B005/117
US20100010367 | A1 | 2010-01-14 | System and methods for monitoring during anterior surgery | 2005-12-30 | The present invention involves a system and methods for nerve testing during anterior surgery, including but not limited to anterior disc replacement surgery, nucleus replacement, and interbody fusion. | 1. A system for conducting nerve testing during surgical procedures employing an anterior approach to the lumbar region, comprising: a surgical accessory capable of delivering an electrical stimulation signal to a nerve lying anterior to the spine;a sensor configured to detect neuromuscular responses evoked by the stimulation signal; anda control unit communicably linked to the stimulation accessory and the sensor. 2. The system of claim 1 and further, wherein the nerve testing is conducted during surgical procedures including at least one of total disc replacement, nucleus replacement, and interbody fusion. 3. The system of claim 1 and further, wherein at least one of the surgical accessory and sensor are adapted for use in at least one of a trans-peritoneal approach, retroperitoneal approach, and a minimally invasive laparoscopic approach. 4. The system of claim 1 and further, wherein the control unit is configured for at least one of (a) directing emission of the stimulation signal from the surgical accessory, (b) receiving and characterizing the neuromuscular response detected by the sensor, and (c) identifying a relationship between the stimulation signal and the neuromuscular response to complete the nerve test. 5. The system of claim 4 and further, wherein the sensor is configured to detect at least one of an EMG voltage output and pressure change and wherein the neuromuscular response is characterized by the magnitude of at least one of the voltage output and pressure change. 6. The system of claim 5 and further, wherein the magnitude of the EMG voltage output is characterized by a peak-to-peak amplitude. 7. The system of claim 5 and further, wherein the relationship identified is the threshold stimulation current necessary to evoke a threshold neuromuscular response, the threshold neuromuscular response being defined by a predetermined magnitude. 8. The system of claim 7 and further, wherein the nerve testing conducted includes at least one of nerve detection during anterior surgical access and pathology monitoring during nerve retraction. 9. The system of claim 8 and further, wherein the nerve is the hypogastric plexus. 10. The system of claim 9 and further, wherein the targeted muscle includes at least one of the bladder sphincter and the anal sphincter. 11. The system of claim 10 and further, wherein the sensor is coupled to a urinary catheter for deployment to the bladder sphincter. 12. The system of claim 11 and further, wherein the sensor contacts the bladder sphincter when the urinary catheter is inserted into the bladder. 13. The system of claim 12 and further, wherein the sensor is an EMG electrode having a generally annular shape for positioning around the exterior surface of the catheter. 14. The system of claim 13 and further, wherein the EMG electrode is fixed in position on the urinary catheter using at least one of a biocompatible adhesive, crimping, and an interference fit. 15. The system of claim 12 and fur... | - | A61B005/0488
US07647237 | B2 | 2010-01-12 | Communication station and software for interfacing with an infusion pump, analyte monitor, analyte meter, or the like | 2002-06-26 | A communication station is for use with a medical device (such as an infusion pump) and a processing device (such as a computer). The communication station includes a housing, a medical device interface coupled to the housing, a processing device interface coupled to the housing and a processor coupled to the housing. The device interface interfaces with the medical device, and the processing device interface interfaces with the processing device. The processor provides a communication path between the medical device and the processing device such that programming and instructions may be communicated from the processing device to the medical device and data may be transferred from the medical device to the processing device. The communication station may be combined with a system that is capable of generating reports either locally or remotely. In addition, the medical device interface may be a cradle that is configurable to attach to different shaped medical devices. | 1. A method of monitoring at least a continuous glucose monitor and an insulin infusion pump used to treat a medical condition for a patient, the method comprising: generating a patient data file;downloading data from at least the continuous glucose monitor and the insulin infusion pump used with the patient, wherein the data includes a plurality of blood glucose levels obtained over a time period;recording the downloaded data in the patient data file;integrating the data from at least the continuous glucose monitor and the insulin infusion pump to present contemporaneous information showing the relationship between at least the continuous glucose monitor and the insulin infusion pump;presenting the integrated data in a user chosen report format; anddisplaying a first plurality of the blood glucose levels from data obtained over a first defined time period and second plurality of the blood glucose levels from data obtained over a second defined time period on the same scale of a single one of the first and second defined time periods within the user chosen report format, with the first plurality of blood glucose levels from the first defined time period shown in an overlaid fashion relative to the second plurality of blood glucose levels from the second defined time period, and with the first and second plurality of blood glucose levels shown as extending over the single one of the defined time periods, the first and second time periods are fully separated in time. 2. The method of claim 1, wherein the user chosen report format temporally displays insulin administration compared to blood glucose levels. 3. The method of claim 1, wherein the user chosen report format displays percentages of blood glucose levels above and below glucose goals. 4. The method of claim 1, wherein the user chosen report format contains information for both bolus and basal insulin administered to the patient. 5. The method of claim 1, wherein the step of presenting the data further comprises calculating a percentage of total insulin delivery in the form of a holus injection over a period of time. 6. The method of claim 1, wherein the step of presenting the data further comprises calculating an average blood glucose level over a period of time. 7. The method of claim 6, wherein the step of presenting the data further comprises calculating a standard deviation. 8. The method of claim 1, wherein the step of presenting the data further comprises transforming the data into graphical waveform data, wherein the graphical waveform data represents the time at which the data occurs. 9. The method of claim 1, wherein each defined time period is the time period of a single day. 10. The method of claim 1, wherein each defined time period is the time period of a single week. 11. The method of claim 1, wherein displaying blood glucose levels from multiple defined time periods comprises displaying a continuous line graphical representation of a blood glucose level for each one... | MiniMed, Inc. | A61B005/00
US07641622 | B2 | 2010-01-05 | Guide wire | 2007-10-04 | A guide wire includes an elongated flexible wire body and a tubular member disposed so as to cover the outer periphery of a portion of the wire body on the distal end side. The tubular member includes a lamination portion having an inner layer and an outer layer made of a resin material, and a reinforcing member disposed between the inner layer and the outer layer and formed by winding a wire material spirally. The reinforcing member reinforces the lamination portion. | 1. A guide wire comprising: an elongated flexible wire body possessing an outer surface;a tubular member disposed in covering relation to the outer periphery of a portion of the wire body on a distal end side of the wire body;the tubular member comprising a distal portion and a proximal portion;the tubular member comprising a lamination portion and at least one wire member;the lamination portion comprising an inner layer made of resin material and an outer layer made of resin material;the at least one wire member being positioned between the inner layer and the outer layer;the tubular member being fixed to the wire body;the tubular member having an inner surface spaced from the outer surface of the wire body, with a gap between the inner surface of the tubular member and the outer surface of the wire body; andthe gap between the inner surface of the tubular member and the outer surface of the wire body extending from the distal portion of the tubular member to the proximal portion of the tubular member. 2. The guide wire according to claim 1, wherein the wire member is a reinforcing member that reinforces the lamination portion. 3. The guide wire according to claim 2, wherein the reinforcing member comprises a distal-most end portion and a proximal-most end portion, the distal-most end portion of the reinforcing member being positioned distally of a distal-most end of the lamination portion so that the distal-most end portion of the reinforcing member is exposed outside the lamination portion and/or the proximal-most end portion of the reinforcing member being positioned proximally of a proximal-most end of the lamination portion so that the proximal-most end portion of the reinforcing member is exposed outside the lamination portion. 4. The guide wire according to claim 1, wherein the wire body possesses a distal end portion and a proximal end, the distal end portion of the wire body being a tapering portion in which an outer diameter of the tapering portion gradually decreases in a distal end direction, and the tapering portion possessing a proximal end that is axially positioned relative to the tubular member so that the proximal end of the tapering portion is encircled by the tubular member. 5. The guide wire according to claim 1, wherein the tubular member possesses a distal end that is fixed to the wire body by way of a fixing member that connects the wire body and the tubular member. 6. The guide wire according to claim 1, wherein the at least one wire member extends distally beyond the lamination portion, the wire member possessing a distal end that is fixed to the wire body by way of a fixing member that connects the wire body and the tubular member. 7. The guide wire according to claim 1, wherein the wire member is a helical coil. 8. The guide wire according to claim 7, wherein the helical coil comprises a plurality of adjacent windings, at least some of the adjacent windings being spaced from one another. 9. The guide wire acco... | Terumo Kabushiki Kaisha | A61B005/00 A61M025/00
US08372435 | B2 | 2013-02-12 | Modular monolayer coatings for selective attachment of nanoparticles to biomolecules | 2009-06-22 | Nanoparticles are functionalized for use as bio-imaging probes using a novel, modular approach. Particle surface modification is based on a phosphonate monolayer platform on which was built a multi-segmented, multi-functional film: the first segment provided hydrolytic stability, the second aqueous suspendability, and the third, selectivity for cell attachment. In vitro imaging experiments visualized nanoparticle-cell surface binding. Peptide-derivatized nano-particles were not displaced from cells by soluble peptide. Methods for coating the host particles and use of rare earth ion-doped particles in imaging methods and photodynamic therapy methods are also disclosed. | 1. Nanoparticles having the structure:wherein R is a covalently attached biomolecule conjugate and M is a nanoparticle selected from the group consisting of lanthanum, yttrium, lead, zinc, cadmium, beryllium, magnesium, calcium, barium, aluminum, radium, silicon and germanium. 2. The nanoparticles of claim 1, the nanoparticle M further comprising ions of one or more rare earth elements. 3. The nanoparticles of claim 2, wherein the rare earth element ions are ions of one or more rare earth elements selected from the group consisting of samarium, gadolinium, lanthanum, yttrium, europium, cerium, terbium, dysprosium, holmium, erbium, thulium, ytterbium and lutetium. 4. The nanoparticles of claim 2, wherein the rare earth ions comprise a combination of ytterbium ions and one or more ions selected from the group consisting of erbium, holmium and thulium ions. 5. The nanoparticles of claim 1, wherein the nanoparticle M comprises an oxide of one or more elements selected from the group consisting of lanthanum, yttrium, lead, zinc, cadmium, magnesium, calcium, barium, aluminum, silicon and germanium. 6. The nanoparticles of claim 1, wherein the nanoparticle M comprises yttrium oxide (Y2O3). 7. The nanoparticles of claim 1 wherein the nanoparticle M comprises yttrium oxide doped with europium ions. 8. The nanoparticles of claim 6 wherein the nanoparticle M is doped with luminescent europium ions. 9. The nanoparticles of claim 1, wherein the nanoparticles comprise one or more biomolecular probes and the nanoparticle M comprises ions of one or more rare earth elements. 10. The nanoparticles of claim 9, wherein the biomolecular probes are selected from the group consisting of proteins and peptides that bind to biological markers. 11. A method for making a nanoparticle having the structure:wherein R is a covalently attached biomolecule conjugate and M is a nanoparticle selected from the group consisting of lanthanum, yttrium, lead, zinc, cadmium, beryllium, magnesium, calcium, barium, aluminum, radium, silicon and germanium, comprising attaching on the nanoparticle M an alkyl or aryl phosphonate monolayer having a terminus reactive for cross-coupling, reacting the nanoparticle with a cross-coupling agent appropriate for binding a conjugate to a surface of the nanoparticle, and attaching a biomolecule conjugate operable to bind a receptor of interest. 12. The method according to claim 11 comprising reacting the nanoparticle M with p-nitrophenyl chloroformate to provide a p-nitrophenyl carbonate derivatized nanoparticle that can react with amines. 13. The method according to claim 12 comprising binding tri(ethylene glycol) monoamine to the p-nitrophenyl carbonate derivatized nanoparticle. 14. The method according to claim 11, wherein the nanoparticle M is yttria oxide. 15. The method of claim 11, wherein the nanoparticle M is doped with ions of one or more rare earth elements. 16. The method according to claim 11 wherein the phosphonate monolayer... | Princeton University | A61K009/51 C01F017/00
US08105270 | B2 | 2012-01-31 | Measuring properties of an anatomical body | 2009-05-12 | A medical device includes a sensor that is configured to measure a property of an outer layer of an anatomical body surface. The sensor includes a source probe configured to stimulate a local surface of the outer layer of an anatomical body surface. The sensor also includes a detector configured to measure a response of the outer layer resulting from the source probe stimulation. A controller coupled to the source probe and the sensor drives the source probe using a tailored stochastic sequence and determines the property of the outer layer using the measured response received from the detector. The sensor can be used with medical devices, such as drug delivery devices including microneedle transport devices and needleless injection devices. | 1. A medical device comprising: a sensor that measures a property of an outer layer of an anatomical body surface, the sensor including:a source probe that stimulates a local surface of the outer layer of an anatomical body surface; anda detector that measures a response of the outer layer resulting from the source probe to stimulation; anda controller coupled to the sensor, wherein the controller drives the source probe using a tailored stochastic sequence and determines the property of the outer layer using the measured response received from the detector. 2. The device of claim 1 wherein the properties are determined with system identification techniques. 3. The device of claim 2 wherein the body surface is modeled as a second order mechanical system. 4. The device of claim 1 wherein the body surface is an internal body surface. 5. The device of claim 1 wherein the body surface is the skin of a subject. 6. The device of claim 1 further comprising a servo-controller coupled to a delivery device for delivering a pharmaceutical, the servo-controller adjusting the delivery characteristics of the delivery device based on the surface properties. 7. The device of claim 1 wherein the source probe comprises a voice coil. 8. The device of claim 1 wherein the detector comprises an accelerometer detecting displacement of the body surface. 9. The device of claim 1 wherein the detector comprises an linear differential variable transducer detecting displacement of the body surface. 10. The device of claim 9 wherein the detector further comprises a strain gauge measuring a static displacement of the body surface. 11. The device of claim 1 further comprising a drug injection device coupled to the sensor, the drug injection device injecting a drug into an anatomical body in response to the determined property of the outer layer. 12. The device of claim 11 wherein the drug injection device comprises a needleless injector. 13. The device of claim 1 further comprising a servo-controller coupled to a delivery device for delivering a formulation, the servo-controller adjusting the delivery characteristics of the delivery device based on the surface properties. | Massachusetts Institute of Technology | A61M031/00
US08618032 | B2 | 2013-12-31 | Polymer having visibility in magnetic resonance image and surface lubricity and medical device | 2007-09-21 | Provided are a polymer which, when wetted, enables a medical apparatus for use in MRI diagnosis and therapy to be easily visible under MRI and, simultaneously, develops surface lubricity, and a medical apparatus coated with the polymer. The polymer including a copolymer composed of a substance for shortening relaxation time of a nuclear species capable of being detected by magnetic resonance, a moiety having a reactive group, and a moiety for developing lubricity is used to coat the medical apparatus, whereby excellent peel resistance, high lubricity, visibility under magnetic resonance, easy applicability, and high safety can be obtained. | 1. A medical apparatus, comprising a substrate surface that is coated with a coating formed from a block copolymer comprising a moiety having a reactive group which is an epoxy group; and a moiety for developing lubricity comprising a hydrophilic monomer or a repetition of said monomer, wherein a substance for shortening relaxation time of a nuclear species capable of being detected by magnetic resonance is chemically bound to an epoxy group of the block copolymer,wherein the block copolymer is directly attached to the substrate surface of the medical apparatus,wherein the molar ratio of the moiety having a reactive group to the moiety for developing lubricity is 1:3 to 1:19,wherein said copolymer, when wetted, develops surface lubricity and MRI visibility,wherein either one of said moieties and said substance for shortening said relaxation time are covalently bound to each other either directly or through a linker molecule, andwherein said linker molecule is at least one selected from the group consisting of a diamine, amino-alcohol, dialkylenetriamine, aminocarboxylic acid, lactam, diol, hydroxycarboxylic acid, dicarboxylic acid, diisocyanate, diisothiocyanate, aminoisocyanate, aminoisothiocyanate, hydroxylisocyanate, hydroxylisothiocyanate, carboxylisocyanate, and carboxylisothiocyanate. 2. The block copolymer according to claim 1, wherein said substance for shortening the relaxation time of a nuclear species capable of being detected by magnetic resonance includes an ion or particle which is paramagnetic or superparamagnetic. 3. The block copolymer according to claim 1, wherein said hydrophilic monomer is at least one selected from the group consisting of: acrylamide and its derivatives; acrylic acid, methacrylic acid and their derivatives; maleic anhydride, maleic acid and their derivatives; and vinylpyrrolidone. 4. The block copolymer according to claim 1, wherein said moiety having a reactive group comprises a monomer having said reactive group in its molecule or a repetition of said monomer. 5. The block copolymer according to claim 1, wherein said copolymer is a block copolymer including a block comprising a monomer having a reactive group in its molecule, and a block comprising a hydrophilic monomer. 6. The medical apparatus according to claim 1, wherein said medical apparatus, when wetted, develops surface lubricity and MRI visibility. 7. A medical apparatus, comprising a substrate surface that is coated with a coating formed from a block copolymer comprising a moiety having a reactive group which is an epoxy group, and a moiety for developing lubricity comprising a hydrophilic monomer or a repetition of said monomer, wherein a substance for shortening relaxation time of a nuclear species capable of being detected by magnetic resonance is chemically bound to the moiety for developing lubricity,wherein the block copolymer is directly attached to the substrate surface of the medical apparatus,wherein the molar ratio of the moiety h... | Terumo Kabushiki Kaisha | C10M145/14 C04B035/04 A61B005/05
US09155476 | B2 | 2015-10-13 | Large surface area temperature sensing device | 2009-03-18 | A temperature probe for monitoring temperatures of a surface of a tissue or organ within the body of a subject includes a section with a substantially two-dimensional arrangement and a plurality of temperature sensors positioned across an area defined by the substantially two-dimensional arrangement. Such an apparatus may be used in conjunction with procedures in which thermal techniques are used to diagnose a disease state or treat diseased tissue. Specifically, a temperature probe may be used to monitor temperatures across an area of a surface of a tissue or organ located close to the treated tissue to prevent subjection of the monitored tissue or organ to potentially damaging temperatures. | 1. A temperature monitoring system, comprising: a temperature probe configured for introduction against an interior surface of an esophagus, the temperature probe including: an elongate member including: a distal portion for insertion into an interior of an esophagus of a subject and shaped to have a substantially two-dimensional arrangement having: a length that coincides with a remainder of a length of the elongate member;a width of about 10 mm to about 30 mm, the width exceeding a width of the elongate member; anda thickness of about a same thickness of the elongate member and any additional elements carried by the distal portion of the elongate member, the distal portion configured to be placed against and to conform to a shape of an interior surface of the esophagus without substantially deforming the shape of the interior surface and/or without substantially displacing tissue of the esophagus; anda proximal portion adjacent to an opposite end of the elongate member from the distal portion, the proximal portion configured for communication with apparatus external to a body of the subject; anda plurality of temperature sensors carried by the distal portion so as to be arranged across an area array and simultaneously positioned against the internal surface of the esophagus when the distal portion is positioned against the internal surface of the esophagus, with: at least two sensors spaced laterally (x-axis) apart from each other in the substantially two-dimensional arrangement by a lateral (x-axis) distance that exceeds a width of the elongate member; andat least two sensors spaced longitudinally (y-axis) apart from each other a longitudinal (y-axis) distance that is at least as great as the lateral (x-axis) distance; anda temperature display system in communication with the temperature probe, the temperature display system including: a display element for showing temperatures at locations that are visually arranged in a manner that corresponds to a physical arrangement of the plurality of temperature sensors carried by the distal portion of the elongate member of the temperature probe. 2. The temperature monitoring system of claim 1, wherein a combined width of the substantially two-dimensional arrangement of the distal portion of the elongate member of the temperature probe and the plurality of temperature sensors carried by the distal portion is configured to reside within the interior of the esophagus without substantially deforming a cross-sectional shape of the interior of the esophagus. 3. The temperature monitoring system of claim 1, wherein the distal portion of the elongate member of the temperature probe includes a single element. 4. The temperature monitoring system of claim 3, wherein at least a portion of the distal portion of the elongate member of the temperature probe is configured to have a substantially two-dimensional arrangement comprising a serpentine configuration. 5. The temperature monitoring system of claim 4, wh... | CIRCA Scientific, LLC | A61B005/00 A61B005/01 A61B019/00 A61B018/00
US09555252 | B2 | 2017-01-31 | Systems for providing neural markers for sensed autonomic nervous system activity | 2009-11-04 | In various method embodiments, a neural activity signal is sensed, a feature from the sensed neural activity signal is extracted, and a neural marker for the extracted feature is created. The neural marker includes information regarding the extracted feature. Various device embodiments comprise a port to receive a neural activity signal, and a feature extractor adapted to receive and process the neural activity signal to produce a neural marker that includes information for the neural activity signal. Various device embodiments comprise a display, a memory adapted to store a neural marker associated with a sensed neural activity signal, and a controller adapted to communicate with the memory and the display to provide a representation of the neural marker on the display. | 1. A device for providing neural markers for sensed autonomic nervous system (ANS) activity, comprising: a port to receive a neural activity signal representative of the ANS activity; anda feature extractor adapted to receive and process the neural activity signal to extract a feature from the neural activity signal and produce an ANS neural marker for the extracted feature, wherein the ANS neural marker includes information for the extracted feature, wherein the information for the extracted feature includes an event correlated with the ANS neural marker; anda memory connected to the feature extractor and configured to store, for each ANS neural marker, both an ANS neural marker label and an event identification for the event correlated with the ANS neural marker,wherein the feature extractor includes a filter for filtering the neural activity signal to extract the feature, and the filter is configured to provide a logarithmic average to the neural activity signal. 2. A device for providing neural markers for sensed autonomic nervous system (ANS) activity, comprising: a port to receive a neural activity signal representative of the ANS activity; anda feature extractor adapted to receive and process the neural activity signal to extract a feature from the neural activity signal and produce an ANS neural marker for the extracted feature, wherein the ANS neural marker includes information for the extracted feature, wherein the information for the extracted feature includes an event correlated with the ANS neural marker; anda memory connected to the feature extractor and configured to store, for each ANS neural marker, both an ANS neural marker label and an event identification for the event correlated with the ANS neural marker,wherein the feature extractor includes a filter for filtering the neural activity signal to extract the feature, and the filter is configured to average neural activity signals aligned to an index event. | Cardiac Pacemakers, Inc. | A61N001/37 A61B005/04 A61N001/36 A61N001/362 A61N001/372 A61B005/00
US20100049236 | A1 | 2010-02-25 | METHOD FOR TESTING AN AGENT FOR STROKES IN HUMANS WITH A NON-HUMAN STROKE MODEL | 2009-08-19 | An agent for treating strokes in humans is tested in a non-human subject by selecting an agent for testing and by preparing a selected non-human subject. The preparation includes: inducing a stroke event by advancing a microwire through the arterial system of the subject to a selected intracranial target position, inserting a microcatheter along the microwire and delivering an embolic device to the target position, occluding the artery at the target position by deploying the embolic device, verifying the occlusion and repositioning the embolic device if needed. After a predetermined occlusion interval, reperfusion of the subject is simulated by removing the embolic device and commencing therapy with the selected agent. At appropriate intervals, the effect of the conducted course of therapy is assessed non-invasively until terminal evaluation. In particular aspects, the method involves occluding the middle cerebral artery through an access achieved via the basilar artery. | 1. A method for testing, in a non-human mammal, an agent for therapeutic treatment of stroke in a human, comprising the steps of: selecting the therapeutic agent for treatment and determining a course of therapy using the therapeutic agent;selecting the non-human mammal;inducing a stroke event in the non-human animal through the substeps of: advancing a microwire through the arterial system of the non-human mammal to a selected intracranial target position;inserting a microcatheter along the microwire and delivering an embolic device to the target position;occluding the artery at the target position by deploying the embolic device, verifying the occlusion and repositioning the embolic device if needed; andwithdrawing the microcatheter from the target position to an intermediate arterial location, leaving the embolic device in place; andafter a pre-determined interval, simulating reperfusion by removing the embolic device;conducting the determined course of therapy using the selected therapeutic agent; andassessing, after an appropriate interval, the effect of the conducted course of therapy on the non-human mammal. 2. The method of claim 1, wherein: the substep of advancing the microwire comprises the steps of: advancing a microwire through the basilar artery of the non-human mammal;advancing the microwire from the basilar artery into and through a selected one of either the left or the right posterior communicating artery; andadvancing the microwire from the selected left or right posterior communicating artery into the corresponding left or right middle cerebral artery, which is the selected intercranial target position. 3. The method of claim 2, wherein: the non-human mammal is a canine. 4. The method of claim 3, wherein: in the stroke-inducing step, the substep of advancing the microwire through the basilar artery is preceded by the substeps of: introducing the microwire into the femoral artery of the non-human animal; andguiding the microwire to the basilar artery from the femoral artery through the aorta, the selected one of the left and right subclavian arteries, the corresponding vertebral artery, the spinal ramus artery and the anterior spinal artery, which becomes the basilar artery. 5. The method of claim 4, wherein: In the microcatheter withdrawing step, the microcatheter is withdrawn at least to the origin of the basilar artery. 6. The method of one of claim 5, wherein: the stroke-inducing step comprises, during the occlusion by the embolic device, the substeps of: performing at least one arteriogram to confirm occlusion of the ipsilateral middle cerebral artery and opacification of the contralateral hemisphere. 7. The method of claim 1, wherein: the embolic device is a platinum coil. 8. The method of claim 6, wherein: the M1 section of the selected middle cerebral artery remains occluded for about one hour. 9. The method of claim 2, wherein: in the stroke-inducing step, the substep of advancing the microwire through the basi... | The Ohio State University Research Foundation | A61M029/00 A61B005/00
US20100049024 | A1 | 2010-02-25 | COMPOSITE MATERIAL FOR IMPLANTABLE DEVICE | 2009-10-30 | Devices suitable for implantation in a body of a host and systems and methods for their manufacture are provided. The implantable devices include a composite material formed at least from a matrix material and hollow gas-filled beads. In preferred embodiments, the composite material includes a polymeric matrix mixed with hollow air-filled glass beads, which are mixed and cured to form at least a portion of the body of the implantable device. Implantable devices including this composite material have decreased weight and/or overall density as compared to implantable devices without the beads incorporated therein, which is believed to improve the acceptance and function of the implantable device in vivo. Additionally, implantable devices concerned with transmitting and receiving via RF are believed to achieve improved RF performance due to a reduced dielectric constant provided by the incorporation of beads within the composite material. | 1. A method for forming a composite material suitable for implantation in a host, the method comprising: providing a matrix material having a first dielectric constant;providing a plurality of hollow gas-filled beads having a second dielectric constant, wherein the second dielectric constant is lower than the first dielectric constant;mixing the hollow gas-filled beads and the matrix material to obtain a mixture; andcuring the mixture, whereby a composite material with a third dielectric constant is formed, wherein the third dielectric constant is lower than the first dielectric constant and higher than the second dielectric constant. 2. The method of claim 1, wherein the gas-filled beads are air-filled beads. 3. The method of claim 2, wherein the implantable device comprises electronics and an antenna for radiating or receiving a RF transmission, and wherein the step of curing comprises curing the mixture proximal to the antenna. 4. The method of claim 3, wherein the step of curing comprises at least partially encapsulating the electronics and the antenna with the composite material. 5. The method of claim 1, wherein the implantable device comprises a glucose sensor. | DexCom, Inc. | A61B005/145 H01P011/00
US20100030312 | A1 | 2010-02-04 | METHOD AND APPARATUS FOR LEAD LENGTH DETERMINATION | 2008-07-31 | A method and apparatus for positioning a lead within an anatomy is disclosed. The apparatus can be tracked relative to the anatomy to determine the position of selected portions of the anatomy. The positions of the portions of the anatomy can be used in determining an appropriate lead length to be positioned within the anatomy. Additionally, a determination of a lead length can be based upon statistical or acquired data of a selected population, according to an appropriate method. | 1. A lead system operable to be positioned in a selected volume, comprising: a distal tip of a lead body operable to be moved within the volume;a tracking device moveable relative to the distal tip and substantially along an axis of the lead body;a localizing system operable to track a position of the tracking device within the volume;a processor operable to determine a first location and a second location of the tracking device within the volume to determine a first dimension by tracking the tracking device;wherein the distal tip is operable to be fixed within the volume and the tracking device is operable to be tracked relative to the fixed distal tip. 2. The lead system of claim 1, further comprising: a display device operable to display image data of the volume;wherein the processor is operable to display a first icon representing the first location relative to the image data and a second icon representing the second location relative to the image data. 3. The lead system of claim 1, further comprising: a memory system including a relationship of a dimension between the first location and the second location and a length of the lead;wherein the processor is operable to access the memory system based upon the determined first location and the second location and output a determined length of the lead. 4. The lead system of claim 1, wherein the lead body includes a sheath extending a length from the distal tip;wherein the tracking device is operable to move a distance along the length. 5. The lead system of claim 1, wherein the tracking device is operable with the location system to determine a spatial location of a first volume location and a second volume location;wherein a dimension between the first volume location and the second volume location is operable to be determined;a measuring device operable to measure the length of the lead body with the volume related to the dimension. 6. The lead system of claim 5, wherein the lead body includes a sheath; wherein the measuring device is fixedly positioned relative to a proximal end of the sheath. 7. The lead system of claim 1 further comprising: a measuring device fixedly placed at a position on the lead body; anda positioning member moveable with the distal tip and moveable relative to the distal tip;wherein the positioning member is operable to be moved relative to the distal tip and measured with the measuring device. 8. The lead system of claim 7, further comprising: a case assembly operable to be positioned within the volume and connected to a proximal end of the lead body;wherein the case is operable to contain a pacing system operable to be programmed to deliver a signal to the distal tip. 9. The lead system of claim 1, further comprising: a display device operable to display a tracked position of the tracking device;wherein an icon representing a tracked position of the tracking device is operable to be superimposed on image data of the volume. 10. A method of positioning a lea... | - | A61N001/00 A61B005/05
US20100030052 | A1 | 2010-02-04 | ANALYTE SENSORS COMPRISING PLASTICIZERS | 2008-07-31 | A method and analyte-measurement device adapted to determine the concentration of an analyte in a fluid from a measurement site within a body, including an animal body, a mammalian body, or a human body. The invention further relates to the use of partially or fully implantable glucose-monitoring devices, including transcutaneous or subcutaneous glucose-monitoring devices. Devices, systems and kits making use of the aforementioned method are provided as well. | 1. An analyte sensor comprising: a working electrode comprising a sensing layer disposed proximate thereto;a counter electrode; anda plasticizer disposed proximate to the sensing layer. 2. The analyte sensor of claim 1, wherein at least a portion of the electrochemical analyte sensor is adapted to be subcutaneously positioned in a subject. 3. The analyte sensor of claim 1, further comprising a membrane disposed over the sensing layer. 4. The analyte sensor of claim 3, wherein the membrane comprises the plasticizer. 5. The analyte sensor of claim 1, wherein the sensing layer comprises the plasticizer. 6. The analyte sensor of claim 1, wherein the plasticizer comprises an alkyl citrate. 7. The analyte sensor of claim 6, wherein the alkyl citrate is triethyl citrate, acetyl triethyl citrate, tributyl citrate, acetyl tributyl citrate, trioctyl citrate, acetyl trioctyl citrate, trihexyl citrate, acetyl trihexyl citrate, butyryl trihexyl citrate, trihexyl o-butyryl citrate or trimethyl citrate. 8. The analyte sensor of claim 1, wherein the plasticizer comprises triethyl citrate. 9. The analyte sensor of claim 4, wherein the plasticizer comprises from about 0.005% to about 8% by weight of the total formulation of the membrane. 10. The analyte sensor of claim 4, wherein the plasticizer comprises from about 1% to about 40% by weight of the total formulation of the membrane. 11. The analyte sensor of claim 4, wherein the plasticizer comprises from about 15% to about 30% by weight of the total formulation of the membrane. 12. The analyte sensor of claim 5, wherein the plasticizer comprises from about 0.1% to about 40% by weight of the total formulation of the sensing layer. 13. The analyte sensor of claim 1, wherein the sensing layer comprises a glucose-responsive enzyme. 14. The analyte sensor of claim 1, wherein the sensing layer comprises a redox mediator. 15. The analyte sensor of claim 13, wherein the redox mediator comprises a ruthenium-containing complex or an osmium-containing complex. 16. The analyte sensor of claim 13, wherein the sensor is a glucose sensor. 17. The analyte sensor of claim 13, wherein the sensor is an in vivo sensor. 18. The analyte sensor of claim 13, wherein the sensor is an in vitro sensor. 19-41. (canceled) | - | A61B005/145 G01N027/26
US20100030095 | A1 | 2010-02-04 | Pulse Acoustic Analysis System for the Diagnostic of Cardiovascular Disease | 2008-07-30 | This invention provides a pulse acoustic analysis system for detecting disease. The present invention further provides a method for detecting a dynamic pressure change in a blood dynamic system in human. | 1. A pulse acoustic analysis system for detecting a cardiovascular disease or condition, comprising: (a) a miniature acoustic sensor;(b) a tool for fixing the miniature acoustic sensor;(c) a signal amplifier connected to the miniature acoustic sensor;(d) a digital-to-analog converter connected to signal amplifier; and(e) a device, having a data collecting and processing program, which is connected to the digital-to-analog converter. 2. The system of claim 1, wherein the disease is a cardiovascular disease. 3. The system of claim 1, wherein the detection is directed to a dynamic state of blood output of a visceral organ. 4. The system of claim 3, wherein the visceral organ is a heart or a blood-vessel-related system. 5. The system of claim 1, wherein the device, having a data collecting and processing program, is a computer. 6. The system of claim 1, wherein the data collecting and processing program performs Fourier transformation. 7. The system of claim 1, wherein the tool is a wrist band which fixing the miniature acoustic sensor on a radial artery. 8. The system of claim 1 which further comprises an oscilloscope connected to the signal amplifier. 9. A method for detecting a dynamic pressure change in a blood dynamic system in a subject, comprising the following steps: (a) fixing a pulse acoustic sensor, which consists of the miniature acoustic sensor and the tool of claim 1, on an artery to detect a pulse acoustic signal;(b) amplifying the pulse acoustic signal by a signal amplifier;(c) retrieving the signal by a digital-to-analog converter and storing the retrived signal in a computer;(d) conducting a waveform analysis of the retrieved signals by a data collecting and processing program to generate a result of amplitude and frequency; and(e) analyzing correlation between the result and the subject's health. 10. The method of claim 9, wherein the artery is radial artery, carotid artery or dorsalis pedis. 11. The method of claim 9, wherein the analysis comprises original waveform analysis, pulse acoustic waveform analysis and frequency spectrum analysis. 12. The method of claim 11, wherein the original waveform and the pulse acoustic waveform are displayed on a screen. 13. The method of claim 9, wherein the data collecting and processing program performs Fourier transformation. 14. The method of claim 11, wherein the original waveform analysis reveals a systole and a diastole of the heart and an on-off switch of cardiac valves. 15. The method of claim 11, wherein the frequency of the pulse acoustic waveform analysis reveals information of organs in the subject. 16. The method of claim 1, wherein the amplitude of the pulse acoustic waveform analysis reveals blood supply information of organs in the subject. 17. The method of claim 11, wherein the original waveform and the pulse acoustic waveform are output simultaneously, and compared, further with an electrocardiograph and an arterial blood pressure graph, to identify dynamic in... | Sustineo Biotechnology | A61B005/024
US20100010338 | A1 | 2010-01-14 | Implantable Medical Device Orientation Detection Utilizing an External Magnet and a 3D Accelerometer Sensor | 2008-09-30 | A method and device for detecting the implanted orientation of an implantable medical device (IMD) in a patient. IMD includes an accelerometer for measuring acceleration signals in three orthogonal directional axes. A y-axis orientation of IMD is determined from the measured accelerometer signals using a gravitational force analysis. IMD includes a magnetic sensor that senses a varying magnetic field exerted on the magnetic sensor from an external magnet moved along a medial-lateral direction with respect to IMD. The z-axis orientation of IMD is determined from the location of the external magnet where the magnetic field exerted on the magnetic sensor is greatest. Based on a known relationship between the accelerometer and magnetic sensor, an orthogonal transformation calculation is performed on the y-axis and z-axis orientations to yield the x-axis orientation. The implanted orientation of IMD with respect to the patient is thus known and used to compensate accelerometer measurements. | 1. A method of determining the orientation of an implantable medical device, comprising: determining a y-axis orientation of the device using accelerometer signals from a 3D accelerometer positioned within the device using gravity as an external force;determining an z-axis orientation of the device based on properties of a magnetic field between a magnetic sensor positioned within the device and an external magnet; anddetermining an x-axis orientation of the device from an orthogonal transformation calculation using the previously determined y-axis and z-axis orientations. 2. The method of claim 1, further comprising determining the y-axis orientation when a patient in which the implantable medical device is implanted is in an upright posture. 3. The method of claim 1, further comprising: monitoring the accelerometer signals from the 3D accelerometer to detect certain low frequency periodic signals signifying that a patient in which the implantable medical device is implanted is involved in an upright activity. 4. The method of claim 1, further comprising determining the z-axis orientation by: moving the external magnet and magnetic sensor with respect to each other between so that the external magnet moves between lateral and medial positions of a patient in which the implantable medical device is implanted;measuring a strength of a magnetic field imparted on the magnetic sensor by the moving external magnet;identifying the location at which the magnetic field between the external magnet and magnetic sensor has a maximum value; anddetermining the z-axis orientation of the device based on the location at which the magnetic field has a maximum value. 5. The method of claim 1, further comprising determining the z-axis orientation by: positioning the external magnet such that the external magnet includes a stationary, known position with respect to the magnetic sensor;measuring a strength of a magnetic field imparted on the magnetic sensor by the external magnet;determining the z-axis orientation of the device based on the strength of the magnetic field imparted on the magnetic sensor and the known position of the external magnet. 6. The method of claim 1, further comprising calibrating the 3D accelerometer using the determined y-axis, z-axis and x-axis orientations. 7. The method of claim 1, further comprising positioning the external magnet in an in-home patient monitoring system so that a patient can compensate the orientation of the device by moving with respect to the in-home patient monitoring system to determine compensate the x-axis orientation. 8. An implantable medical device comprising: a 3D accelerometer sensor configured for measuring acceleration signals in three orthogonal directional axes comprising an x-axis, z-axis and y-axis;a magnetic sensor configured to sense a magnetic field exerted on the magnetic sensor from a moving magnetic field source; anda controller coupled to the 3D accelerometer for receiving acceleration sign... | - | A61B005/05 A61N001/00
US20100016759 | A1 | 2010-01-21 | Device and method for measuring and treating the rigidity and erection of a penis and arterial-venous flows | 2009-05-05 | A measuring method enables the strength of the erector muscles (EM) causing the rigidity of the penis (P) of a patient and the arterial and venous variation ensuring the erection of the penis (P) of the patient to be measured and reinforced by placing a sensor (2, 50) around the penis (P) includes elements to receive the intracavernous pressure (ICP) and pressure variations of the penis (P), and at least one transducer (21, 55) integrated into or offset from the sensor (2, 50) allowing the pressure signal from the elements for receiving the pressure to be transformed into an electrical signal, and by connecting the sensor (2, 50) to a housing (35) including elements for amplifying and recording the signal from the transducer (21, 55), elements for analyzing the signal by a microcontroller allowing the fatigue state of the erector muscles (EM) to be measured and the variations in muscle strength of the erector muscles (EM) to be measured. | 1. A method for measuring the strength of the erector muscles (EM) causing the rigidity of the penis (P) of a patient and for measuring the variations in arterial and venous flows ensuring the erection of the penis (P) of the patient, which comprises: Placing a sensor (2, 50) around the penis (P) comprising means to receive the intracavernous pressure (ICP) and pressure variations of the penis (P), and at least one transducer (21, 55) integrated into or offset from the sensor (2, 50) allowing the pressure signal issued from the means for receiving the pressure to be transformed into an electrical signal,And connecting the sensor (2, 50) to a housing (35) comprising means for amplifying and recording the signal from the transducer (21, 55) and means for analyzing the signal by a microcontroller allowing the state of erector muscle (EM) fatigue to be measured and the variations in muscle strength of the erector muscles (EM) to be measured. 2. The measuring method according to claim 1, further comprising measuring the variations in intracavernous pressure (ICP) of the penis (P) in order to quantify the detumescence, tumescence and rigidity of the penis (P) during both day and night. 3. The measuring method according to claim 2, further comprising measuring the intracavernous pressure (ICP) when the pressure progressively reaches the mean systolic pressure of the penis (P). 4. The measuring method according to claim 1, further comprising measuring the increase in arterial inflow in the corpus cavernosum (CC) and the reduction in venous outflow during the vascular phase, which is the first penis (P) erection phase. 5. The measuring method according to claim 1, further comprising measuring variations in the contraction and rigidity of the erector muscles (EM) surrounding the corpus cavernosum (CC) during the muscular phase, which is the second penis (P) erection phase. 6. The measuring method according to claim 4, further comprising continuously measuring, particularly when the patient is sleeping, the arterial inflow and the venous outflow by measuring the pulse pressure (PP), pulse volume (PV) and intracavernous pressure (ICP). 7. The measuring method according to claim 1, further comprising measuring the Young's modulus of the corpus cavernosum (CC) allowing the compliance of said corpus cavernosum to be evaluated. 8. The measuring method according to claim 2, further comprising measuring and quantifying the cavernosal-venous leak from the intracavernous pressure (ICP) curve. 9. The measuring method according to claim 1, measuring the maximum contraction (Pmax), the pressure gradient (Delta P and mean delta), the surface under the curve resulting in the muscle work of each contraction, the number of peaks and the width of each peak. 10. The measuring method according to claim 1, further comprising measuring the fatigue of the erector muscles (EM) by measuring the angle (α). 11. The measuring method according to claim 1, further comprising ... | - | A61B005/103
US20100004566 | A1 | 2010-01-07 | INTELLIGENT ORTHOTIC INSOLES | 2009-01-08 | An intelligent insole for generating time sensitive information about the pressure on the foot. The insole includes a custom-made, semi-custom or generically sized orthotic component. The orthotic is laminated with a top cover and an intermediate pressure sensor having an array of capacitive pressure sensors. Signal processing equipment may be embedded in the insole or placed locally with the insole as on the side of a shoe. The processor also can connect to a wireless transmitter for relaying the information to a remote site. | 1. An intelligent orthotic insole for insertion into an individual's shoe for broadcasting output signals that represent the pressure exerted by the individual's foot on said insole, said insole comprising: A) an integral structure including: i) a lower portion for engaging the shoe surface and defining a orthotic for the individual, andii) sensor pad means affixed to said lower portion for generating an array of signals containing information about the magnitude and position of pressure thereon,B) a processor responsive to signals from said sensor pad means for generating corresponding output signals, andC) means attached to said processor for transmitting the output signals to a remote location. 2. The intelligent orthotic insole as recited in claim 1 wherein said sensor pad means comprises a capacitive force sensing matrix that produces a three-dimensional pressure map of the forces acting on the insole. 3. The intelligent orthotic insole as recited in claim 2 wherein said capacitive force sensing matrix includes a compressible dielectric material. 4. The intelligent orthotic insole as recited in claim 2 wherein said sensor pad means is formed as a flexible tactile sensor. 5. The intelligent orthotic insole as recited in claim 2 wherein said sensor pad means comprises an embedded capacitive pressure sensor array. 6. The intelligent orthotic insole as recited in claim 2 wherein the sensor pad means comprises embedded discrete capacitive pressure sensors. 7. The intelligent orthotic insole as recited in claim 2 wherein said force sensing matrix comprises up to 200 pressure sensing locations. 8. The intelligent orthotic insole as recited in claim 2 wherein said force sensing matrix comprises between 3 and 20 pressure sensing locations. 9. The intelligent orthotic insole as recited in claim 2 wherein said processor converts the signals from the sensor pad means into additional time-stamped data packets. 10. The intelligent orthotic insole as recited in claim 2 wherein said processor is embedded in said integral structure. 11. The intelligent orthotic insole as recited in claim 10 wherein said transmitting means is embedded in said integral structure. 12. The intelligent orthotic insole as recited n claim 2 additionally comprising an integral cover overlying was lower portion and said sensor pad means. 13. An intelligent orthotic insole for providing time sensitive information about pressure on the foot comprising: A) an integral insole structure including: i) a lower portion for engaging the shoe surface and defining a orthotic for the individual,ii) a cover substantially coextensive with the lower portion,iii) sensor pad means attached to said lower portion and integral therewith for generating an array of signals containing information about the magnitude and position of pressure thereon,B) a processor responsive to signals from said sensor pad means for generating corresponding output signals,C) means attached to said processor for ... | ESOLES, L,L.C. | A61B005/103 A43B013/38 G08B001/08
US20100036266 | A1 | 2010-02-11 | DEVICE AND METHOD FOR DETECTING HEART BEATS USING AIRWAY PRESSURE | 2009-07-31 | Embodiments are directed to a device and method for detecting heart beats and/or monitoring ventilation and/or respiration of a patient using air pressure. In one embodiment, a monitoring device comprises a gas duct and at least one pressure sensor configured to measure pressure in the duct. The duct may be in fluid communication with a patients airways. Based on the sensed pressure, a heart beat of the patient may be detected. In another embodiment, a flow rate of the air expired from the patient may be calculated based on the pressure and a known flow resistance. | 1. A device for obtaining heart beat information of a patient, comprising: a duct configured to allow a gas to pass therethrough;at least one pressure sensor connected to the duct and configured to measure the pressure in the duct; anda processing unit connected to the pressure sensor and operable to generate heart beat information from the pressure signal. 2. The device according to claim 1, wherein the duct is connected to a respiration device. 3. The device according to claim 1, wherein the duct is integrated in a respiration device. 4. The device according to claim 1, further comprising a restriction arranged in the duct, and wherein the pressure sensor is configured to measure a pressure drop across the restriction as a pressure difference between a location in the duct at a first side of the restriction and a known, substantially constant pressure, arranged at a second side of the restriction. 5. The device according to claim 1 further comprising a one-way valve with known flow resistance connected to the duct. 6. The device according to claim 2 wherein the respiration device comprises a respiratory mask or a tracheal tube. 7. The device according to claim 3, wherein the respiration device comprises a bellow. 8. The device according to claim 4, wherein the pressure drop is measured by a single pressure sensor. 9. The device according claim 1, further comprising a restriction arranged in the duct and wherein the pressure sensor is arranged to measure pressure at a location at a first side of the restriction. 10. The device according to claim 1, further comprising a restriction having a first and second side arranged in the duct and wherein the restriction is shaped to ensure that a flow velocity is substantially equal on the first and second sides of the restriction. 11. The device according to claim 1, wherein the pressure sensor is an absolute pressure sensor. 12. The device according to claim 1, wherein the pressure sensor is a differential pressure sensor. 13. The device according to claim 4, wherein the known, substantially constant pressure, is ambient pressure. 14. The device according to claim 1, wherein the processing unit is adapted to calculate flow values from a linear relationship between pressure drop and flow. 15. The device according to claim 1, wherein the processing unit is adapted to calculate respired volume. 16. A method for obtaining heart beat information of a patient, comprising: measuring pressure in a gas duct connected to a patient's airways; andderiving patient heart beat information based on the measured pressure. 17. The method according to claim 16, further comprising providing an elevated pressure in the lungs of the patient. 18. The method according to claim 16, further comprising connecting the duct to a respiration device. 19. The method according to claim 16, wherein the step of measuring pressure in a gas duct comprises measuring a pressure drop across a restriction as the difference bet... | Laerdal Medical AS | A61B005/02
US20100049073 | A1 | 2010-02-25 | Method, System, and Computer Program for Determining if a Subject is Snoring | 2006-06-28 | In one embodiment of the present invention, a method for determining if a subject is snoring is disclosed, including the steps of: determining if the subject is present in a bed; determining if there is snoring; determining if snoring is emitted from the subject; and generating an output signal, wherein the output signal is dependent on if the snoring is emitted from the subject. A computer program arranged to cause a computer to perform the steps of the method is also disclosed. Further, a system for snoring determination is disclosed, comprising a processor, an audio input means providing an audio signal to the processor, and a sensor pad arranged in a bed, the sensor pad providing a sensor pad signal to the processor, wherein the processor comprises devices for determination of presence of a subject from the sensor pad signal; devices for determination of snoring from the audio signal; devices for determination if the subject is emitting the snoring from the sensor pad signal and the audio signal; and an output arranged to provide an output signal according to determinations of the devices for determination of presence, devices for determination of snoring, and devices for determination is the subject is emitting the snoring. | 1. A method for determining if a subject is snoring, comprising the steps of: determining if said subject is present in a bed;determining if there is snoring;determining if snoring is emitted from said subject, comprising the step of processing an audio signal from an audio input means arranged in a vicinity of said bed and a sensor pad signal from a sensor pad of said bed, wherein said processing comprises the step of determining if said audio signal and said sensor pad signal are synchrone;determining if said subject is asleep, comprising measuring physiological parameters of said subject, and processing the measured physiological parameters; andgenerating an output signal, wherein said output signal is dependent on if said snoring is emitted from said subject and if said subject is snoring. 2. The method according to claim 1, wherein said physiological parameters are associated with breathing, heartbeat, or movements, or any combination thereof. 3. The method according to claim 1, wherein said step of determining if a subject is present in said bed comprises the step of detecting a signal from a sensor pad of said bed. 4. The method according to claim 1, wherein said step of determining if there is snoring comprises the steps of: acquiring an audio signal in a vicinity of said bed;discriminating a physiological sound from said audio signal; anddetermining if said physiological sound comprises snoring sounds. 5. The method according to claim 4, wherein said step of determining if said physiological sound comprises snoring sounds comprises the steps of measuring physiological parameters of said subject; andprocessing measured physiological parameters. 6. The method according to claim 5, wherein said physiological parameters are associated with breathing, heartbeat, or movements, or any combination thereof. 7. The method according to claim 1, wherein said processing comprises the step of calculating a correlation between said audio signal and said sensor pad signal. 8. A system for snoring determination comprising a processor, an audio input means providing an audio signal to said processor, and a sensor pad arranged in a bed, said sensor pad providing a sensor pad signal to said processor, wherein said processor comprisesmeans for determination of presence of a subject from said sensor pad signal;means for determination of snoring from said audio signal;means for determination if said subject is emitting said snoring from said sensor pad signal and said audio signal, wherein said means for determination if said subject is emitting said snoring comprises a correlator; andan output arranged to provide an output signal according to determinations of said means for determination of presence, means for determination of snoring, and means for determination if said subject is emitting said snoring. 9. The system according to claim 8, wherein said processor further comprises means for determination if said subject is asleep from said sensor pad ... | - | A61B005/08
US07650180 | B2 | 2010-01-19 | Imaging sensor array and device and method for use therefor | 2004-07-04 | A system and method for receiving an in vivo signal, using for example a receiver, a recorder and an antenna array. The receiver may include for example a switching unit and an amplifier, and may be in proximity to the antenna array. | 1. A method for adjusting operation of an in vivo sensing system, the method comprising: detecting the presence or absence of a receiver connected by cable to a recorder, the receiver receiving image signals from a plurality of antennas placed on a body;upon detecting the presence of the receiver: identifying the type of the receiver;automatically adjusting operation of the recorder according to the type of receiver identified;transferring data received by the receiver to the recorder; andstoring in the recorder data transferred from the receiver; andupon detecting the absence of the receiver, automatically adjusting the operation of the recorder, wherein automatically adjusting the operation of the recorder comprises an operation selected from the group consisting of: not recording data, recording data indicating a receiver is not connected, and stopping to record data. 2. The method according to claim 1, wherein the plurality of antennas comprises a radio frequency antenna. 3. The method according to claim 1, wherein the receiver receives signals selected from a group consisting of: radio frequency signals, control data, and energy. 4. The method according to claim 1, wherein the receiver is to adjust its operation according to the number of the plurality of antennas. 5. The method according to claim 1, wherein the signals are pre-amplified prior to being transferred to the recorder. 6. The method according to claim 1, wherein the strongest signal from the plurality of antennas is selected for recording. 7. The method according to claim 1 wherein the plurality of antennas receive image signals from an in-vivo imaging capsule. | Given Imaging Ltd. | A61B005/05
US07650004 | B2 | 2010-01-19 | Hearing aids and methods and apparatus for audio fitting thereof | 2002-01-16 | A field ready, unsupervised-use ready, method and apparatus for audio fitting a hearing aid is described in a hand held configuration having paired comparisons (hearing selections) stored in and derivable from a memory therein. The paired comparisons are presented one at a time to a user and a preferred selection for each paired comparison is made by a select indicator after the user toggles back and forth between the selections for as many times necessary in determining their preferences. A genetic algorithm converges all the preferences upon a single solution. Crossover and mutation genetic algorithm operators operate on a linear range of indexes representative of parametric values of the pairs. A fully integrated hearing aid having all the above described features incorporated therein is also presented. | 1. An apparatus for fitting a hearing aid by its wearer, comprising: a memory adapted to store a first population comprising a plurality of parent sets;a toggle device adapted to toggle between a pair of the plurality of parent sets;a select indicator for selecting a preferred one set of the pair;a communications link adapted to connect with the hearing aid; anda processor adapted to provide signals to the hearing aid to change operation of the hearing aid based on each parent set, to allow the wearer to select one preferred set for each pair of parent sets, to record a ranking of parent sets, assign probabilities of selection of parent sets, crossover and/or mutate at least one parent set, and replace weakest parent sets with a child set. 2. The apparatus according to claim 1, wherein the processor ranks a hierarchy of the plurality of parent sets. probability of selection by the select indicator to the sets of the plurality of pairs and the another pair. 3. The apparatus according to claim 2, wherein the processor assigns a probability of selection by the select indicator to the sets of the plurality of pairs and the another pair. 4. The apparatus according to claim 3, wherein the genetic algorithm decides which of the plurality of parent sets becomes the first and second sets for each the plurality of pairs. 5. A hearing aid fitted by the apparatus according to claim 1. 6. A method of fitting a hearing aid, comprising the steps: preparing a first population of a plurality of parent sets;presenting a first pair from the parent sets, the first pair comprising a first and second set and being presented with assistance of the hearing aid;selecting a first preference between the first and second sets of the first pair;operating on at least one set of the plurality of parent sets to obtain a child set, the child set being one of a mutation and crossover;replacing one of the plurality of parent sets of the first population with the child set to form a second population;presenting a second pair, the second pair comprising the child set and a third set, the third set being selected from the second population but not being the child set;selecting a second preference between the child set and the third set of the second pair; andconverging on a solution set, the solution set being one of the first, second, third and child sets. 7. The method according to claim 6, further comprising the step of ranking a hierarchy of the plurality of parent sets. 8. The method according to claim 6, further comprising the step of assigning a probability of selection in one of the selecting steps to the first, second, third and child sets. 9. The method according to claim 6, further comprising the step of deciding which set of the plurality of parent sets becomes the first and second sets of the first pair. 10. A method of fitting a hearing aid, comprising the steps: providing a hearing aid having an initial prescription;preparing a population of a plurality of par... | Starkey Laboratories, Inc. | H04R025/00 H04R029/00 A61B005/00 G06N003/12
US07648458 | B2 | 2010-01-19 | Insertion shape detecting probe | 2006-10-10 | An insertion shape detecting probe includes shape detecting elements for generating or detecting magnetic field or magnetic; signal lines connected to the shape detecting elements; supporting members that support the shape detecting elements; inner sheaths into which the signal lines and the supporting members are inserted; and an outer sheath into which the inner sheaths, the shape detecting elements, and at least one portion of the supporting members are inserted. The inner sheaths are each formed from a multiple-lumen tube that has a first lumen formed at a substantially central portion of each inner sheath and second lumens formed around an exterior periphery portion of the first lumen, and the supporting members are elongated wire rods one portion of which is arranged along exterior peripheries of the shape detecting elements in an axial direction and other portion runs through the second lumens. | 1. An insertion shape detecting probe, comprising: plural shape detecting elements that is used to detect a position of an insertion portion for generating magnetic field or detecting magnetic field;plural signal lines that are connected to the plural shape detecting elements;plural supporting members that support the plural shape detecting elements;plural inner sheaths into which the signal lines and the supporting members are inserted, the inner sheaths being each formed by a multiple-lumen tube that has a first lumen formed at a substantially central portion of each inner sheath and plural second lumens formed around an exterior periphery portion of the first lumen, the shape detecting elements and the inner sheaths being linearly arranged, the supporting members being elongated wire rods in which one portion thereof is arranged along exterior peripheries of the shape detecting elements in an axial direction and other portion thereof runs through the second lumens of the inner sheaths; andan outer sheath into which the inner sheaths, the shape detecting elements, and at least one portion of the supporting members are inserted. 2. The insertion shape detecting probe according to claim 1, wherein the supporting members is configured to support the shape detecting elements and maintain linearity so that each supporting member touches the shape detecting elements at the exterior peripheries of the shape detecting elements. 3. The insertion shape detecting probe according to claim 1, wherein the wire rods of the supporting members are formed from metal members surfaces of which are resin coated and which have heat conductivity. 4. The insertion shape detecting probe according to claim 1, wherein the supporting members are formed from shape memory metals with linear shape memory. 5. The insertion shape detecting probe according to claim 1, wherein the supporting members include wire rods that are made from metal members having heat conductivity, and thin wall tubes that cover exterior surfaces of the wire rods. 6. The insertion shape detecting probe according to claim 1, wherein the supporting members are formed from thin tubes that are made from resin members or elastic members. 7. The insertion shape detecting probe according to claim 1, wherein the supporting members are thin tubes inside which cooling mediums are sealed. 8. The insertion shape detecting probe according to claim 1, wherein the supporting members are formed from thin hollow duct members that are connected in loop shapes at a probe distal end portion so that a cooling medium runs through inside the supporting members. 9. The insertion shape detecting probe according to claim 1, further comprising: a protection member that is provided between each shape detecting element and inner sheaths that are adjacent to the shape detecting element to connect each shape detecting element to the inner sheaths that are adjacent to the shape detecting element as well as to protect an exter... | Olympus Corporation | A61B001/00 A61B005/05
US20100004560 | A1 | 2010-01-07 | CONFUSER CROWN SKIN PRICKER | 2009-09-11 | A blood sampling device designed to divert attention from the pain involved in the puncture of the skin tissue by increasing comfort when lancing the skin by affecting the sensation and perception of pain. Projections on the distal end of a skin pricker contact the skin surface to confuse the nerves in the area of the prick to make the prick less noticeable. | 1-18. (canceled) 19. A blood sampling device, comprising: (a) a body portion having a bore and adapted to house a lancet,(b) means for forcing the lancet through the bore in the body;(c) a platform having an aperture, the aperture adapted to receive at least a sharp tip of the lancet and comprising a generally flat surface having an inner boundary and an outer boundary; and(d) projections extending from the flat surface between the inner boundary and the outer boundary and surrounding the aperture to contact skin tissue. 20. The blood sampling device of claim 19, wherein the projections further comprise a pyramidal shape. 21. The blood sampling device of claim 19, wherein the projections encircle the aperture. 22. The blood sampling device of claim 19, wherein the projections further comprise a rounded tip. 23. The blood sampling device of claim 19, further comprising means for retracting the lancet into the body after the lancet is fired. 24. The blood sampling device of claim 19, further comprising a button for firing the lancet. 25. The blood sampling device of claim 19, further comprising means for adjusting the penetration depth of the lancet. 26. The blood sampling device of claim 19, wherein the means for forcing the lancet through the bore in the body comprises a spring. 27. A lancing device for obtaining a blood sample, comprising: (a) a body portion having a bore and adapted to house a lancet; and(b) a nose portion having a frustoconical shape and comprising (1) an aperture adapted to receive at least a sharp tip of the lancet,(2) a platform surrounding the aperture, the platform comprising a generally flat surface having an inner boundary and an outer boundary, and(3) a plurality of projections extending from the flat surface between the inner boundary and the outer boundary and near the aperture to contact skin tissue. 28. The lancing device of claim 27, wherein the projections further comprise a pyramidal shape. 29. The lancing device of claim 27, wherein the projections encircle the aperture. 30. The lancing device of claim 27, wherein the projections further comprise a rounded tip. 31. The lancing device of claim 27, further comprising means for forcing the lancet through the bore in the body. 32. The lancing device of claim 27 further comprising means for retracting the lancet into the body after the lancet is fired. 33. A lancing device, comprising: (a) a body portion having a bore and adapted to house a lancet,(b) means for forcing the lancet through the bore in the body;(c) an aperture adapted to receive at least a sharp tip of the lancet;(d) a platform surrounding the aperture and comprising a generally flat surface having an inner boundary and an outer boundary;(e) projections extending from the platform between the inner and outer boundaries of the platform; and(f) means for retracting the lancet into the body after firing the device. 34. The lancing device of claim 33, wherein the projections further compris... | OWEN MUMFORD, LTD. | A61B005/151 A61B017/34
US20100004516 | A1 | 2010-01-07 | Method and System for Monitoring Physiological Status | 2009-01-23 | A method for monitoring physiological status is applied to a laptop to form a monitoring system. A sensor is disposed on the laptop adjacent to a keyboard of the laptop. The sensor is provided for sensing user's physiological signal and a detection circuit is utilized for converting the physiological signal into a numerical value. Detection software is installed to the laptop and executed for comparing the numerical value with a pre-determined threshold value range and output the numerical value and a comparison result. An alert is issued when the numerical value falls outside the threshold value range. | 1. A method for monitoring physiological status, applied in a laptop, wherein a sensor is disposed on the laptop adjacent to a keyboard of the laptop, the method comprising the following steps: starting and resetting the sensor;giving a threshold value range;detecting continuously by the sensor, to determine whether a physiological signal is input;obtaining the physiological signal by the sensor and converting the physiological signal into a numerical value;comparing the numerical value with the threshold value range to obtain a comparison result; andoutputting the comparison result. 2. The method for monitoring physiological status as claimed in claim 1, further comprising the following step after the step of detecting whether a physiological signal is input: storing the numerical value after obtaining the numerical value. 3. The method for monitoring physiological status as claimed in claim 1, wherein the threshold value range lies in between an upper limit value and a lower limit value. 4. The method for monitoring physiological status as claimed in claim 3, further comprising the following step before the step of outputting the comparison result: determining the comparison result, and issuing an alert if the numerical values of the physiological signal falls outside the threshold value range. 5. The method for monitoring physiological status as claimed in claim 4, wherein the alert is caution sound produced by the laptop. 6. The method for monitoring physiological status as claimed in claim 4, wherein the alert is a caution image displayed by the laptop. 7. The method for monitoring physiological status as claimed in claim 4, wherein the alert is sending the comparison result and a caution message to a remote terminal through a network. 8. The method for monitoring physiological status as claimed in claim 1, further comprising the following step after the step of obtaining a numerical value of the physiological signal: outputting the numerical value. 9. The method for monitoring physiological status as claimed in claim 8, wherein the step of outputting the numerical value is producing speech using the laptop. 10. The method for monitoring physiological status as claimed in claim 8, wherein the step of outputting the numerical value is displaying an image display by the laptop. 11. The method for monitoring physiological status as claimed in claim 8, wherein the step of outputting the numerical value is sending the numerical value to a remote terminal through a network. 12. A system for monitoring physiological status, for monitoring a physiological signal of a user, the system comprising: a laptop executing detection software, for comparing a numerical value of the physiological signal with a threshold value range, outputting the numerical value and a comparison result, and issuing an alert if the numerical value falls outside the threshold value range;a sensor disposed on the laptop, for receiving the physiological signal when the... | Micro-Star Int'l Co., Ltd. | A61B005/00
US07657301 | B2 | 2010-02-02 | Guided radiation therapy system | 2003-11-25 | A system and method for accurately locating and tracking the position of a target, such as a tumor or the like, within a body. In one embodiment, the system is a target locating and monitoring system usable with a radiation delivery source that delivers selected doses of radiation to a target in a body. The system includes one or more excitable markers positionable in or near the target, an external excitation source that remotely excites the markers to produce an identifiable signal, and a plurality of sensors spaced apart in a known geometry relative to each other. A computer is coupled to the sensors and configured to use the marker measurements to identify a target isocenter within the target. The computer compares the position of the target isocenter with the location of the machine isocenter. The computer also controls movement of the patient and a patient support device so the target isocenter is coincident with the machine isocenter before and during radiation therapy. | 1. In radiation therapy using a radiation delivery system having a radiation source that directs a radiation beam to a beam isocenter, an apparatus for locating and tracking a treatment target in a patient comprising: a wireless marker configured to be implanted in the patient at a site relative to the treatment target wherein the marker is energized or excited by an excitation source positioned exterior of the patient;a plurality of sensors in a fixed and known geometry relative to each other that obtain position information about the location and/or orientation of the marker; anda computer operatively coupled to the sensors, the computer having a computer operable medium containing instructions that cause the computer to (a) receive the position information data from the sensors 12 or more times per minute, (b) determine an actual location of the treatment target, and (c) compute a displacement between the treatment target and the beam isocenter while the sensors obtain the position information. 2. The apparatus of claim 1 wherein the instructions contained by the computer operable medium further cause the computer to compute the displacement between the treatment target and the beam isocenter in three dimensions based on the actual location of the treatment target. 3. The apparatus of claim 1 wherein the instructions contained by the computer operable medium further cause the computer to: compute the displacement between the treatment target and the beam isocenter in three dimensions based on the actual location of the treatment target; andprovide the computed displacement between the treatment target and the beam isocenter to at least one of (a) an operator, (b) a control system and (c) a radiation delivery device at least every five seconds while irradiating the patient. 4. The apparatus of claim 1 wherein the instructions contained by the computer operable medium further cause the computer to compute the displacement between the treatment target and the beam isocenter in three dimensions based on the actual location of the treatment target so that the patient can be moved according to the computed displacement while the sensors obtain the position information. 5. The apparatus of claim 1 wherein the instructions contained by the computer operable medium further cause the computer to: compute the displacement between the treatment target and the beam isocenter in three dimensions based on the actual location of the treatment target; andprovide the computed displacement between the treatment target and the beam isocenter to at least one of (a) an operator, (b) a control system and (c) a radiation delivery device at least every five seconds so that the patient can be moved while the sensors obtain the position information to maintain the treatment target in an acceptable range from the beam isocenter. 6. In radiation therapy using a radiation delivery system having a radiation source that directs a radiation beam to a beam isocenter, an a... | Calypso Medical Technologies, Inc. | A61B005/05 A61B019/00 A61N005/10
US07666145 | B2 | 2010-02-23 | Pulse wave transmission detection system | 2003-07-03 | Provided is a pulse wave propagation detection system including an electrocardiographic signal detection unit, and an eyeground image detection unit for detecting an eyeground image in synchronization with an electrocardiographic signal detected through the detection unit, which system further includes a correlation unit for correlating a change in the diameter of an eyeground vein—which diameter is measured, at the optic papilla, by use of an eyeground image synchronized with an arbitrary electrocardiographic signal—with the state of pulse wave propagation through an intracerebral blood vessel or with the state of sclerosis of a capillary artery. | 1. A pulse wave propagation detection system comprising electrocardiographic signal detection means for detecting a first electrocardiographic signal at a first point, and for detecting a second electrocardiographic signal at a second point; andeyeground image detection means for detecting eyeground image data in synchronization with the first time point and the second time point, and for detecting pulse wave propagation in an intracerebral blood vessel on the basis of a change in a diameter of an eyeground vein, the diameter being measured at a target site of the eyeground image data synchronized with the first electrocardiographic signal detected at the first point and the second electrocardiographic signal detected at the second point. 2. A pulse wave propagation detection system according to claim 1, wherein a state of sclerosis of a capillary artery is detected on the basis of a pulse wave diagram prepared based on the change in the diameter of the eyeground vein. 3. A pulse wave propagation detection system according to claim 2, wherein the change in the diameter of the eyeground vein is a change in the diameter of the eyeground vein at an optic papilla. 4. A pulse wave propagation detection system according to claim 2, wherein the change in the eyeground vein diameter is the difference between the diameter of an eyeground vein as measured on the basis of an eyeground image synchronized with an R wave, which is an electrocardiographic signal, and the diameter of the eyeground vein as measured on the basis of an eyeground image synchronized with a T wave, which is an electrocardiographic signal. 5. A pulse wave propagation detection system according to claim 2, wherein the eyeground image detection means detects the eyeground image data, synchronized with the first electrocardiographic signal detected at the first point and the second electrocardiographic signal detected at the second point, by extracting, on a computer display, stationary eyeground images respectively synchronized with the first electrocardiographic signal detected at the first point and the second electrocardiographic signal detected at the second point from a motion eyeground image. 6. A pulse wave propagation detection system according to claim 5, wherein the eyeground image detection means extracts the stationary eyeground images respectively synchronized with the first electrocardiographic signal detected at the first point and the second electrocardiographic signal detected at the second point while displaying the motion eyeground image and an electrocardiogram on the computer display. 7. A pulse wave propagation detection system according to claim 5, wherein the eyeground image detection means comprises executing means for calculating the change in the diameter of the eyeground vein on the basis of the eyeground image data synchronized with an arbitrary electrocardiographic signal. 8. A pulse wave propagation detection system according to claim 7, wherein the ... | BML, Inc. | A61B005/02 G06K009/00
US07657299 | B2 | 2010-02-02 | Automated methods and systems for vascular plaque detection and analysis | 2004-08-21 | Automated methods and systems for the detection and analysis of plaque in one or more regions of a patient's vasculature are described. | 1. A fully automated method of determining whether a blood vessel of a patient's vasculature contains plaque by classifying plaque components, comprising computationally processing at least a first processable data type and a second processable data type obtained using at least one non-invasive medical imaging system in order to analyze at least one cross section of a blood vessel of a patient's vasculature using statistical classifiers developed using statistical modeling to determine if the blood vessel comprises at least one tissue correlated with the presence of plaque, in which event the blood vessel is determined to contain plaque, wherein the fully automated method does not require human intervention and wherein at least each of the first and second processable data types are processable magnetic resonance data generated by an MRI instrument and selected from the group consisting of T1-weighted data, T2-weighted data, PDW-weighted data, TOF-weighted data, and data generated by a combination of two or more of the foregoing. 2. A fully automated method according to claim 1 wherein the MRI instrument generates raw magnetic resonance data from which processable magnetic resonance data are derived. 3. A fully automated method according to claim 2 wherein a plurality of different radio frequency pulse sequence series are employed, wherein each of the different radio frequency pulse sequence series allows raw magnetic resonance data to be generated for a subset of different tissues that may be present in the blood vessel. 4. A fully automated method according to claim 3 wherein the components comprising the processable magnetic resonance data are computationally brought into registration. 5. A fully automated method according to claim 4 wherein registration is accomplished by aligning the components comprising the processable magnetic resonance data about a representation that represents a landmark selected from the group consisting of a vessel lumen centroid and a vessel branch point. 6. A fully automated method according to claim 5 wherein the landmark is selected from the group consisting of a physical landmark and a computational landmark. 7. A fully automated method according to claim 1 wherein the blood vessel comprises a portion of the vasculature supplying blood to an organ selected from the group consisting of a brain and a heart. 8. A fully automated method according to claim 1 wherein the blood vessel is selected from the group consisting of a carotid artery and a coronary artery. 9. A fully automated method according to claim 1 wherein the patient is human. 10. A fully automated method according to claim 1 wherein the processable data are generated by pre-processing raw data generated by the medical imaging system. 11. A fully automated method according to claim 10 further comprising normalizing the processable data prior to computationally processing the processable data. 12. A fully automated method according to claim 1 w... | Ischem Corporation | A61B005/05 A06K009/40
US07668591 | B2 | 2010-02-23 | Automatic activation of medical processes | 2004-08-20 | Systems and methods involve automatic activation, de-activation or modification of therapies or other medical processes based on brain state. A medical system includes a sensor system having one or more sensors configured to sense signals related to the brain state of the patient. A brain state analyzer detects various brain states, including sleep stage and/or brain seizures. A controller uses the brain state detection information to control a medical system configured to perform at least one respiratory or cardiac process. Methods involve sensing signals related to brain state and determining the brain state of a patient based on the sensed signals. At least one respiratory or cardiac medical process is controlled based on the patient's brain state. | 1. A system, comprising: a housing configured to be implanted;a sensor system having one or more sensors configured to sense brain activity;a brain activity detector coupled to the sensor system and configured to determine a brain state based on signals received from the sensor system;a medical system comprising a cardiac rhythm management device and configured to perform at least one respiratory or cardiac process; anda controller coupled to the brain activity detector and the medical system, the controller configured to execute program instruction stored in memory to cause the medical system to activate, de-activate or adjust the at least one cardiac or respiratory process based on the brain state, including executing stored instructions to modify a cardiac rhythm management therapy to address cardiac rhythm disturbances associated with an abnormal brain state based on identification of the abnormal brain state by at least one of the controller and the brain activity detector, wherein at least one of the sensor system and the controller are at least partially housed within the housing. 2. The system of claim 1, wherein the sensor system comprises at least one EEG sensor and the brain activity detector is configured to execute stored program instructions to determine the brain state based on EEG signals sensed by the EEG sensor. 3. The system of claim 1, wherein the sensor system comprises at least one EEG sensor and at least one EMG sensor, and the brain activity detector is configured to execute stored program instruction to determine the brain state based on EEG signals and EMG signals. 4. The system of claim 1, wherein at least one of the brain activity detector and the controller is configured to execute stored program instructions to determine relative sleep deepness, and the controller is configured to execute stored program instructions to decrease intensity of a sleep-disordered breathing therapy delivered by the medical system based on detection of a deeper level of sleep relative to a lighter level of sleep. 5. The system of claim 4, wherein at least one of the brain activity detector and the controller is configured to execute stored program instructions to identify at least REM sleep, sleep stage 3, and sleep state 4, and the controller is configured to execute stored program instructions to at least one of decrease and terminate the sleep-disordered breathing therapy based on detection of at least one of sleep stage 3, sleep stage 4, and sleep stage 4. 6. The system of claim 1, wherein the medical system comprises an external respiratory therapy device configured to deliver a disordered breathing therapy. 7. The system of claim 1, wherein the sensor system is housed within the housing and configured to sense EEG signals. 8. The system of claim 1, wherein the cardiac rhythm management device is configured to deliver a sleep-disordered breathing therapy comprising cardiac stimulation pulses and the controller is configured to ... | Cardiac Pacemakers, Inc. | A61B005/04
US08016775 | B2 | 2011-09-13 | Dual blade lancing test strip | 2009-10-19 | An integrated lancing test strip includes a pair of blade members that each have a lancing tip that are configured to lance skin. A pair of spacer members connect the blade members together such that the blade members define an internal capillary. A test strip is positioned along the internal capillary, and the test strip is configured to test analyte levels in the bodily fluid. During use, the lancing tips form one or more incisions in the skin. The fluid from the incisions is drawn via capillary action through the internal capillary and onto the test strip. | 1. A method, comprising: providing a lancing device that includes at least a pair of blade members that define an internal capillary and a test strip positioned along the internal capillary, wherein the blade members each have a lancing tip, wherein the lancing tips are angled towards one another with the lancing tips being bent relative to the rest of the blade members such that the lancing tips are in closer proximity than the rest of the blade members;lancing one or more incisions in skin with the lancing tips angled towards one another such that the lancing tips are in closer proximity than the rest of the blade members; anddrawing via capillary action between the pair blade members bodily fluid from the one or more incisions onto the test strip through the internal capillary defined by the blade members. 2. The method of claim 1, wherein said drawing the fluid includes touching a drop of the bodily fluid from the one or more incisions in the skin with the blade members. 3. The method of claim 2, wherein said lancing includes removing the blade members from the one or more incisions in the skin before said touching. 4. The method of claim 1, further comprising: wherein the lancing tips are spaced apart before said lancing; andsupporting the lancing tips by contacting the lancing tips together as the lancing tips deflect during said lancing. 5. The method of claim 1, wherein the lancing tips contact one another before said lancing. 6. The method of claim 1, wherein the lancing tips are spaced apart from one another during said lancing. 7. The method of claim 1, wherein the lancing tips contact one another during said lancing. 8. The method of claim 1, further comprising: venting air from the internal capillary through a vent slot located past the test strip along the internal capillary to ensure the bodily fluid covers the test strip before reaching the vent slot during said drawing. 9. The method of claim 8, wherein the vent slot is defined between the test strip and a vent block in which the test strip extends across at least two spacer members. 10. The method of claim 1, further comprising: analyzing the bodily fluid with the test strip. 11. The method of claim 1, further comprising: wherein said drawing the fluid includes touching a drop of the bodily fluid from the one or more incisions in the skin with the blade members;wherein said lancing includes removing the blade members from the one or more incisions in the skin before said touching;wherein the lancing tips are spaced apart before said lancing;supporting the lancing tips by contacting the lancing tips together as the lancing tips deflect during said lancing;venting air from the internal capillary through a vent slot located past the test strip along the internal capillary to ensure the body fluid covers the test strip before reaching the vent slot during said drawing;wherein the vent slot is defined between the test strip and a vent block in which the test strip extends a... | Roche Diagnostics Operations, Inc. | A61B005/00 B65D081/00
US08236282 | B2 | 2012-08-07 | Benzothiazole derivative compounds, compositions and uses | 2009-09-30 | This invention provides benzothiazole derivative compounds, compositions comprising such compounds, methods of preparing such compounds, and methods of using such compounds for detecting amyloid deposit(s) and for diagnosing a disease, disorder or condition characterized by amyloid deposit(s). | 1. A compound of formula Ior a pharmaceutically acceptable salt, hydrate, solvate or prodrug of the compound, wherein: R1 is hydrogen, —OH, —NO2, —CN, —COOR, —OCH2OR, C1-C6 alkyl, C2-C6 alkenyl, C2-C6 alkynyl, C1-C6 alkoxy or halo;R is C1-C6 alkyl;R2 is a radiofluoro;R3 is hydrogen, C1-C6 alkyl, C2-C6 alkenyl or C2-C6 alkynyl; andR4 is hydrogen, C1-C6 alkyl, C2-C6 alkenyl or C2-C6 alkynyl. 2. The compound of claim 1, wherein: R1 is —OH or C1-C6 alkoxy;R2 is 18F;R3 and R4 are independently selected from hydrogen and C1-C6 alkyl. 3. The compound of claim 1 having the following structure: 4. A pharmaceutical composition comprising (i) an effective amount of a compound of claim 1; and(ii) a pharmaceutically acceptable carrier. | University of Pittsburgh—of the Commonwealth System of Higher Education | A61K051/00 A61M036/14
US08123694 | B2 | 2012-02-28 | Electro pneumatic interface for blood pressure system | 2008-07-18 | A blood pressure system that includes a compact monitor housing that contains a pneumatic circuit and an electrical circuit and a processor for inflating and deflating a cuff to provide blood pressure related data. A split connector allows the housing to interface both pneumatically and electronically with one or more independent cuffs and/or with one or more host stations. Each host station contains a dependent cuff and further electrical components to expand and enhance the systems output of related blood pressure data that is generated by either the independent or dependent cuffs. | 1. A blood pressure system comprising: a compact housing that contains a pneumatic circuit that includes a pump for inflating a blood pressure cuff, a valve for deflating said cuff, a pressure sensor for providing cuff pressure data and an electrical circuit having a control means for programming said pneumatic circuit and for processing cuff pressure data;said housing further including a display means for visually presenting said processed data; anda split connector comprising: a female member mounted upon said housing, anda male member mounted upon one or more independent blood pressure cuffs or alternatively one or more host stations, wherein each host station contains a dedicated blood pressure cuff,wherein said female member is adapted to releasably close with said male member to establish a hoseless connection between said pneumatic circuit and said one or more independent cuffs and said one or more host stations and a wireless electrical connection between said electrical circuit and said one or more host stations,wherein said split connector also establishes a wireless electrical connection between said electrical circuit and said one or more independent cuffs when said split connector is closed,wherein said hoseless pneumatic connection provides a leak-tight joint between said female member and said male member when said members are brought together at closure,wherein said wireless electrical connection includes one or more first electrical contacts mounted adjacent to one of said female member and said male member that ride in sliding contact with one or more second electrical contacts encircling the other of said female member and male member when the two receptors are brought together at closure and when said housing is rotated about an axis of said split connector. 2. The blood pressure system of claim 1, wherein said housing further contains a rechargeable battery for providing power to electrical components contained in said housing. 3. The blood pressure system of claim 2, wherein said host station further includes means for recharging said rechargeable battery when said split connector is closed. 4. The blood pressure system of claim 3, wherein said host station contains additional processing means for further storing and processing blood pressure related data from said microprocessor when said split connector is closed. 5. The blood pressure system of claim 4, wherein said host station further includes an enlarged display for visually presenting blood pressure related data. 6. The blood pressure system of claim 5, wherein said host station further includes a bi-directional communication unit for exchanging blood pressure related data with at least one other station. 7. The blood pressure system of claim 1, wherein said control means is a microprocessor that contains storage means for storing blood pressure related data. 8. The blood pressure system of claim 1, wherein said one or more independent cuffs each contain a dete... | Welch Allyn, Inc. | A61B005/02
US08060191 | B2 | 2011-11-15 | Wireless cardiogram signal diagnostic instrument | 2009-03-03 | A wireless cardiogram signal diagnostic instrument includes a band, a signal transmitting device, two sets of connecting pieces, and a conductive assembly. The band has two opposite mounting portions separated by a distance. The two mounting portions are provided with the conductive assembly having a first conductive piece and a second conductive piece. The signal transmitting device is provided with a third conductive piece. The two sets of connecting pieces each have a first connecting piece that is conductive provided on the mounting portion of the band and a second connecting piece that is conductive provided on the signal transmitting device. The first connecting piece is correspondingly assembled with the second connecting piece, so that the signal transmitting device can be detachably assembled on the band. Via the above arrangement, the instrument can be worn on a user's chest to collect and measure the variation values of cardiogram voltages. | 1. A wireless cardiogram signal diagnostic instrument, comprising: a band having a first surface and a second surface opposite to each other, and having two opposite mounting portions separated by a distance;a signal transmitting device having an outer casing, a circuit board provided in the outer casing, and a signal transmitting element, an arithmetic processing element, a first electrode, a second electrode and a third electrode that are electrically connected to the circuit board, the outer casing having a third surface and a fourth surface opposite to each other, the third surface of the outer casing being oriented in the same direction as the first surface of the band, the fourth surface of the outer casing being oriented in the same direction as the second surface of the band;two sets of connecting pieces each having a first connecting piece and a second connecting piece that is conductive, the first connecting pieces being provided on the two mounting portions of the band and electrically coupled to a corresponding first conductive piece and a corresponding second conductive piece respectively, the second connecting pieces being provided on the outer casing of the signal transmitting device and electrically coupled to the corresponding first and second electrodes respectively, the first connecting pieces being assembled with the corresponding second connecting pieces so that the signal transmitting device is detachably assembled on the second surface of the band; anda conductive assembly having a first conductive piece, a second conductive piece, and a third conductive piece, the first and second conductive pieces being nakedly provided on the first surfaces of two mounting portions of the band, the third conductive piece being nakedly provided between the first surfaces of the band and the third surface of the outer casing with the third conductive piece being located between the two mounting portions, the third conductive piece and the third electrode being electrically coupled to each other;wherein the band further comprises a middle portion for integrally connecting the two mountins portions, the third conductive piece is provided on a first surface of the middle portion; anda third connecting piece that is conductive, and a fourth connecting piece that is conductive, the third connecting piece being provided on the outer casing of the signal transmitting device and electrically coupled to the third electrode, the fourth connecting piece being provided in the middle portion of the band and electrically coupled to the third conductive piece, the third connecting piece being assembled with the fourth connecting piece to thereby electrically couple the third conductive piece and the third electrode. 2. The wireless cardiogram signal diagnostic instrument according to claim 1, wherein the third conductive piece is provided on the third surface of the outer casing of the signal transmitting device to be electrically coupled to the third e... | Zentan Technology Co., Ltd. | A61B005/0402 A61B005/0416 A61B005/044
US08278926 | B2 | 2012-10-02 | Method for determining attenuation values of an object | 2009-07-09 | A method for determining attenuation values of an object is disclosed. In at least one embodiment, the method includes stationary positioning of the object, irradiation of the object via a radiation source, measurement of the object's transmission data via a detection system, determination of at least one geometric property of the object on the basis of the transmission data, and assignment of attenuation values to the object on the basis of the geometric property. | 1. A method for determining attenuation values of an object using a medical imaging scanner, the method comprising: stationaryily positioning the object;irradiating the object using a radiation source;measuring transmission data of the object via a detection system of the medical scanner;determining at least one geometric property of the object on the basis of the measured transmission data; andassigning attenuation values to the object on the basis of the at least one determined geometric property. 2. The method as claimed in claim 1, wherein the detection system comprises a plurality of detectors and the geometric property is a diameter of the object, the determination of the diameter comprising: comparing signal intensities of the detectors;identifying the detectors which have detected radiation attenuated by the object on the basis of the signal intensity; anddetermining the diameter on the basis of the geometric position of the identified detectors and the beam path from the radiation source to the respective detector. 3. The method as claimed in claim 2, wherein the assigning of attenuation values to the object comprises: determining at least one material contained in the object;determining attenuation values of the at least one material; andassigning the determined attenuation values to the object. 4. The method as claimed in claim 3, wherein the assigned attenuation values of the object are added to an attenuation map which already contains attenuation values of other objects. 5. The method as claimed in claim 2, wherein the assigned attenuation values of the object are added to an attenuation map which already contains attenuation values of other objects. 6. The method as claimed in claim 1, wherein the assigning of attenuation values to the object comprises: determining at least one material contained in the object;determining attenuation values of the at least one material; andassigning the determined attenuation values to the object. 7. The method as claimed in claim 6, wherein the assigned attenuation values of the object are added to an attenuation map which already contains attenuation values of other objects. 8. The method as claimed in claim 1, wherein the assigned attenuation values of the object are added to an attenuation map which already contains attenuation values of other objects. 9. The method as claimed in claim 1, wherein the object is an arm of a patient. 10. The method as claimed in claim 1, wherein the stationary positioning of the object is effected inside a patient bore of an MR/PET system. 11. The method as claimed in claim 1, wherein a positioning aid is used for stationary positioning. | Siemens Aktiengesellschaft | G01V003/00
US08155411 | B2 | 2012-04-10 | Method, apparatus and computer program for quantitative bifurcation analysis in 3D using multiple 2D angiographic images | 2008-07-22 | A computer-implemented method (and corresponding data processing facility and computer program) for creating a 3D bifurcation healthy model from multiple 2D angiographic images comprises the following: a. Creating a 3-D model based on said images;b. Defining a bifurcation region regarding said 3D model;c. Creating an area curve regarding said region;d. Creating a reference area curve regarding said region whilst reconstructing in 2D;e. Creating a healthy model regarding said area curves; andf. Computing quantitative analysis results regarding said healthy model based on the healthy cross-sectional areas of the bifurcation, obstruction boundaries and the diseased 3D model consisting of the bifurcation centerlines and cross-sections of the lumen on the centerline, the healthy centerline and cross-sections. | 1. A method for analysis of a bifurcated tubular organ using multiple 2D angiographic images of the organ, said method comprising: a. Creating a first 3D model based on said images;b. Defining a bifurcation region regarding said first 3D model;c. Creating an area curve regarding said first 3D model, including c.1 deriving a 2D representation of the bifurcated tubular organ based upon said bifurcation region, said 2D representation defining a bifurcation including a main branch and two side branches,c.2 converting said 2D representation to define two respective vessels formed by splitting of the bifurcation,c.3 measuring various diameters regarding said respective vessels, andc.4 combining 2D quantities regarding said respective vessels;d. Creating a reference area curve regarding the bifurcated tubular organ whilst reconstructing in 2D;e. Creating a 3D healthy model of the bifurcated tubular organ based upon said area curves; andf. Computing quantitative analysis results regarding said 3D healthy model of the bifurcated tubular organ. 2. A method as claimed in claim 1, wherein: both said defining a bifurcation region of b and said creating a reference area curve of d include selecting respective instances of said creating an area curve of c. 3. A method as claimed in claim 1, wherein: said 3D healthy model of the bifurcated tubular organ is based upon said first 3D model, said first 3D model including at least one of centerlines for portions of said bifurcation region, circular cross-sectional areas of said bifurcation region along the centerlines, and obstruction boundaries. 4. A method as claimed in claim 1, wherein: said creating a reference area curve of d includes d.1 Reconstructing a single branch;d.2 Mapping said single branch reconstruction to 2D;d.3 Reconstructing the bifurcation region in 2D; andd.4 Creating an area curve regarding said reconstruction of d.3. 5. A method as claimed in claim 1, wherein: the splitting of the bifurcation of c.2 that defines two respective vessels is effected by the step of defining a virtual edge for each respective vessel. 6. A method as claimed in claim 4, wherein: said reconstructing of d.3 involves processing said images to identify contours of said organ, using said contours to determine a Polygon of Confluence amongst said organ, using said Polygon of Confluence to determine at least one parameter value characterizing geometry of said organ and outputting said parameter value to a user. 7. A method as claimed in claim 6, wherein: at least one reference diameter is derived from a set of annuli with an inner circular edge fitted to an inner edge curve and an outer circular edge fitted to an opposite edge curve. 8. A method as claimed in claim 7, wherein: at least one derived parameter comprises at least one reference diameter value for said bifurcation region, and said reference diameter compensating for damage to the bifurcation region. 9. A method as claimed in claim 8, wherein: said at least... | Pie Medical Imaging B.V. | G06K009/00
US08617068 | B2 | 2013-12-31 | Method and apparatus for assessing sleep quality | 2007-09-27 | Systems and/or methods for assessing the sleep quality of a patient in a sleep session are provided. Data is collected from the patient and/or physician including, for example, sleep session data in the form of one or more physiological parameters of the patient indicative of the patient's sleep quality during the sleep session, a subjective evaluation of sleep quality, etc.; patient profile data; etc. A sleep quality index algorithm, which optionally may be an adaptive algorithm, is applied, taking into account some or all of the collected data. Sleep quality data may be presented to at least the patient, and it may be displayed in any suitable format (e.g., a format useful for the patient to be appraised on the progress of the treatment, a format useful for a sleep clinician to monitor progress and/or assess the effectiveness of differing treatment regimens, etc). | 1. A method of assessing sleep quality of a patient in a sleep session, the method comprising; obtaining sleep session data comprising at least two physiological parameters of the patient indicative of the patient's sleep quality during the sleep session; andcalculating, via at least one processor, a composite sleep quality index from said sleep session data determined over a period comprising two or more sleep sessions by deriving a score index for each sleep session data physiological parameter dependent on subjective patient feedback data relating to previous sleep sessions and calculating said sleep quality index using said score indices;wherein the composite sleep quality index is calculated by adding together all of the score indices after at least one score index is adjusted based on the subjective feedback correspond to healthy sleep of the patient; andcommunicating the composite sleep quality index to at least the patient or a healthcare provider. 2. The method according to claim 1, further comprising visually displaying the composite sleep quality index to at least the patient. 3. The method according to claim 1, wherein the sleep session data includes at least two of sleep duration, sleep fragmentation, sleep structure, Apnea-Hypopnea Index (AHI), sleep latency, and snore for the sleep session. 4. The method according to claim 1, wherein said deriving is dependent on patient profile data. 5. The method according to claim 4, wherein said patient profile data includes at least one of patient age, sex, weight, ethnicity, and sleep disorder data. 6. The method according to claim 1, wherein said calculating is further practiced by performing a relative weighting of said sleep session parameters. 7. The method according to claim 6, wherein said relative weighting of said sleep session parameters is dependent on patient profile data. 8. The method according to claim 7, wherein said patient profile data includes at least one of: patient age, sex, weight, ethnicity, and sleep disorder data. 9. The method according to claim 6, wherein said relative weighting of said sleep session parameters is dependent on optimal sleep characteristic data for the patient. 10. The method according to claim 9, wherein the optimal sleep characteristic data for the patient is derived from patient profile data. 11. The method according to claim 9, wherein the optimal sleep characteristic data for the patient is derived from patient feedback data. 12. The method according to claim 1 further comprising: recording and storing subjective patient feedback data electronically. 13. The method according to claim 1, wherein said communicating is further practiced by displaying the index on a display of a sleep disorder treatment apparatus. 14. The method according to claim 1, wherein said communicating is further practiced by communicating the index via a graphical display. 15. The method according to claim 1, wherein said communicating is further practiced by ... | ResMed Limitied | A61M021/00
US08821408 | B2 | 2014-09-02 | Data mining system for noninvasive intracranial pressure assessment | 2007-04-05 | Systems and methods are described for noninvasively assessing an intracranial pressure of a patient. Some embodiments include providing a simulation model with a measured arterial blood pressure of the patient. Some embodiments further include providing the simulation model with a measured cerebral blood flow velocity of the patient. The simulation model correlates arterial blood pressure values, cerebral blood flow velocity values, and intracranial pressure values. Some embodiments further includes determining an intracranial pressure of the patient based on the simulation model. Some embodiments further includes creating an output data set indicative of the determined intracranial pressure. | 1. A system, for determining an estimated intracranial pressure (ICP) of a first patient, comprising: a processing module that: receives a first input based on a first plurality of measurements of a first physiological parameter of the first patient, the first physiological parameter correlating with ICP;receives a second input based on a second plurality of measurements of a second physiological parameter of the first patient, the second physiological parameter correlating with ICP;processes the first and second inputs based on an ICP model; andoutputs data indicative of the estimated ICP of the first patient;wherein the ICP model is chosen from a plurality of potential ICP models, each of the potential ICP models corresponding to a respective model patient, other than the first patient, and representing a transform function that transforms (a) time series information indicative of measurements of the first and second physiological parameters of the respective model patient to (b) time series information indicative of measurements of ICP of the respective model patient;wherein each of the plurality of models transforms (a) time series information indicative of measurements of the first and second physiological parameters of an index patient other than the respective model patient to (b) time series information indicative of an estimated ICP of the index patient, wherein a dissimilarity value indicates a dissimilarity between the index patient's estimated ICP and the index patient's observed ICP;wherein each of the potential ICP models has an associated mapping function that maps values of said first and second physiological parameters, of each patient in a group of patients other than the respective model patient, to at least one dissimilarity value corresponding to each of the patients in the group;wherein said ICP model is chosen based on dissimilarity values obtained by applying the mapping functions of each of the potential ICP models to data indicative of the first patient's first and second pluralities of measurements, to determine a dissimilarity value corresponding to each of the potential ICP models. 2. The system of claim 1, wherein a hemodynamic feature vector comprises said values of, or derived values from, said first and second physiological parameters. 3. The system of claim 1, wherein said ICP model is chosen based on dissimilarity values obtained by choosing an optimal dissimilarity value from among the corresponding dissimilarity values, based on at least one criterion. 4. The system of claim 3, wherein each of the corresponding dissimilarity values comprises a scalar value. 5. The system of claim 3, wherein a hemodynamic feature vector comprises derived values from said first patient's said first plurality of measurements and second plurality of measurements. 6. The system of claim 3, wherein the at least one criterion comprises determining a minimum dissimilarity value from among the corresponding dissimilarity values. ... | The Regents of the University of California | A61B005/00 A61B005/03 G05B013/02
US08844525 | B2 | 2014-09-30 | Method and apparatus for detecting and treating heart failure | 2009-06-12 | Devices and systems provide methods of detecting a heart failure condition of a patient that may be based on one or more respiratory parameters of a patient. In an example embodiment, a monitoring device determines one or more heart failure condition indicators based on a measure of the patient respiratory airflow and/or a measure of treatment pressure. Respiratory parameters such as respiration rate, hypopneas, apneas, Cheyne-Stokes breathing patterns or apnea-hypopnea counts may be compared to thresholds that are selected to represent a change in the condition of a heart failure patient such as an onset of a decompensation event. Results of the comparisons may trigger a pressure treatment change and/or one or more warnings or messages to notify a patient or physician of a pending change to the patient's heart failure condition so that the patient may more immediately seek medical attention to treat the heart failure condition. | 1. A device for monitoring a patient to evaluate a heart failure condition of a patient comprising: a patient interface;a non-invasive sensor to generate a signal associated with a characteristic of patient respiration; anda processor, coupled with the sensor, the processor configured to control and store a determination of a heart failure condition change indicator based on the signal, compare previously stored and determined values of the indicator from one or more previous treatment sessions and determine a risk of a heart failure decompensation event occurring for the patient based on the comparison, wherein the processor is further configured to trigger a heart failure decompensation risk message based on the comparison. 2. The device of claim 1 wherein a determination of a heart failure condition change indicator is based on a threshold comparison that involves a respiratory rate determined from the signal of the non-invasive sensor. 3. The device of claim 2 further comprising a sensor to generate a signal representative of a respiratory treatment pressure, and wherein a determined heart failure condition change indicator is based on (a) an additional respiratory parameter determined from the respiration measure and (b) a determined treatment pressure, wherein the additional respiratory parameter is an apnea-hypopnea count or apnea count. 4. The device of claim 1 wherein the processor activates a warning based on the heart failure condition change indicator. 5. The device of claim 4 wherein the warning comprises a light signal or a message. 6. The device of claim 4 wherein the message comprises an electronic message to be transmitted. 7. The device of claim 1 further comprising a transmitter coupled with the processor, wherein the processor is further configured to control transmitting of a message including data representing the heart failure condition indicator. 8. The device of claim 1 further comprising a flow generator coupled to the processor and the patient interface, wherein the processor is configured to control generating a flow of breathable gas at a pressure above atmospheric to the patient interface. 9. The device of claim 1 wherein the processor is configured to determine a heart failure condition change indicator with a plurality of threshold comparisons involving a plurality of respiratory parameters determined from a measure of respiratory airflow and wherein the characteristic of respiration is the measure of respiratory airflow. 10. The device of claim 9 wherein the respiratory parameters comprise data representative of one or more apneas and hypopneas experienced by the patient and data representative of one or more Cheyne-Stokes epochs experienced by the patient. 11. The device of claim 1 wherein the processor is configured to determine the heart failure condition change indicator with a threshold comparison involving an oximetry measure and wherein the characteristic of respiration is the oximetry measure. 1... | ResMed Limited | A61M016/00 A61B005/087 A61B005/0205 A61B005/08 A61M016/06
US08523780 | B2 | 2013-09-03 | Electronic device and system for detecting rejection in transplant recipients | 2008-01-17 | The current invention relates to an implantable device, system, method, and computer readable medium comprising computer executable components for detection of rejection of a cardiac allograft. Acute rejection and chronic rejection are the leading cause of death in heart transplant recipients post transplant. The current invention seeks to address this problem by providing methods, devices, systems and computer readable media comprising computer executable components to monitor the efficacy of immunosuppressant therapy and evaluate the degree of rejection in a cardiac allograft. To provide this functionality, an implantable device comprises sensors which measure physiologic parameters at specified intervals after the transplant procedure. These measured parameters are subsequently compared with baseline parameter data to detect rejection of the cardiac allograft. | 1. A method for detecting cardiac allograft rejection in a subject that has received a cardiac allograft, comprising: determining baseline data regarding the subject and the cardiac allograft;measuring parameters of the cardiac allograft during a measurement interval;determining average parameter values for the parameters during the measurement interval; andcomparing the average parameter values with the baseline data to detect the presence of a cardiac allograft rejection;wherein the parameters comprise: impedance within the chest measured without implanting leads into the tissue of the cardiac allograft, electrical activity generated by the cardiac allograft, at least one voltage generated by the cardiac allograft, rate of change of the at least one voltage generated by the cardiac allograft, and body temperature. 2. The method of claim 1, further comprising providing an implantable device for detecting rejection of the cardiac allograft in the subject, comprising: a plurality of sensors to measure the at least one parameter;a processor in electronic communication with the plurality of sensors, the processor configured to receive at least one measured parameter;a memory component to store the baseline data, data corresponding to the at least one measured parameter and/or data corresponding to the at least one average parameter value; anda power supply component to provide electricity to components of the implantable device. 3. The method of claim 1, further comprising determining a presence of a statistically significant deviation from the baseline data, whereby the presence of the statistically significant deviation is indicative of an acute or a chronic rejection episode. 4. The method of claim 1, further comprising annotating the at least one average parameter value if the at least one parameter value differs by more than one standard deviation from the baseline data. 5. The method of claim 1, further comprising annotating the at least one average parameter value if the at least one parameter value differs by more than two standard deviations from the baseline data. 6. The method of claim 1, wherein collecting the data used to determine the baseline data begins about five days after a cardiac allograft transplant procedure. 7. The method of claim 1, wherein determining the baseline data comprises collecting data for about thirty minutes a day for a period of about one week and determining the average values of the data. 8. The method of claim 1, wherein the measurement interval is selected from the group consisting of about 50, 100, 200, 300, 400, 500, and 1000 contraction cycles of the cardiac allograft. 9. The method of claim 8, wherein the measurement interval is about one hundred contraction cycles of the cardiac allograft. 10. The method of claim 1, wherein the measurement interval is selected from the group consisting of about 0.5, 1, 2, 3, 4, 5, 10, 15, 20, 30, 45, and 60 minutes. 11. The method of claim 1, wherein the imped... | Cedars-Sinai Medical Center | A61B005/0402
US08406863 | B2 | 2013-03-26 | Device for using saved frequency information | 2007-12-31 | The invention relates to a device for applying saved frequency information in the form of modulated magnetic fields and/or modulated light to the body, wherein the device comprises at least one transceiver, one display, one memory, one control electronic, one D/A-converter, and a coil (6) for the application of modulated magnetic fields and/or a laser (7) for the application of modulated light. The invention also relates to a method for tapping signals of the electric potential of the heart and the pulse value and applying saved frequency information by means of a device according to the invention. | 1. A device for applying frequency information saved in a server in the form of modulated magnetic fields and modulated light to the body, the device comprising: at least one transceiver for downloading the saved frequency information from the server,one display,one memory connected to the transceiver for storing the downloaded frequency information,one electronic controller,one D/A-converter connected to the memory for receiving the downloaded frequency information therefrom,a coil connected to the D/A converter for applying magnetic fields modulated by the downloaded frequency information,a laser connected to the D/A converter for applying light modulated by the downloaded frequency information,two sensors for measuring the electrical potential of the heart of the body, anda phototransistor for measuring the pulse of the body. 2. The device as described in claim 1 wherein the device is a cell phone. 3. The device as described in claim 1 wherein the magnetic field of the coil has a fluctuating frequency with a mean value of 9 Hz. 4. The device as described in claim 1 wherein the magnetic field is oriented symmetric transverse to the longitudinal axis of the device. 5. The device as described in claim 1 wherein the magnetic field created by the coil has a magnetic field strength of 0.1 to 70 μT at a measuring distance of 0 m. 6. The device as described in claim 1 wherein the laser is a semiconductor laser-diode which creates light with a fluctuating red wave length with a mean value of 650 nm. 7. The device as described in claim 1 wherein the laser has a power of less than 1 mW. 8. The device as described in claim 1 wherein the frequency information to be applied is available in form of a digital file format. 9. The device as described in claim 1 wherein the sensors are metallic electrodes and the phototransistors and electrodes are each arranged near edges of a cell phone for the application of one finger to each sensor. 10. The device as described in claim 1 wherein the electronic controller of the device processes the signals received from the sensors for the computation of a heartrate-variability-heart-coherence-analysis and shows the result on the display. 11. The device as described in claim 10 wherein the result computed by the electronic controller is compared with a database in the memory of the device, and that on the display of the device a corresponding selection of applicable frequency information and additional further information can be shown, the shown selection of applicable frequency information being downloadable from the server by the transceiver. 12. The device as described in claim 1 wherein the sensors, the memory, the electronic controller, the D/A-converter, the coil or the laser are arranged in a common module connectable to a cell phone or built into the cell phone. 13. The device as described in claim 12 wherein the module further comprises its own current supply. 14. The device as described in claim 12 wh... | - | A61N001/00
US09345418 | B2 | 2016-05-24 | EEG net with transmission capabilities | 2008-03-27 | An electroencephalography system is provided. The electroencephalography system includes a support structure that covers at least predetermined areas on a patient's head, a plurality of electrodes mounted on the support structure so that, the electrodes are distributed around the patient's head and measure usable electrical signals that are representative of electrical activity or activity in a patient's head, a plurality of transmitting wires incorporated into the support structure, a multi-pin connector having an output and an input, each of the plurality of transmitting wires directly connected between a respective one of each of the plurality of electrodes and a corresponding pin of the multi-pin connector so that, electrical signals are sent to the output of the connector and a first transmitter that is capable of being operatively coupled to the output of the connector. | 1. An electroencephalography system comprising: a support structure that is adapted to cover at least predetermined areas on a patient's head when it is mounted on a patient's head;a plurality of electrodes that are mounted on the support structure so that, when the support structure is mounted on a patient's head, the electrodes are distributed around the patient's head and are capable of measuring usable electrical signals that are representative of electrical activity or activity in a patient's head;a plurality of transmitting wires incorporated into the support structure;at least one multi-pin connector having an output and an input, with the plurality of transmitting wires hard-wired into the multi-pin connector and where each of the plurality of transmitting wires is directly connected between a respective one of each of the plurality of electrodes and a corresponding pin of the input of the multi-pin connector so that, when the headpiece is mounted on a patient's head and when the electrodes are detachably coupled to the predetermined portions of a patient's head, electrical signals that are representative of electrical brain activity of the patient are sent to the output of the connector; anda first transmitter that is capable of being operatively coupled to the output of the connector. 2. The system of claim 1, wherein each one of the plurality of electrodes are detachably coupled to the support structure. 3. The system of claim 1, where in the support structure includes at least one strap. 4. The system of claim 1, wherein the first transmitter is adapted to transmit the electrical signals through cellular transmission. 5. The system of claim 1, wherein the first transmitter is mounted directly on the support structure. 6. The system of claim 5, further comprising a second transmitter, the first transmitter being designed to send electrical signals representative of brain activity inside the patient's head to the second transmitter which is designed to retransmit the electrical signals to a remote location where the signals can be shown on a display and viewed by medical personnel. 7. The system of claim 6, wherein the second transmitter is adapted to retransmit the electrical signals to the remote location through cellular transmission. 8. The system of claim 6, wherein the second transmitter is adapted to retransmit the electrical signals to the remote location through the internet. 9. The system of claim 1, wherein the connector is adapted to convey electrical signals representative of an EEG measurement, audio signals, and/or an EKG measurement to its output. 10. The system of claim 1, wherein each one of the plurality of electrodes comprise needle electrodes, disc electrodes or pronged electrodes. 11. The system of claim 1, wherein each one of the electrodes is plated with stainless steel, platinum, ruthenium, silver, silver chloride, tin or gold. 12. The system of claim 1, wherein each electrode comprises a plurality of... | Dignity Health | A61B005/0476 A61B005/0478 A61B005/00
US10650915 | B2 | 2020-05-12 | Clinician decision support system | 2007-10-08 | A bedside patient monitor comprises a clinician support system. The clinician support system comprises at least one of the following modules of: a disease-specified decision module (11), disease-specified treatment module (13) or disease-specified observation module (15). The bedside monitor comprises a first interface (105) for manual patient-related input and a second interface (107) for continuous input of a sensor signal. The clinician is supported to deliver requested information and execute the treatment in a predetermined way supported by the clinician support system. | 1. A bedside patient monitor comprising: a clinician support system,a first interface including an input configured to manually input medical information related to a patient associated with the bedside monitor and a display configured to continuously display physiological values of the patient;a second interface configured to continuously input sensor signals from sensors configured to continuously measure the physiological values of the patient;wherein the clinician support system comprises: at least one disease-specific treatment module which supports the clinician in a treatment process, the disease-specific treatment module being started only by a manual input by a qualified clinician to the first interface, wherein the manual input includes a fingerprint and the disease-specific treatment module uses the fingerprint to verify that a person inputting the manual input is authorized to start the disease-specific treatment module;at least one disease-specific observation support module configured for observation of the continuously measured physiological values of the patient, the disease-specific observation module being activated by a signal generated through deactivation of the disease-specific treatment module;wherein the disease-specific treatment module is deactivated in response to the manual input by the qualified clinician confirming that the disease-specific treatment module is to be deactivated and the disease-specific observation support module started to observe the continuously measured physiological values of the patient; and,wherein the disease-specific observation support module is configured to: in response to the deactivation of the disease-specific treatment module, display a list of observation recommendations of the specified disease on the first interface display. 2. The bedside patient monitor as claimed in claim 1, wherein the clinical support system further includes: a disease-specific decision module which is configured to compare the sensor signals and a manual patient-related input with disease-specific detection guidelines to detect early signs and symptoms of a specified disease and in response to detecting the early signs and symptoms, generating a display on the first interface display indicating the presence of the signs and symptoms of the specified disease. 3. The bedside patient monitor as claimed in claim 2, comprising an indicator for signaling that the disease-specific treatment module is running. 4. The bedside patient monitor as claimed in claim 2, wherein the clinician support system comprises an information delivery tool to deliver requested information requested by at least one of the disease-specific modules, wherein the information delivery tool generates a pop-up window to trigger the manual patient-related input through the first interface if the requested information is not available through the interface for automatic input or connected remote sources. 5. The bedside patient monitor as clai... | KONINKLIJKE PHILIPS N.V. | G16H010/20 G06F019/00
US09833183 | B2 | 2017-12-05 | Body fluid sampling device—sampling site interface | 2009-06-01 | An arrangement for producing a sample of body fluid from a wound opening created in a skin surface at a sampling site includes: a housing, the housing comprising a first opening; a skin interface member disposed in the first opening, the skin interface member comprising an inner member having a second opening, and an outer member at least partially surrounding the inner member and attached to the first opening; and at least one skin-penetration member configured and arranged to project within the second opening. Arrangements having alternatively constructed skin interface members are also described. | 1. An arrangement for producing a sample of body fluid from a wound opening created in a skin surface at a sampling site, the arrangement comprising: a housing, the housing comprising a plate member releasably mountable thereto, the plate member comprising a first opening;a skin interface member disposed in the first opening and attached thereto, the skin interface member comprising an inner member having a second opening, and an outer member at least partially surrounding the inner member, the inner member having a first hardness and the outer member having a second hardness less than the first hardness, and wherein a perimeter portion of the outer member is affixed to the plate member and wherein the outer member is coplanar with the plate member and is configured to flex more than the inner member upon application of force to the skin interface member; andat least one skin-penetration member configured and arranged to project within the second opening. 2. The arrangement of claim 1, wherein the housing is shaped and sized to be hand-held. 3. The arrangement of claim 1, wherein the skin interface member is removable from the first opening. 4. The arrangement of claim 1, wherein the second opening comprises an inner diameter of 3 mm-8 mm. 5. The arrangement of claim 4, wherein the inner diameter of the second opening is 6 mm. 6. The arrangement of claim 1, wherein the inner member and the outer member are integral. 7. The arrangement of claim 1, wherein the outer member is formed from elastomer, silicone rubber, soft plastic, or metal and the inner member is formed from metal, plastic or ceramic. 8. The arrangement of claim 1, wherein at least one of the housing, the inner member and the outer member is constructed so as to emit visible light thereby facilitating use in low-light conditions. 9. The arrangement of claim 1, wherein at least one of the housing, the inner member and the outer member is constructed so as to emit thermal energy. 10. The arrangement of claim 1, wherein the inner member comprises a textured surface for improving tactile feel of the user. 11. The arrangement of claim 1, wherein at least one of the inner member, the outer member, and the housing comprise one or more sensors capable of detecting when the skin interfacing member contacts the skin surface of the user at the sampling site. 12. The arrangement of claim 11, wherein the at least one sensor is capable of detecting the amount of force between the skin interfacing member and the skin surface. 13. The arrangement of claim 12, wherein the arrangement is constructed to provide audible and/or visual feedback to the user based on information derived from the sensor. 14. The arrangement of claim 12, wherein the arrangement is constructed to automatically initiate one or more functions or operations based on information derived from the sensor. 15. The arrangement of claim 14, wherein the arrangement is constructed to automatically initiate an automated samp... | Intuity Medical, Inc. | B65D081/00 A61B005/15 A61B005/151 A61B005/157 A61M005/42
US20100010317 | A1 | 2010-01-14 | SELF-CONTAINED DATA COLLECTION SYSTEM FOR EMOTIONAL RESPONSE TESTING | 2008-07-09 | A system and method is provided for providing self-contained data collection systems for emotional response testing of one or more subjects based, for example, on eye properties of the subjects. More particularly, market research may be conducted through testing emotional responses of one or more subjects at one or more self-contained data collection systems. The emotional responses of the subjects may be tested based on eye properties that may be indicative of subconscious physiological reactions that evidence a given emotional response (e.g., blink rate, eye movement, pupil dilation, etc.). Administering tests at the self-contained data collection systems may include presenting one or more tests stimuli to the subjects (e.g., visual stimuli, auditory stimuli, etc.), wherein the self-contained data collection systems may have various systems operable therein to analyze properties of the subjects' eyes to determine whether and/or how the subjects emotionally respond to the stimuli and/or to transmit the data to another location for analysis. | 1. A self-contained computer-implemented data collection system for emotional response testing, comprising: a calibration module configured to adjust one or more devices used in an emotional response test, wherein the devices are adjusted according to an environment and a subject of the emotional response test;a subject module configured to identify the subject and determine one or more profiles for the identified subject;a test selection module configured to select one or more tests from an emotional response test database based on the profiles for the identified subject;one or more output devices configured to present one or more stimuli to the subject during the selected emotional response test; andone or more input devices configured to receive input from the subject during the emotional response test, wherein at least one of the input devices are configured to measure and process eye data to determine the subject's emotional response to the stimuli presented during the emotional response test. 2. The system of claim 1, wherein the profiles for the identified subject include at least one of an information profile or an emotional response profile. 3. The system of claim 1, wherein the input devices include at least one of an emotion detection sensor, a biometric sensor, a physical attribute sensor, an environment sensor, a distance detection sensor, or another sensor or sensory device. 4. The system of claim 1, wherein the measured and processed eye data includes one or more of pupil dilation, blink rate, eye movement, gaze sequence, or other eye properties. 5. The system of claim 1, further comprising a rewards module configured to dispense one or more awards to the subject upon completion of the emotional response test. 6. The system of claim 1, further comprising a collected data database configured to store data regarding the subject's emotional responses to the stimuli presented during the emotional response test. 7. The system of claim 6, further comprising an analysis results database configured to analyze the data in the collected data database. 8. The system of claim 1, further comprising an interface to a remote supervisory center, wherein the self-contained data collection system and the remote supervisory center periodically exchange information relating to emotional response testing. 9. A self-contained computer-implemented data collection method for emotional response testing, comprising: adjusting one or more devices used in an emotional response test, wherein the devices are adjusted according to an environment and a subject of the emotional response test;identifying the subject to determine one or more profiles for the identified subject;selecting one or more tests from an emotional response test database based on the profiles for the identified subject;presenting one or more stimuli to the subject during the selected emotional response test; andreceiving input from the subject during the emotional response test, where... | - | A61B005/00
US20100030061 | A1 | 2010-02-04 | NAVIGATION SYSTEM FOR CARDIAC THERAPIES USING GATING | 2008-07-31 | An image guided navigation system for navigating a region of a patient which is gated using ECG signals to confirm diastole. The navigation system includes an imaging device, a tracking device, a controller, and a display. The imaging device generates images of the region of a patient. The tracking device tracks the location of the instrument in a region of the patient. The controller superimposes an icon representative of the instrument onto the images generated from the imaging device based upon the location of the instrument. The display displays the image with the superimposed instrument. The images and a registration process may be synchronized to a physiological event. | 1. An image guided navigation system for guiding an instrument through a region of a patient, said image guided navigation system comprising: an anatomical gating device operable to sense a first physiological event and a second physiological event, the first physiological event being different from the second physiological event;an imaging device operable to capture image data of the region of the patient in response to the second physiological event by comparing one characteristic of the first physiological event with respect a first threshold and when a second characteristic of said second physiological event to a second threshold;a tracking device operable to track the position of the instrument in the region of the patient;a controller in communication with said anatomical gating device, said imaging device and said tracking device and operable to register said image data with the region of the patient in response to the first and second physiological events, said controller further operable to superimpose an icon representing the instrument onto the image data of the region of the patient based upon the position tracked by said tracking device; anda display operable to display the image data of the region of the patient with the superimposed icon of the instrument. 2. The image guided navigation system for guiding an instrument through a region of a patient as set forth in claim 1, wherein the first physiological event is the generation of an R-wave during a cardiac cycle and the second physiological event is diastole during a cardiac cycle. 3. The image guided navigation system for guiding an instrument through a region of a patient as set forth in claim 2, the anatomical gating device generates a diastolic detection window for confirming whether the patient is in diastole before capturing image data. 4. The image guided navigation system for guiding an instrument through a region of a patient as set forth in claim 3, wherein the anatomical gating device is operable to confirm diastole during the diastolic detection window by comparing at least one characteristic of the ECG signal to one previously stored characteristic being selected from the group consisting of the presence of an R-wave, slew and DC offset. 5. The image guided navigation system for guiding an instrument through a region of a patient as set forth in claim 3, wherein the diastolic detection window extends from about 62% to about 80% of the duration of the R-R interval following the detection of an R-wave. 6. The image guided navigation system for guiding an instrument through a region of a patient as set forth in claim 3, wherein the diastolic detection window is centered at about 70% of the duration of the R-R interval following the detection of an R-wave during normal sinus rhythms. 7. The image guided navigation system for guiding an instrument through a region of a patient as set forth in claim 3, wherein the diastolic detection window is centered at about 45% of... | - | A61B005/055
US20100030480 | A1 | 2010-02-04 | DEVICE AND METHOD FOR EXAMINATION AND EVALUATION OF A BIOLOGICAL ACTIVE AND/OR BIOLOGICAL ACTIVATABLE SUBSTANCE | 2007-05-04 | The invention relates to a device and method for optical examination and for evaluation of a biological active and/or biological activatable substance (2) by use of infra-red, visible or ultra-violet light, at which the examined, respectively evaluated substance is optically stimulated, so that the substance either emits light (in particular fluorescence analysis), or modifies the wavelength of the used light (in particular raman analysis), whereas the device comprises: An electrically driven light source (3),an optical light guide (4), by which the light emitted by the light source (3), is transferable as radiation to the substance being examined (2),an optical or opto-electronical sensor device (5) dedicated to the substance being examined (2), to receive the light emitted or modified by the biological active and/or biological activatable substance,an optical or opto-electronical spectrometer device (6) dedicated to the sensor device (5), by which the light emitted or modified by the substance (2) is measured regarding to light intensity at not less than one frequency, or regarding to frequency shift, whereas a corresponding measurement signal is passed to an analysis circuitry assigned to the spectrometer device (6), andthe device featuring a measurement surface (11), which in shape and dimensions fits to the biological active and/or biological activatable substance (2) being examined. | 1. Device for optical examination and for evaluation of a biological active and/or biological activatable substance (2) by use of infra-red, visible or ultra-violet light, at which the examined, respectively evaluated substance is optically stimulated, so that the substance either emits light (in particular fluorescence analysis), or modifies the wavelength of the used light (in particular raman analysis), whereas the device comprises: an electrically driven light source (3),an optical light guide (4), by which the light emitted by the light source (3), is transferable as radiation to the substance being examined (2),an optical or opto-electronical sensor device (5) dedicated to the substance being examined (2), to receive the light emitted or modified by the biological active and/or biological activatable substance,an optical or opto-electronical spectrometer device (6) dedicated to the sensor device (5), by which the light emitted or modified by the substance (2) is measured regarding to light intensity at not less than one frequency, or regarding to frequency shift, whereas a corresponding measurement signal is passed to an analysis circuitry assigned to the spectrometer device (6), andthe device comprising a measurement surface (11), which in shape and dimensions fits to the biological active and/or biological activatable substance (2) being examined. 2. Device as claimed in claim 1, characterized in that the biological active and/or biological activatable substance (2) is a biological tissue, i.e. a tissue of human, plant or animal origin. 3. Device as claimed in claim 1 or 2, characterized in that at least a portion of the optical light guide (4) and the sensor device (5) are placed in a housing (12), said housing comprising a contact area (13) on or in the upper portion of the housing, with the contact area (13) including the measurement surface (11), or at least partially correlating with the measurement surface (11), on said contact area (13) a part of the body of the substance being examined (2), in particular a hand, a forearm, a leg or a foot is applied, whereas an alignment device (14) is assigned to the contact area (13) on or in the upper portion of the housing for aligning the body part of the test subject with the contact area (13). 4. Device as claimed in claim 1, 2 or 3, characterized in that the optical light guide (4) features a light emission segment (8), which forms, together with the sensor device (5), a structural entity (9), said entity (9) comprising the measurement surface (11) on its upper or outer surface. 5. Device as claimed in one or more of the preceding claims, characterized in that the alignment device (14) comprises a recess (14) or a bulge on or in-the upper portion of the housing (12), the width of said alignment device (14) approximating the width of a middle finger. 6. Device as claimed in one or more of the preceding claims, characterized in that the measurement surface (11) is of almost convex shape... | - | G01N021/64 G01N021/00 G06F019/00 G01J003/00 A61B005/00
US20100016743 | A1 | 2010-01-21 | Identifying Groups of Patients with Similar Physiological Characteristics and Risk Profiles | 2009-07-16 | The invention relates in part to methods for partitioning a plurality of patients into risk profile groups comprising the steps of: recording a physiological signal from each patient of a plurality of patients; segmenting the physiological signal into a plurality of components for each patient of a plurality of patients; grouping the components into a plurality of information classes for each patient of a plurality of patients; assigning a representation to each information class for each patient of a plurality of patients; and grouping the patients in response to the representations of their respective information classes. | 1. A method of partitioning a plurality of patients into risk profile groups comprising the steps of: recording a physiological signal from each patient of a plurality of patients;segmenting the physiological signal into a plurality of components for each patient of a plurality of patients;grouping the components into a plurality of information classes for each patient of a plurality of patients;assigning a representation to each information class for each patient of a plurality of patients; andgrouping the patients in response to the representations of their respective information classes. 2. The method of claim 1 wherein the representation is a numerical value. 3. The method of claim 1 wherein the representation is a symbol. 4. The method of claim 1 wherein the representation is a waveform. 5. The method of claim 4 wherein the waveform is a prototype (archetype) waveform. 6. The method of claim 4 wherein the waveform is a centrotype waveform 7. The method of claim 1 wherein the physiological signal is an ECG and the equivalent time portion is a heartbeat. 8. The method of claim 7 wherein the step of grouping comprises: measuring an electrocardiographic mismatch between each pair of patients in the plurality of patients;assigning each patient of the plurality of patients to a respective cluster of a plurality of clusters;grouping clusters in response to the electrocardiographic mismatch until the number of clusters reaches a predefined minimum; andassigning risk outcomes to each of the clusters. 9. The method of claim 8 wherein the grouping of clusters is performed hierarchically. 10. A method of partitioning a plurality of patients into risk profile groups comprising the steps of: recording an ECG signal from each patient of a plurality of patients;dividing each ECG signal into a plurality of heartbeats for each patient of the plurality of patients;grouping heartbeats into clusters for each patient of the plurality of patients;assigning a representation to each cluster for each patient of the plurality of patients;assigning a value to differences in the representation of clusters for each respective ecg signal for each respective patient of the plurality of patients; andgrouping the patients in response to the values of their respective differences in ECG signals. 11. A method of partitioning a plurality of patients into risk profile groups comprising the steps of: recording an physiological signal from each patient of a plurality of patients;dividing each physiological signal into a plurality of equivalent time portions for each patient of the plurality of patients;grouping time portions into clusters for each patient of the plurality of patients;assigning a representation to each cluster for each patient of the plurality of patients;assigning a value to differences in the representation of clusters for each respective physiological signals signal for each respective patient of the plurality of patients; andgrouping the patients in r... | - | A61B005/0402
US20100023082 | A1 | 2010-01-28 | INDIVIDUALIZED MORPHOLOGY FEATURE EVALUATION AND SELECTION FOR DISCRIMINATION IN IMPLANTABLE MEDICAL DEVICES | 2009-07-22 | An apparatus comprises an implantable sensor, which provides a plurality of physiologic sensor signals of a subject, and a processor. The processor includes a feature module and a detection module. The feature module is configured to identify a feature in the sensor signals and to determine a measure of quality of the feature in the sensor signals. The detection module is configured to perform a morphology analysis of a subsequent portion of at least one of the sensor signals using the feature when the measure of quality of the feature satisfies a quality measure threshold. | 1. An apparatus comprising: an implantable sensor configured to provide a plurality of physiologic sensor signals of a subject;a processor, communicatively coupled to the implantable sensor, wherein the processor includes: a feature module, configured to: identify a feature in at least one of the sensor signals; anddetermine a measure of quality of the feature, wherein the measure of quality includes at least one of: a measure of dispersion of the feature;a measure of regularity of a shape of a sensor signal segment comprising the feature;a number of zero-value crossings in a gradient of the sensor signal segment comprising the feature; ora number of inflection points determined in the feature; anda detection module, configured to perform a morphology analysis of a subsequent portion of at least one of the sensor signals using the feature when the measure of quality of the feature satisfies a quality measure threshold. 2. The apparatus of claim 1, wherein the measure of quality includes the measure of dispersion, and wherein the detection module is configured to select a particular feature for use in the morphology analysis when the measure of dispersion of the feature is less than a dispersion measure threshold. 3. The apparatus of claim 1, wherein the measure of quality includes the measure of regularity of the shape of the sensor signal segment comprising the feature, and wherein the detection module is configured to select a particular feature for use in the morphology analysis when the measure of regularity of the shape of the sensor signal segment comprising the feature exceeds a regularity measure threshold. 4. The apparatus of claim 1, wherein the feature module is configured to identify a peak value of at least one of the sensor signals,wherein the measure of quality includes the number of zero-value crossings in a gradient of the sensor signal segment comprising the feature, the gradient obtained near a time of occurrence of the feature, andwherein the detection module is configured to select a particular feature for use in the morphology analysis when the number of zero-value crossings is less than a threshold number. 5. The apparatus of claim 1, wherein the measure of quality includes the number of inflection points determined in the feature, and wherein the detection module is configured to select a particular feature for use in the morphology analysis when the number of inflection points is less than a threshold number. 6. The apparatus of claim 1, wherein when the feature module provides an indication that the measure of quality ceases to satisfy the quality measure threshold the detection module is configured to perform at least one of: discontinuing using the feature in the morphology analysis; orchanging a morphology threshold used in the morphology analysis. 7. The apparatus of claim 1, wherein the feature module is configured to identify a specified primary feature and a secondary feature different from the primary featu... | - | A61N001/365 A61B005/00
US20100012124 | A1 | 2010-01-21 | REBREATHER RESPIRATORY LOOP FAILURE DETECTOR | 2009-07-08 | A rebreather safety monitoring device comprising a carbon dioxide sensor provided with a gas sampler adapted for sampling a gas in a rebreather breathing loop from a location between an inhale one-way valve on a rebreather mouthpiece and a carbon dioxide scrubber and providing the obtained gas sample to the carbon dioxide sensor, and a means to provide alarms or warnings based on the level of the expired carbon dioxide. The present invention detects a wide range of failures of a rebreather by measurement of the expired carbon dioxide level and application of that level to trigger alarms, provide loop shut-off or provide safety warnings. | 1. A rebreather safety monitoring device comprising: a carbon dioxide sensor provided with a gas sampler, wherein the gas sampler is adapted for sampling a gas in a rebreather breathing loop from a location between an inhale one-way valve on a rebreather mouthpiece and a carbon dioxide scrubber and providing an obtained gas sample to the carbon dioxide sensor; andan interface to one selected from the group consisting of an alarm and a warning, wherein the interface is based on a carbon dioxide reading. 2. A monitor according to claim 1, wherein the carbon dioxide sensor is calibrated using a carbon dioxide level in human expired gas. 3. A monitor according to claim 1, wherein the carbon dioxide sensor is fitted with a hydrophobic membrane. 4. A monitor according to claim 1, wherein the carbon dioxide sensor comprises a dual channel infra-red absorption sensor. 5. A monitor according to claim 1 wherein the carbon dioxide sensor is powered up with a low periodic duty cycle, but where an alarm system is powered even when the carbon dioxide sensor is powered down. 6. A monitor according to claim 1 wherein the carbon dioxide reading is powered up and a reading taken when a change in a user's respiratory rate is detected. 7. A monitor according to claim 1 further comprising a pressure sensor that is applied to compensate the carbon dioxide reading for changes in ambient pressure. 8. A monitor according to claim 1 further comprising a helium sensor that is applied to compensate the carbon dioxide reading for changes in the partial pressure of helium. 9. A monitor according to claim 1, wherein the carbon dioxide sensor is a dual channel infra-red absorption sensor with a circuit to drive an infra-red light source with greater power when under ambient pressure or in a presence of gases having a high thermal capacity, than in air at one atmosphere pressure, thus stabilising a spectrum of an emitted light under a range of pressures or in the presence of helium. 10. A monitor according to claim 1, wherein the carbon dioxide sensor comprises a dual channel infra-red absorption sensor with an infra-red emitter isolated from the effects of pressure and helium such that the emitted spectrum does not change by more than 50% under a range of operating pressures or helium the monitor covers. 11. A monitor according to claim 1, wherein an alarm level is used to drive shut a valve that closes the breathing loop. 12. A monitor according to claim 1, wherein an alarm level is used to drive shut a valve that closes the breathing loop at the mouthpiece and switches a user to an alternative gas source. 13. A monitor according to claim 1, wherein an absence of carbon dioxide, or a presence of carbon dioxide at a partial pressure lower than that normally expired by a human, is applied to trigger one selected from the group consisting of a warning or an alarm level. 14. A monitor according to claim 1, wherein a presence of carbon dioxide at a partial pressure hig... | - | A61M016/20 A61B005/097 G08B017/10 G08B021/00
US20100030111 | A1 | 2010-02-04 | AUTOMATIC MINIATURE INJECTOR AND SAMPLE-TAKER DEVICE FOR MEDICAL USE | 2007-11-22 | An automatic and miniaturized injector and sample-taker device (1) for medical use comprises an enclosure (2) and at least one needle (15) movable relative to said enclosure (2). The injector and sample-taker device includes a detector system (12) suitable for detecting the presence of a blood vessel in a predetermined volume; a control unit (7) suitable for controlling the displacement of the needle as a function of information provided by said detector system (12); and at least one reservoir (13) for containing a substance for injection or to be taken. | 1. An automatic and miniaturized injector and sample-taker device (1; 70) for medical use comprising an enclosure (2; 73) and at least one needle (15; 81) that is movable relative to said enclosure (2; 73), the device being characterized in that it comprises: a detector system (12; 88) suitable for detecting the presence of a blood vessel in a predetermined volume;a control unit (7; 77) suitable for causing the needle to move as a function of the information supplied by said detector system (12; 88); andat least one reservoir (13; 79) for containing a substance to be injected or a sample to be taken. 2. An injector and sample-taker device according to claim 1, characterized in that it includes means (20; 85) for putting the needle into contact with an anesthetic substance. 3. An injector and sample-taker device according to claim 1, characterized in that it includes means (20; 85) for putting the needle into contact with a disinfectant substance. 4. An injector and sample-taker device according to any one of claims 2 to 3, characterized in that the means (20; 85) for putting the needle into contact with one of the substances comprise an element suitable for spraying said anesthetic or disinfectant substance on the needle. 5. An injector and sample-taker device according to claim 4, characterized in that the element suitable for spraying said substance on the needle comprises a micro-electromechanical system. 6. An injector and sample-taker device according to claim 1, characterized in that it includes means (11, 14; 80, 86) suitable for pushing said substance for injection out from the reservoir in order to perform the injection. 7. An injector and sample-taker device according to claim 6, characterized in that the means (11, 14; 80, 86) suitable for pushing said substance for injection out from the reservoir are pneumatic. 8. An injector and sample-taker device according to claim 6, characterized in that the means (11, 14; 80, 86) suitable for pushing said substance for injection out from the reservoir comprise a micro-electromechanical system. 9. An injector and sample-taker device according to claim 1, characterized in that it includes a micro-electromechanical system (11, 16; 86, 82) suitable for moving the needle relative to the enclosure. 10. An injector and sample-taker device according to claim 1, characterized in that it includes means (10; 87) suitable for analyzing a sample. 11. An injector and sample-taker device according to claim 10, characterized in that it includes means (9; 78) suitable for transmitting a signal representative of the analysis to an external appliance. 12. An injector and sample-taker device according to claim 10, characterized in that it includes means (9; 78) suitable for transmitting a signal representative of the analysis to another injector and sample-taker device (1; 70). 13. An injector and sample-taker device according to claim 1, characterized in that it includes signaling means (8; 76) suitab... | Eveon | A61M005/31 G01N001/10 A61B005/15
US07643865 | B2 | 2010-01-05 | Autonomous in-vivo device | 2004-12-30 | The present invention may provide a system, method and device containing a motor. The motor may propel an in-vivo device by rotating a propeller situated for example outside of an outer shell of such device with an electromagnetic field generated by for example a plurality of electrical coils situated for example within such device. | 1. An autonomous in-vivo device, comprising: a capsule;said capsule comprising a rotor, being a propeller, rotatably connected to a shaft;a plurality of electric coils positioned around said rotor;a controller to control current flow through said electric coils causing the rotor to rotate and causing the rotor to tilt towards at least one of said electric coils thereby providing direction and a forward or backward thrust of said device. 2. The device as in claim 1, comprising a duct to channel liquid to said rotor. 3. The device as in claim 1, wherein said rotor includes at least one permanent magnet. 4. The device as in claim 1, wherein said plurality of electric coils are enclosed within an outer shell of said device. 5. The device as in claim 4, wherein said outer shell has a tows shaped end. 6. The device as in claim 1, wherein said plurality of electric coils are mounted on a flexible circuit board. 7. The device as in claim 1, wherein the rotor shaft is attached to a ball and socket joint. 8. A method of propelling an in-vivo device, comprising: providing an autonomous in-vivo capsule imaging device within a gastro-intestinal tract of a subject;generating an electromagnetic field from a plurality of electric coils positioned within said device around a propeller of said device;rotating said propeller in response to changes in said field to provide forward thrust or backward thrust for said device; andtilting said propeller towards at least one of said electric coils in response to changes in said field to provide a direction of said forward or backward thrust. 9. The method as in claim 8, comprising channeling fluids in an in-vivo area to said propeller. 10. The method as in claim 8, wherein said tilting comprises changing strength of current flow through a first coil relative to strength of current flow through a second coil where the first coil is positioned on a first side of said propeller and the second coil is positioned on a second side of said propeller. 11. The method as in claims 10, comprising transmitting a signal from an external operator to said in-vivo device to change said strength of said current in said first coil. 12. The method as in claim 8, comprising changing a direction of a thrust from said propeller. 13. A system comprising: an in-vivo capsule imaging device comprising: a sensor for generating data;a rotor, being a propeller, providing forward thrust or backward thrust when caused to rotate;a plurality of electric coils positioned around said rotor;a controller to control current flow through said plurality of electric coils causing the rotor to rotate and tilt towards at least one of said plurality of electric coils where said tilting provides a direction for a forward or backward thrust;a transceiver for receiving data from the sensor and transmitting said data; andan external transceiver for receiving data from the device and transmitting signals to the device. 14. The system as in claim 13, wherein... | Given Imaging Ltd. | A61B005/05
US07653226 | B2 | 2010-01-26 | Flexible generation of digitally reconstructed radiographs | 2004-09-29 | A system and corresponding method for flexible generation of digitally reconstructed radiograph (DRR) images on a graphics processing unit (GPU) are provided, the system including a processor, an imaging adapter in signal communication with the processor for receiving volumetric data, an integration unit in signal communication with the processor for integrating the volumetric data into incremental line integrals, and a modeling unit in signal communication with the processor for modeling composite line integrals from the incremental line integrals; and the corresponding method including receiving three-dimensional (3D) volumetric data into a graphics pipeline, integrating the 3D volumetric data into incremental line integrals along a plurality of viewing rays, modeling composite line integrals from the incremental line integrals, and adding the incremental line integrals for each composite line integral to form pixels of the DRR image. | 1. A method for a graphics processing unit (GPU) to generate digitally reconstructed radiograph (DRR) images, the method performed by said GPU comprising the steps of: receiving three-dimensional (3D) volumetric data into a graphics pipeline, wherein said 3D volumetric data is stored as a 3D texture;integrating the 3D volumetric data into incremental line integrals along a plurality of viewing rays, wherein values of said incremental line integrals are stored as textures within the GPU;modeling composite line integrals from the incremental line integrals;adding the incremental line integrals for each composite line integral to form pixels of the DRR image; andcorrecting for non-uniform sampling distances due to the projection of a proxy-geometry by means of weighting factors,wherein each incremental line integral along a viewing ray is a weighted sum of rendered geometric elements, which are surface-textured by means of 3D texturing, and wherein line integral fragments are stored as textures within said GPU. 2. A method as defined in claim 1, further comprising: interpolating the texture values of the rendered geometric elements; andconverting the interpolated texture values into total linear attenuation coefficients through post-classification on the GPU. 3. A method as defined in claim 2, further comprising: modeling at least one of Compton or photoelectric interactions by using an effective energy of a radiation source; andconverting the interpolated texture values into total linear attenuation coefficients through at least one of function evaluation, table lookup and texture lookup to generate images reflecting at least one of Compton interaction and photoelectric interaction. 4. A method as defined in claim 1, further comprising at least one of simulation or correction of geometric distortions. 5. A method as defined in claim 4 wherein the correction is computed through at least one of function evaluation or texture lookup. 6. A method as defined in claim 1, further comprising at least one of simulation or correction of intensity distortions. 7. A method as defined in claim 6 wherein the correction is computed through at least one of function evaluation or texture lookup. 8. A method as defined in claim 1, further comprising: deformation of the volumetric dataset to support at least one of deformable registration or deformable visualization of the deformed dataset. 9. A method as defined in claim 8 wherein the deformation is generated through at least one of function evaluation or texture lookup. 10. A method as defined in claim 1, further comprising: off-screen rendering to at least one of textures or a P-buffer to make computed DRR images available for further processing and achieve independence from limitations of display devices. 11. A method as defined in claim 10 wherein the further processing comprises computation of an image gradient. 12. A method as defined in claim 1, further comprising simulating the scattering effect o... | Siemens Medical Solutions USA, Inc. | G06K009/00 G06K009/40 A61B005/00 A61B006/00
US07662101 | B2 | 2010-02-16 | Therapy control based on cardiopulmonary status | 2004-08-31 | Methods and systems provide an approach to therapy control based on assessment of a patient's cardiopulmonary status. Conditions sensed via sensors of an external respiratory therapy device are used to assess a patient's cardiopulmonary status. The respiratory therapy device sensors may be utilized alone or in combination with other sensors to determine cardiopulmonary status of a patient. Therapy delivered to the patient is controlled based on the cardiopulmonary status assessment. For example, therapy delivered to the patient may be initiated, terminated, and/or modified based on the assessed cardiopulmonary status of the patient. Cardiopulmonary status assessment, therapy control, or both, are performed by an implantable device. | 1. A medical therapy control system, comprising: an external respiratory therapy device comprising a first sensor system configured to sense a first set of one or more physiological conditions and a therapy unit configured to provide positive airway pressure to a patient;an implantable cardiac therapy device comprising a second sensor system configured to sense a second set of one or more physiological conditions and a therapy unit configured to deliver electrical stimulation to the patient's heart;a cardiopulmonary disorder processor coupled to the first and the second sensor systems and configured to detect rhythm breathing disorders, including at least apnea, and non-rhythm breathing disorders, including one or more pulmonary vasculature disorders, based on changes in the first set and the second set of sensed physiological conditions, the cardiopulmonary disorder processor configured to identify a cardiopulmonary disorder affecting a patient as a breathing rhythm disorder or a non-rhythm breathing disorder; anda therapy controller coupled to the cardiopulmonary disorder processor and configured to control a therapy delivered to the patient based on the identified cardiopulmonary disorder, wherein at least one of the cardiopulmonary disorder processor and the therapy controller are disposed within the implantable cardiac therapy device;wherein the cardiopulmonary disorder processor is further configured to detect one or more infectious pulmonary disorders based on the first set and the second set of sensed physiological conditions and to identify the cardiopulmonary disorder as the breathing rhythm disorder, the non-rhythm breathing disorder, or an infectious pulmonary disorder;wherein the one or more infectious pulmonary disorders comprise one or more of bronchitis, pneumonia, bronchiolitis, tuberculosis, and bronchiectasis; andwherein the cardiopulmonary disorder processor is further configured to identify the infectious pulmonary disorder as bronchitis, pneumonia, bronchiolitis, tuberculosis, or bronchiectasis. 2. A medical therapy control system, comprising: an external respiratory therapy device comprising a first sensor system configured to sense a first set of one or more physiological conditions and a therapy unit configured to provide positive airway pressure to a patient;an implantable cardiac therapy device comprising a second sensor system configured to sense a second set of one or more physiological conditions and a therapy unit configured to deliver electrical stimulation to the patient's heart;a cardiopulmonary disorder processor coupled to the first and the second sensor systems and configured to detect rhythm breathing disorders, including at least apnea, and non-rhythm breathing disorders, including one or more pulmonary vasculature disorders, based on changes in the first set and the second set of sensed physiological conditions, the cardiopulmonary disorder processor configured to identify a cardiopulmonary disorder affect... | Cardiac Pacemakers, Inc. | A61B005/08
US07650175 | B2 | 2010-01-19 | Method and device for detecting fault in a blood oxygen sensor | 2005-12-22 | A device for measuring blood oxygen content comprises a blood oxygen sensor interface, a signal processing unit, and a plurality of voltage sampling points. A method for detecting fault in a blood oxygen sensor comprises the steps of: connecting the blood oxygen sensor to the signal processing unit; providing voltage signals by the signal processing unit via a pair of testing terminals, for alternately driving two LEDs connected in inverse-parallel with each other in the blood oxygen sensor; sampling the voltages of the pair of testing terminals, respectively, and determining the fault relevant to the LEDs based on the relationship of the voltages; receiving current signals from a photodiode in the blood oxygen sensor by the signal processing unit via another pair of terminals, converting the received current signals into voltage signals, and outputting positive voltage signals via a reverser and a follower, respectively; and sampling the positive voltage signals respectively for determining the fault relevant to the photodiode. Thus, the fault of the sensor can be detected and determined by employing a plurality of sampling points without the tissue to be measured of the human body, thereby the output error due to the faults of the sensor can be avoided, and the device can be maintained conveniently. | 1. A method for detecting fault in a blood oxygen sensor, which is used in a device for measuring blood oxygen content, wherein the method comprises the steps of: a. connecting the blood oxygen sensor to a signal processing unit;b. alternately outputting, from the signal processing unit, voltage signals of reverse polarities via a pair of testing terminals, for alternately driving the two LEDs connected in inverse-parallel with each other in the blood oxygen sensor;c. sampling voltages to ground of the pair of testing terminals respectively and generating a group of comparison voltages V3 and V4 by the device for measuring blood oxygen content each time the signal processing unit outputs a voltage signal of one polarity; andd. determining the faults relevant to the LEDs that occur in the blood oxygen sensor by comparing the voltages V3 and V4. 2. The method for detecting fault in a blood oxygen sensor according to claim 1, wherein in the step d, the method for determining the faults by comparing V3 and V4 comprises: 1) if V3 always equals to V4, a third core is in short-circuit connection with a fourth core, or a first LED or a second LED is broken through;2) if V3 always doesn't equal to V4, and neither of V3 and V4 equals to zero, no fault relevant to the LEDs occurs in the blood oxygen sensor;3) if V3 always equals to zero, the third core is in short-circuit connection with the ground; and4) if V4 always equals to zero, the fourth core is in short-circuit connection with the ground. 3. The method for detecting fault in a blood oxygen sensor according to claim 2, wherein if it is determined in step d that no fault relevant to the LEDs occurs in the blood oxygen sensor, the method further comprises the steps of: e. receiving current signals from a photodiode in the blood oxygen sensor by the signal processing unit via another pair of terminals; converting the current signals flowing through the pair of terminals into voltage signals via two current-voltage converters which are connected to one of the pair of terminals, respectively; then outputting positive voltage signals via a reverser and a follower, respectively;f. sampling the positive voltage signals output from the reverser and the follower, respectively, and generating another group of comparison voltages V1 and V2 by the device for measuring blood oxygen content; andg. determining the faults relevant to the photodiode that occur in the blood oxygen sensor on the basis of the comparison voltages V1 and V2 by the device for measuring blood oxygen content. 4. The method for detecting fault in a blood oxygen sensor according to claim 3, wherein in the step g, the method for determining the faults based on two groups of data of V1 and V2 that are obtained in adjacent two samplings is: 1) if always V1=V2=0, the photodiode and the signal transmitting lines thereof are in an open-circuit or a short-circuit condition, that is, the first core or the second core is in an open-circuit condition,... | Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | A61B005/1455 G01L025/00
US08524154 | B2 | 2013-09-03 | System and method for pseudo-continuous measurement of metabolite concentrations in a mammalian body | 2007-11-08 | A metabolite monitoring system comprising a microdialysis probe including a semi-permeable membrane and a probe flow path passing from an inlet through a sensing volume adjacent to said semi-permeable membrane to an outlet, a fluid delivery device for delivering dialysate to said inlet; and a metabolite monitoring system associated with said outlet for monitoring a concentration of at least one metabolite in said dialysate from said microdialysis probe, wherein said fluid delivery device is configured to deliver a pulsed flow of said dialysate to said inlet. | 1. A metabolite monitoring system comprising: (a) a microdialysis probe including a semi-permeable membrane and a probe flow path passing from an inlet through a sensing volume adjacent to said semi-permeable membrane to an outlet;(b) a fluid delivery device for delivering dialysate to said inlet; and(c) a metabolite monitoring system associated with said outlet for monitoring a concentration of at least one metabolite in said dialysate from said microdialysis probe,wherein said microdialysis probe has an operative volume, and wherein said fluid delivery device is configured to deliver a pulsed flow of said dialysate to said inlet such that each pulse of said dialysate delivered to said inlet has a volume less than said operative volume of said microdialysis probe, wherein said fluid delivery device includes:(a) a dialysate delivery arrangement for generating said pulsed flow of dialysate; and(b) a reagent delivery arrangement for supplying a reagent for mixing with the dialysate from said microdialysis probe,wherein said reagent delivery arrangement and said dialysate delivery arrangement each includes a fluid reservoir acted upon by a piston to deliver the fluid, wherein said reagent delivery arrangement and said dialysate delivery arrangement are mechanically linked so as to maintain a fixed ratio between a quantity of dialysate delivered and a quantity of reagent delivered, and wherein said fluid reservoirs of said reagent delivery arrangement and said dialysate delivery arrangement are substantially concentric fluid reservoirs, said substantially concentric fluid reservoirs being at least partially enclosed between opposing overlapping piston elements such that relative axial displacement of said overlapping piston elements generates delivery of said dialysate and said reagent in said fixed ratio. 2. The metabolite monitoring system of claim 1, wherein said fluid delivery device includes: (a) a pressurized reservoir containing a quantity of said dialysate;(b) at least one valve deployed for selectively opening and closing a supply flow path from said pressurized reservoir to said inlet; and(c) a control system associated with said at least one valve and configured to selectively open said at least one valve to allow pulsed flow of said dialysate to said inlet. 3. The metabolite monitoring system of claim 2, wherein said fluid delivery device further includes: (a) a flow restrictor deployed in said supply flow path; and(b) a pressure measurement arrangement associated with said control system and deployed to measure a pressure drop across at least part of said flow restrictor. 4. The metabolite monitoring system of claim 1, wherein a first of said reagent delivery arrangement and said dialysate delivery arrangement includes a flow channel system having at least one control, and wherein closure of said flow control valve stops fluid delivery from the other of said reagent delivery arrangement and said dialysate delivery arrangement due to s... | G-Sense Ltd. | G01N021/00 F04F001/00 A61M001/14 B67D007/72 G01N031/00
US08370080 | B2 | 2013-02-05 | Methods and systems for determining whether to trigger an alarm | 2008-10-10 | According to embodiment, systems and methods for processing a physiological measurement and generating alarms based on the measurement are provided. Multiple features of a single physiological measurement may be concurrently monitored to generate alarms. One or more of the features may be based on a trend of the physiological measurement. One or more of the features may be based on a wavelet transform of the physiological measurement. Different features may be used in different combinations to lower the percentage of false alarms while still recognizing valid alarm events. | 1. A method for determining whether to trigger an alarm, the method comprising: using a processor for: receiving an SpO2 signal;performing a continuous wavelet transform of the SpO2 signal to generate a scalogram;determining features of the scalogram; anddetermining whether to trigger an alarm based at least in part on the determined features. 2. The method of claim 1, wherein the features comprise features indicative of signal quality. 3. The method of claim 1, wherein the features comprise features indicative of a probe condition. 4. The method of claim 1, wherein the features comprise a moving average of the scalogram and a slope of the scalogram. 5. The method of claim 4, wherein the determining comprises triggering an alarm if the moving average of the scalogram is below a threshold and the slope of the scalogram is non-positive. 6. The method of claim 1, further comprising: determining a characteristic of the SpO2 signal in parallel with the determining features of the scalogram, wherein the determining whether to trigger an alarm comprises determining whether to trigger an alarm based at least in part on the features and the characteristic. 7. The method of claim 6, wherein: the characteristic of the SpO2 signal comprises a moving average of the SpO2 signal; andthe determining whether to trigger an alarm comprises determining whether to trigger an alarm based at least in part on whether the moving average of the SpO2 signal is below a threshold. 8. The method of claim 1, further comprising: generating an alarm based at least in part on the determining whether to trigger an alarm. 9. A system capable of determining whether to trigger an alarm, the system comprising: a sensor capable of generating a photoplethysmograph (PPG) signal;an SpO2 signal generator coupled to the sensor and capable of generating an SpO2 signal based at least in part on the PPG signal;a wavelet transform generator coupled to the SpO2 signal generator and capable of performing a continuous wavelet transform of the SpO2 signal to generate a scalogram;a plurality of feature analyzers, each feature analyzer coupled to the wavelet transform generator and capable of determining a feature of the scalogram and providing an output; andan alarm generator coupled to the outputs of the plurality of feature analyzers and capable of triggering an alarm based at least in part on the outputs. 10. The system of claim 9, wherein the features comprise features indicative of signal quality. 11. The system of claim 9, wherein the features comprise features indicative of a probe condition. 12. The system of claim 9, wherein the plurality of feature analyzers comprises: a first feature analyzer capable of determining a moving average of the scalogram; anda second feature analyzer capable of determining a slope of the scalogram. 13. The system of claim 12, wherein the alarm generator is capable of generating an alarm if the output of the first feature analyzer indicates that the ... | Nellcor Puritan Bennett Ireland | G01N033/49 G06F015/00 G11C017/00
US08790260 | B2 | 2014-07-29 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2009-10-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. In a glucose sensing system comprising a glucose sensor configured for producing an output representative of glucose concentration in a host and processing circuitry configured to process the output, a method for processing data from the output, the method comprising: generating a glucose sensor data point from the glucose sensor output;detecting a signal loss associated with the glucose sensor data point;filtering the glucose sensor data point with a filter whereby a filtered glucose sensor data point is generated; andadjusting, using a processor, the filtered glucose sensor data point to reduce effects associated with the signal loss. 2. The method of claim 1, wherein filtering comprises pass band or stop band filtering. 3. The method of claim 1, wherein filtering comprises recursive filtering. 4. The method of claim 1, wherein filtering comprises infinite impulse response filtering. 5. The method of claim 1, further comprising converting the adjusted sensor data point into a substantially real-time estimated glucose value; and outputting information representative of the substantially real-time estimated glucose value via a user interface. 6. The method of claim 1, further comprising receiving a data stream generated by the continuous glucose sensors continuously during a time period of continuous glucose sensing of the glucose sensor; detecting a signal loss associated with the data stream; and generating an estimated data stream by filtering the data stream continuously in real time and adjusting at least a portion of the filtered data stream associated with the detected signal loss in real time to reduce effects associated with the detected signal loss, wherein the data stream includes the glucose sensor data point. 7. The method of claim 6, further comprising selectively outputting via a user interface glucose concentration information based on the data stream or the estimated data stream responsive to a detection of signal artifacts in the data stream. 8. The method of claim 1, further comprising outputting via a user interface information representative of a substantially real-time glucose concentration of a host, the substantially real-time information based at least in part on the adjusted data point. 9. A glucose sensor system for monitoring a glucose concentration in a host, the system comprising: a continuous glucose sensor configured to produce a data stream; anda processor configured to receive the data stream, wherein the processor is configured to generate an estimated data stream based at least in part on applying the data stream to a recursive filter and adjusting the filtered data stream to reduce effects associated with a signal loss. 10. The system of claim 9, wherein the processor further comprises a pass band filter, and wherein the processor is configured to generate the estimated data stream based additionally on applying the data stream to the pass band filter. 11. The system of claim 10, wherein the proce... | DexCom, Inc. | A61B005/00
US08399021 | B2 | 2013-03-19 | Preparation of poly(ethylene oxide)-coated nanoparticles of maghemite | 2007-02-07 | The invention relates to a material consisting of an aqueous dispersion of PEG-coated nanoparticles of maghemite. The process for preparing said material consists in polymerizing ethylene oxide in the presence of nanoparticles of maghemite in an aqueous dispersion. The polymerization is initiated using reactive groups which are capable of initiating the polymerization of ethylene oxide without the addition of a catalyst, and which are attached to the nanoparticles of maghemite by means of a coupling agent comprising an alkoxysilyl group and one or more of said reactive groups. Applications: contrast agent for MRI. | 1. A method of preparing an aqueous dispersion of maghemite nanoparticles coated with polyethylene oxide (PEO), the method comprising: preparing modified maghemite nanoparticles by modifying maghemite nanoparticles with a coupling agent comprising an alkoxysilyl group and at least one reactive R group, the reactive R group capable of initiating polymerisation of ethylene oxide to polyethylene oxide in the absence of an additional catalyst;adding ethylene oxide to said modified maghemite nanoparticles under anhydrous conditions;polymerizing the ethylene oxide in the presence of said modified maghemite nanoparticles to prepare PEO-coated nanoparticles; anddispersing the PEO-coated nanoparticles in an aqueous solution. 2. The method according to claim 1, wherein preparing modified maghemite nanoparticles comprises: preparing an aqueous dispersion of maghemite nanoparticles; andadding the coupling agent to the aqueous dispersion of maghemite nonoparticles. 3. The method according to claim 1, wherein the at least one reactive R group is a hydroxyl group or a hydroxyl group precursor. 4. The method according to claim 1, wherein the alkoxysilyl group comprises alkyl groups having 1 to 4 carbon atoms. 5. The method according to claim 3, wherein the alkoxysilyl group comprises alkyl groups having 1 to 4 carbon atoms. 6. The method according to claim 1, wherein the alkoxysilyl group is a trialkoxysilyl group. 7. The method according to claim 2, wherein the alkoxysilyl group is a trialkoxysilyl group. 8. The method according to claim 3, wherein the alkoxysilyl group is a trialkoxysilyl group. 9. The method according to claim 1, wherein preparing the modified maghemite nanoparticles comprises: preparing a dispersion of maghemite nanoparticles in an organic solvent; andadding the coupling agent to the organic dispersion of maghemite nanoparticles. 10. The method according to claim 1, comprising: introducing into a reactor the modified maghemite nanoparticles and an apolar anhydrous solvent under an inert atmosphere;progressively adding the ethylene oxide to the reactor under a partial vacuum while agitating, the reactor being maintained at a temperature between −35° C. to −10° C. for a duration of 15 to 30 minutes, to produce a reaction mixture;maintaining the reaction mixture for a duration of between 10 hours and 100 hours at a temperature between 40° C. and 100° C., to prepare PEO-coated nanoparticles; andextracting the PEO-coated nanoparticles by Centrifugation, washing the nanoparticles, and then dispersing the nanoparticles in water. 11. The method according to claim 1, wherein preparing modified maghemite nanoparticles comprises: preparing an aqueous dispersion of maghemite nanoparticles;adding the coupling agent to the aqueous dispersion of maghemite nanoparticles;modifying the maghemite nanoparticles with said coupling agent by hydrolysing the alkoxy groups and the R groups of the coupling agent; andextracting and drying the modified nanopa... | Centre National de la Recherche Scientifique (C.N.R.S.) Universite de Haute Alsace | A61K009/14
US20100030049 | A1 | 2010-02-04 | Subcutaneous Glucose Electrode | 2009-08-31 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably three or four-layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer is overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited (third layer). An outer (fourth) layer is biocompatible. | 1. An electrochemical sensor comprising: one or more non-corroding metal or carbon electrodes;a sensing layer comprising an enzyme coupled to each electrode; anda biocompatible layer comprising a biocompatible hydrogel chemically bound to the sensing layer of each electrode. | Abbott Diabetes Care Inc. | A61B005/1473
US20100020309 | A1 | 2010-01-28 | BLOOD MONITORING SYSTEM | 2009-10-06 | The present invention is directed towards apparatuses and methods for the automated measurement of blood analytes and blood parameters for bedside monitoring of patient blood chemistry. Particularly, the current invention discloses a programmable system that can automatically draw blood samples at a suitable programmable time frequency (or at predetermined timing), can automatically analyze the drawn blood samples and immediately measure and display blood parameters such as glucose levels, hematocrit levels, hemoglobin blood oxygen saturation, blood gases, lactate or any other blood parameter. | 1. A system for periodically monitoring at least one parameter of blood comprising: a fluid channel system connectable to an access device configured to access blood in a patient;at least one pump configured to draw blood from a patient and transport blood through the fluid channel system;a plurality of test substrates contained in a housing;a dispenser system configured to dispense a portion of a blood sample from the fluid channel system to at least one of the plurality of test substrates;a first sensor configured to analyze the blood sample on at least one of the plurality of test substrates and produce a first sensor signal indicative of a first blood parameter;a second sensor positioned near the fluid channel system and configured to provide a second sensor signal corresponding to an optical property of a fluid in the fluid channel system, and indicative of a second blood parameter;wherein information corresponding to the first sensor signal and information corresponding to the second sensor signal are used to provide an improved measurement of a patient fluid parameter; anda controller configured to control the system to automatically periodically draw a blood sample from a patient into the fluid channel system, to dispense a sample from the patient line to at least one of the plurality of test substrates, to sense an optical property with the second sensor, and to determine a measurement of a patient fluid parameter. 2. The system of claim 1 wherein the second sensor comprises an optical sensor. 3. The system of claim 1 wherein the first blood parameter comprises a glucose level. 4. The system of claim 3 wherein the first blood parameter comprises a whole blood glucose level. 5. The system of claim 1 wherein the second blood parameter comprises a blood gas level. 6. The system of claim 1 wherein the second blood parameter comprises a hemoglobin oxygenation level. 7. The system of claim 1 wherein the second blood parameter comprises a hematocrit level. 8. The system of claim 1 wherein the second blood parameter comprises a glucose level. 9. The system of claim 1 wherein the second blood parameter comprises one or more of a blood glucose level, ketone level, hemoglobin level, lactate level, electrolyte level, blood gas level, cholesterol level, or bilirubin level. 10. The system of claim 1 wherein the first blood parameter comprises one or more of a blood glucose level, ketone level, hemoglobin level, lactate level, electrolyte level, blood gas level, cholesterol level, or bilirubin level. 11. A system for periodically monitoring at least one parameter of blood comprising: a fluid channel system connectable to an access device configured to access blood in a patient;at least one pump configured to draw blood from a patient and transport blood through the fluid channel system;a plurality of test substrates contained in a housing;a dispenser system configured to dispense a portion of a blood sample from the fluid channel system to at... | - | G01N033/48 A61B005/155
US20100030116 | A1 | 2010-02-04 | TARGETING DEVICE | 2007-09-19 | A targeting device (20) for use in hip surgery to allow the position of the centreline of the femoral neck to be located, the device comprising: (a) one or more location sensor device (22a, 22b, 23a, 23b), wherein at least part of each location sensor device is movable to a position inside the body that is over the femoral neck, the device or devices being suitable for obtaining data relating to the location of the surface of the femoral neck with respect to a reference position, in both the anterior-posterior and superior-inferior planes; (b) a guide sleeve (32), having a bore capable of receiving a guide wire; and (c) a computer device with software for receiving data from the location sensor device or devices and processing this data so as to determine the location of the centreline of the femoral neck; wherein the guide sleeve can be moved in the anterior, posterior, medial and lateral directions, so that it can be moved so as to be in line with the centreline of the femoral neck as determined by the computer device. | 1. A targeting device for use in hip surgery to allow the position of the centreline of a femoral neck to be located in a body, the device comprising: at least one location sensor device, wherein at least part of each location sensor device is movable to a position inside the body that is over the femoral neck, the at least one sensor device being suitable for obtaining data relating to a location of a surface of the femoral neck with respect to a reference position, in both the anterior-posterior and superior-inferior planes;a guide sleeve, having a bore capable of receiving a guide wire; anda computer device with software for receiving data from the at least one location sensor device and processing the data so as to determine the location of the centreline of the femoral neck;wherein the guide sleeve can be moved in the anterior, posterior, medial and lateral directions, so as to be in line with the centreline of the femoral neck as determined by the computer device. 2. The targeting device of claim 1, wherein the at least one location sensor device is electromechanical. 3. The targeting device of claim 1, wherein the part of each location sensor device that is movable to a position inside the body that is over the femoral neck is positioned inside the body over the femoral neck and radially spaced therefrom. 4. The targeting device of claim 1, wherein the at least one location sensor device comprises at least two location sensor devices, with at least one of the location sensor devices operable to obtain data in the anterior-posterior plane and at least another one of the location sensor devices operable to obtain data in the superior-inferior plane. 5. The targeting device of claim 4, wherein the at least one sensor device comprises at least four location sensor devices. 6. The targeting device of claim 5, wherein at least two of the location sensor devices are operable to obtain data in the anterior-posterior plane and at least two of the location sensor devices are operable to obtain data in the superior-inferior plane. 7. The targeting device of claim 1, wherein the at least one location sensor device comprises at least two sets of location sensor devices, wherein each set includes one or more location sensor devices, wherein at least a first part of each location sensor device is movable to a position inside the body that is over the femoral neck and wherein a second part of each location sensor device is movable so as to be in contact with the femoral neck surface, wherein each location sensor device includes a sensor for obtaining information about the location of the second part that is movable so as to be in contact with the femoral neck surface, such that each set of location sensor devices is suitable for obtaining data relating to the location of the surface of the femoral neck with respect to a reference position, in both the anterior-posterior and superior-inferior planes. 8. The targeting device of claim 7, wherein the s... | - | A61B005/103 A61B017/58
US20100036220 | A1 | 2010-02-11 | Small Volume In Vitro Sensor and Methods of Making | 2009-09-29 | A sensor utilizing a non-leachable or diffusible redox mediator is described. The sensor includes a sample chamber to hold a sample in electrolytic contact with a working electrode, and in at least some instances, the sensor also contains a non-leachable or a diffusible second electron transfer agent. The sensor and/or the methods used produce a sensor signal in response to the analyte that can be distinguished from a background signal caused by the mediator. The invention can be used to determine the concentration of a biomolecule, such as glucose or lactate, in a biological fluid, such as blood or serum, using techniques such as coulometry, amperometry, and potentiometry. An enzyme capable of catalyzing the electrooxidation or electroreduction of the biomolecule is typically provided as a second electron transfer agent. | 1.-11. (canceled) 12. An integrated piercing and analyte testing device, comprising: a housing defining an aperture, wherein the housing comprises:a piercing member adapted to pierce and retract through the aperture and pierce a site on a subject to cause blood to flow therefrom;an analyte sensing member configured to receive a sample of blood from the site on the subject through the aperture and electrochemically determine the concentration of the analyte in the sample; anda display for displaying the concentration of the analyte in the sample. 13. The device of claim 12, wherein the piercing member comprises a releasably engageable lancet. 14. The device of claim 12, wherein the analyte sensor member comprises a releasably engageable electrochemical analyte test strip. 15. The device of claim 12, wherein the device comprises a single disposable cartridge comprising at least one lancet and at least one analyte test strip. 16. The device of claim 12, wherein the device further comprises sample transport member for transporting the sample from the site on the subject through the aperture to the analyte sensing member. 17. The device of claim 16, wherein the transport member comprises a pump or a vacuum-producing element. 18. The device of claim 14, wherein the electrochemical analyte test strip comprises a working electrode and a counter electrode, wherein the working electrode and the counter electrode are separated by a closest distance no greater than 1000 μm. 19. The method of claim 14, wherein the electrochemical analyte test strip comprises an analyte-responsive enzyme and a diffusible redox mediator disposed proximate to the working electrode. 20. The method of claim 19, wherein the analyte-responsive enzyme is glucose dehydrogenase or glucose oxidase. 21. The method of claim 12, wherein the electrochemical analyte test strip is configured to receive no more than about 1 μl of the blood. 22. The method of claim 12, wherein the determining of the concentration of the analyte in the sample comprises determining the concentration of analyte using coulometry. 23. The method of claim 12, wherein the determining of the concentration of the analyte in the sample comprises determining the concentration of analyte using amperometry. 24. The method of claim 12, wherein the determining of the concentration of the analyte in the sample comprises determining the concentration of analyte using potentiometry. 25. The device of claim 34, wherein the analyte is glucose or a ketone. 26. An integrated piercing and glucose testing device, comprising: a housing defining an aperture, wherein the housing comprises:a piercing member comprising a lancet adapted to pierce and retract through the aperture and pierce a site on a subject to cause blood to flow therefrom;an glucose sensing member comprising a electrochemical glucose test strip configured to receive a sample of blood from the site on the subject through the aperture and electrochemically d... | - | A61B005/157
US20100016692 | A1 | 2010-01-21 | SYSTEMS AND METHODS FOR COMPUTING A PHYSIOLOGICAL PARAMETER USING CONTINUOUS WAVELET TRANSFORMS | 2008-10-03 | According to embodiments, systems and methods for computing a physiological parameter are provided. The physiological parameter may be calculated using a continuous wavelet transform technique as well as using a non-continuous wavelet transform technique. More than one value for the physiological parameter may be calculated using various techniques. The values may be evaluated to select a desired value, or an average or weighted average of the values may be computed to generate a desired value. | 1. A method for determining a physiological parameter from a physiological signal, comprising: receiving at least one physiological signal;calculating a first value for a physiological parameter based at least in part on the at least one physiological signal using a continuous wavelet transform technique;calculating a second value for the physiological parameter based at least in part on the at least one physiological signal using a non-continuous wavelet transform technique;analyzing the first value and the second value; anddetermining, based at least in part on the analysis of the first value and the second value, a desired value for the physiological parameter. 2. The method of claim 1 wherein the at least one physiological signal comprises a photoplethysmogram signal. 3. The method of claim 1 wherein analyzing the first value and second value comprises considering at least one of the group consisting of: an expected range of values for the physiological parameter, historical information, patient information, a statistical measure, noise associated with the signal, a confidence indicator. 4. The method of claim 1 further comprising: assigning a first weight to the first value; andassigning a second weight to the second value;wherein determining, based at least in part on the analysis of the first value and the second value, the desired value for the physiological parameter comprises calculating a weighted average of the first value and the second value. 5. The method of claim 4 wherein the first weight and the second weight are based on at least one of the group consisting of: an expected range of values for the physiological parameter, historical information, patient information, a statistical measure, noise associated with the signal, a confidence indicator. 6. The method of claim 1 wherein the physiological parameter comprises one of the group consisting of: blood oxygen saturation, pulse rate, respiration rate, blood pressure, and respiration effort. 7. The method of claim 1 wherein the continuous wavelet transform technique comprises: performing a continuous wavelet transform of the at least one physiological signal;generating at least one scalogram based at least in part on the continuous wavelet transform; andanalyzing features in the at least one scalogram. 8. The method of claim 7 wherein analyzing features in the at least one scalogram comprises a technique selected from the group consisting of: following a ridge, generating a Lissajous figure based on amplitude values of two scalograms, and determining a ratio of an amplitude value of one scalogram to an amplitude value of another scalogram. 9. The method of claim 1 wherein the non-continuous wavelet transform technique comprises a technique selected from the group consisting of: a time domain technique and a spectral technique. 10. A system for determining a physiological parameter from a physiological signal, the system comprising: a sensor configured to generate at least... | Nellcor Puritan Bennett Ireland | A61B005/08 A61B005/1455 A61B005/02 A61B005/021 A61B005/024
US20100042013 | A1 | 2010-02-18 | SYSTEM AND APPARATUS FOR WIRELESS HIGH-FREQUENCY TEMPERATURE ACQUISITION AND ANALYSIS | 2009-08-15 | We disclose a system, an apparatus, and a method for high frequency temperature monitoring and analysis. According to a disclosed embodiment the system comprises: (a) a wireless temperature acquisition and logging device especially designed for multi-day, high-frequency, and high-resolution temperature sampling; and (b) an analysis system implemented in a digital computer with one or more processors in order to analyze and characterize said temperature using a plurality of methods including complexity analysis techniques such as Lempel-Ziv complexity, Approximate Entropy, Sample Entropy, Multiscale Entropy, and Detrended Fluctuation Analysis; and other statistical time-series analysis techniques. According to one embodiment the temperature monitoring system is designed to capture the dynamic aspects of temperature in order to enable researchers and clinicians to study temperature regulation, thermal physiology, and clinical thermometry. | 1. A system for taking temperature measurements and analyzing said measurements, comprising: (a) a wireless temperature acquisition and logging device for acquiring, storing, and wirelessly transmitting said temperature measurements; and(b) an analysis system implemented in a digital computer with one or more processors in order to analyze and characterize said temperature using one or more time-series analysis methods. 2. The system of claim 1, wherein said wireless temperature acquisition and logging device is especially designed for multi-day, high-frequency, and high-resolution temperature measurements. 3. The system of claim 2, wherein said wireless temperature acquisition and logging device comprises one or more temperature sensors, one or more signal conditioners, one or more signal amplifiers, an analog-to-digital converter, a microcontroller, a memory, a wireless communication module, a power circuit, and a power source. 4. The system of claim 3, wherein said memory is substantially equivalent to a flash memory, said wireless communication module is substantially equivalent to Bluetooth, and said power circuit includes a dc-to-dc converter. 5. The system of claim 2, wherein said time-series analysis methods include complexity analysis techniques. 6. The system of claim 5, wherein said complexity analysis techniques include calculating an Approximate Entropy metric of said temperature. 7. The system of claim 5, wherein said complexity analysis techniques include calculating a Sample Entropy metric of said temperature. 8. The system of claim 5, wherein said complexity analysis techniques include calculating a Lempel-Ziv metric of said temperature. 9. The system of claim 5, wherein said complexity analysis techniques include calculating a Detrended Fluctuation metric of said temperature. 10. A system for taking temperature measurements, analyzing said measurements, and characterizing a patient's temperature time-series to help make clinical decisions regarding a patient's clinical condition, comprising: (a) a wireless temperature acquisition and logging device especially designed for multi-day, high-frequency, and high-resolution temperature measurements for acquiring, storing, and wirelessly transmitting said temperature measurements; and(b) an analysis system implemented in a digital computer with one or more processors in order to analyze and characterize said temperature using one or more complexity analysis methods and generate an output temperature complexity metric, said output temperature complexity metric is automatically compared against one or more normality thresholds to help make said clinical decisions regarding said patient's condition. 11. The system of claim 10, wherein said complexity analysis techniques include calculating an Approximate Entropy metric of said temperature. 12. The system of claim 10, wherein said complexity analysis techniques include calculating a Sample Entropy metric of said temperature. 13... | INNOVATEC SL | A61B005/00 G06F015/00 G01K001/00
US20100022894 | A1 | 2010-01-28 | Intravascular Measurement | 2009-07-24 | An implantable sensor for measuring fluidic parameters with a surface acoustic wave transponder for detecting vasomotor quantities and one retaining stent attached to each of two opposite ends of the surface acoustic wave transponder. | 1. An implantable sensor for measuring fluidic parameters comprising a surface acoustic wave transponder for detection of vasomotor quantities and one retaining stent attached to each of two opposite ends of the surface acoustic wave transponder. 2. The implantable sensor according to claim 1, wherein the surface acoustic wave transponder is embodied as a bending beam. 3. The implantable sensor according to claim 1, wherein the surface acoustic wave transponder is provided with a biocompatible coating. 4. The implantable sensor according to claim 3, wherein the biocompatible coating is polyurethane. 5. The implantable sensor according to claim 1, wherein at least one of the retaining stents comprises an antenna or is designed to function as an antenna. 6. The implantable sensor according to claim 5, wherein the antenna is a frame antenna and wherein at least one retaining stent spans the frame antenna on expansion of the at least one retaining stent. 7. The implantable sensor according to claim 5, wherein the at least one retaining stent has two electrically insulated sections designed to function as halves of a half-wave dipole antenna. 8. The implantable sensor according to claim 1, wherein the retaining stents are self-expanding stents. 9. The implantable sensor according to claim 1, wherein the retaining stents are balloon-expanded stents. 10. The implantable sensor according to claim 1, wherein the retaining stents are at least partially made of a biocorrodible material. 11. The implantable sensor according to claim 10, wherein the biocorrodible material is a magnesium alloy. 12. The implantable sensor according to claim 1 which has a marker. 13. The implantable sensor according to claim 1, wherein the surface acoustic wave transponder has one or more reflectors. 14. The implantable sensor according to claim 13, wherein at least one first reflector is arranged on a carrier of the surface acoustic wave transponder with a first distance from a pair of interdigital converters and with a second distance from one end of the surface acoustic wave transponder opposite the pair of interdigital converters, the second distance being greater than the first distance. 15. The implantable sensor according to claim 14, wherein the second distance is at least twice as great, five times as great or ten times as great as the first distance. | BIOTRONIK VI PATENT AG | A61B005/02
US07654955 | B2 | 2010-02-02 | Apparatus and methods for assessing metabolic substrate utilization | 2004-11-19 | Apparatus and methods for assessing metabolic substrate utilization are described. In one embodiment, a processor-readable medium includes code to receive data indicative of a plasma glucose concentration, a plasma free fatty acid concentration, and a respiratory quotient of a subject. The processor-readable medium also includes code to calculate, based on the data, respective values of a set of metabolic parameters. The set of metabolic parameters includes a first metabolic parameter and a second metabolic parameter. The value of the first metabolic parameter is indicative of whether the subject has a predisposition towards oxidation of a first type of metabolic substrate or a second type of metabolic substrate, and the value of the second metabolic parameter is indicative of the subject's responsiveness to a change in availability of the first type of metabolic substrate or the second type of metabolic substrate. | 1. A processor-readable medium, comprising: code to receive data indicative of a blood glucose level, a blood free fatty acid level, and a respiratory quotient of a subject; andcode to calculate, based on the data, a value of a metabolic parameter for the subject, the value of the metabolic parameter being indicative of whether the subject has a predisposition towards one of carbohydrate oxidation and fat oxidation. 2. The processor-readable medium of claim 1, wherein the data includes a first set of measurement results and a second set of measurement results, the first set of measurement results being indicative of the blood glucose level, the blood free fatty acid level, and the respiratory quotient of the subject at a first measurement time, the second set of measurement results being indicative of the blood glucose level, the blood free fatty acid level, and the respiratory quotient of the subject at a second measurement time, the first measurement time being different from the second measurement time. 3. The processor-readable medium of claim 1, wherein the code to calculate the value of the metabolic parameter includes code to perform a regression analysis on the data. 4. The processor-readable medium of claim 1, further comprising: code to compare the value of the metabolic parameter with a baseline value. 5. The processor-readable medium of claim 1, wherein the metabolic parameter is a first metabolic parameter, the processor-readable medium further comprising: code to calculate, based on the data, a value of a second metabolic parameter for the subject, the value of the second metabolic parameter being indicative of the subject's responsiveness to a change in carbohydrate availability. 6. The processor-readable medium of claim 1, wherein the metabolic parameter is a first metabolic parameter, the processor-readable medium further comprising: code to calculate, based on the data, a value of a second metabolic parameter for the subject, the value of the second metabolic parameter being indicative of the subject's responsiveness to a change in fat availability. 7. A computer-implemented method, comprising: applying a set of measurements to a subject to produce a set of measurement results for the subject, the set of measurements being configured to evaluate a blood glucose level, a blood free fatty acid level, and a respiratory quotient of the subject;based on the set of measurement results, using a computer to determine whether the subject has a predisposition towards one of carbohydrate oxidation and fat oxidation; anddisplaying the determination on a user-interface. 8. The method of claim 7, wherein the set of measurements is configured to measure the blood glucose level, the blood free fatty acid level, and the respiratory quotient of the subject after an overnight fast. 9. The method of claim 7, wherein the set of measurements includes a first set of measurements and a second set of measurements, the first set of measurements b... | Entelos, Inc. | A61B005/00
US07660615 | B2 | 2010-02-09 | Insertion set for transcutaneous sensor with connector lock mechanism | 2007-11-09 | An improved insertion set is provided for transcutaneous placement of a sensor such as a glucose sensor at a selected site within the body of a patient. The insertion set comprises a mounting base defining an upwardly open channel for receiving and supporting a flexible thin film sensor, in combination with a cap assembled with said mounting base to capture and retain a proximal end of the sensor within said channel. The sensor further includes a distal segment with sensor electrodes thereon which protrudes from the mounting base for transcutaneous placement, wherein the sensor distal segment is slidably carried by a slotted insertion needle fitted through the assembled base and cap. Placement of the insertion set against the patient's skin causes the insertion needle to pierce the skin to carry the sensor electrodes to the desired subcutaneous site, after which the insertion needle can be slidably withdrawn from the insertion set. The mounting base further includes a fitting and related snap latch members for mated slide-fit releasable coupling of conductive contact pads on a proximal end of the sensor to a cable connector for transmitting sensor signals to a suitable monitoring device. | 1. A connector for connecting a sensor set to a monitor device, wherein the sensor set includes: a mounting base adapted for mounting onto a patient's skin, the mounting base including a connector fitting generally at a rear end of the mounting base, wherein the connector fitting of the sensor set includes a tubular element having a central bore formed therein for pass through reception of a portion of the sensor, and wherein the mounting base further includes including a lock mechanism for releasable coupling to the connector; anda sensor having at least one sensor electrode thereon for generating at least one electrical signal representative of at least one patient condition, the sensor including at least one contact pad conductively coupled to the at least one sensor electrode, the sensor being coupled to the mounting base;wherein the connector is adapted for interengageable coupling with the lock mechanism on the mounting base of the sensor set,wherein the connector includes at least one connector contact that is electrically coupled to the at least one contact pad of the sensor when the mounting base of the sensor set is coupled to the connector,wherein the at least one connector contact receives the at least one electrical signal; andwherein the tubular element of the sensor set is for mating with a corresponding tubular recess formed in the connector, the connector fitting of the sensor set further having a recessed seat to support and retain the sensor, the recessed seat defined to fit within a recess in the connector. 2. The connector of claim 1, wherein the mounting base of the sensor set and the connector have releasably interengageable snap fit latch members that form the lock mechanism. 3. The connector of claim 1, wherein the connector fitting includes a generally D-shaped key formed as a rearward extension of the tubular element, the D-shaped key defining a recessed seat for receiving and supporting the sensor, and wherein the D-shaped key is for mating with a corresponding D-shaped key recess formed in the connector. 4. The connector of claim 3, wherein the recessed seat of the connector fitting is formed at ramp angle relative to a central axis of the bore, and the connector has a corresponding ramp formed in the D-shaped key recess of the connector. 5. The connector of claim 1, wherein the mounting base of the sensor set further includes an upwardly open channel formed therein, and the channel extending from the connector fitting to a front end of the mounting base; wherein the sensor of the sensor set is a flexible sensor having a head formed generally at a proximal end thereof, a proximal segment extending from said head and a distal segment having the at least one sensor electrode thereon, the head including the at least one contact pad conductively coupled to the electrode; andthe sensor set further comprising:a protective cannula having the sensor distal segment and at a portion of said sensor proximal segment received ... | Medtronic MiniMed, Inc. | A61B005/00 H01R004/48 A61M005/178
US07654957 | B2 | 2010-02-02 | Apparatus for physical measurements of the eye | 2006-11-20 | Utilization of a contact device placed on the eye in order to detect physical and chemical parameters of the body as well as the non-invasive delivery of compounds according to these physical and chemical parameters, with signals being transmitted continuously as electromagnetic waves, radio waves, infrared and the like. One of the parameters to be detected includes non-invasive blood analysis utilizing chemical changes and chemical products that are found in the conjunctiva and in the tear film. A transensor mounted in the contact device laying on the cornea or the surface of the eye is capable of evaluating and measuring physical and chemical parameters in the eye including non-invasive blood analysis. The system utilizes eye lid motion and/or closure of the eye lid to activate a microminiature radio frequency sensitive transensor mounted in the contact device. The signal can be communicated by wires or radio telemetered to an externally placed receiver. The signal can then be processed, analyzed and stored. Several parameters can be detected including a complete non-invasive analysis of blood components, measurement of systemic and ocular blood flow, measurement of heart rate and respiratory rate, tracking operations, detection of ovulation, detection of radiation and drug effects, diagnosis of ocular and systemic disorders and the like. | 1. An ophthalmic system comprising: a tonometer for measuring intraocular pressure;at least one of a keratometer for measuring corneal curvature, a biometer for measuring size of the eye and a device for measuring ocular rigidity;an electrical device altering the measured intraocular pressure based upon the at least one of the measured corneal curvature, the measured size of the eye and the measured ocular rigidity to produce a corrected intraocular pressure; anda display identifying the corrected intraocular pressure. 2. The ophthalmic system according to claim 1, further comprising a memory device. 3. The ophthalmic system according to claim 1, wherein the electrical device is a processor. | - | A61B003/16 A61B005/00
US08391953 | B2 | 2013-03-05 | Energy transfer amplification for intrabody devices | 2009-10-06 | Apparatus for driving current in a power circuit of a medical device inserted into a body of a subject includes a power transmitter, which is adapted to generate, in a vicinity of the body, an electromagnetic field having a predetermined frequency capable of inductively driving the current in the power circuit. A passive energy transfer amplifier, having a resonant response at the frequency of the electromagnetic field is placed in proximity to the medical device so as to enhance the current driven in the power circuit by the electromagnetic field. | 1. An invasive medical apparatus, comprising: A catheter, having a distal end, configured to be inserted into a heart of a subject, the catheter comprising a wireless position sensor fixed adjacent to the distal end of the catheter to wirelessly transmit position signals indicative of a position of the catheter within the heart, the wireless position sensor comprising a power coil, a power circuit, and a sensing coil, where the power coil is configured to be driven inductively by a radio-frequency (RF) electromagnetic field to provide operating energy to the wireless position sensor;A signal processing unit external to the subject for determining position and orientation coordinates of the wireless position sensor based on the position signals transmitted by the power coil of the wireless position sensor;A power transmitter configured to generate the RF electromagnetic field in a vicinity of the subject, and dependent upon the spatial position and orientation of the sensing coil, cause currents to flow in the sensing coil, where the power circuit converts the currents into position signals which are transmitted by the power coil to the signal processing unit;A passive energy transfer amplifier configured to be placed on a part of the subject near the heart to enhance inductive driving of the power coil of the wireless position sensor, by amplifying the RF electromagnetic field and causing a greater current flow in the power coil of the wireless position sensor. 2. The apparatus according to claim 1, wherein the passive energy transfer amplifier is adapted to be placed on a chest of the subject, adjacent to the heart. 3. The apparatus according to claim 1, wherein the catheter further comprises one or more electrodes for sensing electrical activity within the heart. | Biosense Webster, Inc. | A61B005/05
US08480684 | B2 | 2013-07-09 | Radiation therapy tattooing system for patient positioning | 2009-03-09 | The present invention provides a three section skin marking system combining target placement, lancing, marking, cleanup and protective bandaging comprising a positioning target, an ink application pad and a bandage pad, wherein the ink application pad is attached to the positioning target and the bandage pad is attached to the positioning target, wherein the positioning target is an appliqué with a top side and a bottom side and a peripheral border of material with a center portion of the appliqué removed and free of material, wherein the bottom side has an adhesive for adhering to skin of a patient; wherein the application pad comprises an absorbent pad surrounding micro-encapsulated or macro-encapsulated ink for marking the patient's skin; and wherein the bandage pad comprises an absorbent pad and an adhesive strip. | 1. A three section skin marking system combining target placement, lancing, marking, cleanup and protective bandaging comprising a positioning target, an ink application pad and a bandage pad, wherein the ink application pad is attached to the positioning target and the bandage pad is attached to the positioning target, wherein the positioning target is an appliqué with a top side and a bottom side and a peripheral border of material with a center portion of the appliqué removed and free of material, wherein the bottom side has an adhesive for adhering to skin of a patient;wherein the application pad comprises an absorbent pad surrounding micro-encapsulated or macro-encapsulated ink for marking the patient's skin; andwherein the bandage pad comprises an absorbent pad and an adhesive strip. 2. The three section skin marking system of claim 1 wherein the ink application pad and the bandage pad are attached with perforations and are detachable. 3. The three section skin marking system of claim 1 wherein each of the three sections is sterilized. 4. The three section skin marking system of claim 1 wherein the bandage pad is an antibiotic coated pad. 5. The three section skin marking system of claim 1 wherein the ink application pad includes ink in the form of at least one selected from the group consisting of thimble, pellet, drop, powder ink, wet blister, dry blister, wet ink pad and invisible ink. | - | A61B006/04
US08521248 | B2 | 2013-08-27 | Fiber-optic probe | 2009-08-19 | In a fiber-optic probe for intravascular measurements, e.g. oxygen saturation measurements, the fiber-optical core has only two fibers. A single fiber core is also possible. A reinforcement fiber improves stiffness, kink resistance and overall strength of the probe. The reinforcement fiber is arranged essentially parallel to the core fibers. The reinforcement fiber may also be wound around the core in a helical manner thus improving the mechanical properties to an even higher degree. The outside of the sheath is coated with an antithrombogenic coating for reducing the danger of clots forming at the surface. The reinforcement fiber may be made of carbon, metal, ceramics or aramide. | 1. A fiber-optic probe for intravascular measurements comprising a fiber-optic probe for intravascular measurements said fiber optic probe comprising core with no more than two optical fibers made of a first material for conducting electro-magnetic radiation from a proximal end of said probe to a distal end of said probe, wherein said fiber-optic probe comprises at least one reinforcing fiber made of a second material different from said first material of which the fiber-optic core is made, said second material selected from the group consisting of carbon, metal, ceramics and aramide, andwherein said fiber-optic probe assembly further comprises an intravascular catheter, said catheter comprising a probe lumen having a proximal opening and a distal opening and being adapted to accommodate said fiber-optic probe,wherein said fiber-optic probe assembly further comprises a sheath made of synthetic sheath material and disposed around said core and said reinforcing fiber so as to be in direct contact with said reinforcing fiber, andwherein the axial length of said probe from said proximal end to said distal end is longer than the length of said probe lumen from said proximal opening to said distal opening. 2. The fiber-optic probe assembly according to claim 1, wherein said fiber-optic core is composed of a single optical fiber. 3. The fiber-optic probe assembly according to claim 1, wherein said fiber-optic core is composed of two optical fibers. 4. The fiber-optic probe assembly according to claim 1, wherein said fiber-optic core is split into two single core fibers with at least one of said at least one reinforcing fibers disposed between said single core fibers. 5. The fiber-optic probe assembly according to claim 1, wherein at least one of said reinforcing fibers is wound around at least part of the core in a helical manner. 6. The fiber-optic probe assembly according to claim 1, further comprising a coupling device for coupling said fiber-optic probe to an electro-magnetic radiation source and to an electromagnetic radiation detector, said coupling device being adapted for guiding efferent electro-magnetic radiation from said electro-magnetic radiation source to said fiberoptic core and for guiding afferent electro-magnetic radiation from said fiber-optic core to said electro-magnetic radiation detector. 7. The fiber-optic probe assembly according to claim 1, further comprising a locking device for locking said fiber-optic probe to a fitting counterpart. 8. The fiber-optic probe assembly according to claim 7, wherein said locking device is permanently fixed to said sheath at a predetermined axial position thereof. 9. The fiber-optic probe assembly according to claim 1, wherein said sheath comprises at least one visible mark indicating an axial position for facilitating defined axial placement. 10. The fiber-optic probe assembly according to claim 1, wherein said sheath is coated with a coating material different from said sheath material... | PULSION Medical Systems SE | A61B005/1455
US08216152 | B2 | 2012-07-10 | Finger arterial elasticity measuring program, finger arterial elasticity measuring device and finger arterial elasticity mesauring method | 2008-02-12 | PROBLEMS TO BE SOLVED: A finger arterial elasticity measuring program, a finger arterial elasticity measuring device and a finger arterial elasticity measuring method are provided for making it possible to measure an elasticity index of a finger artery in accordance with a pulse wave of a finger artery without measuring a blood pressure in evaluating the elasticity of the finger artery related to the degree of arterial sclerosis in an easy and least expensive manner.;MEANS FOR SOLVING THE PROBLEMS: A computer functioning by a finger arterial elasticity measuring program includes a pulse wave data memory unit 42, a cuff pressure data memory unit 43, a normalized pulse volume (NPV) calculating unit 52 for dividing an amplitude (ΔI) of an alternate component of a pulse wave by an average value of direct components (I) to calculate NPV, a relative cuff pressure (Pr) calculating unit 53 for calculating a difference between a cuff pressure at the maximum amplitude of the pulse wave and each cuff pressure as Pr, an FEI calculating unit 54 for calculating a linear regression slope (n) of the following expression (1) as an elasticity index of a finger artery: ln(NPV)=ln(bn)−n·Pr . . . (1) where b is a constant. | 1. A non-transitory, computer-readable medium having program instructions embodied thereon for measuring an elasticity of a finger artery in accordance with a pulse wave of a finger artery, wherein the instructions, when executed by a processor, causes the processor to perform a plurality of steps comprising: storing pulse wave data of a finger artery in a pulse wave data memory unit;storing a cuff pressure of a cuff pressing the finger artery in a cuff pressure data memory unit;calculating a normalized pulse volume (NPV) in an NPV calculating module that acquires an amplitude (ΔI) of an alternating component of a pulse wave and an average value of direct components (I) of the pulse wave from said pulse wave data memory unit and divides said amplitude (ΔI) by the average value of said direct components (I) to calculate NPV;calculating a relative cuff pressure (Pr) in a Pr calculating module that acquires a cuff pressure from said cuff pressure data memory unit and calculates a difference between a cuff pressure at a maximum amplitude of a pulse wave and each cuff pressure as Pr; andcalculating a linear regression slope (n) in a finger elasticity index calculating module using the following expression as an elasticity index of the finger artery (FEI):ln(NPV)=ln(b·n)−n·Pr, where b is a constant. 2. A finger arterial elasticity measuring device which measures an elasticity of a finger artery in accordance with a pulse wave of a finger artery, comprising: a pulse wave data memory unit which stores pulse wave data of a finger artery;a cuff pressure data memory unit which stores a cuff pressure of a cuff pressing the finger artery;a normalized pulse volume (NPV) calculating unit which acquires an amplitude (ΔI) of an alternating component of a pulse wave and an average value of direct components (I) of the pulse wave from said pulse wave data memory unit and divides said amplitude (ΔI) by the average value of said direct components (I) to calculate NPV;a relative cuff pressure (Pr) calculating unit which acquires a cuff pressure from said cuff pressure data memory unit to calculate a difference between a cuff pressure at a maximum amplitude of a pulse wave and each cuff pressure as Pr; anda finger elasticity index calculating unit which calculates a linear regression slope (n) of the following expression as an elasticity index of the finger artery (FEI):ln(NPV)=ln(b·n)−n·Pr, where b is a constant. 3. A method for measuring finger arterial elasticity which measures an elasticity of a finger artery in accordance with a pulse wave of a finger artery, comprising: acquiring pulse wave data of a finger artery from a pulse wave detecting means;storing the pulse wave data in a pulse wave data memory unit;acquiring cuff pressure data from a finger pressing means,storing the cuff pressure data in a cuff pressure data memory unit;calculating a normalized pulse volume (NPV) in a processor by acquiring an amplitude (ΔI) of an alternating component of a pulse wave... | Sapporo Medical University | A61B005/02
US08428674 | B2 | 2013-04-23 | Apparatus for spectrometric based oximetry | 2007-11-14 | A spectrophotometric sensor assembly for non-invasive monitoring of a blood metabolite within a subject's body tissue is provided that includes a pad, a light source, and a light detector The light source is operative to emit light signals of a plurality of different wavelengths. The light detector is operative to detect light emitted by the light source and passed through the subject's body tissue. The light detector is at least partially enclosed in EMI shielding. In some embodiments, the light detector and EMI shielding are disposed in a detector housing that encloses the light detector and shielding. The housing is aligned with a detector aperture disposed in the pad. | 1. A near infrared spectrophotometric sensor assembly for non-invasive monitoring of blood oxygenation levels in a subject's body tissue, said assembly comprising: a pad;at least one light source operative to emit light signals of a plurality of different wavelengths through a source aperture within the pad;at least one light detector operative to detect light emitted by the light source and passed through the subject's body tissue;electromagnetic interference shielding disposed around at least a portion of the light detector;a detector housing including a base and a cap, wherein the base includes a well sized to receive at least a portion of the light detector, and the cap is sized to receive a remainder of the light detector and shielding not received within the base well; andwherein the base and the cap mate to enclose the light detector and shielding, and the housing is positioned relative to the pad so that the light detector is aligned with a light detector aperture disposed within the pad;wherein the base includes a window comprising an optically transparent material that allows light to pass there through and be sensed by the light detector. 2. The sensor assembly of claim 1, wherein the cap includes a cap well sized to receive the remainder of the light detector and electromagnetic shielding not received within the base well. 3. The sensor assembly of claim 2, wherein the base and the cap attach to one another. 4. The sensor assembly of claim 3 wherein the cap is hingedly attached to the base. 5. The sensor assembly of claim 1 wherein the cap is hingedly attached to the base. 6. The sensor assembly of claim 1, wherein the detector housing further includes a port for receiving an electrical conductor. 7. The sensor assembly of claim 1, wherein the base is attached to the pad. 8. The sensor assembly of claim 1, wherein the cap is attached to the pad. 9. The sensor assembly of claim 1, wherein the detector housing has an interior cavity formed by the base and the cap, and wherein the light detector and electromagnetic shielding are positionally fixed within the interior cavity. 10. The sensor assembly of claim 9, wherein there is a press fit between the interior cavity and the light detector and electromagnetic shielding. 11. The sensor assembly of claim 1, wherein the electromagnetic shielding includes a solid shield disposed around at least a portion of the light detector. 12. The sensor assembly of claim 11, wherein the electromagnetic shielding further includes an optically transparent electromagnetic shielding disposed around at least a portion of the light detector. 13. The sensor assembly of claim 1, wherein the electromagnetic shielding further includes an optically transparent electromagnetic shielding disposed around at least a portion of the light detector. 14. The sensor assembly of claim 1, wherein the at least one light detector includes a first light detector and a second light detector, and each of said light de... | CAS Medical Systems, Inc. | A61B005/1455
US08200337 | B2 | 2012-06-12 | Method of selecting specific skeletal muscle highly sensitive to human psychological state, device for selecting specific skeletal muscle, method of evaluating stress during work, and system for evaluating stress during work | 2007-11-13 | The selecting method and apparatus for selecting a specific skeletal muscle affected by a psychological state of a person from skeletal muscles of the person apply at least a first stimulus to the person to bring the person into an uncomfortable state, acquire muscle activity information of each of the skeletal muscles when the first stimulus is applied to the person and select the specific skeletal muscle based on the muscle activity information of each of the skeletal muscles by the first stimulus. The stress evaluating method and system acquire the muscle activity information of the person during an operation for the selected specific skeletal muscle selected by the selecting method and apparatus, and evaluate a degree of stress placed on the person during the operation based on the acquired muscle activity information. | 1. A method of selecting a specific skeletal muscle affected by a psychological state of a person from a plurality of skeletal muscles of the person in order to evaluate a stress on the person during an operation, comprising: a step of applying at least a first stimulus directly to at least one sense organ of the person with an external stimulus applying means to bring the person into an uncomfortable state, the first stimulus being different from a stimulus applied to the person during the operation in which the stress is evaluated;a step of acquiring muscle activity information of each of the plurality of skeletal muscles of the person when the first stimulus is applied to the person; anda step of selecting the specific skeletal muscle based on the muscle activity information of each of the plurality of skeletal muscles of the person when the person is brought into the uncomfortable state by the first stimulus. 2. The method of selecting the specific skeletal muscle according to claim 1, further comprising: a step of acquiring the muscle activity information of each of the plurality of skeletal muscles of the person when the person is in a relaxing state, prior to the step of selecting the specific skeletal muscle,wherein a muscle having a larger amount of change in the muscle activity information when the person is brought into the uncomfortable state by the first stimulus with respect to the muscle activity information in the relaxing state as compared with a set first threshold value is selected as the specific skeletal muscle from the plurality of skeletal muscles of the person in the selecting step. 3. The method of selecting the specific skeletal muscle according to claim 2, wherein: a plurality of stimuli including the first stimulus for bringing the person into the uncomfortable state and a second stimulus for bringing the person into a comfortable state are sequentially applied in the step of applying the stimulus to the person; anda muscle having the larger amount of change in the muscle activity information when the person is brought into the uncomfortable state by the first stimulus with respect to the muscle activity information in the relaxing state as compared with the set first threshold value and a smaller amount of change in the muscle activity information when the person is brought into the comfortable state by the second stimulus with respect to the muscle activity information in the relaxing state as compared with a set second threshold value is selected as the specific skeletal muscle in the selecting step. 4. The method of selecting the specific skeletal muscle according to claim 1, wherein: the muscle activity information of each of the plurality of skeletal muscles of the person is acquired for the each stimulus when the each stimulus is applied to the person, while subjective evaluation information of the person indicating the psychological state of the person at this time is acquired, in the step of acquiring the m... | The Yokohama Rubber Co., Ltd. | A61N001/00
US08181881 | B2 | 2012-05-22 | Stretchable wristband with RFID chip | 2009-08-19 | The present invention is an article of apparel, such as a wristband, that includes an radio frequency identification (RFID) device. The wristband is at least partially formed from an elastomeric material that allows the wristband to stretch while being drawn over the user's hand. The RFID device is positioned on or within the wristband. Once affixed, the RFID device remains an integral part of the wristband. The wristband may be used in any available RFID application, including a contactless payment system. | 1. A wristband comprising: an inner surface, an outer surface, a top edge and a bottom edge, wherein said bottom edge and said top edge are approximately equidistant from one another along the circumference of said wristband;a groove positioned in said inner surface;an RFID chip and antenna positioned in said groove; andan overmold portion molded into said groove and over said RFID chip and antenna such that said inner surface and said outer surface are approximately equidistant from one another along said circumference of said wristband. 2. The wristband of claim 1, wherein said wristband is made of one or more of thermoplastic polyurethane, styrene-ethylene/butylene-styrene copolymer, copolyester, copolyamide, thermoplastic rubber, thermoplastic vulcanite, acrylonitrile butadiene styrene, polycarbonate, and nylon. 3. The wristband of claim 1 wherein the depth of said groove is approximately one-half of the distance between said inner surface and said outer surface. 4. The wristband of claim 1 wherein the circumference of said wristband is between about 18 cm and 22 cm. 5. The wristband of claim 1 wherein said wristband is not made of silicone. 6. The wristband of claim 1 wherein said overmold portion is made of polypropylene. 7. The wristband of claim 1 wherein the RFID chip is ISO/IEC 14443 compliant. 8. The wristband of claim 1 wherein said wristband is circular and does not require snaps, hooks, velcro, or other fasteners to be worn on the wrist. 9. A wristband comprising: an inner surface, an outer surface, a top edge and a bottom edge, wherein said bottom edge and said top edge are approximately equidistant from one another along approximately 75% of the circumference of said wristband, and wherein one or more scallops are located adjacent to said the end of said approximately 75% of the circumference of said wristband;an RFID chip and antenna affixed to the outer or inner surface of said wristband;an overlay material affixed over said RFID chip and said antenna such that said RFID chip and said antenna are secured to said wristband. 10. The wristband of claim 9, wherein said wristband is made of one or more of thermoplastic polyurethane, styrene-ethylene/butylene-styrene copolymer, copolyester, copolyamide, thermoplastic rubber, thermoplastic vulcanite, acrylonitrile butadiene styrene, polycarbonate, and nylon. 11. The wristband of claim 9 wherein the circumference of said wristband is between about 18 cm and 22 cm. 12. The wristband of claim 9 wherein said wristband is not made of silicone. 13. The wristband of claim 9 wherein said overlay material is made of polypropylene. 14. The wristband of claim 9 wherein the RFID chip is ISO/IEC 14443 compliant. 15. A wristband comprising: an inner surface, an outer surface, a top edge and a bottom edge, wherein said bottom edge and said top edge are approximately equidistant from one another along approximately 75% of the circumference of said wristband, and wherein said wristband is... |  | G06K019/06
US09538919 | B2 | 2017-01-10 | System and method for improved ischemia and acute myocardial infarction detection | 2008-07-09 | A system and method are provided for the detection of a heart-related condition by obtaining information in real-time when a condition is initially identified as potentially occurring. A physical exercise and recovery episode is initially detected from physiological signals sensed in a patient. Once detected, a HR-ST segment deviation hysteresis analysis is performed in an implantable medical device (IMD) from certain physiological signals over portions of the exercise and recovery episode to identify the probability that a certain condition is occurring. Once a desired level of probability that the heart-related condition has been detected exists, data utilized in the analysis can be transmitted remotely for clinical review and confirmation of the device's detection of the condition. The patient may be prompted to answer questions related to symptoms that patient is experiencing through an input device in order further confirm the probability that the condition is occurring in the patient. | 1. A method comprising: detecting an exercise and recovery episode in a patient using an implantable medical device (IMD);performing an analysis of heart rate (HR)-ST segment hysteresis over the exercise and recovery episode using the IMD; anddetermining whether the patient is experiencing a particular heart-related condition based on the results of the HR-ST segment hysteresis analysis. 2. The method of claim 1, wherein the heart-related condition includes at least one of an ischemia, unstable angina and an acute myocardial infarction, the method further comprising determining whether the patient is experiencing a particular heart-related condition by computing a probability that the particular heart-related condition is occurring and determining whether the computed probability exceeds a predefined detection threshold. 3. The method of claim 1, wherein the HR-ST segment hysteresis analysis includes at least one of (i) a detailed analysis of an initial recovery period after exercise ceases, (ii) a determination of a prevailing direction of HR-ST hysteresis during the exercise and recovery episode, and (iii) a determination of an average magnitude of HR-ST hysteresis during the exercise and recovery episode. 4. The method of claim 1, further comprising collecting additional information from a patient relating to symptoms being experienced by the patient for use in confirming the detection of the particular heart-related condition. 5. The method of claim 4, wherein the additional information is collected by querying the patient with symptom-related questions after a certain level of probability that the patient is experiencing a particular heart-related condition is reached from the HR-ST segment hysteresis analysis. 6. The method of claim 1, wherein upon a determination that the patient is experiencing a particular heart-related condition based on the results of the HR-ST segment hysteresis analysis, the method further comprising comparing data used in making such determination against at least one of historical data for the patient and baseline data for confirming the detection of the particular heart-related condition. 7. The method of claim 1, further comprising detecting the exercise and recovery episode by sensing at least one physiological signal and deriving a corresponding physiological parameter from each sensed signal. 8. The method of claim 7, wherein the physiological signals comprise at least one of a cardiac electrical signal and a respiratory signal. 9. The method of claim 7, wherein the at least one physiological signal comprises a patient activity signal generated by a sensor. 10. The method of claim 9, wherein the sensor comprises an accelerometer. 11. The method of claim 1, further comprising sensing an electrocardiogram (ECG) signal for deriving Q-wave and T-wave information from the ECG signal for use in confirming the detection of the particular heart-related condition. 12. The method of claim 1, further comprisin... | Medtronic, Inc. | A61B005/04 A61B005/00 A61B005/0452 A61B005/053 A61N001/362 A61N001/368
US20100016731 | A1 | 2010-01-21 | Hemodynamic Detection of Circulatory Anomalies | 2009-04-06 | The preferred indicator is an injection of indocyanine green dye which is detected and quantified by causing it to fluoresce at a sensor location, for example, at the human ear. Quantification is carried out by using cardiac output procedures and where called for, the Valsalva Maneuver is monitored at a monitor/controller giving visual cues to the patient and operator as to the quality of the necessary exhalation pressure maintenance. | 1. A method for detecting the presence of a right-to-left pulmonary shunt in a patient, the improvement comprising the steps: providing an elongate air conduit having proximal end configured to be positioned within the mouth of the patient to receive exhaled breath, and having a distal end configured for air pressure secure attachment with a pressure transducer;providing an indicator delivery system having an outlet located in a vein of the patient in blood flow communication with the right side of the heart of the patient and actuateable by an operator to inject a predetermined amount of metabolically compatible indicator into the vein for travel to the right side of the heart of the patient;providing an indicator sensor located with respect to arterial vasculature and having an output corresponding with the instantaneous concentration of indicator at such vasculature; andproviding a monitor/controller having a display and supporting the pressure transducer in attachment with the air conduit distal end and having a pressure output signal, responsive to such pressure output signal reaching and at least maintaining a predetermined corresponding pressure threshold value to provide a visual representation thereof at the display and commence a test timing, responsive to the test timing to cue the operator to actuate the indicator delivery system, responsive to the sensor output to determine whether a shunt is present. 2. The method of claim 1 in which: the elongate air conduit is provided as a flexible polymeric tube promoting a buffering of air pressure variations therein. 3. The method of claim 1 in which: the monitor/controller is provided with a display with a visible dynamic component response to the pressure output signal to derive a visible cue relating such pressure output with the threshold value. 4. The method of claim 1 in which: the monitor/controller is provided with a visible dynamic component present as a dynamic bar graph. 5. The method of claim 1 in which: the monitor/controller is provided being responsive to a pressure output signal representing a drop in air pressure dropping below the pressure threshold value to create an aural warning. 6. The method of claim 1 in which: the monitor/controller is provided to respond to the maintenance of the pressure output signal at least representing the pressure threshold value and to the commencement of test timing to provide an audible cue to the operator to actuate the indicator delivery system. 7. The method of claim 6 in which: the indicator delivery system is provided to include a flow sensor responsive to identify the commencement and termination of fluid flow through the system; andthe monitor/controller is provided to respond to said commencement and termination of fluid flow and derive an audible alarm to the operator where the time interval of indicator injection is excessive. 8. The method of claim 1 in which: the monitor/controller is provided as being responsive to the pre... | Cardiox Corporation | A61B005/05 A61B005/097 A61B005/02 A61B006/00
US20100036267 | A1 | 2010-02-11 | ANALYSIS DEVICE AND ANALYZING METHOD FOR AUTONOMIC NERVE STATE | 2009-08-06 | A detecting device for detecting an autonomic nervous state of an examinee includes a measuring unit, an analog-to-digital converting (ADC) unit and a digital processing unit. The measuring unit measures a vital sign of the examinee to output a vital sign signal. The ADC unit is coupled to the measuring unit for receiving and converting the vital sign signal from analog to digital to output a digitalized vital sign signal. The digital processing unit is coupled to the ADC unit for receiving and processing the digitalized vital sign signal to obtain a detecting result. The detecting result represents the autonomic nervous state and organic states of the examinee. | 1. A detecting device for detecting an autonomic nervous state of an examinee, comprising: a measuring unit measuring a vital sign of the examinee to output a vital sign signal;an analog-to-digital converting (ADC) unit coupled to the measuring unit for receiving the vital sign signal, and converting the vital sign from analog to digital to output a digitalized vital sign signal; anda digital processing unit coupled to the ADC unit for receiving the digitalized vital sign signal and processing the digitalized vital sign signal to obtain a detecting result, the detecting result representing the autonomic nervous state and organic states of the examinee. 2. The detecting device according to claim 1, wherein the digital processing unit performs a frequency calculation and a frequency domain transform on the digitalized vital sign signal to obtain a vital sign time-frequency spectrum as the detecting result, and in the vital sign time-frequency spectrum, a sympathetic nervous state and a parasympathetic nervous state of an autonomic nerve are corresponding to a very low frequency and a low frequency, respectively. 3. The detecting device according to claim 2, wherein the digital processing unit performs a short-time Fourier transform (STFT) or a Hilbert-Huang transform (HHT) on the digitalized vital sign signal to obtain the vital sign time-frequency spectrum. 4. The detecting device according to claim 1, wherein the digital processing unit performs a frequency calculation and an entropy analysis on the digitalized vital sign signal to obtain an entropy analyzing result as the detecting result, and in the entropy analyzing result, a sympathetic nervous state and a parasympathetic nervous state of an autonomic nerve are corresponding to a very low frequency and a low frequency, respectively. 5. The detecting device according to claim 4, wherein after the digital processing unit performs the entropy analysis, a mean square error calculation is carried out to obtain the entropy analyzing result. 6. The detecting device according to claim 1, further comprising: a recording unit storing a plurality of reference results,wherein the digital processing unit is coupled to the recording unit for reading the reference results, and compares the detecting result and the reference results to determine whether the autonomic nervous state is abnormal or improved. 7. The detecting device according to claim 1, further comprising: a body having a concave, wherein the measuring unit is disposed in the concave, and when a finger of the examinee places into the concave and contacts with the measuring unit, the measuring unit measures the pulse of the examinee to output the vital sign signal,wherein the measuring unit is an infrared measurement device. 8. The detecting device according to claim 1, further comprising: a heating unit disposed in the body to warm up the finger of the examinee. 9. An analysis device for analyzing an autonomic nervous state used with a ... | - | A61B005/02
US20100030120 | A1 | 2010-02-04 | Vibrotactile devices for controlled somato sensory stimulus during functionalmagnetic resonance imaging | 2009-10-05 | A functional magnetic resonance imaging (fMRI) compatible magnetomechanical vibrotactile device (MVD) uses wire coils having small oscillatory currents to interact with the large static magnetic field inherent to MRI scanners. The resulting Lorentz forces which are exerted on the MVD can be oriented to generate large vibrations that may be easily converted to translational motions as large as several centimetres. Representative data demonstrate the flexibility of MVDs to generate different well-controlled vibratory and tactile stimuli to activate special proprioceptive and cutaneous somatosensory afferent pathways. | 1.-27. (canceled) 28. A method of providing somatosensory stimuli to a subject when performing functional magnetic resonance imaging, said method comprising the steps of: (a) providing a magnetomechanical vibrotactile device comprising a coil of wire;(b) placing said device adjacent to an organ of a subject, wherein said organ is positioned within a magnetic field of a magnetic resonance imager;(c) applying a time-varying current to said coil, wherein a force caused by an interaction of said magnetic field with said time-varying current produces vibratory motion of said device;(d) coupling said vibratory motion to said organ to stimulate somatosensory receptors of said subject. 29. The method of claim 28, further comprising varying at least one of a temporal, frequency, and amplitude characteristics of said current. 30. The method of claim 28, wherein said characteristic of said current is varied to preferentially stimulate a specific class of somatosensory receptor. 31.-35. (canceled) 36. The method of claim 28, wherein said device further comprises a former, and wherein said coil is wound around said former. 37. The method according to claim 36 wherein said former is substantially cylindrical. 38. The method according to claim 36 wherein said former is substantially spherical. 39. The method according to claim 38 wherein said device further comprises two additional coils wound around said former, said first coil and said two additional coils being oriented in three orthogonal directions. 40. The method of claim 39, further comprising varying a current provided to at least one of the said three wire coils to achieve vibratory motion in a selected orientation. 41. The method of claim 28, further comprising indirectly coupling said vibratory motion to said organ using a coupling means, said coupling means comprising a proximal end and a distal end, wherein said proximal end is connected to said device, and said distal end comprises a contacting member adapted to couple said vibratory motion to said organ. 42. The method of claim 41, further comprising adapting said coupling means to convert rotational vibratory motion of said device into translational vibratory motion of said contacting member. 43. The method of claim 41, further comprising adapting said coupling means to convert rotational vibratory motion of said device into tapping vibratory motion of said contacting member. 44. The method of claim 41, wherein said contacting member further comprises a textured surface. 45. The method of claim 41, wherein said contacting member further comprises a brush. 46. The method of claim 28, wherein said organ is selected from the list comprising skin, bones, tendons, and eyes. 47. The method of claim 28 wherein said step of placing said device adjacent to an organ of said subject comprises integrating said device with an apparel article worn by said subject. 48. The method of claim 47 wherein said apparel article is a glove. 49. The me... | - | A61H001/00 A61B005/05
US20100022874 | A1 | 2010-01-28 | Image Guided Navigation System and Method Thereof | 2009-07-23 | An image guided navigation system comprises a memory, a locator, a processor and a display. The memory stores a plurality of CT images and a software program. The locator is capable of indicating a direction to a surgical area, and the indicated direction of the locator is defined as a first direction. The processor is electrically connected to the memory and the locator. At least one corresponding image corresponding to the first direction is obtained from the plurality of CT images by the processor executing the software program. The at least one corresponding image comprises at least one simulated fluoroscopic image. The display is capable of showing the at least one corresponding image. | 1. An image guided navigation system for a surgery, wherein a plurality of CT images of a surgical area of a patient is obtained before the surgery, the image guided navigation system comprising: a memory for storing the plurality of CT images and a software program;a locator for indicating a direction to the surgical area, wherein the indicated direction of the locator is defined as a first direction;a processor electrically connected to the memory and the locator, wherein at least one corresponding image corresponding to the first direction is obtained from the plurality of CT images by the processor executing the software program, the at least one corresponding image comprising at least one simulated fluoroscopic image; anda display capable of showing the at least one corresponding image. 2. The image guided navigation system as claimed in claim 1, wherein the at least one corresponding image is changed by adjusting the first direction indicated by the locator to the surgical area. 3. The image guided navigation system as claimed in claim 1, wherein the at least one simulated fluoroscopic image comprises a viewing fluoroscopic image; the viewing fluoroscopic image is on a plane substantially vertical to the first direction. 4. The image guided navigation system as claimed in claim 3, wherein the at least one corresponding image further comprises at least one multiplanar reconstruction (MPR) image; the at least one MPR image is on a plane along the first direction, and a normal line of the plane is substantially vertical to the first direction. 5. The image guided navigation system as claimed in claim 4, wherein the at least one MPR image comprises a transverse section image. 6. The image guided navigation system as claimed in claim 5, wherein the at least one MPR image further comprises a longitudinal section image; the longitudinal section image is substantially orthogonal to the transverse section image. 7. The image guided navigation system as claimed in claim 6, wherein the at least one simulated fluoroscopic image further comprises a lateral fluoroscopic image; the lateral fluoroscopic image is obtained by taking a simulated X-ray photograph from a lateral side of the patient according to a designated position indicated by the first direction. 8. The image guided navigation system as claimed in claim 7, wherein the viewing fluoroscopic image, the lateral fluoroscopic image, the transverse section image, and the longitudinal section image are simultaneously and respectively shown on four different parts of the display. 9. The image guided navigation system as claimed in claim 7, wherein the viewing fluoroscopic image, the lateral fluoroscopic image, the transverse section image, and the longitudinal section image are switchable and selectively shown on the display. 10. The image guided navigation system as claimed in claim 1, wherein the display can simultaneously show a plurality of the at least one corresponding images. 11. The... | - | A61B005/05 A61B006/00
US20100022907 | A1 | 2010-01-28 | Methods based on fluctuations in cortical synchronization | 2009-05-14 | A method of identifying a site of injury in a pediatric patient with brain injury is provided which comprises obtaining an electroencephalogram signal on the patient. The identification of phase synchrony within the signal is indicative of the site of brain injury. Electroencephalogram signals may also be used to determine prognosis of a pediatric patient in a coma in which an increase in the temporal variability of phase synchronized EEG signals over time is indicative of an improvement in the patient. | 1. A method of identifying a site of brain injury in a pediatric patient comprising the steps of: a) obtaining an electroencephalogram signal on the patient; andb) evaluating phase synchronization of the signal obtained, wherein a determination of synchrony in the signal of at least 2 standard deviations greater than the mean synchrony of a control is indicative of a site of brain injury. 2. The method as defined in claim 1, wherein the electroencephalogram signal is obtained over a time period of about 15-45 minutes. 3. The method as defined in claim 2, wherein the time period is about 25-35 minutes. 4. The method as defined in claim 1, wherein the electroencephalogram signal is recorded at an acquisition rate of at least about 150 Hz. 5. The method as defined in claim 1, wherein phase synchronization is evaluated in a frequency range of about 1 to 30 Hz. 6. The method as defined in claim 5, wherein the frequency range is about 1 to 5 Hz. 7. A method of prognosis in a pediatric patient in a coma comprising the steps of: a) obtaining a first electroencephalogram signal on the patient subsequent to the onset of coma;b) obtaining a second electroencephalogram signal on the patient at a time subsequent to obtaining the first signal;c) evaluating phase synchronization patterns in the first and second signals, or in a portion of the first and second signals; and.d) calculating the temporal variability of the synchronized patterns, wherein increased variability in the signal over time is indicative of an improvement in the patient. 8. The method as defined in claim 7, wherein the first electroencephalogram signal is obtained at a time following onset of coma of up to about 5 days. 9. The method as defined in claim 7, wherein the second electroencephalogram signal is obtained at a time subsequent to the first electroencephalogram signal of up to about 7 days. 10. The method as defined in claim 7, wherein the first electroencephalogram signal is obtained within about 24-48 hours following onset of coma, and the second electroencephalogram signal is obtained within about 48-72 hours following onset of coma. 11. The method as defined in claim 7, wherein an increase in the temporal variability of the synchronized patterns in the range of about 10%-100% over time is indicative of an improvement in the patient. 12. The method as defined in claim 7, wherein the first and second electroencephalogram signals are each obtained over a time period of about 15-45 minutes. 13. The method as defined in claim 12, wherein the time period is about 25-35 minutes. 14. The method as defined in claim 7, wherein the first and second electroencephalogram signals are each recorded at an acquisition rate of at least about 150 Hz. 15. The method as defined in claim 7, wherein phase synchronization is evaluated at a frequency range of about 1 to 30 Hz. 16. The method as defined in claim 15, wherein the frequency range is about 1 to 5 Hz. 17. A method of monitoring ... | The Hospital for Sick Children | A61B005/0476
US20100030117 | A1 | 2010-02-04 | SYSTEMS AND METHODS FOR MONITORING ERECTILE FUNCTION AND DIAGNOSING ERECTILE DYSFUNCTION | 2008-02-02 | Methods and systems for monitoring erectile function or dysfunction provide for determining concentration data of biochemical compounds in a subject's erectile tissues; and analyzing the concentration data to yield values indicative of the subject's erectile function. The concentration of biochemical compounds may be measured using Near Infrared Spectroscopy (NIRS). The biochemical compounds may comprise at least one compound from the group consisting of Hemoglobin, Oxygenated Hemoglobin, Cytochromes and Myoglobin. | 1. A method for monitoring erectile function in a mammal, the method comprising: non-invasively monitoring a concentration of one or more biochemical compounds in erectile tissues of the mammal; and,detecting an erection of the erectile tissues based on changes in the monitored concentration. 2. A method according to claim 1 wherein the one or more biochemical compounds are selected from the group consisting of hemoglobin, oxygenated hemoglobin, cytochromes and myoglobin. 3. A method according to claim 1 wherein non-invasively monitoring comprises transmitting infrared light into the erectile tissues, receiving infrared light that has passed through the erectile tissues, and determining the concentration of the one or more biochemical compounds by performing near-infrared spectrometry using the received infrared light. 4. A method according to claim 3 wherein the infrared light comprises light in the wavelength range of 700 nm to 950 nm. 5. A method according to claim 3 wherein the erectile tissues comprise erectile tissues of a penis of the mammal, transmitting infrared light comprises holding an infrared light source against skin of the penis and receiving infrared light comprises holding an infrared light detector against the skin of the penis. 6. A method according to claim 5 wherein the received infrared light comprises light that has been back-scattered within the penis. 7. A method according to claim 5 wherein the received infrared light comprises light that has been transmitted through the penis. 8. A method according to claim 1 further comprising computing a first derivative of the concentration as a function of time. 9. A method according to claim 1 wherein the biochemical compound comprises oxygenated hemoglobin and the method comprises detecting an onset of an erection by identifying a pattern comprising a decrease in the concentration of the oxygenated hemoglobin from a baseline level followed by an increase in the concentration of the oxygenated hemoglobin to above the baseline level. 10. A method according to claim 1 wherein the biochemical compound comprises oxygenated hemoglobin and the method comprises establishing a baseline level for the concentration of the oxygenated hemoglobin in the absence of an erection. 11. A method according to claim 10 comprising detecting an onset of an erection based on changes in the concentration of oxygenated hemoglobin and detecting a peak concentration of the oxygenated hemoglobin during the erection. 12. A method according to claim 11 comprising measuring a first half-time between the onset of the erection and a time corresponding to the peak concentration of oxygenated hemoglobin. 13. A method according to claim 11 comprising detecting an end time of the erection based on changes in the concentration of oxygenated hemoglobin subsequent to the onset and measuring a second half-time between a time corresponding to the peak concentration of oxygenated hemoglobin and the end time of t... | Urodynamix Technologies Ltd. | A61B005/103
US20100049016 | A1 | 2010-02-25 | Noninvasive Transdermal Systems for Detecting an Analyte in a Biological Fluid and Methods | 2009-07-08 | The present invention relates to noninvasive transdermal systems comprised of a noninvasive transdermal patch and a reflectometer. The noninvasive transdermal patches are comprised of a wet chemistry component and a dry chemistry component. The wet chemistry component is a liquid transfer medium in the form of a gel layer for the extraction and liquid bridge transfer of the analyte of interest from the biological fluid within or beneath the skin to the dry chemistry component. The dry chemistry component is a reagent system for interacting with the analyte of interest (glucose) to generate a color change. The reflectometers include a modulated light source for emitting light to illuminate a target surface which possesses a certain color and shade of color for detection by an optical detector. The output signal is processed for determining a corresponding quantity of quality measurement. | 1. A noninvasive transdermal system for detecting an analyte in interstitial fluid extracted from or underneath the skin of a subject, said noninvasive transdermal system comprising: (a) a noninvasive transdermal patch comprising a target surface having a dry chemistry component for interacting with the analyte to generate color or shade of color at said target surface, said dry chemistry component having a sensitivity which enables it to detect the analyte extracted from interstitial fluid, anda wet chemistry component for transferring the analyte from the interstitial fluid in or underneath the skin to said dry chemistry component in an amount sufficient so that said dry chemistry component can detect the analyte; and(b) a reflectometer comprisinga modulated light source for emitting light to illuminate said target surface which possesses a certain color and shade of color following interaction with the analyte;an optical detector for detecting light that is reflected from said target surface and generating a first output indicative of detected light;means for processing the first output to generate a feedback signal for application to the optical detector to compensate for any shift in the first output resulting from the detection of ambient light by the optical detector, and differentially amplify the first output to generate a second output; anda detector for synchronously demodulating the second output to generate a substantially steady DC output voltage that is indicative of the color or shade of color at said target surface. 2. A noninvasive transdermal system for detecting an analyte in a biological fluid extracted from or underneath the skin of a subject, said noninvasive transdermal system comprising: (a) a noninvasive transdermal patch comprising a target surface having a dry chemistry component for interacting with the analyte to generate color or shade of color at said target surface, a dry chemistry component for interacting with the analyte to detect the analyte, said dry chemistry component having a sensitivity which enables it to detect the analyte extracted from interstitial fluid, and a wet chemistry component for transferring in about 15 minutes or less the analyte from the interstitial fluid in or underneath the skin to said dry chemistry component in an amount sufficient, so that said dry chemistry component can interact with the analyte to generate color or shade of color at said target surface for detecting the analyte, said wet chemistry component consisting essentially of a gel and a skin permeate enhancer;(b) a reflectometer comprising a light source for emitting light to illuminate said target surface which possesses a certain color and shade of color;an optical detector circuit for detecting light that is reflected from the target surface and generating a substantially steady DC output voltage that is indicative of the color or shade of color at said target surface;a stored look-up table or mathematical formula corr... | - | A61B005/00
US20100029261 | A1 | 2010-02-04 | Virtual wireless data cable method, apparatus and system | 2009-02-04 | A system, method and apparatus to establish a complete and symmetrical virtual wireless cable for data transmission in such a way that devices, attached to both ends of the system, could use it as if it is a simple regular data cable. | 1. A virtual wireless cable system for data transmission for controlling a plurality of digital data signals transmitted to one or more remotely located devices operatively effective to carry out controllable function upon receipt of a command delivered by a control signal from a first device, wherein said transmission comprises delivering over the air wirelessly a digital audio and/or visual content signal from said first device to a said remotely located device in compressed and optimized format and wherein said digital audio and/or visual content signal is compressed by one or more algorithms effective for compression and over the air delivery of said content and said digital audio and/or visual content is optimized by one or more algorithms effective for optimization of said content for said over the air wireless delivery. 2. The virtual wireless cable system of claim 1 wherein said system is implemented in place of a serial cable employing RS-232 standards and/or a USB cable. 3. The virtual wireless cable system of claim 1 employing RF cellular telephony system. 4. The virtual wireless cable system of claim 1 wherein said system provides for data transmission rates of a minimum of 9600 bit/s. 5. The virtual wireless cable system of claim 1 wherein said system transmits digital data signals in the form of data packets. 6. The virtual wireless cable system of claim 1 wherein said first device is a cell phone or a land line phone. 7. The virtual wireless cable system of claim 1 wherein said remote device is operated by voice recognition. 8. The virtual wireless cable system of claim 1 wherein said remotely located device is employed for access to a secured area. 9. The virtual wireless cable system of claim 1 wherein said system transmits a biometric data signals to a remotely controlled device for access to a secured area or information. 10. The virtual wireless cable system of claim 1 wherein said remotely controlled device is an automobile diagnostics system effective to provide automobile diagnostic information on demand. 11. The virtual wireless cable system of claim 1 wherein said remotely controlled device is a health care diagnostics system effective to provide human and/or animal health diagnostic information on demand. 12. The virtual wireless cable system of claim 1 wherein said remotely controlled device is a musical sound system. 13. The virtual wireless cable system of claim 1 wherein said remotely controlled device is a movie or video or television system. 14. The virtual wireless cable system of claim 1 wherein said remotely controlled device is a vehicle ignition, a watercraft ignition, a heating and/or air conditioning system for a vehicle or watercraft. 15. The virtual wireless cable system of claim 1 wherein said remotely controlled device is selected from window and door controls, telephone systems, computers, security cameras, security systems, intercom systems, driveway sensors, mailbox sensors, automated sh... | - | H04M003/00 G01M017/007 G06F017/00 G06Q030/00 G10L015/00 A61B005/00
US20100049007 | A1 | 2010-02-25 | Integrated physiological sensor apparatus and system | 2008-08-20 | A physiological sensor apparatus, system and method for determining a physiological characteristic, comprising providing at least one physiological sensor that is adapted to measure at least one physiological characteristic at a target measurement site on a subject's body, heating an extended tissue region on the subject's body, whereby blood perfusion of the tissue region is enhanced, and measuring at least one physiological characteristic at the target measurement site with the physiological sensor during or within a predetermined period after heating the extended tissue region. In one embodiment, the sensor system includes at least one temperature algorithm that is adapted to adjust the heat applied to the extended tissue region based on the body's response to the heat stimuli. | 1. An integrated physiological sensor system, comprising: a plurality of physiological sensors, said plurality of physiological sensors including at least a first physiological sensor adapted to measure pulse amplitude at a target measurement site on a subject's body and a second physiological sensor adapted to monitor electrical impulses associated with said subject's heart function; andmeans for heating a tissue region on said subject's body, whereby blood perfusion of said tissue region is enhanced, said tissue region including said target measurement site and extending beyond said target measurement site. 2. The system of claim 1, wherein said system includes a third sensor adapted to monitor blood pressure. 3. An integrated physiological sensor system, comprising: a plurality of physiological sensors, said plurality of sensors including at least a first physiological sensor adapted to measure pulse amplitude at a first target measurement site on a subject's body and a second physiological sensor adapted to monitor electrical impulses associated with said subject's heart function;means for heating a tissue region on said subject's body, whereby blood perfusion of said tissue region is enhanced, said tissue region including said first target measurement site and extending beyond said first target measurement site;a first heat sensor adapted to monitor skin surface temperature of said tissue region; anda processor in communication with said heating means and said first heat sensor, said processor including at least one algorithm for regulating said heating means based on a physiological response of said subject's body to heating of said tissue region. 4. The system of claim 1, wherein said first physiological sensor comprises a pulse oximetry sensor, said pulse oximetry sensor having a signal-to noise ratio. 5. The system of claim 4, wherein said processor includes stored demographic pulse amplitude and skin surface temperature data, and wherein said algorithm is adapted to compare first pulse amplitude measured by said first physiological sensor at said first target measurement site and first skin surface temperature of said tissue region measured by said first heat sensor to said stored demographic amplitude and skin surface temperature data, and adjust the heat provided by said heating means based on said comparison, whereby said signal-to-noise ratio of said first physiological sensor is optimized. 6. The system of claim 5, wherein said demographic pulse amplitude and skin surface temperature data includes measured pulse amplitudes and skin surface temperatures of a second tissue region on a plurality of second subjects during heating of a first tissue region on said second subjects' body, said second subjects' second tissue region being close to, but independent of said second subjects' first tissue region. 7. The system of claim 4, wherein said system includes a third physiological sensor adapted to measure pulse amplitude at a seco... | - | A61B005/00
US20100036277 | A1 | 2010-02-11 | ANIMAL TEMPERATURE MONITOR AND MONITORING METHOD | 2003-07-06 | The invention provides apparatus 2, 12 for determining a condition-related temperature of an animal 1 having a tail 3. The apparatus comprises a temperature sensor 5 and means for determining and providing as an output a condition-related temperature of the animal from an output signal of the temperature sensor. The temperature sensor (or a component 16 bearing the temperature sensor) is secured to the animal so that the temperature sensor is positioned between the tail and left and right buttocks 13, 14 of the animal when the tail is in a specified normal position. The apparatus can be adapted for monitoring of condition-related temperature of several animals, and for remote monitoring. | 1-28. (canceled) 29. A method for determining a condition-indicating temperature of an animal having a tail, the method including the steps of: (a) securing a temperature sensor to an animal having a tail so that when the tail is in a specified normal position close to the animal's buttocks the temperature sensor is positioned between the tail and left and right buttocks of the animal; and(b) determining a condition-indicating temperature from an output signal of the temperature sensor. 30. The method of claim 29 wherein when the tail is in the said normal position the sensor is positioned at least partially within a cleft between the said buttocks of the animal. 31. The method of claim 29 wherein the temperature sensor or a component bearing the temperature sensor is secured to the tail of the animal. 32. The method of claim 29 further including the steps of: (a) also providing on the animal's tail a motion sensor or attitude sensor arranged to indicate movement of the tail from the specified normal position;(b) monitoring an output signal of the motion or tilt sensor; and(c) accepting temperatures derived from output of the temperature sensor only when it can be inferred from output of the motion or attitude sensor that the tail has been at least approximately in the specified normal position long enough for the condition-related temperature to be adequately accurate. 33. The method of claim 29 including the step of analyzing temperatures or temperature-related signals derived from output of the temperature sensor to detect temperatures that are not representative of said condition-indicating temperature, and accepting only temperatures that are representative of said condition-indicating temperature. 34. The method of claim 29 including the step of applying to a temperature as indicated by or estimated from output of the said temperature sensor an experimentally determined correlation between that temperature and a temperature as determined by placement of a thermometer in a cavity of the animal, preferably an ear, the anus, the rectum, the vagina or the mouth. 35. The method of claim 29 including the step of transmitting a temperature or a temperature-related output signal derived from the temperature sensor to a monitoring means remote from the temperature sensor. 36. The method of claim 35 wherein the monitoring means is remote from the animal. 37. The method of claim 35 wherein the transmission occurs in response to an interrogation signal transmitted from a remote location. 38. The method of claim 29 wherein the temperature sensor is one of a plurality of temperature sensors each fitted to a member of a plurality of animals and a temperature of each animal of said plurality is able to be determined and monitored. 39. The method of claim 38 wherein either the temperature or a temperature-related signal is transmitted from a specific animal to a monitoring location in response to an interrogating signal that is transmitted from t... | - | A61B005/01 A01K029/00 A61B005/11
US07657071 | B2 | 2010-02-02 | Systems, devices, and methods for bundle segmentation in diffusion tensor magnetic resonance imaging | 2006-04-10 | Certain exemplary embodiments can comprise a method that can comprise automatically causing a representation of body tissue to be rendered. The representation of the body tissue can comprise a plurality of voxels located in an interior region of the body tissue. Each of the plurality of voxels can have a negative value of an energy change function. | 1. A method comprising: via an information device comprising a graphical user interface, said information device communicatively coupled to a magnetic resonance imaging (MRI) machine:automatically causing a representation of a brain fiber bundle to be rendered, said representation of said brain fiber bundle comprising a plurality of voxels located in an interior region of said brain fiber bundle, each of said plurality of voxels having a negative value of ΔE(x) when ΔE(x) is calculated via:Δ⁢⁢E⁡(x)=log⁢σ1σ2+12⁢d2⁡(I1⁡(x),I0⁡(x))σ12-12⁢d2⁡(I2⁡(x),I0⁡(x))σ22+∇d⁢I1⁡(x)2-∇d⁢I2⁡(x)2+Δ⁢⁢a⁡(x)where:ΔE(x) is an energy change function; σ1 is a variance of tensor distances for voxels located in said interior region;σ2 is a variance of tensor distances for voxels located in an exterior region;I0 is an initial approximation of a tensor image;I1 is a tensor image defined in said interior region;I2 is a tensor image defined in said exterior region;d is a distance measurement between a pair of tensors;∇d is an intrinsic gradient under a metric of a tensor image function Δa(x) is a change in surface area caused by promoting a particular voxel to said interior region; andx represents three dimensional coordinates of a point in Euclidean space, the point directly related to said particular voxel. 2. The method of claim 1, further comprising: promoting a voxel to said interior region based upon a negative value of ΔE(x). 3. The method of claim 1, further comprising: not promoting a voxel to said interior region based upon a positive value of ΔE(x). 4. The method of claim 1, further comprising: initializing said interior region to a single voxel located inside said brain fiber bundle. 5. The method of claim 1, further comprising: defining a band comprising neighboring voxels that approximately surround said interior region. 6. The method of claim 1, further comprising: computing a mean tensor, said mean tensor a local average of tensors inside said interior region. 7. The method of claim 1, further comprising: automatically updating statistical parameters σi according to equationsσ12=∫Ω⁢⁢ⅆ2⁢(I1⁡(x),I0⁡(x))⁢H⁡(x)⁢ⅆx∫Ω⁢H⁡(x)⁢ⅆxσ22=∫Ω⁢⁢ⅆ2⁢(I2⁡(x),I0⁡(x))⁢(1-H⁡(x))⁢ⅆx∫Ω⁢(1-H⁡(x))⁢ⅆx,where: Ω denotes an image domain (union of said interior region and said exterior region);d is a distance measurement between two tensors;H(x)={1, xεΩ1; 0, xεΩ2};Ω1 denotes said interior region; andΩ2 said exterior region. 8. The method of claim 1, further comprising: repeatedly computing ΔE(x) for voxels in a band at least partially surrounding said interior region until all voxels associated with a negative energy change have been promoted to said interior region. 9. The method of claim 1, further comprising: automatically updating I1 based upon a local average of tensors inside said interior region. 10. The method of claim 1, further comprising: automatically updating I2 based upon a local average of tensors inside said exterior region. 11. The method of claim 1, further com... | Siemens Medical Solutions USA, Inc. | G06K009/00 A61B005/05
US07657309 | B2 | 2010-02-02 | Measuring human heart muscle viability using myocardial electrical impedance | 2004-05-28 | A method for detecting a quantitative measure of a physiologic state of a human myocardium or coronary artery. Implementations of the method detect the extent of change of myocardial electrical impedance from a mean baseline value to provide diagnosis of the extent of ischemia, stenosis, tissue rejection, and reperfusion and the effectiveness of cardioplegia and ischemia pre-conditioning as well as the general effectiveness of coronary bypass surgery as measured by post-operative reperfusion. Electrodes are attached to the myocardium, baseline measurements of the mean myocardial electrical impedance are stored and the variance of the myocardial electrical impedance and a baseline value of mean myocardial electrical impedance are computed from the baseline measurements and stored. Mean myocardial electrical impedance values are periodically measured between each electrode pair over an interval of time and stored. After the mean myocardial electrical impedance changes from the computed baseline value by at least the measured variance, the extent of change in the myocardial physiologic state is diagnosed as a continuous, smooth, function of the extent of change, or rate of change, of the periodically measured myocardial electrical impedance from the baseline value. | 1. A method for detecting a quantitative measure of a pathophysiologic state of a human myocardium or coronary artery of an individual, the method comprising: a) attaching at least one electrode pair to the myocardium;b) recording baseline measurements of the mean myocardial electrical impedance and computing the variance of the myocardial electrical impedance between each electrode pair;c) computing a baseline value of mean myocardial electrical impedance from the baseline measurements;d) periodically measuring mean myocardial electrical impedance values between each electrode pair over an interval of time and storing data representing the impedance values as a function of time; ande) after the mean myocardial electrical impedance changes from the computed baseline value by at least the measured variance, diagnosing the extent of change in the myocardial pathophysiologic state as a continuous, smooth, function of the extent of change, or rate of change, of the periodically measured myocardial electrical impedance from the baseline value. 2. A method in accordance with claim 1 wherein: a) the pathophysiologic state is the ischemia of a portion of the myocardium; andb) after the mean myocardial electrical impedance between the electrode pairs rises above a value equal to the arithmetic sum of the baseline myocardial electrical impedance and the variance, myocardial ischemia severity is diagnosed as a continuous, smooth, increasing function of the extent of the rise of the mean myocardial electrical impedance above the baseline value. 3. A method in accordance with claim 1 wherein: a) the pathophysiologic state is the extent of stenosis pre-existing in a coronary artery;b) each electrode pair is attached to the myocardium in the region of the myocardium perfused by the coronary artery;c) the coronary artery is occluded proximally after recording the baseline measurements; andd) after the mean myocardial electrical impedance between the electrode pair rises above a value equal to the arithmetic sum of the baseline myocardial electrical impedance and the variance, the extent of stenosis pre-existing in the coronary artery is diagnosed as a continuous, smooth, decreasing function of the extent of rise of the mean myocardial electrical impedance above the baseline value. 4. A method in accordance with claim 3 wherein the continuous, smooth, decreasing function is substantially: % stenosis=−2.89×% MEI+410.044,wherein% stenosis is the percent pre-existing blockage in the coronary artery and% MEI is the increase of the mean myocardial electrical impedance above the baseline value expressed as a percent. 5. A method in accordance with claim 1 wherein: a) the state is the reperfusion of a portion of the myocardium; andb) after the mean myocardial electrical impedance between the electrode pairs declines below a value equal to the arithmetic difference of the baseline myocardial electrical impedance and the variance, the myocardial reperfusion level is d... | The Ohio State University | A61B005/05
US07649351 | B2 | 2010-01-19 | Method for interference suppression in a measuring device | 2005-02-11 | The present invention describes a method enabling one to shield a device that measures weak biomagnetic signals from strong magnetic interference fields. The measurement sensors are provided with a feedback compensation loop, the difference signal of which is obtained from the measurement sensors themselves. As the actuator of the feedback function, one or more coils are responsible for eliminating, the external interference fields in the region of the sensors. Difference signals can be generated as a linear combination from the signals of two or more sensors. In the control logic, the SSS method can be used to numerically separate the biomagnetic signal being measured from the signals produced by the sources—compensation coils and interference sources—disposed outside the measurement region. The interference suppression can be enhanced by placing the assembly of sensors and the actuators within a magnetically shielding room. | 1. A method for shielding an assembly of sensors that measures biomagnetic signals from external interferences, the assembly of sensors comprising at least one sensor, characterised in that the method comprises the steps of: providing the assembly of sensors with at least one magnetic feedback, wherein the difference signal of one feedback is obtained from at least one sensor of the aforementioned assembly of sensors;producing in the region of the assembly of sensors a magnetic field that compensates for external interferences with at least one actuator disposed outside the assembly of sensors; and the difference signal that produces a compensation voltage in the actuator is a linear combination of the signals of one or more sensors; andusing the SSS method to separate the biomagnetic useful signal being measured from the signals (originating from outside the measurement region), which signals are produced by the actuators and the external interferences. 2. The method as defined in claim 1, characterised in that the method further comprises the steps of: forming a difference signal as a linear combination from the signals of two or more sensors of the assembly of sensors; andfeeding back the difference signal to the actuators, the compensating magnetic field being formed as a linear combination from the magnetic field produced by at least one actuator. 3. The method as defined in claim 2, characterised in that the method further comprises the step of: selecting the difference signals to be obtained as a linear combination and the actuators to be used so that as the feedback is switched on, the external interference signal is minimised. 4. The method as defined in claim 1, characterised in that the method is used in a magnetoencephalographic device (MEG). 5. The method as defined in claim 4, characterised in that the measurement sensor is coupled to two feedback loops, the first of them being a channel of an MEG device that operates in a flow-locked state and the second a feedback loop that controls the compensating actuator. 6. The method as defined in claim 1, wherein in the SSS method, a magnetic field that has been registered using a multi-channel measuring device is analysed in a geometry in which the interesting source is disposed in measurement volume V1, the sensors measuring the field or the components thereof outside volume V1 in volume V2, and the sources of the magnetic interferences and the actuators outside volume V1 and V2 in volume V3, which can be infinite, characterised in that the method further comprises the steps of: parametrising the magnetic field produced by the interesting sources disposed in volume V1 in volume V2 as a sum of elementary fields, each of which is irrotational, sourceless and finite outside volume V1 so that a presentation of the desired accuracy is achieved for the parametrised magnetic field in volume V2;parametrising the sum magnetic field produced by the interference sources and compensating actuato... | Elekta AB (publ) | G01R033/00 A61B005/05
US20100049264 | A1 | 2010-02-25 | Diagnostic indicator and PSG interface for a closed loop neuromodulator | 2009-08-21 | This document discusses, among other things, a system and method for providing diagnostic information from a closed loop neuromodulator, configured to provide sleep therapy for a patient, to a user or a polysomnograph (PSG). A first portion of the diagnostic information can be received at a diagnostic indicator register of the closed loop neuromodulator and an LED driver can be switched using the first portion of the diagnostic information. A second portion of the diagnostic information can be received at a PSG interface of the closed loop neuromodulator and provided to the PSG for display. | 1. A closed loop neuromodulator interface for providing diagnostic information from a closed loop neuromodulator, configured to provide sleep therapy for a patient, to a user or a polysomnograph (PSG), the interface comprising: a diagnostic indicator register coupled to a data bus of the closed loop neuromodulator, the diagnostic indicator register configured to receive a first portion of the diagnostic information and to switch an LED driver of a plurality of LED drivers in response to the first portion of the diagnostic information; anda PSG interface configured to receive a second portion of the diagnostic information and to provide the second portion of the diagnostic information to the PSG. 2. The interface of claim 1, including a digital airflow buffer coupled to an inhalation LED driver of the plurality of LED drivers, the digital airflow buffer configured to receive a digital airflow signal indicative of a respiration state of the patient from a hardened sensor interface of the closed loop neuromodulator; and wherein the inhalation LED driver is configured to illuminate an inhalation LED when the digital airflow signal is indicative of an inhalation respiration state of the patient. 3. The interface of claim 2, wherein the closed loop neuromodulator is configured to detect a sleep related disorder using the digital airflow signal and to provide a first stimulus to the patient in response to the detected sleep disorder event, the first stimulus configured to not interrupt the detected sleep disorder event; and wherein the closed loop neuromodulator is configured to provide, following the first stimulus, a second stimulus to the patient in response to the detected sleep disorder event, the second stimulus having more energy than the first stimulus and configured to interrupt the detected sleep disorder event. 4. The interface of claim 1, including a stimulus signal buffer coupled to the PSG interface, the stimulus signal buffer configured to receive a stimulus signal, indicative of the sleep therapy, from a stimulus generator of the closed loop neuromodulator and to provide the stimulus signal to the PSG interface. 5. The interface of claim 1, including a driver signal buffer coupled to the PSG interface, the driver signal buffer configured to receive a driver signal from a stimulus escalator of the closed loop neuromodulator and to provide the driver signal to the PSG interface. 6. The interface of claim 1, including an analog airflow signal buffer coupled to the PSG interface, the analog airflow signal interface configured to receive an analog airflow signal indicative of respiration effort of the patient from a hardened sensor interface of the closed loop neuromodulator and to provide the analog airflow signal to the PSG interface. 7. The interface of claim 1, including digital airflow signal buffer coupled to an exhalation LED driver of the plurality of LED drivers, the digital airflow signal buffer configured to receive a digital... | Dymedix Corporation | A61N001/00 A61B005/08
US07657437 | B2 | 2010-02-02 | Method for conducting prescription drug co-payment plans | 2003-06-27 | A credit card-based prescription benefits plan utilizes an electronic means for telecommunication to rapidly adjudicate prescription claims. The adjudication process includes a third-party claims processor interposed the pharmacy and the patient's credit card clearinghouse to ensure that the subscriber receives all benefits available under the plan. The third-party claims processor may also provide patient counseling and advocacy by performing patient-specific drug regimen reviews to check for potential adverse drug reactions and drug interactions. | 1. A method of providing prescription benefits to subscribers under a subscriber co-pay plan via an electronic communications network to ensure that the subscriber receives a discounted prescription price, the method comprising the steps of: a) identifying a sponsor for a prescription benefits plan;b) designing the prescription benefits plan to include a formulary of preferred medications, and to define prescription medications that are covered by the prescription benefits plan;c) identifying at least one pharmacy that: i) participates in the prescription benefits plan,ii) receives prescriptions for medication and prescription benefit claims from subscribers, andiii) fulfills the prescriptions for a discounted prescription price;d) identifying at least one eligible subscriber having a valid credit or debit account with a creditor or institution and wishing to participate in the prescription benefits plan;e) enrolling each eligible subscriber who wishes to participate in the prescription benefits plan and generating an electronic file for each of the eligible subscribers;f) generating an eligibility profile for each subscriber and storing said eligibility profile in an electronic memory;g) establishing a prescription benefits manager to administer the prescription benefits plan;h) establishing a prescription claims processor that: i) receives claims over the electronic communications network for prescription benefits from a participating pharmacy;ii) adjudicates the claims in accordance with the subscriber's eligibility profile; andiii) upon approval of a claim, charges the subscriber's credit or debit account for accepted claims via the electronic communications network, the amount of the charge being equal to the discounted price of the prescription minus a predetermined co-payment which is tendered by the subscriber to the pharmacy at the time of fulfilling the prescription;i) making a first payment via the electronic communications network by the creditor to the prescription benefits manager, the first payment being equal to the amount charged to the subscriber's account minus a first predetermined service fee to be retained by the creditor;j) making a second payment via the electronic communications network by the prescription benefits manager to the pharmacy, the second payment being equal to a predetermined portion of the discounted prescription price; andk) making a third payment via the electronic communication network by the prescription benefits manager to the claims processor, the third payment being equal to a second predetermined service fee, wherein the second payment and the third payment are funded by the first payment and whereby the sum of the pre-determined co-payment and the amount charged to the subscriber's account is the discounted prescription price paid by the subscriber. 2. The method of claim 1 wherein the claims processor further: a) performs a drug regimen review prior to approval of a claim; andb) advises the pharma... | Omnicare, Inc. | G06Q010/00 G06Q030/00 G06Q050/00 G06Q040/00 A61B005/00 G06F019/00
US07643861 | B2 | 2010-01-05 | Technique for design, and placement, of a subdermal Ag—Ag/Cl biopotential electrode | 2006-01-17 | A chronic subdermal silver-silver/chloride (Ag—Ag/Cl) electrode for the recording of biopotentials is disclosed. The electrode is comprised of an Ag—Ag/Cl distal end that is placed subdermally by various methods. The recording electrode is made of solid, multi-stranded or a ball of pure silver and the whole length of the electrode is insulated except for the distal end which is used for voltage recordation. The electrode can be placed subdermally by an introducer apparatus such as hypodermic or sewing needle, a split-cannula, a staple, or a suture. Once placed, the introducer apparatus can be removed leaving the recording wire/element in place. The recording electrode can be placed at any location on the body to record the local biopotential (EEG, EOG, EMG, EKG, etc). It can be left in place for hours, days, or weeks, as it will record without any further adjustment. | 1. A wire electrode for measuring biopotentials in a subdermal layer, said electrode comprising a pure multi-stranded or solid silver wire, said wire being insulated and having an uninsulated distal end, the uninsulated distal end of said wire being at least 3 mm in length and coated with silver/chloride, a proximal end of said wire being operatively coupled to a biopotential recording device. 2. The wire electrode of claim 1, wherein said proximal end of said wire further comprises a connector for operatively coupling same to said biopotential voltage recording device. 3. The wire electrode of claim 2, wherein said connector is insulated and sealed with heat-shrink tubing. 4. The wire electrode of claim 2, wherein said connector is soldered or crimped to the proximal end of the wire. 5. The wire electrode of claim 1, wherein the pure multi-stranded or solid insulated silver wire has a length of 2″ and an overall diameter of about 0.0045″ or 0.114 mm. 6. The wire electrode of claim 1, wherein the distal end of the wire is 3 mm in length. 7. The wire electrode of claim 1, wherein the distal end of the wire is coated with silver/chloride by placing the distal end of the wire in chlorine bath with a silver anode with a voltage potential of 1.5 Volts for about 15 minutes. 8. The wire electrode of claim 1, wherein said wire is chlorided with a silver salt to enable recordation of biopotential voltage at the site of placement for EEG, EGG, EKG, EMG applications. 9. The wire electrode of claim 1, wherein the distal end of the wire is 4 mm in length. 10. The wire electrode of claim 1, wherein said proximal end of the wire is operatively coupled to a biopotential recording device via an extension wire. 11. An electrode device for introducing a wire electrode in a subdermal layer and measuring biopotentials therein, said electrode device comprising: a wire electrode comprising a pure multi-stranded or solid silver wire, said wire being insulated and having an uninsulated distal end, the uninsulated distal end of said wire being at least 3 mm in length and coated with silver/chloride, a proximal end of said wire being operatively coupled to a biopotential recording device, an introducer needle for introducing the distal end of the wire in the subdermal layer, said introducer needle comprising a needle hub and a hollow needle extending therefrom, said hollow needle being at least 25 gauge, whereby the wire is hooked into a needle lumen of the hollow needle at the distal end of the wire and the wire is folded back along the length of the hollow needle to carry the distal end of the wire into the subdermal layer. 12. The electrode device of claim 11, wherein said hollow needle is ⅝″ in length. 13. The electrode device of claim 11, wherein said hollow needle has a sharp end lumen and a slanted bevel for permitting penetration of the skin and for allowing said hollow needle to carry the distal end of the wire to the subdermal layer. 14. The electrod... | John Richard Ives Susan Elizabeth Phelan | A61B005/04
US07643873 | B2 | 2010-01-05 | Exercise data apparatus | 2006-04-28 | Exercise data apparatus for use on the body of a user during exercise, includes an electrocardiogram detector for detecting electrocardiogram data of the user during exercise; an attachment device for attaching the electrocardiogram detector onto the chest of the user; a motion detector for detecting body motion data indicative of steps of the user during exercise; a processor for calculating exercise data based on the electrocardiogram data detected by the electrocardiogram detector and the body motion data detected by the motion detector, the exercise data including at least one of heart rate, speed and distance, and calorie consumption of the user; and an output device for informing the user of the exercise data calculated. | 1. An exercise data apparatus for use on the body of a user during exercise, comprising: an electrocardiogram detector for detecting electrocardiogram data of a user during exercise;attachment means for attaching the electrocardiogram detector onto the chest of the user;a motion detector for detecting body motion data indicative of steps of the user during exercise;a processor for calculating exercise data based on the electrocardiogram data detected by the electrocardiogram detector and the body motion data detected by the motion detector, the exercise data comprising at least one of heart rate, speed and distance, and calorie consumption of the user, wherein the processor is programmed to calculate the calorie consumption based on weight, fitness level, and heart rate of the user; andan output device for informing the user of the exercise data calculated, wherein the calorie consumption is calculated according to the formula:Calorieconsumption=TI*Weight*(instant⁢⁢HR/maximum⁢⁢HR)-AF,wherein HR is the heart rate of the user,TI and AF are Training Index and Activity Factor, respectively, and are related to VO2 percentage,VO2 is oxygen consumption of the user, VO2max is maximum rate of oxygen consumption by the user, and VO2 percentage is the percentage of VO2max being consumed by the user, and when VO2 percentage is 50% to 60%, TI=6 and AF=400,when VO2 percentage is 60% to 70%, TI=8 and AF=580,when VO2 percentage is 70% to 80%, TI=12 and AF=1,000,when VO2 percentage is 80% to 90%, TI=18 and AF=1,720, andwhen VO2 percentage is 90% to 100%, TI=22 and AF=2,260. 2. The exercise data apparatus as claimed in claim 1, wherein the motion detector comprises a pedometer. 3. The exercise data apparatus as claimed in claim 2, wherein the pedometer is physically associated with the electrocardiogram detector for attachment onto the chest of the user by the attachment means. 4. The exercise data apparatus as claimed in claim 2, wherein the pedometer is one of a pendulum sensor and a piezoelectric sensor. 5. The exercise data apparatus as claimed in claim 1, wherein the motion detector comprises an accelerometer. 6. The exercise data apparatus as claimed in claim 5, wherein the accelerometer is wrist-mounted. 7. The exercise data apparatus as claimed in claim 5, wherein the accelerometer is one of an integrated circuit sensor and a piezoelectric sensor. 8. The exercise data apparatus as claimed in claim 1, comprising a sensor unit, a data unit separate from the sensor unit, and a signal link between the sensor unit and the data unit, wherein: the sensor unit comprises the electrocardiogram detector, the attachment means, and the motion detector; andthe data unit comprises the processor, the output device, and individual attachment means for attachment onto a wrist of the user. 9. The exercise data apparatus as claimed in claim 8, wherein the motion detector comprises a pedometer. 10. The exercise data apparatus as claimed in claim 8, wherein the output ... | IDT Technology Limited | A61B005/0452 A61B005/0402 A61B005/5083
US07648463 | B1 | 2010-01-19 | Monitoring device, method and system | 2006-12-03 | A monitoring device (20) and method (200) for monitoring the health of a user is disclosed herein. The monitoring device (20) preferably includes eyewear (25) with an optical sensor (30), a digital storage and processing device (35) with a display member (40) and a control component (43), and a connection cable (45). The monitoring device (20) preferably displays the following information about the user: pulse rate; calories expended by the user of a pre-set time period; target zones of activity; time; and distance traveled. | 1. A method of a providing a real-time heart rate for a user, the method comprising: generating a signal corresponding to the flow of blood through an anterior ethmoidal artery of the user, the signal generated from an optical sensor in proximity to the anterior ethmoidal artery of the user, the optical sensor integrated on eyewear worn by the user, the eyewear comprising a lens support portion, a pair of lenses positioned within the lens support portion, a first temporal member attached and extending from the lens support portion, a second temporal member attached and extending from the lens support portion a first nose support attached to the lens support portion and a second nose support attached to the lens supporting portion, wherein the optical sensor comprises a light source integrated into the first nose support and a photodiode integrated into the second nose support;transmitting the signal from the optical sensor through a connection cable to a digital storage and processing device;processing the signal at the digital storage and processing device using an adaptive algorithm to obtain a real-time heart rate for the user, the processing comprising deactivating a processor for a deactivation period, activating the optical sensor during the deactivation period to generate a signal corresponding to the flow of blood through the anterior ethmoidal artery of the user, reactivating the processor after the deactivation period to generate real-time heart rate data of the user from the signal generated by the optical sensor, processing the real-time heart rate data of the user for analysis of a plurality of real-time vital signs of the user including calories expended by the user, and for real-time display of a plurality of the user's vital signs; anddisplaying the real-time heart rate and a plurality of the user's vital signs on a display member disposed on an exterior surface of the digital storage and processing device, the display of the real-time heart rate and the plurality of user's vital signs controlled by the user using a control component extending from the digital storage and processing device. 2. The method according to claim 1 wherein the digital storage and processing device is a MP3 player. 3. The method according to claim 1 wherein the optical sensor comprises a green LED light source which generates a wavelength at approximately 567 nanometers and a light-to-voltage converter. 4. A monitoring device for monitoring the health of a user, the monitoring device comprising: eyewear comprising a lens support portion, a pair of lenses positioned within the lens support portion, a first temporal member attached and extending from the lens support portion, a second temporal member attached and extending from the lens support portion a first nose support attached to the lens support portion and a second nose support attached to the lens supporting portion;an optical sensor comprising a light source integrated into the first nose suppor... | Impact Sports Technologies, Inc. | A61B005/02
US08679017 | B2 | 2014-03-25 | Biological information measurement apparatus and method | 2009-07-23 | A biological information measurement apparatus includes an electromagnetic wave applying unit configured to apply, to a living body, a first electromagnetic wave and a second electromagnetic wave having a frequency different from a frequency of the first electromagnetic wave; a reflected wave receiver configured to receive a first reflected wave corresponding to the first electromagnetic wave and a second reflected wave corresponding to the second electromagnetic wave; a correlation value configured to calculate unit calculating a correlation value between the first and second reflected waves; a correlation value evaluating unit configured to determine whether the correlation value satisfies a given condition; and a biological information measuring unit configured to measure biological information based on the first or second reflected wave when the correlation value evaluating unit determines that the correlation value satisfies the given condition. | 1. A biological information measurement apparatus comprising: an ultrasound applying unit configured to apply ultrasound to a living body;an electromagnetic wave applying unit configured to apply, to a living body, a first electromagnetic wave and a second electromagnetic wave having a frequency different from a frequency of the first electromagnetic wave, while the ultrasound is being applied to the living body;a reflected wave receiver configured to receive a first reflected wave corresponding to the first electromagnetic wave and a second reflected wave corresponding to the second electromagnetic wave;a correlation value calculating unit configured to calculate a correlation value between the first and second reflected waves;a correlation value evaluating unit configured to determine whether the correlation value satisfies a given condition;a biological information measuring unit configured to measure biological information based on at least one of the first and second reflected waves when the correlation value evaluating unit determines that the correlation value satisfies the given condition; anda calculation count checking unit configured to count correlation value calculations performed by the correlation value calculating unit and determine whether a calculation count is greater than a given calculation count threshold,wherein the electromagnetic wave applying unit continuously changes a frequency applied to a subject by applying an output signal having a different frequency from a signal previously applied until the correlation value between the first and second reflected waves exceeds a correlation value threshold,the correlation value calculating unit calculates the correlation value and stores, in a correlation value storage unit, the calculated correlation value and frequencies of the two electromagnetic waves corresponding to the calculated correlation value,the ultrasound applying unit changes a frequency of ultrasound to be applied to the living body when the calculation count checking unit determines that the calculation count is greater than the calculation count threshold, andwherein the electromagnetic wave applying unit obtains frequencies corresponding to the highest correlation value from the correlation value storage unit and applies two electromagnetic waves having the obtained frequencies to the living body when the calculation count checking unit determines that the calculation count is greater than the calculation count threshold. 2. The biological information measurement apparatus according to claim 1, wherein the correlation value evaluating unit determines whether the correlation value is greater than the correlation value threshold, andthe biological information measuring unit measures the biological information when the correlation value evaluating unit determines that the correlation value is greater than the correlation value threshold. 3. The biological information measurement apparatus according to claim 1, ... | Fujitsu Limited | A61B008/00
US08577439 | B2 | 2013-11-05 | System and method for positioning electrodes on a patient body | 2007-09-03 | A system for positioning electrodes on a patient body includes an image capturing system, a memory device, a processing system and an indicator system. The image capturing system generates an actual image of the patient body. The memory device stores a reference image, the reference image including a reference body and a reference position on the reference body. The processing system compares the actual image and the reference image and for determining an electrode position on the patient body by matching the reference position on the reference body. The indicator system indicates the electrode position on the patient body. | 1. System for positioning electrodes on a body of a patient, the system comprising: a bio-impedance electode;an image capturing system for generating an actual image of the patient body;a memory device for storing a reference image, the reference image comprising a reference body and a reference position on the reference body;a processing system for comparing the actual image and the reference image and for determining a position of the bio-impedance electrode on the patient body in the actual image, based on the comparison, by matching the reference position in the reference image to a corresponding position on the patient body in the actual image; andan indicator system for indicating the bio-impedance electrode position on the patient body based on the determined bio-impedance electrode position in the actual image, wherein the indicator system comprises a radiation pointer system for projecting an indicator projection image on the bio-impedance electrode position on the patient body. 2. The system according to claim 1, wherein the indicator system further comprises a mirror positionable such that the patient is enabled to see the indicator projection image on the patient body. 3. The system according to claim 1, the image capturing system comprising a stereo camera system for generating a stereoscopic actual image of the patient body, and wherein the reference image is a stereoscopic image. 4. The system according to claim 3, wherein the processing system determines spatial distances on a surface of the patient body using the stereoscopic actual image and the stereoscopic reference image, enabling determination of the bio-impedance electrode position on the patient body when a position or posture of the patient body in the actual image differs from a position or posture of the reference body in the reference image. 5. The system according to claim 1, wherein the system further comprises a patient body holder for holding the patient body in a predetermined patient posture corresponding to a reference posture of the reference body in the reference image. 6. The system according to claim 1, wherein the indicator system comprises a display device for displaying the actual image of the patient body and the electrode position on the patient body. 7. Method for positioning a bio-impedance electrode on a patient body, the method comprising: (a) capturing an actual image of the patient body;(b) providing a reference image, the reference image comprising a reference body and a reference position on the reference body;(c) comparing the actual image and the reference image for determining a bio-impedance electrode position on the patient body in the actual image, based on comparison, by matching the reference position in the reference image to a corresponding position on the patient body in the actual image; and(d) indicating the bio-impedance electrode position on the patient body, based on the actual image, by projecting a spot of visible radiati... | Koninklijke Philips N.V. | A61B005/04
US08406496 | B2 | 2013-03-26 | Method and system for left ventricle detection in 2D magnetic resonance images | 2009-07-16 | A method and system for left ventricle (LV) detection in 2D magnetic resonance imaging (MRI) images is disclosed. In order to detect the LV in a 2D MRI image, a plurality of LV candidates are detected, for example using marginal space learning (MSL) based detection. Candidates for distinctive anatomic landmarks associated with the LV are then detected in the 2D MRI image. In particular, apex candidates and base candidates are detected in the 2D MRI image. One of the LV candidates is selected as a final LV detection result using component-based voting based on the detected LV candidates, apex candidates, and base candidates. | 1. A method for left ventricle (LV) detection in a 2D magnetic resonance imaging (MRI) image, comprising: detecting a plurality of LV candidates in the 2D MRI image;detecting a plurality of apex candidates in the 2D MRI image;detecting a plurality of base candidates in the 2D MRI image; andselecting one of the plurality LV candidates using component-based voting based on each of the plurality of LV candidates, the plurality of apex candidates, and the plurality of base candidates, wherein said step of selecting one of the plurality LV candidates using component-based voting based on each of the plurality of LV candidates, the plurality of apex candidates, and the plurality of base candidates comprises: determining votes for each of the plurality of LV candidates based on other LV candidates of the plurality of LV candidates,determining votes for each of the plurality of LV candidates based on the apex candidates,determining votes for each of the plurality of LV candidates based on the base candidates, andselecting as a detection result the one of the plurality of LV candidates having a largest number of votes. 2. The method of claim 1, wherein said step of detecting a plurality of LV candidates in the 2D MRI image comprises: detecting the plurality of LV candidates using marginal space learning (MSL). 3. The method of claim 1, wherein said step of detecting a plurality of LV candidates in the 2D MRI image comprises: detecting a plurality of position candidates for the LV in the 2D MRI image using a trained position detector;generating a plurality of position-orientation hypotheses from each of the plurality of position candidates;detecting a plurality of position-orientation candidates from the plurality of position-orientation hypotheses using a trained position-orientation detector;generating a plurality of position-orientation-scale hypotheses from each of the plurality of position-orientation candidates; anddetecting the plurality of LV candidates from the plurality of position-orientation-scale hypotheses using a trained position-orientation-scale detector. 4. The method of claim 3, wherein each of the position detector, the position-orientation detector, and the position-orientation-scale detector are trained using a probabilistic boosting tree (PBT). 5. The method of claim 2, wherein the plurality of apex candidates and the plurality of base candidates are detected using MSL. 6. The method of claim 1, wherein said step of determining votes for each of the plurality of LV candidates based on other LV candidates of the plurality of LV candidates comprises, for each LV candidate: determining a vote for the LV candidate from each of the other LV candidates located within a certain distance of the LV candidate. 7. The method of claim 1, wherein said step of determining votes for each of the plurality of LV candidates based on the apex candidates comprises, for each LV candidate: determining a predicted position of the apex for the LV cand... | Siemens Aktiengesellschaft | G06K009/00
US20100041970 | A1 | 2010-02-18 | IMPLANTABLE MEDICAL DEVICE | 2007-11-27 | An implantable medical device has an oxygen sensor adapted to measure the level of oxygen in oxygenized blood, and to generate an oxygen measurement signal in dependence of the level of oxygen. The oxygen sensor is adapted to perform measurements inside the heart, of blood entering the left atrium of a patient's heart. The obtained oxygen measurement signal is compared to a predetermined threshold level and an indication signal is generated in dependence of the comparison. The, indication signal is indicative of the lung functionality of the patient. | 1.-23. (canceled) 24. An implantable medical device comprising: a housing configured for in vivo implantation in a living subject;an oxygen sensor that measures a level of oxygen in oxygenated blood and that generates an oxygen measurement signal dependent on said level of oxygen;a carrier for said oxygen sensor configured to place said oxygen sensor at a location inside the heart of the subject to cause said oxygen sensor to measure said level of oxygen in blood entering the left atrium of the heart of the subject; anda comparator in said housing and in communication with said oxygen sensor to receive said oxygen measurement signal therefrom, said comparator comparing said oxygen measurement signal to a predetermined threshold and emitting a comparator output signal, as a result of the comparison, that is indicative of long functionality of the subject. 25. A device as claimed in claim 24 wherein said carrier is an electrode lead connected to said housing, said electrode lead having a distal end and said optical sensor being disposed approximately at said distal end of said electrode lead. 26. A device as claimed in claim 25 wherein said optical sensor comprises a photodiode and a light-emitting diode, said light-emitting diode being disposed at a distal tip of said electrode lead, facing a distal direction of said electrode lead. 27. A device as claimed in claim 24 wherein said carrier is configured to place the optical sensor in the left atrium of the heart of the subject. 28. A device as claimed in claim 27 wherein said carrier is configured to place said optical sensor in the left atrium of the heart by proceeding through the right atrium via the septal wall between the right atrium and the left atrium. 29. A device as claimed in claim 24 wherein said oxygen sensor comprises a light source that, during a measurement session, emits light at a first wavelength, and a photodetector and at least one type of photo luminescent molecules embedded in a carrier element that is selectively permeable to oxygen, at least a portion of said carrier element being configured to be in contact with arterial blood of the subject, and said light source and said photodetector being optically connected to said carrier element, said photo luminescent molecules emitting light in response to excitation by the light emitted by said light source, and said photo luminescent molecules reacting with oxygen to alter characteristics of the light emitted from the photo luminescent molecules, and said photodetector being configured to detect the light emitted from the photo luminescent molecules and to generate photodetector output signals indicative of concentration of oxygen in said arterial blood. 30. A device as claimed in claim 24 wherein said carrier comprises an electrode lead connected to said housing and having a distal end, said oxygen sensor being disposed approximately at said distal end of said electrode lead, and said electrode lead comprising a distal t... | St. Jude Medical AB Corproation | A61B005/1459
US20100016746 | A1 | 2010-01-21 | PERSONAL ALERTING DEVICE FOR USE WITH DIAGNOSTIC DEVICE | 2008-07-15 | A personal altering device and method of generating an alert of an event detected by an implanted diagnostic device are provided. The method can include transmitting an event signal to the personal alerting device when one or more of the following occurs: ischemia is detected, a deviation from the baseline waveform is detected, a life-threatening trend begins, or symptoms are indicated. The method can also include generating an alarm signal from the personal alerting device, and the alarm signal can include one or more of a visual message, a light, a vibration, and a sound. In addition, the method can include prompting further action to be taken. | 1. A method of generating an alert of an event detected by an implanted diagnostic device, the method comprising: providing a personal alerting device;transmitting an event signal to the personal alerting device from the implanted diagnostic device when at least one of ischemia is detected, a deviation from a baseline electrocardiogram waveform is detected, a life-threatening trend begins, and symptoms are indicated;storing the event signal in the personal alerting device, the event signal including an event onset timestamp, an event type, and an event electrocardiogram waveform;generating an alarm signal from the personal alerting device, the alarm signal including at least one of a visual message, a light, a vibration, and a sound; andprompting further action to be taken through the personal alerting device. 2. The method of claim 1 and further comprising transmitting an event signal from an implanted diagnostic device that is positioned subcutaneously. 3. The method of claim 1 wherein transmitting an event signal to the personal alerting device from the implanted diagnostic device when ischemia is detected includes detecting a new elevated ST segment in the patient's electrocardiogram. 4. The method of claim 1 and further comprising providing a button on the personal alerting device to at least one of acknowledge and silence the alarm signal. 5. The method of claim 1 and further comprising providing a patient activator button on the personal alerting device that allows an event signal to be triggered when symptoms occur. 6. The method of claim 1 and further comprising automatically relaying at least one of the event signal and the alarm signal to an emergency medical system. 7. The method of claim 1 and further comprising initiating at least one of enhanced recording functions and enhanced diagnostic functions in the implanted diagnostic device when the event signal is transmitted to the personal alerting device. 9. The method of claim 1 and further comprising instructing the implanted diagnostic device to perform further diagnostic tests after transmitting the event signal. 10. The method of claim 1 and further comprising initiating therapeutic mitigations after receiving the event signal. 11. The method of claim 1 and further comprising storing patient information in the personal alerting device, the patient information including at least one of contact information, baseline electrocardiogram waveforms, historical data, periodic data, and episode data. 12. The method of claim 1 and further comprising displaying resuscitative response instructions on the personal alerting device. 13. The method of claim 1 and further comprising relaying the event signal via short message service. 14. The method of claim 1 and further comprising tailoring the personal alerting device by at least one of modifying thresholds and updating baseline electrocardiogram waveforms. 15. A personal alerting system that generates an alert of an event, the per... | - | A61B005/0402
US20100030097 | A1 | 2010-02-04 | METHOD TO DETERMINE THE ATTRIBUTES ASSOCIATED WITH A BRAND OR PRODUCT | 2006-12-22 | A method of evaluating characteristics of a brand or product including the steps of: (a) presenting the brand or product to the subject during a first period; (b) determining brain activity of the subject during the first period; (c) presenting neutral visual and/or audio material to a subject during a second period; (d) determining a reference level of brain activity of the subject during the second period; and (e) evaluating attributes associated by the subject with the brand or product by determining differences in brain activities between the first and second periods. | 1. A method of evaluating characteristics of a brand or product including the steps of: (a) presenting the brand or product to the subject during a first period;(b) determining brain activity of the subject during the first period;(c) presenting neutral visual and/or audio material to a subject during a second period;(d) determining a reference level of brain activity of the subject during the second period; and(e) evaluating attributes associated by the subject with the brand or product by determining differences in brain activities between said first and second periods. 2. A method as claimed in claim 1 including the step of displaying steps (a) and (c) as a presentation sequence. 3. A method as claimed in claim 2 wherein there are a plurality of steps (a) in each presentation sequence and wherein blank periods are presented between successive steps (a). 4. A method as claimed in claim 3 wherein the first periods have a duration of 0.5 to 5 seconds, the second periods have a duration of 10 to 60 seconds and the blank periods have a duration of 0 to 5 seconds. 5. A method as claimed in claim 3 wherein the second period is at the end of said sequence. 6. A method as claimed in claim 1 wherein steps (a) to (d) are presented to a plurality of subjects and step (e) includes the steps of averaging the differences in brain activities of the subjects. 7. A method as claimed in claim 1 wherein step (a) is carried out by displaying the brand or product on a video screen. 8. A method as claimed in claim 1 wherein steps (b) and (d) are carried out by determining gamma or high frequency EEG or MEG activity. 9. A method as claimed in claim 1 wherein steps (b) and (d) are carried out by detecting EEG or MEG activity in the frequency range 8 to 13 Hz. 10. A method as claimed in claim 1 wherein steps (b) and (d) are carried out by assessment of the phase of steady state visually evoked potentials (SSVEP) in EEG signals obtained from the subject or subjects or by assessment of steady state visually evoked responses (SSVER) in MEG signals obtained from the subject or subjects. 11. A method as claimed in claim 1 wherein steps (a) and (c) include the steps of placing electrodes at scalp sites to obtain output EEG signals which enable assessment of: visual attention to detail of the brand or product;visual attention to global features of the brand or product;desirability of the brand or product;emotional intensity associated with the brand or product;long term memory encoding associated with the brand or product;engagement with the brand or product; and/or attraction associated with the brand or product. 12. A method as claimed in claim 11 including the step of applying a sinusoidally varying visual flicker stimulus to each subject during steps (a) and (c) to thereby enable calculation of Fourier coefficients from said output signals to thereby enable calculation of said SSVEP amplitudes and/or phase differences. 13. A method as claimed in claim 12 where... | - | A61B005/0484
US20100030048 | A1 | 2010-02-04 | Subcutaneous Glucose Electrode | 2009-08-31 | A small diameter flexible electrode designed for subcutaneous in vivo amperometric monitoring of glucose is described. The electrode is designed to allow “one-point” in vivo calibration, i.e., to have zero output current at zero glucose concentration, even in the presence of other electroreactive species of serum or blood. The electrode is preferably three or four-layered, with the layers serially deposited within a recess upon the tip of a polyamide insulated gold wire. A first glucose concentration-to-current transducing layer is overcoated with an electrically insulating and glucose flux limiting layer (second layer) on which, optionally, an immobilized interference-eliminating horseradish peroxidase based film is deposited (third layer). An outer (fourth) layer is biocompatible. | 1. An electrochemical sensor comprising: one or more non-corroding metal or carbon electrodes;a sensing layer comprising an enzyme coupled to each electrode; anda biocompatible layer comprising a biocompatible hydrogel chemically bound to the sensing layer of each electrode. | Abbott Diabetes Care Inc. | A61B005/1473
US20100036228 | A1 | 2010-02-11 | SUTURE SLEEVE AND A METHOD FOR IMPLANTING ONE OR MORE ELECTRICAL LEADS INTO A VEIN | 2007-03-26 | In a suture sleeve and implantation method for one or more implantable leads, the suture sleeve is adapted to be inserted into a vein to secure and protect the one or more leads from damage when a suture thread is positioned and tied around the vein in the region over the suture sleeve to secure the suture sleeve and prevent bleeding from the vein. The suture sleeve has two or more lead receiving through holes into each of which a medical implantable lead may be inserted. The suture sleeve is formed such that at least all of the holes except one are provided with sealing means, which easily can be broken or removed when inserting a lead into the hole such that, when inserted into a vein, bleeding is prevented from a through hole of the sleeve by the sealing means even if no lead is positioned in the hole. The invention also relates to a method for implanting one or more electrical leads into a vein. | 1-14. (canceled) 15. A suture sleeve for one or more implantable medical leads, comprising: a suture sleeve body configured for insertion into a vein;said suture sleeve body comprising at least two through-holes each configured to receive a medical lead therein;at least one of said through-holes having a seal therein, said seal being frangible or removable to allow insertion of a medical lead into that through-hole; andsaid suture sleeve body being configured to be inserted into a vein to protect the medical leads therein from damage and being configured to be secured in the vein by a suture thread tied around said suture sleeve body to prevent bleeding from the vein, and said seal preventing bleeding from the through-hole containing the seal even when no lead is in that through-hole. 16. A suture sleeve as claimed in claim 15 wherein seal comprises a thin membrane. 17. A suture sleeve as claimed in claim 15 wherein said sleeve body is formed of an elastic material, and wherein each through-hole that comprises a seal therein has a slot proceeding in a longitudinal direction of that through-hole, the material around said slot being deflected for widening the slot and inserting a medical lead sideways into the through-hole through the widened slot, said seal being located in said slot and being frangible or removable upon opening and widening of the slot. 18. A suture sleeve as claimed in claim 15 wherein all of said through-holes comprise a seal. 19. A suture sleeve as claimed in claim 15 wherein said one of said through-holes without said seal is unbroken and is configured for pre-assembly on a medical lead. 20. A suture sleeve as claimed in claim 15 wherein said suture sleeve body has a generally circular cross-section. 21. A suture sleeve as claimed in claim 15 wherein said suture sleeve body comprises at least one circumferential groove proceeding around an exterior of the sleeve body and configured to receive suture thread therein. 22. A suture sleeve as claimed in claim 15 wherein said sleeve body has a truncated conical shape. 23. A method for implanting one or more electrical leads into a vein by means of a suture sleeve, which is adapted to secure and protect the one or more leads from damage when a suture thread is tied around the vein in the region over the suture sleeve to secure the suture sleeve and prevent bleeding from the vein, comprising the steps of: providing a suture sleeve comprising two or more lead receiving through holes;providing at least each hole except one with a sealing means, which easily can be broken or removed;inserting one or more electrical leads into the suture sleeve and into the vein, while removing the sealing means in the respective hole; andpositioning the suture sleeve into the opening in the wall of the vein and tie a suture thread around the vein in the region of the suture sleeve. 24. A method according to claim 23, comprising the further step of: providing a suture sleeve comprising sealing ... | - | A61B005/04
US20100023080 | A1 | 2010-01-28 | IMPLANTABLE MYOCARDIAL ISCHEMIA DETECTION, INDICATION AND ACTION TECHNOLOGY | 2009-10-02 | One embodiment enables detection of MI/I and emerging infarction in an implantable system. A plurality of devices may be used to gather and interpret data from within the heart, from the heart surface, and/or from the thoracic cavity. The apparatus may further alert the patient and/or communicate the condition to an external device or medical caregiver. Additionally, the implanted apparatus may initiate therapy of MI/I and emerging infarction. | 1. A method of monitoring a heart for evidence of myocardial ischemia/infarction (MI/I), comprising: monitoring the heart using a pacemaker device implanted into a subject, the pacemaker device including a pacing electrode and a plurality of electrical sensors adapted to sense an electrogram signal, the electrical sensors being positioned at locations including the heart and vascular system, wherein only one of the electrical sensors adapted to sense an electrogram signal is positioned in the heart; anddetermining whether evidence of MI/I is present in the subject, using the electrogram signal;wherein at least some of the electrical sensors used in determining whether evidence of MI/I is present in the subject are also used to sense a need to pace the heart. 2. The method of claim 1, further comprising alerting the subject when a determination is made that there is evidence of MI/I present in the subject. 3. The method of claim 1, further comprising wirelessly communicating from the device implanted into the subject to a device external of the subject. 4. The method of claim 1, wherein the pacemaker device implanted into the subject is a first device, the method further comprising: wirelessly communicating from the first device to a second device, the second device external to the subject; andcommunicating from the second device to a third device using an internet communication scheme. 5. The method of claim 1, further comprising initiating therapy within the subject after detecting evidence of MI/I in the subject. 6. The method of claim 1, further comprising providing a supply of a medicine as part of the device implanted into the subject, and supplying a dose of the medicine to the subject after the device makes a determination that there is evidence that MI/I is present in the subject. 7. A method of monitoring a heart for evidence of myocardial ischemia/infarction (MI/I), comprising: monitoring the heart using a pacemaker device implanted into a subject, the pacemaker device including a pacing electrode and plurality of electrical sensors adapted to sense an electrogram signal, wherein all the electrical sensors adapted to sense an electrogram signal are positioned within the heart; anddetermining whether evidence of MI/I is present in the subject, using the electrogram signal;wherein at least some of the electrical sensors used in determining whether evidence of MI/I is present in the subject are also used to sense a need to pace the heart. 8. The method of claim 7, further comprising alerting the subject when a determination is made that there is evidence of MI/I present in the subject. 9. The method of claim 7, further comprising wirelessly communicating from the device implanted into the subject to a device external of the subject. 10. The method of claim 7, wherein the device implanted into the subject is a first device, the method further comprising: wirelessly communicating from the first device to a second device, the second ... | - | A61N001/362 A61B005/0402
US20100049035 | A1 | 2010-02-25 | Brain image segmentation from ct data | 2005-08-25 | The brain structure is extracted from CT data based on thresholding and brain mask propagation. Two thresholds are determined: a high threshold excludes the high intensity bones, while a low threshold excludes air and CSF. Brain mask propagation uses the spatial relevance of brain tissues in neighbouring slices to exclude non-brain tissues with similar intensities. | 1. A method for generating a segmented brain image from a 2-dimensional slice computed tomography (CT) scan data set, comprising the steps of: (a) choosing a reference slice of said CT data, and for said reference slice: determining a region of interest; determining a low threshold value from intensity values of said reference slice within said region of interest; and determining a high threshold value from intensity values of said reference slice within said region of interest; and (b) for each slice in said data set: determining a region of interest; performing a binarization of said slice components by use of said low threshold value and said high threshold value to give foreground connected components; and excluding those foreground connected components that do not satisfy a spatial relevance criterion with reference to an adjacent slice. 2. A method according to claim 1, wherein said foreground connected components are those components having an intensity value falling between said low threshold value and said high threshold value. 3. A method according to claim 2, wherein said spatial relevance criterion is based on the number of foreground connected pixels in said slice being greater than a proportion of foreground connected pixels in said adjacent slice. 4. A method according to claim 3, wherein said excluding step includes determining brain candidate components from said foreground connected components by excluding those foreground connected components that are less than a predetermined distance from the skull defined as a brain mask boundary before applying said spatial relevance criterion. 5. A method according to claim 4, wherein said head mask boundary is determined with reference to . . . (p. 8) 6. Apparatus for generating a segmented brain image, comprising: (a) a computed tomography (CT) scanner producing a CT scan data set; (b) a processor: generating 2-dimensional slice data from said data set; for a reference slice of said CT data: determining a region of interest, determining a low threshold value from intensity values of said reference slice within said region of interest, and determining a high threshold value from intensity values of said reference slice within said region of interest; and for each slice in said data set: determining a region of interest, performing a binarization of said slice components by use of said low threshold value and said high threshold value to give foreground connected components; and excluding those foreground connected components that do not satisfy a spatial relevance criterion with reference to an adjacent slice; and(c) a display device to display the non-excluded foreground connected components in each said slice as said segmented brain image. 7. Apparatus according to claim 6, wherein said processor determines said foreground connected components to be those components having an intensity value falling between said low threshold value and said high threshold value. 8. Apparatus acco... | - | A61B005/05 G06K009/00
US20100004703 | A1 | 2010-01-07 | Energy Medicine Treatment Methods | 2009-03-19 | An energy medicine treatment method includes ascertaining a health dysfunction of a patient, the patient exhibiting a head point(s) on a surface of the cranium and the head point(s) being biologically active in a manner conducive to treatment of the dysfunction. Without identifying a specific location of the head point(s), the method includes stimulating a sufficiently large portion of the cranium to clear the head point(s) and to enable the patient's body to improve proper function of the dysfunction. | 1. An energy medicine treatment method comprising: ascertaining a health dysfunction of a patient, the patient exhibiting a head point on a surface of the cranium and the head point being biologically active in a manner conducive to treatment of the dysfunction; andwithout identifying a specific location of the head point, stimulating a sufficiently large portion of the cranium to clear the head point and to enable the patient's body to improve proper function of the dysfunction. 2. The method of claim 1 wherein ascertaining the health dysfunction comprises using allergy skin scratch or blood test techniques, checking for symptoms characteristic of allergic reactions, using lactose intolerance hydrogen breath test techniques, or checking for symptoms characteristic of lactose intolerance. 3. The method of claim 1 wherein a health symptom of the patient appears to be caused by the health dysfunction and clearing the head point reduces severity of the symptom. 4. The method of claim 1 wherein clearing the head point comprises clearing multiple head points simultaneously. 5. The method of claim 1 wherein stimulating the cranium comprises touching, tapping, massaging, or using some device. 6. The method of claim 1 wherein the health dysfunction is selected from the group consisting of a fungal infection, a fungal infection in a sinus, a Candida infection, elevated metals or chemicals, a digestive enzyme deficiency, a digestive HCl deficiency, a scar/organ nerve energy disruption, muscle dysfunction, vertebral subluxation, and a posterior or postero-lateral spinal disc. 7. The method of claim 1 wherein the patient's body exhibits a health stressor point corresponding to the health dysfunction and clearing the head point comprises touching the stressor point while stimulating the cranium. 8. The method of claim 7 wherein the health dysfunction comprises a bacterial, viral, fungal, or parasitic infection, a digestive enzyme deficiency, a digestive HCl deficiency, a scar/organ nerve energy disruption, a dysbiotic root canal, or organ dysfunction. 9. The method of claim 7 further comprising ascertaining an organ with which the health dysfunction is associated by using by using MRT, the patient's body exhibiting an organ point and clearing the head point comprising touching the organ point while touching the stressor point and stimulating the cranium. 10. The method of claim 9 wherein the health dysfunction comprises a bacterial, viral, fungal, or parasitic infection, a dysbiotic root canal, or organ dysfunction. 11. The method of claim 1 further comprising ascertaining an organ with which the health dysfunction is associated by using MRT, the patient's body exhibiting an organ point and clearing the head point comprising touching the organ point while stimulating the cranium. 12. The method of claim 1 wherein ascertaining the health dysfunction comprises touching a test vial during MRT. 13. The method of claim 12 further comprising ascertainin... | - | A61N001/00 A61N002/00 A61B005/00
US20100041966 | A1 | 2010-02-18 | SLEEP ASSISTANT SYSTEM AND METHOD THEREOF | 2009-08-11 | A sleep assistant method for a sleeper in an environment is provided. The method includes the steps of monitoring a bio-condition of the sleeper to collect a bioinformation of the sleeper; performing a sleep coach mode, which analyzes the bioinformation and provides a sleep knowledge to the sleeper based on the analyzed results; and performing a sleep environment adjusting mode, which adjusts the environment based on the bioinformation. | 1. A sleep assistant method for a sleeper in an environment, comprising: monitoring a bio-condition of the sleeper to collect a bioinformation of the sleeper;performing a sleep coach mode, which analyzes the bioinformation and provides a sleep knowledge to the sleeper based on the analyzed results; andperforming a sleep environment adjusting mode, which adjusts the environment based on the bioinformation. 2. A method according to claim 1, wherein: the bioinformation includes at least one selected from a group consisting of a heart beat, a body temperature, a blood pressure, a skin conductivity and a respiration rate; andthe sleep knowledge includes at least one selected from a group consisting of a direction from a doctor, a prescription by a doctor, a sleep medical knowledge, a medical treatment knowledge, a pre-sleep action and a sleep skill. 3. A sleep environment adjusting system for a sleeper in an environment having an environment parameter, comprising: a sensor sensing the sleeper to obtain a bioinformation;an environment adjusting device disposed in the environment; andan electronic processor electrically connected with the sensor and the environment adjusting device, and controlling the environment adjusting device based on the bioinformation. 4. A system according to claim 3, wherein the sensor is disposed on the sleeper, and the electronic processor has a sleep environment adjusting mode controlling the environment adjusting device to adjust the environmental parameter based on the bioinformation and a preference of the sleeper. 5. A system according to claim 3, wherein the sensor is a biosensor, and the sleeper is one of a person ready to sleep and a person falling asleep. 6. A system according to claim 3, wherein the bioinformation comprises at least one selected from a group consisting of a heart beat, a body temperature, a blood pressure, a skin conductivity and a respiration rate. 7. A system according to claim 3, wherein the environmental parameter comprises at least one selected from a group consisting of a visual effect, an acoustic volume, a temperature, a humidity and an air condition. 8. A system according to claim 3, wherein the environment adjusting device comprises at least one selected from a group consisting of an air conditioner, an illuminating device, an audio device, a wake-up device and a combination thereof. 9. A system according to claim 3, wherein the electronic processor is one selected from a group consisting of a personal computer, a notebook computer and a server, and is electrically connected with the sensor and the environment adjusting device via one of a wireless and a wired connections. 10. A system according to claim 3, wherein the sleep environment adjusting mode is based on a fuzzy algorithm program. 11. A sleep coach apparatus for a sleeper, comprising: a sensor sensing the sleeper to obtain a bioinformation; andan electronic processor electrically connected with the sensor, and providing ... | National Taiwan University | A61B005/00
US20100049027 | A1 | 2010-02-25 | ELECTRODE BELT FOR CARRYING OUT ELECTRODIAGNOSTIC PROCEDURES ON THE HUMAN BODY | 2009-10-29 | An electrode belt is provided for carrying out electrodiagnostic procedures on the human body. The electrode belt includes a belt (2), which is formed at least partially of an elastic material and surrounds the body of a test subject. A plurality of electrodes, are mechanically connected to the belt and are in flat contact with the body of the test subject. At least one contact passes from the electrodes through the belt (2) and is connected to a connection element (11), which is connected to a lead each of a multicore cable (6). Such an electrode belt offers marked advantages in terms of handling along with increased safety from movement artifacts. | 1. An electrode belt for carrying out electrodiagnostic procedures on the human body, the electrode belt comprising: a plurality of connection elements;a multicore cable, said plurality of connection elements being positioned along said multicore cable, one connection element being located at a spaced location from another connection element, each connection element being connected to a lead of said multicore cable;a belt formed at least partially of an elastic material and having a length for surrounding the body of a test subject, said belt being movable from a relaxed state to a stretched state, said belt having a first length in said relaxed state, said belt having a second length in said stretched state, said second length being greater than said first length;a plurality of electrodes connected to the belt for being in flat contact with the body of the test subject, each electrode being positioned along said belt at a spaced location from an adjacent electrode, said elastic material being provided between adjacent electrodes with each electrode moving relative to said adjacent electrode when said belt moves from said relaxed state to said stretched state, whereby said electrodes carry out an electroimpedance tomography procedure in said stretched state; anda plurality of contact means each for mechanically and electrically connecting and mechanically and electrically disconnecting from a respective one of said connection elements, each contact means extending from a respective one of said electrodes through said belt to a position located outside of said belt, each of said connection elements having a retaining means for mechanically and electrically connecting to one of said contact means, said retaining means being flexible to generate a snap in retaining function as said contact means is inserted into said retaining means, whereby each contact means is mechanically and electrically connected to said multicore cable via said retaining means. 2. An electrode belt in accordance with claim 1, wherein said multicore cable has a multicore cable length, said multicore cable length being thirty percent greater than said first length. 3. An electrode belt in accordance with claim 1, wherein said multicore cable has a plurality of multicore cable portions, each multicore cable portion extending between one of said connection elements and an adjacent connection element, each multicore cable portion having a multicore cable portion length, each of said electrodes and an adjacent electrode defining a first distance between electrodes in said relaxed state, said first distance between electrodes being less than said multicore cable portion length. 4. An electrode belt in accordance with claim 3, wherein one or more of said multicore cable portions extends in a meandering pattern when said belt is in said relaxed state, one or more of said multicore cable portions extending substantially parallel to said belt when said belt is in said stretched state... | Drager Medical AG & Co. KG | A61B005/04
US20100049022 | A1 | 2010-02-25 | Methods for improving performance and reliability of biosensors | 2009-10-19 | The present invention relates to a predictive-kinetic method for use with data processing of a sensor-generated signal, as well as, microprocessors and monitoring systems employing such a predictive-kinetic method. Data from a transient region of a signal is used with suitable models and curve-fitting methods to predict the signal that would be measured for the system at the completion of the reaction. The values resulting from data processing of sensor response using the methods of the present invention are less sensitive to measurement variables. | 1. A method for measuring glucose present in a subject, said method comprising: (A) transdermally extracting a sample comprising glucose from the subject using a sampling system that is in operative contact with a skin or mucosal surface of said subject;(B) obtaining a measured signal over time, comprising a measured signal response curve, from the extracted glucose, wherein said measured signal is specifically related to the amount or concentration of glucose, and said measured signal response curve comprises kinetic and equilibrium regions;(C) using (i) a mathematical model comprising selected parameters, wherein said model describes the measured signal response curve, and said mathematical model is selected from the group consisting of a first order process, combined first order and zero order process, a parallel multiple first order process, a flux process, and an nth order process, and (ii) an error minimization method, to iteratively estimate values of the parameters using said model and error minimization method to fit a predicted response curve to said measured signal response curve, wherein (a) the error minimization method provides a calculated error based on differences between said predicted and measured signal response curves, and (b) said estimating is iteratively performed until the calculated error between the predicted and measured signal response curves falls within an acceptable range or until no further statistically significant change is seen in the calculated error, at which time iterative estimation of the parameters is stopped, said iterative estimation and error minimization results in a predicted response curve corresponding to said measured signal response curve, said predicted response curve yields a predicted end-point value and a measurement correlated to the amount or concentration of the glucose. 2. The method of claim 1, wherein said measured signal response curve comprises a measurement of current over time, or measurement of charge over time. 3. The method of claim 2, wherein said measured signal response curve comprises a measurement of current over time, said predicted end-point value is an estimated signal at equilibrium, where said predicted end-point value provides a predicted final background value, and said measurement correlated to the amount or concentration of glucose corresponds to an area under the predicted response curve. 4. The method of claim 3, wherein said area under the predicted response curve is obtained by integration of the predicted response curve. 5. The method of claim 4, wherein before said integration is performed said final background value is used to perform a background subtraction correction of the predicted response curve and said measurement correlated to the amount or concentration of glucose corresponds to an area under the predicted response curve. 6. The method of claim 4, wherein the end-point value of the integrated predicted response curve is converted to an amount o... | Animas Technologies, LLC. | A61B005/1477 A61B005/145 G06F019/00
US20100049010 | A1 | 2010-02-25 | METHOD AND DEVICE FOR MEASURING PHYSIOLOGICAL PARAMETERS AT THE WRIST | 2009-10-01 | A wrist-mounted device for measuring at least one physiological parameter of a subject. The present invention enables such a measurement to preferably be transformed into clinically useful information about the subject. Such information may then optionally be sent to medical personnel, for example at a contact and/or monitoring center, through a gateway device. The gateway device preferably communicates with the wrist-mounted device of the present invention through a wireless communication channel. | 1-92. (canceled) 93. A device for measuring at least one physiological parameter of a subject, comprising: (a) a fastening article for being fastened to the wrist of a first hand of the subject; and(b) a measuring device for measuring at least one physiological parameter of the subject, said measuring device configured to be attached to the wrist by said fastening article, said measuring device comprising: (i) at least one sensor for measuring at least oxygen saturation (SpO2) in the blood of the subject;(ii) a processor for receiving a signal from said sensor and for converting said at least one measurement to form medical information; and(iii) a communication unit for at least receiving said medical information from said processor and for at least transmitting said medical information;wherein said measuring device is configured to be attached to the wrist of a first hand of said subject and wherein said sensor is configured to be touched by a finger of a second hand of said subject to perform said SpO2 measurement. 94. A device for measuring at least one physiological parameter of a subject according to claim 93 wherein said sensor comprises two light sources and a photoelectric detector. 95. A device for measuring at least one physiological parameter of a subject according to claim 93 wherein said light sources are LEDs (light emitting diodes). 96. A device for measuring at least one physiological parameter of a subject according to claim 93 wherein one of said light sources emits infrared band light and the other emits red-band light. 97. A device for measuring at least one physiological parameter of a subject according to claim 93 wherein light from said light sources is reflected from said subject onto said photoelectric detector, and wherein said photoelectric detector converts said light to an electronic signal processed by said processor. 98. A device for measuring at least one physiological parameter of a subject according to claim 93 further comprising a flexible support for supporting said sensor. 99. A device for measuring at least one physiological parameter of a subject according to claim 93 wherein further comprising flexible, opaque material for surrounding said sensor. 100. A device for measuring at least one physiological parameter of a subject according to claim 93 additionally comprising a switch mounted on said measuring device and configured to switch on and off at least one of said measuring device, said sensor, said processor and said communication unit. 101. A device for measuring at least one physiological parameter of a subject according to claim 100 wherein said switch is mounted beside said sensor. 102. A device for measuring at least one physiological parameter of a subject according to claim 100 wherein said switch is mounted beneath said sensor and said sensor and switch are configured to operate said switch when said sensor is touched by said finger of said subject. 103. A device for measuring at least... | - | A61B005/00
US07658716 | B2 | 2010-02-09 | Vital signs monitor using an optical ear-based module | 2004-12-07 | The invention provides a system for measuring blood pressure from a patient that includes: 1) an optical module configured to be worn on the patient's ear and comprising at least one optical source and a photodetector; 2) a calibration source configured to make a blood pressure measurement; and, 3) a processing module configured to: i) receive a first signal from the optical module; ii) receive a second signal from the calibration source; iii) process the first and second signals to generate a calibration table; and iv) receive a third signal from the optical module and compare it to the calibration table to determine the patient's blood pressure. | 1. A system for measuring blood pressure from a patient, comprising: an optical module comprising at least one optical source and a photodetector configured to measure a measurement signal;a processing module configured to receive the measurement signal and periodically determine a blood pressure according to a first pre-determined time period by fitting the measurement signal to a mathematical model and, in response, iteratively varying parameters of the mathematical model until it best matches the measurement signal, said processing module further configured to determine a blood pressure calibration signal according to a second pre-determined time period; anda wireless transmitter in electrical communication with the processing module and configured to periodically transmit the blood pressure to a host computer for display and the calibration signal to the host computer for display. 2. The system of claim 1, wherein the system further comprises an ear-worn clip that includes the optical module. 3. The system of claim 2, wherein the ear-worn clip comprises a first LED that emits red radiation, a second LED that emits infrared radiation, and the photodetector. 4. The system of claim 1, wherein the calibration source is a cuff-based blood pressure module that includes a cuff and a pump. 5. The system of claim 4, wherein the cuff is configured to be worn around the patient's arm. 6. The system of claim 1, wherein the wireless transmitter operates on a protocol based on 802.11, 802.15.4, part-15, or a derivative thereof. 7. The system of claim 1, further comprising a host computer comprising a wireless receiver configured to wirelessly receive the blood pressure and the blood pressure calibration signal from the processing module. 8. The system of claim 7, wherein the host computer is configured to cause the blood pressure and the blood pressure calibration signal to be displayed. 9. A system for monitoring a patient's blood pressure values, the system comprising: an optical module comprising at least one optical source, a photodetector and a transmitter for transmitting a measurement signal from the photodetector;a calibration source comprising a cuff-based blood pressure monitor, a first circuit for receiving the signal from the optical module, a second circuit for controlling the cuff-based blood pressure monitor and a wireless transmitter for periodically transmitting a cuffless-based signal from the first circuit and for periodically transmitting a cuff-based signal from the second circuit, the first circuit having a microprocessor programmed to control the second circuit and process the measurement signal from the optical module to periodically determine according to a first, pre-programmed time period a cuffless-based blood pressure value for the patient and to periodically determine according to a second pre-programmed time period a blood pressure calibration signal, the microprocessor programmed to determine the cuffless-based blood... | Triage Wireless, Inc. | A61B005/00
US07645276 | B2 | 2010-01-12 | System for non-invasive heart treatment | 2005-03-14 | The invention provides a non-invasive method for treatment of arrhythmia. In a first aspect, a method for treatment of atrial fibrillation in a heart of a patient comprises directing radiation from outside the patient toward one or more target treatment regions of the heart so as to inhibit the atrial fibrillation. The radiation may induce isolation of a pulmonary vein. | 1. A system for treatment of atrial fibrillation in a heart of a patient, the system comprising: a source of particle beam radiation or x-ray radiation; anda robotic system, wherein the robotic system comprises a computer processor having means for directing the particle beam radiation or x-ray radiation from outside the patient toward one or more target treatment regions in the heart during use, the means for directing comprising means for treating the atrial fibrillation, the robotic system coupled to the source. 2. The system of claim 1, further comprising an imaging system for imaging a first pulmonary vein coupled to the processor, wherein the means for directing the radiation of the processor is configured for targeting the radiation from outside the patient so as to induce isolation of the first pulmonary vein. 3. The system of claim 2, wherein the processor is configured to target the radiation so that it effects bilateral pulmonary vein isolation. 4. The system of claim 2, wherein the processor is configured to target at least one ectopic focus of the heart for treatment with the radiation. 5. The system of claim 2, wherein the processor is configured to target the radiation so that the radiation ablates linear lesions within the heart to inhibit an ectopic electrical pathway, the lesions corresponding to a Maze lesion pattern. 6. The system of claim 2, wherein the processor is configured so that a first portion of the radiation is directed to the patient in a first treatment, and a second portion of the radiation is directed to the patient in a second treatment so that the atrial fibrillation is inhibited in a fractionated radiation procedure allowing assessing of the patient between the first treatment and the second treatment. 7. The system of claim 2, wherein the processor is configured to direct the radiation so as to avoid exceeding a tolerance of tissues adjacent the one or more target regions. 8. The system of claim 1, wherein the robotic system further comprises a robotic arm movably supporting the source and the processor coupled to the robotic arm, the processor configured to effect delivery of the radiation as a series of radiation beams extending toward the heart and movement of the robotic arm between beams so that the beams are directed toward the heart from different angles. 9. The system of claim 8, further comprising an imaging system coupled to the processor, wherein the processor is configured to dynamically register the radiation beams with the one or more regions of the heart using the imaging system. 10. The system of claim 9, wherein the processor is configured to dynamically register so as to compensate for pumping of the heart, movement of the patient, or breathing of the patient. 11. The system of claim 9, wherein the processor is configured to acquire a plurality of target adjustment images from the imaging system during a radiation treatment procedure, and to determine a position of the one or more ... | CyberHeart, Inc. | A61B018/04 A61N005/10 A61B005/05
US07657305 | B2 | 2010-02-02 | Implantable medical device for improved storage of intracardiac electrograms | 2006-11-27 | An implantable medical device provides for improved storage of recorded IEGMs. A sensing stage is connected to an electrode for picking up electric potentials from inside a heart, the time course of said electric potentials representing a heart signal, a control unit connected to said sensing stage is adapted to process a sequence of data points that each represent an amplitude or magnitude A of a time-varying signal at equidistant points of time t, wherein end points of data segments are determined by processing of the sequence of data points. The control unit is adapted to identify end points of data segments by processing of the sequence of data points. | 1. An implantable medical device comprising: at least a sensing stage connected or being connectable to an electrode for picking up electric potentials from inside a heart, the time course of said electric potentials representing a heart signal, said heart signal being sampled in constant sampling intervals anda control unitthat is connected to said sensing stage andthat is adapted to process a sequence of data points that each represent a magnitude A of a time-varying signal at equidistant points of time t, said processing of the sequence of data points includes:determining end points of data segments by a) determination of a first difference quotient D1=(An−An-1)/(tn−tn-1) with respect to an actual data point An at tn and an immediately preceding, second last data point An-1 at tn-1,b) determination of a second difference quotient D2 (An−Ae)/(tn−te) or D2=(An-1−Ae)/(tn-1−te) with respect to an actual data point An at tn or a second last data point An-1 and tn-1, respectively, and a last end point Ae at te, said last end point Ae being a previously determined end point or a first data point of said sequence of data points, wherein te represents the point of time belonging to said end point Ae,c) selecting the second last data point An-1 as a new end point, if the magnitude of the difference between the two difference quotients DD=D1−D2 exceeds a predetermined first threshold value T, DD>T, anddetermining a data segment length L=tn−te between the point of time tn of an actual data point An and the point of time te of a last end point Ae,and selecting An-1 as a new end point, if said data segment length L exceeds a predetermined maximum length Lmax,and storing said selected end points in association with a segment length between each stored endpoint and an immediately preceding end point as a compressed data representation of said time varying signal. 2. The implantable medical device of claim 1, wherein the control unit is further adapted to: determine whether said first and said second difference quotient are of opposite sign andselect the second last data point An-1 as a new end point, if the magnitude of the difference between the two difference quotients DD=D1−D2 exceeds a predetermined second threshold value T2. 3. The implantable medical device of claim 1 or 2, wherein the control unit is further adapted to: determine whether a data segment length L=tn−te between the point of time tn of an actual data point An and the point of time te of a last end point Ae exceeds a predetermined threshold length L1 and whether the magnitude of the difference between the two difference quotients DD=D1−D2 exceeds a predetermined third threshold value T3,and to select the second last data point An-1 as a new end point, if said two conditions are met. 4. The implantable medical device of claim 3, wherein the control unit is adapted to determine the first, and, if applicable, the second and the third threshold value as a predetermined fraction of a peak amp... | BIOTRONIK CRM Patent AG | A61B005/04 A61B005/0402
US08317729 | B2 | 2012-11-27 | Cervical dilation meter | 2009-08-18 | An instrument for measuring cervical dilation can have a pair of arms connected at their proximal ends to an arm pivot or articulating member, the arms being in movable communication with a gauge assembly for measuring the relative distance between the arms at a fixed location near the proximal ends of the arms. The arms can be disposed to apply an outward lateral pressure against the walls of the cervix, thereby engaging the cervix without the need for physical penetration, gripping, or other attachment of the device to the cervical tissue. Continuous outward lateral pressure of the arms against the cervical walls can allow the arms to expand in response to and in concert with expansion and dilation of the cervix. The relative distance between the arms correlates to the diameter of the cervix, such that the correlated measurement indicated on a scale of the gauge means is the measurement of cervical dilation. | 1. An apparatus comprising: first and second arms, comprising respective proximal and distal portions and intermediate regions therebetween, the proximal portions of the first and second arms coupled together, the distal portions of the first and second arms configured to be inserted between opposing lateral walls of a cervix or vagina and measure a cervical dilation, and the intermediate region of the first arm comprising an outwardly bowed first cephalic curve and the intermediate region of the second arm comprising an outwardly bowed second cephalic curve opposing the first cephalic curve,the first and second arms configured to exert an outward force against the opposing lateral walls in an amount sufficient to hold the apparatus in position while measuring cervical dilation, without requiring physical penetration, attachment, or gripping of the cervix or vagina,a concave portion of the first cephalic curve and a concave portion of the second cephalic curve sized and shaped to receive and accommodate a portion of a fetal head therebetween; anda cervical dilation indicator assembly, communicatively coupled to the first and second arms to receive information about the cervical dilation, and comprising an external cervical dilation indicator to provide an indication of the cervical dilation to a user. 2. The apparatus of claim 1, wherein the outward force is of an amount that does not substantially affect the cervical dilation measurement. 3. The apparatus of claim 1, wherein the concave portions of the first and second cephalic curves are sized and shaped to receive a portion of a descending fetal head therebetween during birthing while the first and second arms continue to exert the outward force against opposing lateral walls of the cervix or vagina to hold the apparatus in position while measuring cervical dilation, without requiring physical penetration, attachment, or gripping of the cervix or vagina. 4. The apparatus of claim 1, wherein the first and second arms comprise respective first and second pelvic curves at or near a location between the intermediate regions and proximal portions of the respective first and second arms, such that the respective intermediate regions of the respective first and second arms angle or curve upward from a plane formed by the respective proximal portions of the respective first and second arms at an angle of about 15 degrees to allow placement of the apparatus if the cervix is in a mid or anterior position. 5. The apparatus of claim 1, comprising a spring that is coupled to the first and second arms to provide the outward force and to bias the first and second arms away from each other and against opposing lateral walls of the cervix or vagina. 6. The apparatus of claim 5, wherein the spring is configured to exert the outward force in an amount sufficient to bias the first and second arms against opposing lateral walls of the cervix or vagina to hold the apparatus in position while measuring cervical... | Glenveigh Medical, LLC | A61B005/103 A61B005/117
US08876755 | B2 | 2014-11-04 | Closed loop control system interface and methods | 2009-07-14 | Method and apparatus including calling, retrieving and/or initiating a programmed function in conjunction with execution of one or more commands related to a closed loop control algorithm, receiving one or more data in response to the one or more commands over a data interface, and executing the one or more commands related to the closed loop control algorithm based on the received one or more data are provided. | 1. A system, comprising: a control unit including a memory unit having stored therein a closed loop control algorithm for execution; andan insulin delivery device in signal communication with the control unit for executing one or more medication delivery functions based on one or more signals received from the control unit; wherein the control unit includes a user interface for initiating one or more application programming interface functions associated with one or more of the operations of the insulin delivery device;wherein the one or more application programming interface functions are initialized at a serializer, and the closed loop control algorithm is integrated with the insulin delivery device and is compatible with both the control unit and the insulin delivery device without the closed loop control algorithm being modified; andwherein the insulin delivery device is configured to execute the one or more medication delivery functions associated with the one or more of the initiated application programming interface functions. 2. The system of claim 1, wherein the closed loop control algorithm stored in the memory unit of the control unit includes a plurality of closed loop control algorithms. 3. The system of claim 1, wherein the control unit includes a deserializer. 4. The system of claim 1, wherein the one or more application programming interface functions are related to one or more of analyte monitoring, a frequency of analyte data logging, a diagnosis of an analyte sensor operation, at least one failure mode related to an analyte sensor, and an analyte sensor replacement schedule. 5. The system of claim 1, wherein the one or more application programming interface functions are related to one or more of a medication delivery profile, a medication delivery modification, a safety shut off routine, operational status monitoring, and a data processing mode. 6. The system of claim 1, further comprising an analyte monitoring device in signal communication with the control unit. 7. The system of claim 1, wherein the serializer is configured such that the closed loop control algorithm can be replaced or modified without modifying the control unit or the insulin delivery device. | Abbott Diabetes Care Inc. | A61M031/00 A61B005/145 A61B005/00
US08747315 | B2 | 2014-06-10 | Signal processing for continuous analyte sensor | 2009-09-23 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A method for estimating an analyte value from a continuous analyte sensor, the method comprising: receiving a data stream from the continuous analyte sensor for a first time period, the data stream indicative of a measured analyte concentration of the host over the first time period;estimating, using a processor module, at least one analyte value for a second time period based on the data stream, wherein the step of estimating at least one analyte value comprises selecting an algorithm from a plurality of algorithms based at least in part on an analysis of a rate of change of the measured analyte concentration; andproviding output based on the estimated at least one analyte value. 2. The method of claim 1, the step of selecting an algorithm is additionally based on one or more of an analyte concentration, an acceleration, and an individual historical pattern of the data stream. 3. The method of claim 1, wherein the step of selecting an algorithm comprises employing a plurality of algorithms on the data stream and determining which of the plurality of algorithms best correlates with the data stream. 4. The method of claim 1, wherein the plurality of algorithms comprise an algorithm selected from the group consisting of polynomial regression, autoregressive algorithms, Fourier transform, Wavelet transform, neural network-based mapping, fuzzy logic based pattern matching, and Genetic-Algorithms based pattern matching. 5. The method of claim 1, wherein the step of selecting an algorithm further comprises applying a physiological boundary to the selected algorithm. 6. The method of claim 1, further comprising a step of evaluating a selected algorithm by applying an evaluation function prior to employing the selected algorithm to estimate the analyte value, wherein the evaluation function is selected from the group consisting of a data association function, a curvature formula, and a physiological boundary. 7. The method of claim 1, wherein the second time period is a future time period. 8. A system for estimating an analyte value from a continuous analyte sensor, the system comprising: an input module operatively connected to the continuous analyte sensor that receives a data stream for a first time period, the data stream indicative of a measured analyte concentration of the host over the first time period; anda processor module comprising programming configured to select an algorithm from a plurality of algorithms based on an analysis of an individual historical pattern of the data stream, estimate at least one analyte value using the selected algorithm, and initiate output based on the estimated at least one analyte value. 9. The system of claim 8, wherein the programming configured to select an algorithm from a plurality of algorithms selects the algorithm additionally based on one or more of an analyte concentration, a rate of change, and an acceleration of the data stream. 10. The system of claim 8, wherein the programming configured... | DexCom. Inc. | A61B005/00 A61B005/145
US08478539 | B2 | 2013-07-02 | System and method for neurological activity signature determination, discrimination, and detection | 2009-05-14 | A system and method are provided for automatically correlating neurological activity to a predetermined physiological response. The system includes at least one sensor operable to sense signals indicative of the neurological activity, and a processing engine coupled to the sensor. The processing engine is operable in a first system mode to execute a simultaneous sparse approximation jointly upon a group of signals sensed by the sensor to generate signature information corresponding to the predetermined physiological response. The system further includes a detector coupled to the sensors, which is operable in a second system mode to monitor the sensed signals. The detector generates upon selective detection according to the signature information a control signal for actuating a control action according to the predetermined physiological response. | 1. A system for automatically correlating neurological activity to a predetermined physiological response comprising: at least one sensor operable to sense signals indicative of the neurological activity;a processing engine coupled to said sensor, said processing engine in a first system mode executing a simultaneous sparse approximation comprising Simultaneous Matching Pursuits, jointly upon a group of signals sensed by said sensor to generate signature information corresponding to the predetermined physiological response; and,a detector coupled to said sensors, said detector in a second system mode monitoring the sensed signals and selectively generating according to said signature information a control signal for actuating a control action according to the predetermined physiological response. 2. The system as recited in claim 1, wherein said sensor includes a transducer applied to a subject to acquire electrical signals indicative of the neurological activity. 3. The system as recited in claim 2, further comprising a transducer applied to the subject to acquire electrical muscle activity indicative of the physiological response. 4. The system as recited in claim 1, wherein said processing engine in said first system mode executes Greedy Adaptive Discrimination (GAD) processing upon the group of sensed signals. 5. The system as recited in claim 4, wherein the sensed signals in a group of sensed signals are variably aligned in time. 6. The system as recited in claim 5, further comprising a behavioral cueing unit prompting the physiological response of a subject. 7. The system as recited in claim 6, further comprising a behavioral response detector unit detecting the physiological response of a subject. 8. The system as recited in claim 4, wherein said processing engine generates said signature information based upon a parametric mean representation defined in a multi-dimensional parametric space, said parametric mean representation including a plurality of parametric mean components each independently representing a mean value within one parametric space dimension. 9. The system as recited in claim 4, further comprising an actuation interface unit coupled to the detector for performing the control action responsive to the control signal. 10. A brain-computer interfacing system for automatically correlating neurological activity of a subject to a predetermined physiological response comprising: at least one transducer sensing signals indicative of the neurological activity;a processing engine coupled to said transducer, said processing engine in a system training mode executing a simultaneous sparse approximation comprising Simultaneous Matching Pursuits, upon a collection of signals sensed by said transducer to generate signature information corresponding to the predetermined physiological response; and,a detector coupled to said transducer, said detector in a system utilization mode monitoring the sensed signals and generating upon de... | - | G01N033/48 A61B005/04
US08775093 | B2 | 2014-07-08 | Pattern recognition system for classifying the functional status of patients with pulmonary hypertension, including pulmonary arterial and pulmonary vascular hypertension | 2009-09-25 | A method employing pattern recognition techniques for identifying the functional status of patients with Pulmonary Hypertension is described. This method describes a process by which sets of cardiopulmonary exercise gas exchange variables are measured during rest, exercise and recovery and stored as unique data sets. The data sets are then analyzed by a series of feature extraction steps, yielding a multi-parametric index (MPIPH) which reflects the current functional status of a patient. The method also employs a description scheme that provides a graphical image that juxtaposes the measured value of MPI to a reference classification system. An additional description scheme provides a trend plot of MPI values measured on a patient over time to provide feedback to the physician on the efficacy of therapy provided to the patient. The method will enable physicians to gather, view, and track complicated data using well-understood visualization techniques to better understand the consequences of their therapeutic actions. | 1. A method of pattern recognition for diagnosing the presence of chronic pulmonary hypertension (PH) and evaluating patients with that condition, comprising: (a) classifying the functional status of patients with a chronic condition by conducting a course of evaluation using a system for providing a pattern that characterizes said functional status using a calculated multiparametric set of data that reflects the status of a patient ((multiparametric index (MPIPH));(b) obtaining a plurality of measurements, including a plurality of weighted individual Ranking Parameters (RP), from a patient, said measurements being taken a plurality of times during the course of evaluation wherein said evaluation includes using a cardiopulmonary exercise gas exchange analyzer to perform sub-maximal cardiopulmonary exercise testing bouts;(c) using a computing device to process data obtained in said plurality of measurements each time said measurements are taken to determine a plurality of weighted RPs by multiplying each by a selected weighting factor to obtain the MPIPH and obtain a statistical pattern of MPIPH values based on said measurements being taken a plurality of times to thereby determine the ongoing functional status of said patient;(d) wherein said MPIPH is based on the sum of a plurality of said individual RPs combined with weighing factors (W) and wherein said plurality of RPs are selected at least in part from end tidal CO2 (ET CO2) cardiopulmonary exercise test related measurements;(e) classifying said functional status based on the MPIPH; and(f) wherein calculated statistical values consider a normal value (NV) and cutoff point (COP);(g) wherein said cardiopulmonary exercise bouts, comprise: performing a gas exchange test on a patient, wherein the gas exchange test includes a rest phase, an exercise phase, and a recovery period, wherein the patient performs submaximal or maximal exercise during the exercise phase;obtaining a plurality of physiological measurements, including end tidal CO2 on a breath-by-breath basis during the gas exchange test;determining a first value associated with an average end tidal CO2 value during at least a portion of the rest phase;determining a second value associated with a rise at the beginning of the exercise phase;determining a third value associated with the first end tidal CO2 value that is less than the first value;determining a fourth value associated with nadir of the exercise PetCO2;determining a fifth value that is associated with the last 30 seconds of exercise;determining a sixth value that is associated with the last end tidal CO2 value of recovery;determining a seventh value that is determined by subtracting the second value from the first valuedetermining an eighth value that is obtained by subtracting the time at which the first value occurred from the time at which the third value occurred;determining a ninth value that is associated with the maximum drop in end tidal CO2 obtained by computing the slo... | Shape Medical Systems, Inc. | G06F019/00 A61B005/083 A61B005/091 A61B005/0205 A61B005/00
US09545518 | B2 | 2017-01-17 | Posture state classification for a medical device | 2009-04-30 | Techniques for posture classification of a patient in a coordinate system of a sensor. According to one aspect, a defined vector is obtained from a sensor disposed in a substantially fixed manner relative to the patient. The defined vector is described in a coordinate system of the sensor and without regard to an orientation in which the sensor is disposed in relation to the patient. A detected vector is obtained from the sensor that is described using the coordinate system of the sensor. The detected vector and the defined vector to are used to classify the posture state of the patient without regard to the orientation in which the sensor is disposed in relation to the patient. A response may be initiated by a medical device, which may include adjusting therapy delivery. | 1. A method of classifying a posture state of a patient, comprising: obtaining a defined vector from at least one sensor disposed in a substantially fixed manner relative to the patient, the defined vector being described in a coordinate system of the at least one sensor and without regard to an orientation in which the sensor is disposed in relation to the patient;obtaining a detected vector from the at least one sensor, the detected vector being described in the coordinate system of the at least one sensor and being indicative of the posture state of the patient;comparing the detected vector and the defined vector;classifying the posture state of the patient based on the comparison and without regard to the orientation in which the sensor is disposed in relation to the patient; andinitiating via a medical device an action related to providing care for the patient, the action being based on the posture state classification. 2. The method of claim 1, wherein obtaining the defined vector includes: measuring outputs of the at least one sensor while the patient assumes a defined posture state; andassociating the measured outputs with the defined vector. 3. The method of claim 1, further comprising associating a region in space with the defined vector, and wherein comparing the detected vector and the defined vector further comprises determining whether the detected vector lies within the region in space. 4. The method of claim 3, further comprising: obtaining multiple defined vectors, each being associated with a respective region in space; andwherein comparing the detected vector and the defined vector comprises determining whether the detected vector lies within any of the regions in space associated with the multiple defined vectors. 5. The method of claim 1, further comprising: associating a tolerance with the defined vector; anddetermining a similarity between the defined vector and the detected vector; andwherein comparing the detected vector and the defined vector comprises comparing the similarity to the tolerance. 6. The method of claim 5, wherein the similarity is a distance, and further comprising: mapping the distance to a value that is directly proportional to a degree of similarity between the defined vector and the detected vector; andwherein comparing the detected vector and the defined vector comprises comparing the value to the tolerance. 7. The method of claim 1, further comprising: obtaining multiple defined vectors;associating each of the multiple defined vectors with a respective tolerance;interrelating multiple ones of the tolerances; andwherein comparing the detected vector and the defined vector comprises comparing the detected vector to a region in space described by the interrelated multiple ones of the tolerances. 8. The method of claim 7, wherein interrelating multiple ones of the tolerances comprises using one or more logical functions. 9. The method of claim 7, wherein the region in space is defined by a toroid... | MEDTRONIC, INC. | A61N001/00 A61N001/365 A61B005/11 A61B005/00
US20100016756 | A1 | 2010-01-21 | PRESSURE SENSOR OPERATING IN A FLUID MEDIUM | 2008-07-20 | A pressure sensor and a method of measuring pressure signals in a fluid medium. The pressure sensor comprises a sensing element for transducing pressure signals to electric signals and a mechanical amplifier connected to the sensing element, comprising an immersed appendage and a transmission element. The sensing element is enclosed within a case, and the mechanical amplifier is structured to seal the case and isolate the sensing element from the fluid medium. Pressure signals in the fluid medium cause movements in the appendage that are transmitted via the elongated transmission element to the sensing element that is isolated from the fluid medium. The method comprises the stages: Sensing the pressure signal and transmitting it via a transmission element to a piezoelectric sensing element inside the measuring apparatus. | 1. A pressure sensor operating in a fluid medium, said sensor comprising: a sensing element for transducing a pressure signal to an electric signal, said sensing element enclosed within a case;a mechanical amplifier operatively connected to said sensing element and structured to seal said case and isolate said sensing element from said fluid medium, said mechanical amplifier comprising: an appendage connected externally to said case and immersed in said fluid medium,an elongated transmission element operatively connecting said appendage to said sensing element, such that said pressure signal is transmitted from said appendage to said sensing element;wherein pressure signals in said fluid medium cause movements in said appendage that are transmitted via said elongated transmission element to said sensing element that is isolated from said fluid medium. 2. The pressure sensor of claim 1, wherein said sensing element comprises: an upper plate connected to said elongated transmission element,a piezoelectric plate connected to said upper plate,wherein said upper plate with said elongated transmission element are arranged to amplify the pressure signal to allow an optimal deformation of said piezoelectric plate. 3. The pressure sensor of claim 2, wherein said sensing element further comprising: a lower metal plate connected to said piezoelectric plate such that at least a portion of said piezoelectric plate is enclosed between the upper plate and the lower metal plate,at least one upper connection to said upper plate and at least one lower connection to said lower metal plate,wherein the conductivity of said piezoelectric plate is measured between said upper and lower connections. 4. The pressure sensor of claim 1, wherein said elongated transmission element comprises a pin. 5. The pressure sensor of claim 1, wherein said elongated transmission element comprises a plurality of transmission elements interconnected by joints, wherein said plurality of transmission elements transmit said pressure signal to said sensing element in a range of intensities allowing an optimal analysis of said electric signal. 6. The pressure sensor of claim 1, wherein said appendage is shaped as a plate. 7. The pressure sensor of claim 1, wherein said appendage is shaped to the form of a cap. 8. The pressure sensor of claim 7, wherein said appendage partly surrounding and in close proximity to the front end of said case. 9. The pressure sensor of claim 1, wherein said appendage is shaped as two interconnected parallel plates. 10. The pressure sensor of claim 1, wherein said appendage is an integral part of said case, resulting in sealing said case. 11. The pressure sensor of claim 10, wherein said appendage comprises an inner notch and the distal end of said elongated transmission element is arranged to fit in said notch, such that movements of said appendage are transmitted to said elongated transmission element. 12. The pressure sensor of claim 1, wherein said ap... | - | A61B005/03
US20100022891 | A1 | 2010-01-28 | INTRALUMINAL SPECTROSCOPE WITH WALL CONTACTING PROBE | 2009-09-28 | A spectroscope for detecting vulnerable plaque within a lumen defined by an intraluminal wall includes a probe through which an optical fiber extends. An coupler in optical communication with the fiber is configured to atraumatically contact the intraluminal wall. A light source provides light to the fiber for illuminating the wall and a detector coupled to the fiber receives light from within the wall. | 1-32. (canceled) 33. An apparatus for detecting vulnerable plaque within a lumen defined by an intraluminal wall, the apparatus comprising: a cannula having a longitudinal axis;a plurality of probes extending through the cannula, each probe having an optical fiber extending therethrough, andan atraumatic light-coupler in optical communication with the optical fiber, the coupler being configured to atraumatically contact the intraluminal wall. 34. The apparatus of claim 33, further comprising a spacer ring attached to each of the probes for maintaining the positions of the probes relative to each other. 35. The apparatus of claim 33, further comprising a hub attached to a distal end of each of the probes. 36. The apparatus of claim 35, wherein the distal end of the probe is attached to the hub at an anchor point that is circumferentially offset from a proximal portion of the probe. 37. The apparatus of claim 35, further comprising a spacer ring attached to each of the probes for maintaining the positions of the probes relative to each other. 38. The apparatus of claim 35, wherein each of the probes resiliently assumes a bow shape having a point of inflection between the hub and the cannula. 39. The apparatus of claim 33, wherein each of the probes resiliently assumes a desired shape. 40. The apparatus of claim 33, wherein the atraumatic coupler comprises means for providing optical communication between the optical fiber and the intraluminal wall. 41. The apparatus of claim 33, wherein at least one of the plurality of probes is integral with the cannula. 42. The apparatus of claim 33, wherein the optical fiber is embedded within the cannula. 43-55. (canceled) 56. The apparatus of claim 36, further comprising a control wire, wherein the control wire passes through the cannula, is coupled to the hub, and is operable to deploy the probes or undeploy the probes. 57. The apparatus of claim 33, further comprising a guidewire lumen within the cannula, said lumen configured to accommodate a guidewire for enabling the cannula to be guided to a region of interest. 58. The apparatus of claim 33, further comprising: a light source in optical communication with the optical fibers; anda detector in optical communication with at least one of the optical fibers, said detector configured to detect a light signal of a corresponding atraumatic light coupler. 59. An apparatus for detecting vulnerable plaque within a lumen defined by an intraluminal wall, the apparatus comprising: a cannula having a longitudinal axis;a plurality of probes extending through the cannula, each probe having an optical fiber extending therethrough, andan atraumatic light-coupler in optical communication with the optical fiber, thecoupler being configured to atraumatically contact the intraluminal wall; anda hub attached to a distal end of each of the probes;wherein, in a deployed state, each of the probes resiliently assumes a bow shape having a point of inflection between t... | INFRAREDX, INC. | A61B005/00
US20100015144 | A1 | 2010-01-21 | Methods for dosing an activin-actriia antagonist and monitoring of treated patients | 2009-06-26 | In certain aspects, the present invention provides methods for dosing a patient with an activin-ActRIIa antagonist and methods for managing patients treated with an activin-ActRIIa anatagonist. In certain aspects, the methods involve measuring one or more hematologic parameters in a patient. | 1. A method for managing a patient that has been treated with, or is a candidate to be treated with, an activin-ActRIIa antagonist, the method comprising monitoring in the patient one or more hematologic parameters that correlate with increased red blood cell levels. 2. The method of claim 1, wherein the hematologic parameters are one or more of the following: red blood cell levels, blood pressure, or iron stores. 3. The method of claim 1, the method further comprising administering to the patient the activin-ActRIIa antagonist in an amount that is appropriate to the observed level of said hematologic parameter in the patient. 4. The method of claim 1, wherein, if the patient has blood pressure elevated above baseline or is hypertensive, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 5. The method of claim 1, wherein, if the patient has blood pressure elevated above baseline or is hypertensive, the patient is treated with a blood pressure lowering agent prior to administration of the activin-ActRIIa antagonist. 6. The method of claim 1, wherein if the patient has uncontrolled hypertension, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 7. The method of claim 1, wherein if the patient has a red blood cell level greater than the normal range for patients of similar age and sex, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 8. The method of claim 1, wherein if the patient has a hemoglobin level of greater than 15 g/dl, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 9. The method of claim 1, wherein if the patient has a hemoglobin level greater than 10, 11 or 12 g/dl, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 10. The method of claim 1, wherein if the patient has iron stores that are lower than the normal range for patients of similar age and sex, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 11. The method of claim 1, wherein if the patient has a transferrin saturation of less than 20%, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 12. The method of claim 1, wherein if the patient has a ferritin level of less than 100 ng/ml, dosing with the activin-ActRIIa antagonist is reduced, delayed or terminated. 13. The method of claim 1, wherein if the patient has iron stores that are lower than the normal range for patients of similar age and sex, the patient is treated with an iron supplement prior to dosing with the activin-ActRIIa antagonist. 14. The method of claim 1, wherein if the patient has a transferrin saturation of less than 20%, the patient is treated with an iron supplement prior to dosing with the activin-ActRIIa antagonist. 15. The method of claim 1, wherein if the patient has a ferritin level of less than 100 ng/ml, the patient is treated with an iron supplement prior to dosing with the activin-ActRIIa antago... | Acceleron Pharma Inc. | A61K039/395 C12Q001/02 G01N033/20 A61K038/16 G01N033/72 G01N033/68 A61P019/08 A61P035/00 A61B005/021
US20100036212 | A1 | 2010-02-11 | APPARATUS, METHOD AND SYSTEM FOR DETERMINING A PHYSIOLOGICAL CONDITION WITHIN A MAMMAL | 2009-08-05 | A system, method and apparatus for determining a physiological condition within a mammal based on pH and/or temperature measurements. In one aspect, the invention is directed to an apparatus for measuring a physiological condition within a mammal comprising: a probe for insertion into a body lumen of the mammal, the probe housing a first circuit board operably coupling a pH sensor for measuring pH within the body lumen and generating a pH signal indicative of the measured pH, a memory device storing parametric data unique to the pH sensor, and a first interface connector located at a proximal potion of the probe; a handle for manipulating the probe, the handle housing a second circuit board operably coupling a microprocessor for processing the pH signal and generating an output signal based on the processing of the pH signal, an indicia device for communicating the output signal generated by the microprocessor, and a second interface connector; the probe connected to the handle in a manner that allows the probe and handle to be repetitively engaged and disengaged from each other; and wherein when the probe is connected to the handle, the first and second interface connectors are in operable connection so that the microprocessor can retrieve the parametric data from the memory device of the probe and receive the pH signal. | 1. An apparatus for measuring a physiological condition within a mammal comprising: an elongated probe for insertion into a body lumen of the mammal, the probe comprising a first circuit board operably coupling a pH sensor for measuring pH within the body lumen and generating a pH signal indicative of the measured pH, a memory device storing parametric data unique to the pH sensor and a first interface connector located at a proximal portion of the probe;a handle for manipulating the probe, the handle comprising a second circuit board operably coupling a microprocessor for processing the pH signal and generating an output signal based on the processing of the pH signal, a display device for displaying the output signal generated by the microprocessor, and a second interface connector;the probe connected to the handle in manner that allows the probe and handle to be repetitively engaged and disengaged from each other; andwherein when the probe is connected to the handle, the first and second interface connectors are in electrical connection so that the microprocessor can retrieve the parametric data from the memory device of the probe and receive the pH signal. 2. The apparatus of claim 1 wherein the pH sensor is an ion-sensitive field effect transistor (ISFET). 3. The apparatus of claim 2 wherein the parametric data includes ISFET slope data. 4. The apparatus of claim 2 wherein the parametric data includes first ISFET slope data determined at ambient temperature and second ISFET slope data determined at normal body temperature. 5. The apparatus of claim 4 wherein the first and second ISFET slope data is based on at least three points of pH parametric data. 6. The apparatus of claim 1 further comprising: the probe comprising a temperature sensor for measuring temperature within the body lumen and generating a temperature signal indicative of the measured temperature, the temperature sensor operably coupled to the first circuit board; andwherein the microprocessor processes the temperature signal and generates the output signal based on the processing of the pH signal and/or the temperature signal. 7. The apparatus of claim 6 wherein the pH sensor is an ISFET and the parametric data includes first ISFET slope data determined at ambient temperature and second ISFET slope data determined at normal body temperature. 8. The apparatus of claim 6 wherein the elongated probe has a distal portion and a proximal portion, the pH sensor and the temperature sensor located at the distal portion and the first interface connector is located at the proximal portion. 9. The apparatus of claim 8 wherein the handle further comprises a socket forming a passageway into the handle, the second interface connector aligned with the socket, and wherein when the probe is connected to the handle, the proximal portion of the probe extends into the socket of the handle. 10. The apparatus of claim 9 wherein the probe and the handle are connected together by one of a tig... | - | A61B010/00 A61B005/00
US20100049033 | A1 | 2010-02-25 | MEDICAL DEVICE POSITION DETECTION SYSTEM, MEDICAL DEVICE GUIDANCE SYSTEM, POSITION DETECTION METHOD OF MEDICAL DEVICE GUIDANCE SYSTEM, AND GUIDANCE METHOD OF MEDICAL DEVICE GUIDANCE SYSTEM | 2007-11-12 | A medical device position detection system and a position detection method of a medical device guidance system, which are capable of detecting a direction of the medical device accurately, are provided. Included are a medical device introduced into a subject's body; a magnetic field response part that is disposed in the medical device, responds to a magnetic field by virtue of possessing a magnetization direction, and guides the medical device; a magnetic field generation part forming the magnetic field within the subject's body; a direction detection magnetic field control part generating a direction detection magnetic field from the magnetic field generation part for detecting the direction of the medical device; a response detection part detecting a response of the magnetic field response part due to the direction detection magnetic field; and a direction calculating part calculating the direction of the medical device according to the direction of the direction detection magnetic field and a detection result of the response detection part. | 1. A medical device position detection system, comprising: a medical device introduced into a subject's body;a magnetic field response part that is disposed in the medical device, responds to a magnetic field by virtue of possessing a magnetization direction, and guides the medical device;a magnetic field generation part generating a magnetic field within the subject's body;a direction detection magnetic field control part generating a direction detection magnetic field, from the magnetic field generation part, for detecting a direction of the medical device;a response detection part detecting the response of the magnetic field response part due to the direction detection magnetic field; anda direction calculation part calculating the direction of the medical device according to a direction of the direction detection magnetic field and a detection result of the response detection part. 2. The medical device position detection system according to claim 1, wherein directions of two axes having different directions from each other among three axis directions having different directions from one another in the medical device, are calculated by the response detection part, and a direction of an axis intersecting a plane formed by the two axes is calculated by the direction calculation part. 3. The medical device position detection system according to claim 1, wherein the response detection part includes an image acquisition part acquiring an image of inside the subject's body. 4. The medical device position detection system according to claim 1, wherein the response detection part includes a magnetic force measurement part measuring a force generated in the magnetic field response part. 5. The medical device position detection system according to claim 1, wherein the direction detection magnetic field control part controls the magnetic field generation part to generate a static magnetic field. 6. The medical device position detection system according to claim 5, wherein the direction detection magnetic field control part controls the magnetic field generation part to sequentially generate a plurality of magnetic fields having different directions or intensities from one another, andthe direction calculation part calculates the direction of the medical device according to respective detection results of the response detection part for the plurality of magnetic fields. 7. The medical device position detection system according to claim 1, wherein the direction detection magnetic field control part controls the magnetic field generation part to generate a gradient magnetic field. 8. The medical device position detection system according to claim 2, wherein the medical device has a substantially cylindrical shape,the magnetization direction of the magnetic field response part is substantially perpendicular to a center axis of the substantially cylindrical shape, andthe plane formed by the two axes detected by the response detection part is substanti... | OLYMPUS MEDICAL SYSTEMS CORP. | A61B005/05 A61B001/00
US20100004549 | A1 | 2010-01-07 | SYSTEM AND METHOD OF SERIAL COMPARISON FOR DETECTION OF LONG QT SYNDROME (LQTS) | 2009-09-21 | The present disclosure includes a system and method of detecting LQTS in a patient by comparing a collected set of ECG data from the patient to a plurality of databases of collected ECG data. The plurality of databases will include a database containing previous ECGs from the patient, a known acquired LQTS characteristics database, and a known genetic LQTS characteristics database. Comparing the patients ECG to these databases will facilitate the detection of such occurrences as changes in QT interval from success of ECGs, changes in T-wave morphology, changes in U-wave morphology and can match known genetic patterns of LQTS. The system and method is sensitive to patient gender and ethnicity, as these factors have been shown to effect LQTS, and is furthermore capable of matching a QT duration to a database of drug effects. The system and method is also easily integrated into current ECG management systems and storage devices. | 1. A method of detecting a long QT syndrome in a patient, the method comprising: acquiring a set of ECG data from the patient;comparing the set of ECG data to a set of ECG characteristic data stored in a plurality of databases, wherein the comparing step is configured to identify any characteristics in the set of ECG data that are emblematic of LQTS; andadjusting an interpretation of the set of ECG data when a characteristic is found in the set of ECG data, such that the adjustment corresponds to the characteristic. 2. The method according to claim 1, further comprising making the interpretation prior to the adjusting step. 3. The method according to claim 1, further comprising sending the set of ECG data to a centralized storage system. 4. The method according to claim 1, wherein the plurality of databases includes a previous ECG database, wherein the comparing step includes identifying any significant change in a QT interval. 5. The method according to claim 1, wherein the plurality of databases includes a known acquired LQTS characteristics database, wherein the comparing step includes identifying a significant characteristic in the set of ECG data with respect to a known acquired characteristic. 6. The method according to claim 1, wherein the plurality of databases includes a known genetic LQTS characteristics database, wherein the comparing step includes identifying a significant characteristic in the set of EGC data with respect to a known genetic characteristic. 7. The method according to claim 1, further comprising updating the interpretation in the centralized storage system. 8. The method according to claim 1, further comprising adding the set of ECG data to any of the plurality of databases. 9. The method according to claim 1, further comprising outputting the interpretation to a user. 10. A system for detecting a long QT syndrome in a patient, the system comprising: an ECG acquisition system configured to collect a set of ECG data from the patient;a storage media for storing a computer application;a processor coupled to the ECG acquisition system and the storage media, and configured to execute the computer application, and further configured to receive the set of ECG data from the ECG acquisition system,wherein when the computer application is executed, the set of ECG data is compared to a set of ECG characteristic data in a plurality of databases, any characteristics in the set of ECG data that are emblematic of LQTS are identified, and an interpretation of the set of ECG data is adjusted when a characteristic is identified in the set of ECG data, such that the adjustment corresponds to the characteristic. 11. The system according to claim 10, wherein the interpretation is made prior to the adjusting step. 12. The system according to claim 10, further comprising a centralized storage system wherein the set of ECG data is sent to the centralized storage system. 13. The system according to claim 10, wherein the plurality o... | General Electric Company | A61B005/0452
US07643971 | B2 | 2010-01-05 | Disease simulation system and method | 2006-10-31 | A system and method for predicting the effect of patient self-care actions on a disease control parameter. A future disease control parameter value X(tj) at time tj is determined from a prior disease control parameter value X(ti) at time ti based on an optimal control parameter value R(tj) at time tj, the difference between the prior disease control parameter value X(ti) and an optimal control parameter value R(ti) at time ti, and a set of differentials between patient self-care parameters having patient self-care values SM(ti) at time ti and optimal self-care parameters having optimal self-care values OM(ti) at time ti. The differentials are multiplied by corresponding scaling factors KM. The system includes an input device for entering the patient self-care values SM(ti). A memory stores the optimal control parameter values R(ti) and R(tj), the prior disease control parameter value X(ti), the optimal self-care values OM(ti), and the scaling factors KM. A processor in communication with the input device and memory calculates the future disease control parameter value X(tj). A display is connected to the processor to display the future disease control parameter value X(tj) to a patient. | 1. A simulation system comprising: a processor unit;a server configured to communicate with the processor unit through at least one communication network and run a simulation program stored in a computer readable medium, wherein (i) the server is remotely situated from the processor unit, (ii) the simulation program is configured to generate one or more simulated disease control parameter values and (iii) the disease control parameter values are related to patient self-care of a chronic or persistent condition being managed by a patient;a recording device remotely situated from the server and configured to communicate with the processor unit; anda healthcare provider computer remotely situated from the server, the processor unit and the recording device, wherein the healthcare provider computer is configured to communicate with the processor unit. 2. The simulation system according to claim 1, further comprising a data storage card, wherein the healthcare provider computer and the processor unit are configured to communicate via the data storage card. 3. The simulation system according to claim 1, wherein the healthcare provider computer and the processor unit are configured to communicate via the at least one communication network. 4. The simulation system according to claim 1, wherein the healthcare provider computer is programmed to determine one or more scaling factors from values of physiological parameters of the patient. 5. The simulation system according to claim 4, wherein the healthcare provider computer is programmed to determine the one or more scaling factors from values of physiological parameters of the patient using predetermined indexes. 6. The simulation system according to claim 5, wherein the healthcare provider computer is programmed to adjust the one or more scaling factors based on differences between actual disease control parameter values and optimal disease control parameter values. 7. The simulation system according to claim 6, wherein the one or more scaling factors are customized to an individual patient. 8. The simulation system according to claim 1, wherein the recording device and the processor unit are housed in a single unit. 9. The simulation system according to claim 8, wherein the single unit is configured to communicate with the server via a wireless network. 10. The simulation system according to claim 8, wherein the single unit comprises a cellular phone. 11. The simulation system according to claim 8, wherein the single unit comprises a mobile telephone. 12. The simulation system according to claim 8, wherein the single unit comprises a personal digital assistant (PDA). 13. The simulation system according to claim 1, further comprising a television remote control configured to communicate with the processor unit. 14. The simulation system according to claim 1, wherein the at least one communication network comprises a network type or a combination of two or more network types selected from the... | Health Hero Network, Inc. | G06F019/00 A61B005/00 G01N033/487
US07643858 | B2 | 2010-01-05 | System and method for detection of brain edema using spectrophotometry | 2006-09-28 | Embodiments of the present invention relate to a system and method of detecting or monitoring brain edema in a patient. One embodiment of the present invention includes emitting a first light into the patient's brain tissue at a first wavelength, emitting a second light into the patient's brain tissue at a second wavelength, detecting the first and second lights after dispersion by the brain tissue at a detector, and determining an amount of water proximate the brain tissue based on the detected first and second lights. | 1. A method of detecting or monitoring brain edema in a patient, comprising: emitting a first light into the patient's brain tissue at a first wavelength through a partial bone region of the patient's skull, wherein the partial bone region of the patient's skull is approximately 3 cm behind a coronal suture of the patient's skull, wherein the coronal suture includes a suture between parietal and frontal bones of the patient's skull;emitting a second light into the patient's brain tissue at a second wavelength through said partial bone region;detecting the first and second lights after dispersion by the brain tissue at a detector; anddetermining an amount of water proximate the brain tissue based on the detected first and second lights. 2. The method of claim 1, comprising emitting a third light into the patient's brain tissue at a third wavelength. 3. The method of claim 2, comprising detecting the third light after reflection or dispersion by the brain tissue. 4. The method of claim 3, comprising determining the amount of water proximate the brain tissue based on the detected third light in combination with the detected first and second lights. 5. The method of claim 1, comprising emitting the first and second lights from at least approximately 5 mm away from the detector. 6. The method of claim 1, comprising emitting the first and second lights from approximately 10-30 mm away from the detector. 7. The method of claim 1, comprising emitting the first and second lights from approximately 1-5 mm away from the detector. 8. The method of claim 1, comprising penetrating approximately 16 mm into the patient with the first and second lights. 9. The method of claim 1, wherein the first and second wavelengths are within a 1050-1350 nm spectral region. 10. The method of claim 1, wherein the first and second wavelengths are within a 1500-1800 nm spectral region. 11. The method of claim 1, wherein the first and second wavelengths are within a 2100-2300 nm spectral region. 12. The method of claim 2, comprising emitting the third light into the brain tissue via the partial bone region. 13. A method of detecting or monitoring brain edema in a patient, comprising: emitting a first light into the patient's brain tissue at a first wavelength approximately 1-5 mm away from a detector, wherein the first wavelength is within a 2100-2300 nm spectral region and wherein the first light is emitted through a partial bone region of the patient's skull approximately 3 cm behind a coronal suture of the patient's skull, wherein the coronal suture includes a suture between parietal and frontal bones of the patient's skull;emitting a second light into the patient's brain tissue at a second wavelength approximately 1-5 mm away from the detector, wherein the second wavelength is within the 1500-2300 nm spectral region and wherein the second light is emitted through said partial bone region;detecting the first and second lights after dispersion by the brain tissue at... | Nellcor Puritan Bennett LLC | A61B005/00
US07654964 | B1 | 2010-02-02 | System and method for detecting arterial blood pressure based on aortic electrical resistance using an implantable medical device | 2006-03-16 | Techniques are provided for use with a pacemaker or other implantable medical device for detecting arterial blood pressure. Briefly, the pacemaker detects aortic electrical resistance using sensing/pacing leads. Aortic electrical resistance pertains to the resistance to an electrical current passing through the aorta. The pacemaker then determines the arterial blood pressure of the patient based on the aortic resistance and a predetermined calibration value that relates aortic resistance to arterial pressure. The calibration value is updated monthly based on blood pressure values detected using an external blood pressure sensor employing a blood pressure cuff. Other techniques described herein pertain to the determination of other physiologic parameters such as stroke volume and cardiac output and to the detection of changes in hematocrit. Any of the various physiological parameters may then be used to trigger or control warning signals and responsive therapy. | 1. A method for use with an implantable medical device for detecting blood pressure, the method comprising: detecting a value representative of aortic electrical resistance (R) within the patient using a detector; andin the implantable medical device, determining arterial blood pressure within the patient based on the aortic resistance value using a blood pressure determination system byretrieving at least one calibration value (K) relating aortic resistance to arterial blood pressure within the patient, anddetermining the arterial blood pressure based on the aortic resistance value (R) and the calibration value (K);wherein the calibration value (K) is representative of arterial pressure (P) multiplied by systemic resistance (R). 2. The method of claim 1wherein the implantable medical device includes a right atrial electrode and a device housing electrode positioned generally on opposing sides of the aorta of the patient; andwherein detecting aortic resistance (R) includes measuring electrical resistance between the right atrial electrode and the device housing electrode. 3. The method of claim 1 further including a calibration procedure for calculating K based on known values of patient pressure (P) and systemic resistance (R). 4. The method of claim 3 wherein the device is used in conjunction with an external arterial blood pressure detector and wherein the calibration procedure includes receiving a value representative of true patient blood pressure (P′) from the external detector measured at a time t;detecting a value representative of patient aortic resistance (R′), the value also detected at about time t; andcalculating K based on K=P′R′. 5. The method of claim 3 wherein the calibration procedure is performed periodically to update the value of K. 6. The method of claim 5 wherein the calibration procedure is performed monthly. 7. A method for use with an implantable medical device for detecting blood pressure, the method comprising: detecting a value representative of aortic electrical resistance (R) within the patient using a detector; andin the implantable medical device, determining arterial blood pressure within the patient based on the aortic resistance value using a blood pressure determination system byretrieving at least one calibration value (K) relating aortic resistance to arterial blood pressure within the patient, anddetermining the arterial blood pressure based on the aortic resistance value (R) and the calibration value (K); andwherein determining arterial blood pressure within the patient based on the aortic resistance value (R) and the calibration value (K) is performed by calculating: P=K/R. 8. The method of claim 1 wherein determining arterial blood pressure within the patient based on the aortic resistance value (R) and the calibration value (K) is performed substantially continuously based on newly detected values of aortic resistance. 9. A method for use with an implantable medical device for detecting blood p... | Pacesetter, Inc. | A61B005/02
US07651578 | B2 | 2010-01-26 | Guidewire with polymer jacket and method of making | 2006-06-08 | A guidewire and methods for making and using the same. The guidewire may include a core wire and a polymer jacket attached to the core wire. The guidewire may be manufactured by advancing the core wire, polymer jacket, and a tooling tube through a heated die assembly so as to heat and compress the polymer jacket onto the core wire. Optionally, the tooling tube may be removed from the jacket after advancing the components through the die assembly. | 1. A method for manufacturing a medical guidewire, the method comprising: providing an elongated core member defining an outer surface;disposing an polymer jacket member about at least a portion of the outer surface of the elongated core member, wherein the polymer jacket member comprises a thermoplastic polymer material;providing an elongated tubular member having an undeformed state and including an inner surface defining a lumen, and defining an outer surface;disposing the elongated tubular member about at least a portion of the polymer jacket member, wherein the portion of the polymer jacket member and a portion of the core member are disposed within the lumen;applying heat and compressive force to the outer surface of the elongated tubular member such that the portion of the polymer jacket member is heated to or above a predetermined temperature at which the thermoplastic polymer material softens and such that the softened portion of the polymer jacket member is deformed to attach to the core member;wherein the elongated tubular member is configured to at least partially recover to the undeformed state when the heat and compressive force is removed; andremoving the elongated tubular member from the polymer jacket member. 2. The method of claim 1, wherein the predetermined temperature is at or above the softening point temperature of the thermoplastic polymer material. 3. The method of claim 1, wherein the predetermined temperature is at or above the Tg and at or below the melting point temperature of the thermoplastic polymer material. 4. The method of claim 1, wherein the predetermined temperature is above the Tg of the thermoplastic polymer material. 5. The method of claim 1, wherein the predetermined temperature is below the melting point temperature of the thermoplastic polymer material. 6. The method of claim 1, wherein when heated to or above a predetermined temperature, the thermoplastic polymer material is more visco-elastic than when at temperatures below the predetermined temperature. 7. The method of claim 1, wherein when heated to or above a predetermined temperature, the thermoplastic polymer material is less viscous than when at temperatures below the predetermined temperature. 8. The method of claim 1, wherein the elongated tubular member comprises a polymer material, and wherein when the predetermined amount of heat and compressive force is applied such that the thermoplastic polymer material is heated to or above the predetermined temperature, the material of the elongated tubular member is not as soft as the thermoplastic polymer material. 9. The method of claim 1, wherein the elongated tubular member comprises a polymer material, and wherein when the predetermined amount of heat and compressive force is applied such that the thermoplastic polymer material is heated to or above the predetermined temperature, the material of the elongated tubular member is less visco-elastic than the thermoplastic polymer material. ... | Boston Scientific Scimed, Inc. | B29C049/00 B60J010/00 B29C065/00 B32B037/00 B65C009/25 C09J005/00 A61B005/00 A61M025/00 A61M005/178 A61F011/00
US07650177 | B2 | 2010-01-19 | Medical sensor for reducing motion artifacts and technique for using the same | 2006-08-01 | A sensor for pulse oximetry or other applications utilizing spectrophotometry may be adapted to reduce motion artifacts by fixing the optical distance between an emitter and detector. A flexible sensor is provided with a stiffening member to hold the emitter and detector of the sensor in a relatively fixed position when applied to a patient. Further, an annular or partially annular sensor is adapted to hold an emitter and detector of the sensor in a relatively fixed position when applied to a patient. A clip-style sensor is provided with a spacer that controls the distance between the emitter and detector. | 1. A sensor comprising: a conformable sensor body;a removable stiffening member disposed within an open cavity in the sensor body; andan emitter and a detector disposed on the stiffening member, wherein the emitter and the detector are adapted to operate in reflectance mode, and wherein the sensor body is adapted to hold the emitter and detector at a substantially fixed optical distance relative to one another when the sensor is applied to a patient. 2. The sensor, as set forth in claim 1, wherein the sensor comprises at least one of a blood constituent sensor or a tissue constituent sensor. 3. The sensor, as set forth in claim 1, wherein the sensor comprises at least one of a pulse oximetry sensor or a sensor for measuring a water fraction. 4. The sensor, as set forth in claim 1, wherein the sensor comprises at least one of a carboxyhemoglobin sensor or a methemoglobin sensor. 5. The sensor, as set forth in claim 1, wherein the sensor is adapted for intrauterine use. 6. The sensor, as set forth in claim 1, wherein the stiffening member is embedded within the sensor body. 7. The sensor, as set forth in claim 1, wherein the stiffening member comprises a metal or a rigid polymeric material. 8. A pulse oximetry system comprising: a pulse oximetry monitor; anda pulse oximetry sensor adapted to be operatively coupled to the monitor, the sensor comprising: a conformable sensor body;a removable stiffening member disposed within an open cavity in the sensor body; andan emitter and a detector disposed on the stiffening member, wherein the emitter and the detector are adapted to operate in reflectance mode, and wherein the sensor body is adapted to hold the emitter and detector at a substantially fixed optical distance relative to one another when the sensor is applied to a patient. 9. The system, as set forth in claim 8, wherein the sensor comprises a tissue constituent sensor. 10. The system, as set forth in claim 8, wherein the sensor comprises a sensor for measuring a water fraction. 11. The system, as set forth in claim 8, wherein the sensor comprises at least one of a carboxyhemoglobin sensor or a methemoglobin sensor. 12. The system, as set forth in claim 8, wherein the sensor is adapted for intrauterine use. 13. The system, as set forth in claim 8, wherein the stiffening member is embedded within the sensor body. 14. The system, as set forth in claim 8, wherein the stiffening member comprises a metal or a rigid polymeric material. 15. A method of manufacturing a pulse oximetry sensor, comprising: providing a conformable sensor body adapted for use on a patient's digit;providing a removable stiffening member disposed within an open cavity in the sensor body; andproviding an emitter and a detector disposed on the stiffening member, wherein the emitter and the detector are adapted to operate in reflectance mode, and wherein the sensor body is adapted to hold the emitter and detector at a substantially fixed optical distance relative to o... | Nellcor Puritan Bennett LLC | A61B005/1455
US07654962 | B2 | 2010-02-02 | Radiation stress non-invasive blood pressure method | 2004-06-28 | The invention determines the energy dispersion via acoustic, electromechanical or other related physiological signals collected from a patient that lies down or otherwise engages a discritized sensing array. Signals are monitored over a range of frequencies and collected in the time domain or frequency domain. A computing machine determines the energy from the signal measured over various elements of the array and calculates the momentum flux. Blood pressure is determined directly from the momentum flux calculation. | 1. A radiation stress, non-invasive vital sign monitoring method comprising: providing one or more discretized sensor arrays,engaging the one or more discretized sensor arrays,measuring and collecting discretized acoustic or electromechanical or physiological signals with the discretized sensor arrays,transmitting discretized signals to a receiving and computer device,the receiving and computer deviceproduces time series data from various discretized sensor array signals,calculates energy spectrum from the time series data,determines variance of each discretized sensor array, andcalculates a value for vital signs of a patient. 2. The method of claim 1, wherein the vital signs are average, mean, systolic and diastolic arterial blood pressure. 3. The method of claim 1, wherein the vital signs are hypertension. 4. The method of claim 1, further comprising lying on, standing on, or otherwise contacting the discretized sensor arrays. 5. The method of claim 1, wherein the collecting the discretized acoustic or electromechanical physiological signals is performed over a range of frequencies. 6. The method of claim 1, wherein the collecting the discretized acoustic or electromechanical physiological signals is performed over a single frequency. 7. The method of claim 1, wherein the collecting acoustic or electromechanical physiological signals further comprises collecting data in a time domain or frequency domain. 8. The method of claim 1, wherein the calculating a value for vital signs is performed with non-time series methods for determining energy at various array points or a combination of array points. 9. The method of claim 1, wherein the transmitting of discretized signals comprises transmitting discretized signals via wire, fiber optics or wirelessly. 10. The method of claim 1, further comprising providing continuous, real-time monitoring of a patient's vital signs. 11. The method of claim 1, further comprising calculating the momentum flux from data gathered from the discretized sensor arrays. 12. The method of claim 11, further comprising calculating a patient's vital signs from the momentum flux. 13. The method of claim 1, wherein the one or more discretized sensor arrays are not attached to the patient. 14. A radiation stress, non-invasive vital sign monitoring device comprising: one or more discretized sensor arrays for measuring and collecting discretized acoustic or electromechanical signals from a patient,a surface on the one or more discretized sensor arrays for engaging a patient,a transmission system for transmitting data collected by the one or more discretized sensor arrays,a receiving device for receiving the transmitted data from the one or more discretized sensor arrays, anda computing device connected to the receiving device for calculating values of vital signs of the patient by producing time series data from various discretized sensor array signals,calculating energy spectrum from the time series data, anddetermin... | Hoana Medical, Inc. | A61B005/02 A61B005/00
US07661200 | B2 | 2010-02-16 | Method and device for determining a person's motions | 2005-03-30 | A sensor worn by a person measures movements of the person. Movements of orientation are not expressed by rotations leading from a stationary reference point to a reference point associated with the sensor, but by angles separating axes or planes of the stationary reference point and axes associated with the wearer, thereby making it possible to better identify his/her movements and his/her activity. The sensor can be applied to the monitoring of patients. | 1. A method for determining movement of a person, comprising: preliminarily calibrating a sensor that is installed on the person by recording a preliminary orientation of the sensor while the person is at a predetermined orientation with respect to a stationary reference;performing a repeated computation of an orientation of the sensor with respect to the stationary reference, expressed as a combination of rotations from the stationary reference, from signals of the sensor; andcorrecting the combination of rotations so as to express rotations of the sensor bound to the person, which includes a vertical axis, a front-back horizontal axis and a left-right horizontal axis, as another combination of rotations of the person that include a side inclination of the person, a front-back inclination of the person and an azimuth rotation of the person. 2. The method according to claim 1, wherein the sensor includes accelerometers and magnetometers. 3. The method according to claim 1, wherein the sensor includes gyrometers. 4. The method according to claim 1, wherein angles of orientation of the sensor are Euler angles. 5. The method according to claim 1, wherein angles of orientation of the sensor are expressed by quaternions. 6. The method according to claim 1, wherein the preliminary calibrating includes storing an orientation of the person with respect to the stationary reference point and storing an orientation of the sensor with respect to the stationary reference point while the person is at the predetermined orientation, andthe correcting includes converting the combination of rotations of the sensor into the another combination of rotations of the person that include the side inclination of the person, the front-back inclination of the person and the azimuth of the person based the orientation of the person with respect to the stationary reference point and the orientation of the sensor with respect to the stationary reference point. 7. A device configured to determine a movement of a person, comprising: a memory; anda processing apparatus configured to preliminarily calibrate a sensor that is installed on the person by causing a preliminary orientation of the sensor while the person is at a predetermined orientation with respect to a stationary reference to be recorded in the memory,perform a repeated computation of an orientation of the sensor with respect to the stationary reference, expressed as a combination of rotations from the stationary reference, from signals of the sensor, andcorrect the combination of rotations so as to express rotations of the sensor bound to the person, which includes a vertical axis, a front-back horizontal axis and a left-right horizontal axis, as another combination of rotations that include a side inclination of the person, a front-back inclination of the person and an azimuth rotation of the person. | Commissariat a l'Energie Atomique | A61B005/103
US07668584 | B2 | 2010-02-23 | Interface apparatus for passive tracking systems and method of use thereof | 2003-08-11 | An interface apparatus (20) for tracking by a tracking system (40) of an object(s) in space for position and orientation and for interacting with the tracking system (40). The interface apparatus (30) comprises passive detectable devices (12,14,16,22) trackable for position by the tracking system (40). A mounting device (10,24) receives the passive detectable devices (12,14,16,22) in a known geometry, and is secured to the object(s) such that a position and orientation of the object(s) is calculable by the tracking system (40) as a function of a tracking of the known geometry of the passive detectable devices (12,14,16,22). One of the passive detectable devices (22) is displaceable with respect to the object(s). A displacement of the passive detectable device (22) with respect to the object(s) is detectable to initiate an interaction with the tracking system (40) while maintaining the tracking of the object(s). | 1. A passive optical interference apparatus for tracking by a tracking system of an object in space for position and orientation and for interacting with the tracking system, the passive optical interface apparatus comprising: at least three passive detectable devices trackable for position by the tracking system; anda mounting device for receiving the at least three passive detectable devices in a known geometry and adapted for being secured to the object such that a position and orientation of the object is calculable by the tracking system as a function of a tracking of the known geometry of the at least three passive detectable devices, at least a first of the at least three passive detectable devices being secured to the mounting device by a pivot joint, the pivot joint operatively connecting said first of the at least three passive detectable devices for pivotal movement between a first position and a second position with respect to the object, a pivotal displacement of said first of the at least three passive detectable devices to any one of said first position and said second position with respect to the object being detectable to initiate an interaction with the tracking system while maintaining the tracking of the object. 2. The passive optical interface apparatus according to claim 1, wherein the apparatus comprises four of said passive detectable devices trackable for position by the tracking system, a second, a third and a fourth of the four passive detectable devices being positioned in said known geometry and the first of the four passive detectable devices being displaceable with respect to the known geometry, a displacement of the first of the four passive detectable devices with respect to the known geometry being detectable to initiate an interaction with the tracking system. 3. The passive optical interface apparatus according to claim 2, wherein the first of the four passive detectable devices is displaceable between said first position and said second position with respect to the known geometry, a displacement of the first of the four passive detectable devices from the first position to the second position being interpreted by the tracking system as a first type of interaction, a displacement of the first of the four passive detectable devices from the second position to the first position being interpreted by the tracking system as a second type of interaction. 4. The passive optical interface apparatus according to claim 2, wherein the first of the four passive detectable devices is biased to said first position by the mounting device, a displacement of the first of the four passive detectable devices away from the first position being detectable to initiate an interaction with the tracking system. 5. A method for interaction between a tracking system tracking a position and orientation of an object in space and a handler of the object, the object having at least three passive detectable devices with at least one of s... | Orthosoft Inc. | A61B005/05
US08131340 | B2 | 2012-03-06 | Detector unit for arrangement within a field generating unit of an MR device | 2007-09-19 | A detector unit for arrangement in a field generating unit of a magnetic resonance device has an RF transmission/reception system for transmitting RF pulses into, or receiving magnetic resonance signals from, an examination volume of the field generation unit. The RF transmission/reception system surrounds a patient tunnel at a radial distance from a tunnel axis thereof, and is divided into two sub-systems located at a distance from each other along the direction of the tunnel axis, so as to form a substantially annular cavity or interstice therebetween. | 1. A detection unit configured for arrangement in a field generation unit of a magnetic resonance apparatus, said detection unit comprising: a radio-frequency transmission/reception system that transmits RF pulses and/or receives magnetic resonance signals, said radio-frequency transmission/reception system being configured to surround a patient tunnel, having a longitudinal tunnel axis, of the field generation unit at a radial distance from the tunnel axis;said radio-frequency transmission/reception system comprising two sub-systems spaced apart from each other along a direction of said tunnel axis and forming a substantially annular interstice between said two sub-systems; anda radio-frequency shield that radially outwardly shields said radio-frequency transmission/reception system, said radio-frequency shield comprising two sub-system shield segments and an interstice shield segment disposed successively along said direction of said tunnel axis, said sub-system shield segments respectively covering said sub-systems of said radio-frequency transmission/reception system and said interstice shield segment at least partially covering said interstice in said radio-frequency transmission/reception system, said interstice shield segment adjoining said sub-system shield segments in a radio-frequency-tight configuration along said direction of said tunnel axis, and said radio-frequency shield comprising two rings extending in a radial direction relative to said tunnel axis, said two rings respectively having radially outer ends that respectively adjoin ends of the sub-system shield segments axially facing toward each other along said tunnel axis. 2. A detection unit as claimed in claim 1 wherein said interstices shield segments adjoins the radially outer ends of said rings. 3. A detection unit as claimed in claim 1 wherein said rings respectively also have radially inner ends, and wherein said interstice shield segment adjoins the respective radially inner ends of said rings. 4. A detection unit as claimed in claim 1 wherein said rings respectively also have radially inner ends, and wherein said interstice shield segment adjoins said rings between said radially inner and radially outer ends relative to said tunnel axis. 5. A detection unit as claimed in claim 1 wherein said interstices shield segments adjoins the radially outer ends of said rings, and wherein said radio-frequency transmission/reception system is a first radio-frequency transmission/reception system, and wherein said detection unit further comprises a second radio-frequency transmission/reception system disposed radially within said interstices shield segment in said interstice of said first radio-frequency transmission/reception system. 6. A detection unit as claimed in claim 5 wherein said second radio-frequency transmission/reception system is operable simultaneously with at least one of the sub-systems of said first radio-frequency transmission/reception system. 7. A detection... | Siemens Aktiengesellschaft | A61B005/05
US07988647 | B2 | 2011-08-02 | Assessment of medical conditions by determining mobility | 2009-03-16 | A mobility assessment detects abnormalities in motion using a stagger algorithm. The detected abnormalities are compared to known norms for a particular test to determine if the abnormalities are normal or not. The comparison can be made for abnormalities associated with different diseases and illnesses to classify whether or not the condition exists and its phase. The results may be used to generate a treatment regime. | 1. A computer implemented system for assessing the mobility of a subject, said system comprising: a camera to record motion of said subject;a motion sensor to receive an output from said camera and determine movement of said subject, said motion sensor generating a data stream representative of such movement;a database; a first computer implemented assessment engine including a computer generated prompt to prompt an operator to provide data related to a test being performed by said subject,an input to permit an operator to provide additional data in response to said prompt for recordal with said data stream in a database;said system further including a second computer implemented assessment engine including a comparator to permit a comparison between motion represented by said data stream and representative movements recorded in said database to indicate whether said output is within said known norm, an allocator operable upon an indication of said comparator to allocate said motion to either a known norm or outside a known norm, and a deviation function to compare motion represented by said data stream with a representative movement and to detect abnormalities in such motion, and;an output to indicate to an operator a classification of the mobility of said subject. 2. A system according to claim 1 wherein said database is organized to provide a pair of subsets of said database, each of which provides a respective known norm derived from the contents of the database, said comparator comparing said abnormalities to said known norms provided by said subsets of said database. 3. A system according to claim 2 wherein said allocator is operable to add said abnormalities to one of said databases. 4. A system according to claim 3 wherein said comparator implements a stagger algorithm to detect abnormalities. 5. A system according to claim 4 wherein said comparator incorporates a fuzzy logic decision engine to implement said stagger algorithm. 6. A system according to claim 2 including a mobility condition database containing records of abnormalities of prior assessments to permit the said output to be compared to prior assessments. 7. A system according to claim 6 wherein records of said mobility condition database include categories of conditions to permit said output to be compared to selected conditions and a classification of said condition obtained. 8. A system according to claim 7 including treatment database to permit said classification to generate a treatment regime. 9. A method of assessing mobility of a subject using a computer implemented system having a camera to record motion of said subject; a motion sensor; a database; a first computer implemented assessment engine and a second computer implemented assessment engine, said method comprising the steps of: recording motion of said subject with said camera,utilising said motion detector to determine movement of said subject from an output from said camera,utilising said first assessm... | - | A61B005/103 A61B005/117
US08290568 | B2 | 2012-10-16 | Method for determining a property map of an object, particularly of a living being, based on at least a first image, particularly a magnetic resonance image | 2009-01-09 | It is disclosed a system and method (12) for determining a property map (82) of an object, particularly a human being, based on at least a first image (84), particularly an magnetic resonance (MR) image, of the object. In the method (12), a structure of reference pairs is defined in a first step (96), wherein each reference pair (16-26) comprises at least two entries (62). The first entry represents a property value, particularly an attenuation value. The second entry (62) preferably represents a group of image points (67) belonging together, which is extracted particularly from MR images (28) and comprises an interesting image point corresponding to the property value. In another step (98) of the method (12) a plurality of training pairs (16-26) is provided. A structure of the training pairs (16-26) corresponds to the structure of reference pairs, and the entries of respective training pairs (16-26) are known. In another step (100) of the method (12), an assignment between the first entries and the other entries (62-66) of the training pairs (16-26) is determined by machine learning, thus allowing prediction of a property value (88) corresponding to an arbitrary point (90) of the first image (84). | 1. A method for determining a property map of an object based on at least a first magnetic resonance (MR) image of the object, comprising: defining a structure of reference pairs, wherein each reference pair has at least first and second entries, wherein each of said first entries represents an unknown property value, and wherein each of said second entries represents a group of image points belonging together, which are extracted from said first MR image recorded by a scanner, wherein the group of image points includes at least one interesting image point which corresponds to one of the unknown property-values, wherein each unknown property value represents an attenuation value;providing a plurality of training pairs, wherein a structure of the training pairs corresponds to the structure of said reference pairs, and wherein first and second entries of the respective training pairs are known; anddetermining an assignment between the first entries and the second entries of the training pairs by machine learning, and predicting an attenuation value for an arbitrary image point which corresponds to the at least one interesting image point of the first MR image by applying the determined assignment to one of the second entries of the reference pairs including said arbitrary image point. 2. The method of claim 1, wherein the structure of said reference pairs additionally comprises third or additional entries, which are taken into account during the step of determining the assignment. 3. The method of claim 2, wherein the third or additional entries are selected from a group consisting of: a registered coordinate;another group of image points belonging together of at least a second image, wherein the other group of image points of the at least second image shows at least partially an identical object patch like the group of image points of the first image; anda class affiliation. 4. The method of claim 3, wherein the registered coordinate is defined by one of a coordinate of a point equivalent to the interesting point in a reference object, and one or more distances of a point equivalent to the interesting point relative to anatomically characteristic positions within the reference object. 5. The method of claim 3, wherein a similarity between the registered coordinate and a position-dependent distance measure is determined, which is dependent on properties of a reference object as well as a position of an equivalent point within the reference object. 6. The method of claim 3, wherein the at least second image is one of a positron emission tomography (PET) image and another MR image recorded by a scanner. 7. The method of claim 3, wherein at least one of the at least first MR image and the at least second image are pre-processed by the extraction of features. 8. The method of claim 7, wherein the pre-processing is performed a Fourier/Wavelet transformation, SIFT, edge detectors, or contrast amplification. 9. The method of claim 3, wherein addit... | Eberhard-Karls-Universitat Tubingen Universitatsklinikum | A61B005/05
US08708934 | B2 | 2014-04-29 | Reorientation of patient posture states for posture-responsive therapy | 2009-04-30 | In general, the disclosure relates to medical devices and systems configured to deliver posture-responsive therapy to a patient. For example, a medical system may detect a posture state of a patient and deliver therapy based at least in part on the detected patient posture state. In some examples, the system comprises a posture state sensor configured to generate first posture sensor data when a patient is in a first posture state. The system may further comprise a processor configured to define posture state reference data corresponding to the first posture state based at least in part on the first posture sensor data, and redefine the posture state reference data based on second posture data generated by the posture state sensor when the patient is in the first posture state. | 1. A method comprising: receiving first posture sensor data from a posture state sensor when a patient occupies a first posture state during a first period of time;defining posture state reference data corresponding to the first posture state based at least in part on the first posture sensor data; andredefining the posture state reference data based on second posture sensor data from the posture state sensor when the patient subsequently occupies the first posture state during a second period of time after the first period of time, wherein at least one of the receiving, defining, and redefining is performed at least in part via one or more processors. 2. The method of claim 1, wherein the posture state reference data defines at least one of a posture volume or posture area. 3. The method of claim 2, wherein the posture volume comprises a posture cone, and wherein the posture state reference data includes a posture reference coordinate vector and a tolerance angle that define the posture cone. 4. The method of claim 2, wherein the posture state reference data comprises a posture reference coordinate vector and at least one cosine value. 5. The method of claim 2, wherein the posture volume comprises a posture donut, and wherein the posture state reference data includes a posture reference coordinate vector and at least one of a range of cosine values or range of angles that define the posture donut. 6. The method of claim 5, wherein the posture state reference data includes a plurality of posture state reference vectors associated with the posture donut. 7. The method of claim 6, wherein the posture state reference coordinate vector comprises a virtual posture state reference coordinate vector, the virtual posture state reference vector defined based at least in part on the plurality of posture state reference vectors associated with the posture donut. 8. The method of claim 1, further comprising associating therapy information with the defined posture state reference data, and subsequently associating at least some of the therapy information with the redefined posture state reference data. 9. The method of claim 8, wherein the therapy information comprises therapy objectification information and parameter information. 10. The method of claim 9, wherein the parameter information comprises different therapy parameter values for therapy delivered when the patient occupies different posture states. 11. The method of claim 8, wherein the therapy information comprises one or more adjustments to posture state reference data made subsequent to defining the posture state reference data. 12. The method of claim 8, wherein the therapy information comprises a tolerance value defined for the posture state reference data prior to redefining the posture state reference data, wherein redefining the posture state reference data comprises redefining a posture reference coordinate vector, andwherein subsequently associating the therapy information with t... | Medtronic, Inc. | A61B005/00
US08631713 | B2 | 2014-01-21 | Device for measuring compressive force of flexible linear body | 2007-08-30 | A measurement device capable of sensing presence of an obstacle within a vessel from outside the vessel in operating a linear body inserted in the vessel is provided. According to this measurement device, an easy-to-use measurement device capable of measuring compressive force in a direction of longitudinal axis applied to the linear body that has a sheath thicker than the linear body, such as a wire having a coil for embolizing an aneurysm attached at a tip end, without lowering measurement accuracy, can be provided. In addition, a length of a restraint portion is minimized, to achieve a smaller size of the measurement device. Moreover, the measurement device capable of measuring compressive force in the direction of longitudinal axis applied to the linear body regardless of magnitude of buckling load of the linear body can be provided, and the same measurement device is applicable to linear bodies of various materials, which leads to cost efficiency. | 1. A measurement device measuring compressive force applied to a linear body having flexibility in a direction of longitudinal axis, comprising: a main body including a through hole through which said linear body passes, wherein said main body includes a fixed curved portion having an inner space to allow bending of said linear body disposed therein and said linear body is bent in a direction within said through hole inside said fixed curved portion of said main body when said compressive force is applied to said linear body;a sensor for detecting a degree of bending;a conversion circuit for converting detected said degree of bending into said compressive force applied to said linear body, based on predetermined correlation between said degree of bending and compressive force; anda groove is disposed within said main body in an inner wall of said through hole along said through hole. 2. The measurement device according to claim 1, wherein said groove is formed along said inner wall of said through hole located on an inner side of bending of said linear body within said through hole in which said linear body is bent, andsaid through hole is disposed such that the distance between said inner wall of said through hole located on an outer side of bending of said linear body and said inner wall of said through hole located on the inner side of bending of said linear body is greater than a sum of a width of said groove and a diameter of said linear body to form said inner space within the fixed curved portion of said main body in which said linear body is bent. 3. The measurement device according to claim 1, incorporated in medical equipment for use. 4. The measurement device according to claim 1, attached to a training simulator simulating a human body for use. 5. A measurement device measuring compressive force applied to a linear body having flexibility in a direction of longitudinal axis, comprising: a main body including a through hole through which said linear body passes, wherein said main body includes a fixed curved portion having an inner space to allow bending of said linear body disposed therein and said linear body is bent in a direction within said through hole inside said fixed curved portion of said main body when said compressive force is applied to said linear body;a sensor for detecting a degree of bending;a conversion circuit for converting detected said degree of bending into said compressive force applied to said linear body based on predetermined correlation between said degree of bending and compressive force, wherein:said through hole includes restraint portions restricting movement of said linear body in a direction other than said direction of longitudinal axis, at opposing end portions of said through hole, andsaid linear body is in parallel to said restraint portions outside a port of said main body through which said linear body passes, when said linear body passes through said through hole and no external force other ... | National University Corporation Nagoya Institute of Technology NTN Corporation | G01L001/24
US08788009 | B2 | 2014-07-22 | High impedance signal detection systems and methods for use in electrocardiogram detection systems | 2009-07-17 | A biomedical sensor system is disclosed that includes a high impedance conductive electrode having an electrode impedance of at least about 20 kΩ/sq-mil, and a dielectric material on a first side of the electrode for receiving a discharge of an electrical signal from the dielectric material responsive to the presence of a time varying signal adjacent a second side of the dielectric material that is opposite the first side. | 1. A biomedical sensor system comprising a high impedance conductive electrode having an electrode impedance of at least about 20 kΩ/sq-mil, and a dielectric material on a first side of said electrode for receiving a discharge of an electrical signal from the dielectric material responsive to the presence of a time varying signal adjacent a second side of said dielectric material that is opposite the first side. 2. The biomedical sensor system as claimed in claim 1, wherein said dielectric material is an adhesive. 3. The biomedical sensor system as claimed in claim 1, wherein said dielectric material includes multiple electrodes on the first side thereof. 4. The biomedical sensor system as claimed in claim 1, wherein said conductive electrode includes a conductive polymer. 5. The biomedical sensor system as claimed in claim 1, wherein said conductive electrode includes any of aluminum, silver, silver chloride, or a conductive graphite. 6. The biomedical sensor system as claimed in claim 1, wherein said conductive electrode is printed onto a substrate. 7. The biomedical sensor system as claimed in claim 1, wherein said conductive electrode includes a conductive carbon coating. 8. The biomedical sensor system as claimed in claim 1, wherein said biomedical sensor system include a monitor system that is coupled to the conductive electrode via a conductive path, and wherein a resistance of the conductive path between the conductive electrode and the monitor system is at least about 1 Ω/sq-mil. 9. The biomedical sensor system as claimed in claim 1, wherein the high impedance conductive electrode and the dielectric material have a combined thickness of less than about 250 microns. 10. A biomedical sensor system comprising a first conductive electrode and a second conductive electrode that is adjacent the first conductive electrode, said first and second conductive electrodes being provided in contact with an anisotropic signal receptive material that is contiguous with both the first conductive electrode and the second conductive electrode and wherein the contiguous receptive material provides that a biomedical signal from a first side of the signal receptive material is detectable by the first conductive electrode that opposes the biomedical signal but is not detectable by the second conductive electrode that does not oppose the biomedical signal irrespective of the thickness of the signal receptive material. 11. The biomedical sensor system as claimed in claim 10, wherein said system further includes an array of conductive electrodes in contact with the signal receptive material such that the signal receptive material is contiguous with each of a plurality of electrodes in the array of conductive electrodes. 12. The biomedical sensor system as claimed in claim 10, wherein said signal receptive material is an adhesive that exhibits dielectric dispersion. 13. The biomedical sensor system as claimed in claim 10, wherein each of said first and... | FLEXcon Company, Inc. | A61B005/0402 H01B001/00 A61B005/0478 A61B005/04
US09596998 | B2 | 2017-03-21 | Ballistocardiographic sensor system with a sensor arrangement and method of ballistocardiographic detection of body movements | 2008-02-18 | Sensor arrangement, comprising at least one ballistocardiographic sensor for detecting body movements of a subject and at least one noise sensor for detecting vibration, method of ballistocardiographic detection of body movements, ballistocardiographic sensor system and uses of the sensor arrangement. It is suggested to pickup an external noise signal using an inertia sensor that is vibration isolated from a subject under ballistocardiographic examination. | 1. A sensor arrangement, comprising: at least one ballistocardiographic sensor configured to detect body movements of a subject;a structure configured to support a body of the subject;at least one noise sensor configured to detect vibration; anda mechanical mounting affixed to the structure, the mounting being configured to: vibrationally isolate the noise sensor from the body,protrude the noise sensor from the structure to a working position, andretract the noise sensor to a stowed position. 2. The sensor arrangement according to claim 1, wherein the ballistocardiographic sensor is arranged at the structure. 3. The sensor arrangement according to claim 1, wherein the noise sensor is a seismic sensor which is arranged in close proximity to the ballistocardiographic sensor and/or to the structure. 4. The sensor arrangement according to claim 1, wherein the noise sensor and the ballistocardiographic sensor are arranged such that vibrations which do not result from the body movements of the subject are detected by both the noise sensor and the ballistocardiographic sensor. 5. A ballistocardiographic sensor system comprising: the sensor arrangement according to claim 1; anda data processor configured to process signals from the noise sensor and the ballistocardiographic sensor. 6. The sensor arrangement according to claim 1, wherein said sensor arrangement is configured for monitoring one or more of heart rate, respiration rate and activity of a subject. 7. The sensor arrangement according to claim 1, wherein said sensor arrangement is configured for surveillance of a person in a care environment. 8. The sensor arrangement according to claim 1, wherein said sensor arrangement is configured for monitoring arrhythmia. 9. The sensor arrangement according to claim 1, wherein said sensor arrangement is configured for assessing sleep quality. 10. The sensor arrangement according to claim 1, wherein said sensor arrangement is configured for surveillance of congestive heart failure patients. 11. A sensor arrangement comprising: at least one ballistocardiographic sensor for detecting body movements of a subject;a structure supporting a body of the subject, the structure including an undercarriage and a fixing brake;a mechanical mounting affixed to the structure; andat least one noise sensor movably attached to the structure via the mechanical mounting for detecting vibration;wherein the noise sensor is vibrationally isolated from the body, and wherein the mechanical mounting is operationally connected to the brake such that the noise sensor is protruded from the structure upon the fixing brake being applied. 12. A method of ballistocardiographic detection of body movements of a subject, comprising: detecting a ballistocardiographic signal from the body of the subject with a ballistocardiographic sensor,with a mounting, mechanically protruding a noise sensor from a structure that supports the body of the subject to a working position and retracting ... | KONINKLIJKE PHILIPS N.V. | A61B005/103 A61B005/117 A61B005/024 A61B005/11 A61B005/113 A61B005/08 A61B005/00 A61B005/0205
US20100004555 | A1 | 2010-01-07 | APPARATUSES AND METHODS FOR EXTRACTING CHEMICALS FROM THE ORAL CAVITY AND BREATH | 2009-03-10 | The present disclosure generally pertains to chemical extraction apparatuses and methods for extracting volatile, semi-volatile, and non-volatile chemicals from a user's oral cavity. In one exemplary embodiment, a chemical extraction apparatus is composed of absorbent material, which is positioned within the oral cavity of a user for a period of time. Volatile, semi-volatile, and non-volatile chemicals are extracted from the breath and saliva of the user. By keeping the absorbent material in the oral cavity for an extended period of time, such as several minutes or hours depending on the types of materials selected, even trace levels of a chemical can be concentrated in the absorbent material thereby enabling conventional analytical techniques to detect the chemical. | 1. A chemical extraction apparatus, comprising: a support element configured to secure the chemical extraction apparatus to at least one tooth of a user; anda sampling element coupled to the support element, the sampling element composed of an absorbent material for absorbing or adsorbing volatile chemicals from an oral cavity of the user. 2. The chemical extraction apparatus of claim 1, further comprising an arm for coupling the sampling element to the support element. 3. The chemical extraction apparatus of claim 1, wherein the support element is configured such that the chemical extraction apparatus is secured to the at least one tooth when the at least one tooth is positioned between the support element and the sampling element such that the support element and the support element press against the at least one tooth. 4. The chemical extraction apparatus of claim 1, wherein the support element is configured to wrap around the at least one tooth to secure the chemical extraction apparatus to the at least one tooth. 5. The chemical extraction apparatus of claim 1, wherein the absorbent material is hydrophobic. 6. The chemical extraction apparatus of claim 1, wherein the absorbent material is composed of polydimethylsiloxane. 7. The chemical extraction apparatus of claim 1, wherein the sampling element is positioned behind a front tooth of the user when the chemical extraction apparatus is secured to the at least one tooth. 8. The chemical extraction apparatus of claim 1, wherein the support element comprises an elastic band. 9. The chemical extraction apparatus of claim 8, wherein the elastic band is configured to wrap around the at least one tooth to secure the chemical extraction apparatus to the at least one tooth. 10. The chemical extraction apparatus of claim 1, wherein the sampling element has a length less than 0.4 inches and a width less than 0.08 inches. 11. A method, comprising the steps of: inserting a chemical extraction apparatus into an oral cavity of a user, the chemical extraction apparatus having an absorbent material;securing the chemical extraction apparatus to at least one tooth of the user; andabsorbing volatile compounds from breaths of the user into the absorbent material. 12. The method of claim 11, further comprising the steps of: analyzing the absorbent material; andidentifying at least one of the volatile compounds based on the analyzing step. 13. The method of claim 12, further comprising the step of determining a concentration of the at least one volatile compound based on the analyzing step. 14. The method of claim 11, further comprising the step of cutting the absorbent material from the chemical extraction apparatus. 15. The method of claim 11, wherein the chemical extraction apparatus comprises a support element, an arm, and a sampling element, wherein the sampling element is composed of the absorbent material, wherein the sampling element is coupled to the support element via the arm, and wherein ... | - | A61B005/097 A61B005/00
US20100022861 | A1 | 2010-01-28 | IMPLANTABLE OPTICAL HEMODYNAMIC SENSOR INCLUDING AN EXTENSION MEMBER | 2008-07-28 | An implantable medical device includes an optical hemodynamic sensor comprising at least one optical emitter and at least one detector. In some examples, the at least one optical emitter may be optically coupled to at least one light transmission member that extends from a housing of the medical device. In addition, in some examples, the at least one detector may be optically coupled to at least one light transmission member that extends from the housing of the medical device. In other examples, an optical emitter and/or detector of a hemodynamic sensor may be carried by an elongated member that extends from a housing of the IMD. The elongated member may electrically couple the optical emitter and/or detector to a controller or other components within the IMD housing. | 1. An implantable medical system comprising: an implantable housing;an extension member extending from the housing;an optical emitter that emits light; anda detector that generates a signal indicative of an intensity of light emitted by the optical emitter and transmitted to the detector, the signal indicating a hemodynamic characteristic of a patient,wherein at least one of the optical emitter or the detector are carried by the extension member. 2. The implantable medical system of claim 1, wherein the extension member comprises an electrical conductor electrically coupled to the at least one of the emitter or the detector. 3. The implantable medical system of claim 2, wherein the extension member comprises a flexible circuit. 4. The implantable medical system of claim 2, further comprising a controller that controls an operation of the optical emitter and the detector, wherein the electrical conductor electrically couples the at least one of the optical emitter or the detector to the controller. 5. The implantable medical system of claim 1, wherein the extension member is movable relative to the housing. 6. The implantable medical system of claim 1, wherein the extension member is substantially rigid. 7. The implantable medical system of claim 1, wherein the optical emitter comprises at least one of a light emitting diode, a laser diode or a vertical cavity surface emitting laser. 8. The implantable medical system of claim 1, wherein the extension member has a smaller cross-sectional size than the housing. 9. The implantable medical system of claim 1, wherein the optical emitter is electrically coupled to the extension member and carried by the extension member. 10. The implantable medical system of claim 1, wherein the detector is electrically coupled to the extension member and carried by the extension member. 11. The implantable medical system of claim 1, wherein the extension member comprises a first extension member, the system further comprising a second extension member, wherein the optical emitter is carried by the first extension member and the detector is carried by the second extension member. 12. The implantable medical system of claim 11, wherein the first and second extension members are substantially parallel. 13. The implantable medical system of claim 1, further comprising a processor that receives the signal generated by the detector and determines a hemodynamic characteristic of a patient based on the signal. 14. The implantable medical system of claim 13, wherein the hemodynamic characteristic comprises at least one of a blood oxygen saturation level, blood flow, a hematocrit level or tissue perfusion. 15. An implantable medical system comprising: an implantable housing;an optical emitter that emits light;a detector that generates a signal indicative of an intensity of light emitted by the optical emitter and transmitted to the detector;a flexible circuit electrically coupled to at least one of the optical emit... | Medtronic, Inc. | A61B005/1459 A61B005/026 A61B005/00
US20100030063 | A1 | 2010-02-04 | SYSTEM AND METHOD FOR TRACKING AN INSTRUMENT | 2008-07-31 | A system for tracking an instrument relative to an anatomical structure is provided. The system can include at least one tracking device, which can be coupled to the instrument. The system can also include a shape sensor coupled to the instrument that can determine a shape of the instrument. The system can include a tracking system that can track a position of the at least one tracking device relative to the anatomical structure. The system can further include a navigation system that can determine a position and shape of the instrument relative to the anatomical structure based on the position of the at least one tracking device determined by the tracking system and the shape of the instrument as sensed by the shape sensor. | 1. A system for tracking an instrument relative to an anatomical structure comprising: at least one tracking device coupled to the instrument;a shape sensor coupled to the instrument that determines a shape of the instrument;a tracking system that tracks a position of the at least one tracking device relative to the anatomical structure; anda navigation system that determines a position and shape of the instrument relative to the anatomical structure based on the position of the at least one tracking device determined by the tracking system and the shape of the instrument as sensed by the shape sensor. 2. The system of claim 1, further comprising: an imaging device that is operable to acquire an image of the anatomical structure. 3. The system of claim 2, further comprising: a display that displays the image of the anatomical structure superimposed with an icon of the instrument at a location that corresponds to the position of the instrument relative to the anatomical structure, and displays the shape of the instrument. 4. The system of claim 3, wherein the instrument is an elongated instrument, and includes a proximal end and a distal end, and the shape sensor is able to determine the shape of the instrument from a region proximate the distal end to a region proximate the proximal end. 5. The system of claim 4, wherein the instrument is selected from the group comprising: catheters, basket catheters, balloon catheters, leads, guidewires, sheaths, endoscopes, ablation catheters, arthroscopic systems, orthopedic implants, spinal implants, deep-brain stimulator (DBS) probes, drug delivery systems, mapping catheters, drill bits, stylets, trocars, screws or combinations thereof. 6. The system of claim 4, wherein the at least one tracking device comprises a plurality of tracking devices, with at least one of the plurality of tracking devices coupled to the proximal end of the instrument, at least one of the plurality of tracking devices coupled to the distal end of the instrument, and at least one of the plurality of tracking devices coupled between the proximal end and the distal end of the instrument, and the shape sensor is located proximate to the at least one tracking device coupled to the proximal end, the at least one tracking device coupled to the distal end and the at least one tracking device coupled between the proximal end and the distal end. 7. The system of claim 6, wherein the navigation system outputs a notification message to the display if a position of the distal end of the instrument determined from the tracking of the at least one tracking device at the distal end of the instrument does not substantially correspond to a position of the distal end of the instrument determined from the shape sensor. 8. The system of claim 6, wherein the at least one tracking device comprises at least one optical tracking device to track at least one degree of freedom information. 9. The system of claim 6, wherein the at least one tracking d... | Medtronic, Inc. | A61B005/05
US20100041971 | A1 | 2010-02-18 | IMPLANTABLE ANALYTE SENSOR | 2009-10-29 | An implantable analyte sensor including a sensing region for measuring the analyte and a non-sensing region for immobilizing the sensor body in the host. The sensor is implanted in a precisely dimensioned pocket to stabilize the analyte sensor in vivo and enable measurement of the concentration of the analyte in the host before and after formation of a foreign body capsule around the sensor. The sensor further provides a transmitter for RF transmission through the sensor body, electronic circuitry, and a power source optimized for long-term use in the miniaturized sensor body. | 1. An electrochemical analyte sensor for measuring an analyte concentration, the sensor comprising: a sensor body comprising electronic circuitry encapsulated within the sensor body; anda plurality of electrodes that extend from an outer surface of the sensor body to the encapsulated electronic circuitry, wherein the electrodes are mechanically and electrically connected and aligned to the electronic circuitry prior to encapsulation within the sensor body. 2. The sensor according to claim 1, wherein the electrodes are swaged to the electronic circuitry. 3. The sensor according to claim 1, wherein the electrodes are welded using a technique selected from the group consisting of spot welding, ultrasonic welding, and laser welding. 4. The sensor according to claim 1, wherein the electrodes and electronic circuitry are encapsulated in the sensor body by a molding process. 5. The sensor according to claim 1, wherein the sensor body comprises a water vapor permeable material. 6. The sensor according to claim 5, wherein the electronic circuitry is spaced from the water vapor permeable sensor body, such that water vapor penetration within a fixed distance from the electronic circuitry is inhibited. 7. The sensor according to claim 6, wherein the electronic circuitry is spaced from the water vapor permeable sensor body by epoxy. 8. The sensor according to claim 7, wherein the electronic circuitry is spaced from the water vapor permeable sensor body by a glass tube. 9. The sensor according to claim 7, wherein the electronic circuitry is spaced from the water vapor permeable sensor body by parylene. 10. The sensor according to claim 7, wherein the electronic circuitry is spaced from the water vapor permeable sensor body by one or more hermetic containers. 11. The sensor according to claim 1, wherein the sensor body comprises a substantially seamless exterior with the electrodes extending through the sensor body to an outer surface thereof. 12. An implantable analyte sensor, comprising: electronics encapsulated within a water vapor permeable body, wherein the electronics comprise a microprocessor module and an RF module that has an RF transceiver with a phase-locked loop, and wherein the microprocessor module is programmed to initiate re-calibration of the phase-locked loop responsive to detection of off-frequency shift. 13. An electrochemical glucose sensor comprising a three-electrode system, the sensor comprising: an electrochemical cell comprising a working electrode, reference electrode, and counter electrode; anda potentiostat configured to control a potential between the working electrode and the reference electrode, wherein an allowable range for the counter electrode voltage is set sufficiently wide such that the glucose sensor is configured to react with other reducible species when oxygen becomes limited and sufficiently narrow to ensure the circuitry does not allow excessive current draw or bubble formation to occur when in use. 14. ... | DexCom, Inc. | A61B005/1468 H01R043/00
US20100004552 | A1 | 2010-01-07 | METHOD AND DEVICE FOR THE DETERMINATION OF BREATH FREQUENCY | 2007-11-28 | The present invention provides a method for determining the respiratory rate of a patient comprising the steps of determining at least two time dependent respiratory signals by at least two different methods as well as a determining of a respiratory rate on the basis of the at least two time dependent respiratory signals. In this connection, the resulting respective instantaneous respiratory rates fi(n) (i=1, 2, . . . ) are determined from the at least two time dependent respiratory signals si(t) (i=1, 2, . . . ) and an average respiratory rate f(n) determined by a weighted averaging of the respiratory rates fi(n) (i=1, 2, . . . ) is produced. In this averaging, the weightings ki(n) (i=1, 2, . . . ) of the individual respiratory rates fi(n) (i=1, 2, . . . ) depend on a difference between the respective respiratory rates fi(n) (i=1, 2, . . . ) and an estimate fi(n) which is determined on the basis of at least two respiratory signals si(t) (i=1, 2, . . . ). An apparatus for carrying out the method is likewise provided. | 1. A method of determining the respiratory rate of a patient comprising the steps: determining at least two time dependent respiratory signals si(t) (i=1,2, . . . ) by at least two different methods;determining the resulting respective instantaneous respiratory rates fi(n) (i=1,2, . . . ) from the at least two time dependent respiratory signals si(t) (i=1, 2, . . . )determining an average respiratory rate f(n) by a weighted averaging of the respiratory rates fi(n) (i=1,2, . . . ), characterized in thatthe weightings ki(n) (i=1,2, . . . ) of the individual respiratory rates fi(n) (i=1,2, . . . ) depend on a difference between the respective respiratory rates fi(n) (i=1,2, . . . ) and an estimate fs(n) determined on the basis of at least two respiratory signals si(t) (i=1,2, . . . ). 2. A method in accordance with claim 1, wherein the estimate fs(n) is determined on the basis of a preceding, already determined average respiratory rate f(n−1). 3. A method in accordance with claim 1, wherein the estimate fs(n) is determined, in particular for initialization by a combination of rate information from at least two time dependent respiratory signals si(t) (i=1,2, . . . ) or by forming an average of the current respiratory rates fi(n) (i=1, 2, . . . ). 4. A method in accordance with claim 1, wherein the at least two time dependent respiratory signals si(t) (i=1,2, . . . ) are determined from measured physiological signals. 5. A method in accordance with claim 4, wherein the measured physiological signals form a selection from the following signals: bioimpedance signal;heart rate variability signal;photoplethysmographic signal (PPG signal);statistical source signal of the ECG;pulse wave transit time signal (PTT signal). 6. A method in accordance with claim 4, wherein the determination of the at least two time dependent respiratory signals si(t) (i=1,2, . . . ) takes place from the measured physiological signals by a band pass filter. 7. A method in accordance with claim 6, wherein the band pass filter allows frequencies in a range from approx. 0.12 Hz to 0.42 Hz to pass. 8. A method in accordance with claim 1, wherein the determination of the respective instantaneous respiratory rates fi(ki), in particular of fhr(m), famp(n) and fpt(k), from the time dependent respiratory signals si(t) (i=1,2, . . . ) takes place by determining the time indices tmax(ki), in particular tmax(m), tmax(n) and tmax(k), of the maxima of the time dependent respiratory signals si(t) (i=1,2, . . . ). 9. A method in accordance with claim 8, wherein a consistency check takes place for the time indices ki (i=1, 2, . . . ), in particular m, n and k, of the instantaneous respiratory rates. 10. A method in particular in accordance with claim 1, wherein a consistency check of the respiratory rates fi(n) (i=1,2, . . . ) is carried out. 11. A method in accordance with claim 10, wherein the consistency check takes place by a comparison of the respiratory rates fi(n) (i=1,2, . . . )... | FRESENIUS MEDICAL CARE DEUTSCHLAND GMBH | A61B005/08
US20100030035 | A1 | 2010-02-04 | Fuzzy system for cardiovascular disease and stroke risk assessment | 2008-08-04 | The present invention relates to a fuzzy system that provides a measure of global cardiovascular risk based on the risk factors associated with cardiovascular disease and stroke. The fuzzy system is of multiple input single output (MISO) wherein the inputs are the values of risk factors and the output is the measure of global cardiovascular risk. | 1. A fuzzy system for measuring cardiovascular disease and stroke risk with reference to type II diabetic patients, comprising one or more risk factors; anda fuzzy engine having single input single output fuzzy sub-engines equal in number to said risk factors whereby each risk factor is inserted into a corresponding sub-engine, a fusion module for accepting outputs from said subengines and delivering a global risk measurement. 2. The fuzzy system for measuring cardiovascular disease and stroke risk with reference to type II diabetic patients in claim 1, wherein said risk factors can be one or more selected from the group consisting of gender, age, systolic blood pressure, diastolic blood pressure, body mass index (BMI), waist circumference, total cholesterol, low-density lipoprotein cholesterol, high density lipoprotein cholesterol, triglycerides, homocysteine, fasting glucose, diabetes mellitus, serum creatine, creatine clearance, smoking status (yes or no), history of smoking, medication use, vascular disease, fibrinogen, HbA1c, APOE Gene polymorphism, pulsatility indes, spectral broading index (SBI) of transcranial doppler ultrasound waveform, EDV ratio, carotid intima-media thickness (IMT), coronary artery calcium (CAC) score, flow-mediated endothelial vasodilatation, family history of CHD &amp; stroke, and diet. 3. The fuzzy system for measuring cardiovascular disease and stroke risk with reference to type II diabetic patients in claim 1, wherein said fuzzy sub-engine comprise fuzzy inferences that can be Mamdani-style or Sugeno-style. 4. A fuzzy system for measuring cardiovascular disease and stroke risk with reference to type II diabetic patients, comprising one or more risk factors; anda fuzzy engine having one or more single input single output fuzzy sub-engines, one or more multiple input single output fuzzy sub-engines, a fusion module for accepting outputs from said sub-engines, wherein said risk factors can correspond to a single input single output fuzzy sub-engine or two or more risk factors can correspond to the same multiple input output sub-engine. 5. The fuzzy system for measuring cardiovascular disease and stroke risk with reference to type II diabetic patients in claim 4, wherein said risk factors can be one or more selected from the group consisting of gender, age, systolic blood pressure, diastolic blood pressure, body mass index (BMI), waist circumference, total cholesterol, low-density lipoprotein cholesterol, high density lipoprotein cholesterol, triglycerides, homocysteine, fasting glucose, diabetes mellitus, serum creatine, creatine clearance, smoking status (yes or no), history of smoking, medication use, vascular disease, fibrinogen, HbA1c, APOE Gene polymorphism, pulsatility index, spectral broadening index (SBI) of transcranial doppler ultrasound waveform, EDV ratio, carotid intima-media thickness (IMT), coronary artery calcium (CAC) score, flow-mediated endothelial vasodilatation, family history of CHD &amp; s... | The Hong Kong Polytechnic University | A61B005/00 G06Q050/00
US20100010365 | A1 | 2010-01-14 | Apparatus for analyzing brain wave | 2009-05-05 | To increase intensity of the brain wave signal for detection. Provided is an apparatus for analyzing a brain wave which is installed on a vehicle comprising: a detection unit for detecting the brain wave signal, and separating and analyzing the detected brain wave signal; a discrimination unit for generating a control signal according to an intensity of the brain wave signal analyzed by the detection unit; a processing control unit for controlling subsequent processing according to a type of each of the plurality of control signals produced by the discrimination unit; at least one brain wave signal induction unit for generating a graphic inducing a predetermined type of the brain wave; and a display for displaying the generated graphic according to the received signal from the brain wave signal induction unit. | 1. An apparatus for analyzing a brain wave which is installed on a vehicle comprising: a detection unit for detecting the brain wave signal, and separating and analyzing the detected brain wave signal;a discrimination unit for generating a control signal according to an intensity of the brain wave signal analyzed by the detection unit;a processing control unit for controlling subsequent processing according to a type of each of the plurality of control signals produced by the discrimination unit;at least one brain wave signal induction unit for generating a graphic inducing a predetermined type of the brain wave; anda display for displaying the generated graphic according to the received signal from the brain wave signal induction unit. 2. The apparatus according to claim 1, wherein the discrimination unit generates control signal by repeating at least two stages in a hierarchy, each corresponding to at most six control signals. 3. The apparatus according to claim 2, wherein the each generation of the plurality of control signal at least two stages is performed by at least two of the brain wave signal induction unit different from each other. 4. The apparatus according to claim 1, further comprising an input unit for controlling a start of an operation of the discrimination unit. 5. The apparatus according to claim 1, wherein the detection unit is installed on one of a ceiling above a seat and a headrest of the seat. 6. The apparatus according to claim 1, wherein the detection unit detects the brain wave without contacting user's head. 7. The apparatus according to claim 1, further comprising a storage unit for storing pattern of the brain wave signal, wherein the discrimination unit compares the pattern of brain wave signal stored in the storage unit and the brain wave signal detected by the detection unit, and discriminates between a brain wave caused by a primary task such as driving a vehicle and a brain wave caused by controlling an electronic device for a task other than the primary task. 8. The apparatus according to claim 1, further comprising a storage unit for storing pattern of brain wave signal inappropriate for the primary task, wherein the discrimination unit compares the pattern of brain wave signal stored in the storage unit and the brain wave signal detected by the detection unit, and detects abnormality which includes a state of driving under an influence of alcohol, a state of working under the influence of alcohol, a state of sleeping while driving, a state of sleeping while working, a state of driving while looking aside for a long period, and a state of working while looking aside for a long period. 9. The apparatus according to claim 1, wherein the detection unit comprises an individual difference determination/correction unit for determining and correcting the detected brain wave signal while considering an individual difference thereof to thereby generate a signal used for control. 10. The apparatus according to ... | Hitachi, Ltd. | A61B005/0476
US20100016739 | A1 | 2010-01-21 | Method of Assessing Blood Volume Using Photoelectric Plethysmography | 2009-09-28 | A method and system for assessing blood volume within a subject includes generating a cardiovascular waveform representing physiological characteristics of a subject and determining blood volume of the subject by analyzing the cardiovascular waveform. The step of analyzing includes generating a first trace of the per heart-beat maximums of the cardiovascular waveform, which is representative of the systolic pressure upon the cardiovascular signal, generating a second trace of the per heart-beat minimums of the cardiovascular waveform, which is representative of the diastolic pressure upon the cardiovascular signal, and comparing the respective first trace and the second trace to generate an estimate of relative blood volume within the subject. In accordance with an alternate method of analyzing harmonic analysis is applied to the cardiovascular waveform, extracting a frequency signal created by ventilation and applying the extracted frequency signal in determining blood volume of the subject. | 1-65. (canceled) 66. A method enabling detection of physiological changes in a subject, the method comprising: generating a cardiovascular waveform representing physiological characteristics of a subject; andanalyzing the cardiovascular waveform, wherein the analyzing the cardiovascular waveform includes: identifying a first component of the cardiovascular waveform related to the amplitude of the cardiovascular waveform; andcalculating an index responsive to changes in the first component of the cardiovascular waveform; wherein index variations are reflective of physiological changes in the subject. 67. The method of claim 66, further comprising means for displaying the index. 68. The method of claim 66, wherein a processor is used for the step of analyzing the cardiovascular waveform. 69. The method of claim 66, wherein the first component is based on differences between corresponding per heart-beat maximums and per heart-beat minimums of the cardiovascular waveform. 70. The method of claim 66, wherein the index is responsive to changes in differences between corresponding per respiratory-cycle maximums and per respiratory-cycle minimums of the first component. 71. The method of claim 66, wherein the analyzing the cardiovascular waveform further includes normalizing the first component of the cardiovascular waveform, wherein the index is responsive to changes in the normalized component of the cardiovascular waveform. 72. The method of claim 66, wherein the index is responsive to changes in a plurality of components of the cardiovascular waveform. 73. The method of claim 66, further comprising monitoring the physiological changes in the subject based on the index. 74. The method of claim 66, wherein the physiological changes in the subject are changes in blood volume of the subject. 75. The method of claim 66, further comprising monitoring fluid responsiveness based on the index variations. 76. The method of claim 66, further comprising detecting at least one of (i) hypovolemia and (ii) dehydration, based on the index variations. 77. The method of claim 66, wherein the cardiovascular waveform is a photoelectric plethysmographic signal. 78. The method of claim 66, wherein the cardiovascular waveform is analyzed in one of: (i) the time domain, (ii) the frequency domain and (iii) the joint time-frequency domain. 79. A method for detecting physiological changes in a subject, the method comprising: generating a cardiovascular waveform representing physiological characteristics of a subject; andanalyzing the cardiovascular waveform, wherein the analyzing the cardiovascular waveform includes: identifying a first component of the cardiovascular waveform;identifying a second component of the cardiovascular waveform;wherein one of the first and second components of the cardiovascular waveform is related to the amplitude of the cardiovascular waveform;comparing the first and second components of the cardiovascular waveform relative to one ano... | YALE UNIVERSITY | A61B005/02 A61B005/0402
US20100049039 | A1 | 2010-02-25 | System and Method of Treating Skin Imperfections Including a Diagnostic Tool | 2009-11-10 | A system and method of treating skin imperfections including a diagnostic tool is disclosed. In a particular embodiment the system includes a diagnostic tool to scan a preselected area of skin of a client to detect at least one skin imperfection and an associative module to associate a skin value to the client. Further, the system includes that a first unadulterated ingredient of alpha lipoic acid, dimethylaminoethenol, vitamin A, vitamin C, and vitamin E, or any combination thereof, is selected based on the skin value, and a second unadulterated ingredient of retinol, alpha and beta hydroxyl acids, or any combination thereof, is selected based on the skin value. In addition, a customized product is compounded using selected first and second unadulterated ingredients and the product is compounded manually by a technician at a point of sale. | 1. A method of treating skin imperfections, the method comprising: scanning a preselected area of skin of a client using a diagnostic tool;detecting a first edge and a second edge of at least one skin imperfection within the preselected area using the diagnostic tool;detecting a surficial depth of the at least one skin imperfection using the diagnostic tool;determining a width of the at least one skin imperfection using a spatial difference between the first edge relative to the second edge;calculating a volume of the at least one skin imperfection using the depth and width of the at least one skin imperfection;associating a skin value to the client based on a total volume of the at least one skin imperfection and a size of the preselected area of the skin of the client;manually selecting a first unadulterated ingredient from a plurality of unadulterated ingredients based on the skin value;manually selecting a second unadulterated ingredient from the plurality of ingredients based on the skin value; andcompounding the first unadulterated ingredient and the second unadulterated ingredient to form a customized product at the point of sale, wherein the compounding is manually completed by a technician. 2. The method of claim 1, wherein the diagnostic tool includes a photodetector and a laser that emits coherent light, the photodetector responsive to the electromagnetic reflections from the skin. 3. The method of claim 1, further comprising treating a cause of the at least one skin imperfection of the client by applying a portion of the customized product to an affected area of the at least one skin imperfection at the point of sale. 4. The method of claim 1, further comprising aiming the diagnostic tool at the preselected area of skin and activating a trigger. 5. The method of claim 4, wherein the diagnostic tool further comprising a memory for storing the skin value of the client. 6. The method of claim 5, further comprising displaying the skin value to the client. 7. The method of claim 5, further comprising referencing the diagnostic tool at a start of each scanning session. 8. The method of claim 1, wherein the first unadulterated ingredient is one of alpha lipoic acid, dimethylaminoethenol, vitamin A, vitamin C, and vitamin E. 9. The method of claim 1, wherein the second unadulterated ingredient is one of retinol, alpha and beta hydroxyl acids. 10. A method of treating skin imperfections, the method comprising: manually selecting a first unadulterated ingredient;manually selecting a second unadulterated ingredient; andcompounding the first unadulterated ingredient and the second unadulterated ingredient to form a customized product at the point of sale, wherein the compounding is manually completed by a technician. 11. A system of treating skin imperfections, the system comprising: a diagnostic tool to scan a preselected area of skin of a client to detect at least one skin imperfection;an associative module to associate a skin value t... | - | A61B006/00 A61B005/103
US07643862 | B2 | 2010-01-05 | Virtual mouse for use in surgical navigation | 2006-05-15 | A method of performing a surgery is provided including a surgical navigation system having a tracking system, computer and monitor placed outside of a sterile field. An input pad and a tracking array attachable to a surgical instrument or bone are placed into the sterile field along with a probe having a probe array. The tracking array and the probe array are acquired by the tracking system and a virtual mouse is activated by positioning the probe relative to the input pad, thereby causing a mouse input to the computer with the virtual mouse. | 1. A method of performing a surgery, comprising: providing a surgical navigation system having a tracking system, computer and monitor placed outside of a sterile field;placing a tracking array into the sterile field, the tracking array including a frame and being attachable to a surgical instrument being used by a surgeon to perform the surgery or a bone, wherein the frame of the tracking array has an input pad disposed on its outer surface;placing a probe having a probe array within the sterile field;acquiring the tracking array and the probe array with the tracking system;activating a virtual mouse by positioning the probe relative to the input pad; andcausing a mouse input to the computer with the virtual mouse. 2. The method of claim 1, wherein activating the virtual mouse with the probe comprises positioning a tip of the probe substantially near a pad marker located on the input pad to cause the mouse input. 3. The method of claim 2, wherein the mouse input comprises identifying a function on the monitor with a pointer. 4. The method of claim 3, further comprising moving the tip of the probe away from the pad marker to cause a second mouse input. 5. The method of claim 4, wherein the second mouse input comprises selecting the function on the monitor. 6. The method of claim 3, further comprising occluding the probe array to cause a second mouse input. 7. The method of claim 6, wherein the second mouse input comprises selecting the function on the monitor. 8. The method of claim 1, wherein activating the virtual mouse with the probe comprises moving the probe substantially near the input pad to cause a pointer on the monitor to correspondingly move. 9. The method of claim 1, wherein activating the virtual mouse with the probe comprises moving a tip of the probe along a substantially flat surface of the input pad to cause the mouse input. 10. The method of claim 9, wherein the mouse input comprises moving a pointer to correspondingly position a cursor on the monitor over an icon. 11. The method of claim 10, further comprising moving the tip of the probe away from the pad marker to cause a second mouse input. 12. The method of claim 11, wherein the second mouse input comprises selecting the function displayed by the monitor. 13. The method of claim 10, further comprising occluding the probe array to cause a second mouse input. 14. The method of claim 13, wherein the second mouse input comprises selecting the function displayed by the monitor. 15. The method of claim 1, further comprising moving the probe away from the input pad to manipulate a corresponding object on the computer monitor, the corresponding object being a human anatomy image. 16. A surgical navigation system, comprising: a computer having a monitor and surgical navigation utilities software;a tracking system coupled to the computer and establishing a measurement field;a tracking array recognizable by the tracking system when exposed to the measurement field, the t... | Biomet Manufacturing Corporation | A61B005/05
US07660617 | B2 | 2010-02-09 | Electrical impedance tomography using a virtual short measurement technique | 2005-11-10 | Typical electrical impedance tomography (“EIT”) systems apply current to a body under test and measure the induced voltages on the surface and then through an inverse process, reconstruct an approximation to the electrical conductivity in the interior. The EIT system described herein employs a new approach to evaluating the interior electrical conductivity by using a virtual short that allows both current and voltages to be measured at various locations on the body under test. The architecture of this system is described, distinguishability of embedded conductive anomalies is evaluated, typical error sources associated with a virtual short measurement is discussed in the context of electrical impedance tomography, and a new imaging system is discussed. | 1. An electrical impedance tomography (“EIT”) method for evaluating interior features of a region of a body under test with a plurality of electrodes, said method comprising: attaching a plurality of electrodes to the region of the body to form an electrode array surrounding the region, wherein each electrode of the plurality of electrodes comprises an electrode pad with at least one voltage lead and one current lead;designating a first one of the plurality of electrodes as a source electrode, and a second one of the plurality of electrodes as a sink electrode;establishing a virtual short between said source electrode and all other electrodes of the plurality of electrodes other than said sink electrode, wherein the virtual short is established by maintaining the source electrode and the all other electrodes at a predetermined voltage;applying a test current to said source electrode;measuring a sense voltage between said source electrode and said sink electrode; andrepeating the designating, establishing, applying, and measuring steps for different source and sink electrodes in the plurality of electrodes; andreconstructing an image of the region of the body under test based on the applied test current and the measured sense voltage. 2. The EIT method according to claim 1, wherein the repeating of said designating, establishing, applying, and measuring steps is carried out for all possible combinations of source and sink electrodes. 3. The EIT method according to claim 1, wherein reconstructing the image of the region of the body under test based on the applied test current and the measured sense voltage comprises: obtaining current data indicative of said test current;obtaining voltage data indicative of said sense voltage, wherein the sense voltage corresponds to the test current; andperforming a reconstruction procedure, based upon said current data and said voltage data, to obtain electrical characteristics of the interior features of the body under test. 4. The EIT method according to claim 3, image wherein obtaining the electrical characteristics of the of the interior features of the region of the body under test comprises calculating the impedance values from the stored current data and the stored voltage data. 5. The EIT method according to claim 3, wherein said reconstruction procedure determines an impedance for a section of the body under test, said section being located between said source electrode and said sink electrode. 6. An electrical impedance tomography (“EIT”) method for evaluating interior features of a region of a body under test, said method comprising: attaching a plurality of electrodes to form an electrode array surrounding the region of the body under test, wherein each electrode of the plurality of electrodes comprises an electrode pad with at least one voltage lead and one current lead;controlling a switching network to designate, from said plurality of electrodes, a source electrode and a sink electrode;mainta... | The Boeing Company | A61B005/05
US07657303 | B2 | 2010-02-02 | Guided radiation therapy system | 2003-12-22 | A system and method for accurately locating and tracking the position of a target, such as a tumor or the like, within a body. In one embodiment, the system is a target locating and monitoring system usable with a radiation delivery source that delivers selected doses of radiation to a target in a body. The system includes one or more excitable markers positionable in or near the target, an external excitation source that remotely excites the markers to produce an identifiable signal, and a plurality of sensors spaced apart in a known geometry relative to each other. A computer is coupled to the sensors and configured to use the marker measurements to identify a target isocenter within the target. The computer compares the position of the target isocenter with the location of the machine isocenter. The computer also controls movement of the patient and a patient support device so the target isocenter is co-incident with the machine isocenter before and during radiation therapy. | 1. An apparatus for preparing a patient for radiation treatment in a radiation device that directs a radiation beam to a beam isocenter, the apparatus comprising: a wireless marker configured to be implanted in the patient relative to a target wherein the marker is energized or excited by an excitation source positioned exterior of the patient;a movable patient support having a platform and a drive system to move the platform, wherein the platform is configured to situate the patient so that the marker can be sensed by a plurality of sensors in a fixed and known geometry relative to each other and the radiation beam can irradiate the patient; anda computer including computer operable instructions that cause the computer to perform the method of (a) receiving position information of the marker provided by the sensors, (b) determining a location of the marker relative to a reference frame using the position information, and (c) computing a displacement between the location of the marker and a desired location for the marker, wherein the target is located at a desired situs in the reference frame when the marker is at the desired location for the marker. 2. An apparatus for preparing a patient for radiation treatment in a radiation device that directs a radiation beam to a beam isocenter, the apparatus comprising: a wireless marker configured to be implanted in the patient relative to a target wherein the marker is energized by an external excitation source;a movable patient support having a platform and a drive system to move the platform, wherein the platform is configured to situate the patient so that the marker can be sensed by a plurality of sensors in a fixed and known geometry relative to each other and the radiation beam can irradiate the patient; anda computer including computer operable instructions that cause the computer to perform the method of (a) receiving position information of the marker provided by the sensors, (b) determining a location of the target relative to a reference frame using the position information of the marker, and (c) computing a displacement between the location of the target and a desired location of the target in the reference frame. 3. An apparatus for preparing a patient for radiation treatment in a radiation device that directs a radiation beam to a beam isocenter, the apparatus comprising: a wireless marker configured to be implanted in the patient relative to a target wherein the marker is configured to generate a signal and is not physically connected with an outside energy source;a movable patient support having a platform and a drive system to move the platform, wherein the platform is configured to situate the patient so that the marker can be sensed by a plurality of sensors in a fixed and known geometry relative to each other and the radiation beam can irradiate the patient; anda computer including computer operable instructions that cause the computer to perform the method of (a) receiving positio... | Calypso Medical Technologies, Inc. | A61B005/05 A61N005/10 A61B019/00
US07651466 | B2 | 2010-01-26 | Pulse contour method and apparatus for continuous assessment of a cardiovascular parameter | 2005-07-11 | A cardiovascular parameter such as cardiac output is estimated from a current pressure waveform data set without needing to directly measure blood flow or arterial compliance. The general shape of an input flow waveform over one beat-to-beat cycle is assumed (or computed), and then the parameters of a flow-to-pressure model, if not pre-determined, are determined using system identification techniques. In one embodiment, the parameters thus determined are used to estimate a current peripheral resistance, which is used not only to compute an estimate of the cardiovascular parameter, but also to adjust the shape of the input flow waveform assumed during at least one subsequent beat-to-beat cycle. Another embodiment does not require computation of the peripheral resistance and still another embodiment computes a flow estimate from an optimized identification of the parameters defining the assumed input flow waveform. | 1. A method for determining a cardiovascular parameter equal to or derivable from cardiac output (CO) comprising: inputting a current pressure waveform data set corresponding to arterial blood pressure over a current pressure cycle;determining defining parameters of an assumed input flow waveform as a function of a peripheral resistance value determined for at least one previous pressure cycle;determining model parameters of a model of a relationship between the assumed input flow waveform and the current pressure waveform data set;computing a current peripheral resistance value as a function of the model parameters; andcomputing an estimate of the cardiovascular parameter as a function of the current peripheral resistance value and the current pressure waveform data set. 2. A method as in claim 1, further comprising determining the defining parameters of the assumed input flow waveform also as a function of shape characteristics of the current pressure waveform data set. 3. A method as in claim 2, in which the assumed input flow waveform is a series of component waveforms, with one component waveform per pressure cycle. 4. A method as in claim 3, in which: the defining parameters include duration and amplitude; andthe duration of the component waveform for the current pressure cycle is set at least approximately equal to a time interval between systole onset and systole in the current pressure waveform data set. 5. A method as in claim 4, further comprising: estimating a diastolic time constant as a product of a sampling rate at which the pressure waveform data set is derived and a function of a model feedback parameter;estimating an arterial compliance value as a ratio of the diastolic time constant and the peripheral resistance value;estimating a systolic time constant from chosen points in the current pressure waveform data set;computing an aortic characteristic resistance value as a ratio of the systolic time constant and the arterial compliance value;setting the amplitude of the component waveform for the current pressure cycle to be inversely proportional to the square of a function of at least one aortic characteristic resistance value. 6. A method as in claim 5, further comprising: computing the mean of a plurality of aortic characteristic resistance values, which will include at least one aortic characteristic resistance value estimated for a previous cycle;setting the amplitude of the component waveform for the current pressure cycle to be inversely proportional to the square of the product of the mean and a calibration constant. 7. A method as in claim 6, further comprising setting the amplitude of the component waveform for the current pressure cycle to be inversely proportional to the square of the product of the mean, the calibration constant, and the arterial compliance value. 8. A method as in claim 3, in which the assumed input flow waveform comprises a train of square-wave signals, each forming a respective one of the co... | Edwards Lifesciences Corporation | A61B005/02
US07654965 | B2 | 2010-02-02 | Method and system for processing electrocardiograms | 2006-06-09 | In a method and system for processing an electrocardiogram (ECG), digital ECG data are received. The digital ECG data evinces a plurality of a patient's heartbeats detected during an ECG. Digital annotation data are generated representing marking to be shown on an ECG tracing of the plurality of heartbeats. An annotated ECG tracing image is produced using the digital ECG data and the digital annotation data. | 1. An ECG processing management system for processing electrocardiograms (ECGs), comprising: means for receiving protocol data for a first clinical trial;means for receiving and storing ECG data for a plurality of ECGs associated with the first clinical trial;means for processing said ECGs to receive interval measurements for said ECGs;means for processing said ECGs after measurement to receive a cardiological evaluation for each of said ECGs;means for receiving protocol data for a second clinical trial for use with said ECGs associated with the first clinical trial, and for receiving modified protocol data for said first clinical trial; andmeans for identifying ECGs from said plurality of ECGs associated with the first clinical trial that require re-measurement and reevaluation based on said protocol data for the second clinical trial and modified protocol data for said first clinical trial,wherein ECGs identified by the identifying means to require re-measurement and reevaluation are automatically reprocessed by said means for processing said ECGs to receive interval measurements and said means for processing said ECGs after measurement to receive a cardiological evaluation as needed to comply with the protocol data for the second clinical trial or modified protocol data for the first clinical trial. 2. The ECG processing management system of claim 1, further comprising: means for receiving ECO data for a second plurality of ECGs for evaluation after receipt of said protocol data for the second clinical trial; andmeans processing said second plurality of ECGs in parallel to obtain interval measurements and cardiac evaluations in accordance with said protocol data for said first clinical trial and said protocol data for said second clinical trial. 3. The ECG processing management system of claim 1, further comprising: means for receiving modified demographic data associated with said ECG data for at least one of the plurality of ECGs;means for comparing said modified demographic data against protocol data to determine if said ECG requires re-measurement and reevaluation. 4. The ECG processing management system of claim 1, further comprising: means for automatically generating a report relating to evaluations of said ECGs. 5. The ECG processing management system of claim 1, wherein said means for receiving protocol data for said first clinical trial identifies a sponsor of said first clinical trial, retrieves a protocol template associated with said sponsor and customizes said protocol template for said first clinical trial. 6. The ECG processing management system of claim 1, further comprising means for determining if additional demographic or patient data is required for said ECGs based on said protocol data for the second clinical trial and modified protocol data for said first clinical trial, wherein ECGs for which it is determined additional demographic or patient data is required are made available to means for processing said ECGs to ... | eResearchTechnology, Inc. | A61B005/0452
US07641616 | B2 | 2010-01-05 | Blood pressure monitoring system and method of having an extended optical range | 2008-03-25 | A system and method for monitoring blood pressure of a wearer has an inflatable arm cuff that is selectably inflatable to differing air pressures that incorporates a fabric having both a light transmission property and a light reflection property when the fabric is illuminated with light having wavelength(s) in the range from about 400 to about 2200 nanometers. A radiation source and a detector are attached to the fabric in relative positions such that the reception of incident radiation by the detector is directly affected by a change in the amount of light transmitted through the fabric relative to the amount of light reflected by the fabric as the fabric stretches in response to motion in the body of a wearer due to changes in the flow of blood through an artery disposed beneath the fabric occurring in consonance with variations in the air pressure within the inflatable cuff. | 1. A method for monitoring blood pressure comprising the steps of: inflating an inflatable cuff to modify the flow of blood through an artery located in the portion of the body of a wearer over which the cuff is disposed;disposing in or over the inflatable cuff a fabric having an external surface and in which the amount of light transmitted through gaps between yarns forming the fabric relative to the light reflected by yarns forming the fabric changes when the fabric stretches;using a radiation source mounted to the fabric to direct radiation onto the external surface of the fabric, wherein said radiation source directs radiation onto the fabric in the range from about 400 to about 2200 nanometers;using a radiation detector responsive to incident radiation in the range from about 400 to about 2200 nanometers that is mounted to the fabric to detect the amount of light reflected by the yarns forming the fabric or transmitted through gaps between the yarns forming the fabric; andrecording the pressure in the cuff in response to the signal output from the detector. 2. The method of claim 1, wherein recording the pressure comprises measuring the pressure in the cuff when the output of the detector is at a minimum value and measuring the pressure in the cuff when the signal from the detector lies within a predetermined range of a baseline signal value. 3. The method of claim 1, wherein the source and detector are attached to the fabric in relative positions such that the reception of incident radiation by the detector is directly affected by a change in the amount of light transmitted through the gaps between the yarns of the fabric relative to the amount of light reflected by the yarns of the fabric as the fabric stretches in response to motion in the body of a wearer due to changes in the flow of blood through an artery disposed beneath the fabric occurring in consonance with variations in the air pressure within the inflatable cuff. 4. The method of claim 1, wherein the fabric is disposed over the cuff. 5. The method of claim 1, wherein the fabric forms at least a portion of a patch. 6. The method of claim 1, wherein the radiation source is separable from the fabric. 7. The method of claim 1, wherein the radiation detector is separable from the fabric. 8. The method of claim 1, the fabric has a first and a second side; and wherein the radiation source and the radiation detector are mounted on the same side of the fabric. 9. The method of claim 1, wherein the fabric comprises reflective yarns and stretchable yarns. 10. A method for monitoring blood pressure of a wearer comprising the steps of: inflating an inflatable cuff disposed over a portion of the body of a wearer to differing air pressures thereby to modify the flow of blood through an artery located in the portion of the body of a wearer over which the cuff is disposed;disposing on, over or in the inflatable cuff a patch at least a portion of which is formed from a fabric having an e... | Textronics Inc. | A61B005/02
US08105247 | B2 | 2012-01-31 | Device for monitoring size of luminal cavity | 2008-07-25 | A system for measuring the size of a gastric lumen includes a gastric banding device, an antenna positioned outside of the patient's body, and a receiver in operative communication with the antenna. The banding device comprises two orthogonal ultrasonic transceiver modules, a microprocessor, and a transmitter in electrical communication with the microprocessor. The microprocessor is in operative communication with each of the two transceiver modules. The transmitter transmits signals to a receiver positioned outside of a patient's body. Another system for measuring the size of a gastric lumen includes a gastric banding device having a first coiled conductor. The system further comprises a circuit external to the banding device, the circuit comprising a second coiled conductor. The circuit further comprises a tunable frequency generator and a spectrum analyzer, the circuit being tunable to allow a resonant frequency. | 1. A system for measuring a size of a gastric lumen of a patient, the system comprising: a circuit external to a gastric banding device and to the patient, the circuit comprising a first coiled conductor, a tunable frequency generator, and a spectrum analyzer,wherein the circuit is tunable to allow a resonant frequency,wherein the frequency generator generates a current in the first coiled conductor,wherein the size of the gastric lumen is determined by a change in resonant frequencies measured by the spectrum analyzer in the first coiled conductor when in the presence of a second coiled conductor in the gastric banding device, andand wherein the size of the gastric lumen is determined without modifying the gastric banding device. 2. The system of claim 1, further comprising the gastric banding device, wherein the banding device has an inner side and an outer side, the inner side being closer to the gastric lumen than the outer side, the second coiled conductor being positioned on the inner side. 3. The system of claim 1, further comprising a heating control element, the heating control element being in electrical communication with the second coiled conductor. 4. The system of claim 1, further comprising at least one coil holder for orienting the first coiled conductor with the first coiled conductor. 5. The system of claim 1, wherein the size of the gastric lumen is determined by a change in resonant frequencies measured by the spectrum analyzer in the first coiled conductor when in the presence of a second coiled conductor in the gastric banding device and after the patient has swallowed a liquid. 6. The system of claim 5, wherein the liquid is a paramagnetic solution. 7. A system for measuring a size of a gastric lumen, the system comprising: a circuit external to a gastric banding device, the circuit comprising a first coiled conductor, a tunable frequency generator, and a spectrum analyzer,wherein the circuit is tunable to allow a resonant frequency,wherein the frequency generator generates a current in the first coiled conductor, andwherein the spectrum analyzer measures resonant frequency values in the first coiled conductor; anda computer comprising a computer-readable storage medium comprising instructions that, when executed, cause the computer to:calculate the size of the gastric lumen based on the difference between a first resonant frequency value of the first coiled conductor and a second resonant frequency value of the first coiled conductor when in the presence of a second coiled conductor in the gastric banding device. 8. The system of claim 7, further comprising the gastric banding device. 9. The system of claim 7, wherein the frequency generator is a variable frequency oscillator. 10. The system of claim 7, wherein the instructions that cause the computer to calculate the size of the gastric lumen based on the difference between a first resonant frequency value of the first coiled conductor and a second resonant frequ... | Buchwald O'Dea LLC | A61B005/103 A61B005/117
US08556811 | B2 | 2013-10-15 | Method and system for determining a cerebrovascular autoregulation state of a patient | 2008-01-04 | A method of diagnosing cerebrovascular autoregulation in a patient includes measuring blood pressure of the patient, measuring, non-invasively, venous oxygen content of the patient's brain substantially simultaneously with the measuring blood pressure, correlating the blood pressure and the venous oxygen content measurements in a time domain, and determining a cerebrovascular autoregulation state of the patient based on the correlating the blood pressure and the venous oxygen content measurements. | 1. A method of diagnosing cerebrovascular autoregulation in a patient, comprising: measuring blood pressure of said patient;measuring, non-invasively, venous oxygen content of said patient's brain substantially simultaneously with said measuring blood pressure;correlating said blood pressure and said venous oxygen content measurements in a time domain; anddetermining, using a data processing unit, a cerebrovascular autoregulation state of said patient based on said correlating said blood pressure and said venous oxygen content measurements. 2. A method of diagnosing cerebrovascular autoregulation in a patient according to claim 1, wherein said measuring blood pressure comprises low-pass filtering said blood pressure measurements to allow measurement of only time-dependent blood pressure variations having a frequency lower than a threshold frequency. 3. A method of diagnosing cerebrovascular autoregulation in a patient according to claim 1, wherein said measuring blood pressure comprises sampling values of said patient's arterial blood pressure at a substantially periodic sampling rate, wherein said substantially periodic sampling rate is selected to have a value to provide low-pass filtering of said arterial blood pressure to allow measurement of only time-dependent blood pressure variations having a frequency lower than a threshold frequency. 4. A method of diagnosing cerebrovascular autoregulation in a patient according to claim 3, wherein said measuring, non-invasively, venous oxygen content comprises sampling values of said patient's venous oxygen content synchronously with said sampling values of said patient's blood pressure. 5. A method of diagnosing cerebrovascular autoregulation in a patient according to claim 4, wherein said correlating said arterial blood pressure and said venous oxygen content measurements comprises calculating a linear correlation coefficient having a value that is indicative of said autoregulation state of said patient. 6. A method of diagnosing cerebrovascular autoregulation in a patient according to claim 5, further comprising displaying said linear correlation coefficient as a function of arterial blood pressure to provide a pattern that is indicative of said autoregulation state of said patient. 7. A method of diagnosing cerebrovascular autoregulation in a patient according to claim 5, further comprising measuring an intracranial pressure of said patient and displaying said linear correlation coefficient as a function of a cerebral perfusion pressure of said patient to provide a pattern that is indicative of said autoregulation state of said patient, said cerebral perfusion pressure being calculated as a difference between said arterial blood pressure and said intracranial pressure. 8. A method of diagnosing cerebrovascular autoregulation in a patient according to claim 1, wherein said measuring, non-invasively, venous oxygen content of said patient's brain comprises measuring differential absorption of n... | Johns Hopkins University | A61B005/00
US08273032 | B2 | 2012-09-25 | Physiological parameter monitoring with minimization of motion artifacts | 2008-07-30 | An implanted system includes at least two optical sensors implanted proximate to an artery of a patient such that one optical sensor is upstream of another optical sensor. Arterial pulses of the patient may be detected based on electrical signals from at least one of the optical sensors. In addition, electrical signals from the optical sensors may be used to minimize the effects of motion artifacts on the detection of arterial pulses. For example, a detected pulse may be determined to be a spurious pulse if the optical sensors indicate the occurrence of the pulse within a predetermined range of time. As another example, a first optical sensor signal may be shifted in time relative to a second optical sensor signal, and the signals may be correlated. An arterial pulse may be detected at a time at which a peak or trough amplitude value of the correlated signal is observed. | 1. A method comprising: receiving, with a processor, a first electrical signal from a first optical sensor implanted within a patient;receiving, with the processor, a second electrical signal from a second optical sensor implanted within the patient;detecting, with the processor, an arterial pulse of the patient based on a first signal characteristic of the first electrical signal;detecting, with the processor, a second signal characteristic of the second electrical signal indicative of the arterial pulse; anddetermining, with the processor, whether the detected arterial pulse is a spurious pulse based on a time delay between the first and second signal characteristics. 2. The method of claim 1, wherein the first and second optical sensors are implanted along an artery of the patient. 3. The method of claim 2, wherein the second optical sensor is located downstream of the first optical sensor relative to a direction of blood flow in the artery. 4. The method of claim 1, wherein the first and second optical sensors are implanted along different arteries of the patient. 5. The method of claim 1, wherein the first and second optical sensors are coupled to a common housing. 6. The method of claim 5, wherein the housing defines a surface and the first and second optical sensors are coupled to the surface of the common housing. 7. The method of claim 1, wherein the first and second optical sensors comprise separate housings that are movable relative to each other. 8. The method of claim 1, wherein determining whether the detected arterial pulse is a spurious pulse comprises determining whether the first and second signal characteristics occurred within a predetermined range of time of one another, and if the first and second characteristics occurred within the predetermined range of time, determining that the detected arterial pulse is the spurious pulse. 9. The method of claim 8, wherein the predetermined range of time comprises less than about 0.5 milliseconds. 10. The method of claim 1, wherein the first and second signal characteristics comprise at least one of a peak amplitude, a trough amplitude or a threshold crossing. 11. The method of claim 1, wherein the first and second signal characteristics comprise threshold crossings of the first and second signals, respectively, and wherein detecting the arterial pulse comprises: normalizing the first electrical signal; anddetecting a first threshold crossing of the normalized first electrical signal, wherein the first threshold crossing indicates the arterial pulse; andwherein detecting the second signal characteristic of the second electrical signal indicative of the arterial pulse comprises: normalizing the second electrical signal;detecting a second threshold crossing of the normalized second electrical signal, wherein the second threshold crossing indicates the arterial pulse. 12. The method of claim 11, wherein determining whether the arterial pulse is a spurious pulse comprises determi... | Medtronic, Inc. | A61B005/02
US08090177 | B2 | 2012-01-03 | Methods for detection and characterization of atypical vessels in cervical imagery | 2008-08-01 | The present invention discloses a method for the detection of atypical vessels in digital cervical imagery. A pre-processing stage is applied to enhance the contrast of blood vessel features compared to the surrounding tissue. Next, a segmentation stage is applied to identify regions of interest for atypical vessels using texture and gradient information. Finally, a post-processing stage (1) identifies other clinically relevant features in the cervical imager, and removes these features from the region of interest; and (2) uses color, size, and shape information to further refine the region of interest to eliminate false positives and determine a final region of interest. This automated method of atypical vessel detection is especially useful for diagnostic purposes such as cervical cancer detection. | 1. A method comprising: obtaining color images of the cervix wherein said images contain color components, atypical vessels, and features unrelated to atypical vessels;pre-processing to enhance contrast of said images to increase visualization of blood vessels from surrounding tissue creating a red component enhanced image and a contrast enhanced red component image, wherein said pre-processing comprises the steps ofsuppressing high intensity components by filtering an image from said images by a to hat filter, and subtracting the to hat-filtered image from said image to create a smoothed image;creating a contrast-enhanced image by filtering the smoothed image with a to hat filter,filtering the smoothed image with a bottom hat filter,adding the top hat-filtered smoothed image to, and subtracting the bottom hat-filtered smoothed image from said smoothed image;augmenting the red component of said contrast enhanced image to create said red component enhanced image;creating a contrast enhanced red component image by filtering the red component enhanced image with a to hat filter,filtering the red component enhanced image with a bottom hat filter,adding the top hat-filtered red component enhanced image to, and subtracting the bottom hat-filtered red component enhanced image from said red component enhanced image;segmenting said red component enhanced image and said contrast enhanced red component image to identify regions of interest, wherein said segmenting comprises the steps ofapplying a bottom-hat filter to said red component enhanced image to extract regions containing maximum texture;performing region-based segmentation on said regions containing maximum texture to create a maximum texture image;applying an edge detection algorithm to said contrast enhanced red component image to detect edges, andfilling areas surrounded by said edges in said contrast enhanced red component image to create a maximum gradient image; andcalculating the intersection of said maximum texture image and said maximum gradient image to locate said regions of interest;post-processing to eliminate false-positives of atypical vessels and determine a final region of interest from said regions of interest, wherein said post-processing comprises the steps of identifying and removing said features unrelated to atypical vessels in said regions of interest; andlocating a final region of interest using known color, shape, and size information for said atypical vessels; and whereby said method detects said atypical blood vessels in digital images of tissue or organs for cancer diagnosis. | STI Medical Systems, LLC | G06K009/00
US08268919 | B2 | 2012-09-18 | Polymer for detecting fingerprint, method of producing the same, composition for detecting fingerprint and method of detecting fingerprint using the same | 2007-09-28 | It is intended to provide a polymer for detecting a fingerprint, which can be easily handled in a scene of the fingerprint detection as a substitute for conventionally employed 2-cyanoacrylate (a monomer), and a composition for detecting a fingerprint. It is also intended to provide a method of detecting a fingerprint whereby a fingerprint can be detected at a high sensitivity regardless of the color or conditions of a specimen and wherein the specimen can be well restored. The present polymer and composition for detecting a fingerprint and method of detecting a fingerprint using the same make it possible to definitely detect a fingerprint by using fluorescence even from an almost white specimen such as a shopping bag or an aluminum foil from which a fingerprint can be hardly detected by the existing methods. | 1. A composition for detecting a fingerprint, which comprises a 2-cyanoacrylate polymer polymerized by use of a polymerization initiator other than one represented by formula (1), and a fluorescent colorant represented by the formula (1):wherein R1 and R2 each represents an alkyl group with 1-3 carbon atoms, and R3 represents an aldehyde group or an alkyl or alkenyl group that has 1-6 carbon atoms and has an aldehyde group at a terminal thereof, whereby when detecting said fingerprint, the polymer is depolymerized. 2. A composition for detecting a fingerprint according to claim 1, wherein the polymerization initiator other than one represented by the formula (1) is water, methanol or an amine. | Toagosei Co., Ltd. | C08G065/26 C08F002/00 A61B005/103 A61B005/117
US08262714 | B2 | 2012-09-11 | Techniques for selecting signal delivery sites and other parameters for treating depression and other neurological disorders, and associated systems and methods | 2009-02-27 | The present disclosure is directed generally to techniques for selecting signal delivery sites and other signal delivery parameters for treating depression and other neurological disorders, and associated systems and methods. A method in accordance with a particular embodiment includes obtaining first imaging information corresponding to a first region of a patient's brain, the first imaging information being based at least in part on functional characteristics of the first region. The method further includes obtaining second imaging information corresponding to a second region of the patient's brain, the second region being a subset of the first region, the second imaging information being based at least in part on functional or structural characteristics of the second region. A target neural population is then selected based at least in part on the second imaging information. The method still further includes applying an electromagnetic signal to the target neural population to improve a patient function. | 1. A method for treating a patient, comprising: obtaining first imaging information corresponding to a first region of a patient's brain, the first imaging information being based at least in part on functional characteristics of the first region;obtaining second imaging information corresponding to a second region of the patient's brain, the second region being a subset of the first region, the second imaging information being based at least in part on functional or structural characteristics of the second region;selecting a target neural population based at least in part on the second imaging information; andapplying an electromagnetic signal to the target neural population to improve a patient function;wherein obtaining first imaging information includes obtaining first imaging information of the patient's middle frontal gyrus. 2. The method of claim 1 wherein obtaining second imaging information includes obtaining second imaging information having a higher resolution than the first imaging information. 3. The method of claim 1 wherein obtaining second imaging information includes obtaining at least one of fractional anisotropy information and diffusion tensor imaging information. 4. The method of claim 1 wherein obtaining first imaging information includes obtaining first imaging information using fMRI techniques. 5. The method of claim 4 wherein obtaining first imaging information includes obtaining first information on a real-time or approximately real-time basis, and wherein obtaining the second information includes using spectroscopy techniques while the patient is present. 6. The method of claim 1 wherein obtaining second imaging information includes obtaining information including a thickness value of gray matter in the second region. 7. The method of claim 1, further comprising identifying the patient as having depression, and wherein: obtaining first imaging information includes obtaining functional magnetic resonance imaging information while engaging the patient in a task selected to trigger an emotional response in the patient, the first imaging information having a first resolution; and whereinobtaining second imaging information includes obtaining diffusion tensor imaging information having a second resolution greater than the first resolution, and wherein the method further comprises obtaining a fractional anisotropy level for the patient, and, based at least in part on the fractional anisotropy level, determining whether the patient is a suitable candidate for receiving an electromagnetic signal; and whereinapplying an electromagnetic signal includes applying an electrical signal from an electrode positioned within the patient's skull cavity and outside a cortical surface of the patient's brain. 8. A method for treating a patient, comprising: identifying a patient as having depression;obtaining first imaging information corresponding to a first region of a patient's brain, the first imaging information being obtained by ... | Advanced Neuromodulation Systems, Inc. | A61N005/055 A61N001/36
US20100040765 | A1 | 2010-02-18 | Lifting and Preserving Bloody Impressions for Law Enforcement | 2009-10-19 | A lifting strip containing titanium dioxide has been successful in the lifting and enhancement of bloody impressions from nonporous, semi porous and porous surfaces. The lifting strip is best when activated with an alcohol such as ethanol or methanol. When the lifting strip has dried after application it can easily be removed from the surface lifting the bloody impression onto a contrasting white background for examination. This method is safe, easy, and effective on several surfaces of contrasting colors. | 1. A method of lifting bloody impressions from a surface including the steps of: providing a backing sheet having titanium dioxide disposed thereon;applying a liquid to the sheet to wet the titanium dioxide;applying the wetted sheet to a bloody impression; andlifting the backing sheet to remove at least an identifiable portion of the impression from the surface. 2. The method of claim 1 wherein the liquid includes water-miscible alcohol. 3. The method of claim 2 wherein the alcohol is applied to the side of the sheet that is to be applied to the impression. 4. The method of claim 2 further including the step of eliminating any pooling of the alcohol before applying the wetted sheet to the impression. 5. A strip material for pressing against a bloody impression on a surface and precisely lifting the bloody impression from the surface and preserving it on the material, the strip material comprising: a porous polymer membrane strip having a coating layer on at least one side of the membrane,the membrane being permeable to water and water miscible alcohols,the coating comprising particles of titanium dioxide in a size and amount for wetting with water or a water miscible alcohol and then being pressed against the bloody impression to transfer and receive the impression onto the coating and, after drying of the wetting liquid, lifting from the surface. 6. The strip material of claim 5 where in the coating layer comprises a polymeric binder for fixing the titanium dioxide particles to the membrane strip. 7. The strip material of claim 5 wherein the coating layer comprises an organic acid for fixing blood from the impression to the coating layer. 8. The strip material of claim 7 in which the acid comprises 5-sulfosalicylic acid or its dihydrate. 9. The assembly of claim 5 wherein the impression contact layer includes a surfactant. 10. The assembly of claim 9 wherein the surfactant includes propylene glycol and p-tert-octylphenoxy polyethoxyethyl alcohol. 11. The assembly of claim 5 wherein the impression contact layer includes a polymer. 12. The assembly of claim 11 wherein the polymer includes polyvinyl alcohol. 13. The assembly of claim 5 wherein the membrane strip has a predetermined size large enough to lift a variety of bloody impressions. 14. The assembly of claim 5 wherein the polymeric support layer includes a polymer transfer membrane having pores. 15. The assembly of claim 14 wherein the polymer transfer membrane includes nylon. 16. A lifting strip assembly for lifting and preserving bloody impressions found on a variety of surfaces including: a lifting strip including a support layer and an impression bonding layer carried on the support layer, the bonding layer including a material adapted to bond with the material in a bloody impression, and to lift and preserve at least a portion of the impression intact; and a transparent protective layer hingedly attached to the lifting strip whereby the protective layer can be disposed ov... | - | A61B005/117 B05D003/02
US07648468 | B2 | 2010-01-19 | Method and apparatus for penetrating tissue | 2002-12-31 | A body fluid sampling system for use on a tissue site includes an electrically powered drive force generator. A penetrating member is operatively coupled to the force generator. The force generator moves the member along a path out of a housing having a penetrating member exit, into the tissue site, stops in the tissue site, and withdraws out of the tissue site. A cartridge houses the penetrating member. The cartridge has first and second seals coupled to the penetrating member to maintain a sterile environment around a portion of the penetration member prior to penetrating member actuation. | 1. A body fluid sampling system for use on a tissue site, the system comprising: an electrically powered drive force generator;a penetrating member operatively coupled to said electrically powered drive force generator and to a processor, said electrically powered drive force generator moving said penetrating member along a path out of a housing having a penetrating member exit, into said tissue site, stopping in said tissue site, and withdrawing out of said tissue site, the processor causing an electronic braking force to be applied to the penetrating member to reduce a speed of the penetrating member to achieve a desired amount of final skin penetration depth;a disposable housing said penetrating member, said disposable having a first and a second seal coupled to said penetrating member to maintain a sterile environment around a portion of the penetration member prior to penetrating member actuation; an analyte detecting member coupled to a sample chamber, the analyte detecting member being configured to determine a concentration of an analyte in a body fluid using a sample that does not exceed a volume of 1 μL of a body fluid disposed in the sample chamber;a penetrating member sensor coupled to said electrically powered force generator and provide depth penetration control of a penetrating member from a known start position, the penetrating member sensor being coupled to a processor with a memory that stores profiles which execute control instructions for the penetrating member driver, for application of a launching force to the penetrating member and control a withdraw force to the penetrating member so that the penetrating member moves in a second direction away from the target tissue, the processor coupled to an amplifier to modulate power from a power supply to the electrically powered drive force generator and through a transducer measure location of the penetrating member with the penetrating member sensor; andwherein in the first direction the penetrating member moves toward the target tissue at a speed that is different than a speed at which the penetrating member moves away from the target tissue. 2. The system of claim 1, wherein the penetrating member sensor is further configured to provide an indication of velocity of a penetrating member. 3. The system of claim 1, wherein the driver is a voice coil drive force generator. 4. The system of claim 1, wherein the driver is a rotary voice coil drive force generator. 5. The system of claim 1, wherein the processor includes a memory for storage and retrieval of a set of penetrating member profiles utilized with the penetrating member driver. 6. The system of claim 5, wherein the processor is utilized to monitor position and speed of a penetrating member as the penetrating member moves in a first direction. 7. The system of claim 5, wherein the processor is utilized to adjust an application of force to a penetrating member to achieve a desired speed of the penetrating member. 8. The... | Pelikon Technologies, Inc. | A61B005/00
US20100049029 | A1 | 2010-02-25 | PIEZOELECTRIC MAGNETIC RESONANCE ELASTOGRAPH (MRE) DRIVER SYSTEM | 2008-08-20 | An array of two or more piezoelectric drivers generates shear waves in a region of interest of a human undergoing a MRE test. The use of the array of drivers allows for better diagnosis of disease of the humans or animals. | 1. A phased array driver for a magnetic resonance elastograph system comprising: a first driver having a piezoelectric element;a second driver having a piezoelectric element;wherein the first driver and the second driver are arranged to produce shear waves in a subject. 2. The driver of claim 1, wherein each piezoelectric element comprises one of a MRI compatible material and a PVF2 material. 3. The driver of claim 1, wherein a shape of the first element is different than a shape of the second element. 4. The driver of claim 1, further comprising: a converter that receives an optical signal and converts the optical signal into an electrical signal for use by the first driver and the second driver. 5. The driver of claim 1, wherein the subject is a human. 6. The driver of claim 5, wherein the driver array is used to diagnosis a medical condition of the human subject. 7. The driver of claim 1, wherein the phase array driver operates according to a driver signal, wherein the driver signal is produced by a signal component that comprises: a signal generator that forms the signal;an oscilloscope that displays the signal and facilitates monitoring of the signal; andan amplifier that increases power of the signal. 8. The driver of claim 7, wherein the signal component comprises a single container. 9. The driver of claim 7, further comprising: a converter that converts the electrical signal to an optical signal. 10. The driver of claim 1, further comprising: a connector that secures the driver array to the subject. 11. A magnetic resonance elastrography system comprising: a magnetic resonance imaging (MRI) system that scans a subject; anda phased array of drivers that produce shear waves in a region of the subject from a signal, wherein each of the drivers in the array comprises a piezoelectric element. 12. The system of claim 11, wherein an operation of the array of drivers is controlled by the MRI system. 13. The system of claim 11, further comprising: a signal component that generates a signal that operates the phased array. 14. The system of claim 13, wherein the signal component comprises: a signal generator that forms the signal;an oscilloscope that displays the signal and facilitates monitoring of the signal; andan amplifier that increases power of the signal. 15. The system of claim 14, wherein the signal component comprises a single container. 16. The system of claim 13, further comprising: an fiber optic wire that is connected between the signal component and the phased array. 17. The system of claim 16, wherein the signal component further comprises: a converter that converts an electrical signal from the signal component into a light signal. 18. The system of claim 16, wherein the phased array further comprises: a converter that converts a light signal from the fiber optic wire into an electrical signal that is useable by the phased array. 19. The driver of claim 11, wherein the subject is a human. 20. The driver of claim 1... | Hong Kong Applied Science and Technology Research Institute Co., Ltd. | A61B005/055
US20100049014 | A1 | 2010-02-25 | Sensor Inserter Assembly | 2009-11-05 | An analyte monitor includes a sensor, a sensor control unit, and a display unit. The sensor control unit typically has a housing adapted for placement on skin and is adapted to receive a portion of an electrochemical sensor. The sensor control unit also includes two or more conductive contacts disposed on the housing and configured for coupling to two or more contact pads on the sensor. A transmitter is disposed in the housing and coupled to the plurality of conductive contacts for transmitting data obtained using the sensor. The display unit has a receiver for receiving data transmitted by the transmitter of the sensor control unit and a display coupled to the receiver for displaying an indication of a level of an analyte, such as blood glucose. An inserter having a retractable introducer is provided for subcutaneously implanting the sensor in a predictable and reliable fashion. | 1. (canceled) 2. An insertion kit, comprising: a sensor and an introducer needle, wherein the sensor and the introducer needle are configured and arranged with a shape, such that the sensor is retained in the introducer needle by a friction fit such that motion of the sensor with respect to the introducer needle is minimized prior to sensor insertion; andan applicator configured to cause the needle to be inserted through a skin of a host, wherein the applicator is further configured to position at least a portion of the sensor in contact with a biological fluid of the host. 3. The kit of claim 2 wherein the sensor is retained in a lumen of the introducer needle. 4. The kit of claim 3 wherein the sensor is retained within the lumen. 5. The kit of claim 2 wherein the sensor is retained within the introducer needle. 6. The kit of claim 2 wherein the sensor is retained in the introducer needle by the friction fit of a contact point between the sensor and the introducer needle. 7. The kit of claim 2 wherein the motion is longitudinal motion. 8. The insertion kit of claim 2, wherein the applicator is further configured to substantially completely withdraw the introducer needle from the host after the sensor is transcutaneously positioned in fluid contact with the biological fluid of the host. 9. The insertion kit of claim 2, wherein the applicator includes a spring biased mechanism. 10. The insertion kit of claim 2, wherein the shape is a bend. 11. An insertion kit, comprising: a sensor and an introducer needle, wherein the sensor and the introducer needle are configured and arranged with a shape, such that the sensor is retained within a lumen of the introducer needle by a friction fit of multiple contact points separated by non-contact points between the sensor and the lumen of the introducer needle, such that longitudinal motion of the sensor with respect to the introducer needle is minimized prior to sensor insertion; andan applicator configured to cause the needle to be inserted through a skin of a host, wherein the applicator is further configured to position at least a portion of the sensor in contact with a biological fluid of the host. 12. An analyte sensor system, comprising: an analyte sensor and an applicator comprising an introducer needle, wherein the applicator is configured to position at least a portion of the sensor in contact with a biological fluid of a host by insertion of, and subsequent withdrawal of, the introducer needle through the skin of the host, wherein the analyte sensor is configured and arranged to be held by a friction fit of the surface of the sensor with the introducer needle in order to maintain the sensor longitudinally within the introducer needle prior to withdrawal of the introducer needle from the host and minimize motion of the sensor during withdrawal of the introducer needle from the host. 13. The system of claim 12 wherein the analyte sensor is configured and arranged to be held by the friction ... | Abbott Diabetes Care Inc. | A61B005/145
US07668580 | B2 | 2010-02-23 | Electrode for living body and device for detecting living signal | 2006-07-17 | A biosignal measurement device includes an electrode and a signal processing member. The electrode includes an insulation sheet having a hole, a device contact portion provided on the top surface of the insulation sheet and a body contact portion provided on the bottom surface of the insulation sheet, the device contact portion and the body contact portion electrically connected to each other via the hole. The signal processing member includes an externally exposed terminal to make surface contact with the device contact portion, an analog signal processing unit, an A/D signal converter and a digital signal processing unit. Also, the device contact portion and the body contact portion are formed of a material which is both conductive and adhesive. Accordingly, the signal processing member may be directly attached. Noise may be reduced. Also, a biosignal may be accurately measured. | 1. A biosignal measurement device for making contact with a living body and measuring a biosignal, comprising: a signal processing member comprising an externally exposed terminal, an analog signal processing unit processing an analog signal transmitted from the terminal, an analog-to-digital (A/D) signal converter converting the analog signal into a digital signal, and a digital signal processing unit processing the converted digital signal; andan electrode member comprising an insulation sheet having a hole, a device contact portion contacted with a top surface of the insulation sheet to make surface contact with the terminal of the signal processing member and a body contact portion contacted with a bottom surface of the insulation sheet and electrically connected with the device contact portion, the device contact portion and the body contact portion formed of a conductive material,wherein the hole integrally connects the device contact portion and the body contact portion, and the device contact portion and the terminal make surface contact by flat contact surfaces formed on the device contact portion and the terminal and are electrically connected to each other via a wide contact area. 2. The device of claim 1, wherein the signal processing member further comprises a wired/wireless transmitter externally transmitting a signal processed by the digital signal processing unit. 3. The device of claim 1, wherein the device contact portion and the body contact portion are formed of a material which is both conductive and adhesive, and the device contact portion attaches to the signal processing member via adhesive properties and maintains an electrical connection state with the terminal of the signal processing member. 4. The device of claim 3, wherein the device contact portion and the body contact portion are provided by placing a material, which is fluidic conductive and adhesive, around the hole of the insulation sheet and molding the fluidic conductive and adhesive material into a solid material which is both conductive and adhesive. 5. The device of claim 1, wherein an adhesive material is applied around the body contact portion on the bottom of the insulation sheet. | Samsung Electronics Co., Ltd. | A61B005/04
US07657444 | B2 | 2010-02-02 | Distance-treatment through public network | 2001-10-16 | A method and system of providing biological distance-treatment through a public network includes at least a treatment instrument which is electrically connected with an information connection system for providing a treatment for a registered user, wherein a treatment information data package sent from a service provider via the information connection system through the public network to provide digital treatment signals to control the treatment, wherein the treatment information data package is selected from the treatment information database based on a treatment request sent from the information connection system to the service provider and the health information profile of the registered user in the service provider. | 1. A method of providing distance-treatment for registered users through Internet, comprising the steps of: (a) establishing an information connection system comprising a computer, a visual signal producer and an audio signal producer, wherein said information connection system is arranged to be capable of communicating with a service provider through said Internet;(b) requiring said registered user to login said service provider through said Internet, wherein each of said registered user is able to present a diagnosis record and a health information profile for said information connection system;(c) communicating a treatment request to said information connection system on the basis of said diagnosis record of said registered users through said Internet, wherein after said verification of said registered user, said service provider recognizes said registered user and admits said registered user to make said treatment request to said service provider at said information connection system through said Internet, wherein in responsive to said treatment request of said registered user, said service provider sends said service provider sends said respective registered user a treatment page which includes a list of said health problems and diseases that said registered user suffers, treatment opinions from doctors, recommendation of beneficial foods and activities for each of said listed health problems and diseases of said registered user, recommended biological treatments with respect to said listed health problems and diseases that said registered user suffers respectively, and information of suggested treatment instrument for executing each recommended biological treatment;(d) based on said treatment request and said health information profile for said registered user, formulating a treatment information data package from a treatment information database provided by said service provider;(e) commanding said computer to perform treatment by digital treatment signals based on said treatment information data package, wherein said treatment is initiated and operated by said information connection system on said registered user, wherein said treatment is selected from a group consisting of an audio and visual treatment to said registered user via said audio device and said monitor respectively, and is based on said treatment data package, wherein said digital treatment signals are decoded into analog treatment signals which are sent to said computer to program and control said treatment of said registered user with said treatment instrument; and(f) sending a responsive health information of said registered user to the service provider for controlling and adjusting properties of said digital treatment signals of said treatment information data package; and(g) communicatively connecting a second treatment instrument with said information connection system, wherein said second treatment instrument further includes at least one addition treatment option as ... | - | A61B005/00
US07658740 | B2 | 2010-02-09 | Surgical attachment instrument and method | 2004-08-17 | A surgical attachment instrument and method according to which a housing is adapted to be coupled to a motor. A tubular member is coupled to the housing so that an end of the tubular member extends from the housing at a predetermined distance. The predetermined distance may be selectively adjusted by a repeatable increment. | 1. A surgical attachment instrument comprising: a housing adapted to be coupled to a motor;a tubular member coupled to the housing so that an end of the tubular member extends from the housing at a predetermined distance; andmeans for selectively adjusting the predetermined distance by a repeatable increment;wherein the adjusting means comprises a coupling connecting the tubular member and the housing, a through-opening formed in the coupling, and a ball extending through the through-opening and engaging the tubular member;wherein at least two positioning detents are formed in the tubular member and longitudinally spaced at a distance corresponding to the repeatable increment for engagement by the ball;wherein the adjusting means further comprises a keeper spring disposed in the housing and engaged with the coupling for applying an inwardly-directed radial force to the ball;wherein the adjusting means further comprises a positioning channel formed in the tubular member and extending between the two positioning detents so that the tubular member moves relative to the housing in a longitudinal direction upon application of a longitudinally-directed force to the tubular member;wherein application of the longitudinally-directed force overcomes the inwardly-directed radial force so that the ball disengages from one of the two positioning detents and engages the positioning channel during the longitudinal movement of the tubular member;wherein continued application of the longitudinally-directed force causes the ball to disengage from the positioning channel and engage the other of the two positioning detents. 2. The surgical attachment instrument of claim 1 wherein further longitudinal movement of the tubular member is prevented when the ball engages the other of the two positioning detents. 3. The surgical attachment instrument of claim 1 wherein the engagement of the ball with the other of the two positioning detents provides acoustic and tactile feedback indicating that the ball has engaged the other of the two positioning detents. 4. The surgical attachment instrument of claim 1 further comprising at least two locking detents formed in the tubular member and longitudinally spaced at a distance corresponding to the repeatable increment, and at least two locking channels formed in the tubular member, each locking channel transversely-extending between each positioning detent and each corresponding locking detent. 5. The surgical attachment instrument of claim 4 wherein the tubular member rotates so that the ball disengages from one of the two positioning detents and engages the corresponding locking channel. 6. The surgical attachment instrument of the claim 5 wherein the disengagement of the ball from the one of the two positioning detents and the corresponding engagement of the ball with the corresponding locking channel provides acoustic and tactile feedback indicating that the ball has disengaged from the one of the two positioning detents.... | Medtronic, Inc. | A61B005/00
US07667603 | B2 | 2010-02-23 | Embedding items with RFID tags for tracking and calibration | 2006-04-13 | RFID tags of very small size are embedded in products or composed of products in a manufacturing process. The system employs different read and write modes to enable auto-tracking of material, some assembly, assembly and component items through various stages of the manufacturing process. As each item passes special predetermined points in the manufacturing process, the embedded tag is activated and placed in track mode. The tag transmits its ID and a track count representing the number of stations passed. The tag's track count is incremented and the updated track count is stored in non-volatile memory in the tag. The tags can be programmed so that once the count exceeds a predetermined count, a status bit is set in the tag's memory indicating that the item has been completely through the manufacturing process. Thus, the system can determine whether an item or product has been completed. After manufacture the same RFID tag can be used for tracking, inventory and item authentication. | 1. A method for tracking progress of products or components of products as to operations performed in a manufacturing process, comprising: attaching an RFID tag to each of a plurality of the products or components, the RFID tag being of the passive type without onboard battery and which derives power to operate from an interrogating RF signal from a reader,at each of a plurality of stations in the manufacturing process, sending an interrogating RF signal to a product or component bearing an RFID tag,the interrogating RF signal at a particular station containing a triggering signal that, when the RF signal powers up the tag, causes the tag to respond with the tag's unique ID and with a track count representing a number of stations which the tag has passed through up to that point prior to the particular station,and the RFID tag incrementing its track count to reflect presence at the particular station, and storing the incremented track count in a non-volatile memory before the tag discharges. 2. The method of claim 1, wherein the incrementing of the track count is triggered by the receipt of a second discrete triggering powering-up signal by the RFID tag. 3. The method of claim 1, wherein the triggering signal comprises the interrogating RF signal being of much higher power than used in reading of the RFID tag for other, or normal, purposes such as inventory track, item calibration or item authentication. 4. The method of claim 1, wherein the RFID tag on a product or component carries a calibration for the particular product or component, in a case where calibration will vary from component to component, and wherein the triggering signal causes the tag to respond with the component's calibration as well as with the tag's unique ID and a track count. 5. The method of claim 4, wherein the products or components are glucose test strips, and wherein the calibration data relate to the calibration of the glucose test strip for the measurement of blood sugar. 6. The method of claim 4, wherein the calibration carried by the tag is used to represent the performance of an integrated circuit. 7. The method of claim 4, wherein the products or components are glucose test strips or dry reagent test strips, wherein the calibration data relate to the calibration of the glucose test strip for the measurement of blood sugar, and further including, following the manufacturing process, using the calibration data later at a post-manufacturing facility, by reading the calibration data and applying the data for calibration of a blood sugar test on a living individual. 8. The method of claim 4, further including use of the calibration of the product or component in a post-manufacturing procedure wherein the product is used with its calibration known, in order to produce an accurate result for the product or component. 9. The method of claim 1, further including, prior to completion of the manufacturing process, leaving the RFID tag attached to the product or comp... | Tagent Corporation | G08B013/14
US07664544 | B2 | 2010-02-16 | System and methods for performing percutaneous pedicle integrity assessments | 2004-04-30 | The present invention involves systems and related methods for performing percutaneous pedicle integrity assessments involving the use of neurophysiology. | 1. A method for performing percutaneous pedicle integrity assessments, comprising the steps of: (a) percutaneously introducing an insulation member to a pedicle target site;(b) establishing electrical communication between a K-wire and an interior of a pedicle screw pilot hole by bringing said K-wire into contact with a tap member disposed through said insulation member, said tap member having a longitudinal lumen for receiving and passing said K-wire therethrough;(c) applying a stimulation signal to said K-wire; and(d) monitoring to assess whether nerves adjacent said pedicle are innervating as a result of the step of applying said stimulation signal to said K-wire. 2. The method of claim 1 and further, wherein the step (d) of monitoring to assess whether nerves adjacent said pedicle are innervating is performed by an EMG monitoring system that detects a voltage response from muscles innervated by said nerves. 3. The method of claim 2 and further, wherein said EMG monitoring system is in electrical communication with said stimulation element and directs said application of said stimulation signal. 4. The method of claim 3 and further, comprising the additional step of determining a relationship between said stimulation signal and said voltage response to indicate said pedicle integrity. 5. The method of claim 4 and further, wherein said relationship is the threshold stimulation current level required to evoke a predetermined voltage response value. 6. The method of claim 4 and further, wherein said EMG system is configured to communicate said relationship to a user. 7. A method for performing percutaneous pedicle integrity assessments, comprising the steps of: (a) percutaneously introducing to a pedicle target site an insulation member comprising a tubular element dimensioned to receive a tap member and a K-wire, the tap member having a longitudinal lumen for receiving and passing said K-wire therethrough to establish electrical communication therebetween;(b) establishing electrical communication between said K-wire and an interior of a pedicle screw pilot hole;(c) applying a stimulation signal to said K-wire; and(d) monitoring to assess whether nerves adjacent said pedicle are innervating as a result of the step of applying said stimulation signal to said stimulation element. 8. The method of claim 7 and further, wherein the step (d) of monitoring to assess whether nerves adjacent said pedicle are innervating is performed by an EMG monitoring system that detects a voltage response from muscles innervated by said nerves. 9. The method of claim 8 and further, wherein said EMG monitoring system is in electrical communication with said stimulation element and directs said application of said stimulation signal. 10. The method of claim 9 and further, comprising the additional step of determining a relationship between said stimulation signal and said voltage response to indicate said pedicle integrity. 11. The method of claim 10 and furthe... | NuVasive, Inc. | A61B005/04 A61B005/05
US07657293 | B2 | 2010-02-02 | Method for monitoring viability of tissue flaps | 2005-09-08 | Methods and apparatus for assessing the viability of tissue such as flap tissue are disclosed. According to one aspect of the present invention, a method for assessing the viability of flap tissue includes obtaining an oxygen saturation level associated with a first location on the flap tissue, determining whether the oxygen saturation level is less than a first level, and identifying the first location as having a poor blood supply if the oxygen saturation level is less than the first level. | 1. A method for assessing viability of flap tissue, the method comprising: placing a sensor head of an optical imaging system in contact with the flap tissue at a first location, wherein the sensor head causes light to be reflected off of the flap tissue to identify an oxygen saturation level;obtaining the oxygen saturation level associated with the first location on the flap tissue;determining whether the oxygen saturation level is less than a first level; andidentifying the first location as having a poor blood supply if it is determined that the oxygen saturation level is less than the first level,wherein the optical imaging system comprises:a probe comprising: a connection interface, the connection interface being adapted to allow the probe to be coupled to a plurality of light sources and at least one photodetector, wherein the light sources and the photodetector are external to the probe; andthe sensor head comprising a first source fiber, a second source fiber, and a first detector fiber, the first source fiber and the second source fiber being arranged to be coupled to the plurality of light sources via the connection interface, and the first detector fiber being arranged to be coupled to the photodetector via the connection interface; anda beam combiner, external to the probe, coupled to the first source fiber and the plurality of light sources via the connection interface. 2. The method of claim 1 wherein the first level is approximately 50 percent oxygen saturation level, wherein the percent oxygen saturation level is a percentage of hemoglobin that is bound to oxygen. 3. The method of claim 1 further including: trimming the flap tissue at the first location if the first location is identified as having the poor blood supply. 4. The method of claim 1, wherein the sensor head includes a first row of two source fibers and a second row of two detector fibers and wherein the first row of two source fibers is not in-line with the second row of two detector fibers. 5. The method of claim 4 wherein the first row is parallel to the second row. 6. The method of claim 5 wherein a first distance between a first source fiber of the first row and a second source fiber of the first row is the same as a second distance between a first detector fiber of the second row and a second detector fiber of the second row. 7. The method of claim 5 wherein a first distance between a first source fiber of the first row and a second source fiber of the first row is not the same as a second distance between a first detector fiber of the second row and a second detector fiber of the second row. 8. The method of claim 1, wherein the sensor head includes a first row having a first source fiber and a second source fiber and a second row having a first detector fiber and a second detector fiber and the first source fiber is arranged nonsymmetrically with respect to the second source fiber, first detector fiber, and second detector fiber. 9. A method for charact... | ViOptix Inc. | A61B005/00
US07654901 | B2 | 2010-02-02 | Video game system using bio-feedback devices | 2007-05-08 | A video game system using one or more bio-feedback devices for the monitor and transmission of physiological data of a participant of the video game system to the video game system where the data, along with conventional controller data is incorporated into game play provided by software through the video game system. A headset electrically coupled to a video game system, the headset including one or more bio-feedback devices for transmitting participant physiological data to the video game system for real-time incorporation into game play. | 1. A headset for use with a video game system, the headset comprising: a first bio-feedback device for transmitting to a video game system physiological data of a participant thereof comprising a respiratory voice sensor, wherein the respiratory voice sensor comprises a thermocouple;one or more second bio-feedback devices for transmitting to a video game system physiological data of a participant thereof, wherein the one or more second bio-feedback devices comprise at least one of an electroencephalogram electrode or a galvanometer;a communication link for electrically coupling the headset and the video game system; andconductive wiring for electrically coupling the first and second bio-feedback devices with the communication link. 2. The headset of claim 1, wherein the electroencephalogram electrode is configured to measure the electrical activity of surface cortical neurons emitted by a participant of the video game system. 3. The headset of claim 1, wherein the thermocouple comprises a polyvinylidine fluoride thermocouple. 4. The headset of claim 1, wherein the galvanometer comprises at least two conductors configured to measure a change in resistance of skin of a participant of the video game system, wherein the at least two conductors are in physical communication with the skin. 5. The headset of claim 1, wherein the headset further comprises: an ear piece comprising a speaker for transmitting an audio component from the video game system to the participant; anda microphone for transmitting verbal communication from the participant to the video game system. 6. A video game system comprising: a) a control deck;b) a headset configured to transmit a first set of physiological data of a participant, wherein the headset is electrically coupled with the control deck; andc) a controller configured to transmit conventional controller data inputted by the participant and a second set of physiological data of the participant, wherein the controller is electrically coupled with the control deck; andwherein the control deck comprises:a circuit board for directing the processing of the data transmitted by the headset and the controller;an analog to digital converter electrically coupled with the circuit board;a serial to parallel converter electrically coupled with the circuit board;a software input module; anda microprocessor electrically coupled with the circuit board and the software input module, wherein the microprocessor comprises executable instructions to alter game play of the video game based on the physiological data of the participant. 7. The video game system of claim 6, wherein: the video game system further comprises software; andthe software is a computer readable medium comprising executable instructions for providing a video game operable through the control deck. 8. The video game system of claim 7, wherein the headset and the controller transmit data to the control deck for processing into a form interpreted by the software suc... | - | A63F009/24 A61B005/00
US08203340 | B2 | 2012-06-19 | Magnetic resonance method and apparatus for generating a perfusion image | 2008-08-11 | In a magnetic resonance method and apparatus for generating perfusion images, a perfusion series of magnetic resonance perfusion images is acquired that includes tag images and at least one control image, that are grouped in pairs. From each pair an initially processed perfusion image is obtained, such as by subtraction. Each initially processed image is subjected to a quality control review by analysis with respect to at least one image quality criterion. Any initially processed image that does not satisfy the quality criterion is rejected. Only initially processed images that satisfy the quality criterion are combined to form a resultant magnetic resonance perfusion image. Artifacts in the resultant perfusion image are thereby reduced or avoided. | 1. A method for generating a magnetic resonance perfusion image, comprising the steps of: from a perfusion series of magnetic resonance images comprising a plurality of tag images and at least one control image, pairing said tag images and said at least one control image to produce a plurality of pairs and generating, from each pair, an initially processed magnetic resonance perfusion image;subjecting each initially processed image to a quality control review by analyzing each initially processed image with respect to at least one image quality criterion, and rejecting any of said initially processed images that do not satisfy said quality criterion, thereby obtaining a plurality of acceptable magnetic resonance perfusion images; andautomatically combining only said acceptable magnetic resonance perfusion images to obtain, as an output, a resultant magnetic resonance perfusion image. 2. A method as claimed in claim 1 wherein each of said initially processed images exhibits an intensity distribution, and employing a mean and standard deviation of said intensity distribution as said quality criterion. 3. A method as claimed in claim 1 wherein each of said initially processed images exhibits an energy content, and employing the root mean square of said energy content as said quality criterion. 4. A method as claimed in claim 1 wherein each of said initially processed images comprises edge features, and implementing an edge detection technique as said quality criterion. 5. A method as claimed in claim 4 comprising implementing said edge detection technique with a grad operator. 6. A method as claimed in claim 4 comprising implementing said edge detection technique with a Laplace operator. 7. A method as claimed in claim 1 comprising employing a reference image as said quality criterion, and comparing each initially processed image to said reference image. 8. A method as claimed in claim 1 comprising completely automatically implementing said quality control review. 9. A method as claimed in claim 1 comprising implementing said quality control review with user interaction. 10. A method as claimed in claim 1 comprising implementing said quality control review offline after acquisition of said series. 11. A method as claimed in claim 1 comprising implementing said quality control review in real time with acquisition of said series. 12. A method as claimed in claim 11 comprising displaying a representation of said quality criterion during acquisition of said series. 13. A method as claimed in claim 11 comprising generating a humanly perceptible indicator if a predetermined number of rejections occur in said quality control review. 14. A method as claimed in claim 1 comprising generating said initially processed magnetic resonance perfusion image from each pair by subtracting the control image from the tag image in each pair. 15. A method as claimed in claim 1 comprising dynamically updating said image quality criterion during acquisition of... | Siemens Medical Solutions USA, Inc. | G01V003/00
US08805482 | B2 | 2014-08-12 | System and method for signal quality indication and false alarm reduction in ECG monitoring systems | 2008-07-28 | An electrocardiogram (ECG) monitoring system configured to determine an index that comprises a set of indicators, where each indicator is indicative of a level of signal quality in an ECG lead. Further, the ECG monitoring system assigns a first indicator from the set of indicators to a first ECG lead signal and communicates the first indicator to a user. | 1. An electrocardiogram (ECG) monitoring system configured to: determine noise from a first ECG lead signal, wherein the determination of the noise comprises subtracting a signal template corresponding to a noise-free cardiac signal from the first ECG lead signal;assign a signal quality indicator from a set of possible signal quality indicators to the first ECG lead signal based on the determined noise; andcommunicate the assigned signal quality indicator to a user. 2. The ECG monitoring system of claim 1 wherein the set of possible signal quality indicators is derived from one of a number scale, a color spectrum, and an audible signal. 3. The ECG monitoring system of claim 1 further configured to: assign an additional indicator from the set of possible signal quality indicators to multiple ECG lead signals, wherein the additional indicator is indicative of an overall quality of the multiple ECG leads; andcommunicate the additional indicator to the user. 4. The ECG monitoring system of claim 1 further configured to remove a pacemaker-induced signal from the first ECG lead signal. 5. The ECG monitoring system of claim 1 wherein the determined noise is substantially free of single-subject normal ECG amplitude variance contributions. 6. The ECG monitoring system of claim 1 wherein the determined noise is substantially free of signal contributions from premature ventricular contractions and ventricular ectopic beats. 7. The ECG monitoring system of claim 1, wherein the determination of noise is configured to account for subject-to-subject variances. 8. The ECG monitoring system of claim 1 further configured to communicate to the user that a lead failure event is likely to occur based on the a trend associated with the assigned signal quality indicator. 9. The ECG monitoring system of claim 1 further configured to determine that an alarm generated by the ECG monitoring system is a false alarm, wherein the determination that the alarm is a false alarm is based on the assigned signal quality indicator. 10. The ECG monitoring system of claim 1 wherein the signal template comprises a plurality of beat templates. 11. The ECG monitoring system of claim 10 wherein the plurality of beat templates are substantially free of artifact-causing noise. 12. The ECG monitoring system of claim 1 wherein the determination of the noise comprises determining fiducial points of the first ECG lead signal. 13. The ECG monitoring system of claim 12 wherein the determination of the noise comprises aligning the signal template with the first ECG lead signal using the determined fiducial points. 14. The ECG monitoring system of claim 1 wherein the assignment of the assigned signal quality indicator is based on analyzing the determined noise with an exponential threshold function nonlinearity. 15. An electrocardiogram (ECG) monitoring system configured to: determine noise from a first ECG lead signal in two or more frequency segments, where the determination of the ... | General Electric Conpany | A61B005/04
US09180299 | B2 | 2015-11-10 | First time right placement of a DBS lead | 2007-11-16 | The disclosure is directed to a deep brain stimulation (DBS) lead having a distal end for providing therapeutic electrical stimulation to tissue in a stimulation target area of a patient's brain, comprising an array of one or more stimulation elements and sensing elements located at the distal end of the lead; each of the one or more stimulation elements is capable of providing electrical stimulation to the brain tissue in the target area; and each of the one or more sensing elements is capable of detecting electrical signals produced by nerve cells within the brain; wherein after the first implantation of the lead into the brain along a trajectory that is pre-determined by non-surgical procedures, the array of stimulation and sensing elements is capable of facilitating the location of the target area and the determination for each of the stimulation elements of the required stimulation parameters needed to provide the therapeutic stimulation to the brain tissue in the stimulation target area, without requiring any additional implantations of the lead after the first implantation. A stimulation system including the DBS lead and a pulse generator is disclosed. Also disclosed is a method of providing therapeutic DBS to brain tissue using the lead. | 1. A stimulation lead having a distal end for providing therapeutic electrical stimulation to tissue in a target area, the lead comprising: an array of a plurality of stimulation elements and a plurality of sensing elements located at the distal end of the lead, the array including columns of two or more of the plurality of stimulation elements, the columns extending down a longitudinal length of the stimulation lead and spaced from one another around the circumference of the stimulation lead, at least one of the plurality of stimulation elements being adjacent to and surrounded by two or more of the plurality of sensing elements, each of the plurality of sensing elements configured to detect electrical signals, at least one of the plurality of sensing elements being positioned between two or more of the plurality of stimulation elements, a number of the plurality of sensing elements being greater than a number of the plurality of stimulation elements, the plurality of sensing elements comprising a plurality of sets of sensing elements, each of the sets of sensing elements comprising two or more sensing elements, each of the plurality of stimulation elements being directly associated with one of the sets of sensing elements, the association being based on a geometric relation between which sets of sensing elements surround which ones of the plurality of stimulation elements, and each set of sensing elements is associated with only one stimulation element that is surrounded by respective set of sensing elements; anda controller configured to receive signals from one or more of the plurality of sensing elements, derive a spatial location and boundary of the target area based on locations of the sensing elements that sensed the signal, and determine that one or more of the plurality of stimulation elements that are surrounded by the respective sets of the sensing elements whose sensing elements sensed the signal lie in the target area and should provide the therapeutic electrical stimulation that covers the spatial location and the boundary of the target area, and control delivery of the therapeutic electrical stimulation via the determined stimulation elements. 2. The stimulation lead of claim 1, wherein the plurality of stimulation elements are metal electrodes. 3. The stimulation lead of claim 1, wherein the plurality of sensing elements are metal microelectrodes or capacitive sensing elements. 4. The stimulation lead of claim 1, wherein the plurality of stimulation elements and the plurality of sensing elements are distributed in an array circumferentially on the surface of the lead. 5. The stimulation lead of claim 1, wherein the plurality of sensing elements are each individually capable of detecting one or more of electrical signals, action potentials, field potentials, biochemical signals and neurotransmitters. 6. The stimulation lead of claim 1, wherein the area of one of the plurality of sensing elements is smaller than the area of tw... | Medtronic Bakken Research Center B. V. | A61N001/36 A61N001/05
US08515549 | B2 | 2013-08-20 | Associating therapy adjustments with intended patient posture states | 2009-04-30 | The disclosure described techniques for associating therapy adjustments with intended patient posture states. The techniques may include receiving a patient therapy adjustment to a parameter of a therapy program that defines electrical stimulation therapy delivered to the patient, identifying a posture state of the patient, and associating the patient therapy adjustment with the posture state when the patient therapy adjustment is within a range determined based on stored adjustment information for the identified posture state. | 1. A method comprising: receiving a patient therapy adjustment to a parameter of electrical stimulation therapy delivered to the patient;identifying a posture state of the patient; andassociating, by a processor, the patient therapy adjustment with the posture state when the patient therapy adjustment is within a range determined based on stored adjustment information for the identified posture state. 2. The method of claim 1, further comprising associating the therapy adjustment with the identified posture state when the therapy adjustment is outside of the range based on confirmation that the therapy adjustment is associated with the identified posture state. 3. The method of claim 2, further comprising receiving the confirmation from the patient. 4. The method of claim 3, further comprising prompting the patient to confirm or reject the association via a user interface. 5. The method of claim 1, wherein the adjustment information includes previous patient therapy adjustments for the program and the posture state, the method further comprising determining the range to be inclusive of at least a subset of the previous patient therapy adjustments for the program and the posture state. 6. The method of claim 5, further comprising setting a minimum value of the range to a lowest value of the previous patient therapy adjustments and setting a maximum value of the range to a highest value of the previous patient therapy adjustments. 7. The method of claim 5, wherein the range includes an upper threshold determined by an average of highest values of the previous patient therapy adjustments, and a lower threshold determined by an average of lowest values of the previous patient therapy adjustments. 8. The method of claim 5, further comprising determining the range based at least in part on a weighted average of previous patient therapy adjustments, wherein more recent therapy adjustments are weighted more highly than less recent therapy adjustments. 9. The method of claim 1, wherein identifying the posture state comprises identifying the posture state if the posture state does not change for a predetermined period of time following sensing of the posture state. 10. The method of claim 1, further comprising updating the range based on a newly received patient therapy adjustment. 11. A system comprising: a user interface that receives a patient therapy adjustment to a parameter of electrical stimulation therapy delivered to the patient; anda processor that:identifies a posture state of the patient; andassociates the patient therapy adjustment with the posture state when the patient therapy adjustment is within a range determined based on stored adjustment information for the identified posture state. 12. The system of claim 11, wherein the processor associates the therapy adjustment with the identified posture state when the therapy adjustment is outside of the range based on confirmation that the therapy adjustment is associated with the iden... | Medtronic, Inc. | A61B005/00 A61B005/02 A61B005/103 A61B005/11 A61N001/00 A61N001/05 A61N001/372
US09186093 | B2 | 2015-11-17 | Compositions and methods for screening cardioactive drugs | 2009-08-20 | We have developed a reduced Drosophila heart preparation in which dissection of the fly heart removes nervous system input and reveals its inherent myogenic activity, which can be preserved for several hours. High speed image capture combined with computer-based analytical packages allows us to generate the equivalent of M-mode traces obtained from ultrasounds of human hearts. | 1. A method of analyzing heart movement comprising: (a) producing a plurality of optical images of the heart as it beats; (b) using a Frame Brightness algorithm to identify changes in overall light intensity in said plurality of optical images to determine whether the heart is in a contracted or relaxed condition; (c) detecting movement of the heart by comparing intensity changes in individual pixels from one of said plurality of said optical images to another of said optical images; and (d) identifying a pause between contraction and relaxation movements of the heart. 2. The method of claim 1 further comprising using a Changing Pixel Intensity algorithm to detect said movement of the heart by comparing intensity changes in individual pixels from one of said plurality of said optical images to another of said optical images. 3. The method of claim 2, wherein the output from the Frame Brightness algorithm is used to inform the Changing Pixel Intensity algorithm whether the heart is in the contracted or relaxed condition. 4. The method of claim 1, further comprising producing a movie of the heart as it beats, wherein the movie comprises said plurality of optical images. 5. The method of claim 2, further comprising determining a member selected from the group consisting of: heart period, heart rate, diastolic interval, systolic interval, standard deviation of heart period, standard deviation of heart rate, standard deviation of diastolic interval, standard deviation of systolic interval, percent fractional shortening of heart contraction, arrhythmia index, velocity of heart muscle contraction, and directionality of heart muscle contraction. 6. The method of claim 2 wherein the heart is selected from the group consisting of a Drosophila heart, a zebrafish larval heart, a mouse embryo heart, and a human heart. 7. The method of claim 6 wherein the heart is an excised, denervated Drosophila heart that retains myogenic activity. | THE REGENTS OF THE UNIVERSITY OF CALIFORNIA BURNHAM INSTITUTE FOR MEDICAL RESEARCH | A61B006/00 A61B005/107 G01N033/50 A61B005/00 A61B008/08
US08888680 | B2 | 2014-11-18 | Method and apparatus for foreign matter detection for blood content sensors | 2008-07-07 | Apparatus and methods for detecting foreign material obstructing a blood content sensor window to facilitate higher data accuracy from such sensors, in accordance with one embodiment of the invention for detecting foreign matter with the blood content sensor when the blood content sensor is not contacted with living tissue, an illuminator for blood content detection and the blood content sensor share a common outer window. When the window is at a distance from tissue and in absence of foreign matter on the window, all or most of the emitted light will be dispersed and little or no light will be reflected back into the sensor. Conversely, when the window is at a distance from tissue and foreign matter is present on the window, a substantial portion of the emitted light will be reflected back into the sensor indicating the presence of the foreign matter. | 1. A method comprising the steps of: illuminating a region of living tissue within a cavity of a body; determining if a blood content sensor is in contact with said illuminated region of tissue; detecting reflected light by said blood content sensor if it is determined that said sensor is not in contact with said illuminated tissue; wherein if the determining step determines that said blood content sensor is in contact with said illuminated region of tissue, then detecting interacted light for determining an indicator of blood content of said illuminated tissue and determining if foreign matter exist on said sensor based on said detected reflected light and validating blood content data based on a previous determination of the existence of foreign matter on said sensor. 2. The method of claim 1 further comprising the step of guiding an endoscope in said cavity based on said blood content data. 3. The method of claim 2 wherein the illuminating step further comprises the step of providing illumination of said tissue to enable visible observation of said tissue with said endoscope. 4. The method of claim 3, wherein the light used for illuminating said tissue for visible observation is within a different wavelength range then the light wavelength range used for the detecting step. 5. The method of claim 4 wherein the light used for determining blood content is a different wavelength then the light used for foreign matter detection. 6. The method of claim 4, wherein the light used for illuminating tissue for visible observation is sequential emission of red, green and blue light. 7. The method of claim 4, wherein the light used for illuminating tissue for visible observation is broadband light. 8. The method of claim 3 wherein the light used for illuminating said tissue and for determining blood content are alternatively emitted. 9. A method for determining foreign matter on a blood content sensor comprising the steps of: illuminating mucosal tissue within a cavity of a body with polarized light in one wavelength range;determining if a blood content sensor is in contact with said illuminated mucosal tissue;detecting interacted light within a first wavelength range by said blood content sensor-to measure blood content within said illuminated tissue if it is determined that the blood content sensor is in contact with said illuminated mucosal tissue;detecting reflected light within a second wavelength range by said blood content sensor if it is determined that the blood content sensor is not in direct contact to measure blood content; anddetermining if foreign matter exist on said blood content sensor based on said detected light from the second wavelength range. 10. An apparatus for determining foreign matter on a blood content sensor comprising: an illuminator configured to illuminate mucosal tissue within a body cavity;a blood content sensor configured to detect reflected and/or interacted light from the mucosal tissue and generate a correspo... | Olympus Medical Systems Corp. | A61B001/04 A61B001/06 A61B001/12 A61B005/00 A61B001/00 A61B005/07
US08388597 | B2 | 2013-03-05 | Bag system with device for associating sampling receptacles | 2009-10-27 | The invention includes a bag system for collecting a biological fluid, in particular blood. The system includes a device for collecting the fluid which is in fluid communication with at least one fluid collection bag, and a device for sampling the fluid to be collected which includes at least one sampling receptacle. The sampling device includes a device for transferring the fluid from the bag system into the sampling receptacle or receptacles, in which the transfer device is provided with a device for associating the receptacle or receptacles. The transfer device may be arranged to allow the support of the receptacle or receptacles in a standby position, the guidance of the receptacle or receptacles to a transfer position and, after transfer, the dissociation of the receptacle or receptacles from the bag system. | 1. A bag system for collecting a biological fluid comprising: a collection device;a fluid collection bag in fluid communication with the collection device; anda sampling device in fluid communication with the collection device, the sampling device including:at least one sampling receptacle;a transfer device having an associating device having a hollow guide open at a front part to allow introduction of the at least one sampling receptacle and one or more flexible projections on an internal face of the hollow guide,wherein the one or more flexible projections are arranged so as to be deformable by sliding the at least one sampling receptacle inside the hollow guide, and wherein at least one of the one or more flexible projections is reversibly deformable from a position inclined towards the front part of the hollow guide to a position inclined towards a rear part of the hollow guide by contact of the at least one sampling receptacle while sliding inside the hollow guide in a front to rear direction, and from a position inclined towards the rear of the hollow guide to a position inclined toward the front of the hollow guide by contact of the at least one sampling receptacle while sliding inside the hollow guide in a rear to front direction, andwherein the associating device is operable to support the at least one sampling receptacle in a standby position, guide the at least one sampling receptacle to a transfer position, and allow dissociation of the at least one sampling receptacle from the bag system. 2. The system of claim 1, wherein the biological fluid comprises blood. 3. The system of claim 1, further comprising: a first tube to establish fluid communication between the collection device and the fluid collection bag; anda second tube to establish fluid communication between the collection device and the sampling device. 4. The system of claim 3, wherein the sampling device further comprises a sampling bag connected to the downstream end of the second tube. 5. The system of claim 1, wherein the receptacle has a body having a first diameter and further comprises a closure element having a second diameter greater than the first diameter. 6. The system of claim 1, wherein the transfer device further comprises: a hollow needle in fluid communication with the bag system,wherein the hollow needle passes through the rear part of the hollow guide so that a downstream part of the needle extends inside the hollow guide and an upstream part of the needle extends outside the hollow guide, and wherein the hollow needle is operable to perforate a closure element of the sampling receptacle, so as to place the downstream part of the needle inside the sampling receptacle. 7. The system of claim 6, wherein the one or more flexible projections are arranged so as to permit sliding of the receptacle inside the guide between the standby position at a distance from the needle and the transfer position. 8. The system of claim 6, further comprising a set of pro... | Macopharma | A61B019/00 A61B005/155 A61M005/32
US09192328 | B2 | 2015-11-24 | Signal processing for continuous analyte sensor | 2009-09-23 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A method for estimating analyte values from a continuous analyte sensor, the method comprising: receiving a data stream from a continuous analyte sensor for a time period; andestimating, using a processor module, one or more analyte values for a future time period, wherein estimating comprises predicting one or more analyte values for the future time period based on the data stream, applying one or more physiological parameters to the predicted one or more analyte values, and obtaining the estimated one or more analyte values based on the application of the one or more physiological parameters to the predicted one or more analyte values. 2. The method of claim 1, wherein the one or more physiological parameters include a predetermined rate of change. 3. The method of claim 2, wherein the predetermined rate of change is at least 4 mg/dL/min. 4. The method of claim 1, wherein the one or more physiological parameters include a predetermined acceleration. 5. The method of claim 4, wherein the predetermined acceleration is at least 0.1 mg/dL/min2. 6. The method of claim 1, further comprising providing output based on the estimated one or more analyte values. 7. The method of claim 6, wherein providing output comprises displaying information on a display device based on the estimated one or more analyte values. 8. The method of claim 6, wherein providing output comprises alarming responsive to the estimated one or more analyte values meeting one or more criteria. 9. The method of claim 1, wherein estimating is responsive to an event trigger. 10. The method of claim 1, wherein the application of the one or more physiological parameters to the predicted one or more analyte values comprises determining physiological boundaries based on the one or more physiological parameters and determining whether the predicted one or more analyte values fall within the physiological boundaries. 11. The method of claim 10, wherein only predicted values that fall within the physiological boundaries are used to obtain the one or more estimated values. 12. A system for estimating analyte values from a continuous analyte sensor, the system comprising: an input module operably connected to a continuous analyte sensor and configured to receive a data stream from the continuous analyte sensor for a time period; anda processor module comprising programming configured to estimate one or more analyte values for a future time period, wherein the programming configured to estimate is configured to predict one or more analyte values for the future time period based on the data stream and obtain the estimated one or more analyte values based on an application of one or more physiological parameters to the predicted one or more analyte values. 13. The system of claim 12, wherein the one or more physiological parameters include a predetermined rate of change. 14. The system of claim 13, wherein the predetermined rate of change is at least 4 mg/dL/min. 15. The system of... | DexCom, Inc. | A61B005/1468 A61B005/1486 A61B005/1495 A61B005/145 A61B005/00
US20100010359 | A1 | 2010-01-14 | DETECTING PROLONGED MYOCARDIAL REPOLARIZATION INDICATIVE OF CARDIAC CONDITION | 2009-03-09 | The present invention comprises methods and apparatus for detecting prolonged myocardial repolarization as an indicator of cardiac conditions, including without limitation, transmural ischemia. In some embodiments, without limitation, the present invention comprises methods and apparatus to detect prolonged repolarization using electrocardiographic and electrophysiological tools and measurements. | 1. A method for detecting prolonged myocardial repolarization as an indicator of transmural ischemia or infarction in a mammalian subject, comprising the steps of: obtaining baseline electrocardiogram data derived from the subject;performing an electrocardiogram on the subject after the baseline data was derived from the subject; andcomparing the baseline electrocardiogram data with data from the electrocardiogram to determine whether the time interval of myocardial repolarization is increased. 2. The method of claim 1, wherein the comparing step comprises comparing QT interval data in the baseline data with QT interval data from the electrocardiogram. 3. The method of claim 1, wherein the comparing step is automated in whole or in part. 4. The method of claim 1, wherein the comparing step is manual in whole or in part. 5. The method of claim 1, wherein the subject is a human. 6. The method of claim 5, wherein the human subject is suspected of having suffered myocardial ischemia after the baseline electrocardiogram data was derived from the subject. 7. A method for detecting prolonged myocardial repolarization as an indicator of transmural ischemia or infarction in a mammalian subject, comprising the steps of: obtaining baseline cardiac electrical data derived from the subject;measuring cardiac electrical activity in the subject after the baseline data was derived from the subject; andcomparing the baseline data with data from the measurement of the cardiac electrical activity to determine whether the time interval of myocardial repolarization is increased. 8. The method of claim 7, wherein the comparing step comprises comparing activation recovery interval data in the baseline data with activation recovery interval data from the measurement of cardiac electrical activity. 9. The method of claim 7, wherein the comparing step comprises comparing monophasic action potential data in the baseline data with monophasic action potential data from the measurement of cardiac electrical activity. 10. A method of detecting myocardium at risk or myocardial viability by measurement of changes in myocardial repolarization in a mammalian subject, comprising the steps of: performing an electrocardiogram on the subject to obtain baseline electrocardiogram data;performing a clinical procedure on the subject;monitoring the subject's electrocardiogram during the clinical procedure to obtain electrocardiogram data; andcomparing the baseline electrocardiogram data with the electrocardiogram data from the clinical procedure to determine whether the time interval of myocardial repolarization is changed during the procedure. 11. The method of claim 10, wherein the subject is a human. 12. A method of detecting myocardium at risk or myocardial viability by measurement of changes in myocardial repolarization in a mammalian subject, comprising the steps of: measuring baseline cardiac electrical data in the subject;performing a clinical procedure on the subject;mea... | - | A61B005/0452
US20100016688 | A1 | 2010-01-21 | Optical methods for real time monitoring of tissue treatment | 2009-02-20 | The present invention comprises methods and systems/devices for non-invasively measuring baseline collagen and collagen changes during treatment of tissue, e.g., denaturation by the application of RF energy, through linear dichroism, circular dichroism, or birefringence. The invention optionally uses polarization sensitive optical measurements to discriminate between denaturation of unidirectionally oriented strands of collagen, such as a ligament or tendon, and denaturation of planar collagen surfaces, such as the dermal layer of the skin or collagen in joint capsules. | 1.-29. (canceled) 30.) A method of monitoring a change in a collagen structure in a tissue, the method comprising: exposing the tissue to a first light;measuring a first reflected light which is reflected from the collagen structure;exposing the tissue or the collagen structure to one or more treatments, which treatments can alter the collagen structure, thereby producing a treated collagen structure;exposing the tissue to a second light;measuring a second reflected light which is reflected from the treated collagen structure;comparing the first reflected light and the second reflected light, thereby monitoring the change in the collagen structure. 31.) The method of claim 30, further wherein the tissue comprises a second collagen structure. 32.) The method of claim 31, wherein the first and the second collagen structures are monitored simultaneously or sequentially. 33.) The method of claim 31, wherein the first collagen structure is closer to the light and/or closer to the treatment than the second collagen structure. 34.) The method of claim 30, wherein the collagen structure comprises dermal collagen, mucosal collagen, synovial collagen, a tendon, a ligament, a fascia, or an aponeourosis. 35.) The method of claim 30, wherein the tissue comprises skin, a fascia, an aponeurosis, a tendon, a ligament, a capsule, a vascular wall, a vaginal wall, an introitus, or a urethra. 36.) The method of claim 30, wherein the treatment comprises application of physical energy, application of radio frequency waves, application of ultrasound, application of heat, application of cold, or application of a cosmeceutical. 37.) The method of claim 30, wherein the first and second light and the first and second reflected light comprise polarized light. 38.) The method of claim 37, wherein comparing the light comprises comparing the polarization of the light 39.) The method of claim 37, wherein the polarized light is linear polarized light. 40.) The method of claim 37, wherein the polarized light is circular polarized light. 41.) The method of claim 30, wherein the light comprises infra red light. 42.) The method of claim 30, wherein the first and second light comprises a wavelength of about 800 to about 1100 nm. 43.) The method of claim 30 wherein the light comprises UV light. 44.) The method of claim 30, wherein comparing the first and second reflected lights comprises determination of DoLP, θ, DoHP, DoVP, or angular dependence of the birefringence of the biological structure relative to an input polarization. 45.) A system or device for monitoring a change in one or more collagen structures in a tissue, the system or device comprising: a light source component configured to emit light to the tissue;one or more light polarizer components configured to polarize light from the light source or to polarize light reflected from the tissue;one or more lens components configured to focus light from the light source or to focus light reflected from the tissu... | Alpha Orthopaedics, Inc. | A61B005/1455 A61N007/00 A61B018/00 A61M031/00
US20100036221 | A1 | 2010-02-11 | Noninvasive Method to Estimate Variation of Blood Glucose Levels Using Metabolic Measurements | 2008-01-10 | A method to estimate the variation of blood glucose concentration in a patient without blood sampling. The method can be implemented by measuring metabolic parameters including heat dissipation by conduction, percentage oxygen content of expired air and volume per minute of expired air. These parameters are used to calculate estimated blood glucose variation. | 1-8. (canceled) 9. A method for estimating blood glucose variation in a patient based on metabolic parameters, comprising: measuring raw metabolic parameters of the patient, wherein the raw metabolic parameters comprise conduction heat loss of an extremity of the patient, oxygen content of air expired by the patient, and a rate of expiration of air expired by the patient;converting the raw metabolic parameters into metabolic features of the patient; andcalculating an estimated blood glucose variation for the patient based on the metabolic features. 10. The method of claim 9, wherein conduction heat loss is measured via a heat flow sensor. 11. The method of claim 10, wherein the heat flow sensor comprises a Data Harvest EasySense Advanced Datalogger™. 12. The method of claim 9, wherein oxygen content is measured via an oxygen analyzer. 13. The method of claim 12, wherein the oxygen analyzer comprises a Teledyne AX300 Oxygen Analyzer with R-17 MED Oxygen Sensor™. 14. The method of claim 9, wherein the rate of expiration of air is measured via a gas speed-measuring device. 15. The method of claim 14, wherein the gas speed-measuring device comprises a Vitalograph Micro Spirometer™. 16. The method of claim 9, wherein the calculating step comprises use of a classification model, wherein the classification model takes the metabolic features as inputs and provides the estimated blood glucose variation as an output. 17. The method of claim 16, wherein the classification model classifies the estimated blood glucose variation into one of up to five classes. 18. The method of claim 9, wherein the calculating step comprises use of a multiple regression analysis, wherein the multiple regression analysis is used to calculate the estimated blood glucose variation from the metabolic features. 19. An apparatus comprising one or more memory units having computer-useable instructions encoded thereon for performing a method of classifying blood glucose variation of a patient based on metabolic parameters, the method comprising: receiving raw metabolic parameters of the patient;converting, in a first computer process, the raw metabolic parameters into metabolic features of the patient;computing, in a second computer process, variations between the metabolic features of the patient and stored metabolic features of the patient; andclassifying, via a third computer process, the blood glucose variation of the patient based on the computed variations. 20. The apparatus of claim 19, wherein the raw metabolic parameters comprise conduction heat loss of an extremity of the patient, oxygen content of air expired by the patient, and a rate of expiration of air expired by the patient. 21. The apparatus of claim 20, wherein the metabolic features comprise heat dissipation and a resting metabolic rate. 22. The apparatus of claim 19, wherein the third computer process comprises a regression analysis. 23. The apparatus of claim 22, wherein the third computer process c... | - | A61B005/145
US20100036378 | A1 | 2010-02-11 | THERMAL IMAGING FEEDBACK FOR OPTIMIZING RADIO FREQUENCY ABLATION THERAPY | 2007-10-19 | The invention relates to methods and systems for monitoring and regulating radiofrequency ablation therapy in order to maximize effectiveness of treatment. The invention uses an imaging scanner to provide feedback regarding location and extent of a treated volume. The feedback is used as input data for control of intensity, duration, and/or placement of radiofrequency treatment. Control of treatment parameters is automatic and/or modulated by an operator. | 1. A method for monitoring and regulating radiofrequency (RF) ablation therapy, the method comprising: providing an image of a target volume and surrounding tissue using an imaging scanner, and inserting an RF probe into the target volume;generating an RF current to heat the target volume near at least one tip of the RF probe, and using imaging data as feedback; andaltering at least one parameter selected from the group consisting of: RF power, exposure time, and position of the RF probe, in response to an imaging feedback signal transmitted to an RF power generator and/or displayed to an operator. 2. The method according to claim 1, wherein the scanner comprises an ultrasound scanner and/or a magnetic resonance scanner. 3. The method according to claim 1, wherein the image comprises an ultrasound real-time image and/or a magnetic resonance image. 4. The method according to claim 1, wherein a boundary of the target volume is determined by an automatic method and/or a manual method. 5. The method according to claim 1, wherein the image guides inserting or adjusting placement of the RF probe in the target volume. 6. The method according to claim 1, wherein altering the parameter further comprises computing a temperature elevation in the target volume. 7. The method according to claim 6, wherein the temperature elevation is derived from the image. 8. The method according to claim 7, wherein the temperature elevation is measured through temperature-dependent changes in imaging properties of tissue. 9. The method according to claim 1, wherein altering the parameter further comprises computing an estimated accumulated thermal dose. 10. The method according to claim 9, wherein the estimated accumulated thermal dose is derived from the image. 11. The method according to claim 1, wherein altering the parameter further comprises estimating a location or a volume of a thermally affected region, to calculate an estimated coagulated volume (ECV). 12. The method according to claim 11, wherein the ECV is derived from the image. 13. The method according to claim 1, wherein altering the parameter further comprises comparing the ECV to the target volume. 14. A system for monitoring and regulating radiofrequency (RF) ablation therapy comprising: an imaging scanner;a RF probe for inserting into a predetermined target volume;a radiofrequency power generator for providing power to the RF probe;a feedback signal from the scanner; anda feedback mechanism, wherein the feedback signal initiates a feedback event, thereby monitoring and regulating RF ablation therapy. 15. The system according to claim 14, wherein the imaging scanner comprises an ultrasound scanner and/or a magnetic resonance scanner. 16. The system according to claim 14, wherein the imaging scanner computes at least one of the following: temperature elevation in the target volume, an accumulated thermal dose in the target volume, and an estimated coagulated volume (ECV), to produce a resultin... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B018/14 A61B005/055 A61B008/12
US20100002233 | A1 | 2010-01-07 | IMAGING OF TURBID MEDIUM | 2007-10-24 | The invention relates to imaging of a turbid medium, for example in connection with optical mammography. A device for imaging a turbid medium (20) is disclosed, the device comprising: a holder (20) arranged for receiving the turbid medium and a matching fluid (21); one or more radiation sources (24) and one or more photodetectors (25). The matching fluid is a vapor with one or more optical properties of the matching fluid substantially matching the corresponding one or more optical properties of the turbid medium. In an embodiment, the matching fluid (21) is a composite vapor comprising at least two components. | 1. A device for imaging a turbid medium (1, 20), the device comprising: a holder (20) arranged for receiving the turbid medium and a matching fluid (7, 21, 53);one or more radiation sources (3, 24) for irradiating the turbid medium and the matching fluid;one or more photodetectors (4, 25) for measuring the intensity of the radiation;wherein the matching fluid is a vapor with one or more optical properties of the matching fluid substantially matching the corresponding one or more optical properties of the turbid medium. 2. The device according to claim 1, wherein the matching fluid (7, 21, 53) is a composite vapor comprising at least two components (50, 51). 3. The device according to claim 1, wherein the vapor comprises a first scattering component (51) dissolved in droplets of a second component (50). 4. The device according to claim 2, wherein the matching fluid comprises a component with a transport mean-free path, ltra, below 3 millimeter. 5. The device according to claim 3, wherein the size of the droplets of the second component is larger than the transport mean-free path of the scattering component of the first component. 6. The device according to claim 3, wherein the ratio between the refractive index of the first scattering component and the second component is larger than 1.5. 7. The device according to claim 3, wherein the first scatting component is titanium dioxide and the second component is water. 8. The device according to claim 3, wherein a dye is added to the second component. 9. The device according to claim 1, further comprising a nebulizer (27) and wherein the vapor is generated by the nebulizer. 10. The device according to claim 1, further comprising a device (28) for generating sound waves for randomizing the position of the particles in the vapor. 11. The device according to claim 1, wherein the radiation source irradiates the turbid medium at a selected wavelength such that at the selected wavelength one or more selected optical properties of the matching fluid substantially matching the corresponding optical properties of the turbid medium. 12. The device according to claim 1, further comprising a processing unit for deriving an image of the turbid medium from the measured intensities. 13. The device according to claim 1, wherein the one or more of the optical properties are one or more attenuation coefficients. 14. The device according to claim 1, wherein the one or more of the optical properties are such that the scattering and absorption properties are higher than those of water. 15. The device according to claim 1, wherein the vapor is a mist or smoke. 16. A method of imagining a turbid medium, the method comprising: arranging (60) in a holder the turbid medium and a matching fluid;irradiating (61) the turbid medium and the matching fluid with one or more radiation sources;measuring (62) the intensity of the radiation by one or more photodetectors;wherein the matching fluid is selected as a vapor with... | Koninklijke Philips Electronics N.V. | A61B005/00 G01N021/00
US07660620 | B2 | 2010-02-09 | Timing techniques for magnetic resonance imaging | 2003-09-29 | The invention is directed to techniques in which magnetic resonance imaging (MRI) is coordinated with the operation of an implantable medical device (IMD). By using an IMD to sense conditions, MRI can be improved because the sensed conditions can accurately define timing for application of electromagnetic radiation bursts. Moreover, by applying stimulation pulses specifically to coordinate the electromagnetic radiation bursts, the MRI may also be improved. | 1. A method of performing magnetic resonance imaging (MRI) on tissue of a patient having an implantable medical device (IMD) with a telemetry unit communicating timing information as to operational conditions of the IMD, comprising: receiving the timing information communicated from the (IMD);applying a plurality of discrete MRI electromagnetic radiation bursts to the tissue of the patient, wherein said applied radiation bursts are synchronized with operational conditions of the IMD based upon the received timing information;imaging the tissue upon a display, whereby the imaged tissue is displayed in a substantially common state during each of said plurality of radiation bursts, further comprising:stimulating a patient with the IMD based upon the timing information to produce a specified cardiac rhythm;imaging the heart tissue based on timing information during application of the stimulation applied to the patient by the IMD; andat least one of storing and displaying at least some of said plurality of images when the heart tissue is at common points during the produced cardiac rhythm. 2. A method according to claim 1, further comprising sensing cardiac activity with the IMD and wherein the timing information includes timing of sensed cardiac activity. 3. A method according to claim 1, further comprising sensing physiologic conditions of the patient with the IMD. 4. A method according to claim 1, further comprising applying cardiac pacing pulses using the IMD and wherein the timing information defines timing of cardiac pacing stimulation pulses applied to the patient by the IMD. 5. The method of claim 1, wherein stimulation based on the timing information is intended to induce an arrhythmia. 6. The method of claim 1, wherein the IMD is a pacemaker and wherein the timing information defines a timing of a cardiac cycle. 7. The method of claim 1, wherein performing the MRI includes applying one or more electromagnetic radiation bursts based on the timing information. 8. The method of claim 1, wherein performing the MRI includes applying one or more gradient magnetic fields based on the timing information. 9. An apparatus comprising: means for receiving timing information from an cardiac-based implantable medical device (IMD) related to timed operating conditions of the IMD;means for performing magnetic resonance imaging (MRI) of heart tissue by applying electromagnetic radiation based upon the received timing information; andmeans for synchronizing application of electromagnetic radiation with the received timing information, whereby the heart tissue being imaged is placed in a substantially common state during each application of radiation, wherein the IMD comprises means for applying cardiac stimulation pulses and wherein the timing information defines timing of cardiac stimulation pulses applied to the patient by the IMD to produce a specified cardiac rhythm and wherein the MRI comprises:means for imaging the cardiac tissue during applicat... | Medtronic, Inc. | A61B005/05
US07668342 | B2 | 2010-02-23 | Method of bioimage data processing for revealing more meaningful anatomic features of diseased tissues | 2005-09-09 | The present invention discloses a method for generating elevation maps or images of a tissue layer/boundary with respect to a fitted reference surface, comprising the steps of finding and segmenting a desired tissue layer/boundary; fitting a smooth reference surface to the segmented tissue layer/boundary; calculating elevations of the same or other tissue layer/boundary relative to the fitted reference surface; and generating maps of elevation relative to the fitted surface. The elevation can be displayed in various ways including three-dimensional surface renderings, topographical contour maps, contour maps, en-face color maps, and en-face grayscale maps. The elevation can also be combined and simultaneously displayed with another tissue layer/boundary dependent set of image data to provide additional information for diagnostics. | 1. A method for generating image maps from a 3D image data set, said data set obtained from imaging a sample using an imaging device, said sample having at least one boundary surface associated therewith, said method comprising the steps of: identifying a subset of the image data corresponding to the boundary surface;fitting a substantially smooth reference surface to the subset of image data;calculating the distance between points on the reference surface to points of interest within the sample identified from the 3D image data set;generating an image map of the calculated distances; anddisplaying the image map. 2. A method as recited in claim 1, wherein the points of interest correspond to points on said boundary surface. 3. A method as recited in claim 1, wherein the points of interest correspond to points on a different boundary surface. 4. A method as recited in claim 1, wherein the sample is skin. 5. A method as recited in claim 1, wherein the sample is a blood vessel. 6. A method as recited in claim 1, wherein the sample is an eye. 7. A method as recited in claim 6, wherein the boundary surface is the cornea. 8. A method as recited in claim 6, wherein the boundary surface is the retina. 9. A method as recited in claim 1, wherein the 3D image data set is obtained using an OCT system. 10. A method as recited in claim 1, wherein the 3D image data set is obtained from a confocal microscope. 11. A method as recited in claim 1, wherein the 3D image data set is obtained from an ultrasound scanner. 12. A method for generating information from a 3D image data set, said 3D data set obtained from imaging a sample using an imaging device, said sample having at least one boundary surface associated therewith, said method comprising the steps of: identifying a subset of the image data corresponding to the boundary surface;fitting a substantially smooth reference surface to the subset of image data;calculating the distance between points on the reference surface to points of interest within the sample identified from the 3D image data set; andgenerating a 2D data set of the calculated distances, wherein the identifying, fitting, calculating and generating steps are performed in a processor. 13. A method as recited in claim 12, further including the step of generating an image map from the 2D data set. 14. A method as recited in claim 12, wherein the 3D image data set is obtained using an OCT system. 15. A method as recited in claim 14, wherein the sample is an eye. 16. A method as recited in claim 15, wherein the points of interest correspond to points on said boundary surface. 17. A method as recited in claim 15, wherein the points of interest correspond to points on a different boundary surface. 18. A method as recited in claim 15, wherein the boundary surface is the cornea. 19. A method as recited in claim 15, wherein the boundary surface is the retina. 20. A method as recited in claim 12, wherein regions of the boundary surface tha... | Carl Zeiss Meditec, Inc. | G06K009/00 A61B005/103 G01B009/02
US07667624 | B2 | 2010-02-23 | Methods and apparatus for clinical data compression | 2007-01-22 | A method for compressing clinical data includes monitoring a clinical parameter of a patient using a digital electronic monitoring device, wherein the clinical parameter is a function of time. A processor is utilized to locate exceptional values of the clinical parameter in accordance with predetermined criteria. The method further includes electronically encoding and communicating diagnostically significant information concerning the exceptional values of the clinical parameters to a clinical database repository (CDR) and information concerning normal values of the clinical parameters to the CDR, wherein the communicated exceptional values of clinical parameters and the communicated normal values of clinical parameters are encoded using different methods. | 1. A method for compressing clinical data, said method comprising: selecting a first encoding method and a second encoding method, the first encoding method and the second encoding method being fixed during a monitoring session of a first patient, the second encoding method providing a compression ratio that is higher than a compression ratio provided by the first encoding method;monitoring a clinical parameter of the first patient during the monitoring session using a digital electronic monitoring device, wherein the clinical parameter is a function of time;locating exceptional values of the clinical parameter in accordance with predetermined criteria using a processor;electronically encoding diagnostically significant information concerning the exceptional values of the clinical parameters using the first encoding method;electronically encoding information concerning normal values of the clinical parameters, wherein the normal values of clinical parameters and the predetermined criteria are encoded using the second encoding method; andcommunicating the differently encoded exceptional values of clinical parameters, the normal values of clinical parameters, and the predetermined criteria to a clinical database repository (CDR). 2. A method in accordance with claim 1 wherein electronically encoding information concerning normal values of the clinical parameter comprises run length encoding the information concerning the normal values of clinical parameters. 3. A method in accordance with claim 1 wherein electronically encoding diagnostically significant information comprises encoding a time series of values of the monitored clinical parameter. 4. A method in accordance with claim 1 wherein the digital electronic monitoring device is an electrocardiograph machine. 5. A method in accordance with claim 4 wherein the clinical parameter is the electrocardiogram signal. 6. A method in accordance with claim 5 wherein electronically encoding information concerning normal values of the clinical parameter comprises encoding a length of time during which the electrocardiogram signal is within normal values according to the predetermined criteria and electronically encoding diagnostically significant information comprises encoding the time varying electrocardiogram signal during a period within a length of time during which the electrocardiogram signal is outside of the normal values. 7. A method in accordance with claim 1 wherein the digital electronic monitoring device is one of a heart rate monitor, a sleep apnea monitor, and a vital sign monitor. 8. A method in accordance with claim 1 further comprising storing the differently encoded communicated exceptional values of clinical parameters, the communicated normal values of clinical parameters, and the predetermined criteria in the CDR. 9. A method in accordance with claim 8, further comprising, for at least a second patient: monitoring said clinical parameter of the second patient;electronically e... | General Electric Company | H03M007/30 A61B005/04
US07643872 | B2 | 2010-01-05 | System and method for detecting cardiac ischemia based on T-Waves using an implantable medical device | 2007-03-15 | A technique is provided for detecting episodes of cardiac ischemia based on an examination of the total energy of T-waves. Since cardiac ischemia is often a precursor to acute myocardial infarction (AMI) or ventricular fibrillation (VF), the technique thereby provides a method for predicting the possible onset of AMI or VF. Briefly, the technique integrates internal electrical cardiac signals occurring during T-waves and then compares the result against a running average. If the result exceeds the average by some predetermined amount, ischemia is thereby detected and a warning signal is provided to the patient. The maximum slope of the T-wave is also exploited. Techniques are also set forth herein for reliably detecting T-waves, which help prevent P-waves from being misinterpreted as T-waves on unipolar sensing channels. The T-wave detection technique may be used in conjunction with ischemia detection or for other purposes. | 1. In an implantable medical device for implant within a patient, a system comprising: a T-wave detection subsystem operative to detect individual T-waves;a T-wave energy integration subsystem operative to determine a total energy for each of the individual T-waves; anda cardiac ischemia detection subsystem operative to: a) calculate a running average of the total energies of the detected T-waves;b) calculate the difference between the total energy of a last detected T-wave and the running average;c) compare the energy difference to a predetermined energy threshold value;repeat a), b) and c) for a plurality of cardiac beats; anddetect cardiac ischemia when the difference exceeds the predetermined energy threshold value for a predetermined number of cardiac beats. 2. The system of claim 1 wherein the detection subsystem is operative to detect termination of an episode of ischemia when the energy difference between the total energy and the running average of the total energies is less than the predetermined energy threshold value for a predetermined number of cardiac beats. 3. The system of claim 1 further comprising a T-wave slope determination subsystem operative to determine a maximum slope of individual T-waves and wherein the cardiac ischemia detection subsystem is operative to: d) calculate a running average of the maximum slopes of the detected T-waves;e) calculate the difference between the maximum slope of a last detected T-wave and the slope running average;f) compare the slope difference to a predetermined slope threshold value;repeat d), e) and f) for a plurality of cardiac beats; anddetect cardiac ischemia when the slope difference exceeds the predetermined slope threshold value for a predetermined number of cardiac beats. 4. The system of claim 3 wherein the detection subsystem is operative to detect termination of an episode of ischemia when the difference between the maximum slope and the running average slope is less than the slope threshold value for a predetermined number of cardiac beats. 5. A method comprising: detecting a plurality of individual T-waves;calculating a total energy for each of the individual T-waves;a) calculating a running average of the total energies of the detected T-waves;b) calculating the difference between the total energy of a last detected T-wave and the running average;c) comparing the energy difference to a predetermined energy threshold value;repeating a), b) and c) for a plurality of cardiac beats; andindicating cardiac ischemia when the energy difference exceeds the predetermined energy threshold value for a predetermined number of cardiac beats. 6. The method of claim 5 wherein upon detection of ischemia: determining the difference between the total energy and the average of the total energies;comparing the energy difference to the energy threshold value; andindicating termination of ischemia when the energy difference is less than the energy threshold value for a predetermined number of car... | Pacesetter, Inc. | A61B005/0452 A61N001/36
US20100049265 | A1 | 2010-02-25 | EMI/ESD hardened sensor interface for a closed loop neuromodulator | 2009-08-21 | This document discusses, among other things, apparatus and mthiods for protecting an input and electronics of a closed loop neuromodulator. Electromagnetic interference (EMI) can be reduced using an EMI filter. Energy surges can be dissipated using a electrostatic discharge (ESD) device. An analog and a digital output can be produced using respiration information received at an input, each output indicative of resiration effort. | 1. A hardened sensor interface for protecting an input or electronics of a closed loop neuromodulator, configured to optimize sleep therapy for a patient, from electromagnetic interference (EMI) or electrostatic discharge (ESD), the interface comprising: an input configured to receive respiration information from a respiration sensor;an EMI filter coupled to the input and a grounded metal shield, the EMI filter configured to reduce EMI between the input and the electronics;an ESD protection device coupled to the grounded metal shield, the ESD protection device configured to protect the electronics from an energy surge; andwherein the closed loop neuromodulator is configured to receive respiration information from the input and to produce separate analog and digital respiration effort outputs indicative of respiration effort of the patient using the received respiration information. 2. The interface of claim 1, wherein the closed loop neuromodulator is configured to detect a sleep disorder event using the digital respiration effort output. 3. The interface of claim 2, wherein the closed loop neuromodulator is configured to provide a first stimulus to the patient in response to the detected sleep disorder event, the first stimulus configured to not interrupt the detected sleep disorder event; and wherein the closed loop neuromodulator is configured to provide a second stimulus to the patient in response to the detected sleep disorder event, following the first stimulus, the second stimulus having more energy than the first stimulus and configured to interrupt the detected sleep disorder event. 4. The interface of claim 1, wherein the EMI filter includes an external EMI filter and an internal EMI filter, wherein the internal EMI filter is configured to suppress electromagnetic interference generated from the closed loop neuromodulator. 5. The interface of claim 4, wherein the internal EMI filter is enclosed within the grounded metal shield. 6. The interface of claim 5, wherein the internal EMI filter is configured to attenuate interference signals having a frequency in a range from about 2 kHz to about 500 MHz. 7. The interface of claim 5, wherein the external EMI filter is located outside the grounded metal shield. 8. The interface of claim 5, wherein the external EMI filter is configured to attenuate interference signals having a frequency in a range from about 100 MHz to about 10 GHz. 9. The interface of claim 5, wherein the external EMI filter includes ferrite beads. 10. The interface of claim 1, wherein the ESD protection device includes a plurality of interleaved vias connecting first and second ground planes, wherein the closed loop neuromodulator include a printed circuit board, and wherein the first ground plane is disposed on a first side of the printed circuit board and the second ground plane is disposed on a second side of the printed circuit board, the second ground plane opposite the first ground plane. 11. A system for prot... | Dymedix Corporation | A61N001/00 A61B005/00 A61B005/08
US20100041960 | A1 | 2010-02-18 | HEALTH MANAGEMENT DEVICE | 2009-08-13 | A health management device is disclosed in embodiments of the present invention. The health management device comprises at least one sensor, a microprocessor and a memory unit. The sensor measures the status of a target and generates a sensing signal. The microprocessor, coupling to the sensor, processes the sensing signal and generates a data signal. The memory unit, coupling to the microprocessor, generates a measurement data according to the data signal and stores the measurement data. Wherein, at least one health management software is embedded in the memory unit and computes the data signal to obtain the measurement data. | 1. A health management device, comprising: at least one sensor for measuring physiology information and generating a sensing signal;a microprocessor coupling to the sensor for processing the sensing signal and generating a data signal; anda memory unit coupling to the microprocessor for generating a measurement data according to the data signal and for storing the measurement data;wherein at least one health management software is embedded in the memory unit and is executed on at least a compatible platform to compute and analyze the data signal to obtain the measurement data according to the operation by a user. 2. The device according to claim 1, wherein the memory unit is a swappable/swap device. 3. The device according to claim 1, wherein the health management software is proprietary native software or cross-platform software and the health management software performs one task selected from the group consisting of the following or combination thereof: data statistics, management, and operation. 4. The device according to claim 1, wherein the measurement data is selected from the group consisting of the following or combination thereof: the average of measurement, the trend chart of measurement. 5. The device according to claim 1, further comprising: a transmission interface that is a means of communication among the sensor, the microprocessor, and the memory unit. 6. The device according to claim 1, wherein measurement data is outputted via a transmission interface. 7. The device according to claim 6, wherein the transmission interface is a universal serial bus (USB) transmission interface and the transmission interface comprises at least one mass storage class. 8. The device according to claim 6, wherein measurement data is outputted via a transmission interface to a computer system or a compatible platform. 9. The device according to claim 1, wherein the memory unit is a flash memory. 10. A health management device, comprising: at least one sensor for measuring physiology information and generating a sensing signal;a microprocessor coupling to the sensor for processing and analyzing the sensing signal and generating a data signal;a memory unit coupling to the microprocessor for generating a measurement data according to the data signal and for storing the measurement data wherein at least one health management software is embedded in the memory unit and is executed on a compatible platform to compute and analyze the data signal to obtain the measurement data according to the operation by a user; anda transmission interface coupling to the memory unit for outputting the measurement data. 11. The device according to claim 10, wherein the sensing signal comprises one of the information selected from the group consisting of the following or combination thereof: blood sugar data, blood pressure data, and body weight data. 12. The device according to claim 10, wherein the memory unit is a swappable/swap device. 13. The device accordi... | - | A61B005/00
US07641639 | B2 | 2010-01-05 | Body fluid suction reservoir | 2003-11-26 | The present invention provides a body fluid suction reservoir including a flexible bag, two plates for supporting the bag, and a spring put between the two plates. The body fluid suction reservoir further includes a lock having a hook and a release part and consisting of a member different from the two plates, by which lock the spring is held in a state compressed between the two plates, and the held state is releasable. The lock has flexibility and is held by one of the plates so as to be deformable from a first state in which the hook is engaged with the other plate to a second state in which the hook is not engaged, and the lock is urged to take the first state. | 1. A body fluid suction reservoir comprising: a bag element having a fluid collecting port for taking in body fluid;a first plate and a second plate arranged in the bag element so as to face each other;an elastic member disposed between the first plate and the second plate; anda lock body attached to the second plate, whereinthe lock body comprises a hook part and an operation part for displacing the hook part,the hook part is adapted to engage with a hook engagement surface of the first plate to fix the first plate and the second plate, so that the elastic member is compressed and held between the first plate and the second plate,the operation part is operated to displace the hook part in parallel with a surface of the second plate, so that fixing of the first plate and the second plate are released to generate negative pressure in the bag elements,the first plate comprises a first shell part forming a part of an outline of the body fluid suction reservoir, and a core part provided on a side of the shell part facing the second plate,the second plate comprises a second shell part forming a part of the outline of the body fluid suction reservoir,at least a part of the lock body is located on a side of the second shell part facing the core part, so that the elastic member is compressed and held by the core part and at least the part of the lock body, andthe core part and the lock body comprise material having higher hardness than the first shell part and the second shell part. 2. The body fluid suction reservoir according to claim 1, characterized in that the lock body is elastically deformable, a part of the lock body is fixed to the second plate, thereby the operation part is operated so as to displace at least the hook part in parallel with the surface of the second plate, and the hook part returns to an initial position after the operation. 3. The body fluid suction reservoir according to claim 1, characterized in that at least one of the first plate and the second plate comprises a cutout portion, and the operation part is located in the cutout portion or in correspondence with the cutout portion, so as to be positioned inside an outline of the body fluid suction reservoir. 4. The body fluid suction reservoir according to claim 1, wherein the first plate comprises a first shell part forming a part of an outline of the body fluid suction reservoir, and a first core part provided on a side of the first shell part facing the second plate,the second plate comprises a second shell part forming a part of the outline of the body fluid suction reservoir, and a second core part provided on a side of the second shell part facing the first plate,the hook engagement surface is configured by a surface of the first core part located on an opposite side to the second plate, andthe elastic member is compressed and held by the first core part and the second core part. 5. The body fluid suction reservoir according to claim 1, wherein the first core part and... | Sumitomo Bakelite Company, Ltd. | A61M001/00 F01B019/00 A61B005/00
US08501159 | B2 | 2013-08-06 | Magnetic nanoparticles for the application in hyperthermia, preparation thereof and use in constructs having a pharmacological application | 2007-12-18 | There are described nanoparticles of magnetic metal oxides employable in constructs consisting in polymer particles possibly also incorporating pharmacologically active substances. | 1. A hyperthermically effective composition, the composition comprising: magnetic nanoparticles functionalised with ethyl-12-(hydroxyamino)-12-oxododecanoate, a polymer optionally containing a pharmacologically active molecule, and an external protecting layer of surface agents wherein said pharmacologically active molecule is chosen among: antitumour agents, antimicrobial agents, anti-inflammatory agents, immunomodulators, molecules acting on the central nervous system or those capable of marking the cells so as to allow their identification with the normal means of diagnostic detection. 2. The composition according to claim 1 wherein said pharmacological active molecule is connected to the polymer or dispersed in it. 3. The composition according to claim 1 wherein said magnetic nanoparticles are spinels and oxides of the MIIMIII2O4 type, in which MII is selected from the group consisting of Fe, Co, Ni, Zn, and Mn, and MIII is selected from the group consisting of: Fe and Cr a nanometric form. 4. The composition according to claim 3 wherein said magnetic nanoparticles are selected from the group consisting of cobalt ferrite, magnetite and maghemite. 5. The composition according to claim 1 wherein said polymer is a water insoluble polymer selected from the group consisting of polyesters, polyamides, polyanhydrides, polyorthoesters, peptides, polyamineamides, and insoluble organic molecules. 6. The composition according to claim 1 wherein said surface agents are selected from the group consisting of polyelectrolytes, polypeptides, water-soluble proteins, block copolymers, modified polyethylene glycols, modified polysaccharides, phospholipids, polyamineamides, globular proteins, human serum proteins, and pluronics block copolymers. 7. The composition according to claim 1 wherein the composition has an average diameter in the range between 50 and 300 nm. 8. The composition according to claim 7, wherein the magnetic nanoparticles comprise cobalt ferrite. 9. The composition according to claim 1 wherein the composition has an average diameter in the range between 30 and 100 nm. 10. The composition according to claim 1 wherein said magnetic nanoparticle comprises cobalt ferrite, and said polymer is a block polymer. 11. A hyperthermically effective composition, the composition comprising: magnetic nanoparticles functionalised with a bifunctional compound, and a polymer optionally containing a pharmacologically active molecule wherein the bifunctional compound is ethyl-12-(hydroxyamino)-12-oxododecanoate. 12. The composition according to claim 11 wherein said pharmacological active molecule is connected to the polymer or dispersed in it. 13. The composition according to claim 11 wherein said magnetic nanoparticles are spinels and oxides of the MIIMIII2O4 type, in which MII is selected from the group consisting of Fe, Co, Ni, Zn, and Mn, and MIII is selected from the group consisting of Fe and Cr a nanometric form. 14. The composition accordin... | Colorobbia Italia S.p.A. | A61B005/055
US08986217 | B2 | 2015-03-24 | Mapping catheter as well as mapping catheter apparatus and associated method | 2009-10-26 | The invention relates to a mapping catheter for determination of data of an area of an organ embodied as a flat surface, especially of the heart, to be presented graphically, with at least one thermosensor essentially aligned in the direction of the longitudinal axis of the mapping catheter for determination of temperature-related data which is arranged at a tip of the mapping catheter being provided in the distal area of the mapping catheter for introduction into the organ. | 1. A mapping catheter apparatus for acquiring data along an area of an organ of a patient, comprising: a catheter tube comprising a distal area having a tip configured for introduction into and positioning along the organ;a thermosensor array formed in a grid structure adapted to be folded within the catheter tube during introduction and positioning along the organ and adapted to be unfolded outward and pushed forward of the tip to achieve contact with tissues of the organ on reaching a position along the organ intended for measurement, wherein the grid structure allows the passage of blood therethrough when positioned along the organ and comprises a net or a mesh that unfolds into a three-dimensional structure to contact a surface of the organ;wherein thermosensors in the array are distributed along a plurality of points of the grid structure across the surface area and not solely along a circumference thereof, including distributed in a direction pointing forward of the tip of the catheter tube when the grid structure is unfolded therefrom, to achieve contact with tissues simultaneously at a number of points in the area of the organ,wherein the thermosensors are arranged on the grid structure at a plurality of different radial distances from a given point along a longitudinal axis extending through the catheter tube and extending forward of the tip upon unfolding of the grid structure,wherein the thermosensors further provide temperature related, image forming data corresponding to a thermal gradient detected simultaneously across the area of the organ of the patient where the grid structure achieves contact with tissues;a control device, coupled to receive the image forming data from the thermosensors of the mapping catheter, configured to generate an image therefrom for display of a graphical presentation of thermal information over the area of the organ; andan image output device coupled to the control device via a data connection to display the graphical presentation as a multi-dimensional thermomap as a function of position data along the area of the organ. 2. The mapping catheter apparatus as claimed in claim 1, further comprising one or more devices in an area of at least one thermosensor for cooling down or heating up a part of the area of the organ. 3. The mapping catheter apparatus as claimed in claim 1, further comprising a sensor for determining data selected from the group consisting of: electrophysiological data of the patient, physiological data of the patient, biometric data of the patient, and position data of the mapping catheter. 4. The mapping catheter apparatus as claimed in claim 1, wherein the control device processes the temperature related data in real time. 5. The mapping catheter apparatus as claimed in claim 1, wherein the control device detects a position or an orientation of the mapping catheter based on a position sensor system of the mapping catheter. 6. The mapping catheter apparatus as claimed in claim 1, ... | Siemens Aktiengesellschaft | A61B005/01 G01K013/00 A61B005/042 A61B005/06 A61B005/00 A61B018/14 A61B018/00
US08311621 | B2 | 2012-11-13 | Device for judging degree of awakening and method for judging degree of awakening | 2007-12-04 | A device for judging a degree of awakening and a method for judging a degree of awakening which can more reliably detect weak sleepiness of people in action. The device for judging a degree of awakening in accordance with the present invention comprises a heartbeat sensor for acquiring a heartbeat signal from a driver and an ECU for detecting sleepiness of the driver by processing the heartbeat signal. In the ECU, a heartbeat signal preprocessing section for acquiring a heartbeat period time series from the heartbeat signal, a feature amount extracting section for acquiring a heartbeat fluctuation low frequency component power from the heartbeat period time series, a sleepiness detecting section for judging whether sleepiness occurs in the driver or not according the heartbeat fluctuation low frequency component power, a stimulus timing setting section for setting a timing for imparting a stimulus for removing sleepiness to the driver according to the heartbeat fluctuation low frequency component power, and a sleepiness removing stimulus output section for imparting the stimulus to the driver at thus set timing. | 1. A device for judging a degree of awakening, the device comprising: signal acquiring means for acquiring a biological signal of a person in action;index acquiring means for acquiring a physiological index indicating a strength of a state of acting against sleepiness from the biological signal;judging means for judging the degree of awakening of the person according to the physiological index;stimulus providing means for imparting a stimulus for raising the degree of awakening of the person; andtiming setting means for setting a timing for providing the stimulus before a predetermined time passes after the physiological index exceeds a predetermined value, wherein the biological signal is a heartbeat signal, andthe physiological index comprises an amplitude spectral power of a heartbeat fluctuation low frequency component acquired from the heartbeat signal. 2. The device for judging a degree of awakening according to claim 1, wherein the physiological index is correlated with a sympathetic activity. 3. The device for judging a degree of awakening according to claim 1, wherein the judging means judges that the degree of awakening is lowered when the amplitude spectral power of the heartbeat fluctuation low frequency component is greater than a first predetermined value. 4. The device for judging a degree of awakening according to claim 3, wherein the judging means judges that the degree of awakening is further lowered when the amplitude spectral power of the heartbeat fluctuation low frequency component is greater than both the first predetermined value and a second predetermined value. 5. The device for judging a degree of awakening according to claim 1, wherein the timing setting means provides the stimulus according to the physiological index indicating the strength of the state of acting against the sleepiness. 6. The device for judging a degree of awakening according to claim 1, wherein the physiological index corresponds to a time between a time when a heartbeat fluctuation low frequency component power exceeds a predetermined value and a time when the heartbeat fluctuation low frequency component power becomes lower than the predetermined value or at a time when the heartbeat fluctuation low frequency component power attains a local minimum. 7. A device for judging a degree of awakening, the device comprising: signal acquiring means for acquiring a biological signal of a person in action;index acquiring means for acquiring a physiological index indicating a strength of a state of acting against sleepiness from the biological signal;judging means for judging the degree of awakening of the person according to the physiological index;stimulus providing means for imparting a stimulus for raising the degree of awakening of the person; andtiming setting means for setting a timing for providing the stimulus before the amplitude spectral power of the heartbeat fluctuation low frequency component attains the nearest local minimum after exceed... | Toyota Jidosha Kabushiki Kaisha | A61B005/0402
US08419653 | B2 | 2013-04-16 | Spinal access and neural localization | 2009-07-16 | A method for locating neural tissue in a patient body may involve: advancing a probe along a natural tissue interface between the neural tissue and another tissue in the body, the probe having a first surface oriented toward the neural tissue and a second surface oriented away from the neural tissue; delivering a first electrical current to a first electrode along the first surface of the probe; delivering a second electrical current to a second electrode along the second surface of the probe; and verifying that the first surface of the advanced probe remains oriented toward the neural tissue and the second surface remains oriented away from the neural tissue by monitoring neural response to the first and second electrical currents. | 1. A device for detecting if neural tissue is above or below a given region of a device, the device comprising: a probe having a proximal rigid elongate shaft and a distal end, the distal end being flexible such that it may be articulated to form more than one curved shape and the distal end being substantially flat and having an axial length, a width, and a thickness, wherein the axial length is greater than the width and the width is greater than the thickness, the distal end further having a first upper and second lower laterally oriented surfaces, wherein the upper surface is opposite and separated by the thickness from the lower surface;a first signal transmitter configured to deliver a neural localization signal disposed along the first upper surface of the probe;a second signal transmitter configured to deliver a neural localization signal disposed along the second lower surface of the probe,wherein by monitoring a response to the neural localization signals from the first upper surface and the second lower surface, the orientation of the neural tissue relative to the upper surface and the lower surface of the probe can be safely verified. 2. The device of claim 1, wherein the first and second signal transmitters comprise first and second electrodes. 3. The device of claim 1, wherein the distal portion of the probe is configured to facilitate passage of the distal portion of the probe into an intervertebral foramen of a spine. 4. The device of claim 1, further comprising: a wire for pushing or pulling to position the distal end of the probe; and at least one tissue removal device for advancing along or being pulled into position by the wire to remove tissue from the body. 5. The device of claim 1, further comprising at least one sizing probe configured to perform at least one of the functions of assessing an amount of space between tissues before advancing the probe, increasing the amount of space between tissue before advancing the probe, and assessing the amount of space between tissues after performing a tissue modification procedure. 6. A device for determining if neural tissue is above or below a given region of a device in a human body, the device comprising: a elongate probe having a proximal rigid elongate shaft portion and a distal portion, the distal portion being flexible such that it may be articulated to form more than one curved shape and the distal portion being substantially flat and having an axial length, a width, and a thickness, wherein the axial length is greater than the width and the width is greater than the thickness, the distal portion further having a first upper surface separated by the thickness from an opposite second lower surface;a first electrode coupled with the first surface of the probe at or near the distal portion and connectible with at least one energy source configured to deliver increasing levels of energy to the first upper surface to determine a first stimulation level at which the neural ti... | Baxano, Inc. | A61B005/05 A61B005/04 A61N001/00
US08150531 | B2 | 2012-04-03 | Associating therapy adjustments with patient posture states | 2009-04-30 | A therapy adjustment received from a patient is associated with a sensed patient posture state. Thereafter, a stability indication can be associated with the sensed posture state to indicate that a therapy adjustment specific to the sensed posture state was received and implemented. In addition, therapy parameter values associated with related patient posture states are updated based on the therapy adjustment. In some examples, only the related posture states that are not associated with respective stability indications are updated based on the therapy adjustment. The absence of a stability indication indicates that the patient has not inputted a therapy adjustment specific to the respective related posture state. As a result, the therapy parameter values associated with the related posture states that are not associated with stability indications may float until the patient adjusts the therapy parameter value associated with the respective related posture state. | 1. A method comprising: receiving user input indicating an adjustment to a first therapy parameter value that defines therapy delivered to a patient by a medical device;associating the adjustment with a sensed posture state of the patient;associating the sensed posture state with a stability indication that indicates the user has inputted the adjustment specific to the sensed posture state, and storing the stability indication within a memory; andbased on the adjustment, automatically updating a second therapy parameter value associated with at least one patient posture state related to the sensed posture state and not associated with a respective stability indication. 2. The method of claim 1, wherein an absence of an association between the at least one patient posture state related to the sensed postured state and a respective stability indication indicates the user has not inputted an adjustment specific to the at least one patient posture state related to the sensed posture state. 3. The method of claim 1, further comprising: starting a posture search timer having a search period upon receiving the user input; andstarting a posture stability timer having a stability period, wherein associating the adjustment with the sensed posture state occurs only when the sensed posture state begins within the search period and does not change during the stability period. 4. The method of claim 3, wherein the search period is between about thirty seconds and about thirty minutes and the stability period is between about thirty seconds and about thirty minutes. 5. The method of claim 3, wherein the search period is about three minutes and the stability period is about three minutes. 6. The method of claim 1, wherein the adjustment to the first therapy parameter value indicates a net change to the first therapy parameter value, and updating the second therapy parameter value comprises adjusting the second therapy parameter value based on the net change. 7. The method of claim 6, wherein updating the second therapy parameter value comprises adjusting the second therapy parameter value based on the net change comprises updating the second therapy parameter value to maintain a predetermined ratiometric balance between the first and second therapy parameter values. 8. The method of claim 1, wherein updating the second therapy parameter value comprises adjusting the second therapy parameter value to be substantially equal to the first therapy parameter value. 9. The method of claim 1, wherein updating the second therapy parameter value comprises: determining a lowest, highest or average therapy parameter value of a group of therapy parameter values associated with the related posture states and associated with stability indications in a memory; andupdating the second therapy parameter value based on the lowest, highest or average therapy parameter value. 10. The method of claim 9, wherein the related posture states comprise related lying posture states ... | Medtronic, Inc. | A61N001/08
US08126549 | B2 | 2012-02-28 | Cardiac protection system and method | 2008-07-15 | A cardiac ischemic protection system and method for conditioning a patient's heart is provided. The method can include detecting acute myocardial infarction, angina pectoris, silent ischemia, or stunning and providing closed-loop dyssynchronous pacing to the patient's heart to precondition and/or postcondition the patient's heart in order to reduce ischemic damage. | 1. A method of providing cardiac ischemic protection to a patient's heart, the method comprising: detecting at least one of acute myocardial infarction, angina pectoris, silent ischemia, and stunning;providing closed-loop dyssynchronous pacing to the patient's heart to at least one of precondition and postcondition the patient's heart in order to reduce ischemic damage,wherein the dyssynchronous pacing is configured to produce discordant motion of the ventricular walls of the patient's heart; andpacing the right ventricular apex with an atrioventricular delay with a dual chamber ventricular pacemaker,wherein the atrioventricular delay is about 30 milliseconds. 2. The method of claim 1 and further comprising alerting the patient to seek treatment after detecting at least one of acute myocardial infarction, angina pectoris, silent ischemia, and stunning. 3. The method of claim 1 and further comprising alerting medical personnel of the patient's condition after detecting at least one of acute myocardial infarction, angina pectoris, silent ischemia, and stunning. 4. The method of claim 1 and further comprising detecting acute myocardial infarction by detecting ST segment changes. 5. The method of claim 1 and further comprising providing pacing to precondition the patient's heart by providing one to ten periods of about five minutes of pacing and intermittent periods of sinus rhythm of about five minutes. 6. The method of claim 1 and further comprising providing continuous pacing to the patient's heart. 7. The method of claim 1 and further comprising manually activating the dyssynchronous pacing. 8. A cardiac ischemic protection system that conditions a patient's heart, the system comprising: a detector that detects at least one of acute myocardial infarction, angina pectoris, silent ischemia, and stunning; anda pulse generator that provides closed-loop dyssynchronous pacing to the patient's heart to at least one of precondition and postcondition the patient's heart in order to reduce ischemic damage,wherein the dyssynchronous pacing is configured to produce discordant motion of the ventricular walls of the patient's heart,wherein the pulse generator provides pacing to the right ventricular apex with an atrioventricular delay with a dual chamber ventricular pacemaker, andwherein the atrioventricular delay is about 30 milliseconds. 9. The system of claim 8 and further comprising an implantable pacing device that includes the detector. 10. The system of claim 8 and further comprising a patient alerting device that alerts the patient to seek treatment after the detector detects at least one of acute myocardial infarction, angina pectoris, silent ischemia, and stunning. 11. The system of claim 10 wherein the patient alerting device alerts medical personnel of the patient's condition after detecting at least one of acute myocardial infarction, angina pectoris, silent ischemia, and stunning. 12. The system of claim 8 wherein the detector detects ... | Medtronic, Inc. | A61N001/362
US09066663 | B2 | 2015-06-30 | Heart monitor | 2007-05-22 | A heart monitor for use as a non-invasive screening tool for identifying potential atrial fibrillation in patients comprises a sensor for producing an output waveform of the patient's actual sinus rhythm, a processing unit arranged to store the normalized waveform of an ideal sinus rhythm and to compare the actual and ideal sinus rhythm waveforms and to produce an output dependant on the difference on a display. The value of the output is indicative of whether the patient is atrial fibrillation or other cardiac arrhythmia. | 1. A monitor for monitoring the heart of a patient, the monitor comprising a sensor for producing a first waveform of the patient's actual heart rhythm, a processing unit that is configured to generate first data pertaining to the first waveform of the patient's actual heart rhythm and to store second data pertaining to a second waveform of a simulated ideal heart rhythm that is simulated using an idealized model, said second data representing a peak mathematical function which is independent of fundamental frequency and which comprises a series of decaying pointed peaks of fixed width separated by a frequency, f, which is equal to the frequency of the peak of greatest magnitude, wherein the processing unit processes the first data as well as the second data for the first and second waveforms and produces an output dependent on differences between the first and second waveforms. 2. A monitor as claimed in claim 1, in which the sensor is arranged to produce the first waveform of the patient's actual heart rhythm by sensing the electrical activity of the heart of the patient to produce an electrocardiogram. 3. A monitor as claimed in claim 1, in which the sensor is arranged to produce the first waveform of the patient's actual heart rhythm by sensing the pulsatile flow of blood through a part of the patient's body to produce a plethysmogram. 4. A monitor as claimed in claim 3, in which the sensor comprises a light emitter and a light detector arranged to detect a change in the transmissivity of said part of the patient's body as blood flows therethrough in a pulsatile manner. 5. A monitor as claimed in claim 4, in which the light emitter is arranged to emit light having a wavelength of 500 to 1100 nm. 6. A monitor as claimed in claim 1, in which said processing unit is arranged to compare frequency spectrums of said first and second waveforms. 7. A monitor as claimed in claim 6, in which the second waveform has the same or similar fundamental frequency as the first waveform. 8. A monitor as claimed in claim 6, in which the processing unit is operable to determine the fundamental frequency of the waveform of the patient's actual heart rhythm and to compare the frequency spectrum of the first waveform with the frequency spectrum of the idealized waveform having the same or similar fundamental frequency. 9. A monitor as claimed in claim 1, in which said processing unit is arranged to store a model of the frequency spectrum of said second waveform. 10. A monitor as claimed in claim 1, in which the peak mathematical function is area normalized, such that the area beneath and between frequency limits of interest is 1. 11. A monitor as claimed in claim 1, in which the processing unit is arranged to produce a patient mathematical function representative of the frequency spectrum of said first waveform. 12. A monitor as claimed in claim 11, in which the patient mathematical function is produced using a Fourier transformation of said first wavefor... | MELYS AFS LIMITED | A61B005/024
US08764667 | B2 | 2014-07-01 | Method and system for monitoring sleep | 2008-03-06 | A method of monitoring sleep of a sleeping subject is disclosed. The method comprises determining a phase shift of input radiofrequency signals received from the subject during sleep relative to output radiofrequency signals transmitted to the subject during sleep, calculating cardiac output based on the phase shift, and using the cardiac output for identifying sleep apnea events. | 1. A method of monitoring sleep of a sleeping subject using output radiofrequency signals transmitted to the subject during sleep and input radiofrequency signals received from the subject during sleep, the method comprising: determining a phase shift of the input radiofrequency signals relative to the output radiofrequency signals to provide processed signals;applying to said processed signals a dynamically variable filter characterized by a frequency band which is dynamically adapted in response to a change in a physiological condition of the subject;using a cardiac output calculator, calculating cardiac output based on said filtered processed signals; andusing said cardiac output for identifying sleep apnea events. 2. The method of claim 1, further comprising supplementing said cardiac output with blood oxygen content, wherein said identification of said sleep apnea events is based on said cardiac output and said blood oxygen content. 3. The method of claim 1, further comprising inputting or estimating hemoglobin concentration of the sleeping subject, wherein said identification of said sleep apnea events is based on said cardiac output and said hemoglobin concentration. 4. The method of claim 3, further comprising estimating total oxygen delivery and generating a wakening signal when an estimated value is below a predetermined threshold. 5. The method of claim 4, wherein said identification of said sleep apnea events is further based on said estimated value of said total oxygen delivery. 6. The method of claim 4, further comprising controlling positive airway pressure delivered to the sleeping subject based on said estimated value. 7. The method of claim 1, further comprising reducing or eliminating amplitude modulation of said input radiofrequency signals so as to provide input radiofrequency signals of substantially constant envelope. 8. The method of claim 1, further comprising mixing said output radiofrequency signals and said input radiofrequency signals so as to provide a mixed radiofrequency signal, and filtering out a portion of said mixed radiofrequency signal so as to substantially increase a signal-to-noise ratio of a remaining portion of said mixed radiofrequency signal. 9. The method of claim 1, wherein said cardiac output is calculated using a linear relationship between said filtered processed signals and said cardiac output. 10. The method of claim 1, wherein said physiological condition is a heart rate of the subject. 11. The method of claim 10, wherein a lower frequency bound characterizing said filter is about 0.9*(HR/60) Hz at all times, wherein said HR is said heart rate in units of beats per minute. 12. The method of claim 10, wherein an upper frequency bound characterizing said filter is about 6+1.5*[(HR/60)−1] Hz at all times, wherein said HR is said heart rate in units of beats per minute. 13. A method of monitoring sleep of a sleeping subject, comprising: transmitting output radiofrequency signals to the ... | Cheetah Medical, Inc. | A61B005/02
US20100022868 | A1 | 2010-01-28 | ACTIVE MRI COMPATIBLE AND VISIBLE iMRI CATHETER | 2006-09-13 | A catheter for use with magnetic resonance imaging including a catheter body, a radio frequency antenna disposed in a distal end of the catheter body, a first ultrasonic transducer disposed in the distal end of the catheter body and being electrically connected to the radio frequency antenna, an acoustic waveguide disposed in the catheter body and extending from the distal end to a proximal end of the catheter body, the acoustic waveguide being acoustically matched to the first ultrasonic transducer. The catheter also has a second ultrasonic transducer disposed in the proximal end of the catheter body. | 1. A catheter for use with magnetic resonance imaging, comprising: a catheter body;a radio frequency antenna disposed in a distal end of said catheter body;a first ultrasonic transducer disposed in said distal end of said catheter body and being electrically connected to said radio frequency antenna;an acoustic waveguide disposed in said catheter body and extending from said distal end to a proximal end of said catheter body, said acoustic waveguide being acoustically matched to said first ultrasonic transducer; anda second ultrasonic transducer disposed in said proximal end of said catheter body. 2. A catheter for use with magnetic resonance imaging according to claim 1, wherein radio frequency signals received by said radio frequency antenna are converted to corresponding acoustic signals by said first ultrasonic transducer and transmitted acoustically along said acoustic waveguide to be converted to electrical signals by said second ultrasonic transducer. 3. A catheter for use with magnetic resonance imaging according to claim 1, wherein electrical signals input to said second ultrasonic transducer are converted to acoustic signals that are transmitted along said acoustic waveguide to said first ultrasonic transducer to be converted to electrical signals which are then converted to radio frequency illumination signals by said radio frequency antenna. 4. A catheter for use with magnetic resonance imaging according to claim 1, wherein said acoustic waveguide is an optical fiber. 5. A catheter for use with magnetic resonance imaging according to claim 4, wherein said optical fiber has a cladding-to-core ratio of about 3:1. 6. A catheter for use with magnetic resonance imaging according to claim 4, wherein said optical fiber has a fused silica core. 7. A catheter for use with magnetic resonance imaging according to claim 4, wherein said optical fiber has a core doped with titanium oxide. 8. A catheter for use with magnetic resonance imaging according to claim 7, wherein said core is doped with about 7.5% titanium oxide. 9. A catheter for use with magnetic resonance imaging according to claim 4, wherein said optical fiber has a core doped with B2O3. 10. A catheter for use with magnetic resonance imaging according to claim 9, wherein said core of said optical fiber is doped with about 5% B2O3. 11. A catheter for use with magnetic resonance imaging according to claim 4, wherein said optical fiber has an outer diameter of about 600 μm. 12. A catheter for use with magnetic resonance imaging according to claim 1, wherein said radio frequency antenna is a coil antenna. 13. A catheter for use with magnetic resonance imaging according to claim 12, wherein said coil antenna is made from at least one of platinum-iridium, insulated copper and insulated gold. 14. A device for use with magnetic resonance imaging, comprising: an insertion component adapted to be inserted into at least one of a body cavity and an internal region of a patient;a radio ... | THE GOVERNMENT OF THE U.S.A., AS REPRESENTED BY THE SECRETAR, DEPT. OF HHS | A61B005/055 A61B008/14
US20100016736 | A1 | 2010-01-21 | Estimating Aortic Blood Pressure from Non-Invasive Extremity Blood Pressure | 2009-07-13 | Methods and a computer program product for using a circulatory measurement on an extremity of a particular subject to derive an aortic blood pressure for that subject. A model is constructed that maps a peripheral cardiovascular waveform to a central cardiovascular waveform on the basis of a plurality of model parameters. A time record is obtained using a non-invasive blood pressure sensor disposed at a solitary position periphery of the cardiovascular system of a subject. The time record is then transformed to obtain a plurality of test central blood pressure waves, with a single test central blood pressure wave is based on each of a plurality of sets of values of the model parameters. An optimum set of values of the model parameters is then selected, based on a specified criterion applied to the plurality of test central blood pressure waves, so that the aortic circulatory waveform of the subject can be obtained. | 1. A method for deriving an aortic circulatory waveform of a particular subject, the method comprising: constructing a model that maps a peripheral cardiovascular waveform (CW) to a central cardiovascular waveform on the basis of a plurality of model parameters;acquiring a time record of the peripheral CW with a non-invasive blood pressure sensor disposed at a solitary peripheral point;transforming the time record of the peripheral CW to obtain a plurality of test central blood pressure waves, one test central blood pressure wave based on each of a plurality of sets of values of the model parameters;electing an optimum set of values of the model parameters based on a specified criterion applied to the plurality of test central blood pressure waves; andobtaining the aortic circulatory waveform of the subject based on the elected set of values of the model parameters. 2. The method of claim 1, wherein the elected set of values of the model parameters is specific to the particular subject. 3. The method of claim 1, wherein each of the plurality of sets of values corresponds to successive putative transit times between the central CW and the peripheral CW. 4. The method of claim 3, wherein electing an optimum set of values of the model parameters includes electing an optimum system order corresponding to an optimum transit time. 5. The method of claim 3, wherein each putative transit time between the central CW and the peripheral CW corresponds to an order of a generalized auto-regressive moving average (ARMA) model. 6. The method of claim 1, wherein the specified criterion is a criterion applied to the plurality of test central blood pressure waves in the time domain. 7. The method of claim 6, wherein the specified criterion is a minimum norm of a second time derivative of the test central blood pressure wave. 8. The method of claim 1, wherein the model is an affine model. 9. A computer program product for use on a computer system for establishing an aortic circulatory waveform of a subject, the computer program product comprising a computer usable medium having computer readable program code thereon, the computer readable program code including: memory for storing a model that maps a peripheral cardiovascular waveform (CW) at a peripheral point to a central cardiovascular waveform on the basis of a plurality of model parameters;an input for receiving a time record of a peripheral CW from a non-invasive blood pressure sensor;a module for calculating a succession of sets of values for each of the plurality of model parameters;computer code for transforming the time record of the peripheral CW to obtain a plurality of test central blood pressure waves, one test central blood pressure wave based on each of the set of model parameters;a software module for selecting an optimum system order corresponding to an optimum transit time based on a specified criterion applied to the plurality of test central blood pressure waves; andan output for displa... | MASSACHUSETTS INSTITUTE OF TECHNOLOGY | A61B005/021
US07660618 | B2 | 2010-02-09 | Magnetic resonance imaging device | 2003-06-05 | An MRI apparatus having a pulse sequence of dynamic MRA performs a pulse sequence for monitoring arrival of a contrast agent at a blood vessel of interest. The pulse sequence is substantially the same as the imaging sequence except that the phase-encode number and/or the slice-encode number is small. A time-series image reconstructed by using data measured under the condition of applying gradient magnetic fields with a low spatial resolution is displayed to ascertain that the contrast agent has arrived at the blood vessel of interest. When the contrast agent arrives at the blood vessel of interest, the phase-encode and slice-encode are added to continue the substantial measurement pulse sequence without changing the pulse sequence itself. The first image of the substantial measurement is produced by using the data acquired at the time of arrival of the contrast agent in monitoring. Thus, the arrival time of the contrast agent can be reliably monitored with high time-resolution and the substantial measurement can be performed with most suitable timing to obtain an excellent blood vessel image. | 1. A magnetic resonance imaging apparatus comprising an imaging means for applying high-frequency magnetic fields and gradient magnetic fields to an object in a static magnetic field, in accordance with a pulse sequence for dynamic measurement for continuously obtaining a plurality of time-series images, and for measuring NMR signals emitted from the object, a signal processing means for forming images of a desired tissue of the object, from the NMR signals, a display means for displaying the images and a control means for controlling the imaging means and the signal processing means; wherein said imaging means is provided with a two-dimensional monitoring mode in which a desired slab of the object is measured using a pulse sequence for the dynamic measurement under a condition of applying gradient magnetic fields with a low spatial resolution, and is provided with a three-dimensional measurement mode in which the same slab is measured using the same pulse sequence under a condition of applying gradient magnetic fields with a high spatial resolution, andsaid control means includes mode switching means for switching from the two-dimensional monitoring mode to the three-dimensional measurement mode, andsaid switching means switches the two-dimensional monitoring mode to the three-dimensional measurement mode with desired timing when the two-dimensional monitoring mode is performed, andwherein the dynamic measurement performed by said imaging means is blood imaging for observing a change of blood flow using a contrast agent, where slice encode is omitted in the two-dimensional monitoring mode to obtain two-dimensional images and the slice encode is added in the three-dimensional measurement mode to obtain three-dimensional images,wherein said control means controls said signal processing means to reconstruct images immediately after the three-dimensional measurement mode begins, using data including data acquired in the pulse sequence performed without the slice encode just before the three-dimensional measurement. 2. The magnetic resonance imaging apparatus of claim 1, wherein data of the time-series images is three-dimensional data and is transformed to a two-dimensional projected image to be displayed on said display means. 3. The magnetic resonance imaging apparatus of claim 1, wherein said mode switching means has an input means for mode switching, and the two-dimensional monitoring mode is switched to the three-dimensional measurement mode by directly inputting a switching instruction to said mode switching means. 4. The magnetic resonance imaging apparatus of claim 1, wherein the control means sends an instruction to the mode switching means to indicate starting time of the three-dimensional measurement mode based on signal intensity of data acquired when the pulse sequence is applied without slice encode. 5. The magnetic resonance imaging apparatus of claim 1, wherein the apparatus is provided with means for extracting reference data fro... | Hitachi Medical Corporation | A61B005/055
US07655119 | B2 | 2010-02-02 | Meter for use in an improved method of reducing interferences in an electrochemical sensor using two different applied potentials | 2004-10-29 | The present invention is directed to an improved meter that utilizes a method of reducing the effects of interfering compounds in the measurement of analytes and more particularly to a method of reducing the effects of interfering compounds in a system wherein the test strip utilizes two or more working electrodes. In one embodiment of the present invention, a meter is described which applies a first potential to a first working electrode and a second potential, having the same polarity but a greater magnitude than the first potential, is applied to a second working electrode. The meter then measures the generated current and utilizes a predetermined algorithm to correct the measured current to compensate for the presence of interfering compounds in the sample. | 1. A meter for use with a test strip in the detection of analytes, said meter comprising: a first connector for connecting said meter to a first working electrode on a test strip;a second connector for connecting said meter to a second working electrode on said test strip;a common connector for connecting said meter to a reference electrode on said test strip;a first voltage source connected between said first connector and said common connector; anda second voltage source connected between said second connector and said common connector; wherein said first and second voltage sources generate the first potential at said first connector and the second potential at said second connector when said test strip is inserted into said meter and a sample is applied to said test strip, wherein said first and second voltages have the same polarity;wherein said meter measures a first current value at said first connector and a second current value at said second connector at a predetermined time after said test strip is inserted and sample applied;wherein the meter is configured to generate a preselected first potential at the first connector using the first voltage source, the first potential sufficient to ensure a magnitude of current at the first working electrode that is diffusion limited and insensitive to applied potential;wherein the meter is also configured to generate a preselected second potential at the second connector using the second voltage source, the magnitude of the second potential being greater than the magnitude of the first potential; and wherein the meter is configured to calculate a corrected value of analyte current using the following formula:A1=W2-YW1X-Ywhere: A1 is said corrected analyte current; W1 is the current at said first connector measured at said predetermined time; W2 is the current measured at said second connector at said predetermined time; X is an analyte dependent voltage effect factor; and Y is an interfering compound dependent voltage effect factor. | Lifescan Scotland Limited | A61B005/0496 C12Q001/54 G01N027/26
US07654970 | B2 | 2010-02-02 | Cervical dilation measurement apparatus | 2008-01-16 | The present invention provides a medical device for measuring cervical dilation, where the medical device is positionable about a hand having first and second fingers, with each finger having a tip and a side surface. The medical device may include a housing, a first extension element movably coupled to the housing, a second extension element movably coupled to the housing, and a dilation indication mechanism to measure a distance between the first and second extension elements. The medical device may also include a first lateral pressure sensor positionable about a side surface of the first finger, a second lateral pressure sensor positionable about a side surface of the second finger, a third pressure sensor positionable about a tip of the first finger, and a fourth pressure sensor positionable about a tip of the second finger. | 1. A medical device for measuring cervical dilation comprising: a substantially spherical base housing;a first elongated rod movably coupled to the base housing;a second elongated rod movably coupled to the base housing, wherein each of the first and second elongate rods are movable about the base housing in at least two planes of motion;a dilation indication mechanism coupled to at least one of the base housing, first elongate rod, and second elongate rod to provide a measurement of a distance between the first and second elongated rods; anda first pressure measurement sensor positionable about a side surface of a finger;a second pressure measurement sensor positionable about a side surface of a finger;a third pressure measurement sensor positionable about a tip surface of a finger;a fourth pressure measurement sensor positionable about a tip surface of a finger; and wherein the substantially spherical base housing, first elongate rod, second elongate rod, first pressure sensor, second pressure sensor, third pressure sensor, and fourth pressure sensor are disposed within a sterile lubricious sheath. 2. The medical device according to claim 1, wherein the dilation indication mechanism includes a processor in electrical communication with at least one of the base housing, first elongated rod, and second elongated rod. 3. The medical device according to claim 2, wherein the processor is in electrical communication with at least one of the pressure measurement sensors. | Intrapartum, LLC | A61B005/103 A61B005/117
US07641618 | B2 | 2010-01-05 | Capacitance-type pressure sensor and heart beat / respiration measuring device using the same | 2005-03-18 | A heart beat/respiration measuring device comprising a sheet-like capacitance-type pressure sensor adapted to be pressed against the human body, and a measuring circuit for measuring a heart rate and/or respiration rate from the output of the sensor. The capacitance-type pressure sensor includes a sheet-like dielectric body elastically deformable in all directions and a pair of conductive clothes with stretchability disposed on opposite sides of the dielectric body. The measuring circuit comprises a resonant circuit wherein the capacitance-type pressure sensor serves as an oscillation capacitor, and a calculation processing circuit for detecting variations in the oscillation frequency of the resonant circuit and calculating the heart rate and/or respiration rate based on the frequency component or components of heart beats and/or respiration included in the variations. | 1. A capacitance-type pressure sensor comprising two sheet-like dielectric bodies elastically deformable in all directions superposed on each other with one conductive cloth interposed therebetween, and two conductive clothes disposed on opposite sides of the two sheet-like dielectric bodies stacked together, the two conductive clothes being electrically connected to each other, wherein a plurality of through holes or projections and depressions are formed on each of the two sheet-like dielectric bodies, and the through holes or projections and depressions provided on one sheet-like dielectric body and the through holes or projections and depressions provided on the other sheet-like dielectric body are staggered. 2. A heart beat/respiration measuring device comprising: a sheet-like capacitance type pressure sensor adapted to be pressed against a human body, anda measuring circuit for measuring a heart rate and/or respiration rate from the output of the sensor, the capacitance-type pressure sensor comprising:two sheet-like dielectric body elastically deformable in all directions superposed on each other with one conductive cloth interposed therebetween, andtwo conductive clothes with stretchability disposed on opposite sides of the two dielectric body stacked together, the two conductive clothes being electrically connected to each other, the measuring circuit comprising:a resonant circuit wherein the capacitance-type pressure sensor serves as an oscillation capacitor, anda calculation processing circuit for detecting variations in the oscillation frequency of the resonant circuit and calculating the heart rate and/or respiration rate based on the frequency component or components of the heart beats and/or respiration included in the variations, wherein,a plurality of through holes or projections and depressions are formed on each of the two sheet-like dielectric bodies, and the through holes or projections and depressions provided on one sheet-like dielectric body and the through holes or projections and depressions provided on the other sheet-like dielectric body are staggered. 3. A capacitance-type pressure sensor comprising a first conductive cloth with stretchability and flexibility, a second conductive cloth with stretchability and flexibility opposed to the first conductive cloth, and a sheet-like dielectric body with stretchability and flexibility provided on the side of the first conductive cloth opposed to the second conductive cloth, wherein the sensor has space between the second conductive cloth and the sheet-like dielectric body, and a location and size of the space varies a capacitance of the capacitance-type pressure sensor. 4. The heart beat/respiration measuring device comprising: the capacitance-type pressure sensor according to claim 3, adapted to be pressed against a human body, anda measuring circuit for measuring a heart rate and/or respiration rate from the output of the sensor, wherein the measuring circuit comprises:a ... | Sanyo Electric Co., Ltd. | A61B005/08
US20100010336 | A1 | 2010-01-14 | METHOD AND SYSTEM FOR DIAGNOSIS OF NEUROPSYCHIATRIC DISORDERS INCLUDING ATTENTION DEFICIT HYPERACTIVITY DISORDER (ADHD), AUTISM, AND SCHIZOPHRENIA | 2009-07-23 | A method and system for medical imaging of neuropsychiatric disorders including attention deficit hyperactivity disorder (ADHD), autism, and schizophrenia. Noninvasive, in vivo methods identify novel brain molecular biomarkers of normal neurodevelopment in order to determine molecular underpinnings of abnormal neurodevelopment. The described brain molecular biomarkers will aid in the presymptomatic diagnosis of neuropsychiatric disorders which begin in childhood and adolescence, such as ADHD, autism, and schizophrenia. | 1. A method for diagnosing attention deficit hyperactivity disorder (ADHD), autism or schizophrenia disease in a human, comprising: imaging a brain of a human with a medical imaging process, wherein the medical imaging process is a multinuclear magnetic resonance process based on comparison of intensities obtained from relevant 31P or 1H nuclear magnetic resonance signals obtained from normal human child hosts and known abnormal human child hosts with cognitive impairment due to attention deficit hyperactivity disorder (ADHD), autism or schizophrenia with signal information obtained from a new human child patient with an unknown cognitive condition;determining a first signal intensity for synaptic phosphomonoesters (sPME) of the human brain;determining a second signal intensity for synaptic phosphodiesters (sPDE) of the human brain;determining a third signal intensity for synaptic vesicles phosphodiesters (iPDE) of the human brain;determining a fourth signal intensity for phosphocreatine (PCr) of the human brain;determining a fifth signal intensity for myo-inositol of the human brain; anddetermining whether a conclusion of cognitively impaired cognitive impairment due to attention deficit hyperactivity disorder (ADHD), autism or schizophrenia is suggested using the determined first, second, third, fourth and fifth signal intensities. 2. The method of claim 1 wherein an increased sPME level, a decreased sPDE level, a decreased iPDE level, a decreased PCr level and a decreased myo-inositol level antedate abnormal brain structural changes for determining attention deficit hyperactivity disorder (ADHD), autism or schizophrenia. 3. A method for diagnosing attention deficit hyperactivity disorder (ADHD), autism or schizophrenia disease in a human, comprising: imaging a brain of a human with a medical imaging process, wherein the medical imaging process is a multinuclear magnetic resonance process based on comparison of intensities obtained from relevant 31P or 1H nuclear magnetic resonance signals obtained from normal human child hosts and known abnormal human child hosts with cognitive impairment due to attention deficit hyperactivity disorder (ADHD), autism or schizophrenia with signal information obtained from a new human child patient with an unknown cognitive condition; anddetermining whether a conclusion of cognitively impaired cognitive impairment due to attention deficit hyperactivity disorder (ADHD), autism or schizophrenia is suggested using a first signal intensity for synaptic phosphomonoesters (sPME) of the human brain, a second signal intensity for synaptic phosphodiesters (sPDE) of the human brain, a third signal intensity for synaptic vesicles phosphodiesters (iPDE) of the human brain, a fourth signal intensity for phosphocreatine (PCr) of the human brain, and a fifth signal intensity for myo-inositol of the human brain 4. The method of claim 3 wherein an increased sPME level, a decreased sPDE level, a decreased iPDE level, a decreas... | - | A61B005/055
US07662105 | B2 | 2010-02-16 | Systems and methods for determining respiration metrics | 2005-12-14 | Systems, devices and methods provide for acquiring respiration information. A respiration information device includes timer circuitry to time a plurality of shorter time apertures and a plurality of longer time apertures. A respiration sensor, which may be implemented as a transthoracic impedance sensor, is configured to generate a signal indicative of patient respiration. For each aperture of the plurality of shorter time apertures and for each aperture of the plurality of longer time apertures, an estimated characteristic of the respiration is determined. Respiration metrics are developed using one or both of the estimated respiration characteristics of the shorter time apertures and the estimated respiration characteristics of the longer time apertures. | 1. A respiration information device, comprising: timer circuitry configured to time a plurality of shorter time apertures and a plurality of longer time apertures;a respiration sensor configured to generate a signal indicative of patient respiration; andrespiration information circuitry configured to determine, for each aperture of the plurality of shorter time apertures and for each aperture of the plurality of longer time apertures, an estimated value of a respiration characteristic based on the respiration, the respiration information circuitry comprising:measurement circuitry configured to measure the respiration characteristic for each breath cycle of each aperture; anda respiration processor configured to determine the estimated value of the respiration characteristic for each aperture based on the measurement of the respiration characteristic and configured to determine a maximum respiration value from the estimated values determined for the plurality of shorter time apertures and to determine a minimum respiration value from the estimated values determined for the plurality of longer time apertures. 2. The device of claim 1, wherein the respiration information circuitry is configured to determine one or more respiration metrics using one or both of the estimated values of the respiration characteristic determined for the shorter time apertures and the estimated values of the respiration characteristic determined for the longer time apertures. 3. The device of claim 1, wherein at least one of the timer circuitry, the respiration sensor, and the respiration information circuitry includes an implantable component. 4. The device of claim 1, wherein the respiration sensor comprises a transthoracic impedance sensor. 5. The device of claim 1, wherein the respiration characteristic is a respiration rate. 6. The device of claim 2, further comprising a trending unit configured to develop a trend using at least one of the estimated values of the respiration characteristic determined for the plurality of shorter time apertures, the estimated values of the respiration characteristic determined for the plurality of longer time apertures, and the one or more respiration metrics. 7. The device of claim 2, further comprising a diagnostics unit configured to detect a presence of a disease or disorder using at least one of the estimated values of the respiration characteristic determined for the plurality of shorter time apertures, the estimated values of the respiration characteristic determined for the plurality of longer time apertures, and the one or more respiration metrics. 8. The device of claim 7, wherein the disease or disorder comprises heart failure. 9. The device of claim 1, further comprising a therapy unit configured to deliver therapy based on a presence of a disease or disorder. 10. The device of claim 9, wherein the therapy comprises a pacing therapy. 11. A medical system, comprising: timer circuitry configured to time a plurality... | Cardiac Pacemakers, Inc. | A61B005/08 A61B005/05 A61N001/00
US07744541 | B2 | 2010-06-29 | Cerebral vascular reactivity monitoring | 2008-07-29 | A monitoring device may include logic to receive arterial blood pressure data associated with a patient and receive tissue hemoglobin data associated with the patient. The logic may also calculate a linear correlation between the arterial blood pressure data and the tissue hemoglobin data. The correlation may be used to assess vascular reactivity. | 1. A method, comprising: receiving, by a processing device, arterial blood pressure data associated with a patient;receiving, by the processing device, tissue hemoglobin data associated with the patient;calculating, by the processing device, a mean value for the arterial blood pressure data every first period of time;calculating, by the processing device, a mean value for the tissue hemoglobin data every first period of time; andcalculating, by the processing device, a linear correlation between the mean value of the arterial blood pressure data and the mean value for the tissue hemoglobin data. 2. The method of claim 1, wherein the calculating a linear correlation comprises: comparing a predetermined number of the mean values of the arterial blood pressure data and the tissue hemoglobin data, andgenerating a hemoglobin volume value every second period of time. 3. The method of claim 2, further comprising: outputting, by the processing device, a plurality of linear correlation values, the plurality of linear correlation values corresponding to a continuous hemoglobin volume index. 4. The method of claim 2, further comprising: plotting, by the processing device, the hemoglobin volume values against arterial blood pressure or a cerebral perfusion pressure associated with the patient. 5. The method of claim 4, further comprising: determining, by the processing device based on the plotting, a range of at least one of arterial blood pressure or cerebral perfusion pressure at which autoregulation for the patient is functioning properly. 6. The method of claim 2, wherein the calculating a linear correlation comprises calculating a Pearson coefficient. 7. The method of claim 6, wherein the calculating a Pearson coefficient comprises comparing paired samples of the arterial blood pressure and the tissue hemoglobin data taken over overlapping periods of time and repeating the comparing at least once every 60 seconds to generate a linear correlation coefficient at least once every 60 seconds. 8. The method of claim 1, further comprising: comparing, by the processing device, the linear correlation to a pressure reactivity value; anddetermining, by the processing device, a range of at least one of arterial blood pressure or cerebral perfusion pressure for the patient based on a result of the comparing. 9. A monitoring device, comprising: a processor configured to: receive arterial blood pressure data associated with a patient,receive tissue hemoglobin data associated with the patient, andcalculate a linear correlation between mean values of the arterial blood pressure data and mean values of the tissue hemoglobin data. 10. The monitoring device of claim 9, wherein the processor is further configured to calculate a linear correlation value every first period of time, and wherein a plurality of linear correlation values comprise an index that corresponds to a hemoglobin volume index of the patient over a period of time. 11. The monitoring device of cl... | Raba Equity Partners II, LLC The Johns Hopkins University | A61B005/00
US07662110 | B2 | 2010-02-16 | Devices for collecting blood and administering medical fluids | 2003-07-30 | Novel devices which can be used to both collect blood samples from and administer medical fluids to a patient on a repeated and continual basis using one rather than multiple needle insertions. The device typically includes a main tubing segment confluently connected to a cannula for insertion in the patient's vein. A syringe port and a volumeter for collecting blood branch separately from the main tubing segment. The device is used to collect blood by attaching an empty blood collection syringe to the syringe port, inserting the cannula in the patient's vein, allowing passive flow of blood from the main tubing segment into the volumeter under intrinsic venous blood pressure and capillary action, and then facilitating active flow of blood from the volumeter into the blood collection syringe by extending the syringe plunger. The device may be used to administer medical fluids to the patient through the main tubing segment from a medical fluid syringe or catheter attached to the syringe port. | 1. A device for collecting blood from and administering medical fluids to a patient, comprising: a main tubing segment for conveying the blood and the medical fluids;an indicator unit and an access port disposed in bidirectional fluid communication with said main tubing segment in branched relationship to each other at a tubing branch, said indicator unit adapted for indicating blood volume;wherein said indicator unit has a fluid-sealed first end disposed in fluid communication with and proximate said main tubing segment, a second end distal to said main tubing segment relative to said first end, an air flow pathway defined through said indicator unit between said first end and said second end and a bidirectional liquid flow pathway defined in coinciding relationship with said air flow pathway between said first end and said second end;a clamp operably engaging said main tubing segment and adapted to selectively block and unblock flow of the fluids in both directions through said main tubing segment; andat least one air-permeable and liquid-impermeable membrane provided at said second end of said indicator unit. 2. The device of claim 1 further comprising a blood volumeter provided in said indicator unit. 3. The device of claim 2 wherein said blood volumeter is a volumeter chamber. 4. The device of claim 3 wherein said indicator unit is disposed in removable fluid communication with said main tubing segment. 5. The device of claim 1 wherein said indicator unit is disposed in removable fluid communication with said main tubing segment. 6. The device of claim 5 further comprising a blood volumeter provided in said indicator unit. 7. A device for collecting blood from and administering medical fluids to a patient, comprising: a main tubing segment for conveying the blood and the medical fluids;a blood volumeter having a fluid-sealed first end disposed in fluid communication with and proximate said main tubing segment and a second end distal to said main tubing segment relative to said first end and an access port disposed in fluid communication with said main tubing segment in branched relationship to said blood volumeter, said blood volumeter adapted for indicating blood volume and said blood volumeter and said access port defining branched bidirectional fluid flow pathways;wherein an air flow pathway is defined through said blood volumeter between said first end and said second end and a bidirectional liquid flow pathway coincides with said air flow pathway between said first end and said second end;a clamp operably engaging said main tubing segment and adapted to selectively block and unblock flow of the fluids in both directions through said main tubing segment; andat least one air-permeable and liquid-impermeable membrane disposed in fluid communication with said blood volumeter at said second end of said blood volumeter and allowing bidirectional fluid movement between said blood volumeter and said access port. 8. The device of claim 7 ... | One Stick, LLC | A61B005/00 A61B005/02 A61B019/00 B65D081/00 A61M037/00 A61M029/00 A61M031/00 G01N001/00
US07657300 | B2 | 2010-02-02 | Registration of human anatomy integrated for electromagnetic localization | 2002-03-21 | A method for use during a procedure on a body. The method generates a display representing relative positions of two structures during the procedure. The method comprises the steps of storing an image data set in memory, the image data set representing the position of the body based on scans taken of the body prior to the procedure; reading the image data set stored in the memory, the image data set having a plurality of data points in known relation to a plurality of reference points for at least one of the two structures; placing one or more magnetic field sensors in known relation to the reference points of the two structures; generating a magnetic field; detecting the magnetic field with the magnetic field sensors; ascertaining the locations of the sensors based upon the magnetic field detected by the sensors and processing the locations of the sensors to generate a displaced image data set representing the relative position of the two structures during the procedure; and generating a display based on the displaced image data set illustrating the relative position of the two structures during the procedure. | 1. A device for use in a system for displaying relative positions of two structures during a procedure on a body, the system including a scanning device for scanning the body to generate a scan image, said device comprising: a base having a top side and a bottom side operable to be attached to the body via the bottom side of said base;a fiducial marker operable to be removably attached to the top side of said base; anda sensor formed within and as one structure with said base, wherein the fiducial marker is spaced a predetermined distance from the body when the base is attached to the body, the predetermined distance being greater than a slice thickness used for obtaining the scan image creating a barrier that inhibits the fiducial marker from blending with the body in the scan image. 2. The device as defined in claim 1 wherein attachment of said base to the body is selected from a group consisting of adhesive, implanting, clamping, stapling and screwing. 3. The device as defined in claim 1 wherein said fiducial marker is operable to be represented by at least one data point in an image data set of the body. 4. The device as defined in claim 3 wherein said image data set of the body is formed by use of a device selected from a group consisting of MR, CT, ultrasound, fluoro and PET. 5. The device as defined in claim 1 wherein said fiducial marker has a unique geometrical shape formed by at least three distinct non-collinear points operable to enable a determination of both a location and orientation of said fiducial marker. 6. The device as defined in claim 1 wherein said fiducial marker is spherical in shape to enable at least a determination of a location of said fiducial marker. 7. The device as defined in claim 1 wherein said sensor is selected from a group consisting of an electromagnetic sensor, an acoustic sensor, an optical sensor and an inertial sensor. 8. The device as defined in claim 1 wherein said sensor is an electromagnetic sensor having at least two coils positioned perpendicular to one another to provide both a location and orientation of said sensor. 9. The device as defined in claim 1 wherein said sensor is an electromagnetic sensor having at least one coil operable to identify at least a location of said sensor. 10. The device as defined in claim 1 wherein said sensor is selected from a group consisting of a five degree of freedom sensor and a six degree of freedom sensor. 11. The device as defined in claim 1 wherein communication with said sensor is selected from a group consisting of cable, radio wave, acoustic and optical. 12. The device as defined in claim 1 wherein the two structures includes body elements. 13. The device as defined in claim 1 wherein the two structures include a body element and a medical instrument. 14. The device as defined in claim 1 wherein said sensor is operable to be positioned at both a known location and orientation relative to said fiducial marker. 15. The device of claim 1 wherein... | Medtronic Navigation, Inc. | A61B005/05
US07647097 | B2 | 2010-01-12 | Transcutaneous implant | 2004-12-20 | Devices, systems and methods are disclosed for a neural access device that includes an implant which transcutaneously exits the skin of a patient and provides transport of signals between a sensor implanted in a patient and an external device. The transcutaneous implant has integrated features to provide reduced risk of injury due to mechanical forces as well as electrostatic discharge energy applied to the external portion of the device. Transcutaneous devices which provide wireless communication between a sensor and an external device are also disclosed. | 1. A neural access device for implanting in a patient comprising: a sensor comprising a plurality of electrodes to allow for chronic access of cells;an electrical conduit comprising a plurality of conductors, each conductor electrically connected to one or more of the electrodes;a connector assembly comprising: a first portion configured to be placed above a surface of a patient's skin;a second portion configured to be placed beneath the surface of the patient's skin;an attachment member for securing the second portion beneath the surface of the patient's skin; andat least one signal communication port configured to transport a signal to and/or from one or more of the electrodes; andan electrostatic discharge protector comprising a removable cap covering the at least one signal communication port and configured to protect the device from electrostatic dischargewherein the signal communication port is configured to transmit cellular activity detected from one or more of the electrodes and electrical signals to one or more of the electrodeswherein the cellular activity is detected from one or more of the electrodes and the electrical signals are transmitted to at least one of the one or more electrodes from which cellular activity is detected. 2. The device of claim 1, wherein the cap comprises an electrically insulating material. 3. The device of claim 1, wherein the cap comprises an electrically conductive material. 4. The device of claim 1, wherein the cap is attached to the connector assembly by a hinge. 5. The device of claim 4, wherein the hinge is spring-loaded. 6. The device of claim 1, wherein the cap comprises at least one filament attached to the cap. 7. The device of claim 6, wherein the at least one filament is designed to resemble human hair. 8. The device of claim 1, wherein the electrostatic discharge protector comprises an integrated circuit element configured to absorb electrical energy to prevent undesired electrical energy from reaching one or more of the electrodes. 9. The device of claim 1, wherein the electrostatic discharge protector comprises at least two redundant elements each configured to prevent undesired electrical energy from reaching one or more of the electrodes. 10. The device of claim 1, wherein the sensor is configured to detect multicellular signals emanating from the central nervous system. 11. The device of claim 10, wherein the multicellular signals comprise one or more of neuron spikes, electrocorticogram signals, local field potential signals, and electroencephalogram signals. 12. The device of claim 1, wherein the sensor is configured to detect multicellular signals from clusters of neurons and provide signals midway between single neuron and electroencephalogram recordings. 13. The device of claim 1, wherein the sensor is configured to transmit electrical signals to the central nervous system. 14. The device of claim 1, wherein the sensor is configured to both transmit and receive electrical... | BrainGate Co., LLC | A61B005/04
US07643860 | B2 | 2010-01-05 | Portable instrument for measuring a physiological quantity, including a device for illuminating the surface of an organic tissue | 2004-05-20 | A portable instrument for measuring a physiological quantity arranged to contact with the surface of organic tissue includes, essentially arranged in the same plane: an illumination device including at least one light source with an illumination surface for subjecting a portion of organic tissue to light emission in at least one wavelength range; and a detection device distant from the illumination device for detecting intensity of light emission produced by the illumination device after propagation in the organic tissue. The illumination device includes an optical element forming a guide coupled to the at least one light source for guiding light emission from the source by total internal reflection in a substantially parallel direction to the surface of the organic tissue and for distributing light emission into several illumination zones on the surface of the organic tissue over a substantially broader area than the illumination surface of the light source. | 1. A portable instrument, for measuring a physiological quantity, adapted to be brought into contact with a surface of an organic tissue, the portable instrument including, essentially arranged in a same plane: an illumination device including at least one light source with a determined illumination surface for subjecting a portion of the organic tissue to a light emission in at least one determined wavelength range; anda detection device distant from said illumination device for detecting intensity of the light emission produced by said illumination device after propagation in the organic tissue,wherein said illumination device further includes an optical element forming a guide having at least two portions at different orientations in the plane coupled to said at least one light source for guiding the light emission from said source by total internal reflection in a substantially parallel direction to the surface of the organic tissue to be illuminated and for distributing the light emission into several determined illumination zones on the surface of the organic tissue over a substantially broader area than the illumination surface of said light source. 2. The instrument according to claim 1, wherein said detection device has a determined geometrical contour and wherein each illumination zone is determined to be at a substantially constant distance from the closest point of said geometrical contour of the detection device. 3. The instrument according to claim 1, wherein said optical element forming a guide has a branched structure, or an at least partially cellular structure, or a branched and at least partially cellular structure. 4. The instrument according to claim 3, wherein said optical element forming a guide has an essentially honeycombed structure. 5. The instrument according to claim 1, wherein said detection device includes one or several quasi-punctual photoreceptors essentially arranged in the same plane as said optical element forming a guide and wherein said optical element forming a guide has a structure whose geometry essentially coincides with an arc of a circle or a circle of predetermined radius around the position of each photoreceptor. 6. The instrument according to claim 1, wherein said optical element is a solid optical guide provided with reflective micro-prismatic structures for redirecting the light emission produced by said light source in the direction of the organic tissue surface to be illuminated, said optical element having a first face, or lower face, directed towards said organic tissue to be illuminated, a second face, or upper face, opposite said first face, and lateral faces joining said first and second faces, oriented substantially perpendicularly to the organic tissue surface to be illuminated,each of said micro-prismatic structures including at least a first facet arranged on said second face, or on said lateral faces, or on said second face and on said lateral faces, to redirect said light emissio... | Asulab S.A. | A61B005/1455
US07662119 | B2 | 2010-02-16 | Multiple function ratcheting orthotic device | 2008-07-21 | An orthotic device for a joint of a human body at which a body appendage pivots includes a first member able to extend along a posterior portion of the appendage on a first side of the joint and a second member able to extend along a posterior portion of the appendage on a second side of the joint opposite the first side. A connection joining the first member and the second member provides a series of defined positions spaced angularly about an axis at which positions the connection can be alternately locked and released, allows pivoting, selective ratcheting and positioning of the second member relative to the first member about the axis, and can restrict pivoting of the second member relative to the first member within a desired angular range of motion. Notably, this orthotic device uniquely combines multiple joint functions into one convenient device which allows it to be used at different stages of therapy without being replaced or requiring additional equipment. | 1. An orthotic device for a joint of a human body at which a body appendage pivots, comprising: a first member able to extend along a posterior portion of the appendage on a first side of the joint;a second member able to extend along a posterior portion of the appendage on a second side of the joint opposite the first side; anda connection joining the first member and the second member and providing at least one of a series of defined positions spaced angularly about an axis at which positions the connection can be alternately locked and released, and pivoting of the second member relative to the first member about the axis within a desired angular range of motion by adjusting a flexion limiting lock nut relative to an extension limiting lock nut on a bolt aligned with the axis. 2. The orthotic device of claim 1 wherein: the connection alternately can be locked to prevent flexion of the appendage, and the connection permits flexion and extension of the appendage about the desired range of motion. 3. The orthotic device of claim 1 wherein the appendage is an arm and the joint is an elbow, and the first member extends along a forearm and the second member extends along an upper arm. 4. The orthotic device of claim 1 wherein the appendage is a leg and the joint is a knee, and the first member extends along a thigh and the second member extends along a calf. 5. An orthotic device for a joint of a human body at which a body part pivots, comprising: a first member having a first member cuff able to extend along a first posterior portion of the body part on a first side of the joint;a second member secured to the first member for pivoting about an axis, the second member having a second member cuff able to extend along a second posterior portion of the body part on a side of the joint opposite the first side, and the second member including ratchet teeth; anda ratchet supported for movement relative to the ratchet teeth and including a latch, a spring for urging the latch toward engagement with the ratchet teeth, and a lockout pin for alternately securing the latch against movement toward engagement with the ratchet teeth and for releasing the latch to engage the ratchet teeth, wherein the lockout pin includes a shank that can engage the first member such that the latch is prevented from engaging the ratchet teeth to allow contracture of the joint. 6. The orthotic device of claim 5, further comprising: mutually aligned adjuster holes formed on the ratchet, wherein the shank extends through the adjuster holes and the shank is formed with a screw thread; anda lockout hole formed on the first member, engageable by the shank and located such that the latch is prevented from engaging the ratchet teeth when the shank extends into the lockout hole. 7. The orthotic device of claim 5, further comprising: mutually aligned adjuster holes formed on the ratchet, wherein the shank extends through the adjuster holes and the shank is formed with a screw thread;a... | Anatomical Concepts, Inc. | A61F005/00
US08979888 | B2 | 2015-03-17 | Paranasal ostium finder devices and methods | 2009-07-30 | Devices and methods for locating sinus ostia and positioning a guide wire within the sinus ostia. The subject devices include a shaft having a distal end, a proximal end, a curved region located between the distal and proximal ends, and an interior channel, an extensible and retractable guide wire movably mounted within the interior channel and a probe tip joined to the guide wire. Certain devices further include expandable portions for engaging and treating body anatomy. | 1. A device for locating and dilating a natural ostium of a maxillary sinus, the device comprising; (a) an elongate shaft assembly, comprising: (i) a substantially rigid straight proximal portion,(ii) a curved distal portion extending from a distal end of the straight proximal portion,(iii) an atraumatic distal tip at the end of the curved distal portion, wherein the curved distal portion has a size and shape to allow passage of the distal portion into a nasal cavity to position the atraumatic distal tip within or near a maxillary sinus ostium, and(iv) an inflation lumen passing through at least part of the shaft assembly; and(b) at least one expandable dilator coupled with the shaft assembly, wherein the expandable dilator extends from the distal end of the straight proximal portion to a distal end of the curved distal portion and encompasses the curved distal portion therebetween, wherein the expandable dilator is in fluid communication with the inflation lumen; wherein the atraumatic distal tip comprises a ball-shaped tip for enhancing tactile feedback. 2. The device as in claim 1, wherein the expandable dilator comprises an inflatable balloon. 3. The device as in claim 2, wherein at least a portion of the balloon is disposed proximal to a curve in the distal portion of the shaft assembly. 4. The device as in claim 3, wherein the balloon extends from proximal to the curve to distal to the curve along the shaft assembly. 5. The device as in claim 2, wherein the dilator comprises two balloons, one disposed proximal to a curve of the distal portion of the shaft assembly and another disposed distal to the curve. 6. The device as in claim 2, wherein the balloon expends asymmetrically outward away from an inner radius of curvature of the curved distal portion of the shaft assembly. 7. The device as in claim 1, further comprising a handle including an actuator, wherein the dilator is advanceable and retractable along the shaft assembly using the actuator. 8. A device for locating and dilating a natural ostium of a maxillary sinus, the device comprising; (a) an elongate shaft assembly, comprising: (i) a substantially rigid straight proximal portion,(ii) a curved distal portion extending from a distal end of the straight proximal portion,(iii) an atraumatic distal tip at the end of the curved distal portion, wherein the curved distal portion has a size and shape to allow passage of the distal portion into a nasal cavity to position the distal tip within or near a dilation site in the nasal cavity, and(iv) an inflation lumen passing through at least part of the shaft assembly; and(b) at least one expandable dilator coupled with the shaft assembly, wherein the expandable dilator extends from the distal end of the straight proximal portion and encompasses the curved distal portion, wherein the expandable dilator is in fluid communication with the inflation lumen; wherein the atraumatic distal tip comprises a ball-shaped tip for enhancing tactile fe... | Acclarent, Inc. | A61M029/00 A61B001/233 A61B017/24 A61B017/22 A61B017/32 A61M025/10
US08093896 | B2 | 2012-01-10 | Uniform magnetic field spherical coil for MRI | 2007-09-21 | The present invention relates to a spherical design for a magnetic resonance imaging (MRI) coil to provide a compact and light-weight highly uniform magnetic field for a variety of medical imaging applications. A preferred embodiment of the invention provides a transportable MRI system in which the spherical electromagnet can be worn like a helmet by patients for diagnosis and assessment of traumatic brain injuries, for example. | 1. A method of magnetic resonance imaging (MRI) comprising: using a spherical coil assembly to generate a magnetic field within the coil assembly;positioning an object to be imaged within the spherical coil assembly;using an RF assembly positioned relative to the spherical coil assembly, the RF assembly generating a magnetic resonance signal; andacquiring a magnetic resonance image of the object within a region of interest. 2. The method of claim 1 further comprising providing a second spherical coil that is concentrically mounted within the spherical coil assembly. 3. The method of claim 2 further comprising providing a third spherical coil. 4. The method of claim 3 further comprising operating the first, second and third controls to control three orthogonal magnetic field components within the coil assembly. 5. The method of claim 3 further comprising applying a third magnetic field having a third orientation to the region of interest. 6. The method of claim 1 further comprising forming an image of brain tissue. 7. The method of claim 1 further comprising actuating the coil with a controller. 8. The method of claim 1 further comprising providing the coil assembly including a fluxball. 9. The method of claim 1 further comprising imaging a region of interest in a small animal or plant material. 10. The method of claim 1 further comprising applying a first magnetic filed having a first orientation to a region of interest with a first coil assembly and applying a second magnetic field having a second orientation to the region of interest that is orthogonal to the first orientation with a second coil assembly. 11. The method of claim 1 further comprising separating portions of the coil assembly to provide an opening for insertion of an object to be imaged. 12. The method of claim 1 further comprising providing a coil assembly having an inner radius in a range of 8 cm to 50 cm. 13. The method of claim 1 further comprising applying a gradient field with the coil assembly with a pulse sequence to image an object. 14. A magnetic resonance imaging (MRI) system comprising: a first magnetic field generating system providing a magnetic field within a spatial region in which material to be imaged is located, the generating system including a spherical coil assembly;a control system connected to the spherical coil assembly;an RF electromagnetic field generating and receiving system positioned relative to the spherical coil assembly, the RF system generating magnetic resonance (MR) data in response to magnetic resonance within the material;a data processing system that receives and processes the collected MR data, the processing system including a controller that generates a plurality of pulse parameters. 15. The system of claim 14 wherein the processing system is programmed to process image data. 16. The system of claim 14 wherein the processing system generates an actuating signal to actuate an activation magnetic field. 17. The system of cla... | Massachusetts Institute of Technology | G01V003/00
US08114028 | B2 | 2012-02-14 | Analysis system for determining an analyte in a body fluid and disposable integrated sample acquisition and analysis element | 2009-08-27 | An analyte determination analysis system with a disposable integrated sample acquisition and analysis element is disclosed. A piercing element has a tip for generating a bodily puncture wound and a test strip having a body fluid sample receiving zone to perform an analysis on a body fluid sample. A reusable analysis instrument comprises a coupling unit that couples the integrated sample acquisition and analysis element to a drive that drives a piercing element in a piercing movement comprising two phases: a propulsion phase wherein the piercing element moves on a puncture path in a puncture direction and a retraction phase wherein the piercing element moves opposite to the puncture direction after reaching a reversal point. A measuring and analysis unit measures a variable characteristic for the determination of the analyte and determines a desired analysis result on the basis of the measurement. The sample receiving zone of the test strip forms a sample contact. The piercing element, located parallel to the test strip, moves on a movement path relative to the test strip a puncture path. The piercing element, located adjacent to the sample contact, has a capillary channel with a sample inlet for body fluid entry after the piercing and a sample outlet for body fluid exit. The piercing element movement has a sample transfer position wherein the sample outlet is adjacent to the sample receiving zone such that the body fluid may be transferred from the capillary channel through the sample outlet to the sample receiving zone. | 1. An analysis system for determining an analyte in a body fluid, comprising: a disposable integrated sample acquisition and analysis element comprising, a flat test strip having a front end and a sample receiving zone to receive a body fluid sample for analysis, wherein the sample receiving zone is located on a flat side of the test strip to form a sample contact side, wherein the sample receiving zone is adjacent to the front end of the test strip in a puncture direction, anda piercing element located parallel to the test strip and adjacent to the sample contact side of the test strip, the piercing element comprising, a tip for generating a bodily puncture wound, anda capillary channel comprising, a sample inlet for body fluid entry after the piercing into the capillary channel, anda sample outlet for body fluid exit from the capillary channel,wherein the piercing element is movable relative to the test strip on a movement path during at least a part of a puncture movement and is positioned in the puncture direction behind the sample receiving zone in such a manner that the sample receiving zone is located between the piercing element and the front end of the test strip,a reusable analysis instrument comprising, a coupling unit adapted to couple the integrated sample acquisition and analysis element to a drive for driving the piercing element in a piercing movement, wherein the piercing movement comprising, a propulsion phase in which the piercing element is moved on a puncture path in the puncture direction,a retraction phase in which the piercing element is moved opposite to the puncture direction after reaching a reversal point of the piercing movement, anda sample transfer position, wherein the sample outlet is adjacent to the sample receiving zone of the test strip such that body fluid may be transferred from the capillary channel through the sample outlet to the sample receiving zone, anda measuring and analysis unit for measuring a measurement variable characteristic for the determination of the analyte and for determining a desired analysis result on the basis of the measurement. 2. The analysis system according to claim 1, wherein the measuring and analysis unit comprises, an optical measuring apparatus for measuring a photometric measurement with an optically measurable measurement variable which is characteristic for the determination of the analyte. 3. The analysis system according to claim 1, wherein the sample receiving zone has a test field comprising, at least one absorbing layer,a sample receiving surface on the sample contact side of the test strip, anda detection surface on the side of the test strip opposite to the sample contact side. 4. The analysis system according to claim 1, wherein the tip of the piercing element is moved during the propulsion phase of the piercing movement beyond the front end of the test strip in the puncture direction. 5. The analysis system according to claim 1, further comprising a collection ... | Roche Diagnostics Operations, Inc. | A61B005/00
US08185325 | B2 | 2012-05-22 | Multi-functional, discrete and mutually exclusive method for determining concentrations of gases in a gaseous mixture | 2009-02-10 | Method and device for detection and quantitative and qualitative analysis of components in a gaseous mixture distinguished by high selectivity and high resolution.;Method allows to distinguish the influence of individual gases, by themselves or in a mixture, on the microstructure of a sensor's sensitive layer and utilizing the variations of measured parameters to analyze and derive the characteristics of gases, for example, the concentration of a gas or multiple gases in a mixture.;As an example, the method could be utilized in medicine for non-invasive detection of the blood glucose level in diabetics.;Device realizing the method is described. | 1. A method for determining a concentration of a component of a sample, the method comprising: [1] trapping a volume of the sample in a trapping chamber;[2] passing at least a portion of the trapped volume of sample into a sensing chamber, wherein the sensing chamber includes or is in connection with a sensor;[3] detecting a change in current as a result of interaction of the sample with a sensor;[4] using at least Mathieu's Equation with a value for each of its initial conditions, to determine when a value associated with the detected current is in a stable domain;[5] when the value associated with the detected current is not in a stable domain, varying one or more of the initial conditions of Mathieu's Equation, and repeating the determining until the detected current is in a stable domain; and[6] determining the concentration of the component of the sample using the value associated with the detected current. 2. The method of claim 1, wherein the sample is a gas. 3. The method of claim 1, wherein the sample is a liquid. 4. The method of claim 1, wherein the value associated with the detected current is a value of the detected current. 5. The method of claim 1, wherein a temperature of the sensor is set to a different temperature from that as used during a previous determining of the concentration of the component of the sample. 6. The method of claim 1, wherein a pressure of the trapped volume of sample is varied from a pressure of the trapped volume of sample as measured during a previous determining of the concentration of the component of the sample. 7. The method of claim 1, wherein Mathieu's Equation is of a form D^2x/dt^2+lambda^2(1+mu*f(t))x=0 where lambda represents a numeric constant, where mu represents another numeric constant, and where t represents time. 8. The method of claim 1, wherein the trapping of the volume of the sample in the trapping chamber includes trapping the volume of the sample for a sufficient time so that a substantial portion of the trapped volume of the sample is isolated from an influence of an atmospheric temperature and isolated from an influence of an atmospheric pressure. 9. The method of claim 1, wherein the trapping of the volume of the sample in the trapping chamber includes trapping the volume of the sample sufficiently so as to avoid a measurable amount of contamination. 10. The method of claim 1, wherein a stable domain is a numeric domain where values can be predicted in accordance with the said Mathieu's Equation, and wherein an unstable domain is a numeric domain where a resulting value is unreliable. 11. The method of claim 1, wherein the determining the concentration of the component of the sample includes the use of parametric vibrations when the value associated with the detected current is approximately at a border between a stable domain and an unstable domain. 12. The method of claim 1, wherein the method further comprises using a unique identifier of, or a particular value assoc... |  | G01N031/00 G11C017/00
US08273019 | B2 | 2012-09-25 | Clinical investigation data logging with contextual time shifting | 2008-07-22 | A health monitoring device that includes a data logger module that collects health information from a subject or clinician during specific time ranges that are part of a health monitoring schedule. The collected health information and/or the time ranges are determined in part by the content of the collected health information. The device records the collected health information along with a time stamp. The device also includes a clock and a time shift mode that allows a user to adjust the clock to a specified point in the health monitoring schedule so that the device operates at the specified point in the same manner as if the device had reached the specified point without intervention by the user to adjust the clock. | 1. A health monitoring device comprising: a clock, wherein the value of the clock comprises an operating date-time of the health monitoring device as the device executes a health monitoring schedule based on a study definition that defines a planned date-time for scheduled elements and events of interest in the study definition to occur and wherein the clock value determines an operational state of the health monitoring device;a data logger module including logic for: collecting a first set of health information from a subject or clinician during a first time range and a second set of health information from the subject or clinician during a second time range, at least one of the second set of health information and the second time range being determined in part by a content of the first set of health information, the first and second time ranges being part of the health monitoring schedule;recording the first and second sets of health information in a computer readable medium; andincluding a time stamp with the recorded health information in the computer readable medium;a user interface for receiving the sets of health information from the subject or the clinician;a time shifting module including logic to permit, in a time shifting mode, an adjustment of the clock from a first event of interest in the health monitoring schedule, corresponding to a first operational state of the device, to at or relative to a second event of interest in the health monitoring schedule, corresponding to a second operational state of the device, the device operating at the second event of interest in the health monitoring schedule in the same manner and collecting health information from the subject or clinician in the same manner as if the device had reached the second event of interest without intervention by the time shifting module and resulting in activation of the second operational state of the health monitoring device;wherein the time shifting module further includes logic to prevent adjustment of the date and/or time to alter operational state of the health monitoring device and move to the second event of interest when the device is not in the time shifting mode;wherein adjustment of the clock from the first point in the health monitoring schedule to the second event of interest in the health monitoring schedule changes an operational state of the device without modification to the study definition; anda display having a screen for displaying information corresponding to the operational state of the device. 2. The device of claim 1, wherein the logic for collecting a first set of health information and a second set of health information includes logic to collect diary entries. 3. The device of claim 1, wherein the logic for collecting a first set of health information during a first time range and a second set of health information during a second time range includes logic to determine first and second time ranges that correspond to at least one of a spe... | PHT Corporation | A61B005/00
US20100016677 | A1 | 2010-01-21 | Devices for and Methods of Analyzing a Physiological Condition of a Physiological Subject Based on A Workload Related Property | 2007-03-20 | A method of influencing a physiological condition of a physiological subject, the method comprising detecting a workload distribution/activity in a brain of the physiological subject, the workload distribution/activity being indicative of the physiological condition, and selectively modifying the workload distribution/activity in the brain of the physiological subject to thereby modify the physiological condition. | 1. A device for influencing a physiological condition of a physiological subject, the device comprising a detection unit adapted for detecting a workload related property in a brain of the physiological subject, the workload related property being indicative of the physiological condition, wherein the detection unit is adapted for detecting or localizing one or more capacity bottlenecks in the brain of the physiological subject; anda modification unit adapted for selectively modifying the workload related property in the brain of the physiological subject to thereby modify the physiological condition so that the one or more capacity bottlenecks are reduced by performing one or more of the following measures: eliminating or decreasing too much strength of patterns of unwanted behavior or unwanted activity;redirecting processing from special parts of the brain to other parts;strengthening associations with patterns of wanted behavior or activity;selectively modifying unwanted activity;taking away problems that cause excessive processing needs. 2. The device of any one of claims 1, 14, and 21, adapted for influencing, as the physiological condition, at least one of the group consisting of a mental condition, a mental disorder, a mental disease, headache, migraine, mental fitness, aggressive behavior, violence behavior, crime behavior, obsessive behavior, Dyslexia, ADHD, an eye-focusing problem, concentration, memory, depression, fatigue, autism, Dyscalculation, whiplash, dizziness, drug addiction, gambling addiction, and a trauma. 3-4. (canceled) 5. The device of any one of claims 1, 14, and 21, wherein the detection unit is adapted for detecting at least one of the group consisting of a location, a shape, a size, a temporal property, and a number of the one or more capacity bottlenecks. 6. The device of any one of claims 1, 14, and 21, wherein the detection unit is adapted for detecting a functional coupling between portions of the brain of the physiological subject. 7. The device of claim 1, wherein the modification unit is adapted for selectively modifying the workload related property in the brain of the physiological subject by taking at least one measure of the group consisting of supplying a medication to the physiological subject, supplying a perceivable stimulus to the physiological subject, reducing the activity in one or more selected portions of the brain of the physiological subject, increasing the activity in one or more selected portions of the brain of the physiological subject, and temporarily disabling a sense of the physiological subject. 8. (canceled) 9. The device of claim 1, wherein the modification unit is adapted for selectively modifying the workload related property to thereby balancing a workload distribution over different portions of the brain. 10. The device of claim 1, comprising a database unit adapted for storing predetermined data indicative of the physiological conditions and/or of other physiological condi... | - | A61B005/00 A61K031/167
US20100036207 | A1 | 2010-02-11 | Relative response system including reprogramming capability for autonomous or interrelated stimulus and sensor systems for measuring biological response data relative to either an absolute reference and/or relative to other biological response | 2009-07-24 | A wireless system of independent CPU pods for initiating, collecting, and optionally processing, absolute and/or relative response data of biological systems to stimulus and then using wireless means to transmit the data to a plurality of external communications (e.g. human interface device) via a bridge pod. The system is set up to perform a plurality of sequences of stimulus and then response measurement according to programming through the wireless network which can be changed on each CPU pod via re-programming at any time to create new configurations and metrics. The sequences may execute on each CPU pod(s) and CPU bridge pod (s) autonomously or the programing can create a plurality of different interrelated sequences depending on system objectives. A plurality of sensors and stimulate devices are connected or integrated to the CPU pods according to the purposes of the programmed stimulate response sequences. This system is beneficial for getting stimulus-response data for any biological activity in a relative of independant response manner making it quite general. Thus, interrelated reactions of any number of biological systems relative to each other can be measured, processed and communicated to any desired device for further processing or display. Thus, either individual performance or entire teams relative performance can be tracked to improve performance in a plurality of ways in a plurality of sports, military, etc. exercises. | 1. A relative response system including reprogramming capability for autonomous or interrelated stimulus and sensor systems for measuring biological response data relative to either an absolute reference and/or relative to other biological response for collecting, and optionally processing, stimulus-response data of biological systems, comprising: a wireless network, for providing data transmission between cpu pod(s) and cpu bridge pod(s) and also, when needed, a transmission path for software reprogramming of cpu pod(s) and cpu bridge pod(s);an integrated wireless negotiation and microprocessor, power source, independently programmable and reprogrammable at any time cpu pod, for executing sequences of stimulus-response according to changeable programming either synchronously or asynchronously and with autonomous or interrelated behavior with other response sensors, other cpu pod(s) or other cpu bridge pod(s) and communicating on the wireless network;a cpu bridge pod, for communication connection from wireless network to the external communication path as well as perform all functions that a cpu pod performs, wirelessly connected to said CPU pod;an external communication path, for download and upload of programming and/or data between the cpu bridge pod and a plurality of external communications (e.g. network, computer, human, etc), electrically connected to said CPU bridge pod;a stimulus response sequence, for software coded instructions for executing a plurality independent and/or interrelated stimulus-response actions for an individual biological system or a group of biological systems in order to measure a plurality of biological response data relative to an absolute reference and/or a reference of one response sensor relative to another;a stimulate device, for stimulating any aspect of a biological system including anything a biological system can sense voluntarily or involuntarily, electrically connected to said CPU pod; anda response sensor, for sensing any aspect of a biological system, electrically connected to said CPU pod. | - | A61B005/16 A61B005/00
US20100022849 | A1 | 2010-01-28 | MEDICAL WORKSTATION WITH INTEGRATED SUPPORT OF PROCESS STEPS | 2009-06-29 | A device (13) is provided for use with a treatment device for treating a patient. The device (13) includes a detection device (17) for detecting at least one transition of the patient's treatment from a first phase of treatment to a second phase of treatment. A signal device (25) for sending at least one signal when the transition is detected. It proposes, furthermore, a workstation equipped herewith. | 1. A device for use with a treatment device for treating a patient, the device comprising: a detection means for detecting at least one phase of treatment and/or a transition of the patient's treatment from a first phase of treatment to a second phase of treatment; anda signal means for sending at least one of a signal related to the detected phase of treatment and a signal when the transition is detected. 2. A device in accordance with claim 1, further comprising: a display means for displaying the phase of treatment in which the patient is at a point in time. 3. A device in accordance with claim 1, further comprising: an input means for selecting possible phases of treatment. 4. A device in accordance with claim 1, further comprising: a display means for displaying the actions associated with a change from one phase to another. 5. A device in accordance with claim 1, further comprising: an input means for changing values that the treatment device seeks to reach or are set in the device for treatment parameters of phases of treatment. 6. A device in accordance with claim 1, further comprising: a data bank means for storing work flows data. 7. A device in accordance with claim 1, with a control means for controlling work flow. 8. A device in accordance with claim 1, further comprising: wherein the device for treatment is a device for anesthetizing the patient. 9. A medical workstation, comprising: a plurality of treating devices and/or monitoring devices for the treatment and/or monitoring of a patient; anda treatment phase detection and signaling device comprising a detection means for detecting at said plurality of treating devices and/or monitoring devices at least one phase of treatment and/or a transition of the patient's treatment from a first phase of treatment to a second phase of treatment; and a signal means for sending at least one of a signal related to the detected phase of treatment and a signal when the transition is detected. 10. A workstation in accordance with claim 9, wherein said treatment phase detection and signaling device further comprises: a display means for displaying the phase of treatment in which the patient is at a point in time. 11. A workstation in accordance with claim 9, wherein said treatment phase detection and signaling device further comprises: an input means for selecting possible phases of treatment. 12. A workstation in accordance with claim 9, wherein said treatment phase detection and signaling device further comprising: a display means for displaying the actions associated with a change from one phase to another. 13. A workstation in accordance with claim 9, wherein said treatment phase detection and signaling device further comprises: an input means for changing values that the treating devices and/or monitoring devices seeks to reach or are set in the device for treatment parameters of phases of treatment. 14. A workstation in accordance with claim 9, wherein said treatment phase detecti... | DRAGER MEDICAL AG & CO. KG | A61B005/00 A61B019/00
US20100016742 | A1 | 2010-01-21 | System and Method for Monitoring, Measuring, and Addressing Stress | 2008-07-19 | In accordance with one embodiment of the present invention, a method for quantifying stress is presented. The method includes interacting with an end user through a stress module and processing heart rate variance data collected during a stress test. The heart rate variance data may include a coherence attribute. The method further includes transmitting the heart rate variance data over a network and to a specified location. In more specific embodiments, the method includes delivering a numerical score to the end user. The score signifies a stress level for the end user. | 1. A method, comprising: interacting with an end user through a stress module;processing heart rate variance data collected during a stress test, wherein the heart rate variance data may include a coherence attribute; andtransmitting the heart rate variance data over a network and to a specified location. 2. The method of claim 1, further comprising: delivering a numerical score to the end user, and wherein the score signifies a stress level for the end user. 3. The method of claim 1, wherein the interacting is performed through a website, a health station, a handheld device, or a personal computer. 4. The method of claim 1, wherein the heart rate variance data is stored and sent to a first destination and then to a second destination over a network. 5. The method of claim 4, wherein the handheld device is a cellular phone, an I-Phone, a PDA, or a handheld stress test tool. 6. The method of claim 1, wherein the test includes designations for Low, Medium, and High stress levels, whereby each stress level includes a respective range of stress scores. 7. The method of claim 1, wherein the test includes sounds that are played for the end user, or messaging provided to the end user. 8. The method of claim 1, further comprising: providing educational materials to the end user after the stress test. 9. The method of claim 1, wherein the stress test includes a color scheme that indicates levels of stress for the end user. 10. The method of claim 1, further comprising: providing a health station as the remote location;capturing stress data from the end user at the health station;storing the stress data; andanalyzing the stress data to determine an intervention plan for the end user. 11. An apparatus, comprising: a stress module interacting with an end user, wherein the stress module processes heart rate variance data collected during a stress test, the heart rate variance data including a coherence attribute, and wherein the stress module delivers a numerical score to the end user, wherein the score signifies a stress level for the end user. 12. The apparatus of claim 11, wherein the score is based on a plurality of time intervals associated with the test. 13. The apparatus of claim 11, wherein the interacting is performed through a website, a health station, a handheld device, or a personal computer. 14. The apparatus of claim 11, further comprising: a memory that stores heart rate variance data, wherein the data is sent to a first destination and then to a second destination over a network. 15. The apparatus of claim 11, wherein the test includes designations for Low, Medium, and High stress levels, whereby each stress level includes a respective range of stress scores. 16. Software for processing data, the software being embodied in a component that includes a processor and a memory, the software when executed by the processor being operable to: interact with an end user via a stress module; andprocess heart rate variance data collected ... | - | A61B005/0402
US20100030060 | A1 | 2010-02-04 | Scope, Vision | 2008-08-04 | The VISION SCOPE allows claustrophobic individuals to enter an open or closed Magnetic Resonance Imaging (MRI) machine without claustrophobic tendencies and view the outside areas while undergoing testing without interfering with the magnetic field created by the MRI machine. | 1. I claim as my invention a Scope, Vision which allows claustrophobic individuals to enter an open or closed Magnetic Resonance Imaging (MRI) machine without claustrophobic tendencies and view the outside areas while undergoing testing. 2. Our research indicates that a minimum of 30% of individuals requiring a closed MRI must be removed from the machine prior to completion of the procedure. 3. The non metallic nature of the mirror sheeting does not affect functionality of the MRI machine. 4. The use of the Scope, Vision alleviates the claustrophobic tendencies for individuals, by allowing them to view the outside areas thus enabling total completion of the MRI. 5. Our research indicates some individuals requiring a closed MRI required medication prior to the MRI procedure. In some cases the use of the Scope, Vision has eliminated the use of medication. 6. The availability lessens the pre-MRI stress of claustrophobic individuals. 7. Use of the Scope, Vision will allow laboratories to be more efficient and effective by scheduling more patients and having less down time in the testing process. 8. Cost of the Scope, Vision is minimal as compared to the machine usage time and cost per individual. | - | A61B005/055
US20100010328 | A1 | 2010-01-14 | PROBES AND SENSORS FOR ASCERTAINING BLOOD CHARACTERISTICS AND METHODS AND DEVICES FOR USE THEREWITH | 2008-07-11 | A probe for use in a patient having a vessel carrying blood to ascertain characteristics of the blood having a cannula adapted to be inserted into the vessel of the patient. The cannula has a length so that when the distal extremity is in the vessel of the patient the proximal extremity is accessible outside of the patient. A gas sensor assembly is carried within the distal extremity of the cannula for determining gas characteristics of the blood in the vessel. A pressure sensor is carried within the distal extremity of the cannula for determining the pressure of the blood in the vessel. | 1. A probe for use in a patient having a vessel carrying blood to ascertain characteristics of the blood, the probe comprising: a cannula adapted to be inserted into a vessel of a patient and having a proximal extremity and a distal extremity, the cannula having a length so that when the distal extremity is in the vessel of the patient the proximal extremity is accessible outside of the patient;a gas sensor assembly carried within the distal extremity of the cannula for determining gas characteristics of the blood in the vessel;a pH sensor assembly carried within the distal extremity of the cannula for determining pH of the blood in the vessel; anda pressure sensor assembly carried within the distal extremity of the cannula for determining pressure of the blood in the vessel. 2. The probe of claim 1, wherein the pressure sensor assembly comprises a solid state pressure sensor. 3. The probe of claim 1, wherein the distal extremity of the cannula comprises a gas permeable material proximate to the gas sensor assembly. 4. The probe of claim 1, further comprising an electrolyte solution disposed within the cannula, and wherein the gas sensor assembly includes electrodes disposed in the electrolyte solution for providing electrical inputs to and electrical outputs from the gas sensor assembly. 5. The probe of claim 1, wherein the gas sensor assembly comprises an oxygen sensor assembly. 6. The probe of claim 5, wherein the gas sensor assembly further comprises a carbon dioxide sensor assembly. 7. The probe of claim 6, further comprising conductor elements connected to the oxygen sensor assembly and to the carbon dioxide sensor assembly and configured to supply electrical outputs to the proximal extremity of the cannula. 8. The probe of claim 5, wherein the oxygen sensor assembly includes: an electrolyte fill solution disposed in the cannula; andfirst, second, and third electrodes disposed in the electrolyte fill solution and configured to provide electrical outputs to the proximal extremity of the cannula. 9. The probe of claim 1, further comprising a flex circuit assembly extending from the proximal extremity of the cannula to the distal extremity of the cannula, wherein the gas sensor assembly, the pH sensor, and the pressure sensor assembly are mounted on the flex circuit assembly. 10. The probe of claim 1, wherein the gas sensor assembly, the pH sensor assembly, and the pressure sensor assembly are attached to conductor elements of a flex circuit assembly that provides a zero profile connector. 11. The probe of claim 2, combined with a module comprising electronics configured to create and digitize analog signals from the pressure sensor assembly and software algorithms configured to provide measurements of systolic and diastolic blood pressure, mean arterial pressure, heart rate, and systemic vascular resistance. 12. The probe of claim 5, wherein the oxygen sensor assembly comprises a working electrode, a counter electrode, and a refere... | - | A61B005/1468
US20100004557 | A1 | 2010-01-07 | METHOD FOR DIAGNOSIS AND TREATMENT OF PROCESSING DIFFICULTIES, INTEGRATION PROBLEMS, IMBALANCES AND ABNORMAL POSTURES | 2009-09-17 | Methods for diagnosis and treatment are provided for a patient having processing difficulties, integration problems, imbalances or abnormal postures. In one embodiment, a patient is identified having retinal signal problems creating at least one negative medical condition. Then, at least a portion of the patient's lacrimal drainage system is blocked or altered. Further, as a result of the diagnosis and treatment, the patient achieves improved retinal signal processing and at least one of the patient's negative medical conditions is improved. Other embodiments include using in combination the Z-Bell Test and/or light blocking or altering devices for diagnosis and treatment. | 1. A method for an auditory test for a patient, comprising the steps of: instructing the patient to close the patient's eyes;selecting a location in front of the patient within a test zone between the patient's shoulders and above the patient's armpits;using a means for generating a sound to create a sound at the selected location;instructing the patient to locate the device generating the sound with at least one of the patient's fingers;scoring the test results and repeating the test as required for different selected locations,analyzing the test results; anddiagnosing sensory linkages of non-visual fibers. 2. The method of claim 1 whereby the means for generating a sound is a bell. | - | A61B005/00
US20100016698 | A1 | 2010-01-21 | INTEGRATED RECEIVER FOR CONTINUOUS ANALYTE SENSOR | 2009-09-23 | A system is provided for monitoring glucose in a host, including a continuous glucose sensor that produces a data stream indicative of a host's glucose concentration and an integrated receiver that receives the data stream from the continuous glucose sensor and calibrates the data stream using a single point glucose monitor that is integral with the integrated receiver. The integrated receiver obtains a glucose value from the single point glucose monitor, calibrates the sensor data stream received from the continuous glucose sensor, and displays one or both of the single point glucose measurement values and the calibrated continuous glucose sensor values on the user interface. | 1. A system comprising: a continuous sensor configured for at least partial insertion into a host, wherein the continuous sensor generates continuous sensor data associated with an analyte concentration in the host;a single point monitor external to the host, wherein the single point monitor generates at least one single point value associated with an analyte concentration of a biological sample received from the host; anda processing unit configured to calibrate the continuous sensor data using the single point value from the single point monitor, wherein the processing unit is in data communication with the continuous sensor and configured to receive and process the continuous sensor data from the continuous sensor, and wherein the single point monitor is physically connected to the processing unit. 2. The system of claim 1, wherein each of the continuous sensor and the single point monitor comprise a sensing membrane comprising an enzyme domain, and wherein the enzyme domain of the sensing membrane of the continuous sensor and the enzyme domain of the sensing membrane of the single point monitor are substantially the same. 3. The system of claim 1, wherein the single point monitor is configured to receive a single use test strip. 4. The system of claim 1, wherein the processing unit further comprises programming to control administration of a constituent to the host. 5. The system of claim 4, wherein the programming to control administration of a constituent to the host is configured to determine at least one of an amount of insulin and a time of insulin administration. 6. The system of claim 1, wherein the processing unit further comprises programming to calculate a rate of change of the continuous sensor data and/or directional trend of the continuous sensor data. 7. The system of claim 6, wherein the processing unit further comprises programming to control administration of a constituent to the host based at least in part on a rate of change of the continuous sensor data and/or a directional trend of the continuous sensor data. 8. The system of claim 1, wherein the processing unit further comprises programming to control administration of a constituent to the host based at least in part on the at least one single point value. 9. The system of claim 1, wherein the processing unit is configured to monitor the continuous sensor data to adaptively determine a time for requesting a biological sample. 10. The system of claim 1, further comprising a user interface associated with the processing unit and the single point monitor, wherein the processing unit is configured to prompt the user, via the user interface, to provide a biological sample to the single point monitor at predetermined intervals. 11. The system of claim 1, further comprising a user interface associated with the processing unit and the single point monitor, wherein the processing unit is configured to prompt the user to provide a biological sample to the single point mo... | DexCom, Inc. | A61B005/1468 A61M005/00
US20100036272 | A1 | 2010-02-11 | METABOLIC MEASURE SYSTEM INCLUDING A MULTIPLE FUNCTION AIRWAY ADAPTER | 2008-01-31 | A system (300) for measuring a metabolic parameter. The system includes an integrated airway adapter (20, 100, 200) capable of monitoring any combination of respiratory flow, O2 concentration, and concentrations of one or more of CO2, N2O, and an anesthetic agent in real time, breath by breath. Respiratory flow may be monitored with differential pressure flow meters under diverse inlet conditions through improved sensor configurations which minimize phase lag and dead space within the airway. Molecular oxygen concentration may be monitored by way of luminescence quenching techniques. Infrared absorption techniques may be used to monitor one or more of CO2, N2O, and anesthetic agents. | 1. A metabolic measurement system (300) comprising: (a) an airway adapter (20, 100, 200, 334) adapted to be coupled in series with a mainstream gas flow comprising: (1) a housing have a bore (34) defined therethrough to carry the mainstream gas flow through the airway adapter,(2) a window (40) defined in the housing providing optical access to the gas flow through the airway adapter, and(3) an opening (234) defined in the housing providing optical access to the gas flow through the airway adapter;(b) a sensor head (22, 334) adapted to be removably attached to the airway adapter, the sensor head comprising: (1) an infrared sensing system adapted to transmit or receive infrared through the window, and(2) a luminescence quenching system adapted to (i) transmit excitation radiation through the opening, (ii) receive emitted radiation from a luminescence material through the opening, or both (i) and (ii);(c) a flow measurement system (310, 352) adapted to be coupled in series with such a mainstream gas flow; and(d) a processor (320, 330) adapted to receive signals from the infrared sensing system, the luminescence quenching system, and the flow measurement system, wherein the processor is adapted to determine a metabolic parameter including oxygen consumption (VO2), carbon dioxide production (VCO2), respiratory quotient (RQ), resting energy expenditure (REE), or any combination thereof. 2. The system of claim 1, wherein (a) the processor is disposed in the sensor head, or (b) the processor is spaced apart from the sensor head and communicates with the infrared sensing system, the luminescence quenching system, or both via a hardwired or a wireless communication link. 3. The system of claim 1, wherein at least a portion of the flow measurement system is disposed in the housing of the airway adapter, the sensor head, or both. 4. The system of claim 3, wherein the flow measurement system comprises: a flow restrictor (76, 136, 138, 216, 218, 220, 222, 224) disposed in the housing and adapted to create a pressure drop across the flow restrictor; anda pressure sensor adapted to measure a pressure associated with the pressure drop. 5. The system of claim 4, wherein the pressure sensor is disposed in the sensor head. 6. The system of claim 5, wherein the processor is spaced apart from the sensor head, and further comprising a hardwired or wireless communication link to communicate signals from the infrared sensing system, the luminescence quenching system, and the pressure sensor to the processor. 7. The system of claim 6, further comprising an output device (316, 322, 324, 332, 340) operatively coupled to the processor for providing the metabolic parameter in a human perceivable format. 8. The system of claim 4, wherein the pressure sensor is spaced apart from the airway adapter, and further comprising a pneumatic tubing (312) communicating one side of the flow restrictor with the pressure sensor. 9. The system of claim 4, wherein the processor is dis... | KONINKLIJKE PHILIPS ELECTRONICS N.V. | A61B005/083 G06F019/00 G01J005/00 G01F001/00 A61B005/08
US07657296 | B2 | 2010-02-02 | Unitary medical sensor assembly and technique for using the same | 2006-07-28 | A sensor assembly is provided that includes a frame having a first portion and a second portion connected by a hinge. An emitter and a detector are disposed on the frame. A coating is provided over the frame, the emitter and the detector to form a unitary sensor assembly. The sensor assembly also includes a resistance-providing component disposed generally about the hinge. In one embodiment, the sensor assembly may be placed on a patient's finger, toe, ear, and so forth to obtain pulse oximetry or other physiological measurements. A method of manufacturing the sensor assembly is also provided as is a method of cleaning a fluid-tight sensor assembly. | 1. A sensor assembly, comprising: a frame comprising a first portion and a second portion connected by a hinge;an emitter disposed on the frame;a detector disposed on the frame;a coating molded over the frame, the emitter, and the detector to form a unitary sensor assembly; anda resistance-providing component disposed generally about the hinge. 2. The sensor assembly of claim 1, wherein the resistance-providing component comprises one or more elastic bands disposed beneath the coating. 3. The sensor assembly of claim 1, wherein the resistance-providing component comprises one or more elastic bands disposed external to the coating. 4. The sensor assembly of claim 1, wherein the resistance-providing component comprises one or more elastic bands disposed within the coating. 5. The sensor assembly of claim 1, wherein the resistance-providing component comprises a region of the coating. 6. The sensor assembly of claim 1, wherein the resistance-providing component comprises a thick region of the coating. 7. The sensor assembly of claim 1, comprising one or more signal transmission structures connected to at least one of the emitter or the detector. 8. The sensor assembly of claim 1, wherein the unitary sensor assembly comprises at least one of a pulse oximetry sensor, a sensor for measuring a water fraction, or a combination thereof. 9. The sensor assembly of claim 1, wherein the frame comprises a single molded part. 10. The sensor assembly of claim 1, wherein the frame comprises two or more parts joined to form the frame. 11. The sensor assembly of claim 1, wherein the resistance-providing component comprises at least one of a polymeric material or a metal. 12. The sensor assembly of claim 1, wherein the coating comprises a thermoplastic elastomer. 13. The sensor assembly of claim 12, wherein the thermoplastic elastomer comprises at least one of a thermoplastic polyolefin, a thermoplastic vulcanizate alloy, thermoplastic polyurethane, silicone, or a combination thereof. 14. The sensor assembly of claim 1, wherein the coating comprises a conformable material. 15. The sensor assembly of claim 1, wherein no moving parts of the frame remain exposed from the coating. 16. The sensor assembly of claim 1, wherein the unitary sensor assembly is configured to move between at least two mechanically stable configurations. 17. The sensor assembly of claim 16, wherein the unitary sensor assembly is configured to move between the at least two mechanically stable configurations when a resistance provided at least partially by the resistance-providing component is overcome. 18. The sensor assembly of claim 1, wherein the sensor assembly is fluid tight. 19. The sensor assembly of claim 1, wherein the coating is chemically bonded to the frame. 20. The sensor assembly of claim 1, comprising a pulse oximeter monitor configured to receive data from the unitary sensor assembly. 21. The sensor assembly of claim 1, comprising a multi-parameter monitor conf... | Nellcor Puritan Bennett LLC | A61B005/145
US07668581 | B2 | 2010-02-23 | Biomagnetic measurement apparatus | 2005-07-19 | A CAM (current arrow map) 71 and another CAM 72 are obtained from magnetocardiogram waveforms measured from both front and back sides of a subject using data at a point of time of an R-wave peak, then the coordinates of the CAM 72 are inverted in both x and y axis directions to obtain a CAM 73. The coordinates of each of the CAMs 71 and 73 are converted to polar coordinates to obtain CAMs 74 and 75 at both front and back sides of the subject. After that, the CAM 74 is adjusted to the CAM 76 in scale to obtain a CAM 76 , and CAMs 74 and 76 are combined. Then, the CAM data at measuring points on each measuring-points-missing radius vector is inserted through an arithmetic processing by interpolation so that CAM data at every measuring point on a planispheric chart 77 are connected to each another consecutively. | 1. A biomagnetic measurement apparatus, comprising: a plurality of magnetic field sensors disposed two-dimensionally to detect a magnetic field from an inspection object respectively;an arithmetic processing unit for collecting said magnetic fields to be measured from first and second directions of said inspection object to process their data arithmetically; anda display device for displaying a result of said arithmetic processing,wherein said arithmetic processing is executed to convert each of the positions of said plurality of magnetic field sensors to polar coordinates represented by a radius vector r and an oblique angle respectively within a first range in which said oblique angle θ is 0°&lt;θ&lt;180° when detecting said magnetic fields from said first direction and each of the positions of said plurality of magnetic field sensors to polar coordinates within a second range in which said oblique angle θ is 180°&lt;θ&lt;360° when detecting said magnetic fields from said second direction, then set said first magnetic field data to correspond to said polar coordinates within said first range and said second magnetic field data to correspond to said polar coordinates within said second range, and combine data of said first and second magnetic fields to obtain data of a third magnetic field at polar coordinates within a third range in which said oblique angle θ is θ=0° and θ=180°. 2. The biomagnetic measurement apparatus according to claim 1, wherein said arithmetic processing for combining magnetic field data includes a first interpolating processing to obtain data of said third magnetic field at said polar coordinate within said third range in which said oblique angle θ is θ=0° with use of said first and second magnetic field data at said oblique angle θ that gets closest to θ=0° and a second interpolating processing to obtain said third magnetic field data at said polar coordinate within said third range in which said oblique angle θ is θ=180° with use of data of said first and second magnetic fields at said oblique angle θ that gets closest to θ=180°. 3. The biomagnetic measurement apparatus according to claim 2, wherein a weight coefficient is used for both of said first and second interpolating processings respectively. 4. A biomagnetic measurement apparatus, comprising: a plurality of magnetic field sensors disposed two-dimensionally to detect a magnetic field from an inspection object respectively;an arithmetic processing unit for collecting said magnetic fields to be measured from first and second directions of said inspection object as first and second magnetic field data to process their data arithmetically; anda display device for displaying a result of said arithmetic processing,wherein said arithmetic processing is executed to convert each of the positions of said plurality of magnetic field sensors to polar coordinates represented by a radius vector r and an oblique angle respectively within a first range in which said oblique a... | Hitachi High-Technologies Corporation | A61B005/00
US07657298 | B2 | 2010-02-02 | System, device, and method for determining a position of an object | 2004-03-11 | A system is disclosed for determining a position and a change in the position of an anatomical structure. The system utilizes a surgical navigation system and a substrate that is capable of being removably mounted to an outer surface of a patient's body. The substrate includes a sensor that is tracked by the surgical navigation system and a positional device that determines the position of an anatomical structure relative to the sensor. The concatenation of the position of the sensor and the relative position of the anatomical structure allows a global position of the anatomical structure to be determined by a computer system and displayed to the user. | 1. A system for determining a position and a change in the position of an anatomical structure, comprising: a surgical navigation system;a substrate including means for removably attaching the substrate to an outer surface of a body, wherein the body includes an anatomical structure;a sensor attached to the substrate that can be tracked by the surgical navigation system to determine a position of the sensor;a magnetic transmitter attached to the substrate;a magnetic sensor comprising means for attachment to the anatomical structure,wherein the magnetic transmitter and the magnetic sensor are utilized to determine a position of the anatomical structure relative to the sensor;a first circuit for calculating a global position of the anatomical structure by concatenating the position of the sensor and the position of the anatomical structure relative to the sensor; anda second circuit for displaying the global position of the anatomical structure on a display unit. 2. The device of claim 1, wherein the sensor is an optical tracking device. 3. The device of claim 2, wherein the optical tracking device includes three LED's. 4. The device of claim 1, wherein the anatomical structure is a bony structure. 5. The device of claim 1, wherein the means for attachment includes an anchor. 6. The device of claim 1, wherein the magnetic sensor is connected to a retrieval device. 7. The device of claim 6, wherein the retrieval device comprises a wire. 8. The device of claim 6, wherein the retrieval device is attached to the substrate. 9. The device of claim 1, further comprising a plurality of magnetic sensors, wherein each of the magnetic sensors comprises means for attachment to the anatomical structure, and wherein the magnetic transmitter and the plurality of magnetic sensors are utilized to determine the position of the anatomical structure. 10. A method for determining a position and a change in the position of an anatomical structure using a surgical navigation system, the method comprising the steps of: providing a surgical navigation system;attaching a substrate in a removable manner to an outer surface of a body, the substrate having an associated sensor and having a positional device for determining a position of the anatomical structure relative to the sensor, where in the positional device includes a magnetic transmitter attached to the substrate and a magnetic sensor comprising means for attachment to the anatomical structure, and wherein the body includes an anatomical structure spaced interiorly from the outer surface, the magnetic transmitter and the magnetic sensor are utilized to determine a position of the anatomical structure relative to the sensor, and the sensor is tracked by the surgical navigation system to determine a position of the sensor;determining the global position of the anatomical structure by concatenating the position of the sensor and the position of the anatomical structure relative to the sensor;tracking the positio... | Stryker Leibinger GmbH & Co. KG | A61B005/05
US07657304 | B2 | 2010-02-02 | Imaging device for radiation treatment applications | 2002-10-05 | A radiotherapy clinical treatment machine is described having a rotatable gantry and an imaging device with articulating robotic arms to provide variable positioning and clearance for radiation treatment applications. According to one aspect of the invention, a first and a second robotic arms are pivotally coupled to the rotatable gantry, allowing the robotic arms to maneuver independently from the rotatable gantry. | 1. An apparatus, comprising: a rotatable gantry including a treatment source;a first arm coupled to the rotatable gantry at a first coupling point, and the arm having a pivot point separate from the first coupling point;an imaging source coupled to a source end of the first arm;a different second arm coupled to the rotatable gantry at a different second coupling point; andan image detector coupled to the imaging end of the second arm. 2. The apparatus of claim 1, wherein the first arm and the second arm extend and maneuver independent from the rotatable gantry. 3. The apparatus of claim 1, wherein the first arm and the second arm extend independently to an imaging position from a retracted position. 4. The apparatus of claim 3, wherein the first arm and the second arm swing laterally into the imaging position from the retracted position. 5. The apparatus of claim 3, wherein the first arm and the second arm swing downward into the imaging position from the retracted position. 6. The apparatus of claim 1, wherein the imaging source is at a different distance from the gantry than the image detector, wherein the imaging source is pivotally positioned to target the image detector, and the image detector is pivotally positioned to capture an imaging beam from the imaging source. 7. The apparatus of claim 1, wherein the image detector provides target volume information. 8. The apparatus of claim 7, further comprising: a computer to position a treatment beam based on the target volume information. 9. The apparatus of claim 1, further comprising: a computer having a mechanical flexure feedback control component to correct a mechanical flexure of the imaging source and a mechanical flexure of the imager. 10. The apparatus of claim 1, wherein the pivot point is at the first coupling point. 11. The apparatus of claim 1, wherein the pivot point pivotally couples the imaging source to the source end of the first arm. 12. The apparatus of claim 1, wherein the pivot point is between the first coupling point and the source end of the first arm. 13. The apparatus of claim 1, wherein the first coupling point for the first arm is a first pivot point, and the pivot point is a second pivot point. 14. The apparatus of claim 13, wherein the different second coupling point for second arm is a third pivot point. 15. The apparatus of claim 1, wherein the second arm is separate from the rotatable gantry. 16. A method to perform a clinical treatment, comprising: independently positioning a first arm and a second arm with respect to where the first and second arm are coupled to a rotatable gantry to receive target volume data, the first arm coupled to the rotatable gantry at a first coupling point, and the first arm having a pivot point separate from the first coupling point, the second arm coupled to the rotatable gantry at a second coupling point, wherein the first arm is positioned from a clearance position and wherein the gantry includes a treatment source;... | Varian Medical Systems, Inc. | A61B005/05 A61B006/00 A61N005/10 H05G001/02
US07668579 | B2 | 2010-02-23 | System and method for the detection of physiologic response to stimulation | 2006-02-10 | The disclosed embodiments relate to a system and method for detecting a response of a patient to an external stimulus. An exemplary method comprises the acts of generating at least one time series of at least one physiologic parameter, detecting the external stimulus, detecting at least one response associated with the stimulus, the response comprising at least one of a pattern and a value, and outputting an indication of the response. | 1. A method for detecting a response of a patient to an external stimulus, comprising: generating a signal indicative of at least one time series of at least one physiologic parameter with an oximeter;detecting the external stimulus and identifying at least one type of stimulus event by analyzing the signal with a processor;detecting at least one response associated with the stimulus using the processor, the response comprising at least one of a pattern and a value; andoutputting an indication of the response from the processor. 2. The method according to claim 1, comprising receiving a signal corresponding to the physiologic parameter from the oximeter at the processor. 3. The method according to claim 1, wherein the at least one response corresponds to a peak value. 4. The method according to claim 1, wherein the value is indicative of a magnitude value. 5. The method according to claim 1, wherein the value is indicative of a magnitude of rise in oxygen saturation. 6. The method according to claim 1 wherein the external stimulus is auditory. 7. The method according to claim 1, wherein the external stimulus comprises an alarm. 8. The method according to claim 7, wherein the value is indicative of a peak value of oxygen saturation that follows the alarm. 9. The method according to claim 1, wherein the value is indicative of a peak value of oxygen saturation that follows the external stimulus. 10. The method according to claim 1, wherein the value is indicative of a magnitude of rise in oxygen saturation that follows the external stimulus. 11. The method according to claim 1, wherein the physiologic parameter is related to a plethesmographic pulse and the value is indicative of a change in a component of the plethesmographic pulse that follows the stimulus. 12. The method according to claim 1, wherein physiologic parameter is related to a plethesmographic pulse and the value is indicative of a rise in the plethesmographic pulse amplitude that follows the stimulus. 13. The method according to claim 1 wherein the external stimulus comprises an alarm due to a fall in oxygen saturation. 14. The method according to claim 1, wherein the value is indicative of a magnitude of rise in oxygen saturation that follows a fall in oxygen saturation. 15. The method according to claim 1, wherein the value is indicative of a peak value of oxygen saturation that follows a fall in oxygen saturation. 16. The method according to claim 1, wherein the value is indicative of a slope of rise in oxygen saturation that follows a fall in oxygen saturation. 17. The method according to claim 1, wherein the value expresses a relationship between two other values. 18. The method according to claim 1, wherein the value expresses a relationship between a fall magnitude and a rise magnitude. 19. The method according to claim 1, wherein the value expresses a relationship between a fall duration and a rise duration. 20. The method according to claim 1, wherein the v... | - | A61B005/1455
US07666146 | B2 | 2010-02-23 | Apparatus, computer system and computer program for determining intrathoracic blood volume and other cardio-vascular parameters | 2005-04-22 | Apparatus, computer system, computer program and storage medium for determining intrathoracic blood volume (ITBV) and other cardio-vascular parameters of a patient by thermodilution measurements. The apparatus includes: temperature influencing means for introducing a traveling temperature deviation to a patient's blood stream, a temperature sensor device for measuring the local temperature of the patient's blood downstream of the temperature influencing means, a computer system adapted to: record local blood temperature measured as a function of time to determine a thermodilution curve, determine global enddiastolic blood volume (GEDV) and intrathoracic thermovolume (ITTV) from the thermodilution curve, determine intrathoracic blood volume (ITBV) as a function of GEDV, ITTV and the airway pressure inside the patient's lungs (P). By determining ITBV not only from GEDV but also a function of ITTV and P enhanced accuracy can be obtained, especially if the patient suffers from pulmonary edema and/or is mechanically ventilated. | 1. An apparatus for determining intrathoracic blood volume (ITBV) and other cardio-vascular parameters of a patient by thermodilution measurements comprising: a) temperature influencing means for provoking an initial local temperature change in the proximity of a first place of a patient's vascular system, thus introducing a traveling temperature deviation to patient's blood stream,b) a temperature sensor device for measuring the local temperature of patient's blood at a second place of patient's vascular system downstream of said first place,c) means for determining a pressure at least approximately representing an airway pressure of said patient,d) a computer system coupled to said temperature sensor device and said means for determining said pressure and adapted to record said patient's local blood temperature measured at said second place as a function of time to determine a thermodilution curve,e) said computer system being further adapted to determine patient's global enddiastolic blood volume (GEDV) and patient's intrathoracic thermovolume (ITTV) from said thermodilution curve,f) said computer system being further adapted to determine patient's intrathoracic blood volume (ITBV) according to the following formula: ITBV=f(GEDV,ITTV,P),ITBV being the intrathoracic blood volume, GEDV being the global enddiastolic blood volume, ITTV being the intrathoracic thermovolume, P being said airway pressure inside patient's lungs. 2. The apparatus according to claim 1, wherein said function f(GEDV, ITTV, P) is selected to be f(GEDV,ITTV,P)=a·GEDV+b+c·ITTV+d·P a being a species dependent parameter, with 1&lt;a&lt;2b being a species dependent parameter, including zeroc being a species dependent parameter, with c≦0d being a species dependent parameter, including zero,with the restriction that c and d may not be zero simultaneously. 3. The apparatus according to claim 1, wherein said function f(GEDV, ITTV, P) is selected to bef⁡(GEDV,ITTV,P)=ac⁢⁢(ITTV-GEDV)(ITTVnorm-GEDVnorm)+d⁢⁢PPnorm⁢GEDV+ba, b, c, and d being species dependent parameters, with 1&lt;a/(c+d)&lt;2, wherein ITTVnorm, GEDVnorm and Pnorm are empirical normal values of ITTV, GEDV and P, respectively. 4. The apparatus according to claim 1, wherein said function f(GEDV, ITTV, P) is selected to beITBV=a·GEDV[c2·(ITTV-GEDV)(ITTVnorm-GEDVnorm)+1]·[d2·PPnorm+1]+b+c1·(ITTV-GEDV)(ITTVnorm-GEDVnorm)+d1·PPnorma, b, c1, c2, d1, d2 being species dependent parameters, with 0.5≦a/(c2+1)(d2+1)≦10, wherein ITTVnorm, GEDVnorm and Pnorm are empirical normal values of ITTV, GEDV and P, respectively. 5. The apparatus according to claim 1 wherein P is set equal to a transmural lung pressure Ptm, being defined as Ptm=ITP−Pmv, ITP being a intrathoracic pressure and Pmv being a microvascular pressure. 6. The apparatus according to claim 1, wherein P is a pressure measured in the airway of a mechanical respirator. 7. The apparatus according to claim 6, wherein P is a positive end expiratory pressure (PEEP). 8. T... | Pulsion Medical Systems AG | A61B005/02
US07647098 | B2 | 2010-01-12 | System and method for prediction of cognitive decline | 2005-10-31 | A system and method for prediction of cognitive decline, comprises an input receiving input data corresponding to brain activity of an individual and a processor coupled to the input for analyzing the input data to obtain a selected set of features, the processor comparing the selected set of features to at least a portion of entries in a database corresponding to brain activity of a plurality of individuals, wherein entries in the database have been separated into a plurality of categories corresponding to one of a degree of cognitive decline and a propensity for future cognitive decline of individuals relating to the entries, the processor determining, based on the comparison, a category most closely corresponding to the selected set of features. | 1. A system for prediction of cognitive decline of a patient during a future time period, comprising: an input receiving input data corresponding to brain activity of a patient; anda processor coupled to the input for analyzing the input data to obtain an input set of values for a selected set of features, the processor being configured to compare the input values to a comparison set of values for the selected set of features derived from at least a portion of entries in a database corresponding to brain activity of a plurality of individuals, wherein entries in the database have been separated into a plurality of categories corresponding to a degree of cognitive decline of the individuals when the data was collected and a degree of cognitive decline of the individuals during a period of time subsequent to the collection of the data, the processor determining, based on the comparison, a category most closely corresponding to the selected set of features. 2. The system according to claim 1, wherein the input data includes at least one of quantitative electroencephalogram data, quantitative magnetoencephalogram data and quantitative event related potential data corresponding to brain activity of the patient. 3. The system according to claim 1, wherein the processor converts the input data to numerical values corresponding to a time series for each of a plurality of regions of the brain from which the input is received. 4. The system according to claim 1, wherein the processor analyzes the input data using one of bi-spectral, spectral analysis, wavelet analysis, chaos analysis, fractal analysis, principal component analysis, independent component analysis, singular value decomposition, microstate distributions, microstate transition probabilities and global field synchrony. 5. The system according to claim 1, wherein the database includes data corresponding to one of the selected set of features, multivariate combinations of the features of the selected set of features from a brain region and a multivariate composite of the selected set of features from a plurality of brain regions. 6. The system according to claim 1, wherein the entries in the database are separated into the plurality of categories using one of descriptive statistics, discriminant analysis, cluster analysis, neural networks, logistic regression and heuristic selection. 7. The system according to claim 1, further comprising means for converting the input data to a series of quantitative values. 8. The system according to claim 7, wherein the means for converting the input data to a series of quantitative values includes one of an analog/digital converter and a digital amplifier. 9. The system according to claim 1, wherein the input data includes one of P300 and mis-match negativity. 10. The system according to claim 1, wherein the input data includes mid-latency auditory EP. 11. The system according to claim 1, wherein the selected set includes a ratio of Theta band brain ... | New York University | A61B005/04
US07645233 | B2 | 2010-01-12 | Powering a guide wire mounted sensor for intra-vascular measurements of physiological variables by means of inductive coupling | 2005-06-07 | The present invention relates to a sensor assembly arranged to be disposed in a body for measuring a physiological variable. A basic idea of the present invention is to provide a sensor and guide wire assembly which communicates with an external unit by an inductive coupling between a first coil provided in the sensor assembly and a second, external coil which is disposed in or connected to the external unit. The external unit is preferably the same unit that is used to numerically or graphically display the physiological variables that are measured with the sensor and guide wire assembly. With the terminology introduced above, it should here be clear that both the signal transfer from the sensor as well as the energy feeding from the external unit to the sensor may be accomplished through the inductive coupling between the first and second coils. | 1. A sensor assembly arranged to be disposed in a body for measuring a physiological variable, the sensor assembly comprising: a sensor element fixedly mounted at a distal portion of a guide wire for measuring the physiological variable and to provide a signal representing the measured physiological variable; anda first coil positioned at the distal portion of the guide wire and arranged to receive energy by inductive coupling with a second coil that is coupled to an energy supply, wherein the second coil and energy supply are adapted to be located outside the body, in order to form a power supply for the sensor element, wherein the second coil is detachably arranged around a core wire at a proximal portion of the guide wire. 2. The sensor assembly in accordance with claim 1, wherein the first coil is a spiral portion available in a guide wire assembly at which the sensor assembly is arranged. 3. The sensor assembly in accordance with claim 1, further being arranged to communicate the signal representing the measured physiological variable via the first coil. 4. The sensor assembly in accordance with claim 1, further comprising: an antenna arranged to communicate the signal representing the measured physiological variable. 5. The sensor assembly in accordance with claim 3, further comprising: a modulator arranged to modulate a carrier signal with the signal representing the measured physiological variable before the modulated carrier signal is communicated. 6. A system comprising the sensor assembly in accordance with claim 1 and a monitoring device, said monitoring device comprising said energy supply. 7. The system of claim 6, wherein said monitoring device further comprises: a demodulator arranged to demodulate the signal representing the measured physiological variable, wherein the monitoring device is further arranged to provide a representation of said measured physiological variable. 8. The sensor assembly in accordance with claim 1, wherein the guide wire has an outer diameter of approximately 0.36 mm. | Radi Medical Systems AB | A61B005/00
US08309941 | B2 | 2012-11-13 | Charged particle cancer therapy and patient breath monitoring method and apparatus | 2009-09-17 | The invention comprises a patient respiration monitoring and/or control method and apparatus used in conjunction with multi-axis charged particle or proton beam radiation therapy of cancerous tumors. The respiration monitoring system uses thermal and/or force sensors to determine where a patient is in a respiration cycle in combination with a feedback signal control delivered to the patient to inform the patient when breath control is required. The resulting breath control is timed with charged particle delivery to the tumor to enhance accuracy, precision, and/or efficiency of tumor treatment. | 1. An apparatus for controlling extraction of charged particles from a synchrotron for irradiation therapy of a tumor of a patient, comprising: a synchrotron configured to accelerate the charged particles, said synchrotron comprising: a radio-frequency cavity system comprising a first pair of blades for inducing oscillation of the charged particles; anda foil configured to yield slowed charged particles from the charged particles having sufficient oscillation to traverse said foil, wherein the slowed charged particles pass through a second pair of blades having an extraction voltage directing the charged particles out of said synchrotron through a Lamberson extraction magnet, said radio-frequency cavity system for inducing oscillation configured and timed using said respiration signal;a respiration sensor generating a respiration signal, said respiration signal corresponding to a variable rate breathing pattern of the patient, said respiration sensor comprising at least one of: a force meter strapped to the patient's chest;a first thermal resistor configured in a first position proximate an exhalation path of the patient; anda second thermal resistor configured in a second position both out of the exhalation path of the patient and in a local room environment common with the patient, wherein said respiration signal is generated using differences between readings from said first thermal resistor and said second thermal resistor,wherein said radio-frequency cavity system for inducing oscillation is timed using said respiration signal to time extraction of the charged particles to the tumor based on said variable rate breathing pattern. 2. The apparatus of claim 1, further comprising: a charged particle controller controlling delivery of the charged particles from the synchrotron at sequential non-periodic intervals, said non-periodic intervals correlating with said variable rate breathing pattern, wherein said breathing pattern comprises sequential respiration cycles varying by at least one second in length. 3. The apparatus of claim 1, further comprising: an X-ray generation source configured to generate X-rays of the patient and the tumor at said non-periodic intervals. |  | H05H007/10
US08309357 | B2 | 2012-11-13 | Analyte-testing device | 2008-01-22 | A device adapted to determine an analyte concentration of a fluid sample using a test sensor. The device comprises a display adapted to display information to a user. The device further comprises at least one user-interface mechanism adapted to allow the user to interact with the device. The device further comprises a body portion including at least one opening formed therein, the at least one opening being of sufficient size to receive the test sensor. The device further comprises a memory adapted to store a plurality of stored analyte concentrations. The device further comprises a processing feature adapted to inhibit the stored analyte concentrations from being displayed on the display. | 1. A method of using a device adapted to determine an analyte concentration of a fluid sample using a test sensor, the method comprising the acts of: providing a device comprising a display, a body portion including at least one opening formed therein, at least one user-interface mechanism positioned on one of the display and body portion, and a memory adapted to store a plurality of stored analyte concentrations;interacting with the at least one user-interface mechanism, the at least one user-interface mechanism being adapted to assist in displaying information regarding a testing procedure and at least one of the plurality of stored analyte concentrations on the display;determining whether a predetermined condition exists, the predetermined condition including a test sensor being positioned in the at least one opening; andif the predetermined condition exists, preventing the plurality of stored analyte concentrations from being displayed on the display during the testing procedure. 2. The method of claim 1, further comprising the act of, if the predetermined condition does not exist, displaying the at least one of the plurality of stored analyte concentrations on the display. 3. The method of claim 1, wherein the device further includes a software component adapted to inhibit the plurality of stored analyte concentrations from being displayed on the display. 4. The method of claim 1, wherein the act of the preventing the plurality of stored analyte concentrations from being displayed on the display includes deactivating the user-interface mechanism. 5. The method of claim 1, further comprising performing a predetermined act, the predetermined act assisting in displaying at least one of the plurality of stored analyte concentrations on the display. 6. The method of claim 1, further comprising the act of, upon removing the test sensor from the at least one opening, allowing the plurality of stored analyte concentrations to be displayed on the display. 7. The method of claim 1, wherein the device further includes a detector positioned proximate to the at least one opening. 8. The method of claim 7, wherein the act of determining whether a predetermined condition exists includes the detector detecting that the test sensor is positioned in the at least one opening. | Bayer HealthCare, LLC | G01N015/06 G01N033/00 G01N033/48
US08285353 | B2 | 2012-10-09 | System for analyzing tissue perfusion using concentration of indocyanine green in blood | 2007-07-11 | The present invention relates, in general, to a system for analyzing tissue perfusion using the concentration of indocyanine green and a method of measuring the perfusion rate using the system and, more particularly, to a system for measuring tissue perfusion by injecting indocyanine green into a living body, detecting variation in the concentration of indocyanine green with the passage of time, and analyzing the detected variation, and a method of measuring the perfusion rate using the system. The present invention provides a method of measuring perfusion in a living body, which enables accurate measurement for respective regions in a wide range from a perfusion rate decreased to less than 10% of normal perfusion to a perfusion rate increased to greater than normal perfusion using the above-described mechanism of ICG in a living body, which cannot be conducted using the conventional technology. | 1. A computer for measuring a rate of tissue perfusion using time-series analysis of fluorescence images of indocyanine green (ICG), comprising: 1) an input module configured to receive signals from a photodetector,2) a numerical conversion module configured to process input signals from the input module into numerical values, said input signals comprising intensities of fluorescence over time in a region of interest,3) a rate calculation module configured to calculate rates of tissue perfusion with using the numerical values and an equation representing dynamics of the ICG, and4) an output module configured to output results of the calculation,wherein the equation representing dynamics of the ICG is eitherFIisc=PAP-1⁢/⁢τ⁢(ⅇ-tτ-ⅇ-Pt)⁢⁢or⁢-In⁢⁢τ-Tmaxτ=InP-PTmax,wherein the FIisc (Fluorescence Intensity ischemia) is an intensity of ICG fluorescence of a measurement target tissue, the P is a rate of tissue perfusion, the A is an intensity of fluorescence obtained from ICG image during the first 1 minute, the τ is t1/2/ln2, the t1/2 is a time at which the intensity of ICG fluorescence is half of a highest value and the Tmax is a time at which the intensity of ICG fluorescence in an ischemic tissue is highest. 2. An apparatus for measuring tissue perfusion, comprising: a stand configured to allow light to pass therethrough, a light source adjustable to be located to radiate light onto an ICG-injected tissue disposed on the stand, a filter adjustable to be located to filter only near infrared ray wavelengths of 800 to 850 nm from fluorescence signals emitted from the living body under the action of the light source, a detector configured to detect fluorescence light passed through the filter, and the computer of claim 1 operably connected to the detector and configured to image light detected by the detector and measure the rates of tissue perfusion through the analysis of the ICG dynamics with the passage of time. 3. A method of measuring a rate of tissue perfusion, comprising: 1) detecting sequentially a concentration of ICG until the concentration of ICG is decreased to a lowest level by injecting the ICG into a living body, radiating light onto the living body using a light source, and measuring intensities of ICG fluorescence generated in the living body with the passage of time after the injection using a photodetector;2) processing the sequential concentration of ICG in the living body detected in the step 1) into numerical data by converting the intensities of ICG fluorescence into numerical values for respective regions and respective times; and3) calculating the rates of perfusion from the numerical values,wherein the perfusion rates are calculated using the following equation:FIisc=PAP-1⁢/⁢τ⁢(ⅇ-tτ-ⅇ-Pt)⁢⁢or⁢-In⁢⁢τ-Tmaxτ=InP-PTmax,wherein the FIisc (Fluorescence Intensity ischemia) is an intensity of ICG fluorescence of a measurement target tissue, the P is a rate of tissue perfusion, the A is an intensity of fluorescence obtained from ICG ... | Korea Advanced Institute of Science and Technology | A61B005/00
US08090436 | B2 | 2012-01-03 | Tissue discrimination and applications in medical procedures | 2009-10-29 | A system and method for discriminating tissue types, controlling the level of therapy to tissue, and determining the health of a known tissue by measuring the characteristics of an electrical signal applied to a conductive element located within or by the tissue. Additionally, the system and method may be used for determining whether the conductive tip of a pedicle probe or pedicle screw is located in one of cortical bone, cancellous bone, and cortical bone near a boundary with soft tissue, whether the conductive tip of a cannula is located adjacent to one of nerve tissue and annulus tissue, and whether the conductive tip of a cathode is located adjacent to one of nerve tissue and prostate gland tissue. | 1. A system for preventing violation of a cortical wall of a pedicle during spine surgery, said pedicle including a cancellous core within said cortical wall, said system comprising: (a) a screw configured for advancement and purchase into said pedicle and having a non-conductive shaft and a conductive tip; and(b) a processor in electrical communication with the screw, the processor configured to direct application of an electrical signal to the conductive tip and determine at least one of the capacitance and impedance characteristics of the tissue adjacent to said conductive tip to determine whether said conductive tip lies within or adjacent cortical bone, cancellous bone, and cortical bone near the boundary with soft tissue. 2. The system of claim 1, wherein said conductive tip comprises a bi-polar electrode. 3. The system of claim 1, wherein the determined characteristics of the tissue further include a phase angle relationship between a voltage and current of the applied electrical signal. 4. The system of claim 1, wherein the applied electrical signal has a range of predetermined frequencies. 5. The system of claim 4, wherein the range of predetermined frequencies is from 400 KHz to 100 MHz. 6. The system of claim 1, wherein the applied signal has a fixed frequency. 7. The system of claim 1, wherein the applied electrical signal has a sliding frequency. 8. The system of claim 1, wherein the applied signal is a sinusoidal signal. 9. The system of claim 1, wherein the applied signal is a square wave signal. 10. The system of claim 1, further comprising an output element for communicating feedback regarding said measured characteristics. 11. The system of claim 1, further comprising an input device for receiving instructions from a user. 12. The system of claim 1, wherein said screw further comprises a conductive head. 13. A method for preventing violation of a cortical wall of a pedicle during spine surgery, said pedicle including a cancellous core within said cortical wall, comprising the steps of: (a) advancing a pedicle screw into at least one of said cancellous core and said cortical wall of said pedicle during spine surgery, said pedicle screw having a conductive tip for measuring impedance;(b) delivering an electrical signal to said conductive tip during step (a);(c) measuring the impedance of at least one of said cancellous core and said cortical wall during step (a); and(d) stopping the advancement of said pedicle screw based on a change in said measured impedance from step (c)to prevent a violation of said cortical wall by the continued advancement of said pedicle screw. 14. The method of claim 13, further comprising the step of communicating feedback to said user. 15. The method of claim 13, wherein the delivery of said electrical signal of step (b) and the measuring of impedance of step (c)are performed by a processor in electrical communication with said pedicle screw. 16. The method of claim 14, wherein the communi... | NuVasive, Inc. | A61B005/053 A61B017/86
US07957794 | B2 | 2011-06-07 | System and method for transducing, sensing, or affecting vaginal or body conditions, and/or stimulating perineal musculature and nerves using 2-way wireless communications | 2009-08-17 | A system for sensing vaginal conditions or for affecting vaginal or body conditions in a mammal. The system includes a portable probe unit and a separate controller unit, both of which include two-way communication devices to both transmit signals to, and receive signals from, each other, wirelessly and in real time. The substantially smooth and sealed outer surface of the cylindrical body of the probe unit has a rounded end and is dimensioned to permit comfortable and repeated insertion into, removal from, and containment entirely within a mammal's vagina. The probe body has a device substantially flush with its outer surface and configured to deliver electrical pulses, receive electrical pulses, sense or sample vaginal conditions and/or deliver medication. Signals to the probe unit comprise control and programming signals acting as a feedback loop configured to start, stop and/or alter activity of the substantially flush device of the probe unit, either automatically or manually, for sensing vaginal conditions or for affecting vaginal or body conditions. | 1. A system for sensing vaginal conditions or for affecting vaginal or body conditions in a mammal, comprising: a portable probe unit, said probe unit comprising a substantially cylindrical body having a substantially smooth and substantially sealed outer surface with a rounded end and so dimensioned as to permit comfortable and repeated insertion into, removal from, and containment entirely within a mammal's vagina; means substantially flush with the outer surface of the body of the probe and configured to accomplish at least one of delivering electrical pulses, receiving electrical pulses, sensing or sampling vaginal conditions, and delivering medication; a programmable microprocessor; a memory; a battery; and two-way communication means with antenna and adapted to both transmit signals to a controller unit and receive signals from said controller unit wirelessly and in real time, even when said probe unit is contained entirely within a mammal's vagina; andthe controller unit comprising two-way communication means adapted to both receive signals from said probe unit and transmit signals to said probe unit wirelessly and in real time, even when said probe unit is contained entirely within a mammal's vagina, wherein said signals to said probe unit comprise control and programming signals acting as a feedback loop configured to start, stop, and/or alter the activity of said substantially flush means of said probe unit, either automatically or manually, for sensing vaginal conditions or for affecting vaginal or body conditions. 2. A system according to claim 1, wherein said substantially flush means of said probe unit comprises means for sampling cervical fluid or other changes in the vaginal environment. 3. A system according to claim 1, wherein said substantially flush means of said probe unit comprises means for sensing at least one of temperature, pH, secretion viscosity, vaginal pathogens and atypical cervical cells. 4. A system according to claim 1, wherein said substantially flush means of said probe unit includes at least one of a conductive band and a sensor transducer. 5. A system according to claim 1, wherein said probe unit is adapted to be programmed to start, stop and/or alter at least one of sensing vaginal conditions, sampling vaginal conditions and delivering medication after a predetermined period of time. 6. A system according to claim 1, wherein said probe unit is less than one inch in diameter and less than four inches long. 7. A system according to claim 1, wherein said probe unit is provided with means for facilitating removal from a vagina. 8. A system according to claim 1, wherein said system is adapted to permit a human user to operate said controller unit while said probe unit is in the human user's vagina. 9. A system according to claim 1, wherein said two-way communication means of at least one of said probe unit and said controller unit includes means for transmitting signals to and/or receiving signals from an... | Athena Feminine Technology | A61B005/04 A61B005/05 A61B005/00 A61B010/00 A61B005/103 A61B005/117
US09636051 | B2 | 2017-05-02 | Detection meter and mode of operation | 2009-06-08 | A method for performing an assay to determine the presence or concentration of an analyte contained in a sample of body fluid by using a device comprising at least one analyte quantification member and a sensor associated therewith, the method includes: applying a first sample to the analyte quantification member; and detecting the presence or absence of an adequate sample volume; wherein upon detection of the absence of an adequate sample volume, initiating a finite timed period, and signaling the user to introduce a second sample of body fluid to the analyte quantification member. Associated arrangements and devices are also disclosed. | 1. A device for performing an assay to determine the presence or concentration of an analyte contained in a sample of body fluid, the device comprising: at least a first analyte quantification member;at least one passageway in fluid communication with the first analyte quantification member;a sensor configured to interrogate at least the first analyte quantification member and to detect the presence or absence of an adequate sample volume;a vacuum source configured to assist in the expression of body fluid from a wound;a non-transitory memory device programmed with instructions to: activate the vacuum source;initiate a first time period and interrogate the first analyte quantification member using the sensor to detect the presence or absence of an adequate sample volume applied to the first analyte quantification member within the first time period; andupon the detection of the absence of an adequate sample volume, initiate a second time period and signal a user to manually milk the wound and introduce an additional sample of body fluid into the at least one passageway to increase the sample volume applied to the first analyte quantification member; anda processor programmed to carry out the instructions on the non-transitory memory device. 2. The device of claim 1, wherein the non-transitory memory device further comprises programmed instructions to interrogate the first analyte quantification member using the sensor during the second time period to detect the presence or absence of an adequate sample volume. 3. The device of claim 2, wherein the non-transitory memory device further comprises programmed instructions to signal the user that the assay is being terminated upon detection of the absence of an adequate sample volume within the second time period. 4. The device of claim 2, wherein the non-transitory memory device further comprises programmed instructions to continue with the assay upon detection of an adequate sample volume during the second time period. 5. The device of claim 4, wherein the non-transitory memory device further comprises programmed instructions to interrogate the first analyte quantification member using the sensor subsequent to the second time period to ascertain the presence or concentration of the analyte. 6. The device of claim 1, wherein the analyte comprises glucose and the body fluid comprises blood. 7. The device of claim 1, wherein the first analyte quantification member comprises an assay pad containing a chemical reagent formulated to produce a color change upon reaction with the analyte. 8. The device of claim 1, wherein the device comprises a plurality of the analyte quantification members. 9. The device of claim 8, wherein the device comprises a removable cartridge containing the plurality of the analyte quantification members. 10. The device of claim 7, wherein the sensor comprises an optical sensor configured to produce signals indicative of the color change. 11. The device of claim 1, wherein... | Intuity Medical, Inc. | G01N033/00 A61B005/145 A61B005/15 A61B005/151 A61B005/157 A61B005/1455 A61B005/1468
US08234131 | B2 | 2012-07-31 | System and method for chronic illness care | 2009-07-14 | In one aspect, the present invention relates to a computerized system programmed for providing care support to at least one patient having at least one chronic illness. In one embodiment, the system includes a medical professional module adapted for receiving, storing, and providing data in communication with at least one medical professional at the point of care, a health coach module adapted for receiving, storing, and providing data in communication with the at least one patient and at least one health coach, a patient module adapted for receiving, storing, and providing data in communication with the at least one patient, and, a public health module adapted for receiving, storing, and providing data in communication with at least one research professional. Each of the medical professional module, health coach module, patient module, and public health module is operatively associated with a corresponding one of the at least one medical professional, at least one health coach, at least one patient, and at least one research professional through a network. | 1. A computerized system programmed for providing care support to at least one patient having at least one chronic illness, comprising a server comprising: (a) a medical professional module adapted for receiving, storing, and providing data in communication with at least one medical professional at the point of care, wherein the medical professional module comprises at least one algorithm for generating at least one care decision for the at least one patient at the point of care, wherein the at least one algorithm is based at least in part on at least one of disease-specific national diagnostic guidelines, disease-specific national treatment guidelines, patient-specific medication plans, and clinical care pathways;(b) a health coach module adapted for receiving, storing, and providing data in communication with the at least one patient and at least one health coach;(c) a patient module adapted for receiving, storing, and providing data in communication with the at least one patient; and(d) a public health module adapted for receiving, storing, and providing data in communication with at least one research professional, wherein the public health module is adapted for generating cost-benefit information associated with at least one treatment decision for a chronic illness selected by at least one of the at least one medical professional and at least one health coach,wherein the server and each of the medical professional module, health coach module, patient module, and public health module are operatively associated with a corresponding one of the at least one medical professional, the at least one health coach, the at least one patient, and the at least one research professional through a network computing system to perform for interacting with at least one user of the system programmed for providing care support to at least one patient having at least one chronic illness. 2. The system of claim 1, wherein the system is further programmed for performing at least one of the acts of: (a) providing data comprising at least one of information from electronic medical records associated with the at least one patient, information associated with at least one treatment compliance program for treating the at least one chronic illness, and information associated with the clinical health condition of the at least one patient;(b) receiving an alert when a proposed medication or proposed medication dosage conflicts with a current medication or current medication dosage associated with the at least one patient;(c) receiving data for supporting patient treatment compliance to chronic illness-specific treatment guidelines selected by at least one of the at least one medical professional and the at least one health coach;(d) providing data through the network from at least one of a health monitoring device, mobile electronic device, and user computer;(e) tracking use of patient-specific curriculum by the at least one patient; (g) tracking communication between ea... | Morehouse School of Medicine | G06Q100/00 G06Q050/00 A61B005/00 G06F019/00
US08255037 | B2 | 2012-08-28 | Cardiac roadmapping | 2008-02-25 | According to an exemplary embodiment of the present invention, a cardiac roadmapping technique is provided, in which a static roadmap is built onto which a view of the intervention device is projected at the correct location. This new viewing mode combines the advantages of both roadmapping, i.e. accurate localization, and the advantages of a motion-free view, i.e. maximum readability. | 1. A method for cardiac roadmapping for examination of an object of interest located in a volume of interest with an examination apparatus, the method comprising the steps of: creating an object of interest map sequence of the volume of interest on the basis of a first image sequence of the volume of interest;creating a device map sequence of the volume of interest on the basis of a second image sequence of the volume of interest;calculating a device map to object of interest map motion estimation to maximize the similarity between the device map sequence and the object of interest map sequence;creating a static reference map of the volume of interest on the basis of a first image sequence of the volume of interest, wherein the reference map is a static roadmap of the object of interest;calculating a reference motion estimation to bring the object of interest map sequence into registration with the static reference map;performing a motion composition to bring the device map sequence into registration with the static reference map, resulting in a navigation sequence; anddisplaying the navigation sequence as an image on a display device. 2. The method of claim 1, wherein the object of interest is a cardiac vessel-tree;wherein the first image sequence are angiograrns of the object of interest;wherein the second image sequence is a life-time sequence acquired during an intervention of the object of interest; andwherein the reference map is a reference vessel-map. 3. The method of claim 1, wherein the motion composition is done by vector interpolation steps, resulting in a sequence of registered device maps. 4. The method of claim 1, wherein the static reference map of the volume of interest is selected by selecting the image of the first image sequence of the volume of interest having maximum visibility, or by selecting the image of the first image sequence of the volume of interest having smallest distance to all others. 5. The method of claim 1, wherein the object of interest map sequence of the volume of interest is subjected to a vessel cycle extraction corresponding to a full cardiac cycle. 6. The method of claim 1, wherein the device map to object of interest map motion estimation comprises a sequence-to-sequence motion estimation and matching operation. 7. The method of claim 1, further comprising a step of superimposing the sequence of registered device maps to the static reference map. 8. A non-transitory computer-readable medium comprising instructions for performing the steps of: creating an object of interest map sequence of the volume of interest on the basis of a first image sequence of the volume of interest;creating a device map sequence of the volume of interest on the basis of a second image sequence of the volume of interest;calculating a device map to object of interest map motion estimation to maximize the similarity between the device map sequence and the object. of interest map sequence;creating a static reference map of... | Koninklijke Philips Electronics N.V. | A61B005/05
US08308652 | B2 | 2012-11-13 | Apparatus, method and system for determining a physiological condition within a mammal | 2009-08-05 | A system, method and apparatus for determining a physiological condition within a mammal based on pH and/or temperature measurements. In one aspect, the invention is directed to an apparatus for measuring a physiological condition within a mammal comprising: a tubular housing extending along a longitudinal axis from a distal end to a proximal end, the housing having an internal cavity; a first opening in the elongated housing forming a first passageway into the internal cavity; an ion-sensitive field effect transistor (ISFET) for measuring pH within a body lumen of the mammal, the ISFET positioned in the internal cavity, the ISFET aligned with the first opening so that at least a portion of the ISFET is exposed via the first opening; and a first seal between the ISFET and the housing forming a hermetic seal about a perimeter of the first opening. | 1. An apparatus for measuring a physiological condition within a mammal comprising: an elongated tubular housing extending along a longitudinal axis from a distal end to a proximal end, the elongated tubular housing having an internal cavity;a first opening in the elongated tubular housing funning a first passageway into the internal cavity;an ion-sensitive field effect transistor (ISFET) for measuring pH within a body lumen of the mammal, the ISFET positioned in the internal cavity, the ISFET aligned with the first opening, so that at least a portion of the ISFET is exposed via the first opening;a first seal between the ISFET and the elongated tubular housing forming a hermetic seal about a perimeter of the first opening;a temperature sensor for measuring temperature within the body lumen;a second opening in the elongated tubular housing forming a second passageway into the internal cavity;the temperature sensor comprising a thermistor covered by a metal cap, the metal cap exposed via the second opening; anda second seal between the metal cap and the elongated tubular housing forming a hermetic seal about a perimeter of the second opening. 2. The apparatus of claim 1 wherein a space exists between the metal cap and the thermistor, the space filled with a thermally conductive and electrically non-conductive epoxy. 3. The apparatus of claim 1 further comprising: a third opening in the elongated tubular housing forming a third passageway into the internal cavity;a diaphragm in the third passageway that is in contact with an electrolyte solution buffered at a known pH stored in a chamber of the internal cavity. 4. The apparatus of claim 3 wherein the first, second and third openings are aligned along an axis that is substantially parallel to the longitudinal axis. 5. The apparatus of claim 4 further comprising a well formed in the outer surface of the housing for holding fluids, the well having a floor, and wherein the first, second and third openings are located in the floor of the well. 6. The apparatus of claim 5 wherein the floor is a substantially planar surface. 7. An apparatus for measuring a physiological condition within a mammal comprising: an elongated tubular housing extending along a longitudinal axis from a distal end to a proximal end; andan ion-sensitive field effect transistor (ISFET) for measuring pH within a body lumen of the mammal, the ISFET located on a distal portion of the elongated tubular housing and having at least a portion of the ISFET exposed for contact with fluid of the body lumen;a memory device located within the elongated tubular housing, the memory device storing parametric data unique to the pH sensor; andwherein the parametric data includes first ISFET slope data determined at ambient temperature and second ISFET slope data determined at normal body temperature. 8. The apparatus of claim 7 wherein the first and second ISFET slope data is based on at least three points of pH parametric data. 9. An apparat... | PH Diagnostics, Inc. | A61B005/00 A61B010/00
US08290561 | B2 | 2012-10-16 | Signal processing for continuous analyte sensor | 2009-09-23 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A method for estimating an analyte value from a continuous analyte sensor, the method comprising: receiving a data stream from the continuous analyte sensor for a first time period, thereby obtaining a measured analyte value;estimating at least one analyte value for a second time period based on the data stream;comparing, using electronic circuitry, the estimated analyte value with the measured analyte value; anddetermining a confidence interval based on the comparing. 2. The method of claim 1, wherein the step of comparing is performed continuously. 3. The method of claim 1, wherein the step of comparing comprises determining a correlation and/or deviation between the estimated analyte value and the measured analyte value. 4. The method of claim 1, further comprising analyzing a variation of one or more estimated analyte values based at least in part on the step of comparing. 5. The method of claim 1, wherein the step of estimating comprises selecting an algorithm from a plurality of algorithms based at least in part on the step of comparing. 6. The method of claim 1, wherein the step of comparing comprises comparing a first matched data pair formed from the measured analyte value and a reference analyte value with a second matched data pair formed from the estimated analyte value and a reference analyte value. 7. The method of claim 1, further comprising a step of providing output based at least in part on the comparing step. 8. The method of claim 7, wherein the output comprises displaying the measured analyte value. 9. The method of claim 7, wherein the output comprises displaying the estimated analyte value. 10. The method of claim 7, wherein the output comprises displaying an estimated analyte value for a future time period. 11. The method of claim 1, wherein the second time period comprises a future time period. 12. The method of claim 1, wherein the second time period is the same as and/or overlaps the first time period. 13. The method of claim 7, further comprising outputting the confidence interval. 14. A system for estimating analyte data from a continuous analyte sensor, the system comprising: an input module operatively connected to the continuous analyte sensor that receives a data stream for a first time period; anda processor module comprising programming configured to estimate at least one analyte value for a second time period based on the data stream, wherein the processor module further comprises programming configured to compare the estimated analyte value with a corresponding measured analyte value and determine a confidence interval therefrom. 15. The system of claim 14, wherein the programming configured to estimate at least one analyte value estimates the at least one analyte value continuously. 16. The system of claim 14, wherein the programming configured to compare the estimated analyte value with a corresponding measured analyte value determines a correlation and/or deviation between the estimated an... | DexCom, Inc. | A61B005/05 A61B005/00
US08538778 | B2 | 2013-09-17 | Methods and systems for integrated health systems | 2009-05-15 | Methods, business methods, and systems are provided herein for integrated healthcare. As the amount of medical information increases rapidly, including information from multiple biomarkers, analysis and management of that information becomes more and more important to extract meaningful conclusions from the information. Statistical and computational methods are described herein that have been created for the methods and systems for integrated healthcare. For example, a computer system is described extracts significance over time of PSA and fPSA biomarker tests for prostate health. | 1. A method of performing a course of medical action for a medical condition of a subject comprising: a) obtaining a first series of PSA values from a subject;b) performing a first dynamic analysis using a computer system, comprising:i) fitting said first series of PSA values to form a first fitted trend;ii) estimating a first characteristic of said first fitted trend, wherein said first characteristic is a first PSA velocity defined as a rate of change in trend PSA; iii) calculating one or more probabilities of the occurrence of one or more medical conditions of said subject using said first characteristic, wherein said one or more medical conditions comprises a prostate medical condition selected from the group consisting of the following: prostatitis due to inflammation prostatitis due to infection prostate cancer, benign prostate hyperplasia, and no prostate disease;c) determining if the results of said first dynamic analysis have reached a first threshold for medical action, wherein said threshold corresponds to a probability that the subject has a medical condition, wherein the probability is based on analysis of PSA velocity population data; andd) performing a first course of medical action for a first medical condition if said results of said first dynamic analysis reaches said threshold, wherein said first course of medical action is treatment with antibiotics and/or anti-inflammatory medication; measurement of prostate volume; continued PSA tests; a free PSA test; a proPSA test; a biopsy, medical imaging; or medical treatment for prostate cancer. 2. The method of claim 1, wherein said subject is a human. 3. The method of claim 1, wherein said computer system comprises a device for network communication, a storage unit, and a processor. 4. The method of claim 3, wherein said computer system comprises a Monte Carlo engine. 5. The method of claim 1, wherein said first dynamic analysis further comprises estimating a second characteristic of said first fitted trend, wherein said second characteristic is a first PSA variation. 6. The method of claim 5, further comprising: e) obtaining a second series of PSA values from said subject, wherein at least one of said second series of PSA values is measured after performing said first course of medical action;f) performing a second dynamic analysis on a computer system, comprising:i) comparing said PSA value measured after performing said first course of medical action to said first series of PSA values and said first fitted trend;ii) estimating a third characteristic of said second series of PSA values, wherein said characteristic is a second PSA velocity or a second PSA variation;g) determining if the results of said second dynamic analysis have reached a second threshold for medical action, wherein said threshold is based on analysis of population data and the results of said first dynamic analysis; andh) performing a second course of medical action for a second medical condition when the re... | Soar BioDynamics, Ltd. | G06Q050/00
US08690769 | B2 | 2014-04-08 | Metabolic energy monitoring system | 2009-04-21 | A metabolic energy monitoring system having one or more physiological measurement platforms and displays enabling the calculation and display of energy balance, kilocalorie energy expenditure and kilocalorie intake is described. In preferred embodiments, the system utilizes one or more on-body monitoring platforms to enable measurement of change in body composition and kilocalorie energy expenditure over a period of time thereby enabling a comparator to calculate net energy balance over this period of time and to calculate kilocalorie intake over this same period of time. Such data may then be displayed on a display device in wireless communication with the on-body monitoring platform to provide the user of the system with useful information and guidance in weight management applications. | 1. A method for determining a metabolic status over a period of time, the method comprising: using a sensor to make at least one determination of a body composition change encompassing said period of time;making at least one determination of a kcal energy expenditure encompassing said period of time; andmaking at least one determination of an energy intake corresponding to said period of time;wherein the determination of the body composition change includes a determination of a fluid mass change, andwherein said determination of the body composition change is utilized by circuitry for a determination of an energy balance corresponding to said period of time, and wherein the determination of the energy balance does not include the fluid mass change, andwherein the determination of said energy balance is utilized in conjunction with the determination of the kcal energy expenditure in making the determination of the energy intake. 2. The method of claim 1 wherein the period of time is hours, days or weeks. 3. The method of claim 1 further comprising trending or patterning the metabolic status over multiple periods of time using the determination of the energy intake. 4. A system for determining a metabolic status of a user over a period of time, the system comprising: a) at least one monitoring platform for making at least one measurement of at least one physiological parameter;b) at least one comparator for making a determination of a body composition change and a kcal energy expenditure over said period of time; andc) and at least one display unit for displaying at least a portion of said determination;wherein the determination of the body composition change includes a determination of a fluid mass change, andwherein said measurements are utilized by the comparator for the determination of the body composition change for said period of time, then subsequently employed by the comparator for an automatic determination of an energy balance which does not include the fluid mass change, andwherein the determination of the energy balance in conjunction with the determination of the kcal energy expenditure enables a determination of a kcal energy intake. 5. The system of claim 4 wherein said monitoring platform is configured to be affixed to a body region and enables ambulatory activities by the user while being affixed over the period of time. 6. The system of claim 4 wherein said monitoring platform and said display unit are incorporated into a handheld device. 7. The system of claim 4 further comprising a data management system in communication with said comparator and said display unit. 8. The system of claim 7 wherein the data management system is capable of presenting one or more weight management aids or tools to the user. 9. The system of claim 4 wherein the display is capable of displaying one or more incentives or rewards to the user in response to determined parameters. 10. A system for automatically determining a body composition cha... | PhiloMetron, Inc. | A61B005/00
US08721559 | B2 | 2014-05-13 | Non-invasive method and device for measuring cardiac output | 2009-01-05 | A system comprising a housing containing a signal generator coupled to an antenna and a dielectric material disposed about the antenna. The device is adapted to generate and direct a plurality of signals towards the heart of the person and measure a magnitude of a signal returned from the heart. The device further comprises a processor to compare differences between a magnitude of a signal propagated and the magnitude of the signal returned off the heart and determine a signal frequency having a maximum return loss value based on those differences. The processor also estimates a change in the amplitude of motion of a portion of a wall of the heart based on the differences between the magnitude of the signal propagated by the device and the magnitude of the signal returned off of the portion of the heart. | 1. A system for measuring cardiac parameters of the heart of a person, comprising: a device comprising a housing containing a signal generator coupled to an antenna and a dielectric material disposed about the antenna, wherein said device is adapted to repeatedly generate a sweep of radio frequency signals at a plurality of frequencies and direct said signals towards the heart of the person and measure magnitudes of signals returned off the heart of the person; anda processor contained within said housing, said processor being adapted to compare magnitudes of the radio frequency signals propagated into the heart of the person and magnitudes of the signals returned from the heart to determine return loss values for each of the radio frequency signals,said processor being adapted to determine a maximum return loss value for each sweep of radio frequency signals and estimate a change in the amplitude of motion of a portion of a wall of the heart based on changes in the maximum return loss values. 2. The system of claim 1, wherein said processor is further adapted to calculate the cardiac output of the person; and further comprising a display disposed about the device adapted to display the measured signals and that can be calibrated to display the estimated stroke volume of the person. 3. The system of claim 2, wherein said processor is further adapted to calculate the cardiac parameters of the person based on an estimated stroke volume of the person. 4. The system of claim 1, wherein said device is adapted to generate a signal at a frequency and magnitude having a determined maximum return loss value into the portion of the heart and receive a portion of the signal propagated into the portion of the heart. 5. The system of claim 1, wherein the dielectric material is selected from the group consisting of ceramics, glass, liquids, gels and polymeric materials. 6. The system of claim 3, wherein said processor is further adapted to estimate the stroke volume of the heart based on a change of the amplitude of motion of the portion of the wall of the heart. 7. The system of claim 6, wherein said processor is adapted to estimate the timing of a QRS complex of the heart based on the change of the amplitude of motion of the portion of the wall of the heart. 8. The system of claim 6, wherein said processor is further adapted to estimate the timing a QRS complex of the heart based on a change of a frequency of a signal having the maximum return loss value. 9. The system of claim 1, wherein the dielectric material has a predetermined thickness and a predetermined dielectric constant such that a substantial portion of a near field component of the radio frequency signals propagated towards the heart of the person is contained within said dielectric material. 10. The system of claim 1, wherein the dielectric material has a predetermined thickness and a predetermined dielectric constant such that substantially an entire near field component of the radio ... | Raytheon Company | A61B005/02 A61B005/05
US08442621 | B2 | 2013-05-14 | Surgical trajectory monitoring system and related methods | 2007-05-17 | Systems and methods for determining a desired trajectory and/or monitoring the trajectory of surgical instruments and/or implants in any number of surgical procedures, such as (but not limited to) spinal surgery, including (but not limited to) ensuring proper placement of pedicle screws during pedicle fixation procedures and ensuring proper trajectory during the establishment of an operative corridor to a spinal target site. | 1. A system for controlling the trajectory of a surgical instrument, comprising: a surgical instrument having a proximal portion, a distal portion, and a longitudinal axis extending between said proximal portion and said distal portion, said distal portion configured for advancement through body tissue to a target surgical site and said proximal portion configured to remain outside said body tissue when said distal portion is advanced to said target surgical site and to accommodate an orientation sensor removably coupled thereto such that said surgical instrument is orientable along a desired trajectory, said surgical instrument operable to at least one of create an access corridor, maintain said access corridor, and form a hole in bone;said orientation sensor operable to determine a first angular relationship in a first plane between said sensor and a reference direction and operable to determine a second angular relationship in a second plane between said orientation sensor and said reference direction, wherein said second plane is orthogonal to said first plane and said orientation sensor is removably coupled to said surgical instrument in a known orientation, wherein said orientation sensor is operable while said distal portion of said surgical instrument is positioned within said body tissue; anda feedback device communicatively coupled to said sensor and configured to communicate information to a user regarding at least one of said determined first and second angular relationships between said orientation sensor and said reference direction;wherein said surgical instrument is further configured to angulate about said distal portion until said feedback device indicates that the angular orientation of said surgical instrument longitudinal axis in said first plane matches the desired instrument trajectory in said first plane and the angular orientation of said surgical instrument longitudinal axis in said second plane matches the desired instrument trajectory in said second plane. 2. The system of claim 1, wherein said orientation sensor comprises at least one accelerometer and said reference direction is gravity. 3. The system of claim 1, wherein said surgical instrument includes at least one of a pedicle hole formation instrument, a tissue dilation instrument, and a tissue retraction instrument. 4. The system of claim 3, wherein said pedicle hole formation instrument includes at least one of a pedicle access needle assembly, a pedicle access probe, a drill, and a tap, wherein said tissue dilation instrument includes at least one of a K-wire and a dilation cannula, and wherein said tissue retraction instrument includes a retractor. 5. The system of claim 1, wherein said sensor is temporarily coupled to said instrument by a connector having a first end for releasably engaging said instrument. 6. The system of claim 1, wherein said feedback device includes at least one of a liquid crystal display (LCD) screen, a light emitting diode (LED) ... | NuVasive, Inc. | A61B005/05
US08715195 | B2 | 2014-05-06 | System and method for accurate placement of a catheter tip in a patient | 2008-04-09 | A system for accurate placement of a catheter tip in a patient, the system including a catheter adapted for placement within a patient, the catheter having a tip at a distal end thereof and having a proximal end which is normally located outside of the patient, a pressure sensor adapted to sense pressure at the tip of the catheter and catheter tip placement location indicating circuitry operative in response to at least an output of the pressure sensor for indicating the location of the catheter tip in the patient. | 1. A system for accurate placement of a catheter tip in a patient, the system comprising: a catheter adapted for placement within a vein of a patient, the catheter having a tip at a distal end thereof and having a proximal end which is normally located outside of the patient;a pressure sensor adapted to sense pressure within the vein of said patient at said tip of said catheter, said pressure sensor being adapted to sense a venous pressure waveform including at least one of a V pressure wave, generated by filling of the right atrium, and an A pressure wave, generated by contraction of the right atrium; andcatheter tip placement location indicating circuitry operative in response to at least an output of said pressure sensor and an ECG signal to provide an indication that said catheter tip is located in said patient at the junction of the superior vena cava (SVC) and the right atrium (RA),an R wave of said ECG signal being employed by said catheter tip placement location indicating circuitry as a fiducial point with respect to which a propagation delay of said venous pressure waveform is measured, said catheter tip placement location indicating circuitry employing said propagation delay to provide said indication that said catheter tip is located at the junction of the superior vena cava (SVC) and the right atrium (RA). 2. A system for accurate placement of a catheter tip in a patient according to claim 1 and wherein said pressure sensor is adapted to sense a heart valve opening/closing signal which propagates at approximately 1540 m/sec. and to sense a pressure wave that is generating by the emptying, filling and contraction of the right atrium which propagates at approximately 2 m/sec. 3. A system for accurate placement of a catheter tip in a patient according to claim 2 and wherein the catheter tip placement location indicating circuitry is operative to distinguish between the heart valve opening/closing signal and the pressure wave that is generating by the emptying, filling and contraction of the right atrium. 4. A system for accurate placement of a catheter tip in a patient according to claim 2 and wherein the catheter tip placement location indicating circuitry is operative for indicating the location of said catheter tip in said patient on the basis of the time relationship of the heart valve opening/closing signal and the pressure wave that is generating by the emptying, filling and contraction of the right atrium sensed by said pressure sensor. 5. A system for accurate placement of a catheter tip in a patient according to claim 1 and wherein said pressure sensor is located at said catheter tip. 6. A system for accurate placement of a catheter tip in a patient according to claim 1 and wherein said pressure sensor is located at said proximal end of said catheter. 7. A system for accurate placement of a catheter tip in a patient according to claim 1 and wherein said indication is based on the change in successive measurements of a pro... | Elcam Medical Agricultural Cooperative | A61B005/0215
US08473032 | B2 | 2013-06-25 | Feature-based registration method | 2009-06-02 | Methods for registering a three-dimensional model of a body volume to a real-time indication of a sensor position that involve analyzing scanned and sensed voxels and using parameters or thresholds to identify said voxels as being either tissue or intraluminal fluid. Those voxels identified as fluid are then used to construct a real-time sensed three-dimensional model of the lumen which is then compared to a similarly constructed, but previously scanned model to establish and update registration. | 1. A method for registering a three-dimensional model of a body volume to a real-time indication of a sensor position, the method comprising: collecting reference data on the body volume;creating the three-dimensional model of said body volume from the reference data;inserting a sensor into a body lumen and recording location data from said sensor including a real-time sensor position;processing said location data recorded by the inserted sensor to form a three-dimensional shape corresponding to space within said lumen, the three-dimensional shape including a plurality of cavity voxels;assigning a value to each cavity voxel of the plurality of cavity voxels encountered by the sensor, the value of each cavity voxel corresponding to a frequency with which each cavity voxel encounters the sensor;adjusting a density of the plurality of cavity voxels in accordance with the value of each cavity voxel;creating clouds of cavity voxels having varying densities that match interior anatomical cavity features;defining a plurality of parameters having predefined thresholds;determining which of the cavity voxels of the plurality of cavity voxels identified by the location data satisfy the predefined thresholds of the plurality of parameters;selecting the cavity voxels satisfying the predefined thresholds of the plurality of parameters; andcomparing said three-dimensional shape including only cavity voxels satisfying the predefined thresholds of the plurality of parameters to said three dimensional model to establish a feature-based registration. 2. The method of claim 1 wherein collecting reference data comprises acquiring a plurality of CT scans. 3. The method of claim 1 wherein creating the three-dimensional model of said body volume involves assembling a plurality of CT scans into a three-dimensional CT model. 4. The method of claim 1 wherein inserting the sensor into the body lumen and recording location data from said sensor comprises placing a catheter having said sensor at its distal tip into the body lumen, and recording position data from said sensor while moving said sensor within at least said body lumen. 5. The method of claim 1 wherein processing said location data comprises de-cluttering said location data. 6. The method of claim 5 wherein processing said location data further comprises digitizing said location data. 7. The method of claim 6 wherein processing said location data further comprises filtering said location data. 8. The method of claim 1 wherein cavity voxels with higher densities are given higher weight in registration than cavity voxels with lower densities. 9. The method of claim 1 wherein assigning values to cavity voxels of said location data involves assigning values based on tissue value Hounsfield numbers. 10. The method of claim 1 wherein defining parameters comprises defining a density range required for each cavity voxel of the plurality of cavity voxels. 11. The method of claim 1 wherein defining parameters comp... | Superdimension, Ltd. | A61B005/05 G06K009/00 G06K009/34
US08246547 | B2 | 2012-08-21 | Packages of apparatus for non-invasive detection of pulse rate and blood flow anomalies | 2007-06-27 | The present invention provides a package for an apparatus for non-invasively monitoring of blood flow of an object, a magnetic field sensing device, and a magnetic source positioning device. | 1. A package for an apparatus that non-invasively monitors blood flow of an object, comprising: a magnetic source that produces a localized, uni-directional, and constant magnetic field;a signal acquisition module with a magnetic sensor, wherein the magnetic sensor is disposed within the localized, uni-directional, and constant magnetic field and detects the modulations of the localized, uni-directional, and constant magnetic field caused by the blood flow in a blood vessel near the skin surface of the object; anda gaseous damping cushion, wherein the magnetic source is disposed within or onto the gaseous damping cushion;such that the gaseous damping cushion enables the modulations of the localized, uni-directional, and constant magnetic field of the magnetic source caused by the blood flow to be amplified and/or propagated, and isolated from external noises. 2. The package of claim 1, further comprising: a signal conditioning module for converting the output of the signal acquisition module with appropriate amplifications; anda digital signal processing module for processing the output signal from the signal conditioning module;thereby pulse rate and blood flow anomaly can be monitored. 3. The package of claim 1, further comprising a display/user interface/alarm module for providing visual or acoustic notification to a user. 4. The package of claim 1, wherein the magnetic source is a permanent magnet. 5. The package of claim 1, wherein the magnetic source is an electromagnet. 6. The package of claim 1, wherein the magnetic source comprises a plurality of electromagnets. 7. The package of claim 6, wherein the strength of the localized, uni-directional, and constant magnetic field produced by the plurality of electromagnets is controlled electronically. 8. The package of claim 1, wherein the magnetic sensor is any magnetic sensor with appropriate sensitivity of detecting the modulation of the magnetic field from the magnetic source. 9. The package of claim 8, wherein the magnetic sensor is a giant magnetoresistance (GMR) based magnetic sensor. 10. The package of claim 8, wherein the magnetic sensor is a tunneling magnetoresistive (TMR) magnetic sensor. 11. The package of claim 8, wherein the magnetic sensor is an anisotropic magnetoresistive (AMR) sensor. 12. The package of claim 1, wherein the signal acquisition module comprises a printed circuit board and two magnetic sensors; wherein the printed circuit board is sandwiched by the two magnetic sensors; and wherein the two magnetic sensors have an orthogonal formation. 13. The package of claim 1, wherein the gaseous damping cushion has a doughnut configuration with a circular chamber; wherein the circular chamber is disposed with means for allowing the embedded magnetic source to move and holding the embedded magnetic source in place. 14. The package of claim 13, wherein the means within the circular chamber comprises a plurality of semi-flexible flaps. 15. The package of claim 1, w... | Nanyang Polytechnic | A61B005/02
US08142369 | B2 | 2012-03-27 | Measurement and use of in-socket residual limb volume change data for prosthetic fitting | 2009-07-27 | Changes in the volume of residual limbs on which prosthetic sockets are worn can be measured based on bioimpedance measurements along one or more segments of the limb. A current at an appropriate frequency (e.g., in the range from 1 kHz to 1 MHz) is injected at two current electrodes that contact the skin of the residual limb. The voltage at the voltage electrodes disposed between the current electrodes is measured and using an appropriate model, the change in the segmented volume of the limb can be determined during periods of different activity and at different times during the day. This information can be used for assessing the fit of the socket and can also provide a feedback signal for automatically controlling volume management devices, to ensure a more comfortable fit when the volume of the limb is changing. | 1. A method for using bioimpedance to measure volumetric changes of a residual limb of a person with a limb amputation over time, while the person is wearing a prosthetic socket on the residual limb, comprising the steps of: (a) injecting an alternating current into tissue of the residual limb, between two longitudinally spaced-apart points along the residual limb;(b) detecting a change in voltage at a plurality of points that are intermediate the two spaced-apart points; and(c) based upon the change in the voltage, using a model for determining a change in the volume of the residual limb over time. 2. The method of claim 1, further comprising the step of controlling a frequency of the alternating current to be within a frequency range from about 1 kHz to about 1 MHz. 3. The method of claim 1, further comprising the step of determining changes in the volume of the residual limb during periods of different types of activity. 4. The method of claim 1, further comprising the step of using the change in volume of the limb to determine if the prosthetic socket should be changed to provide a new prosthetic socket that better fits the residual limb of the person. 5. The method of claim 1, further comprising the steps of using the change in volume of the residual limb to determine a non-essential fluid volume and an essential fluid volume to aid in designing a new prosthetic socket for the residual limb. 6. The method of claim 1, further comprising the step of using the change in volume of the residual limb to determine a cause of a volume control problem of the person. 7. The method of claim 1, further comprising the step of using a signal indicative of the change in volume of the residual limb as a feedback signal to control a device that modifies the volume of the residual limb, to automatically compensate for the change in the volume of the residual limb as the person engages in different activities. 8. The method of claim 1, further comprising the step of using a signal indicative of the change in volume of the residual limb as a feedback signal to control a device that modifies a volume of at least one of the prosthetic socket, and a component disposed in the prosthetic socket that can change an available volume in the prosthetic socket, so as to automatically compensate for the change in the volume of the residual limb as the person engages in different activities. 9. The method of claim 1, further comprising the step of using a signal indicative of the change in volume of the residual limb as a feedback signal to control a device in at least one of a prosthetic foot, a pylon, and an alignment adaptor, to modify a force applied to the residual limb through the prosthetic socket, so as to automatically compensate for the change in the volume of the residual limb as the person engages in different activities. 10. The method of claim 9, further comprising the step of wirelessly transmitting the feedback signal that controls the device to the... | University of Washington | A61B005/1477
US08155264 | B2 | 2012-04-10 | Gated computed tomography | 2007-11-13 | A computed tomography system (100) includes a windowing component (140) that receives an ECG signal that includes a premature heart cycle. The ECG signal is time-synchronized with x-ray projection data of a beating heart. The windowing component (140) either removes or repositions a first reconstruction window within a first heart cycle to correspond to a desired cardiac phase when the premature heart cycle causes the first reconstruction window to correspond to a different cardiac phase, based on available data. A reconstructor (148) that reconstructs projection data corresponding to a plurality of reconstruction windows from different cardiac cycles generates image data indicative of the desired phase of the heart. | 1. A system, comprising: a windowing component that receives an ECG signal that includes a premature heart cycle and previously acquired x-ray projection data of a beating heart, wherein the ECG signal is time-synchronized with the previously acquired x-ray projection data of the beating heart, and wherein the windowing component positions a first reconstruction window within a first heart cycle to correspond to a desired cardiac phase of the previously acquired x-ray projection data when the premature heart cycle causes the first reconstruction window to correspond to a different cardiac phase; anda reconstructor that reconstructs projection data corresponding to a plurality of reconstruction windows from different cardiac cycles to generate image data indicative of the desired phase of the heart. 2. The system of claim 1, wherein the windowing component repositions the first reconstruction window in terms of time relative to a reference signal within the first heart cycle when a first time interval from the reference signal to the first reconstruction window is less than a second time interval from the reference signal to the premature heart cycle. 3. The system of claim 1, wherein the windowing component removes the first reconstruction window when a first time interval from the reference signal to the first reconstruction window is greater than a second time interval from the reference signal to the premature heart cycle. 4. The system of claim 3, further including a recommendation component, wherein the recommendation component recommends a different phase that ensures that the first time interval from the reference signal to the first reconstruction window is less than the second time interval from the reference signal to the premature heart cycle. 5. The system of claim 1, wherein the first heart cycle is the premature heart cycle. 6. The system of claim 5, wherein the windowing component attempts to remove the first reconstruction window on condition that no data gap occurs. 7. The system of claim 5, wherein the windowing component repositions the first reconstruction window in terms of time relative to a reference signal within the first heart cycle when the premature heart cycle is an atrial extrasystole. 8. The system of claim 5, wherein the windowing component repositions the first reconstruction window in terms of time relative to a reference signal in a subsequent heart cycle when the premature heart cycle is a ventricular extrasystole. 9. The system of claim 8, further including a console, wherein an operator provides an input via the console that confirms or rejects the recommended cardiac phase for reconstruction. 10. The system of claim 1, further including a recommendation component, wherein the recommendation component recommends a second phase for reconstruction when the second phase has a relatively higher probability of having sufficient data for reconstruction after removing a reconstruction window due to the prem... | Koninklijke Philips Electronics N.V. | A61B006/00
US08855761 | B2 | 2014-10-07 | Method and system for temperature analysis to provide an early marker of congestive heart failure progress that precedes a patient's symptoms | 2009-07-10 | The temperature of a patient is a significant predictor of death in heart failure patients. Temperature provides a window into the physiology of the patient's underlying condition and may be used as an early marker for CHF exacerbations. The patient's temperature is taken to form a time series of temperature values. In accordance with some embodiments, the time series of temperature values is converted to the frequency domain by, for example, a discrete Fourier Transform. The frequency domain representation then is analyzed for a marker indicative of the worsening condition of the patient. In accordance with other embodiments, the patient's time series of temperature values is analyzed for a marker using, for example, Cosinor analysis. In yet other embodiments, both the time and frequency domain temperature data is analyzed for markers of the patient's worsening medical condition. | 1. A method of predicting worsening of heart failure in a patient, comprising: obtaining, by a processor, time domain temperature data (TDTD) of core body temperature of the patient taken over a time period of more than one day;converting, by said processor, said TDTD to a frequency domain representation (FDR);determining, by said processor, a first marker indicative of worsening of heart failure of the patient by analyzing the FDR;determining, by said processor, a second marker indicative of worsening of heart failure by analyzing the TDTD, wherein analyzing the TDTD comprises performing at least one of Cosinor analysis, approximate entropy, tempogram, day-night mean temperature difference, body temperature daily range, interdaily stability, and time-lag embedding;weighing outputs of at least two analysis techniques to render a total score indicative of the current condition of the patient, wherein the at least two analysis techniques comprise a technique for analyzing the TDTD and another technique for analyzing the FDR; anddetermining that the patient is experiencing worsening heart failure based on the total score. 2. The method of claim 1 wherein said TDTD is selected from the group consisting of temperature taken by a Foley™ urinary catheter, temperature taken by a rectal probe, rectal cavity temperature, bladder temperature, esophagus temperature, trachea temperature. 3. The method of claim 1 wherein obtaining the TDTD comprises acquiring temperature values from at least one temperature sensor or reading temperature values from storage. 4. The method of claim 1 wherein converting said TDTD to FDR comprises performing at least one of Fourier transform, wavelet transform, and Chirplet transform. 5. A system comprising a processor that receives time domain temperature data (TDTD) pertaining to a patient, converts said TDTD to a frequency domain representation (FDR), and determines a first marker indicative of worsening of heart failure of the patient by analyzing the FDR and determines a second marker indicative of worsening of heart failure by analyzing the TDTD, wherein analyzing the TDTD comprises performing at least one of Cosinor analysis, approximate entropy, tempogram, day-night mean temperature difference, body temperature daily range, interdaily stability, and time-lag embedding, wherein the processor weighs outputs of at least two analysis techniques to render a total score indicative of the current condition of the patient, wherein the at least two analysis techniques comprise a technique for analyzing the TDTD and another technique for analyzing the FDR. 6. The system of claim 5 wherein said processor utilizes at least one of Fourier transform, wavelet transform, and Chirplet transform to convert said TDTD to FDR. 7. The system of claim 5 wherein the first marker indicative of worsening of heart failure is determined upon detecting oscillation in the amplitude of the peak at the patient's circadian frequency in the FDR. 8. ... | Texas Heart Institute | A61N001/365 A61B005/01 A61B005/00 G06F019/00
US08945022 | B2 | 2015-02-03 | Puncturing system | 2009-10-12 | A puncturing system for obtaining a sample of body fluid comprises a magazine that comprises a lancet carrier that carries several lancets, a puncturing device comprising a compartment for the magazine, an incremental advancing mechanism for moving the lancets successively to a puncturing position, and a puncturing drive for accelerating one of the lancets that is positioned in the puncturing position in a puncturing motion. The magazine comprises a coupling facility with a receptacle for the lancet carrier. The receptacle is mobile with respect to a magazine housing The lancets that are carried by the lancet carrier can be moved relative to the receptacle by actuating the incremental advancing mechanism The receptacle is adapted to couple the puncturing drive to one of the lancets positioned in the puncturing position for transmitting a driving force generated by the puncturing drive during a puncture to the lancet positioned in the puncturing position. | 1. A puncturing system for obtaining a sample of body fluid, the system comprising: a magazine comprising a lancet carrier that has several lancets;a puncturing device comprising: a compartment for the magazine, anda puncturing drive for accelerating one of the lancets that is positioned in a puncturing position in a puncturing motion;wherein the magazine further comprises a coupling facility with a receptacle defining a recess in which a portion of the lancet carrier is received, said receptacle being mobile with respect to a magazine housing;wherein the lancets of the lancet carrier are configured to be moved relative to the receptacle;wherein the receptacle is adapted to couple the puncturing drive to the lancet positioned in the puncturing position for transmitting a driving force generated by the puncturing drive during a puncture to the lancet positioned in the puncturing position; andwherein the puncturing drive is configured to move the lancet along with the portion of the lancet carrier received in the recess by moving the receptacle in the puncturing motion during the puncture. 2. The puncturing system of claim 1 wherein the lancet carrier is ribbon-shaped. 3. The puncturing system of claim 1 wherein the lancet carrier carries test fields for testing a sample of body fluid obtained from a puncturing wound, said test fields being arranged between the lancets. 4. The puncturing system of claim 1 wherein the coupling facility is configured to couple to the lancet carrier and to move the portion of the lancet carrier received in the recess during a puncture in conjunction with the lancet that is positioned in the puncturing position. 5. The puncturing system of claim 1 wherein the coupling facility is configured to remain coupled to the puncturing drive during movement of the lancet carrier relative to the receptacle. 6. The puncturing system of claim 1 wherein the coupling facility is configured to grasp the lancet carrier in a clamping manner prior to the puncture and to release the lancet carrier after the puncturing motion and a returning motion are completed. 7. The puncturing system of claim 1 wherein the coupling facility comprises a spring element configured to effect a clamping force for clamping the lancet carrier. 8. The puncturing system of claim 1 wherein the coupling facility includes a structured contact surface configured to contact the lancet carrier. 9. The puncturing system of claim 8 wherein the structured contact surface is configured to resist a relative motion of the lancet carrier and the structured contact surface with respect to each other, wherein the amount of resistance depends on the direction of the relative motion. 10. The puncturing system of claim 1 wherein the coupling facility has fibers that are obliquely inclined in the direction of motion during retraction of the lancet in the puncturing position. 11. The puncturing system of claim 10 wherein the coupling facility and the lancet carrier combi... | Roche Diagnostics Operations, Inc. | A61B005/15 A61B005/151 A61B005/155 A61B005/157 A61B005/145
US08275453 | B2 | 2012-09-25 | Method and system for detecting electrophysiological changes in pre-cancerous and cancerous breast tissue and epithelium | 2009-10-23 | A method and system are provided for determining a condition of a selected region of epithelial and stromal tissue in the human breast. A plurality of measuring electrodes are used to measure the tissue and transepithelial electropotential of breast tissue. Surface electropotential and impedance are also measured at one or more locations. An agent may be introduced into the region of tissue to enhance electrophysiological characteristics. The condition of the tissue is determined based on the electropotential and impedance profile at different depths of the epithelium, stroma, tissue, or organ, together with an estimate of the functional changes in the epithelium due to altered ion transport and electrophysiological properties of the tissue. Devices for practicing the disclosed methods are also provided. | 1. A method for determining a condition of a region of ductal epithelial breast tissue comprising: (A) establishing a connection between a first electrode and the epithelial breast tissue comprising: (1) placing over the nipple of a breast a cup having an interior volume, first and second openings, and a current-passing first electrode disposed within the interior volume;(2) placing a source of suction in communication with the first opening;(3) introducing an electroconductive medium within the interior of the cup; and(4) applying suction to the first opening;(B) placing a second electrode in contact with the surface of the breast;(C) establishing a signal between the first and second electrodes;(D) measuring an electrical property between the first and second electrode; and(E) determining the condition of the ductal epithelial tissue based on the measured electrical property between the first and second electrode. 2. The method of claim 1, wherein the electrical property is impedance. 3. The method of claim 1, wherein the electrical property is a DC potential. 4. The method of claim 1, wherein the electroconductive media comprises physiologic saline. 5. The method of claim 4, wherein the physiologic saline further comprises an agent suitable for inducing a challenge test of the epithelial tissue. 6. The method of claim 5, wherein the agent is selected from the group consisting of agonists, antagonists, ionic substitutions, hormones, growth factors and mixtures thereof. 7. The method of claim 6, wherein the agent is a hormone. 8. An apparatus for determining a condition of a region of ductal epithelial breast tissue comprising: (A) a cup having an interior volume, first and second openings, and a current-passing first electrode disposed within the interior volume;(B) a source of suction suitable for communication with the first opening;(C) the cup adapted to receive an electroconductive medium within the interior of the cup;(D) a second electrode adapted for contact with the surface of the breast; wherein: (1) the first electrode is operable for electrical communication with the ductal epithelial tissue; and(2) the second electrode can be placed at a location spaced from said first electrode; andwherein the apparatus is adapted to: (3) establish a signal between the first and second electrodes; and(4) measure at least one electrical property between the first and second electrode so as to determine the condition of the ductal epithelial tissue based on the measured electrical property between the first and second electrode. 9. The apparatus of claim 8, wherein the electrical property is impedance. 10. The apparatus of claim 8, wherein the electrical property is DC potential. 11. The apparatus of claim 8, wherein the electroconductive media comprises physiologic saline. 12. The apparatus of claim 11, wherein the physiologic saline further comprises an agent suitable for inducing a challenge test of the epithelial tissue, said agent sel... | Epi-Sci, LLC | A61B005/05 A61N001/30
US08107730 | B2 | 2012-01-31 | Imaging system sag correction method and apparatus | 2008-08-05 | An apparatus for use with a single modality imaging system configured to generate uncorrected imaging data of a patient, the single modality imaging system includes two gamma cameras and a patient stretcher disposed between the two gamma cameras, the apparatus for compensating for downward stretcher deflection at the extended end of the patient stretcher that occurs during stretcher extension. The apparatus includes a single sag sensor for sensing the downward deflection of the patient stretcher, a subtracting device configured to generate a sag correction factors based on a baseline stretcher height and an input received from the sag sensor, and a compensator configured to modify at least a portion of the uncorrected imaging data to compensate for sag using the using the sag correction factor to generate a unified image. | 1. An apparatus for use with a single modality imaging system configured to generate uncorrected imaging data of a patient, the single modality imaging system includes at least one gamma camera and a patient stretcher disposed between the two gamma cameras, the apparatus comprising: a single sag sensor for sensing the downward deflection of the patient stretcher;a subtracting device configured to generate a sag correction factors based on a stretcher baseline height and an input received from the sag sensor; anda compensator configured to modify at least a portion of the uncorrected imaging data to compensate for sag using the using the sag correction factor to generate a unified image. 2. An apparatus in accordance with claim 1 wherein the single sag sensor is mounted on an imaging system stator. 3. An apparatus in accordance with claim 1 wherein the gamma camera has a field of view that is greater than approximately 15 centimeters. 4. An apparatus in accordance with claim 1 wherein the stretcher baseline height represents a height of the patient stretcher including the patient. 5. An apparatus in accordance with claim 1 wherein the sag sensor is positioned downstream from the imaging system gantry proximate to a distal end of the gamma camera field of view (FOV). 6. An apparatus in accordance with claim 1 wherein the uncorrected imaging data includes a plurality of slices, the subtracting device is configured to generate a sag correction factor for each slice, and the compensator is configured to modify each slice compensate for stretcher sag using the using the sag correction factors. 7. An apparatus in accordance with claim 1 wherein the sag sensor comprises at least one of an ultrasonic sensor, a laser sensor, an acoustical sensor, an optical sensor, a light sensor, and a magnetic sensor. 8. A nuclear medicine imaging system comprising: at least a first gamma camera;a patient stretcher disposed between the first gamma camera;a single sag sensor for sensing the downward deflection of the patient stretcher;a subtracting device configured to generate a sag correction factors based on a stretcher baseline height and an input received from the sag sensor; anda compensator configured to modify at least a portion of the uncorrected imaging data to compensate for stretcher sag using the using the sag correction factor. 9. A nuclear medicine imaging system in accordance with claim 8 wherein the gamma camera has a field of view that is greater than approximately 15 centimeters. 10. A nuclear medicine imaging system in accordance with claim 8 wherein the stretcher baseline height represents a height of the patient stretcher including the patient. 11. A nuclear medicine imaging system in accordance with claim 8 wherein the sag sensor is positioned downstream from the imaging system gantry proximate to a distal end of the gamma camera field of view (FOV). 12. A nuclear medicine imaging system in accordance with claim 8 wherein the uncorrected ... | General Electric Company | G06K009/00
US08118751 | B2 | 2012-02-21 | Devices and methods for accelerometer-based characterization of cardiac function and identification of LV target pacing zones | 2009-03-02 | Systems according to the invention employ an acceleration sensor to characterize displacement and vibrational LV motion, and uses this motion data to characterize the different phases of the LV cycle for analyzing LV function. Systems may identify a target pacing region or regions in the LV or RV using the acceleration sensor by localizing regions of late onset of motion relative to the QRS, or isovolumic contraction, or mitral valve closure, or by pacing of target regions and measuring LV function in response to pacing. Systems further provide an implantable or non-implantable acceleration sensor device for measuring LV motion and characterizing LV function. An implantable myocardial acceleration sensing system (“IAD”) includes at least one acceleration sensor, a data acquisition and processing device, and an electromagnetic, e.g., RF, communication device. The IAD may be integrated into the pacing lead of a CRT device and can operate independently of the CRT IPG. | 1. A device for monitoring cardiac function, comprising: a. a single multi-axis acceleration sensor configured for disposition within the coronary sinus;b. wherein the acceleration sensor is configured to characterize left ventricular function by simultaneously monitoring both cardiac vibration corresponding to mitral regurgitation and left ventricular longitudinal displacement motion;c. a circuit implemented in hardware or software, or both, for receiving signals indicative of left ventricular function including signals corresponding to cardiac vibration caused by mitral regurgitation and signals corresponding to left ventricular longitudinal displacement motion, the circuit further mathematically singly integrating the received signals to obtain velocity information about left ventricular function and the circuit further mathematically doubly integrating the received signals to obtain displacement information about left ventricular function, wherein the single and double integrations are performed using only data from the single multi-axis acceleration sensor. 2. The device of claim 1, wherein the acceleration sensor is designed to detect vibrational frequencies between about 20 Hz and 150 Hz and displacement motion frequencies less than about 20 Hz. 3. The device of claim 1, wherein the acceleration sensor is designed to detect vibrational frequencies related to mitral regurgitation greater than about 150 Hz. 4. The device of claim 1, further comprising an endovascular catheter, and wherein the sensor is disposed on the endovascular catheter. 5. The device of claim 4, wherein the endovascular catheter is a coronary sinus catheter, and further comprising an occlusion balloon coupled to the catheter. 6. The device of claim 4, wherein the endovascular catheter is a coronary sinus catheter, and further comprising a pacing electrode coupled to the catheter guidewire lumen formed in the coronary sinus catheter. 7. The device of claim 1, further comprising an LV pacing lead, and wherein the sensor is disposed on the LV pacing lead. 8. A device for monitoring cardiac function, comprising: a. a catheter including a single three-axis acceleration sensor for disposition within the coronary sinus, wherein the acceleration sensor is configured to characterize left ventricular function by simultaneously monitoring both cardiac vibration corresponding to mitral regurgitation and left ventricular longitudinal displacement motion;b. a guide wire lumen formed within the catheter; andc. a circuit implemented in hardware or software, or both, for receiving signals indicative of left ventricular function including signals corresponding to cardiac vibration caused by mitral regurgitation and signals corresponding to left ventricular longitudinal displacement motion, the circuit further mathematically singly integrating the received signals to obtain velocity information about left ventricular function and the circuit further mathematically doubly integrating... | CardioSync, Inc. | A61B005/04
US08315681 | B2 | 2012-11-20 | Method for noninvasive measurement of glucose and apparatus for noninvasive measurement of glucose | 2009-04-06 | A method for noninvasive measurement of glucose in a tissue of a subject, including the steps of bringing an adaptation device, which has a shape similar to a measurement probe, into contact with a skin part of a subject for stretching the skin part of the subject under a pressure that is higher than a pressure per unit area applied by the measurement probe during the noninvasive measurement, maintaining the contact for a predetermined period of time followed by relieving the contact, bringing the measurement probe into contact with the stretched skin part of the subject for the noninvasive measurement, collecting signals emitted from the subject, and estimating a glucose concentration based on the collected signals. | 1. A method comprising the steps of: inducing a change in glucose metabolism in a nutrient capillary in skin by temperature-changed glycolysis;measuring a change in localized reflectance light signals at a plurality of light source-detector distances and a plurality of wavelengths over a specific time period after skin-probe contact as a function of a time for which a localized reflectance probe is brought into contact with the skin, said temperature-changed glycolysis causing a change with respect to light attenuation, oxygen consumption in a tissue and concentration of a hemoglobin variant;selecting a time window in which a tissue-probe adaptation effect on the signals is minimized and an effect on glycolysis induced by temperature has time dependence, and using a signal measured in the time window for a subsequent calculation;calculating one set of functions based on a plurality of localized reflectance values at the plurality of light source-detector distances and the plurality of wavelengths at a plurality of time intervals in the time window and at least two wavelengths;deriving a calibration relationship between a combination of the calculated functions and a glucose concentration in a living body; andusing the calibration relationship for predicting a glucose concentration in a body fluid in subsequent measurement,wherein at least one function in the plurality of time intervals is associated with a change in oxygen consumption. 2. The method according to claim 1, wherein the calculating step comprises the step of calculating a rate of change in at least one function associated with the time dependence of a temperature change on the localized reflectance values at the plurality of wavelengths and light source-detector distances over the specific time window. 3. The method according to claim 1, wherein the calculating step comprises the step of calculating a degree of change in at least one function associated with a time dependent effect of temperature stimulation on the localized reflectance values at the plurality of wavelengths and light source-detector distances, and the degree of change is calculated for at least one time window and is averaged over adjacent time windows. 4. The method according to claim 1, wherein the wavelengths have different absorption coefficients with respect to oxyhemoglobin and deoxyhemoglobin. 5. The method according to claim 1, wherein data points used in the calculation are present in a time window which starts after a predetermined time from the skin-probe contact. 6. A noninvasive measurement apparatus comprising: a unit which modulates a temperature of a localized reflectance optical probe, when the probe has been brought into contact with skin to a temperature substantially different from a normal temperature of the skin for inducing a change in temperature of a tissue in a vicinity of the probe and up to a depth of the skin surrounded by a skin vascular system;a unit which measures a change in loc... | Toshiba Medical Systems Corporation | A61B005/1455
US08768429 | B2 | 2014-07-01 | Internal bleeding detection apparatus | 2006-12-28 | An EIT system (1) adapted to detect internal bleeding in a body portion, the EIT system (1) comprising a plurality of electrodes (3) adapted in use to extend in a substantially linear orientation across one side only of the body portion and to be applied in electrical contact with the skin of the body portion, a current source adapted to cyclically apply an electric current between one pair of the electrodes (3), a voltage measuring means to measure the voltage across each of the other pairs of the electrodes resulting from the current, a data collection system (2) and a data analysis system (4) to analyze data resulting from the voltages that are measured by the voltage measuring means, wherein the analysis system (4) is configured to obtain quantitative information related to amounts and rates of conductive tissue changes occurring in the body, based on an EIT analysis equivalent to that obtained from data derived from electrodes spaced around the full perimeter of the body portion. Also disclosed is an electrode belt suitable for bioelectrical use and in particular for detection of change of volume of tissue in a body portion. | 1. An electrode belt for use with an Electrical Impedance Tomography (EIT) system, the belt comprising: a core layer comprising a plurality of electrodes spaced along an elongate length of the belt, the electrodes each having a contact face configured to contact and make electrical connection with skin of a portion of a body of a patient under examination;the core layer disposed between a first insulating layer and a second insulating layer, wherein the first insulating layer provides insulation over the electrodes and the second insulating layer does not provide insulation over the electrodes, thereby allowing the electrodes to be exposed on only one side thereof;a first shielding layer disposed over the first insulating layer and a second shielding layer disposed over the second insulating layer, wherein the first and second shielding layers comprise a conductive material thereby providing circumferential active or ground shielding for the belt;a first outer layer disposed over the first shielding layer and a second outer layer disposed over the second insulating layer, the second outer layer including a plurality of apertures corresponding to the electrodes to provide engagement of only one side of the electrodes on the skin of the portion of the body;wherein the core layer, the first and second insulating layers, and the first and second shielding layers are created using a multilayer printing process; andthe belt being configured to, in use, extend substantially from one side of the portion of the body of the patient under examination to another side of the portion of the body of the patient under examination without encircling the body of the patient. 2. An electrode belt as claimed in claim 1, wherein the contact faces are of elongate form, the elongate length of the contact faces being oriented transversely to the elongate length of the belt. 3. An electrode belt as claimed in claim 1, wherein the contact faces are substantially rectangular with a length in the range of 75 mm to 100 mm and a width in the range of 5 to 25 mm. 4. An electrode belt as claimed in claim 1, wherein the belt is flexible to enable the belt to conform to a profile of the body portion upon which the belt is placed to facilitate contact of each electrode with the body portion. 5. An electrode belt as claimed in claim 1, wherein, the portion of the body is an abdomen, and opposing ends of the belt are formed with a curvature to facilitate contact of electrodes in a vicinity of sides of the abdomen when in use. 6. An electrode belt as claimed in claim 1, wherein at least one selected from the electrodes and ends of the belt, are provided with an adhesive surround or an adhesive member to facilitate secure engagement of the one selected from the electrodes and the ends of the belt with the skin. 7. An electrode belt as claimed in claim 1, wherein each of the electrodes has a conductive track connecting each electrode with a printed termination. 8. An electrode b... | E.I.T. Pty Ltd. | A61B005/04
US08560033 | B2 | 2013-10-15 | Magnetically controlled wireless optical oxygen sensor for intraocular measurements | 2009-05-18 | A sensor for intraocular measurements moveable within at least one of an vitreous humor of an eye, an aqueous humor of an eye and an intraocular replacement medium. The sensor includes a magnetic body susceptible to magnetic fields and at least one sensor film. | 1. A sensor for intraocular measurements moveable within at least one of a vitreous humor of an eye, an aqueous humor of an eye and an intraocular replacement medium, the sensor comprising: a magnetic body susceptible to magnetic fields; andat least one sensor film including a supporting matrix and at least one of a luminescence dye, a fluorescence dye and a phosphorescence dye. 2. The sensor as recited in claim 1, wherein the magnetic body includes at least one of a ferromagnetic material, a ferrimagnetic material, a paramagnetic material and a compound with magnetic properties. 3. The sensor as recited in claim 2, wherein the magnetic body includes the compound with magnetic properties which includes at least one of magnetic particles, nanoparticles and super paramagnetic nanoparticles. 4. The sensor as recited in claim 2, wherein the magnetic body is shaped as a sphere, an oblate spheroid, a prolate spheroid, a plane or a cylinder. 5. The sensor as recited in claim 2, wherein the at least one ferromagnetic material, ferrimagnetic material, paramagnetic material and a compound with magnetic properties is selected from at least one of alloys of iron, nickel and cobalt. 6. The sensor as recited in claim 1, wherein the at least one luminescence dye, fluorescence dye and phosphorescence dye has a luminescence effect susceptible to at least one of a physical parameter, a chemical parameter and the concentration of a chemical analyte. 7. The sensor as recited in claim 1, wherein the supporting matrix includes at least one of a polymer, silica gels and sol-gels. 8. The sensor as recited in claim 7, wherein the polymer is polystyrene. 9. The sensor as recited in claim 1, wherein the at least one luminescence dye, fluorescence dye and phosphorescence dye is at least one of an iridium complex and a ruthenium complex. 10. The sensor as recited in claim 1, wherein the sensor film is disposed on a highly reflecting surface of the magnetic body. 11. The sensor as recited in claim 1, wherein a surface of the magnetic body is coated with at least one of a metal, metal alloy and a polymer. 12. The sensor as recited in claim 11, wherein the at least one metal, metal alloy and polymer is at least one of gold, titanium, polystyrene and polypyrrole. 13. A sensor for at least one in-vivo measurement moveable within at least one of a human or an animal body, the sensor comprising: a magnetic body susceptible to magnetic fields; andat least one sensor film having a luminescence activity that depends on at least one of a physical parameter, a chemical parameter and the concentration of a chemical analyte. 14. The sensor according to claim 13, wherein the luminescence activity of the at least one sensor film depends on at least one of a molecular oxygen concentration, a carbon dioxide concentration, a glucose concentration, temperature and pH. | ETH Zurich | A61B005/00
US08435187 | B2 | 2013-05-07 | Identification of disease characteristics using isotope ratios in breath | 2008-10-20 | Methods are disclosed for distinguishing whether an animal is experiencing a bacterial infection or a viral infection. One monitors breath taken from the animal over time to measure the relative amount of a first breath stable isotope to a second breath stable isotope therein over time. A quick change in the isotope ratios within several hours from the likely infection is indicative of a bacterial infection. A delayed change in the isotope ratios, followed by periodic repeated alterations in the ratios, is indicative of viral infection. The methods are particularly efficient when using cavity ringdown spectroscopy for the monitoring. They may be used for monitoring a patient already admitted to a hospital, or for monitoring a patient initially complaining of adverse symptoms, or for triage, or for collectively monitoring a population of animals. | 1. A method of distinguishing whether an animal is experiencing a viral infection, the method comprising: monitoring on at least three days breath taken from the animal over time to measure a relative amount of a first isotope to a second isotope therein over time, wherein the first and second isotopes are selected from the group consisting of a pair of 13C and 12C, a pair of 15N and 14N, a pair of 17O and 16O, and a pair of sulphur isotopes; andbased on a comparison of (i) changes to the relative amount of the first isotope to the second isotope over time, with (ii) a known periodic cyclic pattern of changes of the relative amount of the first isotope to the second isotope when there is viral infection, the known periodic cyclic pattern of changes evidencing periods of increased cell rupture that are time separated by periods of less cell rupture, determining whether viral infection is present;wherein either the absence of such a known periodic cyclic pattern of changes indicates a lack of a viral infection, or the presence of such a known periodic cyclic pattern of changes indicates the presence of a viral infection. 2. The method of claim 1, wherein the first and second isotopes are the pair of 13C and 12C. 3. The method of claim 2, wherein the monitoring step uses cavity ringdown spectroscopy to determine the changes in the relative amount. 4. The method of claim 1, wherein the animal is a mammal. 5. The method of claim 4, wherein the mammal is a human. | Wisconsin Alumni Research Foundation | A61B005/08 G01N033/497 G01N033/00
US08849382 | B2 | 2014-09-30 | Apparatus and display methods relating to intravascular placement of a catheter | 2009-09-10 | An integrated catheter placement system for accurately placing a catheter within a patient's vasculature is disclosed. In one embodiment, the integrated system comprises a system console, a tip location sensor for temporary placement on the patient's chest, and an ultrasound probe. The tip location sensor senses a magnetic field of a stylet disposed in a lumen of the catheter when the catheter is disposed in the vasculature. The ultrasound probe ultrasonically images a portion of the vasculature prior to introduction of the catheter. ECG signal-based catheter tip guidance is included in the integrated system to enable guidance of the catheter tip to a desired position with respect to a node of the patient's heart. Various aspects for visualizing and manipulating display of the ECG signal data acquired via the present system, together with aspects of various ECG sensor configurations, are also disclosed. | 1. A catheter placement system for inserting a catheter into a vasculature of a patient, comprising: a catheter assembly including: a catheter, a distal end of the catheter comprising a conductive band; anda stylet removably received within a lumen of the catheter comprising an ECG sensor component for detecting ECG signals of a node of a heart of the patient when at least a portion of the catheter assembly is disposed within the vasculature of the patient, the stylet including two conductive portions longitudinally separated by an interposed nonconductive portion, wherein the nonconductive portion has a longitudinal length less than a longitudinal length of the conductive band;an ECG signal receiving component that receives ECG signals detected by the ECG sensor component, the ECG signal receiving component being operably connectable to the ECG sensor component; anda connection integrity component that verifies an integrity of an operable connection between the ECG signal receiving component and the ECG sensor component. 2. The catheter placement system as defined in claim 1, further comprising at least one ECG electrode for placement on a skin portion of the patient wherein the ECG signal receiving component is disposed external to the patient. 3. The catheter placement system as defined in claim 2, wherein the at least one external ECG electrode includes a first electrode and a second electrode positioned on a single pad for placement on the skin portion of the patient. 4. The catheter placement system as defined in claim 2, wherein the connection integrity component verifies that the external ECG electrode is operably connected to the skin portion of the patient and that the ECG sensor component of the stylet is operably connected to the ECG signal receiving component. 5. The catheter placement system as defined in claim 4, wherein the connection integrity component applies a first voltage to the ECG sensor component of the stylet and a second voltage different from the first voltage to the external ECG electrode. 6. The catheter placement system as defined in claim 5, further comprising a second external ECG electrode for placement on a skin portion of the patient, and the connection integrity component monitors a resulting voltage to determine whether the external ECG electrode and the second external ECG electrode are properly connected to the skin portion of the patient. 7. The catheter placement system as defined in claim 6, wherein the resulting voltage monitored by the connection integrity component comprises a detected signal from the second external ECG electrode and an average voltage from the first and second voltage. 8. The catheter placement system as defined in claim 2, wherein the connection integrity component comprises a radiating element configured to emit a signal of a known frequency. 9. The catheter placement system as defined in claim 8, wherein the connection integrity component is configured to detect a radiati... | C. R. Bard, Inc. | A61B005/04 A61N001/00 A61B005/06 A61B008/08 A61B019/00 A61B019/08 A61B005/042
US08460201 | B2 | 2013-06-11 | Visualization of stress level cardiac functional analysis results | 2007-07-23 | The invention relates to a system (100) for visualizing a cardiac parameter at a plurality of positions in a myocardium and at a plurality of stress levels, the system comprising a determination unit (110) for determining a value of the cardiac parameter at a position from the plurality of positions in the myocardium and at a stress level from the plurality of stress levels on the basis of stress level cardiac functional data, and a visualization unit (120) for visualizing the determined value of the cardiac parameter by displaying a point in a viewing plane. The visualized points are defined by their polar coordinates in a polar coordinate system in the viewing plane. A radial coordinate of the point visualizes the determined value of the cardiac parameter. An angular coordinate of the point visualizes an angular coordinate of the position in the myocardium in a cylindrical coordinate system. Thus, the system allows easy numerical comparison of local myocardial contractions at different stress level values. | 1. A system for visualizing a cardiac parameter at a plurality of positions in a myocardium and at a plurality of stress levels, the system comprising: a memory unit which receives and stores stress level cardiac functional data, the stress level cardiac functional data including cylindrical coordinates of positions on an outer surface and on an inner surface of a wall of the myocardium at a plurality of phases of a cardiac cycle and at a plurality of stress levels;a determination unit for determining a value of the cardiac parameter at a position from the plurality of positions in the myocardium and at a stress level from the plurality of stress levels on the basis of stress level cardiac functional data; anda visualization unit for visualizing the determined value of the cardiac parameter by displaying a point in a viewing plane, where in a polar coordinate system in the viewing plane; a radial coordinate of the point visualizes the determined value of the cardiac parameter; andan angular coordinate of the point visualizes an angular coordinate of the position in the myocardium in a cylindrical coordinate system. 2. The system as claimed in claim 1, wherein the plurality of positions in the myocardium are included in a slice of the myocardium substantially perpendicular to is cylindrical axis of the cylindrical coordinate system. 3. The system as claimed in claim 2, further including: an indication unit for indicating an angle in the polar coordinate system; anda plot unit for plotting a graph, based on positions from the plurality of positions in the myocardium which have angular coordinates substantially equal to the indicated angle. 4. The system as claimed in claim 2, further including an image display unit for displaying the slice of the myocardium at a stress level from the plurality of stress levels. 5. The system as claimed in claim 1, further including a computation unit for computing an average of the cardiac parameter over a set of positions from plurality of positions in the myocardium at a stress level from the plurality of stress levels. 6. An image acquisition apparatus comprising the system claimed in claim 1. 7. A workstation comprising the system as claimed in claim 1. 8. A method visualizing a cardiac parameter at a plurality of positions in a myocardium and at a plurality of stress levels, the method comprising: determining a value of the cardiac parameter at a position from the plurality of positions in the myocardium and at a stress level from the plurality of stress levels on the basis of stress level cardiac functional data; andvisualizing the determined value of the cardiac parameter by displaying a point in a viewing plane, where in a polar coordinate system in the viewing plane: a radial coordinate of the point visualizes the determined value of the cardiac parameter; andan angular coordinate of the point visualizes an angular coordinate of the position in the myocardium in a cylindrical coordinate system. 9. ... | Koninklijke Philips Electronics N.V. | A61B005/02 A61N001/362 A61N001/00
US08214009 | B2 | 2012-07-03 | Electrode for living body and device for detecting living signal | 2009-10-23 | A biosignal measurement device includes an electrode and a signal processing member. The electrode includes an insulation sheet having a hole, a device contact portion provided on the top surface of the insulation sheet and a body contact portion provided on the bottom surface of the insulation sheet, the device contact portion and the body contact portion electrically connected to each other via the hole. The signal processing member includes an externally exposed terminal to make surface contact with the device contact portion, an analog signal processing unit, an A/D signal converter and a digital signal processing unit. Also, the device contact portion and the body contact portion are formed of a material which is both conductive and adhesive. Accordingly, the signal processing member may be directly attached. Noise may be reduced. Also, a biosignal may be accurately measured. | 1. An electrode for transmission between a biosignal measurement device and a living body, comprising: a sheet member formed of an insulating material and having a hole; andan electrode member formed of a conductive material and including a device contact portion contacted with a top surface of the sheet member and having flat contact surfaces, and a body contact portion on a bottom surface of the sheet member, the device contact portion and the body contact portion electrically connected to each other at the hole,wherein the hole integrally connects the device contact portion and the body contact portion. 2. The electrode of claim 1, wherein the electrode member is formed of a material which is both conductive and adhesive. 3. The electrode of claim 2, wherein the electrode member is provided by placing a material, which is both fluidic conductive and adhesive, around the hole of the sheet member and molding the fluidic conductive and adhesive material into a solid material which is both conductive and adhesive. 4. The electrode of claim 1, wherein an adhesive material is applied around the body contact portion on the bottom surface of the sheet member. 5. The electrode of claim 1, wherein an insulation member is provided around the device contact portion to make surface contact with the biosignal measurement device, and the insulation member is formed of a material which is both nonconductive and adhesive. 6. The electrode of claim 1, further comprising: a device holder formed of a nonconductive material on the device contact portion and adapted to hold an edge of a body or a terminal of the biosignal measurement device. | Samsung Electronics Co., Ltd. | A61B005/04
US08241222 | B2 | 2012-08-14 | Monitoring hemodynamic status based on intracardiac or vascular impedance | 2008-10-31 | This disclosure relates to monitoring intracardiac or vascular impedance to determine a change in hemodynamic status by detecting changes in an impedance parameter over cardiac cycles. An example method includes measuring a plurality of impedance values of a path within a patient over time, wherein the path includes at least one blood vessel or cardiac chamber of the patient, and wherein the impedance values vary as a function of blood pressure within the at least one vessel or chamber, determining a plurality of values of an impedance parameter over time based on the measured impedance values, wherein each of the impedance parameter values is determined based on a respective sub-plurality of the impedance values, comparing at least one of the impedance parameter values to at least one prior impedance parameter value, and identifying a change in a cardiovascular parameter related to the blood pressure based on the comparison. | 1. A method comprising: measuring a plurality of impedance values of a path within a patient over time, wherein the path includes at least one blood vessel or cardiac chamber of the patient, and wherein the impedance values vary as a function of blood pressure within the at least one vessel or chamber;determining a plurality of values of an impedance parameter over time based on the measured impedance values, wherein each of the impedance parameter values is determined based on a respective sub-plurality of the impedance values, determining the plurality of values of the impedance parameter comprising identifying an R-wave for a current cardiac cycle,identifying a maximum impedance and a minimum impedance for the current cardiac cycle,determining a first time interval between the R-wave and the minimum impedance,determining a second time interval between the R-wave and the maximum impedance, anddetermining a ratio of the first time interval and the second time interval;comparing the ratio of the current cardiac cycle to a ratio of a prior cardiac cycle; andidentifying a change in a cardiovascular parameter related to the blood pressure based on the comparison. 2. The method of claim 1, further comprising triggering an alert in response to the change in the cardiovascular parameter. 3. The method of claim 1, further comprising modifying a therapy delivered to the patient in response to the change in the cardiovascular parameter. 4. The method of claim 1, wherein determining a plurality of values of an impedance parameter further comprises identifying a fiducial point of the plurality of impedance values for the current cardiac cycle, and determining a third time interval between the R-wave and the fiducial point, and comparing the third time interval for the current cardiac cycle to a third time interval for a previous cardiac cycle. 5. A non-transitory computer-readable medium comprising instructions for causing a programmable processor to: measure a plurality of impedance values of a path within a patient over time, wherein the path includes at least one blood vessel or cardiac chamber of the patient, and wherein the impedance values vary as a function of blood pressure within the at least one vessel or chamber;determine a plurality of values of an impedance parameter over time based on the measured impedance values, wherein each of the impedance parameter values is determined based on a respective sub-plurality of the impedance values, wherein the instructions for causing the programmable processor to determine the plurality of values of the impedance parameter comprise instructions that cause the programmable processor to: identify an R-wave for a current cardiac cycle,identify a minimum impedance and a maximum impedance for the current cardiac cycle,determine a first time interval between the R-wave and the minimum impedance,determine a second time interval between the R-wave and the maximum impedance, anddetermine a ratio of the first ti... | Medtronic, Inc. | A61B005/02
US07904143 | B2 | 2011-03-08 | Binary logistic mixed model for complex fractionated atrial electrogram procedures | 2008-07-07 | Methods and systems are provided for locating arrhythmogenic cardiac focus. A training set of electrical signal data is acquired and automatically analyzed to identify CFAE locations. The data representing the CFAE locations are and subjected to determinations of medical significance, such as ablation decisions, by a battery of physicians. The decisions and descriptive statistics of the data are fitted to a binary logistic mixed model, in which the physician decisions are treated as random effects. The model is then applied to patient CFAE data to establish estimates of medical significance, such as suitability for ablation. | 1. A method for locating an arrhythmogenic focus in a heart of a living subject, comprising the steps of: obtaining training electrical signal data from respective training locations of a training set of hearts;automatically analyzing said training electrical signal data to identify training complex fractionated electrograms therein;recording determinations of medical significance of a plurality of observers with respect to said training complex fractionated electrograms;generating first estimations of medical significance at said respective training locations by fitting a mixed regression model to said training complex fractionated electrograms and said determinations of medical significance;obtaining patient electrical signal data from respective locations of a patient heart;automatically analyzing said patient electrical signal data to identify patient complex fractionated electrograms in said patient heart;applying said mixed regression model to said patient complex fractionated electrograms to obtain second estimations of medical significance at said respective locations of said patient heart; anddisplaying an indication that one or more of said respective locations of said patient heart are medically significant. 2. The method according to claim 1, wherein said mixed regression model is a binary logistic mixed model having fixed effects and random effects, and wherein said fixed effects comprise a coefficient of variation and a range of said training electrical signal data, and said random effects comprise said first estimations of medical significance. 3. The method according to claim 1, wherein said observers comprise five observers. 4. The method according to claim 1, wherein said determinations of medical significance comprise first decisions to ablate primary ablation locations selected from said respective training locations and second decisions to ablate secondary ablation locations selected from said respective training locations. 5. A computer software product for locating an arrhythmogenic focus in a heart of a living subject, including a computer storage medium in which computer program instructions are stored, which instructions, when executed by a computer, cause the computer to acquire training electrical signal data from respective training locations of a training set of hearts, automatically analyze said training electrical signal data to identify training complex fractionated electrograms therein, record determinations of medical significance of a plurality of observers with respect to said training complex fractionated electrograms, generate first estimations of medical significance at said respective training locations by fitting a mixed regression model to said training complex fractionated electrograms and said determinations of medical significance, obtain patient electrical signal data from respective locations of a patient heart, automatically analyze said patient electrical signal data to identify patient comple... | Biosense Webster, Inc. | A61B005/046
US08285352 | B2 | 2012-10-09 | Systems and methods for identifying pulse rates | 2008-10-10 | According to embodiments, techniques for using continuous wavelet transforms and spectral transforms to identify pulse rates from a photoplethysmographic (PPG) signal are disclosed. According to embodiments, candidate pulse rates of the PPG signal may be identified from a wavelet transformed PPG signal and a spectral transformed PPG signal. A pulse rate may be determined from the candidate pulse rates by selecting one of the candidate pulse rates or by combining the candidate pulse rates. According to embodiments, a spectral transform of a PPG signal may be performed to identify a frequency region associated with a pulse rate of the PPG signal. A continuous wavelet transform of the PPG signal at a scale corresponding to the identified frequency region may be performed to determine a pulse rate from the wavelet transformed signal. | 1. A method for determining a pulse rate from a photoplethysmographic (PPG) signal, comprising: using a processor for: receiving a PPG signal;performing a continuous wavelet transform of the PPG signal to produce a wavelet transformed signal;performing a spectral transform of the PPG signal to produce a spectrum;identifying a first candidate pulse rate of the PPG signal from the wavelet transformed signal;identifying a second candidate pulse rate of the PPG signal from the spectrum; anddetermining a pulse rate from the first candidate pulse rate and the second candidate pulse rate. 2. The method of claim 1 wherein the spectral transform comprises a Fourier transform. 3. The method of claim 1 wherein determining the pulse rate from the first candidate pulse rate and the second candidate pulse rate comprises selecting one of the candidate pulse rates. 4. The method of claim 3 wherein selecting one of the candidate pulse rates comprises selecting the candidate pulse rate based at least in part on its source. 5. The method of claim 3 wherein selecting one of the candidate pulse rates comprises selecting the candidate pulse rate based at least in part on its value. 6. The method of claim 3 further comprising determining confidence values associated with each of the first and the second candidate pulse rates and wherein selecting one of the candidate pulse rates comprises selecting the candidate pulse rate based at least in part on the confidence values. 7. The method of claim 1 wherein determining the pulse rate from the first candidate pulse rate and the second candidate pulse rate comprises combining the candidate pulse rates. 8. The method of claim 7 further comprising determining confidence values associated with each of the first and the second candidate pulse rates and wherein combining the candidate pulse rates comprises combining the candidate pulse rates weighted by the confidence values. 9. A system for determining a pulse rate from a photoplethysmographic (PPG) signal, comprising: a wavelet processor capable of performing a continuous wavelet transform of a received PPG signal to produce a wavelet transformed signal and identifying a first candidate pulse rate of the PPG signal from the wavelet transformed signal;a spectral processor capable of performing a spectral transform of the received PPG signal to produce a spectrum and identifying a second candidate pulse rate of the PPG signal from the spectrum; anda pulse rate discriminator capable of determining a pulse rate from the first candidate pulse rate and the second candidate pulse rate. 10. The system of claim 9 wherein the spectral transform comprises a Fourier transform. 11. The system of claim 9 wherein the pulse rate discriminator is capable of selecting one of the candidate pulse rates. 12. The system of claim 11 wherein the pulse rate discriminator is capable of selecting one of the candidate pulses based at least in part on its source. 13. The system of claim 11 wher... | Nellcor Puritan Bennett LLC | A61B005/024 A61B005/1455 G01J003/433 G01N033/48
US08712509 | B2 | 2014-04-29 | Virtual physician acute myocardial infarction detection system and method | 2008-07-25 | A system and method for virtually detecting a medical condition, such as acute myocardial infarction (AMI), in a patient using holistic diagnostic procedures implemented in medical devices. Physiological parameters in a patient are monitored in an implantable medical device (IMD) to detect deviations from desired characteristics. When severe physiological parameter deviations exist indicating with a desired certainty that the patient is experiencing a medical condition (e.g., AMI), an alert is generated. If only minor deviations from the desired characteristics exist, additional holistic diagnostic procedures are performed virtually for diagnosing whether the patient is likely to be experiencing the medical condition, such as querying the patient through an external device regarding symptoms the patient is experiencing and analyzing the patient's responses to the questions to determine whether the patient is experiencing the medical condition. When diagnosis is still unclear, additional holistic diagnostic procedures can be performed virtually until diagnosis is certain. | 1. A virtual physician system comprising: a medical device implantable in a patient and configured to monitor at least one physiological parameter from sensed physiological signals for deviations from desired characteristics;the implantable medical device configured to: determine whether sufficient deviations are present in the monitored physiological parameter(s) that indicate with a desired level of certainty that the patient is experiencing a certain medical condition;generate an alert when the desired level of certainty that the patient is experiencing the medical condition is determined to exist;determine whether minor deviations are present in the monitored physiological parameter that potentially indicate that the patient is experiencing the medical condition without reaching with the desired level of certainty that the patient is experiencing the medical condition, wherein the minor deviations are not sufficient to indicate with the desired level of certainty that the patient is experiencing the certain medical condition;cause additional holistic diagnostic procedures to be performed for diagnosing whether the patient is experiencing the medical condition when minor deviations are determined;cause, as at least part of the additional holistic diagnostic procedures, an external device to query the patient regarding symptom-related questions and collect responses to the symptom-related questions from the patient; andcause, when the minor deviations are determined to be present, a determination of whether the patient is experiencing the medical condition based on the presence of the minor deviations and the collected responses to the symptom-related questions from the patient. 2. The virtual physician system of claim 1, further comprising the external device, wherein the external device is configured to communicate with the implantable medical device, the external device including an interface for allowing information to be communicated to and from the patient, the implantable medical device further configured for transmitting a command to the external device to query the patient regarding symptom-related questions. 3. The virtual physician system of claim 2, wherein the external device is further configured to collect responses to the symptom-related questions and analyze the responses to the symptom-related questions in view of the detected minor physiological parameter deviations to diagnose whether the patient is experiencing the medical condition. 4. The virtual physician system of claim 3, wherein, when the symptom-related information collected from the patient fails to indicate with the desired probability that the patient is experiencing the medical condition, the external device is further configured to cause at least one additional physiological parameter diagnostic test to be performed to determine whether the patient is experiencing the medical condition. 5. The virtual physician system of claim 4, wherein the implantable medi... | Medtronic, Inc. | A61B005/0452 A61B005/00
US08882669 | B2 | 2014-11-11 | Computer program product, device and method for measuring the arousal of a user | 2008-02-11 | Computer program product for processing GSR (galvanic skin response) signals, which when run on a computer controls the computer to estimate a level of arousal, or at least a change in the level of arousal, of a user, provided with an algorithm that is configured to control a computer to calculate a third or higher central moment of a recorded GSR signal and to derive an estimation of the level of arousal from said third or higher central moment. | 1. A non-transitory computer-readable medium encoded with instructions configured to control a processor of an arousal measuring device to perform a method comprising: receiving galvanic skin response signals in the processor of an arousal measuring device from a galvanic skin response sensor of the arousal measuring device;calculating a third or higher central moment of a galvanic skin response signal, wherein the third or higher order moment comprises at least one of: a skewness and a kurtosis;estimating at least one of: a level of arousal or a change in the level of arousal, wherein said estimation of the level of arousal is based on said calculated third or higher central moment exceeding a corresponding one of a plurality of threshold values, and wherein the estimation of the change in the level of a arousal is determined when one of:an increase in the level of arousal is determined when at least one of: a relatively low skewness or a relatively high kurtosis is calculated, ora decrease in the level of arousal is determined when at least one of: a relatively high skewness or a relatively low kurtosis is calculated; andoutputting a signal to a user perceivable interface when one of: a level of arousal or a change in the level of arousal are estimated. 2. The medium of claim 1, wherein the medium comprises instructions to control the processor to signal to a user perceivable interface when at least one of: said change in the level of arousal is determined, orthe level of arousal is derived. 3. The medium of claim 2, wherein the signal to the user perceivable interface is performed approximately immediately in real-time. 4. The medium of claim 1, wherein the increase in the level of arousal is derived when the calculated skewness is below a threshold of approximately −0.3. 5. The medium of claim 1, wherein the increase in the level of arousal is derived when the calculated kurtosis is above a threshold of approximately 0.3. 6. An arousal measuring device for processing galvanic skin response signals to estimate at least one of: a level of arousal or a change in the level of arousal of a user, said device comprising: a galvanic skin response sensor for providing galvanic skin response signals from a user body;a processing circuit for processing the provided galvanic skin response signals;a storage arrangement for storing at least one of the processed galvanic skin response signals;wherein the processing comprises:calculating a third or higher central moment of a recorded galvanic skin response signal, wherein the third or higher order moment comprises at least one of: a skewness and a kurtosis;estimating at least one of: a level of arousal or a change in the level of arousal, wherein said estimation of the level of arousal is based on said calculated third or higher central moment exceeding a corresponding one of a plurality of threshold values, and wherein said change in the level of arousal is determined when one of:an increase in the lev... | Koninklijke Philips N.V. | A61B005/02
US08452357 | B2 | 2013-05-28 | Non-invasive methods of using spectral information in determining analyte concentrations | 2009-10-08 | A non-invasive method of determining the concentration of an analyte uses Raman or fluorescence spectral information. A high-intensity band of light is applied to one side of skin tissue. The high-intensity light enters the skin tissue and generates a Raman or fluorescence signal. A Raman-generating material or fluorescence-generating material is placed in a location nearest the other side of skin tissue. The Raman-generating or fluorescence-generating material is located generally opposite of the entry of the applied high-intensity light. The Raman or fluorescence signal is collected and the analyte concentration is determined using the collected Raman signal. | 1. A non-invasive method of determining the concentration of an analyte using Raman spectral information, the method comprising the acts of: applying a high-intensity, narrow band of light to a first side of tissue, the high-intensity light entering the tissue at a first side and generating a Raman signal;providing a Raman-generating material adapted to generate Raman signals;reflecting the high-intensity light from the Raman-generating material so as to produce an additional Raman signal that passes towards the first side of the tissue;reflecting the Raman signal generated from the high-intensity light entering the tissue towards the first side of the tissue via the Raman-generating material;collecting the Raman signal generated from the high-intensity light entering the tissue and the additional Raman signal generated from the Raman-generating material; anddetermining the analyte concentration using information from the collected Raman signals. 2. The method of claim 1, wherein the high-intensity light is applied from a monochromatic light source. 3. The method of claim 2, wherein the high-intensity light is applied from a laser-diode source. 4. The method of claim 1, wherein the analyte is glucose. 5. The method of claim 1, wherein the Raman signal is collected at a wavelength range of from about 300 to about 5000 nm. 6. The method of claim 5, wherein the Raman signal is collected at a wavelength range of from about 830 to about 1030 nm. 7. The method of claim 1, wherein the information from the collected Raman signals assists in determining the level of optical scattering and absorbance of the tissue, the information being used to assist in determining the analyte concentration. 8. The method of claim 1, wherein the information from the collected Raman signals assists in determining the level of optical scattering of the tissue, the information being used to assist in determining the analyte concentration. 9. The method of claim 1, wherein the information from the collected Raman signals assists in determining the level of absorbance of the tissue, the information being used to assist in determining the analyte concentration. 10. The method of claim 1 further including a spatial filter for blocking Raman signals in one portion from being collected and for allowing Raman signals in another portion to be collected. 11. The method of claim 10, wherein the spatial filter includes a plurality of apertures. 12. The method of claim 11, wherein determining the analyte concentration further includes using information from the optical scattering of the Raman signals through the plurality of apertures. 13. The method of claim 10, wherein the spatial filter includes at least two sections that are moveable with respect to each other. 14. A non-invasive method of diagnosis using Raman spectral information, the method comprising the acts of: applying a high-intensity, narrow band of light to a first side of tissue, the high-intensity light ente... | Bayer HealthCare LLC | A61B005/1455
US07953472 | B2 | 2011-05-31 | Method and apparatus for treatment of discogenic pain | 2009-09-17 | A method of treating discogenic back pain in a patient includes the steps of acquiring an image of at least a portion of the patient's spine, locating the spinal disc causing the back pain, and identifying a treatment volume encompassing at least a portion of the spinal disc. The method further includes prescribing a radiation dose to be delivered to the treatment volume, providing a radiation therapy machine, positioning the patient in the radiation therapy machine, programming the radiation therapy machine to deliver the radiation dose to the treatment volume, and delivering the radiation dose to the treatment volume. | 1. Non-transitory computer readable media with computer-executable instructions embodied thereon that when executed by a computer control a device for treating back pain, the media comprising: instructions for locating an spinal disc and defining a treatment volume encompassing at least a portion of one or more nerves transmitting impulses that result in the sensation of the back pain based on an image;instructions for creating a treatment plan for delivering a radiation dose to the treatment volume to decrease or stop the transmission of the impulses by the one or more nerves within the treatment volume; andinstructions for configuring the device to deliver the radiation dose to the treatment volume. 2. The non-transitory computer readable media of claim 1, further comprising: instructions for identifying a radiosensitive area proximate to the treatment volume. 3. The non-transitory computer readable media of claim 2, further comprising: instructions for configuring the treatment plan to avoid the radiosensitive area when delivering the radiation dose. 4. The non-transitory computer readable media of claim 3, wherein the radiosensitive area is avoided by delivering the radiation dose through one or more treatment portals extending through non-radiosensitive areas to reach the treatment volume. 5. The non-transitory computer readable media of claim 1, wherein the device for treating back pain is a radiation therapy machine. 6. The non-transitory computer readable media of claim 1, wherein the device for treating back pain comprises a medical imaging device and further comprising instructions for acquiring the image. 7. The non-transitory computer readable media of claim 6, wherein the medical imaging device is at least one of a CT scanner, a MRI scanner, a PET scanner, or a SPECT scanner. 8. The non-transitory computer readable media of claim 1, wherein the treatment volume encompasses at least a portion of the annulus fibrosus of the spinal disc. 9. The non-transitory computer readable media of claim 1, wherein the treatment volume encompasses at least a portion of the posterior wall of the spinal disc. 10. A discogenic back pain treatment device, comprising: a patient support;a radiation source;a radiation delivery mechanism coupled to the radiation source; anda treatment planning computer adapted to control the radiation delivery mechanism, the treatment planning computer programmed to deliver a radiation dose from the radiation source to a treatment volume by executing program instructions, the program instructions comprising: instructions for locating a spinal disc and defining the treatment volume to encompass at least a portion of the spinal disc;instructions for creating a back pain treatment plan for delivering the radiation to the treatment volume encompassing at least a portion of one or more nerves transmitting impulses that result in the sensation of the back pain based on an image; andinstructions for configuring the radiat... | Spinal Generations, LLC | A61B005/05 A61N005/10
US08233958 | B2 | 2012-07-31 | Signal processing for continuous analyte sensor | 2009-10-12 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A method for compensating for a time lag of continuous analyte sensor data, the method comprising: receiving sensor data from a continuous analyte sensor, wherein the sensor data comprises a) a physiological time lag with respect to blood analyte data or reference analyte data, b) a membrane-induced time lag, or c) a computationally-induced time lag;continuously and/or intermittently estimating, using a processor module, analyte values based on the sensor data, wherein estimating comprises compensating analyte values based on a first time lag associated with the sensor data, evaluating analyte values compensated for the first time lag with time-corresponding received analyte values to determine an updated time lag associated with the sensor data, and compensating analyte values based on the updated time lag associated with the sensor data; andcontinuously or intermittently providing, using an output module, an output based on the estimated analyte values. 2. The method of claim 1, wherein the step of estimating analyte values based on the sensor data comprises predicting one or more analyte values for a time period associated with the first or updated time lag, wherein the time period is from about 1 second to about 60 minutes. 3. The method of claim 2, wherein the time period is from about 15 minutes to about 20 minutes. 4. The method of claim 1, further comprising a step of forming at least one matched data pair by matching the first or updated time-lag compensated estimated analyte values with reference data. 5. The method of claim 1, wherein the continuous analyte sensor is continuous glucose sensor. 6. The method of claim 5, wherein the continuous glucose sensor comprises a working electrode at least partially implanted in a host over a time period of continuous implantation. 7. A system for compensating for a time lag of continuous analyte sensor data, the system comprising: an input module operatively connected to the continuous analyte sensor that receives sensor data, wherein the sensor data comprises a) a physiological time lag with respect to blood analyte data or reference analyte data, b) a membrane-induced time lag, and/or c) a computationally-induced time lag; anda processor module comprising programming configured to continuously and/or intermittently estimate one or more analyte values based on the sensor data including a first time lag compensation, evaluate analyte values compensated for the first time lag compensation with time-corresponding received anal e values to determine an updated time lag compensation associated with the sensor data and compensating analyte values based on the updated time lag compensation associated with the sensor data, wherein the processor module further comprises programming configured to continuously or intermittently provide an output based on the time-lag compensated estimated analyte values. 8. The system of claim 7, wherein the step of estimating one or more analyte values based on ... | DexCom, Inc. | A61B005/05 A61B005/00
US08295906 | B2 | 2012-10-23 | MRI guided radiation therapy | 2008-08-20 | Radiation therapy of a lesion within a patient is guided to take into account movement of the lesion caused by respiration and/or cardiac effects by using MRI to image the patient while on the treatment support to obtain a series of images of a location of the lesion within the patient while obtaining data relating to respiration and/or cardiac function of the patient to generate lesion movement data as a function of respiration and/or cardiac function. After the magnet is removed from the bunker, during the radiation therapy, real time data relating to respiration and/or cardiac function of the patient is used to guide the radiation therapy using the lesion movement data obtained during the imaging. | 1. A method for guiding radiation therapy of a patient comprising: providing a treatment support for receiving and supporting the patient during treatment;providing a radiation therapy treatment system for applying radiation therapy to the patient;locating the treatment support and the radiation therapy treatment system in a radiation bunker arranged to prevent escape of the radiation;providing a door on an opening into the bunker;providing an MRI system for imaging the patient including a cylindrical magnet;moving the radiation therapy treatment system in the bunker to a location spaced from the treatment support;locating a patient on the treatment support in the bunker, the patient having a lesion requiring radiation therapy;preparing the patient for radiation therapy on the treatment support;with the radiation therapy treatment system moved to the spaced location, moving the cylindrical magnet from a position outside the bunker through the opening with the door open into a position surrounding the treatment support for obtaining images of the patient while on the treatment support;the bunker, the door and the opening being shaped and arranged to allow the magnet to move into the bunker;while the patient is on the treatment support, using the MRI system to obtain a series of images of a location of the lesion within the patient;moving the magnet of the MRI system away from the treatment support and out of the bunker through the opening, and closing the door so as to allow the radiation therapy;and during the radiation therapy, using the images of the patient to guide the radiation therapy. 2. The method according to claim 1 wherein the images from the MR system in an MR coordinate system are correlated relative to a coordinate system of the radiation therapy by using the treatment support as a common baseline. 3. The method according to claim 1 wherein the radiation therapy is generated by a collimated radiation source which is rotated round the lesion in a manner which controls the application of a required dose of radiation to the lesion while accommodating the shape of the lesion and the movement of the lesion. 4. The method according to claim 1 including, while the patient is on the treatment support during the obtaining of the images, obtaining data relating to respiration and/or heartbeat of the patient, using the series of images to generate lesion movement data relating to a correlation between movement of the location of the lesion and the data relating to the respiration and/or heartbeat and during the radiation therapy, using real time data relating to respiration and/or heartbeat of the patient and guiding the radiation therapy using the lesion movement data correlated to the real time data of the respiration and/or heartbeat. 5. The method according to claim 4 wherein the lesion movement data relating to a correlation between movement of the location of the lesion and the data relating to the respiration and/or heartbeat is obt... | Imris Inc | A61B005/055
US07998161 | B2 | 2011-08-16 | Disposable incision safety lancet | 2008-10-22 | A disposable incision safety lancet is provided, including a casing (2) and a trigger. It is characterized in that the casing (2) is provided inside with an incision blood-taking mechanism composed of a cam (3), a main swing arm (4), a secondary swing arm (5), a blade (1) and a spring (8). The present invention combines the two arc-shaped tracks obtained from the two swing mechanisms, the main swing arm and the secondary swing arm, via the cam and a toggle pin provided on the cam effectively and interlockingly, thus making the point of the blade incise along a V-formed path composed of two arcs of different shape connected with each other. Compared with the prior art single-swing-arm structure of the incision blood-taking mechanism, such a double-swing-arm structure can move with higher accuracy and controllability. An incision of different blood-taking depth and blood-taking width can be designed by means of changing the difference of change of the radius of curvature of the flange of cam (3) as well as changing the swing radius of the main swing arm (4) and the secondary swing arm (5), which is especially suitable for taking blood from infants' heel. | 1. A disposable incision safety lancet, comprising: a casing (2) and a trigger;wherein:the casing (2) is provided inside with an incision blood-taking mechanism composed of a cam (3), a main swing arm (4), a secondary swing arm (5), a blade (1) and a spring (8);the cam (3) is rotatably fixed in the casing (2) via a pivot (6); the spring (8), as a driving element of the incision blood-taking mechanism, is positioned between the cam (3) and the casing (2), and acts on the rotary direction of the cam (3); the trigger, as a triggering control member of the incision blood-taking mechanism, is positioned on the casing (2); in a pre-triggering state, an action portion (10) of the trigger is propped against a lock gate (7) provided on the cam (3), and forces the spring (8) to be in an energy storage state, with a triggering portion (11) of the trigger being stretched out of the casing (2);the main swing arm (4), as the first swing arm, is provided with a locating end (12) and a swing end (13), the locating end (12) being positioned in the casing (2), the swing end (13) being hung in the casing (2); the swing end (13) is provided with a working face (14), which is contactively matched with flange of the cam (3); the flange of the cam (3) is provided, corresponding to this working face (14), with a curved surface (15), whose curvature radius is transient from a first radius (R1) to a second radius (R2); this curved surface (15) forces the main swing arm (4) to make a first swing with rotation of the cam (3), with the distance from center of the locating end (12) itself to the point of the blade (1) as a third radius (R3);the secondary swing arm (5), as the second swing arm, is provided with a locating end (16) and a swing end (17); the locating end (16) of the secondary swing arm (5) is connected with the swing end (13) of the main swing arm (4) fittedly and rotatably via a pin and a hole, and an initial rotation-angle orientation is realized by means of a contactively-propped structure provided mutually and correspondingly in the rotation direction; the blade (1) is fixedly mounted on the swing end (17) of the secondary swing arm (5); the secondary swing arm (5) is provided in the rotary direction with a toggle face (18), corresponding to which is a toggle pin (19) provided on the cam (3); this toggle pin (19) acts on the toggle face (18) on the secondary swing arm (5) with rotation of the cam (3), and forces the secondary swing arm (5) to make a second swing with the distance from center of the locating end (16) itself to the point of the blade (1) as a fourth radius (R4);the blade (1) is located at a blood-taking opening (20) provided on the casing (2); when the incision blood-taking mechanism makes the first swing, the point of the blade (1) is swingingly stretched out of the casing (2) from the blood-taking opening (20) along a first arc-shaped path (A); when the incision blood-taking mechanism makes the second swing, the point of the blade (1) is swi... | - | A61B005/151
US08301219 | B2 | 2012-10-30 | Patient monitoring systems and methods | 2009-07-09 | Various systems and methods are provided for monitoring patient physiological data without the need for adhesives or wires to be connected to a patient. The system is also fully functional without the need for manipulation or intervention by a care provider or medical personnel. In particular, the system can include a pad having a plurality of electrical contacts or electrodes formed thereon in a predetermined pattern. The electrodes can be configured to sense electrical signals produced by a patient's body. A controller in communication with the electrodes can be configured to select and process signals from the electrodes and send the information derived from the electrical signals to an output device for displaying resulting physiological data so that the patient can be monitored. | 1. A system for monitoring patient physiological data, comprising: a pad having a plurality of electrical contacts formed thereon in a predetermined pattern, the electrical contacts being configured to sense electrical signals produced by a patient's body;a controller in communication with the electrical contacts and configured to select at least two electrical contacts, each sensing a processable electrical signal, and to process a differential signal derived from the processable signals; andan output device in communication with the controller for displaying the patient's physiological data based on information processed by the controller;wherein the controller is further configured to actively change the selection of the at least two electrical contacts over time by periodically determining which of the plurality of electrical contacts are in contact with the patient and changing the selection of the at least two electrical contacts based on an assessment of the quality of the sensed electrical signals. 2. The system of claim 1, wherein the plurality of electrical contacts comprise strips of conductive material. 3. The system of claim 2, wherein the strips of conductive material are formed into two mirrored sets of nested L-shaped configurations on the pad. 4. The system of claim 2, wherein the strips of conductive material are selected from the group consisting of a nickel/copper polyester fabric and a cloth fabric containing conductive threads. 5. The system of claim 1, wherein at least one of the plurality of electrical contacts is a devoted active ground configured to enable removal of 60 Hz noise from the signal processed by the controller. 6. The system of claim 1, wherein the plurality of electrical contacts are configured to sense electrical signals through the patient's skin to produce physiological data regardless of patient size or patient position on the pad. 7. The system of claim 1, wherein the controller includes a multiplexing network, at least one amplifier, at least one filter, and at least one microprocessor. 8. The system of claim 1, wherein the controller is battery operated. 9. The system of claim 1, wherein the output device is configured to communicate with the controller wirelessly. 10. The system of claim 1, wherein the pad is selected from the group consisting of a gurney pad, a sheet, a warming blanket, and a hospital gown. 11. The system of claim 1, wherein the electrical signals sensed by the plurality of electrical contacts are electrocardiograph signals and physiological data displayed by the output device is an electrocardiograph. 12. The system of claim 1, wherein the controller is further configured to select the at least two electrical contacts such that the at least two electrical contacts span an organ producing the electrical signals. 13. A system for monitoring patient physiological data, comprising: a pad having an array of electrodes formed thereon in a predetermined configuration, the pad ... | The General Hospital Corporation Massachusetts Institute of Technology | A61B005/04 A61B005/0428
US08183453 | B2 | 2012-05-22 | Interventive-diagnostic device | 2009-04-21 | Apparatus for improving health of a user is provided, including a first sensor, adapted to measure a first physiological variable, which is indicative of a voluntary action of the user. A second sensor is adapted to measure a second physiological variable, which is substantially governed by an autonomic nervous system of the user. Circuitry is adapted to receive respective first and second sensor signals from the first and second sensors, and, responsive thereto, to generate an output signal which directs the user to modify a parameter of the voluntary action. | 1. A method for generating music, comprising: receiving a rhythm signal corresponding to a rhythm of a cyclic physiological activity of a user, the physiological activity having first and second activity phases thereof;analyzing the rhythm signal to determine first and second durations thereof, respectively corresponding to the first and second activity phases;determining first and second new durations responsive to desired changes of the first and second durations of the rhythm signal;generating responsive to the new durations a music signal for presentation to the user, the music signal having first and second music phases thereof respectively corresponding to the first and second activity phases, a duration of each of the music phases expressible as being approximately equal to an integer multiple of a base duration, the integer multiple being less than or equal to four; anddirecting the user to modify durations of the first and second activity phases responsive to the respective durations of the first and second music phases. 2. A method according to claim 1, wherein generating the music signal comprises setting the duration of one of the music phases to be approximately equal to an integer multiple of the other one of the music phases. 3. A method according to claim 1, wherein directing the user to modify the durations comprises directing the user to attempt to perform the first and second activity phases of the physiological activity such that the respective durations thereof are substantially equal to the durations of the first and second music phases. 4. A method according to claim 1, wherein receiving the rhythm signal comprises receiving a motion signal corresponding to an activity of the user selected from the list consisting of: walking, jogging, and running. 5. A method according to claim 1, wherein receiving the rhythm signal comprises receiving a respiration signal corresponding to respiration of the user. 6. A method according to claim 5, wherein receiving the breathing signal comprises receiving an indication of a timing characteristic of inspiratory and expiratory phases of the respiration. 7. A method according to claim 1, wherein determining the new durations comprises determining the new durations responsive to a vasomotor frequency of the user. 8. A method according to claim 7, and comprising measuring a cardiovascular variable of the user and determining the vasomotor frequency responsive thereto. 9. Apparatus for generating music, comprising: a sensor, adapted to receive a rhythm signal corresponding to a rhythm of a cyclic physiological activity of a user, the physiological activity having first and second activity phases thereof;a processor, adapted to analyze the rhythm signal to determine a frequency thereof and to determine a new frequency responsive to a desired change of the frequency of the rhythm signal; andcircuitry, adapted to: generate at the new frequency a music signal for presentation to the user, th... | Intercure Ltd. | G10H001/00
US08233957 | B2 | 2012-07-31 | Sensor module for a catheter | 2008-01-11 | A sensor module (130) for a catheter (110), the sensor module (130) comprising a biofilm detection unit (131) adapted for detecting a characteristic of a biofilm (132) and electric circuitry (135, 800) for providing an output signal indicative of a result of the detection. | 1. A catheter, the catheter comprising: a tubular element enclosing a lumen;a plurality of sensor modules spread out spatially along the length of the tubular element, wherein each of the sensor modules comprises: a biofilm detection unit configured to detect a characteristic of a biofilm, wherein the biofilm detection unit includes a first sensor located at an inner surface of a tubular wall of the catheter and a second sensor located at an outer surface of the tubular wall of the catheter;electric circuitry configured to provide an output signal indicative of a result of the detection,wherein one of the sensor modules comprises a processing unit integrated in a tubular wall of the tubular element sandwiched between the first sensor and the second sensor, and wherein the plurality of sensor modules enable determination of a spatial dependence in the formation of the biofilm. 2. The catheter of claim 1, wherein the processing unit is configured to process the characteristic for generating the output signal. 3. The catheter of claim 2, wherein the processing unit is configured to generate a probability information regarding a reliability of the detected characteristic of the biofilm. 4. The catheter of claim 2, wherein the processing unit is configured to generate the output signal continuously or at regular time intervals. 5. The catheter of claim 2, wherein the processing unit is configured to generate the output signal comprising a unique identifier. 6. The catheter of claim 1, wherein at least one biofilm detection unit comprises an ion sensitive field effect transistor. 7. The catheter of claim 1, wherein at least one biofilm detection unit comprises an impedance sensor. 8. The catheter of claim 7, wherein the impedance sensor is configured to perform a measurement between at least one first electrode integrated in a substrate and at least one second electrode integrated in the substrate or located in a medium surrounding the sensor module. 9. The catheter of claim 7, wherein the impedance sensor is configured to perform a measurement between four electrodes arranged in a configuration with two outer electrodes and two inner electrodes arranged between the two outer electrodes;wherein the impedance sensor is configured to apply an alternating electric signal between the outer electrodes and is configured to measure an electric response signal between the two inner electrodes indicative of the impedance value. 10. The catheter of claim 1, comprising a communication unit configured to transmit, in a wired or wireless manner, the output signal to a base station. 11. The catheter of claim 1, wherein at least a part of each sensor module is monolithically integrated in semiconductor technology. 12. The catheter of claim 1, wherein each of the sensor modules comprises a temperature detection unit configured to detect a temperature signal; wherein the electric circuitry of the sensor module is configured to provide an output signal indica... | NXP B.V. | A61B005/05
US08353846 | B2 | 2013-01-15 | Method and apparatus for tinnitus evaluation | 2007-07-17 | Methods, articles of manufacture, and systems for evaluating tinnitus are disclosed herein. According to various embodiments, a tinnitus evaluation system may include a tinnitus evaluation module configured to perform one or more tinnitus evaluation tests. A tinnitus evaluation test may comprise generating a first single-frequency sound based at least in part on a sound of a patient's tinnitus, and generating a narrow-band sound centered at the frequency of the first single-frequency sound. Tests include a hearing threshold test, a loudness match test, a pitch match test, a bandwidth match test, a minimum masking level test, and a residual inhibition test. | 1. An apparatus for evaluating tinnitus, comprising: at least one stimulus generator configured to generate a single-frequency sound based at least in part on a pitch of a patient's tinnitus, the stimulus generator further configured to generate a first multi-frequency sound centered at the frequency of the single-frequency sound; anda user interface configured to receive input of an indication of which one of the single-frequency sound and the first multi-frequency sound is more similar to a sound of the patient's tinnitus. 2. The apparatus of claim 1, wherein the stimulus generator is configured to generate octave band pass noises and fixed-width multi-frequency noises. 3. The apparatus of claim 1, wherein the stimulus generator comprises one or more devices selected from a group consisting of a programmable oscillator, a programmable state variable filter, a programmable noise source, and an analog multiplier. 4. The apparatus of claim 1, further comprising a computing apparatus interface for coupling the apparatus to a computing apparatus. 5. The apparatus of claim 4, wherein the stimulus generator is configured to generate sounds based at least in part on instructions received from the computing apparatus. 6. The apparatus of claim 1, wherein the user interface is configured to allow the patient to selectively control the sounds generated by the stimulus generator. 7. The apparatus of claim 1, further comprising a signal preamplifier configured to selectively modify sounds generated by the stimulus generator. 8. The apparatus of claim 7, further comprising at least one signal attenuator configured to attenuate the modified sounds. 9. The apparatus of claim 8, wherein the apparatus comprises two signal attenuators including a first signal attenuator configured to output attenuated sounds to a right earphone and further including a second signal attenuator configured to output attenuated sounds to a left earphone. 10. The apparatus of claim 1, wherein the apparatus is configured to perform a bandwidth match test and at least one test selected from the group consisting of a hearing threshold test, a loudness match test, a pitch match test, a minimum masking level test, and a residual inhibition test. 11. A system for evaluating tinnitus, comprising: a tinnitus evaluation module adapted to generate one or more sounds; anda computing apparatus coupled to the tinnitus evaluation module and configured to: cause the tinnitus evaluation module to output a single-frequency sound based at least in part on a pitch of a patient's tinnitus;cause the tinnitus evaluation module to output a first multi-frequency sound centered at the frequency of the single-frequency sound; anddetermine, based on a first input from the patient, which one of the single-frequency sound and the first multi-frequency sound is more similar to a sound of the patient's tinnitus. 12. The system of claim 11, wherein the computing apparatus is further configured to cause th... | Oregon Health & Science University The United States of America as represented by The Department of Veterans Affairs | A61B005/00
US08721558 | B2 | 2014-05-13 | Device for determining the flow rate of a blood flow, and cardiovascular assist device | 2009-07-24 | A device for the in vivo determination of the blood flow rate in a patient's blood vessel includes a microelectrode arrangement provided for placement in the blood vessel, an electrical power source which provides excitation energy having physiologically harmless parameters for obtaining a measured signal, a signal detector for detecting an electrical measured signal resulting from the blood flow in the presence of the excitation energy at measuring electrodes of the microelectrode arrangement, and a signal evaluation device, connected to the signal detector, for determining the blood flow rate on the basis of the measured signal. | 1. An implantable cardiovascular assist device providing in vivo determination of the blood flow rate in a bloodstream of a patient, the device including: a. an elongated electrode line including one or more stimulation electrodes for delivering stimulation pulses to a heart,b. multiple sets of microelectrodes configured for placement in the bloodstream, (1) the sets of microelectrodes being positioned spaced apart on the surface of the electrode line separate and spaced from the one or more stimulation electrodes,(2) each set of microelectrodes including at least one excitation electrode and at least one measuring electrode, each excitation electrode being spaced from one or more of the measuring electrodes by an elongated gap which is narrower than the electrode lengths of the spaced excitation and measuring electrodes, with the gap and the electrode lengths being measured along an axis bisecting the electrodes,c. an electrical power source connected to the excitation electrodes, the power source being configured to provide excitation energy,d. a signal detector wherein: (1) each set of microelectrodes is independently connected to the signal detector, whereby each set delivers independently obtained measured signals to the signal detector, and(2) in the presence of the excitation energy, the signal detector detects an electrical measured signal at the measuring electrodes resulting from the blood flow, ande. a signal evaluation device in communication with the signal detector, the signal evaluation device being configured to determine the blood flow rate on the basis of the measured signal. 2. The device of claim 1 wherein: a. the microelectrodes include at least two measuring electrodes,b. the power source includes an alternating voltage source wherein alternating voltage is provided as excitation energy to the excitation electrodes, andc. the signal detector performs time-dependent measurement of a complex impedance between two or more measuring electrodes. 3. The device of claim 1 wherein the power source includes a pulse generator for generating an excitation pulse sequence, wherein the excitation pulse sequence is provided as excitation energy to the excitation electrodes. 4. The device of claim 1 wherein at least one of the electrodes is switchable to alternate between service as a measuring electrode and an excitation electrode. 5. The device of claim 1 wherein at least one of the excitation electrode and the measuring electrode: a. is bounded by a perimeter, wherein the electrode has a geometric electrode area defined within the perimeter; andb. has a surface area configured to be in contact with the blood flow wherein the surface area is at least twice as large as the geometric electrode area. 6. The device of claim 1 wherein the microelectrodes collectively have a surface area between 300 and 2000 square micrometers. 7. The device of claim 1 wherein: a. the microelectrodes include at least two measuring electrodes,b. each measu... | BIOTRONIK CRM Patent AG | A61B005/02
US08870787 | B2 | 2014-10-28 | Ventricular shunt system and method | 2009-07-24 | A ventricular shunt systems and methods of preventing hydrocephalus are described herein. In one aspect, the ventricular shunt system has at least one pressure sensor that is configured to be selectively electromagnetically coupled to an ex-vivo source of RF energy and is variable in response to the pressure in a patient's ventricle. | 1. A ventricular shunt system comprising: a catheter assembly comprising: a valve having an inlet port and an outlet port, the valve configured to control a rate of fluid flow between the inlet port and the outlet port;a ventricular catheter having a distal end and a proximal end, the proximal end of the ventricular catheter being coupled to the inlet port of the valve;a drainage catheter having a distal end and a proximal end, the proximal end of the drainage catheter being coupled to the outlet port of the valve;at least one pressure sensor mountable on a portion of the catheter assembly comprising a passive electrical resonant circuit that is configured to be selectively electromagnetically coupled to an ex-vivo source of radio frequency (RF) energy; wherein the passive electrical resonant circuit is variable in response to the pressure in a patient's ventricle, wherein each passive electrical resonant circuit, in response to the electromagnetic coupling, is configured to generate an output signal characterized by a frequency that is dependent upon urged movement of a portion of the passive electrical resonant circuit and is indicative of pressure in the patient's ventricle, wherein the passive electrical resonant circuit of the at least one pressure sensor comprises an inductance-capacitance (LC) resonant circuit that comprises an inductor operably coupled to a capacitor; anda sealed reservoir having a distal end positionable proximate the distal end of the ventricular catheter, the sealed reservoir configured to contain a non-compressible fluid, wherein a portion of the distal end of the reservoir is pliable,wherein the at least one pressure sensor is mountable to detect the pressure of the non-compressible fluid within the sealed reservoir, and wherein the at least one pressure sensor is mounted in a wall of the reservoir. 2. The ventricular shunt system of claim 1, wherein the inductor is a coil inductor. 3. The ventricular shunt system of claim 1, wherein the capacitance of the capacitor is variable in response to the pressure therein the patient's ventricle. 4. The ventricular shunt system of claim 1, wherein the capacitor comprises a first and second spaced capacitor elements, wherein at least one of the spaced capacitor elements are mounted in or on a flexible material, and wherein the respective first and second capacitor elements are movable relative to each other in response to pressure changes applied to the flexible material. 5. The ventricular shunt system of claim 1, wherein the inductor is configured to allow inductance in the passive electrical resonant circuit when the pressure sensor is subjected to a time variable electromagnetic field. 6. The ventricular shunt system of claim 1, wherein the inductance of the LC resonant circuit is between about 5 to about 15 micro-Henry. 7. The ventricular shunt system of claim 1, wherein the resonant frequency of the LC resonant circuit is between about 25 to about 45 MHz. 8. The v... | CardioMEMS, Inc. | A61B005/00 A61B005/03 G01R023/12 G01R023/14 H01Q007/00 A61B005/05 G01F015/06 A61M027/00 A61B005/07
US08077297 | B2 | 2011-12-13 | Methods and systems for discriminating bands in scalograms | 2008-10-03 | The present disclosure is directed towards embodiments of systems and methods for discriminating (e.g., masking out) scale bands that are determined to be not of interest from a scalogram derived from a continuous wavelet transform of a signal. Techniques for determining whether a scale band is not of interest include, for example, determining whether a scale band's amplitude is being modulated by one or more other bands in the scalogram. Another technique involves determining whether a scale band is located between two other bands and has energy less than that of its neighboring bands. Another technique involves determining whether a scale band is located at about half the scale of another, more dominant (i.e., higher energy) band. | 1. A method comprising: receiving a signal using a processor;generating a scalogram using the processor based at least in part on a continuous wavelet transform of the signal;determining whether at least one band of scales from the scalogram is not of interest using the processor; anddiscriminating the at least one band of scales if it is determined to be not of interest using the processor. 2. The method of claim 1, wherein the determining is performed at least in part by a mod-max discriminator. 3. The method of claim 1, wherein the determining comprises determining whether the at least one band's energy amplitude is being modulated by another band. 4. The method of claim 3, wherein the determining whether the at least one band's energy amplitude is being modulated by another band comprises: computing a transform of the amplitude of a profile ridge candidate; andcomparing the computed transform to at least one scale value of a source ridge candidate. 5. The method of claim 3, wherein the determining whether the at least one band's energy amplitude is being modulated by another band comprises filtering the signal at one scale of interest. 6. The method of claim 1, wherein the determining comprises determining whether the at least one band is a coupled band. 7. The method of claim 6, wherein the determining whether the at least one band is a coupled band comprises: computing the energy of a source ridge candidate;determining a threshold from the computed energy of the source ridge candidate;computing the energy of a profile ridge candidate; andcomparing the computed energy of the profile ridge candidate to the threshold. 8. The method of claim 6, wherein the coupled band is located on the scalogram between a first band of interest and a second band of interest. 9. The method of claim 1, wherein the determining comprises determining whether the at least one band is a half scale band. 10. The method of claim 1, wherein the discriminating comprises masking the at least one band. 11. A system for discriminating at least one band of scales, the system comprising: a processor capable of: receiving an input signal;generating a scalogram based at least in part on a continuous wavelet transform of the input signal;determining whether at least one band of scales from the scalogram is not of interest; anddiscriminating the at least one band of scales if it determined to be not of interest. 12. The system of claim 11, wherein the processor is coupled to a pulse oximeter. 13. The system of claim 12, wherein the input signal is a photoplethysmograph signal. 14. The system of claim 11, wherein the determining comprises determining whether the at least one band's energy amplitude is being modulated by another band. 15. The system of claim 11, wherein the determining comprises determining whether the at least one band is a coupled band. 16. The system of claim 11, wherein the determining comprises determining whether the at least one band is a half... | Nellcor Puritan Bennett Ireland | G01J003/433
US08469886 | B2 | 2013-06-25 | Signal processing for continuous analyte sensor | 2009-09-23 | Systems and methods for dynamically and intelligently estimating analyte data from a continuous analyte sensor, including receiving a data stream, selecting one of a plurality of algorithms, and employing the selected algorithm to estimate analyte values. Additional data processing includes evaluating the selected estimative algorithms, analyzing a variation of the estimated analyte values based on statistical, clinical, or physiological parameters, comparing the estimated analyte values with corresponding measure analyte values, and providing output to a user. Estimation can be used to compensate for time lag, match sensor data with corresponding reference data, warn of upcoming clinical risk, replace erroneous sensor data signals, and provide more timely analyte information encourage proactive behavior and preempt clinical risk. | 1. A method for processing sensor data from a continuous analyte sensor, comprising: receiving sensor data from the continuous analyte sensor;processing, using a processor module, the sensor data; andproviding output on a display, using an output module, based at least in part on the sensor data, wherein the output comprises the processed sensor data plotted on a graph according to a non-uniform y-axis glucose concentration scale. 2. The method of claim 1, wherein providing output comprises displaying the output in real-time. 3. The method of claim 1, wherein processing the sensor data comprises prospectively calibrating the sensor data, and wherein the sensor data displayed on the non-uniform axis comprises a prospectively calibrated sensor data point. 4. The method of claim 1, wherein the non-uniform y-axis comprises a plurality of different zones simultaneously displayed, and wherein the scale resolutions scale resolution along the y-axis of a first zone of the plurality of zones is different than the scale resolution alone the y-axis of a second zone of the plurality of zones, and wherein the first zone is associated with a hypoglycemia condition of a host and the second zone is associated with a hyperglycemia condition of the host. 5. The method of claim 4, wherein the plurality of different zones further comprises a third zone having different scale resolution along the y-axis that is different than both the first zone and the second zone, and wherein the third zone is associated with a target glucose concentration range. 6. The method of claim 1, wherein the continuous analyte sensor is a continuous glucose sensor. 7. The method of claim 1, wherein the scale of the y-axis is incremented in a plurality of different resolutions. 8. The method of claim 1, wherein the y-axis constantly increases in value. 9. The method of claim 1, wherein the y-axis is incremented in a non-uniform manner according to glycemic clinical to a host. 10. A system for processing sensor data from a continuous analyte sensor, comprising: an input module operatively connected to the continuous analyte sensor so as to receive sensor data from the continuous analyte sensor;a processor module that processes the sensor data; andan output module that provides output based on the sensor data, wherein the output comprises the processed sensor data plotted on a graph according to a non-uniform y-axis glucose concentration scale. 11. The system of claim 10, wherein the sensor data is real-time sensor data. 12. The system of claim 10, wherein the processor module prospectively calibrates the sensor data, and wherein the sensor data plotted on the graph is prospectively calibrated sensor data. 13. The system of claim 10, wherein the non-uniform y-axis comprises a plurality of different zones simultaneously displayed, and wherein the y-axis scale resolution of a first zone of the plurality of zones is different than the y-axis scale resolution of a second zone of the p... | DexCom, Inc. | A61B005/00 A61B005/05
US09198612 | B2 | 2015-12-01 | Determination of neuropsychiatric therapy mechanisms of action | 2008-07-08 | A computer implemented method, apparatus, and computer program product of determining mechanisms of action for therapies. A first set of brain scans for each subject in a plurality of subjects generated at a first time period and a second set of brain scans for each subject generated at a second time period are received. Each subject is diagnosed with a given condition and received a given therapy. A set of changes in the set of brain scans is identified for the each subject based on a comparison of a first set of regions of interest in the first set of scans for each subject with a second set of regions of interest in the second set of scans for each subject. A set of typical changes attributable to the given therapy is identified. A mechanism of action for the given therapy is generated based on the set of typical changes. | 1. A computer implemented method of assessing neuroimaging and medical data to determine mechanisms of action for neuropsychiatric therapies, the computer implemented method comprising: receiving, at a processor from a first non-transitory computer readable storage medium via a network connection, neuroimaging data of a first set of human brain scans for each human subject in a plurality of human subjects diagnosed with a given neuropsychiatric condition generated at a first time period prior to beginning implementation of a neuropsychiatric therapy to treat the given neuropsychiatric condition and a second set of human brain scans for the each human subject in the plurality of human subjects diagnosed with the given neuropsychiatric condition generated at a second time period of a given amount of time after beginning the implementation of the neuropsychiatric therapy to treat the given neuropsychiatric condition, wherein a set of one or more scanning devices generate the neuroimaging data of the first set of human brain scans and the second set of human brain scans, and wherein the neuropsychiatric therapy is received by the each human subject in the plurality of human subjects diagnosed with the given neuropsychiatric condition;automatically searching, using the processor, via the network connection a set of online electronic medical literature sources stored in a second non-transitory computer readable storage medium for portions of medical literature describing the neuropsychiatric therapy used to treat the given neuropsychiatric condition, wherein the set of online electronic medical literature sources comprises at least one of published journal articles, medical study results, medical texts, or pharmaceutical studies;automatically identifying, using the processor, a first set of regions of interest in the first set of human brain scans for the each human subject in the plurality of human subjects diagnosed with the given neuropsychiatric condition generated at the first time period prior to beginning the implementation of the neuropsychiatric therapy to treat the given neuropsychiatric condition and a second set of regions of interest in the second set of human brain scans for the each human subject in the plurality of human subjects diagnosed with the given neuropsychiatric condition generated at the second time period of the given amount of time after beginning the implementation of the neuropsychiatric therapy to treat the given neuropsychiatric condition;identifying, using the processor, a set of changes in the neuroimaging data of the first set of human brain scans and the second set of human brain scans for the each human subject in the plurality of human subjects diagnosed with the given neuropsychiatric condition based on a comparison of the first set of regions of interest for the each human subject with the second set of regions of interest for the each human subject, wherein the set of changes comprises indicators of change occur... | International Business Machines Corporation | A61B005/05 A61B005/00 A61B006/03 A61B006/00 A61B005/055 G06F019/00
US08926528 | B2 | 2015-01-06 | Single-axis sensors on flexible backbone | 2008-08-06 | An apparatus includes a narrow elongate probe is adapted for insertion into the body of a living subject. The probe may be flexible and has a plurality of sensors consisting of single coils of very fine wire wound about a backbone of the probe, which transmit signals proximally via fine connecting wires to a position processor. The position processor analyzes the signals to determine position coordinates at multiple points along the length of the probe. | 1. An invasive medical probe apparatus, comprising: at least one field generator for generating an externally applied magnetic field and for establishing a frame of reference;an elongate flexible probe, having a distal end for insertion into a body of a subject and within the frame of reference;a flexible backbone formed of a solid longitudinally non-perforated material at the distal end of the probe, wherein the flexible backbone is made of a ferromagnetic material, and wherein the flexible backbone lacks a longitudinal lumen;a plurality of sensors comprising single-axis coils spirally disposed around said flexible backbone and carried by said flexible backbone, each single-axis coil being formed of wire, said single-axis coils being fixed at different, respective points about said flexible backbone, said single-axis coils being located on said flexible backbone in a known relation to a reference location on said probe, and when subjected to the externally applied magnetic field, said single-axis coils generate respective signals to determine, six-dimensional translational and orientation coordinates of the single-axis coils, wherein the ferromagnetic material of the flexible backbone increases the gain of the single-axis coils;wherein the elongate flexible probe comprises electrodes positioned on the probe at known fixed locations with respect to at least one of said single-axis coils, the location of said electrodes being derived from the six-dimensional translational and orientation coordinates of the single-axis coils; andsignal processing circuitry, which receives and processes said signals to determine respective locations of said points along a portion of said probe and determines six-dimensional translational and orientational coordinates of the single-axis coils in relation to the frame of reference. 2. The apparatus according to claim 1, wherein said probe has a proximal segment, and wherein said single-axis coils are supported by flexible connecting wires that attach said single-axis coils to said proximal segment. 3. The apparatus according to claim 1, wherein said probe has a distal segment that divides into a plurality of flexible branches, said single-axis coils being distributed on said branches. 4. The apparatus according to claim 1, wherein said single-axis coils are formed of wire that having a diameter in a range of 8 to 70 microns. 5. The apparatus according to claim 1, wherein said single-axis coils are formed of wire that does not exceed 15 microns in diameter. 6. The apparatus according to claim 1, wherein said single-axis coils are formed of wire that does not exceed 10 microns in diameter. 7. The apparatus according to claim 1, wherein a diameter of said probe does not exceed 8 French. 8. The apparatus according to claim 1, wherein said single-axis coils are connected to said signal processing circuitry by respective wires running along said backbone. 9. The apparatus according to claim 1, wherein said backbone ... | Biosense Webster, Inc. | A61B005/00 A61M025/00 A61B005/103 A61B005/117 A61B019/00
US08437828 | B2 | 2013-05-07 | Blood inspection device | 2008-01-16 | A blood inspection device capable of separately discharging a sensor and a filter by using a single discharger means. The blood inspection device has a housing provided with a circular hollow cylindrical body having an opening. A filter including a filter is provided inside the cylindrical body, and a sensor is provided outside the cylindrical body. A body of the discharger is slidable outside the cylindrical body. A first discharge section of the discharge comes into contact with the sensor unit to push out and discharge it. A second discharge section of the discharger comes into contact with the filter to push out and discharge it. | 1. A blood test apparatus comprising: a housing that comprises an open cylindrical body having a cylindrical shape;a sensor unit that includes a blood sensor which analyzes blood and that is detachably attached to the open cylindrical body;a filter unit that includes a filter which protects an interior of the housing and that is detachably attached to the open cylindrical body; andan ejector that is slidably attached to the open cylindrical body, that moves the sensor unit toward an open side of the open cylindrical body to first push out and eject the sensor unit from the open cylindrical body of the housing and that then moves the filter unit toward the open side of the open cylindrical body to push out and eject the filter unit from the open cylindrical body of the housing,wherein the sensor unit comprises a first holder that holds the blood sensor, and the filter unit comprises a second holder that holds the filter, andwherein the ejector comprises more than one ejector including a first ejector for ejecting the sensor unit and a second ejector for eject the filter unit,the first ejector is configured to abut the first holder to push out the sensor unit, andthe second ejector is configured to abut the second holder to push out the filter unit. 2. The blood test apparatus according to claim 1, wherein the first ejector and the second ejector are integrally formed. 3. The blood test apparatus according to claim 2, wherein the sensor unit can be replaced while the filter unit is kept attached to the open cylindrical body. 4. The blood test apparatus according to claim 1, wherein: the housing accommodates a laser emitting apparatus within the housing, the laser emitting apparatus radiates laser light on skin through an interior of the open cylindrical body and punctures the skin;the filter unit is arranged between the laser emitting apparatus and the sensor unit; andthe filter allows the laser light to pass therethrough. 5. The blood test apparatus according to claim 1, wherein: the housing accommodates an optical reader within the housing, the optical reader radiates light on skin through an interior of the open cylindrical body and detects an amount of received reflected light;the filter unit is arranged between the optical reader and the sensor unit; andthe filter allows the light to pass therethrough. 6. The blood test apparatus according to claim 1, wherein: the housing accommodates therein:a laser emitting apparatus that radiates laser light on skin through an interior of the open cylindrical body and punctures the skin; andan optical reader that radiates light on the skin through the interior of the open cylindrical body and detects an amount of received reflected light;the filter unit is arranged between the laser emitting apparatus and the sensor unit, and between the optical reader and the sensor unit; andthe filter allows the laser light and the light to pass therethrough. 7. The blood test apparatus according to claim 1, wherein... | Panasonic Corporation | A61B005/05 A61B005/00 A61B017/32
US08771714 | B2 | 2014-07-08 | Systems and methods for in vivo measurement of interstitial biological activity, processes and/or compositions | 2007-10-31 | Provided are systems and methods for detecting and/or measuring in vivo interstitial biological activity, processes, and or compounds in human or animal subjects. | 1. A method of real-time in vivo monitoring of a biological process, comprising: identifying a compound of interest that is located in the non-cardiac interstitium of the subject, wherein the activity or presence of the compound is related to the in vivo biological process being monitored;positioning a probe within the non-cardiac interstitium of the subject;infusing the interstitium of the subject with a fluorogenic substrate specific for the compound of interest wherein the substrate is infused therethrough the probe;receiving interstitial fluid from the subject with the probe, wherein the interstitial fluid comprises the substrate; anddetecting the received substrate to monitor the biological process in the subject. 2. The method of claim 1, wherein the biological process monitored is death or injury of a cell. 3. The method of claim 1, wherein the non-cardiac interstitium is located in the joint of a subject. 4. The method of claim 1, wherein the biological process is cancer, the method further comprising, prior to positioning the probe within the interstitium of the subject, administering cancer therapy to the subject having cancer wherein the administered therapy causes death or injury of a tumor cell and the subsequent release of identified compound from the dead or injured cell. 5. The method of claim 1 further comprising, prior to positioning the probe within the interstitium of the subject, administering a pharmaceutical agent to the subject wherein the pharmaceutical agent modifies the biological process therein the subject, wherein the fluorogenic substrate is configured to be altered to a detectable state by the biological process, wherein the detectable state of the fluorogenic substrate indicates the modification of the biological process by the administered agent, wherein the modification is used to monitor the therapeutic regimen in subject. 6. The method of claim 5, wherein an increase in detected fluorogenic substrate indicates the modification of the biological process by the administered agent. 7. The method of claim 5, wherein a decrease in detected fluorogenic substrate indicated the modification of the biological process by the administered agent. 8. The method of claim 7, wherein the pharmaceutical agent acts to inhibit the biological process. 9. The method of claim 8, wherein the pharmaceutical agent is Aprotinin. 10. The method of claim 1 further comprising, prior to positioning the probe within the interstitium of the subject, administering a pharmacological agent to the subject to effect a modification of the biological process therein the subject, wherein at least a portion of the fluorogenic substrate is altered in the interstitium of the subject, wherein the received substrate is the altered fluorogenic substrate, anddetermining the modification of the biological process which is indicative of the pharmacological activity of the administered pharmacological agent via detection of the received substrate. 1... | MUSC Foundation for Research Development | A61K009/00 A61B005/0275 C12Q001/34 C12Q001/37 C12Q001/06 G01N033/574 G01N033/573 G01N033/00 A61K051/00
US08938292 | B2 | 2015-01-20 | Estimating cardiovascular pressure and volume using impedance measurements | 2008-10-31 | Techniques for estimating a cardiac chamber or vascular pressure based upon impedance are described. A device or system may measure an impedance between at least two electrodes implanted within or proximate to a cardiovascular system. The device or system may estimate a pressure of an element of the cardiovascular system based on a relationship between impedance and volume of the element, and based on a empirical relationship between the volume and the pressure. The device or system may also estimate the dimension of the element based on the impedance-volume relationship, or other characteristics based on the impedance. Because the impedance measurements may be obtained, in some examples, by using electrodes and leads implanted within the cardiovascular system and coupled to an implantable medical device, a practical estimation of a cardiovascular pressure can be obtained on a chronic basis without requiring the use other invasive sensors, such as micronanometer transducers. | 1. A method comprising: storing in a memory of an implantable medical device a plurality of relationships between volume and pressure, the plurality of relatioships comprising a relationship between volume and pressure in avascular lumen;storing in the memory a plurality of relationships between impedence and volume;measuring an impedance between at least two electrodes implanted within or proximate to a cardiovascular system of a patient;enabling a control unit to select one of the stored plurality of relationships between volume and pressure and one of the stored plurality of relationships between impedence and volume based on at least one selection criterion; andestimating by the control unit, a pressure of an element of the cardiovascular system based on the selected relationship between impedence and volume and the selected relationship between the volume and the pressure, wherein the relationship between the volume and the pressure comprises an elastance function derived from empirical data comprising historical data of the patient. 2. The method of claim 1, wherein the element of the cardiovascular system comprises one of a cardiac chamber or a vascular lumen. 3. The method of claim 1, wherein relationship between the measured impedance and the volume comprises an inversely proportional relationship. 4. The method of claim 1, wherein the relationship between the volume and the pressure comprises a linear approximation of the elastance function for a given volume interval. 5. The method of claim 1, further comprising: selecting the relationship between the volume and the pressure from the plurality of relationships based upon a particular characteristic of the patient. 6. The method of claim 1, wherein measuring the impedance comprises measuring a change in impedance,wherein estimating the pressure comprises estimating a change in pressure based on a relationship between the measured change in impedance and a change in volume of the element of the cardiovascular system and based on a relationship between the change in volume and the change in pressure. 7. The method of claim 1, further comprising: selecting a combination of electrodes as the at least two electrodes based upon the element of the cardiovascular system to be measured. 8. The method of claim 1, wherein measuring the impedance comprises: generating a plurality of electric fields between the at least two electrodes;applying the plurality of electric fields to the element of the cardiovascular system such that the combined electric field is substantially homogeneous over the element; andmeasuring an impedance of a resultant current flow between the at least two electrodes. 9. The method of claim 1, wherein measuring the impedance comprises: measuring a first impedance for a first region;measuring a second impedance for a second region; andsubtracting the second impedance from the first impedance to determine an impedance for the element. 10. The method of claim 1, wherein... | Medtronic, Inc. | A61B005/05 A61N001/365 A61B005/0215 A61B005/053 A61B005/021 A61B005/00
US08073519 | B2 | 2011-12-06 | Systems and methods for replacing signal artifacts in a glucose sensor data stream | 2009-10-14 | Systems and methods for minimizing or eliminating transient non-glucose related signal noise due to non-glucose rate limiting phenomenon such as ischemia, pH changes, temperatures changes, and the like. The system monitors a data stream from a glucose sensor and detects signal artifacts that have higher amplitude than electronic or diffusion-related system noise. The system replaces some or the entire data stream continually or intermittently including signal estimation methods that particularly address transient signal artifacts. The system is also capable of detecting the severity of the signal artifacts and selectively applying one or more signal estimation algorithm factors responsive to the severity of the signal artifacts, which includes selectively applying distinct sets of parameters to a signal estimation algorithm or selectively applying distinct signal estimation algorithms. | 1. In a glucose sensing system comprising a glucose sensor configured for producing an output representative of glucose concentration in a host and processing circuitry configured to process the output, a method for processing data from the output, the method comprising: generating at least one data stream from the glucose sensor output;evaluating the data stream to determine a presence of a signal artifact;determining a severity of the signal artifact over time;continuously generating an estimated data stream;initiating replacement of a portion of the data stream with a time-corresponding portion of the estimated data stream in response to determining that the severity of the signal artifact exceeds a first predetermined severity threshold; andterminating replacement of the data stream with the time-corresponding portion of the estimated data stream in response to determining that the severity of the signal artifact falls below a second predetermined severity threshold. 2. The method of claim 1, wherein the severity of the signal artifact is determined at least in part by determining the duration of the signal artifact. 3. The method of claim 1, wherein the severity of the signal artifact is determined at least in part by determining an amplitude of the signal artifact. 4. The method of claim 1, wherein the first and second thresholds are the same. 5. The method of claim 1, wherein the first threshold is in terms of a duration of time, an amplitude, or a frequency. 6. The method of claim 1, wherein the second threshold is in terms of a duration of time, an amplitude, or a frequency. 7. The method of claim 1, wherein the replaced portion comprises more than one contiguous data point, each data point representative of a discrete estimated glucose concentration value. 8. The method of claim 1, wherein replacing comprises replacing a plurality of contiguous data points of the signal artifact with estimated data points of the estimated data stream. 9. A glucose sensor system for monitoring a glucose concentration in a host, the system comprising: a continuous glucose sensor configured to produce a data stream;a signal artifact detection module configured to detect a presence of a signal artifact in the data stream and determine severity of the signal artifact over time;a signal estimation module configured to perform signal estimation to continuously generate an estimated data stream; anda signal artifact replacement module configured to initiate replacement of a portion of the data stream with a time-corresponding portion of the estimated data stream in response to determining that the severity of the signal artifact exceeds the predetermined severity threshold, and terminate replacement of the data stream with the time-corresponding portion of the estimated data stream in response to determining that the severity of the signal artifact falls below a second predetermined severity threshold. 10. The system of claim 9, wherein the second thre... | DexCom, Inc. | A61B005/05 A61B005/00
US08175684 | B2 | 2012-05-08 | Method for processing images and associated medical imaging system | 2008-07-29 | A system and method of processing images of a region of interest of a patient is provided. The method comprises acquiring a reference image of the region of interest of the patient; during a pullback of an intravascular sensor in the region of interest, triggering simultaneously the steps of: acquiring a data collected by the sensor characteristic of the region of interest; and acquiring a succession of images of the region of interest associated with the location of the intravascular sensor when acquiring the data, respectively. The method further includes registering the succession of images; associating the location of the intravascular sensor relative to the step of acquiring the data collected by the intravascular sensor; and displaying and positioning the data collected by the intravascular sensor on the reference image in correspondence to the location of the intravascular sensor at the respective step of acquiring the data. | 1. A method of processing images of a region of interest of a patient, comprising the steps of: acquiring a reference image of the region of interest of the patient;during a pullback of an intravascular sensor introduced inside the region of interest of the patient, triggering simultaneously the steps of: acquiring a set of non-image data using the intravascular sensor, wherein the non-image data is characteristic of the region of interest, andacquiring a succession of images of the region of interest, each of the succession of images associated with the location of the intravascular sensor when acquiring the non-image data, respectively;registering the succession of images for which the non-image data collected by the intravascular sensor are available;associating the location of the intravascular sensor relative to the step of acquiring the non-image data using the intravascular sensor; anddisplaying values corresponding to the non-image data collected by the intravascular sensor at positions on the reference image corresponding to the location at which the respective non-image data was collected. 2. The method according to claim 1, wherein the step of acquiring the non-image data using the intravascular sensor and the step of acquiring the succession of images are performed at generally the same phase of an electrocardiogram of the patient. 3. The method of claim 1, wherein the step of acquiring the non-image data using the intravascular sensor is performed at a first frequency and the step of acquiring the succession of images is performed at a second frequency, said first frequency being generally greater than the second frequency. 4. The method according to claim 3, wherein the location of the intravascular sensor is time interpolated so that the frequency of the interpolated locations is generally equal to the first frequency. 5. The method according to claim 1, wherein the pullback of the intravascular sensor is implemented by a pullback actuator. 6. The method according to claim 1, wherein the pullback of the intravascular sensor is implemented at a constant speed. 7. The method of claim 1, comprising changing a cadence of data acquisition based on the speed of the pullback of the intravascular sensor. 8. The method of claim 1, wherein the non-image data comprises temperature data. 9. The method of claim 1, further comprising displaying a respective position of the intravascular sensor on the reference image with each of the values corresponding to the non-image data at the positions on the reference image corresponding to the location at which the respective non-image data was collected. 10. A medical imaging system operable to process images of region of interest of a patient, comprising: an image acquisition system operable to acquire a reference image of the region of interest of a patient;an intravascular sensor operable during pullback to sense non-image data characteristic of the region of interest, said sensor configured... | General Electric Company | A61B005/05
US08285365 | B2 | 2012-10-09 | Interventional device for RF ablation for use in RF fields | 2007-08-20 | A interventional device for RF ablation for use in a RF electrical and/or magnetic field especially of a MR imaging system is disclosed, comprising an ablation catheter which is preferably trackable or can be guided or visualized in the image generated by the MR imaging system by means of a MR micro-coil (102), and which is provided with an ablation electrode (101). The interventional device further comprises a transmission path (103) with line segments (104′, 104″) and transformers (105) therebetween, for connecting the MR micro-coil (102) in a differential mode with a MR receiver (108) and for connecting a RF amplifier (107) in a common mode with the ablation electrode (101) for conveying RF ablation power. | 1. An interventional device for RF ablation for use in a RF electrical and/or magnetic field, comprising: an ablation catheter or probe with an ablation electrode for delivering RF ablation power, and an MR/RF micro-coil for receiving MR signals and/or for transmitting RF signals, anda transmission path connected with the ablation catheter or probe, wherein the transmission path comprises at least two line segments which are electrically coupled by at least one coupling element comprising a high pass filter characteristic so that: common mode currents induced in the transmission path by the RF electrical and/or magnetic field of a first frequency to which the transmission path is exposed, are at least substantially blocked,the RF ablation power having a predetermined second RF frequency, which is higher than the first frequency, is at least substantially, or to an extent which is necessary for a desired ablating of tissue, transmitted in the form of common mode currents over the transmission path, andMR signals and/or RF signals which are received and/or transmitted by the MR/RF micro-coil are conveyed over the transmission path in the form of differential mode currents. 2. The interventional device according to claim 1, wherein the coupling element is a transformer and the high pass filter characteristic is realized by a stray capacity of the transformer, the stray capacity representing a first impedance which is matched to a second impedance of the at least two line segments by at least one matching network located on the transmission path. 3. The interventional device according to claim 2, wherein the transformer is realized in PCB technology on a substrate in the form of a first conductive loop and a second conductive loop which overlap each other for inductive coupling. 4. An interventional device for RF ablation for use in a RF electrical and/or magnetic fields, the device comprising: an ablation catheter or probe;an ablation electrode mounted to the catheter or probe to deliver ablation power to tissue to be ablated;an MR/RF micro-coil for receiving MR signals and/or for transmitting RF signals, for active and/or passive tracking, guiding and/or visualization of the catheter or probe in an MR image generated by an MR imaging system;a transmission path extending along the ablation catheter or probe, the transmission path including at least two line segments which are electricaly coupled by a high pass filter which blocks common mode currents induced in the transmission path by the RF electrical and/or magnetic field at a first frequency;a source of RF ablation power having a predetermined second RF frequency which is higher than the first frequency, the source of RF ablation power being connected with the transmission path to transmit the RF ablation power as common mode currents of the second frequency over the transmission path to the ablation electrode;wherein the MR signals and/or RF signals which are received and/or transmitted by t... | Koninklijke Philips Electronics N.V. | A61B005/05
US08265752 | B2 | 2012-09-11 | System and method for assessing atrial electrical stability | 2009-04-28 | A method for identifying a susceptibility of a subject to atrial-rhythm disturbances includes a) placing a plurality of sensors on the subject to measure a physiologic signal of the subject, and b) recording the physiologic signal from the sensor. The physiologic signal includes an atrial electrical activity of the subject. The method includes c) determining a beat-to-beat variability in the atrial electrical activity of the subject. The beat-to-beat variability includes alternans of electrocardiographic waveforms of a predetermined number of a sequence of heart beats. The method includes d) determining a susceptibility to atrial-rhythm disturbances of the subject using the beat-to-beat variability in the atrial electrical activity determined in step c), and e) generating a report of the susceptibility to atrial-rhythm disturbances of the subject. | 1. A method for identifying a susceptibility of a subject to atrial-rhythm disturbances, comprising the steps of: a) placing a plurality of sensors on the subject to measure a physiologic signal of the subject; b) recording the physiologic signal from the sensor, the physiologic signal including an atrial electrical activity of the subject; c) determining a beat-to-beat variability in the atrial electrical activity of the subject, the beat-to-beat variability including alternans of electrocardiographic waveforms of a predetermined number of a sequence of heart beats; d) determining a susceptibility to atrial-rhythm disturbances of the subject using the beat-to-beat variability in the atrial electrical activity determined in step c); and e) generating a report of the susceptibility to atrial-rhythm disturbances of the subject. 2. The method of claim 1, further comprising f) using the report of the susceptibility to determine a therapy plan including at least one of a delivery of a chemical substance and a delivery of electrical energy to the subject. 3. The method of claim 2, wherein the delivery of electrical energy includes delivering electrical energy during an absolute refractory period on alternate beats within the sequence of heart beats. 4. The method of claim 1, wherein step c) includes performing a Fast Fourier Transform to estimate alternans of the electrocardiographic waveforms. 5. The method of claim 1, wherein step c) includes identifying beat-to-beat variability created by at least one of atrial depolarization and repolarization alternans. 6. The method of claim 1, wherein step a) includes arranging the plurality of sensors in a bipolar configuration. 7. The method of claim 1, wherein step a) includes arranging the plurality of sensors in a unipolar configuration. 8. The method of claim 1, wherein step e) includes displaying the susceptibility to atrial-rhythm disturbances of the subject showing a first output showing raw electrocardiogram (ECG) signals of the subject and a second output showing the susceptibility to atrial-rhythm disturbances of the subject. 9. A method for identifying a susceptibility of a subject to atrial-rhythm disturbances, comprising the steps of: a) placing a sensor on the subject to measure a physiologic signal of the subject; b) recording the physiologic signal from the sensor, the physiologic signal including an electrical activity of a heart of the subject; c) determining a beat-to-beat variability in the atrial electrical activity of the subject during a sequence of heart beats; d) determining a susceptibility to atrial-rhythm disturbances of the subject using the beat-to-beat variability in the atrial electrical activity; and e) generating a report of the susceptibility to atrial-rhythm disturbances of the subject. 10. The method of claim 9, further comprising f) determining a therapy location for a delivery of one of a chemical substance and an electrical energy to the subject to compensate fo... | The General Hospital Corporation | A61N001/365
US08290575 | B2 | 2012-10-16 | Biomedical signal morphology analysis method | 2007-11-30 | A way of quantifying the shape of an ECG waveform is disclosed by detecting the JT segment using two Hidden Markov Models and calculating the analytic signal of the JT segment. Parameters calculated from the analytic signal are used as shape descriptors for the JT segment. The shape descriptors may be displayed in a dimensionality-reduced mapping. Templates representing characteristic shapes can be produced by finding cluster centers in the shape descriptor space, and the novelty of new waveforms can be quantified by comparing the position in shape descriptor space of new shape descriptors to a predefined normal training set or to previously encountered waveforms. Novel shape descriptors can be used to retrieve the corresponding waveforms, and templates of such novel shapes can be created by averaging such waveforms, using dynamic time warping to allow for variations in heart rate. The templates can be manually segmented and the manual segmentation propagated back into other waveforms having similar shape descriptors. | 1. A method of analysing a biomedical signal having a repetitive signal feature, comprising the steps of segmenting each of a plurality of said repetitive signal features in the biomedical signal, analysing one or more of the segments to find the values of a plurality of parameters describing the shape of said one or more of the segments, said plurality of parameters constituting components of a shape descriptor vector for that segment, recording the values, and tracking changes in the said values through the biomedical signal; the method further comprising the step of defining normality for said shape descriptor vector, and calculating for each of said plurality of repetitive signal features the difference between its shape descriptor vector and normality. 2. A method according to claim 1 wherein the biomedical signal is an electrocardiogram and the repetitive signal feature is the heartbeat. 3. A method according to claim 2 wherein the step of analysing one or more of the segments comprises analysing the T-wave of the heartbeat, being the segment from the J point to the Toffset. 4. A method according to claim 3 wherein the J point is detected by use of a Hidden Markov Model comprising a state model for the waveform after the J point. 5. A method according to claim 3 wherein the Toffset is detected by use of a Hidden Markov Model comprising state models for the T wave and the waveform after Toffset. 6. A method according to claim 5 wherein Toffset is detected by the steps of: maintaining from beat to beat through the electrocardiogram an estimate of the QT interval based on a relationship between the QT interval and the heart rate, estimating the position of Toffset from the estimated QT interval, and applying the Hidden Markov Model comprising state models for the T wave and the waveform after Toffset to a portion of the electrocardiogram defined with respect to the estimated Toffset position. 7. A method according to claim 6 wherein a Kalman filter is used to maintain said estimate of the QT interval. 8. A method according to claim 1 wherein normality is defined by shape descriptor vectors based on a plurality of repetitive signal features from training set. 9. A method according to claim 8 wherein the step of calculating for each of said plurality of repetitive signal features the difference between its shape descriptor vector and normality comprises calculating the distance in shape descriptor vector space between its shape descriptor vector and the nearest shape descriptor vector based on the training set. 10. A method according to claim 8 wherein the step of calculating for each of said plurality of repetitive signal features the difference between its shape descriptor vector and normality comprises calculating the density in the shape descriptor vector space of points derived from the training set, and calculating the probability of normality of the shape descriptor vector for each of said plurality of repetitive signal features.... | OBS Medical Limited | A61B005/04
US08361000 | B2 | 2013-01-29 | Automated infant massager | 2009-07-17 | An automated infant massager system that monitors an infant's physical, physiological and behavioral response to the massage and can adjust the intensity of the massage and massager duration settings based on the responses of the infant and/or information from a user, preferably autonomously. A processor, which is operatively associated with the infant massager, coordinates overall operations and functions of the infant massager, as well as receiving information regarding the infant's physical, behavioral, and physiological responses to the massage. The user input through the remote receiver transmitter, and/or an associated bar code/radio-frequency identification (RFID) scanner, and/or direct communications with the processor may be employed. In a further embodiment of the present invention, a plurality of infant massaging units are controlled in network fashion by means of a processor and router, thereby permitting one or more infant massagers to be controlled by one or more users and allows one or more automated infant massagers to communicate and interface with electronic medical record (EMR) systems. | 1. An automated infant massager system comprising: at least one automated infant massager configured to massage an infant,an infant support,a plurality of movable massage elements structured to underlie at least a portion of the infant being massaged,a processor for receiving, and/or measuring, infant information directly or indirectly from said infant massager,said infant information including information relating to at least one of the physical, behavioral, and physiological responses of the massaged infant, andsaid processor configured to alter the massage, while the massager is performing a massage, based upon said infant information if it departs from predetermined limits. 2. The automated infant massager system of claim 1 including said processor configured to determine from said infant information at least one of the distressed stated and/or pain level of said massaged infant based upon a comparison with data stored within said processor. 3. The automated infant massager system of claim 2 including said processor configured to make adjustments to the operation of said massager responsive to said infant information with such adjustment being selected from the group of intensity and duration of said massage. 4. The automated infant massager system of claim 3 including said program configured to initiate an alarm if said infant information exceeds predetermined limits. 5. The automated infant massager system of claim 1 including a remote receiver transmitter operatively associated with said infant massager and said processor, andsaid remote receiver transmitter being configured to receive infant information from said infant massager and to responsively deliver control signals to said processor or said infant massager. 6. The automated infant massager system of claim 5 including a user input configured to deliver signals to at least one system component selected from the group consisting of said processor and said remote receiver transmitter. 7. The automated infant massager system of claim 1 including sensors monitoring said infant massager and said infant to provide infant information to said processor regarding at least one condition selected from the group consisting of monitored infant conditions. 8. The automated infant massager system of claim 7 including said processor responsive to receipt of said infant information issuing control signals to said infant massager through a controller. 9. The automated infant massager system of claim 5 including said remote receiver transmitter configured to receive visual and audible infant information regarding said massaged infant. 10. The automated infant massager system of claim 1 including said infant massager being configured to discretely or continuously monitor physical parameters of said infant. 11. The automated infant massager system of claim 1 including said infant massager being configured to discretely or continuously monitor behavioral parameters of said infant. 12. The autom... | - | A61B005/00
US08235911 | B2 | 2012-08-07 | Methods and systems for filtering a signal according to a signal model and continuous wavelet transform techniques | 2008-10-03 | According to embodiments, systems and methods are provided for filtering a signal. A first reference signal may be generated according to a signal model and a second reference signal may be generated by analyzing a continuous wavelet transform of a signal. The first and second reference signals may then both be applied to an input signal to filter the input signal according to the components of both of the reference signals. | 1. A method for filtering an input signal, the method comprising: receiving, by a processor, the input signal;generating, by the processor, a first reference signal based at least in part on a signal model associated with a component of the input signal;generating, by the processor, a second reference signal based at least in part on a continuous wavelet transform of the received input signal; andfiltering, by the processor, the received input signal using at least the first reference signal and the second reference signal. 2. The method of claim 1, wherein generating a first reference signal based at least in part on a signal model comprises: measuring at least two signals of the same or similar type as the received input signal;assuming some mathematical relationship exists between the two measured signals; andminimizing the correlation between a signal portion and a noise portion of at least one of the at least two measured signals. 3. The method of claim 1, wherein generating a second reference signal based at least in part on a continuous wavelet transform of the received input signal comprises generating a scalogram of the received input signal. 4. The method of claim 1, wherein filtering the received input signal comprises filtering the received input signal in wavelet space. 5. The method of claim 1, wherein: at least one of the first and second reference signals comprises a noise component; andfiltering the received input signal comprises applying an adaptive correlation canceller to the received input signal. 6. The method of claim 1, wherein: at least one of the first and second reference signals comprises a signal component; andfiltering the received input signal comprises cross correlating at least one of the first and second reference signals with the received input signal. 7. The method of claim 1 further comprising combining at least some of the first reference signal with at least some of the second reference signal before filtering the received input signal. 8. The method of claim 1 further comprising determining at least one physiological parameter from the filtered input signal. 9. The method of claim 1, wherein filtering the received input signal comprises: filtering the received input signal using the first reference signal to produce a first filtered signal; andfiltering the first filtered signal using the second reference signal to produce a twice filtered signal. 10. The method of claim 1, wherein the input signal comprises at least one photoplethysmograph (PPG) signal. 11. A system for filtering, comprising: a sensor capable of generating an input signal; anda processor capable of: generating a first reference signal based at least in part on a signal model associated with a component of the input signal;generating a second reference signal based at least in part on a continuous wavelet transform of the input signal; andfiltering the input signal using at least the first reference signal and the second referenc... | Nellcor Puritan Bennett Ireland | A61B005/026 G06F019/00 G06F015/00 G11C017/00
US08903472 | B2 | 2014-12-02 | Applications of scaling laws of tree structures | 2008-01-22 | In at least one embodiment, a method for diagnosing vascular disease is provided, the method comprising the steps of obtaining a vessel image showing a vasculature of a vessel, identifying at least two measurements from the vasculature of the vessel, the measurements relating to at least two parameters, calculating a relationship between the at least two parameters from the at least two measurements to generate one or more vasculature data points, and comparing the one or more vasculature data points to data relative to a model vasculature to determine the extent of vascular disease. In an another embodiment, a method for diagnosing vascular disease in a patient's vascular tree is provided, the method comprising the steps of generating a model vascular tree from a minimum energy hypothesis calculation, and comparing the patient's vascular tree with the model vascular tree to determine the extent of vascular disease. | 1. A method for diagnosing vascular disease, the method comprising the steps of: obtaining a vessel image showing a vasculature of a vessel having a vascular disease, the vasculature of the vessel including at least part of a biological tree having a stem and a crown;identifying at least two measurements from the vessel image, the measurements relating to at least two parameters selected from the group consisting of a vessel volume, a vessel length, a vessel diameter, a flow rate of a stem, and crown resistance;calculating at least two cumulative measurements corresponding to the selected at least two parameters, the at least two cumulative measurements selected from the group consisting of a volume of the crown corresponding to the vessel volume parameter, a cumulative vessel length of the crown corresponding to the vessel length parameter, a diameter of a most proximal stem corresponding to the vessel diameter parameter, a flow rate of a most proximal stem corresponding to the flow rate of the stem parameter, and a resistance of the biological tree corresponding to the crown resistance parameter;calculating a relationship between the at least two parameters from the at least two measurements and the at least two cumulative measurements to generate one or more vasculature data points using a processor operably connected to a storage medium having the at least two measurements stored thereon; andcomparing the one or more vasculature data points to data relative to a model vasculature to determine the extent of vascular disease. 2. The method of claim 1, wherein the vessel image is an image selected from the group consisting of an angiograph, a computed tomography scan, and a magnetic resonance image. 3. The method of claim 1, wherein the at least two parameters comprise the vessel volume and the vessel length, and wherein the at least two cumulative measurements comprise the volume of the crown and the cumulative vessel length of the crown. 4. The method of claim 3, wherein the step of calculating a relationship between the at least two parameters is performed using a volume-length relation. 5. The method of claim 3, wherein the step of calculating a relationship between the at least two parameters is performed using a resistance-length and volume relation. 6. The method of claim 1, wherein the at least two parameters comprise the vessel diameter and the vessel length, and wherein the at least two cumulative measurements comprise the diameter of the most proximal stem and the cumulative vessel length of the crown. 7. The method of claim 6, wherein the step of calculating a relationship between the at least two parameters is performed using a diameter-length relation. 8. The method of claim 1, wherein the at least two parameters comprise the flow rate and the vessel diameter, and wherein the at least two cumulative measurements comprise the flow rate of the most proximal stem and the diameter of the most proximal stem. 9. The method of cla... | DTherapeutics, LLC | A61B005/05 A61B005/02 A61B006/00 A61B005/026 A61B005/055 A61B006/03
US09066691 | B2 | 2015-06-30 | Using a continuous wavelet transform to generate a reference signal | 2008-10-03 | According to embodiments, systems and methods for generating reference signals are provided. A signal may be transformed using a continuous wavelet transform. Regions of interest may be selected from a transform or the resulting scalogram that may be used to generate a reference signal to use in filtering the signal or other signals. Cross-correlation techniques may be used to cancel noise components or isolate non-noise components from the signal. A physiological parameter may then be determined from the filtered signal or isolated components in the signal. | 1. A method for generating a noise reference signal, the method comprising: receiving a first input signal;receiving a second input signal, wherein the second input signal is different from the first input signal;performing, using processing equipment, a continuous wavelet transform of the first input signal to produce a first transformed input signal;performing, using processing equipment, a continuous wavelet transform of the second input signal to produce a second transformed input signal;generating, using processing equipment, a first scalogram from the first transformed input signal;generating, using processing equipment, a second scalogram from the second transformed input signal, wherein the first scalogram is generated based at least in part on the second scalogram;identifying, using processing equipment, regions in the first scalogram that exceed a predetermined threshold; andperforming, using processing equipment, an inverse continuous wavelet transform on at least part of the first transformed input signal, based at least in part on the identified regions, to generate the noise reference signal. 2. The method of claim 1, wherein identifying regions in the first scalogram comprises identifying, using processing equipment, at least one noise component in the first scalogram. 3. The method of claim 1, wherein identifying regions in the first scalogram comprises identifying, using processing equipment, at least one non-noise component in the first scalogram. 4. The method of claim 1, wherein the first input signal comprises a photoplethysmograph (PPG) signal. 5. The method of claim 1 further comprising outputting, using processing equipment, the noise reference signal to an adaptive filter. 6. The method of claim 5 further comprising filtering, using processing equipment, the first input signal using at least the adaptive filter and the noise reference signal to produce a filtered input signal. 7. The method of claim 6 wherein filtering the first input signal comprises cross correlating, using processing equipment, the noise reference signal with the first input signal. 8. The method of claim 6 further comprising determining, using processing equipment, a physiological parameter based at least in part on the filtered input signal. 9. The method of claim 1 further comprising processing, using processing equipment, the received first input signal together with the noise reference signal. 10. The method of claim 1 further comprising extracting, using processing equipment, the identified regions from the first scalogram, wherein the noise reference signal is generated from the extracted identified regions. 11. The method of claim 1, wherein the predetermined threshold is adjusted dynamically. 12. The method of claim 11, wherein the predetermined threshold is adjusted dynamically based at least in part on the first scalogram. 13. A system for generating a noise reference signal, comprising: an input capable of receiving a first inpu... | Nellcor Puritan Bennett Ireland | A61B005/1455 A61B005/00 G06K009/64
US08600471 | B2 | 2013-12-03 | Catheter with a sensing region for redox reactions | 2007-06-08 | The present disclosure relates to catheters, for example a catheter which is adapted such that the user is able to determine whether it has been correctly positioned within 5 the subject. Also included in the present disclosure are means for detecting the position of the catheter and combinations of a catheter and processors for determining the positioning of the catheter. The present disclosure also relates to methods of determining the position of a catheter in, for example, a human or animal body. | 1. A catheter comprising a first end portion having a sensing region comprising at least one redox reagent which is in electrical connection with an electrode and which is capable of undergoing a redox reaction when in a selected bodily environment, wherein the rate of the redox reaction is dependent on the selected bodily environment, wherein the catheter is a naso-gastric tube and wherein the redox reagent is Vitamin K1 and further wherein the selected bodily environment is the stomach. 2. The catheter according to claim 1, which comprises a pair of electrodes, wherein the redox reagent is in electrical communication with one of said electrodes. 3. The catheter according to claim 1, wherein the redox reagent is in electrical communication with an information processor which is adapted to detect or measure electrical signals generated by the redox reaction. 4. The catheter according to claim 1, wherein the redox reagent is adapted to undergo a redox reaction at a specific rate only when the sensing region is inserted into the-selected location of a subject's body and not when the sensing region is inserted into an alternate location in the subject's body. 5. The catheter according to claim 1, wherein the redox reagent undergoes a redox reaction at a specific pH range found in a specified bodily environment. 6. The catheter according to claim 1, wherein the redox reaction produces a redox product, and further wherein the sensing region further comprises a regenerating agent which regenerates the redox reagent from the redox product. 7. The catheter according to claim 6, wherein the sensing region further comprises at least one layer comprising the regenerating agent, wherein said agent is capable of being degraded in the selected bodily environment. 8. The catheter according to claim 7, wherein the sensing region comprises a plurality of layers comprising the regenerating agent, wherein the layers are in a stepped arrangement, there being a proximal portion of each layer which is uncoated by a successive layer and each layer, other than the outermost, has a distal portion coated with the next outer layer. 9. The catheter according to claim 6, wherein the regenerating agent comprises a biocompatible polymer. 10. The catheter according to claim 9, wherein the biocompatible polymer comprises gelatin. 11. The catheter according to claim 10, which further comprises a second end portion at a distal end of the catheter, which is adapted to be connected to an information processor for detecting or measuring an electrical signal produced by the redox reaction. 12. The catheter according to claim 10, wherein the gelatin is capable of regenerating Vitamin K1 from its reduced form. 13. The catheter according to claim 1, which comprises a non-sensing region which is electrically insulated. 14. A combination of a catheter according to claim 2 and an information processor for detecting or measuring an electrical signal produced by a redox reaction ... | The University of Hull | A61L029/16 A61B005/05
US08767049 | B2 | 2014-07-01 | Membrane-deformation mapping technique | 2009-07-21 | In accordance with one embodiment of the present disclosure, a method for deformation mapping of a tissue is provided. The method includes utilizing a device to measure transient three-dimensional deformations in a tissue sample. The device comprises a non-contacting, high-speed stereo imaging apparatus and a mechanism for digital image correlation. The method further includes identifying regions of the tissue that are prone to damage based upon the deformations. | 1. A method for deformation mapping of a tissue comprising: utilizing a device to measure transient three-dimensional deformations in a tissue sample comprising a heart valve under physiologic flow conditions which replicate valve opening and closure, the device comprising a non-contacting, high-speed stereo imaging apparatus, a pulsatile pump system configured to provide periodic pulsatile flow of fluid through the heart valve, and a mechanism for digital image correlation, wherein the pulsatile pump system is configured to operate in synchronization with the high speed stereo imaging apparatus; andidentifying regions of the tissue that are prone to damage based upon the deformations. 2. The method as in claim 1, wherein the heart valve comprises a bioprosthetic heart valve. 3. The method as in claim 1, wherein the high speed stereo imaging apparatus comprises two cameras. 4. The method as in claim 1, further comprising marking the surface of the tissue with particles. 5. The method as in claim 4, wherein the particles comprise water-resistant dye particles. 6. A device for deformation mapping a tissue comprising: a non-contacting, high-speed stereo imaging apparatus, a pulsatile pump system, and a mechanism for digital image correlation, wherein the high speed stereo imaging apparatus and the mechanism for digital image correlation are configured to measure transient three-dimensional deformations in a tissue sample comprising a heart valve under physiologic flow conditions which replicate valve opening and closure and identify regions of the tissue that are prone to damage based upon the deformations, and wherein the pulsatile pump system is configured to provide periodic pulsatile flow of fluid through the heart valve and operate in synchronization with the high speed stereo imaging apparatus. 7. The device of claim 6, wherein the high speed stereo imaging apparatus comprises one or more mirrors. 8. The device of claim 6, wherein the high speed stereo imaging apparatus comprises two cameras. 9. The device of claim 8, wherein the cameras are high speed cameras. 10. The device of claim 9, wherein the high speed stereo imaging apparatus comprises an optical window, the tissue sample being capable of being observed through an optical window by the high speed cameras. | University of South Carolina | H04N013/02 G01J005/04
US08174394 | B2 | 2012-05-08 | System for noninvasive determination of analytes in tissue | 2009-09-17 | An apparatus and method for noninvasive determination of analyte properties of human tissue by quantitative infrared spectroscopy to clinically relevant levels of precision and accuracy. The system includes subsystems optimized to contend with the complexities of the tissue spectrum, high signal-to-noise ratio and photometric accuracy requirements, tissue sampling errors, calibration maintenance problems, and calibration transfer problems. The subsystems can include an illumination/modulation subsystem, a tissue sampling subsystem, a data acquisition subsystem, a computing subsystem, and a calibration subsystem. The invention can provide analyte property determination and identity determination or verification from the same spectroscopic information, making unauthorized use or misleading results less likely than in systems that use separate analyte and identity determinations. The invention can be used to control and monitor individuals accessing controlled environments. | 1. An apparatus for initiating an action responsive to an individual entering or leaving or attempting to enter or leave a controlled environment, comprising: a) a spectroscopic system adapted to determine a spectroscopic signal from the response of at least a portion of the individual's tissue to incident electromagnetic radiation, wherein the spectroscopic system comprises: i) an illumination/modulation subsystem;ii) a tissue sampling subsystem;iii) a data acquisition subsystem;iv) an analysis subsystem adapted to determine from the spectroscopic signal one or more analyte properties characteristic of the individual; andv) a calibration subsystem;b) an action initiation system adapted to initiate an action responsive to the one or more analyte properties. 2. An apparatus as in claim 1, wherein the analysis system is further adapted to determine an identity characteristic of the individual from the spectroscopic signal, and wherein the action initiation system is adapted to initiate an action responsive to both the identity characteristic and the one or more analyte properties. 3. An apparatus as in claim 2, where the identity characteristic comprises verification or nonverification of the purported identity of the subject. 4. An apparatus as in claim 3, further comprising an identity input system comprising at least one of: an alphanumeric password entry device, a key input device, a magnetic swipe card reader, a GPS location device, a finger print reader, a voice recognition system, a retina scanner, bar code scanner, or an implanted sensor or material. 5. An apparatus as in claim 2, where the identity characteristic comprises determination of whether the individual is part of a predetermined group, without requiring the individual to indicate a purported identity. 6. An apparatus as in claim 5, where the authorized group consists of only one individual. 7. An apparatus as in claim 1, wherein the one or more analyte properties comprises the presence of alcohol or one or more alcohol byproducts in the individual; the concentration of alcohol or one or more alcohol byproducts in the individual; the presence of a substance of abuse in the individual; the concentration of a substance of abuse in the individual; or a combination of the preceding. 8. An apparatus as in claim 1, wherein the action initiation system is adapted to perform at least one of: communicate the one or more analyte properties to an agent associated with the controlled environment; record the one or more analyte properties; allow or disallow access to the controlled environment based on the one or more analyte properties; or communicate to an agent associated with the controlled environment the result of a comparison of the one or more analyte properties with predetermined values. 9. An apparatus as in claim 1, wherein the portion of the individual's tissue comprises at least one of: the finger, forearm, ear lobe, palm, lip, hand, or skin of the individual. 10. An appa... | TruTouch Technologies, Inc. | G08B013/14 A61B005/1455
US08231556 | B2 | 2012-07-31 | Obtaining baseline patient information | 2009-04-30 | The disclosure relates to a method and system for obtaining baseline patient information. In some examples, a method may include acquiring first patient data, wherein the first patient data comprises at least one of first posture state data indicative of a plurality of posture states of a patient during a first time period or first therapy adjustment data indicative of a plurality of patient therapy adjustments made during the first time period; generating baseline patient information based at least in part on the first patient data; and comparing the baseline patient information to patient information generated based on second patient data. Therapy is not delivered to the patient according to a detected posture state of the patient during the first time period, and therapy is delivered to the patient according to the detected posture state of the patient during the second time period. | 1. A method comprising: acquiring first patient data, wherein the first patient data comprises at least one of first posture state data indicative of a plurality of posture states of a patient during a first time period or first therapy adjustment data indicative of a plurality of patient therapy adjustments during the first time period;generating baseline patient information based at least in part on the first patient data; andcomparing the baseline patient information to patient information generated based on second patient data, wherein the second patient data comprises at least one of second posture state data indicative of a plurality of posture states of the patient during a second time period or second therapy adjustment data indicative of a plurality of patient therapy adjustments during the second time period,wherein therapy is not delivered to the patient according to a detected posture state of the patient during the first time period, and therapy is delivered to the patient according to the detected posture state of the patient during the second time period, wherein the acquiring, generating, and comparing are performed via one or more processors. 2. The method of claim 1, wherein the baseline patient information comprises at least one of baseline posture state information or baseline therapy adjustment information. 3. The method of claim 2, wherein the baseline postures state information comprises at least one of baseline sleep quality information, baseline proportional posture information, or baseline posture state adjustment information. 4. The method of claim 2, wherein the baseline therapy adjustment information comprises a number of therapy adjustments received over the first time period. 5. The method of claim 1, further comprising modifying one or more aspects of the therapy based on the comparison of the baseline patient information and the patient information. 6. The method of claim 5, wherein modifying one or more aspects of the therapy based on the comparison comprises at least one of terminating, adjusting, or initiating one or more aspects of posture-responsive therapy. 7. The method of claim 1, further comprising generating the first posture state data, wherein generating the first posture state data comprises receiving posture state sensor data and detecting the posture state of the patient based on the posture state sensor data. 8. The method of claim 1, wherein comparing the baseline patient information to patient information generated based on the second patient data comprises presenting at least a portion of the baseline patient information and at least a portion of the patient information to a user via a user interface display of an external device. 9. The method of claim 1, wherein comparing the baseline patient information to patient information generated based on the second patient data comprises determining one or more differences between the baseline patient information and the patient information gen... | Medtronic, Inc. | A61B005/00 A61B005/103 A61B005/117 A61N001/00
US08688209 | B2 | 2014-04-01 | Device for ambulatory monitoring of brain activity | 2007-09-19 | A device for ambulatory monitoring of brain activity includes a patch secured to a head of a monitoring subject and an electrode secured to the head of the monitoring subject. The electrode produces a signal indicating brain activity of the monitoring subject. The device further includes a signal processor that receives the signal produced by the electrode and processes the signal to produce data. The signal processor is secured to the monitoring subject. The device further includes a display that displays the data produced by the signal processor, the display being secured to the monitoring subject. The device also includes a power source secured to the monitoring subject, the power source providing power to the electrode, the signal processor, and the display wherein the monitoring subject may ambulate unimpeded by the patch, the electrode, the signal processor, the display and the power source. | 1. A device for ambulatory monitoring of brain activity, comprising: a plurality of electrodes adapted to be operatively secured to a head of a monitoring subject, each electrode producing a signal indicating brain activity of the monitoring subject;a signal processor that receives the signals produced by the plurality of electrodes and processes the signals to produce data, the signal processor adapted to be secured to the monitoring subject and analyzing the signals such that the signal processor determines an occurrence of a brain spike based on a) a convolved signal that exceeds a first fixed threshold and b) coefficients determined from wavelet decomposition that exceed a second fixed threshold;a display that displays the data produced by the signal processor, the display adapted to be secured to the monitoring subject; anda power source adapted to be secured to the monitoring subject, the power source providing power to the plurality of electrodes, the signal processor, and the display,wherein the monitoring subject may ambulate unimpeded by the plurality of electrodes, the signal processor, the display and the power source. 2. The device according to claim 1, further comprising a communication port for transferring the data to a remote device. 3. The device according to claim 1, further comprising a strap adapted to be disposed around the head of the monitoring subject. 4. The device according to claim 3, wherein the plurality of electrodes are integral to the strap. 5. The device according to claim 1, further comprising a patch adapted to be attached to the head of the monitoring subject, wherein the plurality of electrodes are integral to the patch. 6. A method of ambulatory monitoring of brain activity, comprising: operatively securing a plurality of electrodes to the head of the monitoring subject;securing a signal processor to the monitoring subject;securing a display to the monitoring subject;securing a power source to the monitoring subject;providing power from the power source to the plurality of electrodes, the signal processor, and the display;producing signals indicating brain activity of the monitoring subject;transmitting the signals from the plurality of electrodes to the signal processor;analyzing the signals to determine a period of time when an absence or reduction in the number of neural spikes is measured to determine whether a neurological seizure event is imminent,producing data based on the period of time;displaying the data on the display, wherein the monitoring subject may ambulate unimpeded by the plurality of electrodes, the signal processor, the display and the power source. 7. The method according to claim 6, further comprising transferring the data from a communications port to a remote device. 8. The method according to claim 7, further comprising securing a strap around the head of the monitoring subject. 9. The method according to claim 8, wherein the plurality of electrodes are integral to the strap... | Koninklijke Philips N.V. | A61B005/04
US08428705 | B2 | 2013-04-23 | Atrial fibrillation detection | 2009-09-22 | Atrial fibrillation is detected in an electrical signal representative of a beating heart by measuring atrial activity over a time window of three or more beats, measuring beat interval variation over the time window and combining the measures of atrial activity and beat interval variation to produce an indication of an atrial fibrillation condition in the electrical signal. | 1. A computer implemented method comprises: receiving an electrical signal representative of a beating heart;measuring by the computer beat interval variation over a time window of three or more beats in the electrical signal;measuring by the computer atrial activity to produce an atrial activity measure, by: detecting a TQ interval in the electrical signal, the TQ interval defined by an offset of the T wave in the electrical signal and an onset of a subsequent QRS interval in the electrical signal;detecting P-waves in the electrical signal within the detected TQ interval of the electrical signal;classifying the detected P-waves as conducted or non-conducted; andcombining by the computer the atrial activity measure of classified P-waves and beat interval variation measure to produce an indication of an atrial fibrillation condition in the electrical signal. 2. The method of claim 1 wherein the electrical signals are provided from plural electrical channels and the method further comprises: determining whether one or more P-waves are detected during a TQ interval on a channel from the plural channels with a largest expected P-wave; combining a number and times of detected P-waves from that channel when plural P-waves are detected; otherwise,combining a number and times of detected P-waves from an alternate one of the plural channels. 3. The method of claim 1 wherein measuring atrial activity further comprises: computing a measure of the fraction of beat intervals that have one detected P wave. 4. The method of claim 1 measuring atrial activity further comprises determining if more than one P wave was present in a beat interval. 5. The method of claim 1 wherein detecting P waves further comprises: provisionally classifying a P wave as conducted or non-conducted, by: classifying a P-wave closest to the R-wave as conducted if the P-wave is within a time interval of 120 ms&lt;PR&lt;400 ms, otherwiseclassifying the P-wave as non-conducted; andclassifying any other P-waves in the same TQ interval as non-conducted. 6. The method of claim 1 wherein measuring atrial activity further comprises: using information from more than one signal channel representative of a beating heart to detect and classify the P-waves. 7. The method of claim 6 wherein measuring atrial activity further comprises: measuring the atrial activity for a plurality of channels;determining a quality indicator for the measured channels; andselecting one of the measures from the channels based at least in part on the determined quality indicators for the measured channels. 8. The method of claim 1 wherein measuring atrial activity further comprises: removing part or all of a ventricular waveform from the received ECG signal. 9. The method of claim 8 wherein measuring atrial activity further comprises: measuring the atrial activity from a portion of the signal from which part or all of the ventricular waveform was removed. 10. The method of claim 1 wherein measuring atrial activity... | Medicomp, Inc. | A61B005/0402
US08290573 | B2 | 2012-10-16 | Systems and methods for quantifying and providing indicia of ST-segment resolution in an ECG signal | 2009-08-20 | Systems and methods are provided for quantifying and providing indicia of ST-segment resolution in an electrocardiogram (ECG) signal. A receiver acquires an electrocardiogram (ECG) signal that includes an ST-segment. A processor processes the ECG signal to determine values for the ST-segment deviation relative to an isoelectric baseline. A user is allowed to provide a baseline signal to the processor. The processor responds to the baseline signal by marking a baseline ST-segment value corresponding to a baseline time. A user interface displays a linear graphical trend of variations in the measured ST-segment values relative to the baseline ST-segment value. In certain embodiments, the processor detects user-selected trigger events such as post-intervention ST deviation relative to the baseline time and the baseline ST-segment value, and provides indicia of the trigger event. In addition, or in other embodiments, a verbal annunciation of a percent ST-segment resolution is provided. | 1. A system for dynamically monitoring ST-segment resolution in a patient, the system comprising: a receiver to acquire an electrocardiogram (ECG) signal through a plurality of leads, the ECG signal including, for each of a plurality of detected heart beats, an ST-segment between a QRS complex and a T wave;a processor to process the ECG signal to determine values for the ST-segment relative to an isoelectric value;one or more controls to provide a user selected baseline signal to the processor, the processor responding to the baseline signal by marking a baseline ST-segment value corresponding to a baseline time; wherein the one or more controls further allow the user to select criteria for evaluating ST-segment resolution, wherein the processor determines the values for the ST-segment based on the user selected criteria, and wherein the user-selected criteria further comprises one or more statistical algorithms selected from the group comprising mean, standard deviation, variance, skew, and kurtosis; anda user interface in communication with the processor to display a linear graphical trend of variations in the measured ST-segment values relative to the baseline ST-segment value. 2. The system of claim 1, wherein the processor is further configured to: detect a trigger event relative to at least one of the baseline time and the baseline ST-segment value; andprovide indicia of the trigger event. 3. The system of claim 2, wherein the trigger event comprises an ST-segment resolution percentage reaching a threshold value. 4. The system of claim 3, wherein the indicia of the trigger event comprises a marker displayed at a location on the linear graphical trend corresponding to an elapsed time when the ST-segment resolution percentage reaches the threshold, wherein the marker indicates the elapsed time and the ST-segment resolution percentage. 5. The system of claim 3, further comprising an audio component to provide an audible signal for the indicia of the trigger event. 6. The system of claim 5, wherein the audible signal comprises a verbal annunciation that communicates the ST-segment resolution percentage. 7. The system of claim 6, wherein the verbal annunciation further communicates an identification of a lead that acquired the ECG signal. 8. The system of claim 2, wherein the trigger event comprises a lapse of a predetermined amount of time after a user-selected time. 9. The system of claim 8, wherein the indicia of the trigger event comprises a display of a measured percentage ST-segment resolution value on the user interface for one or more leads relative to the lapsed time. 10. A system for dynamically monitoring ST-segment resolution in a patient, the system comprising: a receiver to acquire an electrocardiogram (ECG) signal through a plurality of leads, the ECG signal including, for each of a plurality of detected heart beats, an ST-segment between a QRS complex and a T wave;a processor to process the ECG signal to determine values... | MR Holdings (HK) Limited | A61B005/0452
US07936170 | B2 | 2011-05-03 | RF coil and apparatus to reduce acoustic noise in an MRI system | 2008-08-08 | A radio frequency (RF) coil for a magnetic resonance imaging (MRI) system includes a first end ring section containing a plurality of openings and a second end ring section containing a plurality of openings. A plurality of rungs is disposed between the first end ring section and the second end ring section. Each rung has a first end connected to the first end ring section and a second end connected to the second end ring section. Each rung can also include a plurality of openings. The openings in the end rings and rungs reduces eddy currents and improves RF performance of the RF coil. | 1. A radio frequency (RF) coil for a magnetic resonance imaging (MRI) system, the RF coil comprising: a first end ring section;a second end ring section;a plurality of rungs disposed between the first end ring section and the second end ring section, each rung comprising a first end connected to the first end ring section and a second end connected to the second end ring section;a first plurality of openings located in the first end ring section; anda second plurality of openings located in the second end ring section. 2. An RF coil according to claim 1, wherein each rung further comprises: a first opening located in proximity to the first end of the rung; anda second opening located in proximity to the second end of the rung. 3. An RF coil according to claim 2, wherein each rung further comprises: a slit having a width, a first end connected to the first opening and a second end connected to the second opening; anda capacitor coupled across the width of the slit. 4. An RF coil according to claim 1, wherein the first end ring section further comprises a plurality of capacitors. 5. An RF coil according to claim 1, wherein the second end ring section further comprises a plurality of capacitors. 6. An RF coil according to claim 1, wherein the plurality of rungs are constructed of a material with high electrical conductivity. 7. An RF coil according to claim 1, wherein each opening in the first plurality of openings has a rectangular shape. 8. An RF coil according to claim 1, wherein each opening in the second plurality of openings has a rectangular shape. 9. An RF coil according to claim 1, wherein each opening in the first plurality of openings has a U-shape. 10. An RF coil according to claim 1, wherein each opening in the second plurality of openings has a U-shape. 11. A radio frequency (RF) coil for a magnetic resonance imaging (MRI) system, the RF coil comprising: a first end ring section;a second end ring section; anda plurality of rungs disposed between the first end ring section and the second end ring section, each rung comprising: a first end connected to the first end ring section;a second end connected to the second end ring section;a first opening located at the first end of the rung;a second opening located at the second end of the rung;a slit having a width, a first end connected to the first opening and a second end connected to the second opening; anda capacitor coupled across the width of the slit. 12. An RF coil according to claim 11, further comprising: a first plurality of openings located in the first end ring section; anda second plurality of openings located in the second end ring. 13. An RF coil according to claim 12, wherein the first end ring section further comprises a plurality of capacitors. 14. An RF coil according to claim 12, wherein the second end ring section further comprises a plurality of capacitors. 15. A resonance assembly for a magnetic resonance imaging (MRI) assembly, the resonance assembly com... | General Electric Co. | G01V003/00
US08091261 | B2 | 2012-01-10 | Bands for making adjustable loops | 2007-12-20 | Apparatus for forming a band into a persistent loop. The band incorporating the apparatus has a hole portion which contains a hole, a tongue portion which fits through the hole, and an adhesive attachment area on one or both of the hole and tongue portions. The band is made into a persistent loop by drawing the tongue portion through the hole and attaching the adhesive attachment area to the loop's outside surface. The size of the loop may be adjusted by varying the amount of the tongue portion which is drawn through the hole. A loop may be temporarily preformed by engaging a structure on the tongue portion with the hole. In embodiments in which the band is made of stiff material, the engaging structures may be the edges of the tongue portion and the edges of the hole. The disclosed embodiment of the apparatus is an identification band. | 1. Apparatus for forming a band into a persistent loop, the loop having an inner surface and an outer surface and the apparatus comprising: portions of the band including a hole portion of the band, the hole portion containing a hole and having a first end and a second end;a tongue portion of the band which is attached to the first end of the hole portion and fits through the hole; andan adhesive attachment area in the hole portion, the adhesive attachment area being located between the hole and the second end of the hole portion and having adhesive on the inner surface of the loop,the band being persistently formed into the loop by passing the tongue portion through the hole from the inner surface and attaching the adhesive to the tongue portion. 2. The apparatus set forth in claim 1 wherein: the tongue portion has a length which permits the size of the loop to be varied by varying the amount of the tongue portion which is passed through the hole. 3. The apparatus set forth in claim 1 wherein: the tongue portion includes a structure for mechanically engaging the hole and thereby permitting the band to be temporarily formed into the loop prior to being persistently formed into the loop. 4. The apparatus set forth in claim 3 wherein: the structure engages an edge of the hole. 5. The apparatus set forth in claim 4 wherein: the band is made of a material which has substantially the stiffness of writing paper; andthe tongue portion passes through the hole without being folded or rotated. 6. The apparatus set forth in claim 4 wherein: the structure is an edge of the tongue portion. 7. The apparatus set forth in claim 6 wherein: the structure is an edge of a broadened portion of the tongue portion. 8. The apparatus set forth in claim 1 wherein: the adhesive attachment area includes a tab. 9. The apparatus set forth in claim 8 wherein: when the adhesive attachment area is attached to the band, the tab is folded over and attached to the band's inner surface. 10. The apparatus set forth in claim 1 wherein: the portion of the tongue portion which has passed through the hole is designed to be removed after the adhesive attachment area on the hole portion has been attached to the band. 11. The apparatus set forth in claim 1 further comprising: a plurality of the adhesive attachment areas, the plurality including the adhesive attachment area in the hole portion and a further adhesive attachment area on the tongue portion 12. The apparatus set forth in claim 1 further comprising: an information area on the outer surface of the persistent loop. 13. The apparatus set forth in claim 12 wherein: the information area is on the outer surface of the hole portion of the band. 14. The apparatus set forth in claim 12 wherein: the information area is on the outer surface of the tongue portion of the band. 15. The apparatus set forth in claim 1 wherein: the hole has a form which makes the hole resistant to tearing. 16. A business form comprising: a face ply... | Endur ID Incorporated | A44C005/00 A61B005/103 A61B005/117
US08121664 | B2 | 2012-02-21 | Sensor | 2008-08-20 | A sensor for sensing physiological signal from outside the human body, said sensor comprising an electrode for abutment against the skin; and a transmission conductor electrically connected to said electrode, said transmission conductor formed from a flexible material, wherein in use a substrate is disposed between the skin and said transmission conductor, and said substrate is permeable to moisture and air thereby exposing said transmission conductor to said moisture. | 1. A sensor for sensing physiological signals from outside a human body, said sensor being adapted to be held against a skin of a user during use and comprising: an electrode for abutment against the skin; anda transmission conductor on which there is disposed a flexible substrate, wherein said flexible substrate is moisture permeable and said electrode protrudes through said flexible substrate and said electrode and said transmission conductor are integrally formed from a flexible conductive material and being adapted for wearing on a body of the user, and wherein in use the moisture permeable substrate is disposed between the skin and said transmission conductor and abuts against the skin, thereby exposing said transmission conductor to said moisture and facilitating signal acquisition from said skin. 2. A sensor as claimed in claim 1, wherein said moisture increases conductivity between said transmission conductor and said skin with minimal interference to the electrical connection between said transmission conductor and said electrode during use, thereby facilitating signal acquisition by the transmission conductor from said skin. 3. A sensor as claimed in claim 1, wherein said flexible material is a substantially plastic material including carbon black. 4. A sensor as claimed in claim 3, wherein said substantially plastic material is a rubberized material. 5. A sensor as claimed in claim 1, wherein said sensor has a thickness of about 2-3 mm, a width of about 20-30 mm, and a length of about 120-180 mm. 6. A sensor as claimed in claim 1, in which said transmission conductor is riveted to a terminal that allows for said sensor to be connected to an electronics module. 7. A sensor according to claim 6, further comprising said electronics module, and said electronics module contains means for transmitting, recording or displaying a physiological signal. 8. A sensor as claimed in claim 1, wherein said moisture permeable substrate is a stretchable textile material. 9. A sensor as claimed in claim 1, wherein said flexible substrate is a flexible textile material. 10. A sensor as claimed in claim 1, wherein said sensor comprises an adhesive layer which is disposed between said flexible substrate and said transmission conductor, said adhesive layer being perforated to allow moisture and air to permeate between said flexible substrate and said transmission conductor. 11. A heart rate monitoring belt which comprises at least one sensor as claimed in claim 1. 12. A heart rate monitoring belt which comprises a sensor as claimed in claim 1, and wherein an innermost layer of said belt includes said flexible substrate. 13. An apparel which comprises a sensor as claimed in claim 1, and wherein an innermost layer of said apparel includes said flexible substrate. 14. A sensor as claimed in claim 1, wherein at least a portion of said transmission conductor is permeable to moisture and air. 15. A sensor as claimed in claim 14, wherein said portion ... | Dayton Technologies Limited | A61B005/0408
US08271078 | B2 | 2012-09-18 | Method and apparatus for interfacing with a living object to obtain a biosignal measurement | 2009-06-18 | An apparatus for use in connection with biosignal measurement from the skin of a living object, comprising means for applying a test input impedance to the biosignal measurement, means for determining a response of a biosignal characteristics to the test input impedance, and means for setting a usage input impedance on the basis of the response. | 1. An apparatus to measure a biosignal from skin of a living object, wherein the apparatus comprises a processing device configured to: apply a test input impedance to the biosignal;determine a response of the biosignal to the test input impedance; andadjust a usage input impedance automatically based on the response of the biosignal to the test input impedance such that a quality of the biosignal is maintained during measurement of the biosignal in response to (1) the usage input impedance being applied to the biosignal and (2) changes in a characteristic associated with at least one of the skin and skin-apparatus interface. 2. The apparatus of claim 1, wherein the processing device is configured to: apply at least two different test input impedances to the biosignal; anddetermine a response of the biosignal to the at least two different test impedances. 3. The apparatus of claim 2, wherein the at least two different test input impedances include a first test input impedance sufficiently high ensuring detection of the biosignal, and a second test input impedance lower than the first test input impedance. 4. The apparatus of claim 1, wherein the processing device is configured todetermine an object impedance on the basis of the response of the biosignal to the test input impedance. 5. The apparatus of claim 4, wherein the processing device is configured to set the usage input impedance such that the usage input impedance compensates for the changes in the object impedance in response to the biosignal characteristics. 6. The apparatus of claim 4, wherein the object impedance includes a constant body impedance component and a variable interface impedance component dependent on the electricity conductivity of the skin of the living object. 7. The apparatus of claim 1, wherein the processing device is configured to: periodically repeat applying a test input impedance to the biosignal;determining a response of the biosignal to the test input impedance; andsetting a usage input impedance based on the response. 8. The apparatus of claim 1, wherein the quality of the biosignal includes one or more elements of a group comprising the strength of the biosignal, voltage level of the biosignal, and signal-to-interference ratio of the biosignal. 9. The apparatus of claim 1, wherein the biosignal is an electrocardiographic signal of a human being. 10. The apparatus of claim 1, wherein the processing device is configured to: determine an object impedance, andadjust the usage input impedance based on the object impedance such that the quality of the biosignal is maintained during measurement of the biosignal in response to (1) the usage input impedance being applied to the biosignal and (2) changes in a characteristic associated with at least one of the skin and skin-apparatus interface. 11. The apparatus of claim 10, wherein the processing device is configured to update determination of the object impedance in response to a predetermined change in the ... | Polar Eletro Oy | A61B005/053